Z-1943-2010 | 54366 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | BIOMET MICROFIXATION REF: 91-6103 "1.5MM" SYSTEM HIGH TORQUE (HT), SD, X-DR, SCREW 1.5MM X 3.5MM (WXL) TITANIUM LOT 829760, ***NON STERILE PRODUCT***.Bone screws used along with implantable bone plates for oral, cranio-maxillofacial procedures. | BIOMET MICROFIXATION | 3 | June 30, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling 91-6103 1.5x3.5mm HT SD X-drive screw Lot Number 829760. This package may contain 91-6704 1.5x4mm ThinFlap SD X-drive screw. | DESIGN: Process Design | Biomet Microfixation notified each consignee by email and fax on 10/05/2009. Each consignee was asked to use the included Response Form to provide product disposition. | Contact the recalling firm for information | 91 | 91 units | US distribution: CA, CO, MI, and MO. | Terminated | December 15, 2010 | 168 | 91 | Not_Computer
Z-1940-2010 | 55654 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences, BD FACS Sample Prep Assistant II (SPA II), blood sample preparer for flow cytometry analysis | BD FACS Sample Prep Assistant II (SPA II) | 2 | June 30, 2010 | 2010 | BD Biosciences, Systems & Reagents | Samples prepared on the BD SPA II, SN T0076 produced inaccurate sample results due to inaccurate sample dispensed as a result of the cap piercing probe | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm filed an MDR and implemented a field action consisting of a February 2010 notification letter and product replacement of affected probes. A CAPA for the probe design has been implemented. Effectivity is being tracked via mail and follow up phone call. | Contact the recalling firm for information | 349 | 349 units | Worldwide distribution: USA, Australia, Belgium, China (Hong Kong), Colombia, Jamaica, Korea, Mexico, Singapore and Uruguay. | Terminated | May 18, 2011 | 322 | 349 | Not_Computer
Z-1942-2010 | 55801 | complement c1q, antigen, antiserum, control | DAK | Complement components immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostics BINDAZYME Human C1q binding Circulating Immune Complex EIA Kit, Part#: MK021 | BINDAZYME | 3 | June 30, 2010 | 2010 | Inova Diagnostics Incorporated | The recall was initiated after the binding site notified Inova Diagnostics, inc. that the affected Bindazyme Human C1q binding Circulating Immune Complex EIA Kits may contain an ELISA plate labeled as "tissue Transglutaminase Coated Wells" instead of "C1q Coated Wells". | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Notification letters from INOVA Diagnostics Incorporated were mailed to each of the consignees via Certified mail on October 30, 2009.The notification letters inform the consignees of the reason for recall and provided instructions for final disposition of the recalled kits. In addition, a FAX back Attachment was attached to the recall notice for customers to respond to the recall. |  | 700 | 700 kits | Worldwide distribution: USA, Germany, Italy, Belgium, Romania, Portugal, Argentina, Spain, Sweden, and India. | Terminated | June 30, 2010 | 0 | 700 | Not_Computer
Z-1939-2010 | 55749 | control, pump speed, cardiopulmonary bypass | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | Hospital Grade Power Cord, UL, Part # 66319, 055476, manufactured by Electri-Cord Manufacturing Co. beginning in February 2003. The cords are a detachable unit, supplied with the instruments listed below: Medtronic Bio-ConsoleÂŽ 550 centrifugal pump console, Model #95180, Catalogue #95180, E95180, R95180.Medtronic Bio-ConsoleÂŽ 560 centrifugal pump consolesModel # 560BC, Catalogue #560BC, 560BCS, 560BCS1.Medtronic autoLogÂŽ, Model # ATLG110, Catalogue # ATLG110, ATLG110EATLG110R.Medtronic SequestraÂŽ 1000, Model #SEQ1000, Catalogue # SEQ1000, SEQ1000E, SEQ1000R. | Power cord used on BioConsole, autoLog, Sequestra | 2 | June 30, 2010 | 2010 | Medtronic Cardiovascular Revascularization & Surgical Therap | The potential risks during the use of this power cord include electrical shock, delay in setup and therapy, interruption of therapy, device failure, and fires. Although Medtronic has received no reports of adverse events related to the affected cord, a failure may potentially lead to serious adverse health consequences, depending on the device, and therapy being interrupted. This Field Notificatio | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent on 4/21/10 a Medtronic "Urgent Medical Device Recall Notice" letter dated April 19, 2010. The letter was addressed to their customers and described the problem and the product involved. They recommended to discontinue use of the AC power cords and to destroy them. Requested consignees to fill out the attached Medtronic Recall Certificate (FCA 1005), and return it to Medtronics. |  | 699 | 619 USA, 80 OUS | Worldwide distribution: USA, Latin America, Middle East, Japan, and Canada. | Terminated | December 24, 2011 | 542 | 699 | Battery | Treatment/Delivery/Therapy | Remove | Remove or Replace
Z-1936-2010 | 55885 | tray, start/stop (including contents), dialysis | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Dialysis On/Off Kit, DT11300, Sterile, Micropore Surgical Tape, Single Use Rolls, Port Access Dressing Change Tray, Tri-State Hospital Supply Company Manufacturers of Centurion Healthcare Products Howell, MI | Centurion Port Access Dressing Change Tray | 2 | June 30, 2010 | 2010 | Centurion Medical Products | The 3M micropore tape is under recall due to the tape not performing and needles becoming loose and dislodged. | OTHER/UNDETERMINED: Pending | The firm sent Urgent Recall Product Notices to its two customers dated 3/16/2010. The letter explained that the tape may not adhere properly. The consignees were instructed to remove any kits containing the recalled tape from their inventory, if the tape was used to secure dialysis needles or pressure dressings. If the tape was used for other purposes, the kits could still be used. The recall letter included a reply form for customers to complete and email or fax back. Upon receipt of the form, a Centurion Representative will contact them to retrieve affected product. Customers should contact Matthew Price at 517-546-5400 ext. 1135, if further info is needed. | Stop use and contact the recalling firm for information. | 3144 | 3144 kits | Nationwide Distribution -- AR and OH. | Terminated | February 09, 2011 | 224 | 3312 | Not_Computer
Z-1935-2010 | 55796 | electromyograph, diagnostic | IKN | Diagnostic electromyograph. | Physical Medicine | Neurology | 510(k) | Lumbar Matrix Scan, Diagnostic Electromyography System consisting of the CERSR Electromyography System and Electrode Array; the system consists of two devices. The first is a system of capital equipment referred to as CERSR Electromyography System, which includes a cart. The second device is disposable electrodes arrays, which are single use devices, used with the capital equipment that is adhered to the patient for monitoring the bioelectric signals produced by the muscles. The disposable is referred to as the Electrode Array. CERSR Electromyography System is packaged in a reusable crate maintained by SMI. The Electrode Array is packaged in a sealed foil pouch. Both devices are non-sterile and the packaging is for protection from damage during shipping. | Lumbar Matrix Scan (LMS) | 2 | June 30, 2010 | 2010 | SpineMatrix Inc | A full design review of the CERSR system was conducted in 7/2009. The audit revealed numerous design verification tests that either did not pass or did not meet the design input requirement. One requirement found that it was possible for the physician to think they were reviewing one patients scan, but were actually reviewing a scan from a different patient. This event did not happen; however, | DESIGN: Device Design | On 1/29/2010 all customers were provided a Voluntary Product Recall Notification letter that was sent via U.S. Postal Service Certified Mail and they were contacted by phone to arrange for pick up of the devices. The letter identified the affected product and described the problem. Customers are to check their inventory for the affected product and contact SpineMatrix Customer Service to receive instruction as to how to return the affected product, if the firm did not already contact the customer. Customers must complete the verification form and return it. Questions should be directed to Customer Service. | Contact the recalling firm for information | 1166 | CERSR - 14 units, Electrode Arrays - 1,152 units | Nationwide Distribution -- CA, OH, OR & TX. | Terminated | September 19, 2012 | 812 | 1166 | Not_Computer
Z-1944-2010 | 55186 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Monitor, includes one wall bracket, Model 25022. Stanley Senior Technologies, Lincoln, NE. | Tabs Professional Monitor, Model 25022 | 2 | July 01, 2010 | 2010 | Stanley Security Solutions, Inc. | Product might not indicate low battery alarms or exit alarms when powered by a DC power supply. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | The firm made phone calls beginning 3/26/10 to customers who purchased the 12-volt power supply informing them of the recall and instructing them to determine the number of units that will need to be returned for replacement. Important Safety Notice letters dated 3/25/10 were issued via regular mail on 3/26/10 and also providing the reason for the recall, instructions, and customer service contact information to arrange for return of the units. | Contact the recalling firm for information | 84 | 84 units | Nationwide Distribution | Terminated | September 16, 2011 | 442 | 187 | Battery | Alarm/Message | Replace | Remove or Replace
Z-1903-2010 | 55853 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | LIFEPAK 20 defibrillator/monitor.Product Part numbers: 3202487-xxx, 3202488-xxx, U3202487-000, 3202488-015, and U3202488-000. | LIFEPAK 20 defibrillator/monitor. | 1 | July 01, 2010 | 2010 | Physio Control, Inc. | A failure on the power supply assembly can result in either "No DC power" or "No DC or AC power." A failure of DC (battery) power can result in a delay of defibrillation therapy if no AC (line) power is available and the device will not operate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Firm began mailing Urgent Medical Device Correction notifications to affected customers on May 26, 2010 regarding the power supply to the device. All affected power supplies will be updated; this will be accomplished in two phases. Devices at a higher risk will be updated first in phase 1. The remaining lower risk devices will be updated afterward in Phase 2. Customers are advised to keep the defibrillators in service, and to follow recommended daily Operator Checklist steps, while service updates are scheduled. Customers with questions should contact the firm at 800-442-1142 between 6 am and 4 pm Pacific Time, Monday - Friday. |  | 43001 | 43,001 units | Worldwide distribution. In US, distributed domestically including to gov't accounts; and internationally to the following countries: ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Open |  |  | 43001 | Battery | Device Operation | Upgrade Battery | Repair
Z-1887-2010 | 55786 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Navilyst Medical, Inc., Vaxcel(TM) Port Plastic Standard Port with 9F Polyurethane Catheter (1.8 mm/2.8 mm), REF Catalog No. 45-360, UPN Product No. M001453600, STERILE, Rx ONLY, For single use only --- Implantable Vascular Access System --- (kit contains Huber needle) -- previously distributed by Boston Scientific Corporation. | Vaxcel | 1 | July 01, 2010 | 2010 | Navilyst Medical, Inc | FDA testing determined that non-coring needles included in the Vaxcel Port and Vaxcel Port with PASV Valve may result in coring and resulting septum damage which could lead to patient injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Navilyst Medical, Inc. is recalling Vaxcel Ports and Vaxcel Ports with PASV Valve with valid expiration dates worldwide. Navilyst sent "Urgent Medical Device Recall - Immediate Action Required" letter/recall notification packages on May 13, 2010, via Federal Express to 471 domestic consignees; on May 14, 2010, via e-mail to 14 international consignees and to Boston Scientific Corporation (prior owner of the affected products) via facsimile and Federal Express. On June 3, 2010, Navilyst Medical sent Recall Notification Packages on behalf of Boston Scientific Corp. via Federal Express to 182 domestic consignees. The letter instructs the customer to immediately discontinue use of and segregate recalled product, complete and return the reply verification tracking form to fax# 518-742-4450 and package and return the recalled product to the address provided.If you have any question, please call Navilyst Medical, Inc., Customer Service at 1-800-833-9973 or Michael Duerr at 1-518-742-4571. |  |  |  | Worldwide distribution:USA and countries of Belgium, Brazil, Canada, Colombia, Croatia, Germany, Greece, Hong Kong, Italy, Saudi Arabia, South Korea, Spain, Tunisia, and United Kingdom. | Terminated | August 03, 2012 | 764 |  | Not_Computer
Z-1898-2010 | 55670 | system, nucleic acid amplification test, dna, methicillin resistant staphylococcus aureus, direct specimen | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid Xpert MRSA/SA Blood Culture Test for use with the GeneXpert Dx System, manufactured by Cepheid, Sunnyvale, CA.Medical device, qualitative in vitro diagnostic test intended for the detection of Staphylococcus aureus and methicillin-resistant Staphylococcus aures DNA directly from patient positive blood cultures. It is not intended to monitor treatment for MRSA/SA infections. Sub culturing of positive blood cultures is necessary to recover organisms for susceptibility testing or for epidemiological typing. | Cepheid Xpert MRSA/SA Blood Culture | 1 | July 01, 2010 | 2010 | Cepheid | The firm is receiving reports of inaccurate results, which may result in delay of care for seriously ill patients. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Cepheid, issued a technical bulletin to customers on April 5, 2010 and an updated to the April 5 bulletin dated April 9, 2010 instructing them to perform a reflexive repeat test on any blood culture results initially determined as having SA.If further information on this please contact your Cepheid Sales Representative 1-888-336-2743 or www.CEPHEID.COM. | Contact the recalling firm for information | 5510 | 5,510 units | Worldwide distribution: USA and Europe, Taiwan, and Australia. | Terminated | May 25, 2011 | 328 | 5510 | Not_Computer
Z-1953-2010 | 55776 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Toggleloc Artificial ligament fixation device ZIPLOOP Technology Implant Kit, Sterile, REF 909848, Biomet Sports Medicine Warsaw, IN.Intended for soft tissue to bone fixation. | Toggleloc Artificial ligament fixation device | 2 | July 01, 2010 | 2010 | Biomet, Inc. | The firm has received complaints that the saddle is too long. The saddle is supposed to be between 7/8 and 1 1/8 inches long. The saddle in the complaints was up to 1.9 inches long. | PRODUCTION CONTROLS: Process Control | Biomet Sports Medicine issued an "Urgent Medical Device Recall Notice" dated March 31, 2010 to user facilities and distributors. A second notification containing updated information was issued June 2, 2010. Consignees were instructed to discontinue use of the product and return it to the firm along with a completed Response Form. For further information, contact Biomet at 1-574-371-3755 or 1-574-371-3756, Monday through Friday, 8AM to 5PM. | No consumer action necessary | 2466 | 2466 both products | Worldwide Distribution -- United States, Canada, Australia, The Netherlands, Belgium, Italy, France, New Zealand, Austria, United Kingdom, Chile, Greece, Poland, Denmark, Turkey, Portugal, South Africa and Spain. | Terminated | August 06, 2012 | 767 | 2466 | Not_Computer
Z-1951-2010 | 55771 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OptifluxÂŽ F180NR Dialyzers, Catalog Number: 0500318E. Manufactured by Fresenius Medical Care North America, Ogden, UT 84404. Designed for single use acute and chronic hemodialysis. | Optiflux F180NR Dialyzer | 2 | July 01, 2010 | 2010 | Fresenius Medical Care North America | Hemodialyzer may leak at the header/end cap | PRODUCTION CONTROLS: Process Control | Fresenius Medical notified customers by telephone script beginning May 14, 2010 and follow up letter advising users to examine inventory for recalled lot, discontinue use and contact Fresenius Customer Care 1-800-325-5188 , ext. 6610 for instructions on how to return the recalled product. |  | 34152 | 34,152 units | Nationwide Distribution | Terminated | September 13, 2011 | 439 | 34152 | Not_Computer
Z-1888-2010 | 55824 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | 22ga x 1" Straight Huber NeedleThis product is a small gauge needle intended for single use hypodermic access to implanted ports. | Huber Needle | 1 | July 02, 2010 | 2010 | Multi-Med, Inc. | FDA sample determined the needles to be coring. |  | The firm, Multi Med, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 26, 2010 to their sole consignee Navilyst Medical, Inc., via email and overnight courier with signature notification. The letter describes the product, problem and action to be taken by customer. The customers was instructed to immediately examine their inventory and quarantine product subject to recall, if the product was distributed to others customers, identify and notify them at once and include recall notification letter,. Recall has been initiated due to FDAs sample analysis that these needles may core implanted port septums.If you have any questions, call me, Sue Starkey, at 603-357-8733 or via email at astarkey@multimedinc.com. |  | 5169 | 5169 | Worldwide distribution | Open |  |  | 10169 | Not_Computer
Z-1949-2010 | 55825 | sealant,polymerizing | NBE |  |  | General & Plastic Surgery |  | Coseal Surgical Sealant 4mL, Product Codes: 934071 & 934074.Indicated for use in vascular reconstructions to achieve adjunctive hemostasis. | Coseal Surgical Sealant | 2 | July 02, 2010 | 2010 | Baxter Bioscience | The recall is being conducted as a precautionary measure due to out of specification results at the 18 month and 21 month time points (out of the 24 month licensed shelf life to expiration) for the percent thiol substitution during a Coseal Stability study. The parameter that is out of specification is an indicator of possible failure for the product to gel appropriately. | PRODUCTION CONTROLS: Process Control | The recall notification was initiated on January 14, 2010 with the firm sending a Product Recall letter and attached Customer Reply Form to the affected consignees. The recall letter informed the consignees of the reason for recall and requested them to immediately stop the usage/distribution of the lots noted above and to quarantine any units you may still have in your inventory. The letter provided instruction for product return.The letter also requested the consignees to please forward the recall information as appropriate to other services/facilities/departments that they may have provided this product to. Consignees were instructed to complete theCustomer Reply Form, and fax it as soon as possible to 888-345-4261. Consignees with questions were instructed to contact Stericycle at 877-870-4486. |  | 15262 | 15,262 | Nationwide Distribution | Terminated | December 22, 2010 | 173 | 26440 | Not_Computer
Z-1895-2010 | 55854 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq Two-Channel Infuser; List Number: 16027; Distributed by: Hospira, Inc. Morgan Hill, CA and Manufactured by: Hospira Inc, La Aurora, Heredia, Costa Rica.Intended for the delivery of fluids, solutions, drugs, agents, nutritionals, electrolytes, blood and blood products via parenteral, enteral, intravenous, intra-arterial, subcutaneous, epidural or irrigation routes of administration. | Hospira brand Symbiq TwoChannel Infuser | 1 | July 02, 2010 | 2010 | Hospira Inc | Hospira has received numerous customer complaints of the Symbiq infusion pump's failure to detect air-in-line conditions during operation. | DESIGN: Device Design | Hospira issued a "Urgent: Medical Device Recall Notification" dated June 11, 2010 to update a previously issued Clinical Bulletin. Customers were informed of the issue and instructed on proper mitigation steps to further avoid the issue associated with the affected product. Consignees were asked to complete and return the Reply Form.For further information, contact Hospira Technical Support Operations at 1-800-332-4002, 6AM to 4PM PST, Monday through Friday. |  | 28485 | 28,485 total | Worldwide Distribution -- United States (AL, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SC, SD, TN, TX, VA and WA), and Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and United Arab Emirates. | Terminated | December 03, 2013 | 1250 | 28485 | Not_Computer
Z-1956-2010 | 54371 | screw, fixation, intraosseous | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET 3i GS300 GOLD SLOTTED SCREW-3mm, STERILE, Rx Only. Made in USA. Lots 868517, 865348, and 868515. BIOMET 3i, 4555 Riverside Dr., Palm Beach Gardens, FL. | Gold Slotted Screw | 2 | July 06, 2010 | 2010 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling their Gold Slotted Retaining Screw 3MM, Model Number GS300. Product is Out Of Specification - product was manufactured with a material not specified for use in this product. | OTHER/UNDETERMINED: Pending | Each US customer was contacted by telephone using a prepared phone script. Following telephone contact, Biomet 3i performed a follow-up of each call with a faxed letter. International customers were first notified by e-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected product and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i.Questions or concerns are directed to Kelly Taylor at the firm at 561-776-6700. |  | 495 | 495 screws | Product was distributed to 3 distributors, 14 direct consumer/users. Product was distributed in Japan, Spain, Switzerland, Uruguay, Paraguay, Mexico, France, Netherlands, Australia, and in the US in KY, DC, IL, RI, MA, NJ and TX. | Terminated | August 30, 2010 | 55 | 495 | Not_Computer
Z-1955-2010 | 55774 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida System | SSH880CV Aplio Artida System | 2 | July 06, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction on the SSH-880CV Aplio Artida System due to a problem related to a wrong measurement value display in using with DICOM viewer.When an image is transferred from this system in DICOM transfer and a measurement is performed on a Doppler image in the DICOM viewer, measured values such as VTI (Velocity Time Integral) and PI/RI | DESIGN: Software Design | The software will be replaced. To accomplish this, Toshiba America Medical Systems (TAMS) will notify consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention.For questions or concerns, please contact your Toshiba representative or the firm at (714) 730-5000. |  | 22 | 22 | Nationwide ((CA, GA, IA, LA, MA, MD, MN, MT, PA, PR, OH, OR, TN). | Terminated | March 21, 2012 | 624 | 22 | Software | Display/Image | Software update | Software Update
Z-1971-2010 | 55316 | enzyme linked immunoabsorbent assay, herpes simplex virus, non-specific | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Sure-Vue HSV-2 Rapid Test, Catalogue No: 23-046395.20 test KitIn Vitro Diagnostic test for Herpes Simplex Virus 2 | Biokit and SureVue HSV2 Rapid Test | 2 | July 08, 2010 | 2010 | Biokit U.S.A. Inc. | Negative serum was giving a positive result Herpes Simplex Type 2 (HSV-2). It is possible that misdiagnosis (i.e., positive) could be made. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biokit USA notified US Customers by fax on 2/18/10 of the reason for the kit recall and requested to discard any remaining units. Accounts were advised to retest patients or review test results. Questions are directed to the firm at 800-926-3353. |  | 8 | 8 kits | US (CA, FL, GA, IN, KY, MA, NY, VA) and UK | Terminated | January 26, 2012 | 567 | 200 | Not_Computer
Z-1968-2010 | 55710 | instrument, manual, surgical, general use | MDM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker 18 MM AVS TL Trial Spacer; Catalog number: 48389418.Manufactured by: Stryker Spine SAS, Cestas, France;Distributed in the US by: Stryker Spine, 2 Pearl Court, Allendale, NJ 07401.Intended for use in the thoraco-lumbar spine. | AVS TL Trial Spacer | 2 | July 08, 2010 | 2010 | Stryker Spine | Stryker Spine has received reports of breakage of the distal end of the AVS TL Spacer Trial during the trialling step of the AVS TL procedure. | PRODUCTION CONTROLS: Process Control | Urgent Product Recall letters, dated May 13, 2010 were sent via Federal Express to all direct accounts (branches and hospitals) with a separate letter to surgeons. The letter identified the affected product and also described the issue, potential hazards, and risk mitigations. The letter also asked customers to quarantine any affected product and return to Stryker Spine. A Customer Response Form was provided and customers should contact their Stryker Spine Sales Rep for additional information on replacement product. Questions should be directed to Tiffani Rogers on 201-760-8206. |  | 139 | 139 units | Nationwide Distribution. | Terminated | September 06, 2012 | 791 | 2110 | Not_Computer
Z-1969-2010 | 55721 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Starion Instruments TLS3 35C Thermal Ligating Shears, Electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments Corp, Sunnyvale, CA.Intended use: medical device for use in cutting and cauterization of soft tissue during surgery, as well as cutting of natural or synthetic, non-metallic sutures during surgery. | Starion Instruments TLS3 35C Thermal Ligating Shears | 2 | July 08, 2010 | 2010 | Starion Instruments | Some jaws exhibited lack of ductility and could break if overloaded or used improperly by placing too much tissue in the bottom of the "v" of the jaw. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm called domestic consignees but no letter was sent to domestic cosignees and a notification letter was sent to international consignees.The firm called domestic consignees and sent a notification letter to international consignees. No records were kept of phone calls and a copy of the letter sent to international consignees was not obtained. |  |  |  | Worldwide distribution | Terminated | November 27, 2010 | 142 |  | Not_Computer
Z-1972-2010 | 55750 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Starion Instruments Real Hand ThermaSeal Thermal Ligating Shears, electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments, Sunnyvale, CAIntended use: Medical Device used for cutting and cauterization of soft tissue during surgery, and of natural or synthetic, non-metallic sutures during surgery. | Starion Instruments Real Hand ThermaSeal Thermal Ligating Shears | 2 | July 09, 2010 | 2010 | Starion Instruments | Due to an improper fixture revision, device may be missing a screw at the jaw assembly, resulting in jaw components falling off during use. | PRODUCTION CONTROLS: Process Control | The firm submitted an MDR but determined that there was no risk to health. E-mails were sent on June 26, 2008, to consignees, describing the product, the problem, and action to be taken by consignees. The consignees were ask to provide a list of all unused ThermaSeal devices and their locations, to return the products to Starion from the lots noted and to provide a copy of the attached letter to their sales force and customers if they choose.If you have any questions, please call (408) 522-5200 ext. 326 or (408) 350-9934. | Contact the recalling firm for information | 217 | 217 units | Worldwide distribution. | Terminated | October 27, 2010 | 110 | 217 | Not_Computer
Z-1973-2010 | 55419 | twister, wire | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device, SINGLE-USE Intended use: The device is a sterile, single-use only surgical instrument used to tension and twist USP size 5, 6, and 7 (metric size 7.0, 8.0, and 9.0) stainless steel surgical sutures for sternal closure. The device should not be used on patients with poor sternal bone quality. | TORQ Sternal Closure Device | 2 | July 09, 2010 | 2010 | KARDIUM INC. | A component of the device is installed incorrectly which can result in the component falling out of the device and into the patient during use. | OTHER/UNDETERMINED: Pending | The firm, Kardium, informed it's US distributor on March 30, 2010 by e-mail. The distributor was instructed to return all unused product to Kardium. The units remaining at hospitals will be returned to Kardium by the distributor sale reps. If you have any questions, please contact (604) 248-8891. |  | 420 | 420 units | Worldwide distribution: USA and Canada. | Terminated | October 07, 2010 | 90 | 420 | Not_Computer
Z-1974-2010 | 55598 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit.Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M.A programmable diagnostic computer intended for acquiring, digitizing, storing, displaying and reviewing hemodynamic data for use in hospital cardiac catheterization laboratories. The system is cleared for us with a component physiological data collection unit - the Schiller Argus Pro PB 1000 Monitor unit. | Horizon Cardiology Hemo Component Schiller Argus Pro PB1000 Monitor unit | 2 | July 09, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202.If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager. | Contact the recalling firm for information | 26 | 26 units; 19 in operational use, 8 as emergency replacement units. | Nationwide distribution: CA, CO, MD, NH, NT, NY, TN, and TX . | Terminated | July 21, 2010 | 12 | 26 | Software | Display/Image | Replacement | Remove or Replace
Z-1954-2010 | 55894 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Stellant CT Dual Syringe Kits, Catalog Number SDS-CTP-QFT and SDS-CTP-SPK | Stellant CT Dual Syringe Kits | 2 | July 09, 2010 | 2010 | Medrad Inc | Kits contain a T connector that may be susceptible to cracking or breaking or leaking. This could cause product leakage or, less likely, patient injury. | DESIGN: Component Design/Selection | The recalling firm issued an Urgent Medical Device Field Correction on 5/24/10 requesting customers to check their inventory for the affected lots and complete the fax back form with the amount of product in stock. The manufacturer plans to make arrangements to provide individually packaged low-pressure connector tubing to use with the product in the event that cracking or breaking and leaking of the T connector is observed. This requires vigilance by the customers over the life of all of the devices.For questions concerning this recall, please contact the company 1-800-633-7231, select option 2, and indicate that you are calling about the Stellant CT Dual Syringe Kit field correction. | Contact the recalling firm for information | 1382889 | 1,382,889 kits | nationwide and Worldwide distribution includes AU, BE, BG, BM, BR, CA, CH, CI, CL, CN, CO, CS, DE, DK, EC, EG, FI, FR, GB, GT, HN, IN, IT, JM, JP, KW, KY, MU, MX, MY, NL, OM, PA, PH, PK, PR, SE, SG, TW, YE, ZA. | Terminated | October 13, 2010 | 96 | 1382889 | Not_Computer
Z-1995-2010 | 55206 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit; Part Number: A78803The Access AccuTnl assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of cardiac troponin I (cTnI) levels in human serum and plasma using the Access Immunoassay Systems to aid in the diagnosis and treatment of myocardial infarction and cardiac muscle damage. Cardiac Troponin I determination aids in the risk stratification of patients with unstable angina or non-ST segment elevation acute coronary syndromes with respect to relative risk of mortality, myocardial infarction, or increased probability of ischemic events requiring urgent revascularization procedures. | Access AccuTnI Assay | 2 | July 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because as a follow up to PCA # 12605 dated December 16, 2009, Beckman Coulter has confirmed additional reports that each of the following assays may produce lower than expected results when run in conjunction with Access AccuTnl (REF A78803):(1) Access Testosterone (REF 33560);(2) Access Thyroid Uptake (REF 33810);(3) Access Total T4 (REF 33800).This issue can occ | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, issued a "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with attached Customer Response form dated March 24, 2010 to all customers. The letter describes the issue, impact, action and resolution. In the PCA beckman Coulter recommends the following mitigation to reduce the risk of generating erroneous results: (1) Verify the APF software on the immunoassay system. If the updated APF software listed in the table above was already loaded , skip the remainder of this section and proceed to the Resolution section;(2) If the updated APF software has not yet been loaded, but is available, load the software appropriate for the system, skip the remainder of this section and proceed to the Resolution section;(3) If the updated APF software is not immediately available, please contact Technical Support at 1-800-854-3633 in the United States and Canada. Consignees outside the United States and Canada, contact their local Beckman Coulter Representative and follow the instructions below until the software has been loaded on their system. (a) Access and Access 2 users should batch process Access Testosterone, Thyroid Uptake, or Total T4 samples; (b) For UniCel DxI users, assign Access AccuTnI (REF A78803) to a different reagent pipettor(s) than Access Testosterone, Thyroid Uptake, or Total T4;(c) The Access Testosterone, Thyroid Uptake, or Total T4 results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Abnormally low Access Testosterone, Thyroid Uptake, or Total T4 results from samples tested in conjunction with Access AccuTnI (REF A78803) should be reviewed in the context of other diagnostic tests.Customers were instructed that after loading the updated APF software, the Access AccuTnI assay (REF A78803, test code TropI, test ID 275) must be recalibrated. Recalibrate of the Access Testosterone, T. |  | 32407 | 32,407 units | Worldwide distribution: USA and countries of Algeria, Australia, Bangladesh, Belgium, Bulgaria, Canada, Chile, Czech Republic, France, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Lebanon, Libyan Arab Jamahiriya, Monaco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Poland, Puerto Rico, Romania, Russian Federation, South Africa, Spain, Swaziland, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | May 04, 2012 | 662 | 32407 | Other | Output/Calculation | Instructions/Software update | Software Update
Z-2022-2010 | 54688 | system, applicator, radionuclide, manual | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Angiotech Prostate Stabilization Set.Prostate Stabilization Set. STERILE. Product Number: 500118200. 510k K944522. Qty Dist. - 4,180.MEDICAL DEVICE TECHNOLOGIES, INC. 3600 Southwest 74th Avenue, Gainesville, Florida 32608 U.S.A.Intended Use: Prostate Stabilization Set | Angiotech Prostate Stabilization Set | 2 | July 13, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies doing business as Angiotech has conducted a recall on Angiotech Breast Localization Needles, Soft Tissue Biopsy Needles, Vascular Access Needles, Bone Biopsy Needles, Access Needles, Galactography Kit, Stabilization Needles and Needle Guides for packaging integrity. | OTHER/UNDETERMINED: Pending | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated September 1, 2009, to all customers. on September 2, 2009 by UPS 2nd day delivery with an assigned tracking number. The letter described the product, problem and action to be taken by customers. The customers were instructed to please review your current inventory and segregate the affected lot numbers, complete the attached Customer Acknowledgement Form and return via fax to our Quality Assurance department at 1-352-338-0662 or 1-800-333-0440. If you are a distributor of this product, you are responsible for notifying any customer you may have sold the product to.If you have any questions concerning this notification, please contact Shannon Brooks at (352) 338-0440 ext.355 or Katrenia Williams at ext. 353. |  | 4180 | 4,180 | Worldwide distribution: USA and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | February 25, 2011 | 227 | 305698 | Not_Computer
Z-1924-2010 | 55648 | unit, phacofragmentation | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Constellation Vision System, Model: Constellation Vision System w/Laser, Catalogue Number: 8065751145The Constellation Vision System is an ophthalmic microsurgical system that is indicated for both anterior segment (i.e. phacoemulsification and removal of cataracts) and posterior segment (i.e., vitreoretinal) ophthalmic surgery in addition to the indications included with the optional Next Generation laser. | Constellation Vision System | 1 | July 13, 2010 | 2010 | Alcon Research LTD dba Alcon Laboratories, Inc. | Alcon identified system performance and machine settings that may impact the infusion performance of the CONSTELLATION Vision System, causing unexpected shut down, unresponsive touch screens, and fluidics issues. | DESIGN: Software Design | The firm, Alcon, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 2, 2010, along with training material to all customers. The letter informed the customers of the problem and indicated that Alcon Surgical Sales representative will be contacting them in the near future to follow-up on this training material and to answer any questions you may have. Alcon informed the customers that this action will not require removal of the CONSTELLATION Vision System from their facility. The letter also informed the customers that their Alcon Technical Customer Service Representative will be contacting them when we are authorized to provide you with the software and hardware updates for your system.For assistance and questions regarding the CONSTELLATION Vision System, customers were told they may contact their Alcon Surgical Sales Representative or Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8. | Contact the recalling firm for information | 824 | 824 units total | Worldwide distribution. | Terminated | September 19, 2011 | 433 | 824 | Computer
Z-1999-2010 | 55790 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Microstream O2/CO2 NASAL FILTERLINE ADULT 25UNPart Number: 006912 | Microstream O2/CO2 NASAL FILTERLINE ADULT | 2 | July 13, 2010 | 2010 | Oridion Medical 1987 Ltd. | The Microstream Filterline may not be recognized by the monitor when connected. No CO2 readings will be displayed and no alarms may sound for unmonitored patients. Replacing the filterline may delay treatment. | DESIGN: Process Design | Oridion notified customers by letter on 5/26/10 and requested that they remove the FilterLines from the related production lots from the field and an immediate replacement with FilterLines from subsequent production lots. Questions are directed to the responsible person at the supplier named in the letter, or to Oridion Capnography at 888-674-3466. |  | 2 | 2 units | AZ, CA, FL, HI, LA, , MA, MO, NJ, OH, PA, WA, WI, | Terminated | June 21, 2012 | 709 | 350827 | Not_Computer
Z-1997-2010 | 55544 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems. | PVT681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems | 2 | July 13, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems due to an incorrect sweep direction setting. Images acquired using the endo-cavitary transducer PVT-681MV and cross-sectional images created from these volume images are reversedhorizontally on the Aplio XG (SSA-790A) and the Xario XG ( | OTHER/UNDETERMINED: Pending | Toshiba America Medical Systems (TAMS) will issue notification to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. matter. Questions regarding this letter are directed to the firm at (800) 421-1968 or to your local Toshiba representative. | Contact the recalling firm for information | 27 | 27 | Nationwide | Terminated | February 29, 2012 | 596 | 27 | Not_Computer
Z-1998-2010 | 55675 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | O-Arm Imaging SystemThe O-Arm imaging system is a mobile X-Ray system designed for medical applications in surgical theaters, focusing on 3D imaging but also offering 2D fluoroscopic Imaging. | OArm Imaging system | 2 | July 13, 2010 | 2010 | Medtronic Navigation, Inc. | The possibility that the real time image display can freeze up during use of a 2D continuous fluoroscopy. |  | The firm, Medtronic, Inc., sent an "Urgent Field Safety Notice" letter dated May 10, 2010, to all custmers. The letter describes the product, problem and actions to be taken by customer. The firm recommended that the users of the O-Arm Imaging System reveiw the information provided to ensure patient safety remains uncompromised during use of the product. The customers were instructed to discontinue use and contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200, if the image is not updating as expected. A Medtronic Navigation representative will be contacting you within the coming months to schedule a time to evaluate and service the potentially affected components within your system. A service visit will also be scheduled starting November 2010 to update your system with a software based solution that will prevent this issue from occuring.If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200. | Contact the recalling firm for information | 221 | 221 | Worldwide distribution: USA and countries of Germany, Singapore, South Africa, Poland, Libya, Belgium, Switzerland, Turkey, Italy, Denmark, France, Spain, Finland, Czech Republic, Ireland, Slovakia, the NetherLands, Austria, UAE, Kuwait, Israel, Australia, Sweden, Canada, South Korea, Japan and India. | Terminated | October 07, 2011 | 451 | 221 | Software | Display/Image | Software update | Software Update
Z-1996-2010 | 55895 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-STAT PT/INR CartridgeAbbott Point of Care, Princeton, NJ | iSTAT PT/INR Cartridge | 2 | July 13, 2010 | 2010 | Abbott Point of Care Inc. | A high mean bias (12-16%) was documented in 8 lots of i-STAT PT/INR cartridges. This may lead a clinician to believe that a patient is adequately anticoagulated when they are not. Inadequate anticoagluation therapy may be given and issues with clotting and/or thrombosis could occur. No reports of injury have been received. | OTHER/UNDETERMINED: Pending | Urgent recall notifications were sent on May 13 2010. Customers were instructed to discontinue use and to return the identified lots. Additional questions are directed to Abbott Point of Care Technical Support at 800-366ÂŽ8020, Option 1 or your Abbott Point of Care representative. |  | 44976 | 44976 in US, 3648 in rest of world. | Worldwide distribution: Argentina, Hong Kong, China, Finland, Greece, India, Portugal, Russian Federation, Saudi Arabia, United Arab Emirates, Italy, Spain, Australia, Germany, UK, Sweden, The Netherlands, France, Canada. | Terminated | September 22, 2011 | 436 | 44976 | Not_Computer
Z-2025-2010 | 55867 | antigens, febrile, slide and tube, all groups, salmonella spp. | GNC | Salmonellaspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Difco Salmonella 0 Group A Antigen, catalog #240731, packaged in 5 mL glass vials, labeled in part ***Becton, Dickinson and Company, Sparks, MD 21152*** | Difco Salmonella 0 Group A Antigen, catalog 240731, | 3 | July 15, 2010 | 2010 | Becton Dickinson & Co. | In vitro diagnostic reagent to aid in the diagnosis of Salmonellosis may exhibit decreased or no reactivity. If proper controls are run with the lot in question, it would become apparent that the antigen is non reactive and patient specimen results should not be reported. If control testing is not performed a salmonella-infected patient could go undiagnosed, with possible adverse effects. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recalling firm notified customers by fax and e-mail on 05/19/10. The notification advised that the referenced lot number of Salmonella 0 Group A Antigen may exhibit decreased or no reactivity when testing the patient's serum directly for homologous antibodies by either a slide or tube agglutination test to aid in the diagnosis of Salmonellosis. The recalling firm also advised that they would issue replacements for the referenced lot number under Catalog number 240731, which would not be available until approximately July 15, 2010. Consignees were requested to return enclosed response form.Customers with questions about replacements should contact BD customer service at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 93 | 93 units | Product was sold to distributors nationwide and to affiliates in Belgium and Korea. | Terminated | July 21, 2010 | 6 | 93 | Not_Computer
Z-2027-2010 | 54565 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch SureStep and SureStep Pro Test strips, for use with the OneTouch SureStep Blood Glucose Meters, in 100 count and 50 count packaging distributed by LifeScan Inc., Milpitas, CA, Manufactured by LifeScan Puerto RicoMedical device for quantitative measurement of glucose in whole blood for use in home settings (SureStep). | OneTouch SureStep | 2 | July 17, 2010 | 2010 | Lifescan Inc | Products may provide inaccurate results at readings over 400 mg/dL | OTHER/UNDETERMINED: Pending | The firm, Lifescan, Inc., issued two "Urgent Medical Device Recall" letters dated February 24, 2010 to all customers for the OneTouch SureStep Test Strips and the SureStep Pro Test Strips. The letters describes the product, problem and action to be taken by the customers. The customers were instructed to look at your test strip vials for the recalled lots and check the strips or if you have test strips from the lot, call LifeScan Customer Service at 800-574-6139 between 5 a.m. to 7 p.m. Pacific Time, seven days a week or visit www.SureStep.com. Replacement product will be shipped to you immediately at no charge. The customers were instructed to continue to test your blood glucose, while wafting for replacement. Also, the customers are asked to continue use of the meter with precautions if results over 400 units are observed. For more information, please visit www.SureStep.com or call our Healthcare Professional Line at 888-455-5627. |  | 1014350 | 1,014,350 | Worldwide distribution: USA and countries of Europe, Latin America and Asia. | Terminated | December 08, 2010 | 144 | 9640700 | Not_Computer
Z-2026-2010 | 54521 | motor, surgical instrument, ac-powered | GEY | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ADVANTAGE Turbo 2-Button Shaver. Rx Only, Made In USA, NON STERILE. CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA. | ADVANTAGE Turbo 2Button Shaver | 2 | July 17, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Advantage, 2-Button Turbo Handpiece Catalog Number D9924. These products may self-activate due to moisture intrusion into the housing body. This issue has been isolated to products manufactured between June 7, 2004 and June 1, 2006. | OTHER/UNDETERMINED: Pending | ConMed Linvatech notified affected customers by letter sent by FedEx next day delivery. Customers were asked to complete a response form and indicated if they have the product or have transferred the product. A Service Request number will be assigned for the returned product.Additional questions are directed to the firm's Customer Service department at 800-535-8536. |  | 2373 | 2,373 | Worldwide distribution. Products were distributed domestically throughout the U.S. Also United Arab Emirates, Argentina, Australia, Belgium, Switzerland, Chile, China, Germany, Denmark, Spain, France, United Kingdom, Hong Kong, India, Italy, Japan, Lithuania, Morocco, Netherlands, New Zealand, Singapore, Turkey, South Africa. | Terminated | August 19, 2010 | 33 | 2373 | Not_Computer
Z-2032-2010 | 55575 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | DePuy ASR 100 Acetabular Implant, 46 mm, 999800746, sterile, DePuy International, Ltd, Leeds, United Kingdom. | DePuy ASR 100 Acetabular Implant, 46 mm | 2 | July 17, 2010 | 2010 | Depuy Orthopaedics, Inc. | Notification to clinicians of new revision rate data/information regarding the use of the device. As part of the post-market surveillance, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports. As part of the post-m | OTHER/UNDETERMINED: Pending | DePuy Orthopaedics issued a Field Safety Notice to share the new data with surgeons and to reiterate the importance of correct implant positioning and patient selection in achieving optimal implant performance and survivorshop. For additional information regarding this matter, please contact the DePuy Scientific Information Office at 888 554 2482. |  | 1205 | 1205 | Nationwide | Terminated | May 10, 2012 | 663 | 37834 | Not_Computer
Z-2030-2010 | 55900 | clip, implantable | FZP | Implantable clip. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Medtronic U-CLIP Double Arm Approximation Device, model V100D U-Clip, Catalog Number V100DRFRN2. Sterilized using Irradiation. The U-Clip consists of specially designed vascular clips optionally connected to needles via flexible members. These clips may or may not include coils surrounding a Nitinol wire core. This design allows precise placement of clips prior to closure and facilitation from standard medical and implantable grade materials. Accessories include Clip delivery and removal devices as well as vessel cutting and retention tools.The Coalescent Surgical U-CLIP is intended for endoscopic and non-endoscopic general soft tissue and prosthetic material approximation/attachment and/or ligation and the creation of anastomoses in blood vessels, grafts and other tubular structures; including use in cardiovascular and coronary artery bypass grafting procedures. | Mectronic UCLIP V100D | 2 | July 17, 2010 | 2010 | Medtronic Cardiovascular Revascularization & Surgical Therap | Medtronic received reports involving the V100D U-Clip where during handling, the release mechanism has become separated from the needle assembly. These instances of separation have occurred when pulling the clip through prosthetic material, or when grasping and releasing the clip using robotic instrumentation. While this issue does not impact performance of the clip once implanted, separation of | OTHER/UNDETERMINED: Pending | The firm, Medtronic CardioVascular, sent an "Urgent Medical Device Recall Notice", dated April 2, 2010, to Risk Managers of each affected account. The letter described the issue, identified affected product, and the action to be taken by customers. The customers were ask to immediately discontinue use of affected devices and to quarantine all unused V100D U-Clip inventory. The letter stated that a Medtronic representative will contact the account to arrange return of the V100D U-Clips devices and that credit will be issued upon return of unused devices. Medtronic has stopped manufacture of the product therefore replacements will not be available. Devices can be returned to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428. The customers were ask to please fill out the Medtronic Recall Certificate and fax it to 651-367-2620.If you need additional information, please contact CardioVascular LifeLine Technical Services at 877-526-7890, or you local Medtronic sales representative. | Contact the recalling firm for information | 22946 | 22,946 ( 20,714 US, 2,232 OUS) | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and countries of Australia, Belgium, Canada, China, France, Germany, Greece, Hong Kong, Hungary, Japan, Korea, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | December 17, 2011 | 518 | 22946 | Not_Computer
Z-2029-2010 | 56013 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic StealthStation S7 System | Medtronic StealthStation S7 System | 2 | July 17, 2010 | 2010 | Medtronic Navigation, Inc | Excessive wear at the power cord retractor or incorrect assembly of the power plug may cause sparking, intermittent loss of power, or electrical shock. | DESIGN: Device Design | Customers were notified by letter on Jun 14, 2010, advising them of the problem. Instructions were provided to continue using existing retractable power cords until they could be replaced with non-retractable cords by Medtronic, beginning on Aug 17. Further information is available from the company by calling 800-595-9709. |  | 370 | 370 units | Nationwide, including 3 VA facilities. No military distribution. Foreign accounts in Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Namibia, Netherlands, New Zealand, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, United Arab Emirates. | Terminated | February 18, 2011 | 216 | 370 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-2061-2010 | 56004 | prosthesis, vascular graft, of 6mm and greater diameter | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Gold Woven Double Velour Vascular Branch Graft | Hemashield Gold Woven Double Velour Vascular Branch Graft | 2 | July 22, 2010 | 2010 | Maquet Cardiovascular, LLC | Various Hemashield products may develop glycerol droplets on the inner blister of the package if not stored under proper conditions, which could lead to the prolongation of surgery and potentially compromise patient safety. | DESIGN: Packaging Design/Selection | Letters were sent by Federal Express commencing July 12, 2010.Distribution has to be obtained from Boston Scientific.Questions are directed to the local Sales/Service Representative, or to the firm's Customer Service at 1-888-880-2874. | Contact the recalling firm for information |  | xx | Nationwide Distribution. | Terminated | July 14, 2011 | 357 |  | Not_Computer
Z-2073-2010 | 56139 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Mast Roller Pump, Catalog 50-80-00, for use with Sorin HLM C5 and HLM S5. | Mast Roller Pump | 2 | July 22, 2010 | 2010 | Sorin Group USA, Inc. | Although accessories were designed for use with two models, the HLM C5 and the HLM S5, the submission that included references to the Mast Roller Pump, B-Care5, and the Electrical Venous Line Occluder accessories covered the new HLM S5 system. The three accessories were not cleared for use with the HLM C5. | PREMARKET APPROVAL: No Marketing Application | The two customers who had received the suspect device accessories were identified and were contacted by letter on 6/18/2010. They were told to immediately discontinue use of the accessories and arrange for return of the products. They were also told that they could continue to use the CD5 Machine without the recalled accessories. A Product Recall Response Form was to completed by each consignee. Questions are directed to the firm at 800-221-7943. | Contact the recalling firm for information | 1 | 1 unit | CA, PA, TN. | Terminated | July 21, 2010 | -1 | 3 | Not_Computer
Z-2072-2010 | 56056 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂŽ, ver. 4.0.0, 4.0.1, 4.0.2, ThomoTherapy Incorporated 1240 Deming Way, Madison, WI 53717. The TomoTherapy HI-ART SystemÂŽ is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy (IMRT). The HI-ART System provides precise delivery of radiation to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy HiArt System, ver. 4.0.0, 4.0.1, 4.0.2. | 2 | July 22, 2010 | 2010 | TomoTherapy Incorporated | An issue has been identified, as a result of input received from one TomoTherapy Hi-Art Treatment System customer.With a particular workflow, the Planning Station's Draft and/or final Plan Report can provide incorrect dose statistics. | DESIGN: Software Design | Consignees were sent on 6/23/10 a TomoTherapy "Urgent Field Safety Notice Medical device Correction" letter dated June 22, 1010. The letter addressed the Issue, Product Affected, Recommended Action and Resolution.Prior to generating a draft or final Plan Report for a given plan, the plan must be closed and reopened at least once. This issue will be remedied by Hi-Art software version 4.0.3, scheduled for release in late 2010.Customers with questions were provided with a contact in the letter at the TomoTherapy Customer Interaction Center. The firm can be contacted at 608-824-2800. | Contact the recalling firm for information | 29 | 29 | CA, CO, IL, LA, MI, MO, NY, ND, OH, OK, PA,VA, WI, BELGIUM CANADA, KOREA, MALAYSIA, SWITZERLAND, ITALY, GREAT BRITAIN, FRANCE. | Terminated | December 17, 2011 | 513 | 29 | Software | Output/Calculation | Software update | Software Update
Z-2071-2010 | 56059 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Fathom"-14 Steerable Guidewire, Peripheral Vascular Use, Pre-Shaped 300 cm x 10 cm, REF 50-815, UPN M001508150 Sterilized with ethylene oxide gas, Rx Only, Made in 2405 Orion Circle, West Valley City, UT 84119 USA. | Fathom"14 Steerable Guidewire, Peripheral Vascular Use, PreShaped 300 cm x 10 cm | 2 | July 22, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is initiating a Medical Device Recall regarding specific lots of the FathomTM-14 Steerable Guidewire. For lubricity, the distal portion of these guidewires is coated with hydrophilic polymer and the proximal portion is coated with polytetrafluoroethylene (PTFE). Through internal inspection process, the firm determined that Fathom guidewires from affected lots may exhibit a P | DESIGN: Process Design | Consignees were sent on 6/3/10 an "Urgent Medical Device Recall" letter dated 3 June 2010. The letter was addressed to Risk Manger/Interventional Radiology Manger/Material Manger. The letter described the problem and the product involved. The letter included the "Customer Instructions" which stated to Immediately discontinue use of an segregate recalled product and, requested customers to return the Reply Verification Tracking Form. Questions are directed to the firm at 510-624-2563. |  | 78 | 78 units | AL, AR, CA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MD, MA, MS, NV, NY, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT,VA, WA, WT. | Terminated | December 10, 2011 | 506 | 745 | Not_Computer
Z-2076-2010 | 56084 | tube tracheostomy and tube cuff | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-A-HC-PERC8, G12403. | Cook Medical Ciaglia Blue Dolphin | 1 | July 26, 2010 | 2010 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cook Medical sent an Urgent Product Recall letter dated April 21, 2010, to hospitals and other user facilities receiving product direct from Cook as well as distribution centers receiving direct shipments from Cook. Consignees were directed to cease distribution of product and notify their customers immediately. |  | 28819 | 28819 total, all products | Worldwide distribution: USA, Canada, Mexico, Australia, Chile, Costa Rica, Denmark, Dominican Republic, Guatemala, India, Malaysia, Panama, Peru, Puerto Rico, Republic of Korea, Trinidad and Tobago, and Uruguay. | Terminated | January 23, 2012 | 546 | 28819 | Not_Computer
Z-2100-2010 | 56158 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | 19" Barco MFCD 1219 (Touch Screen) + Low Profile Stand and Rack with LCD ArmAgfa's Computed Radiography Systems with NX2.X software are indicated for use in providing diagnostic quality images to aid the physician with diagnosis. | AGFA Computed Radiography system with NX2.X Workstations | 2 | July 26, 2010 | 2010 | AGFA Corp. | Monitor fell from the mounting bracket used to support the monitor. | OTHER/UNDETERMINED: Pending | The firm, AGFA Healthcare sent an 'URGENT SAFETY NOTICE" dated June 30, 2010, via FED-EX to customers. The letter describes product, problem and actions to be taken by the firm and customers. AGFA will provide a safety check to the customers facility, have a service representative visit , in the near future, their site to inspect the monitor/s and attach a warning label on each monitor checked for preventative action. The customers are ask to distribute the information within their facility to all individuals who need to be aware, to complete the attached URGENT SAFETY NOTICE FEEDBACK FORM as soon as possible and return via fax at 864-421-1664. If you have any questions about this matter, please feel free to contact me via email at Debbie.norris@agfa.com or call me at 864-421-1754. |  | 805 | 805 units | Worldwide distribution: USA and countries of Puerto Rico and Canada. | Terminated | April 28, 2011 | 276 | 805 | Not_Computer
Z-2111-2010 | 56085 | guide, needle, surgical | GDF | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Temno Coaxial Chiba Fine Needle Aspiration; a coaxial introducer needle (20G x 15 cm) with a Chiba biopsy needle (22G x 20 cm); sterile, for single use only sets, 5 sets per case; CareFusion, McGaw Park, IL 60085 USA, Made in Dominican Republic; REF CCH2220For use in soft tissue aspiration such as breast, kidney, liver, lung, thyroid, lymph nodes and other various soft tissue masses. It is not intended for use in the bone. | Temno | 2 | July 28, 2010 | 2010 | Carefusion 2200 Inc | The introducer needle is only 10 cm in length instead of the required 15 cm. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Cardinal Health, on behalf of CareFusion, telephoned the customers on June 15, 2010 and sent a follow-up "URGENT: PRODUCT RECALL" letter on the same date via UPS next day air. The letter describes product, problem, and action taken by customers. The customers were instructed to immediately stop using affected product, return the enclosed acknowledgment form via fax to 847-686-9101 (note: Cardinal Health is required to confirm receipt of this notification from the customers), notify any customers to whom they may have distributed product to, contact CareFusion Customer Service at 800-653-6827 to arrange for credit and return of the affected product, and send the affected product to CareFusion, 1240 Waukegan Rd, Waukegan, IL 60085, Attn: Scott Fitzer, after obtaining an RMA/RGA number from CareFusion.If you have any questions, please contact 800-292-9332. |  | 65 | 65 units | Nationwide distribution: California, Florida, Georgia, Illinois, Kentucky, Massachusetts, New Jersey, New York and Pennsylvania. | Terminated | January 24, 2011 | 180 | 65 | Not_Computer
Z-2109-2010 | 56130 | apparatus, exhaust, surgical | FYD | Air-handling apparatus for a surgical operating room. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MICROTEK MEDICAL, INC, STACKHOUSE LASER RESISTANT SUCTION PROBE, Size: 1/4" x 9" (0.64cm x 22.9cm), REF 24014, QTY 10 EA, STERILE EO, 2015-05, Microtek Medical, Inc. , Columbus, MS 39702, www.microtekmed.com, CE0044, EC REP Microtek Medical B.V. , Zutphen, The Netherlands, Manufactured in the U.S.A. Suction probe for a smoke evacuation device | Stackhouse laser resistant suction probe | 2 | July 28, 2010 | 2010 | Microtek Medical Inc | Product labeled as sterile was distributed without sterilization. | OTHER/UNDETERMINED: Pending | The firm notified its consignees of the problem and requested return of the product by telephone on 05/25/2010. They followed with a letter dated on 06/03/2010 re: Product Removal stating the reason for the recall and confirming the receipt of the recalled product. For further information please contact (662) 244-3220. | No consumer action necessary | 2 | 2 cases of 10 units | Nationwide distribution including the states of Montana and Georgia | Terminated | August 12, 2010 | 15 | 2 | Not_Computer
Z-2108-2010 | 56169 | plasma, coagulation control | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | GEM Check Coag CNTRL Abnormal APTT Whole Blood Controls Package Insert; IVD. 06260060100. 15/pk.Manufactured for Instrumentation Laboratory Company, Lexington, MA 02421-3125.Intended to be used to perform quality control assays using GEM test cartridges on the IL GEM PCL/PCL Plus Portable Coagulation Laboratory. | GEM Check Coag Whole Blood Controls | 2 | July 28, 2010 | 2010 | International Technidyne Corporation | Certain lots of GEM Check Coag Whole Blood Controls contain an error in the package insert where the incorrect Acceptable Performance Ranges were published. | PRODUCTION CONTROLS: Error in Labeling | Urgent: Medical Device Correction Notices, dated July 1, 2010, were sent by Federal Express on July 8, 2010 to all first consignee end users. The letters identified the affected product and asked customers to review and follow all instructions as outlined in the ITC in the communication. The info should be shared with laboratory staff and the notification should be retained as part of the laboratory Quality System documentation. The enclosed response form should be completed and returned to ITC within 5 days. | Contact the recalling firm for information | 101 | 101 kits | Worldwide Distribution -- USA, including states of NY, TN, CA, AR, IL and TX and countries of Australia, Italy, Spain, Germany, Chile, Spain, Malta, Belgium, Austria, and Switzerland. | Terminated | September 23, 2010 | 57 | 196 | Not_Computer
Z-2110-2010 | 56114 | assay, heparin | KFF | Heparin assay. | Hematology | Hematology | 510(k) | RQCHRT HRT Control Level 1 & 2;Whole Blood controls for use with Hemochron Kaolin and Celite HRT Test Tubes.20 vials Whole Blood - 2.0mL; dried20 vials Diluent - 4.0 mLThe Hemochron HRT Whole Blood Quality Controls are intended to be used to perform quality control assays using the Hemochron HRT test tubes. | HEMOCHRON(R) HEPARIN RESPONSE TEST | 2 | July 28, 2010 | 2010 | International Technidyne Corp. | An ingredient in the RQCHRT Whole Blood Controls for use with Hemochron Kaolin and Celite HRT Test tubes may become less potent over time and result in excessive High LQC Failures. | OTHER/UNDETERMINED: Pending | ITC sent "URGENT MEDICAL DEVICE RECALL" letters dated June 24, 2010, via Federal Express to Laboratory Director, Clinical Laboratory Manager, or Point of Care Coordinator consignees/customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to check their inventory for the product, if they have boxes or loose vials to please stop using the product, remove it from inventory and contact ITC Technical Support to arrange return. They were also instructed to complete the attached Customer Account Tracking Form, whether or not they have remaining inventory, and return it by fax, email or mail.If you have any questions, please call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (international), Extension 4707 or email techsupport@itcmed.com. | Contact the recalling firm for information | 38 | 38 in US, 3 Out of US | Worldwide distribution: USA including states of CA, NC, TX, FL, MD, OK, NY, and NJ and country of Italy. | Terminated | September 23, 2010 | 57 | 38 | Not_Computer
Z-2112-2010 | 56090 | electrocardiograph | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | IMPAX CV Results Manager/Results Manager Administration Tool | RM, RMAT | 2 | July 29, 2010 | 2010 | AGFA Corp. | Agfa Service technician had incorrectly modified a report template at one site resulting in erroneous anatomic segment locators. | OTHER/UNDETERMINED: Pending | A letter was sent to the consignee on 6/11/2010 indicating what steps would be taken by Agfa concerning the problem. The letter also included a feedback form which was to be returned to Agfa. The recall was expanded and a "URGENT SAFETY NOTICE" was sent to each additional consignee on January 25, 2012. The letter described the potential issue and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the additional consignees. |  | 358 | 358 | AK, AZ, CA, CO, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, MS, MT, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI and Canada | Terminated | July 02, 2013 | 1069 | 358 | Not_Computer
Z-2113-2010 | 56057 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseSirius Anaesthesia System, gas machine. Blease Beech House, Chiltern Court Asheridge Road, Chesham, Bucks HP5 2PX, EnglandIntended for use in the hospital environment and operating room. | BleaseSirius Anaesthesia Machine | 2 | July 29, 2010 | 2010 | Spacelabs Healthcare, Llc | Potential for the touch screen on the ventilator portion of the device to go blank when touched. Possibility that this failure may also cause mechanical ventilation to stop. | OTHER/UNDETERMINED: Pending | On 06/24/2010, Spacelabs began mailing the URGENT-MEDICAL DEVICE CORRECTION letter to their US consignees and e-mailed the same letter to their foreign consignees.The letter describes the touch screen of the BleaseSirius Anesthesia machine goes blank when touched and the possibility that this failure may cause the mechanical ventilation to stop.The consignees are asked to weigh the benefits versus the risk when deciding to use the device or not until the firm upgrades the device. The letter also instructs the consignees to notify their staff about the possibility of the display failure to prevent any adverse event that may occur. The letter states that their field service engineer will contact consignees to schedule a time to visit and upgrade all affected devices at no cost.The US consignees are advised to call 1-800-522-7025, Select 2, and the foreign consignees are advised to call 1-425-657-7200 x5089 for technical support or questions regarding the recall. |  | 80 | 80 units total | Worldwide Distribution -- USA, including states of IA, GA, LA, NC, TN, and TX and countries of AUSTRALIA, CHILE, CHINA, INDIA, ITALY, PHILIPPINES, MEXICO, POLAND, SINGAPORE, THAILAND, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | December 03, 2010 | 127 | 80 | Software | Display/Image | Software update | Software Update
Z-2118-2010 | 55330 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp,Stryker Orthopaedics Passport A.R. Knee Instrumentation Distal Guide Stand; Catalog number 7650-5005. | Passport A.R. Knee Instrumentation Distal Guide Stand | 2 | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | There is the potential for the posts on the Distal Resection Guide Stand to disassociate from the main body part. | DESIGN: Device Design | Stryker Urgent Product Recall letters dated April 8, 2010, were sent by Federal Express to Stryker Branches/agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. |  | 676 | 676 in US, 1305 internationally. | Worldwide distribution. | Terminated | August 28, 2012 | 757 | 676 | Not_Computer
Z-2117-2010 | 55226 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Dimension Vista LOCI Reaction Vessels, catalog number KS855. For in vitro diagnostic use. | Dimension Vista LOCI Reaction Vessels | 3 | August 02, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Vessel pick up errors and vessel jams when the vessel is being loaded. The error halts the instrument and disrupts processing. Vessel flange is thin. | DESIGN: Process Design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated March 2010 informing users of the potential problem, to discard any remaining product, and if needed, to request a replacement lot. |  | 5544 | 5544 bags | Worldwide distribution: USA, Canada, Europe, Japan, Korea, and New Zealand. | Terminated | August 03, 2010 | 1 | 5544 | Not_Computer
Z-2119-2010 | 55357 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | X-Celerate Instrumentation Punch Thru Tibial Base Plate, Catalog numbers: 3750-0003. 3750-0004, 3750-0005, 3740-0006, 3750-0007, 3750-0009, 3750-0011, 3750-0013; Stryker Howmedica Osteonics Corp Scorpio Total Knee System | Scorpio | 2 | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The locating pins on the Punch Thru Tibial Baseplates could break or disassociate from the main body of the instrument. | DESIGN: Device Design | Stryker Urgent Device Correction letters dated April 5, 2010, were sent to all Stryker branches/agencies, hospital Risk Management and surgeons who have used the product by Federal Express. Questions on the bulletin should be directed to Scott Solis, Product Manager ,Knee Marketing at 201-831-5665. |  | 8276 | 8276 | Worldwide distribution: USA, China, Taiwan, India, Australia, Korea, Singapore, Germany, France, Spain, Italy, Switzerland, Poland, Romania, South Africa, United Kingdom, Netherlands, Sweden, Turkey, Japan, Chile, Brazil, Colombia, Mexico, Panama, Argentina, and Uruguay. | Terminated | August 28, 2012 | 757 | 8276 | Not_Computer
Z-2164-2010 | 55605 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MIKA Speedblock, Size 8-12.The MIKA Speedblock is placed on the distal femur as a guide for making the anterior, posterior and both chamfer cuts. | MIKA Speedblock | 2 | August 02, 2010 | 2010 | Encore Medical, Lp | A crack or complete fracture at the anterior chamfer cut slot may occur on the size 8, 10, and 12 MIKA Speedblocks. | DESIGN: Component Design/Selection | An "Urgent Field Safety Notice" Letter dated May 20, 2010, was sent to the customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to locate, inspect and verify the quantity still in their possession, complete and return the field safety notice response page to Beth Onderlinde at beth.onderlinde@djosurgical.com or fax to 512-834-6313 by June 7, 2010, (completion of form is mandatory), and identify and inform customers that they may have distributed/sold product to, of the notice. The customer was also instructed to please inspect the listed part numbers prior to use and again after use, if a crack is found in the speedblock, contact your Customer Service Representative and be sure to utilize the femoral/tibial impactor (801-01-043) for impaction of the speedblock and reiterate to your physicians that the MIKA mini slaphammer (800-02-290) should be used for speedblock removal. The product will be replaced as it becomes available.If you have any questions, call Director of Commercial Logistics at (512) 834-6330 or International Customer Service at (512) 834-6275. You may also contact our EU Authorized Representative, MDSS GmbH, via email at info@mdss.com or via phone at +49 (0) 511 6262 8630. | Contact the recalling firm for information | 796 | 729 and 67 specials | Worldwide distribution. | Terminated | December 15, 2011 | 500 | 796 | Not_Computer
Z-2140-2010 | 55931 | container, i.v. | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Exacta-Mix EVA Container, Sterile R, 2000 mL, Order No. REF: 140, Baxa Corporation. | ExactaMix | 2 | August 02, 2010 | 2010 | Baxa Corporation | Fill port cap on TPN bags may become detached prior to use, thus compromising the sterility of the fluid pathway. | PRODUCTION CONTROLS: Packaging | Consignees were notified by e-mail and Baxa Bag Product Notice letter on 05/26/2010 with instructions to not use any bags if the fill port cap is detached. Photos of a detached cap and properly attached cap were included. Defective bags could be returned to Baxa for credit. Information available at 800-678-2292. |  | 3306700 | 3,306,700 bags | Worldwide distribution, including USA, Argentina, Australia, Bermuda, Brazil, Canada, China, Colombia, Hong Kong, Malaysia, Mexico, New Zealand, Pakistan, Panama, South Korea, Taiwan, United Kingdom, and Uruguay. | Terminated | August 04, 2011 | 367 | 12612066 | Not_Computer
Z-2133-2010 | 55621 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD DORADO PTA BALLOON DILATATION CATHETERS, PRODUCT CODE # DR135712 | BARD DORADO | 2 | August 02, 2010 | 2010 | Bard Peripheral Vascular Inc | The affected DORADO PTA Balloon Dilatation catheters contain catheter materials that, under certain conditions, are subject to premature breakage allowing the balloon to partially or completely separate from the catheter shaft. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The Bard Peripheral Vascular "URGENT MEDICAL DEVICE RECALL NOTIFICATION" with attached "RECALL AND EFFECTIVENESS CHECK FORM" was sent to the affected consignees beginning May 7, 2010 via FedEx. The recall notice informed the consignees of the reason for recall, the products affected and instructed them not to use or further distribute any affected product.The consignees were also instructed to call the firm's Recall Coordinator Silvia Carrillo at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at silvia.carrillo@crbard.com to facilitate the expedient return of the affected products. Additionally they were asked to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. | Contact the recalling firm for information | 72 | 72 units | Nationwide | Terminated | February 01, 2012 | 548 | 3279 | Not_Computer
Z-2120-2010 | 55367 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp., Stryker Orthopaedics Passport A.R. Knee Instrumentation Anterior Skim Guide; Catalog Number 7650-5003. | Stryker Orthopaedics Passport A.R. Knee Instrumentation | 2 | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | One of the two posts on the Anterior Resection Guide could disassociate from the main body of the part which has been determined to be caused by a fracture weld. | DESIGN: Device Design | Stryker Urgent Product Recall letters dated April 5, 2010, were sent out via Federal Express to Stryker branches/agencies, Hospital Risk Management, hospital Chief of Orthopedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. Questions and/or concerns should be directed to Rich Wolyn at 201-972-2100. |  | 682 | 682 units in US, 1292 units internationally. | Worldwide distribution: USA, Ireland, Germany, Sweden, The Netherlands, France, Italy, Switzerland, South Africa, United Kingdom, Japan, Argentina, Mexico, Miramar (city), China, Taiwan, India, Australia, Korea, New Zealand, Malaysia, Singapore, and Canada. | Terminated | August 28, 2012 | 757 | 682 | Not_Computer
Z-2116-2010 | 55874 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | PowerPark is an accessory to the V or H Universal Stand and comprised of two parts: the stand mounted assembly (P12822) and a docking assembly (P12834). The stand mounted assembly is mounted to the V or H Universal stand and the docking assembly is plugged into a wall circuit. It provides a docking and charging station for the V and H Universal Stands.The product marketing brochure labeled in part: "SonoSite Ltd...21919 30th Drive SE... Bothell, WA 98021, USA". | PowerPark | 2 | August 02, 2010 | 2010 | Sonosite, Inc. | PowerPark system has a potential for loss of power at the corresponding wall outlet and/or circuit due to poor alignment. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 05/27/10, SonoSite began contacting their consignees by telephone. On 06/01/10, SonoSite sent out the URGENT-MEDICAL DEVICE RECALL letter to their consignees. The letter describes the recall product as SonoSite PowerPark. Customers should call the Technical Support at 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall.The letter instructs consignees to stop using the PowerPark immediately and return the product back to SonoSite in exchange for a full refund. The consignees will be contacted by the firm's Sales Representative and/or Clinical Applications Specialist to schedule an appointment to retrieve the device and deliver a refund. | Contact the recalling firm for information | 26 | 26 | Worldwide Distribution -- USA, including states of GA, NM, PA, TX, VA, and WA and country of Australia. | Terminated | April 18, 2011 | 259 | 26 | Battery
Z-2115-2010 | 56052 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂŽ, Version 4.0.0 & 4.0.1. TomoTherapy Incorporated 1240 Deming Way, Madison, WI 53717. Intended to be used as an integrated system for planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy HiArt System, Version 4.0.0 & 4.0.1. | 2 | August 02, 2010 | 2010 | TomoTherapy Incorporated | In some cases, the patient's diagnostic CT image is narrower than the Hi-Art radiotherapy couch image. | DESIGN: Software Design | Consignees were sent on 2/22/10 a TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated Monday, February 22, 2010. The letter was addressed to TomoTherapy Customer and identified the the affected product, and included a description of the Issue, Product Affected, Recommendation Actions and Resolution. The issue will be remedied by software versions 4.0.2 and later. Questions should be directed towards TomoTherapy Customer Interaction Center by e-mail or telephone. | Contact the recalling firm for information | 7 | 7 | Worldwide Distribution -- MI, MO, VA, and WI and country of CANADA. | Terminated | February 10, 2011 | 192 | 7 | Software | Display/Image | Software update | Software Update
Z-2151-2010 | 56037 | heat-exchanger, cardiopulmonary bypass | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Procedure Kit Custom MP4, 4:1, PRELIEF, W/RECIRC, COND / PUMP & TABLE PACKP/N 65503/70881-02 | Terumo Custom Kit | 2 | August 02, 2010 | 2010 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid | PRODUCTION CONTROLS: Equipment Maintenance | Terumo, Ashland, MA notified consignees on 6/8/10, byphone, fax, email and mail to advise distributors and end-users of the required recall actions. Follow-up communication sent by priority overnightmail on 6/9/10, with proof of delivery, specifying the awareness to the issue and necessary correctionactivities should a unit be discovered. Phase 2 initiated on 7/1/10 to sale rep to remove recalled product and perform a product reconciliation to ensure that all affected product is identified, returned to Terumo CVS, and that replacement product is sent immediately. | Contact the recalling firm for information | 48 | 48 packs | Distribution USA and Canada | Terminated | May 07, 2012 | 644 | 589 | Not_Computer
Z-2121-2010 | 56063 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MEDFUSIONÂŽ SYRINGE INFUSION PUMP, MODELS 3010 AND 3010a WITH SOFTWARE VERSIONS 2.0.2, 2.0.3, AND 2.0.4 MedfusionÂŽ Syringe Infusion Pumps are designed for use in critical care areas for the administration of fluids requiring precisely controlled infusion rates. | MEDFUSION SYRINGE INFUSION PUMP, | 2 | August 02, 2010 | 2010 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a correction of MedfusionÂż Syringe Infusion Pumps, Models 3010 and 3010a, with obsolete software versions 2.0.2, 2.0.3, and 2.0.4. Smiths Medical has found that pumps with obsolete software may continue to run past the set volume limit. This can result in an over-delivery  if the syringe is over-filled. Smiths Medical has received no reports of serious patient in | DESIGN: Software Design | Consignees were sent on 6/22/10 a Smith Medical "Urgent Medical Device Correction Notice', dated June 11, 2010. The letter was addressed to Dear Risk/ Safety Managers, Distributors, Clinical/ Biomedical Engineering, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists; and Nursing, Pharmacy,and Anesthesia Professionals. The letter described the problem and product involved. It provided Advise on Action to be Taken by the User which included to 1) Determine the software version of their Medfusion pump 2. Complete and return the Confirmation Form 3. Return the completed Software Update Test Form 4. If unable or unwilling to perform the software upgrade they provided information for returning the pumps to Smiths Medical for the upgrade. For further information please contact Smiths Medical at 651-628-7280. | Contact the recalling firm for information | 4816 | 4816 | Worldwide Distribution: AL, AK, AZ, AR, CA,CO, CT, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV,NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI. and countries of CANADA, MEXICO, NEW ZEALAND, BRASIL, HONG KONG, and GREAT BRITAIN. | Terminated | April 17, 2012 | 624 | 4816 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2166-2010 | 55507 | system, measurement, blood-pressure, non-invasive | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | The Well@Home Monitor by Patient Care technologiesModel 179-2150Intended to be used as a clinical tool for remotely monitoring a patient's health status at home or in another healthcare facility on a spot check basis. | WellHome Monitor | 2 | August 03, 2010 | 2010 | Zoe Medical Incorporated | Touch panel may become unresponsive. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Zoe Medical notified the 2 distributors by e-mail on January 26, 2009 and Feb 6, 2009 of the recall, requesting return of the units for correction. |  | 62 | 62 | Nationwide Distribution -- MA and MD. | Terminated | August 03, 2010 | 0 | 62 | Not_Computer
Z-2172-2010 | 55871 | sirolimus test system | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Architect Sirolimus Reagent Pack; 100 test pack containing 1 bottle (8.0 mL) Anti-Sirolimus Coated Microparticles, 1 bottle (8.0 mL) Sirolimus Acridinium-labeled Conjugate in Citrate Buffer, and 1 bottle (10 mL) Assay Diluent containing saline; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 01L76-25.The Architect Sirolimus assay is chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of sirolimus in human whole blood on the Architect i System, as an aid in the management of renal transplant patients receiving sirolimus therapy. | Architect Sirolimus Reagent Pack | 2 | August 03, 2010 | 2010 | Abbott Laboratories | The Architect Sirolimus assay lot 80162M100 is exhibiting higher than normal reports of barcode read errors. | PRODUCTION CONTROLS: Equipment Maintenance | Abbott sent a "Product Recall" letter dated May 27, 2010, to all customers receiving the affected lot of reagent. The customers were informed of the increased frequency of barcode read errors due to the print quality of the barcode label. The following error codes have been observed for some kits within the lot: 4000 - Unable to read reagent bar code in position (x) on (y) carousel, 0208 - Reagent carousel scan error, and 2011 - Extra reagent bottle detected in position (x) on (y) carousel. The customers were requested to determine if they are currently using or have inventory of lot 80162M100, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside the U.S., please contact your local Customer Service representative. | Contact the recalling firm for information | 875 | 875 kits | Worldwide distribution: USA including Puerto Rico, and countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, United Kingdom, Venezuela and Vietnam. | Terminated | July 15, 2011 | 346 | 875 | Not_Computer
Z-2171-2010 | 55557 | sirolimus test system | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | IMx Sirolimus Reagent Pack; 100 test pack containing 1 bottle (9.7 mL) Anti-Sirolimus Antibody Coated Microparticles, 1 bottle (9.7 mL) Sirolimus Alkaline Phosphatase Conjugate, and 1 bottle (10 mL) 4-Methylumbelliferyl Phosphate, 1.2 mM; Produced by Axis-Shield Diagnostics, Ltd., Dundee, UK for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA; list 5C91-21The IMx Sirolimus assay is an in vitro reagent system for the quantitative determination of sirolimus in human whole blood as an aid in the management of renal transplant patients receiving sirolimus therapy. | IMx Sirolimus Reagent | 2 | August 03, 2010 | 2010 | Abbott Laboratories | IMx Sirolimus reagent lot 802873106 has exhibited an increased frequency of calibration errors. | OTHER/UNDETERMINED: Pending | Abbott sent a "Product Recall" letter dated April 30, 2010, to all customers receiving the affected lot of reagent. The accounts were informed of the increased frequency of calibration errors, including: 157 Check 1 Out of Range, 161 Check 5 Out of Range, and 173 Mode 1 deviation Factor High. The accounts were requested to determine if they are currently using or have inventory of lot 802873106, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside of the U.S., please contact your local Customer Service Representative. |  | 466 | 466 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Austria, Bahrain, Brazil, Canada, Chile, China, Croatia, Cyprus, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Jordan, Lebanon, Lithuania, Netherlands, Norway, Philippines, Poland, Russia, Panama, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE and Venezuela. | Terminated | May 13, 2011 | 283 | 466 | Not_Computer
Z-2170-2010 | 56018 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Basic Care Bed, P1441A (Non-electric), Hill Rom Batesville, IN. Intended for low to moderate acuity patients in the low end acute care area of the hospital. | HillRom Basic Care Bed | 2 | August 03, 2010 | 2010 | Hill-Rom, Inc. | The flexing of the junction of the sleep deck's head and mid sections can create excess pressure on the pivot pin E-ring, which could cause the E-ring to come off. |  | The firm sent an Urgent Field Safety Notice to its customers dated 9/16/2009. The letter identified the affected product, explained the reason for the recall, actions to be taken, and a contact reference. Customers are to locate the affected product, examine the pivot points and ensure the E-ring is present. If the E-ring is missing, then the affected bed is to be taken out of service. Customers' distributors or agents will contact them with a permanent fix. Questions should be directed towards Hill-Rom Technical Support. | Contact the recalling firm for information | 962 | 962 beds | Worldwide Distribution -- Argentina, Bahrain, Bhutan, Brazil, Bulgaria, Chili, China, Columbia, Costa Rica, Ecuador, Egypt, Eritrea, France, Hong Kong, India, Indonesia, Iran, Iraq, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mexico, Morocco, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Syria, Taiwan, Thailand, Trinidad & Tobago, Turkey, United Arab Emirates, United States, Uruguay, Venezuela, and Vietnam. | Terminated | March 16, 2012 | 591 | 11056 | Not_Computer
Z-2165-2010 | 56055 | stimulator, auditory, evoked response | GWJ | Evoked response auditory stimulator. | Neurology | Neurology | 510(k) | SonaMed Clarity System II; evoked response auditory stimulator; SonaMed Corp., Waltham, MA 02154 | SonaMed Clarity System | 2 | August 03, 2010 | 2010 | Natus Medical Inc | There is insufficient evidence to support any claims for safety and efficacy on the SonaMed Clarity devices. | PREMARKET APPROVAL: No Marketing Application | Natus sent Urgent Medical Device Recall letters dated 6/4/10 to all known SonaMed Clarity device accounts, to the attention of the Director of Maternal Child Health and the Risk Manager. The letters referenced the letters previously sent to them by Natus in October 2008 and April 2010 informing them of Natus' concerns about whether the Clarity device could be lawfully used to comply with FDA-related requirements and other pertinent professional guidelines/standards for newborn infant hearing screening, and that they concluded that the Clarity devices should no longer be used. Natus recognized the need for the accounts to continue screening without interruption, and are providing on loan at no-charge handheld screener(s) to facilities that do not have alternate screening devices. The accounts were requested to contact Natus Technical Service group at 800-272-8075 to discuss this loaner device option. Natus requested that the accounts stop using the SonaMed Clarity devices; disable the screening function of their Clarity system(s) by returning to Natus, in the enclosed pre-addressed return envelope, either of the two probes (TPI-830- 8-pin ear probe, or the TPI-815- 7-pin ear probe) they may possess and have used in screening, making sure to return all 7- or 8-pin probes for all of your Clarity systems; and complete the attached "Device Recall Reply Form" and return it in the enclosed self-addressed stamped envelope. If you have any questions concerning any aspect of this recall notice, please call Natus Technical Support at 1-800-272-8075. |  | 284 | 284 units | Nationwide Distribution | Terminated | June 22, 2011 | 323 | 284 | Not_Computer
Z-2173-2010 | 55775 | prosthesis, toe, constrained, polymer | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Sgarlato Labs GAIT (Great Toe Implant Technique) Implant-- Large, Product Number 30-100-03 Large, Manufactured by Sgarlato Med LLC, San Jose, CA | Sgarlato Labs GAIT (Great Toe Implant Technique) Implant | 2 | August 04, 2010 | 2010 | Sgarlato Med, LLC | Products with an expiration date of August, 2009 were labeled with an expiration date of May 2013. Additionally, the outer packages were labeled with an incorrect lot number. The outer and inner packages did not match. | PRODUCTION CONTROLS: Labeling Mix-Ups | Recall initiated in March 2009 via telephone calls. |  | 27 | 27 units | Nationwide Distribution | Terminated | August 19, 2011 | 380 | 27 | Not_Computer
Z-2174-2010 | 55813 | reagents, clostridium difficile toxin | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Catalog 441401, 200 tests/kit, labeled, in part ***BD Diagnostics, 2555 Boul. Du Parc-Technologique, Quebec, Canada G1P4S5***Intended for use as an aid in diagnosis of CDAD. | BD GeneOhm Cdiff Assay | 3 | August 05, 2010 | 2010 | Becton Dickinson & Co. | In vitro diagnostic test kits do not perform to specifications. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recalling firm notified accounts by an Urgent Product Recall letter on 03/05/2020. The letter identified the affected product and explained the reason for recall. It also advised customers to discontinue use of the affected product and discard any remaining inventory and fax back an acknowledgement form. The letter stated that BD will provide replacements for any discarded product. Further assistance can be obtained by contacting BD Technical Service at 1-888-436-3646. |  | 31 | 31 kits | Nationwide Distribution -- IL, IN, PA, WA, CA, VA, NJ, and KS. | Terminated | October 15, 2010 | 71 | 31 | Not_Computer
Z-2175-2010 | 55908 | apparatus, automated blood cell diluting | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks Auto Prep System, Circulating Tumor Cell Kit (epithelial, IVDVeridex, LLCa Johnson & Johnson CompanyIntended for the enumeration of circulating tumor cells (CTC) of epithelial origin in whole blood. | CellTracks Auto Prep System | 2 | August 05, 2010 | 2010 | Veridex, LLC | Complaints of patient sample carryover. | PRODUCTION CONTROLS: Process Control | On 19 May 2010, all consignees were notified, US and foreign, via Urgent Product Correction Notification letters. The letter identified the affected product and explained the reason for recall. Customers were asked to follow the required actions provided in the letter, including completing and returning the attached Confirmation of Receipt form. Questions should be directed to Customer Technical Services at 1-877-837-4339. |  | 101 | 101 Domestically, 87 Internationally | Worldwide distribution -- USA, China, Japan, England, France, Germany, Italy, Spain and Canada. | Open |  |  | 101 | Not_Computer
Z-2177-2010 | 56049 | accessories, implant, dental, endosseous | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Impression set for WN solid abutment H 4.00mm POM/ AI. Ref : 040.380 Kit Y4021.Intended to aid in the fabrication of dental prosthetics, and be used as a accessory with endosseous dental implants. | Impression Set, dental | 3 | August 06, 2010 | 2010 | Straumann Manufacturing Inc. | The label on the tubular packaging for the Straumann Impression set f. WN solid abutment has the incorrect description of RN solid abutment instead of a WN. | PRODUCTION CONTROLS: Labeling Mix-Ups | Field Safety Notice, dated June 15, 2010, issued to consignees states that the 040.380 Impression set f. WN solid abutment Lot Y4021 which you purchased is incorrectly labeled Impression set f. RN solid abutment. Customers were asked to return the affected products and completed Customer Confirmation Form for replacements. Questions should be directed towards Bernard McDonald at 978-747-2514. |  | 163 | 163 | Nationwide Distribution | Terminated | September 22, 2011 | 412 | 163 | Not_Computer
Z-2191-2010 | 49374 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SonoSite TEE, 8-3 MHz Transducer labeled: "Tee Transducer... SonoSite, Inc. 21919 30th Drive SE Bothell, WA 98021".The TEE/8-3 MHz Transducer is designed for 2D, M Mode, color Doppler, pulse wave (PW) Doppler and continuous wave (CW) Doppler by applying ultrasound energy through the esophagus or stomach of the patient into the heart. The TEE transducer is intended to be used on adults only. Backscattered ultrasound energy from the patient's heart is used to form images of the heart to detect abnormalities in structure or motion, to evaluate the velocity of blood flowing within the heart, and to obtain a color depiction of the velocities of blood flowing in the heart. | Transesophegeal Ultrasound Transducer | 2 | August 09, 2010 | 2010 | Sonosite, Inc. | The TEE transesophageal transducer, when connected to a SonoSite M-Turbo Ultrasound System, may exhibit erratic operation of the Scanplane Orientation Indicator on the system display. The indicator may move erratically and not accurately represent the orientation of the TEE transducer array. | OTHER/UNDETERMINED: Pending | On August 21, 2008, SonoSite began calling 15 consignees who have purchased the TEE/8-3 MHz Transducer and an M-Turbo 1.1 system. The firm provided instruction to discontinue use the system to 11 consignees and left voice messages to the remaining four consignees.On August 21, 2008, SonoSite sent out the URGENT-MEDICAL DEVICE CORRECTION letter and the Urgent Field Safety Notice to all affected consignees in the US and in foreign countries.The URGENT-MEDICAL DEVICE CORRECTION letter describes the recall product as TEE/8-5 MHz Transducer, SonoSite Part Numbers P05183 and P06841. The letter informed the consignees that the TEE transesophageal transducers may exhibit erratic operation of the Scanplane Orientation Indicator on the system display when it connected to the SonoSite M-Turbo Ultrasound System. The letter instructs consignees to discontinue use of the TEE/8-3MHz Transducer with the M-Turbo (or Micromaxx system) immediately, while the firm is working to create a software solution to resolve the issue and offers customers a software upgrade within 6-8 weeks.The consignees are also advised to forward the recall notice to others who need to be aware of the issue or where the potentially affected devices have been distributed.Customers should call the Technical Support at 1-1-425-951-1330 or 1-877-657-8118 from Monday to Friday, 6 am-5 pm PST or e-mail to service@sonosite.com for any questions regarding the recall. | Contact the recalling firm for information | 357 | 357 units | Worldwide distribution: USA and countries including Argentina, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Finland, France, Germany, India, Iran, Ireland, Israel, Italy, Korea, Netherlands, New Zealand, Poland, Slovenia, South Africa, Spain, Switzerland, and Thailand. | Terminated | August 24, 2011 | 380 | 357 | Software | Display/Image | Software update | Software Update
Z-2183-2010 | 55139 | stretcher, wheeled, powered | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Stryker Atlas Transport Stretcher with Zoom, Model 660Z, Stryker Medical, Portage, MI |  | 2 | August 09, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The retractable cord stretchers were fitted with power cords, manufactured by Electri-Cord, whose prongs have the potential to fracture inside the molded section of the plug, causing melting and charring. | DESIGN: Component Design/Selection | The firm has sent Urgent - Medical Device Recall letters, dated March 19, 2010, to all direct accounts notifying them of the recall. The letter identified the affected product and described the reason for recall. Customers were asked to locate the stretchers identified in the letter and visually inspect each cord. If the plug shows signs of excessive wear, etc. the stretcher is to be taken out of service immediately and the plug removed. Customers are to call the number provided in order to be placed on the priority replacement list. If the cord does not show signs of damage then the stretcher can remain in service until a replacement cord is available. The enclosed post card should be completed and returned to the firm. A copy of the notice should be forwarded to any users that stretchers have been sold or loaned to. The firm should be provided with the serial numbers of any stretchers that have been disposed of. Questions should be directed towards Stryker at 800-STRYKER. |  | 61 | 61 stretchers | Worldwide Distribution -- USA, Canada, Brazil, Japan, Singapore, Spain, and the Far East. | Terminated | May 02, 2011 | 266 | 4915 | Not_Computer
Z-2188-2010 | 56091 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAWÂŽ HEX SCREW CHARLOTTE " F&A; SYSTEM CONTENTS: 1 EACH Foot and Ankle Screw ... IMPLANT MATERIAL: Stainless Steel Ref: 4013-3520 ... Size 3.5MM, Length 20 MM, Use with ClavaÂŽ 3.5MM - Non-Sterile. Implantation: used in Food and Ankle surgeries. | CLAW HEX SCREWS | 2 | August 09, 2010 | 2010 | Wright Medical Technology Inc | One lot of CLAW Hex Foot and Ankle Screws was incorrectly packaged containing Multi-use Compression Screws. | PRODUCTION CONTROLS: Packaging Process Control | The firm initiated their recall to their sales representatives (distributors) via telephone on 06/07/2010 and followed with a letter dated June 15, 2010 sent via Fed Ex on 06/28/2010. The hospital involved was also notified by letter on 06/29/2010. The letter described the product, the problem,and the action to be taken by the customers. The customers were instructed to immediately return any affected product to Wright as soon as possible, and to complete an attached response form confirming receipt of the notice and return a copy by fax to 901-867-7401. Customers may contact customer service for return instructions and replacement inventory at 800-238-7117.If you have any questions regarding this matter, contact Debby Daurer at 800-874-5630. |  | 8 | 8 units | Nationwide Distribution in States of FL, PA, KS, IN, MO, and CA | Terminated | November 23, 2011 | 471 | 8 | Not_Computer
Z-2190-2010 | 56115 | catheter, embolectomy | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Corporation, QuickCat Extraction Catheter, REF 60090-01, Sterile R, Working Length 145cm, for embolus and thrombus removal from arterial vessels. | QuickCat | 2 | August 09, 2010 | 2010 | Spectranetics Corporation | Guidewire lumen may be occluded, preventing guidewire insertion into the catheter. |  | Consignees were notified by letter on Jul 1, 2010, and told to not use affected lots and to contact Spectranetics to arrange a product exchange. More information is available at 800-231-0978. | Contact the recalling firm for information | 5080 | 5080 units (1778 US, 3302 Non-US) | Worldwide distribution: USA, Argentina, Austria, Belgium, Brazil, Chile, Czech Republic, Finland, France, French Polynesia, Germany, Hungary, Israel, Italy, Netherlands, Poland, Spain, Sweden, Switzerland, and Venezuela. | Terminated | August 20, 2010 | 11 | 5080 | Not_Computer
Z-2192-2010 | 56163 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | REF 87100 87k Arthroscopy Tubing Set. For use with the 87k Arthroscopy Pump (87000). STERILE EO. Rx ONLY. Warning: Do not use this tubing set with any Small Joint Cannula System.CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA. Made in USA.The 87K Arthroscopy Tubing Set is designed for use with the 87K Arthroscopy Pump System. This system can be used in arthroscopic procedures using fluid irrigation for joint distension. This system may be used where gravity systems are currently in use such as knee, shoulder, ankle and elbow arthroscopy. | 87k Arthroscopy Tubing Set | 2 | August 09, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec Largo, FL is recalling 87K Arthroscopy Tubing Set, Product Number 87100, Lots 0904151 to and including 1005201. Product may have a breach in the seal that could potentially compromise the sterility of the contents. | PRODUCTION CONTROLS: Process Control | ConMed Linvatec Corporation sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 6, 2010, via FedEx next day delivery to all customers. The letter described the product, problem and action to be taken by customers within 30 days of receipt of the letter. The customers were asked to PLEASE DISCONTINUE USE OF THE IDENTIFIED PRODUCT WITH LOT NUMBERS BEGINNING WITH 0904151 THROUGH AND INCLUDING 1005201; to immediately check their inventory for the affected lots number and Do Not Use if they have any of the products; to complete and return a Reply Form and to notify their customers if the product was further distributed and return all unused affected product with lot codes starting 0904151 to and including 1005201, to ConMed Corporation, 525 French Road, Utica, NY 13502. On receipt, ConMed Linvatec will issue credit to customers.Should you have any additional questions, please contact our Customer Service department at 800-535-8536 or fax at 727-319-5701 or email: Custerserv1@linvatec.com. | Contact the recalling firm for information | 97830 | 97,830 | Worldwide distribution: USA and Canada. | Terminated | February 24, 2011 | 199 | 97830 | Not_Computer
Z-2193-2010 | 56321 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Crossfire Console, model number 475-000-000, Manufactured by Stryker Endoscopy San JoseMedical device indicated for use in orthopedic and arthroscopic procedures for the knee, shoulder, ankle, elbow, wrist and hip. The system provides abrasion, resection, debridement, and removal of bone and soft tissue through its shaver blade and ablation and coagulation of soft tissue, as well as hemostasis of blood vessels through its electrosurgical probe. Examples of use include resection, ablation and coagulation of torn knee cartilage, subacromial decompression and resection of synovial tissue in other joints. The probe is contraindicated for use in procedures where a nonconductive irrigant is used or with patients having cardiac pacemakers or other electronic implants. | Crossfire Console | 2 | August 09, 2010 | 2010 | Stryker Endoscopy | RF energy or powered shaver handpieces may continue to operate when not intended when used with the iSwitch Wireless Foot switch. | DESIGN: Software Design | Stryker Endoscopy sent an "Urgent: Device Correction" letter dated June 28, 2010, via FedEX Mailer to US customers. A letter dated July 9, 2010, was sent to International sites.The letter describes the product, problem, and action to be taken by customers. The Customers were instructed to follow the included instructions and use the USB dongle to install the correction and to complete and sign the enclosed "acknowledgement of receipt" form,upon completion of the software correction, and fax to (480) 754-8378 or scan in and email to crossfireconsole@stryker.com. The firm noted that-It is important to send the self addressed confirmation forms back to Stryker RAQA. If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at crossfireconsole@stryker.com. |  | 989 | 989 units | Worldwide distribution: USA and countries including Australia, France, Germany, Greece, South Africa, Iberia, Italy, New Zealand, Poland, UK, Switzerland, Canada, China, Japan and the EMEA. | Terminated | March 21, 2011 | 224 | 989 | Software | Physical Safety Hazards | Software update | Software Update
Z-2194-2010 | 48998 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The Philips EasyDiagnost Eleva with DRF. Philips Medical Systems multifunctional x-ray system. | EasyDiagnost Eleva | 2 | August 10, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential for bright artifacts to appear on the patient image from a previous exposure. | OTHER/UNDETERMINED: Pending | On July 9, 2008, Philip Medical System began sending out the URGENT-Device Correction letter to their consignees. This letter describes the recalled product as EasyDiagnost Eleva with DRF system with SW release 1.0.The firm informs the consignees that under special circumstances there is a small possibility that bright artifacts appear on a patient image from a previous exposure.The consignees are advised to take care about the X-ray field limitation. If the detector is not affected by direct radiation, no problem will occur. If direct radiation reaches the detector, the collimation should not be too tight, such that the area of direct radiation is larger than 1.2 inch x 1.2 inch (or 30 mm x 30 mm). The minimum area is needed to detect and prevent memory artifacts.The firm will upgrade any affected products with the Field Change Order (FCO)70600031 and it will be implemented before the end of December 2008. The upgrade software is 1.0.3. The correction would be free of charge and the consignees should contact Philips Call Center at 800-772-9377, #5, #2. |  | 7 | 7 | Product was distributed in AZ, CA, FL and WA. | Terminated | October 27, 2010 | 78 | 7 | Software | Display/Image | Software update | Software Update
Z-2195-2010 | 55787 | nephelometer, for clinical use | JQX | Nephelometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMAGE Immunochemistry Systems Buffer 1 (BUF1), Part Number: 447650 | IMMAGE IMMUNOCHEMISTRY SYSTEM | 2 | August 10, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall of IMMAGE Immunochemistry Systems Buffer 1 (BUF1) due to received reports of increased recovery of control andpatient samples using the lots of IMMAGE BUF1. Some lots of BUF1 have been reported to cause high shifts in recovery of control or patient samples. Quality Control failures may result after replacing BUF1 with a different lot. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter will send a Product Corrective Action (PCA) letter on the week of April 19, 2010 to the affected customers. The letter provided the customers with an explanation of the problem and a work around. | Contact the recalling firm for information | 1309 | 1309 | Worldwide: U.S., Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, France,Germany,Greece,Hong Kong, Iceland,Ireland,Italy,Japan,Liechtenstein,Macao,Malaysia,Netherlands,New Zealand,Norway,Oman,Poland,Qatar,Russian Federation,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tunisia,Turkey andUnited Kingdom, | Terminated | May 22, 2012 | 651 | 1309 | Not_Computer
Z-2202-2010 | 56092 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Intellivue Patient Monitor Model MP70 (M8007A) | Intellivue | 2 | August 10, 2010 | 2010 | Philips Healthcare Inc. | Additional Instruction for Use for Intellivue MP60/MP70 to prevent monitors from becoming loose and detaching from the mounting arm. | DESIGN: Labeling Design | Philips issued on 6/28/10 an Urgent Medical Device Correction notification/Field Safety Notice. Included with the Urgent Medical Device Correction notification/Field Safety Notice is an updated instructions for use and service instructions for these monitors that remind users to follow themounting arm manufacturer's instructions for use when repositioning the monitor and instructing them toperform a periodic check that the monitor is not loose on the arm. |  | 82019 | 82,019 units | Worldwide distribution: USA, Canada, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, Bosnia, BRAZIL, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRI, LITHUANIA, LUXEMBOURG, MACAU, MALAYSIA, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA & MONTE NEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, United Arab Emirates, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, VIRGIN ISLANDS (BR), and YEMEN. | Open |  |  | 93548 | Not_Computer
Z-2204-2010 | 55909 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard CR femoral 57.5 mm left interlok, Sterile, REF 183022, Knee prosthesis component. | Vanguard | 2 | August 10, 2010 | 2010 | Biomet, Inc. | A knee implant labeled as left was opened and the package actually contained a right knee implant. | PRODUCTION CONTROLS: Process Control | On 4/13, 14 and 15 Biomet made phone calls to the consignees to inform them of the recall. The firm sent Urgent Medical Device Recall Notices dated 4/20/2010. The customers were instructed to immediately locate the recalled products and return them to the manufacturer. |  | 5 | 5 | 1 unit in US. 3 units in Biomet distributor inventory, 1 unit in Japan | Terminated | January 19, 2011 | 162 | 5 | Not_Computer
Z-2200-2010 | 56117 | plasma, coagulation control | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | International Technidyne Corporation (ITC) directCheck Whole Blood Quality Controls Plasma, Coagulation Control. Catalog number: DCJLR-N; in vitro diagnostic15 vialsFor performance of Level 1 quality control ACT-LR test cuvettes on the Hemochron Microcoagulation Systems. | directCheck | 2 | August 10, 2010 | 2010 | International Technidyne Corp. | The package insert for the directCheck Microcoagulation Quality Control contains an incorrect acceptable performance range regarding ACT-LR cuvettes | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ITC sent Urgent Medical Device Correction letters on June 25, 2010 by Federal Express to hospital Risk Managers and Lab Directors, Clinical Laboratory Manager or Point-of-Care Coordinators. | Contact the recalling firm for information | 122 | 122 boxes in US, 1 out of US. | Nationwide and one customer in Belgium. | Terminated | September 23, 2010 | 44 | 122 | Not_Computer
Z-2196-2010 | 56157 | clamp, vascular, reprocessed | NMF | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | ASCENT HEALTHCARE SOLUTIONS. Reprocessed Femoral Compression Device, Model 11165. | FemoStop Reprocessed Femoral Compression Device | 2 | August 10, 2010 | 2010 | Ascent Healthcare Solutions, Inc. | The FemoStop Devices may fail to inflate or hold pressure. A separation between the dome and arch base prevents the device from maintaining pressure when inflated. | PRODUCTION CONTROLS: Process Control | Ascent sent an urgent medical device voluntary recall notification Letter and Response Form to each consignee on June 25, 2010. Consignees were requested to immediately examine their inventory of Ascent Reprocessed FemoStop devices with lot numbers located on an attached list; remove and quarantine all affected devices; complete the attached Recall Effectiveness Check Form; ship all affected product back to Ascent using the address indicated in the letter, and fax a completed Recall Effectiveness Check Form to 480-763-5350. |  | 175 | 175 devices | Devices were distributed to 8 consignees (end users/hospitals) located in VA, CA, FL, MI, WA. | Terminated | August 11, 2010 | 1 | 175 | Not_Computer
Z-2197-2010 | 56187 | cerclage, fixation | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | CERCLAGE WIRE, 0.6 MM DIA. X 175 MM LENGTH, SST, REF 71163006, QTY: (6), STERILE R, CE 008, 2020/07, Smith & Nephew, Inc., Memphis, TN 38116, EC REP: Smith & Nephew Orthopaedics GmbH, Tuttlingen, GermanyUsage: Bone fracture fixation | CERCLAGE WIRE | 2 | August 10, 2010 | 2010 | Smith & Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, compromising sterility. | PRODUCTION CONTROLS: Packaging | All affected Smith & Nephew Sales Representatives were notified of problem and the recall was sent via overnight mail on 01/23/2009. All affected Smith & Nephew International Distributors were notified via email and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. |  | 1 | 1 unit | KY, TX, DC, NC, CA, AZ, GA and OH; 3 International: Australia, Malaysia, and Switzerland | Terminated | January 23, 2012 | 531 | 43 | Not_Computer
Z-2205-2010 | 56201 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg.These circuits are intended for use with the MR Conditional LTV 1200 System with the ventilator utilizing an external exhalation valve, as single patient use (SPU) only patient circuits for adult and pediatric patients. The LTV 1200 ventilator is intended to provide continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. The ventilator is suitable for use in institutional, home and transport settings, and is applicable for adult and pediatric patients weighing at least 5 kg (11 lbs.), who require the following types of ventilatory support: -Positive Pressure Ventilation, delivered invasively (via ET tube) or non-invasively (via mask). -Assist/Control, SIMV, CPAP or NPPV modes of ventilation. -Breath types including Volume, Pressure Control and Pressure Support. | Pulmonetic Systems | 2 | August 10, 2010 | 2010 | CareFusion 203, Inc. | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed, immediately upon connection of the patient circuit to the ventilator and turning the ventilator on, the ventilator will begin to autocycle (with increased pressure delivered) with both | PRODUCTION CONTROLS: Process Control | Customers were sent a CareFusion "Urgent Medical Device Recall Notification" dated July 2, 1010. The letter was addressed to Respiratory Care Manger. The letter described the product , problem and advised consignees to remove the affected product from inventory and promptly return it to CareFusion for replacement.Please contact CareFusion Technical Support at 800-754-1914, option 2, upon receipt of this notification to arrange for prompt return and replacement of any the patient circuits in your inventory. | Contact the recalling firm for information | 200 | 200 boxes (5 each) | Nationwide distribution: AL, CA, FL, GA, HI, IL, IN, KS, MA,MI, MN, MT, NJ, NY, NC, OR, PA, RI, TN, TX, UT, VA, WA, and WI. | Terminated | December 17, 2011 | 494 | 200 | Hardware | Treatment/Delivery/Therapy | Return | Remove or Replace
Z-2210-2010 | 48992 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Motorized Height Ceiling Suspensions (MCS) and motorized height movement (actuator) for X-ray systems.Usage: Ceiling suspension unit with an actuator to hold monitors and X-ray shields. | Motorized Height Ceiling Suspensions (MCS) for CV Xray systems | 2 | August 11, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Monitor ceiling carriage may separate from the monitor ceiling suspension. | OTHER/UNDETERMINED: Pending | On 07/02/08, the firm sent a "Urgent - Device Correction" letter to all their consignees. The letter describes a product as X-Ray system with Monitor Ceiling Suspension and motorized height movement. The consignees are informed that the fixation bolts supporting the monitor ceiling carriage may come loose creating a situation which may allow the monitor ceiling carriage to separate from motorized height Monitor Ceiling Suspension and that would cause serious injury to patient and user.The firm has issued a mandatory Field Correction Action to correct the problem and the firm's technicians will inspect all ceiling suspension actuators for loose screws. If loose screws are found the actuator will be replaced.Consignees should call Philips Call Center at 800-722-9377 #5, #4, #1, and Reference FCO72200126. |  | 1111 | 1111 units | United States and Canada. | Terminated | December 02, 2010 | 113 | 1111 | Not_Computer
Z-2208-2010 | 55678 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems AFP QC, Part Number: 33219The Access AFP QC is intended for monitoring system performance of immunoenzymatic procedures for the quantitative measurement of alpha-fetoprotein (AFP) using the Access Immunoassay Systems. | Access AFP QC | 3 | August 11, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed an issue with the Access AFP QC kits (REF 33219) which contain three levels of control material identified as QC1, QC2, and QC3, Beckman Coulter has confirmed that vials of QC1 in the kit lots identified above contain microbial contamination. Customers may see that the results from a contaminated vial of QC1 are within the established refere | OTHER/UNDETERMINED: Pending | A Product Corrective Action (PCA) letter was senton the week of April 19, 2010 to all customers whopurchased the Access Immunoassay Systems AFPQC material, lot numbers referenced above. Theletter provided the customers with an explanation ofthe problem identified and instructed the customersto (1) Please discard all remaining inventory of the three Access AFP QC kit lots identified above,including all vials of QC1, QC2, and QC3.(2) Please review your ac and patient results obtained while using the affected control lots. Interpret Access AFP test results in light of the clinical presentation of the patient, including: symptoms,clinical history, data from other tests and other appropriate information.(3) Please share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. (4) Complete and return the enclosed response formwithin 10 days so we can be assured that you have received this important notification.For customers that need assistance or have any questions regarding this notification, they were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. | Contact the recalling firm for information | 166 | 166 | Nationwide and Canada | Terminated | May 07, 2012 | 635 | 166 | Not_Computer
Z-2206-2010 | 54370 | screw, fixation, intraosseous | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET MICROFIXATION "1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, QTY 1, 1.5 X 4MM TITANIUM. REF 91-1504 1 COUNT and REF 95-1504 - 5 PACK.BIOMET MICROFIXATION 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA, WWW.BIOMETMICROFIXATION.COM.Intended to aid in the alignment and stabilization of bone in the oral cranio-maxillofacial skeletal system. | 1.5x4mm HT XDrive Screw | 3 | August 11, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation has received a complaint from a customer who reports finding a 3.5mm length screw in one 5 pack of 4mm length screws. | OTHER/UNDETERMINED: Pending | Biomet Microfixation is recalling their 1.5x4mm HT X-drive Screws Lot Number 148530 after receiving a customer reported finding one 3.5MM screw in a package labeled as containing 4.0MM screws. Lot 148530 was packaged into one count packages 91-1504 and 5 count packages 95-1504. The 1 and 5 count packages were distributed between 12/10/2009 and 02/01/2010. Beginning on 03/17/2010 customers identified as receiving Lot 148530 HT X-drive Screws were notified of the recall by a combination of an Urgent Medical Device Recall Notice letter, e-mail, fax and telephone calls. Customers were instructed to inspect their inventory and to return any affected products. They were instructed to complete the attached Inventory Reconciliation sheet and to fax it to Biomet Microfixation. Customers were also instructed to contact any customers to whom they may have further distributed the product. Questions should be directed to Christy Cain at 1-800-874-7711 or 904-741-4400, ext. 468. |  | 1247 | 1247 screws | Worldwide Distribution -- USA, including states of FL, GA, IN, LA, MD, MI, MO, NC, NY, OK, TX, and VA and countries of Philippines, Japan, Brazil, and Canada. | Terminated | August 11, 2010 | 0 | 1247 | Not_Computer
Z-2207-2010 | 56031 | maggots, medical | NQK |  |  | General & Plastic Surgery | 510(k) | Monarch Labs Medical Maggots with the accessory LeFlap dressing (previously known as Creature Comforts II), 250-500 per vial, Batch #: MM-090406 with CCII-0901.Maggot therapy is intended to debride serious non-healing (chronic) wounds, and the dressing intended to confine the maggots to the wound. | Medical Maggots with the accessory LeFlap dressing | 2 | August 11, 2010 | 2010 | Monarch Labs, LLC | Monarch Labs initiated the field notification regarding their Medical Maggots from Batch MM090406 with the accessory LeFlap dressing due a customer complaint from a physician of a patient developing cellulitis after using maggots from the affected batch. | OTHER/UNDETERMINED: Pending | The customer recall communication was initiated on May 3, 2009 with Monarch Labs notifying all consignees with a via a faxed Recall Letter with fax-back attachment of the initial compliant they had received and the possible infection that may occur if instructions for use are not followed. All therapists, who used maggots from that batch were instructed to contact the firm immediately if they have experienced any adverse events, such as infection, cellulitis, or any other problem. If not, they were asked to simply return the fax-back attachment , noting "no problem" . |  | 14 | 14 vials | Nationwide Distribution -- CA, IA, KY, IL, TX, SC & UT. | Terminated | December 13, 2012 | 855 | 14 | Not_Computer
Z-2209-2010 | 56118 | pulse generator, permanent, implantable | NVZ |  |  | Cardiovascular |  | Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. | Medtronic Carelink Monitor Model 2490C | 2 | August 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. No patient injuries have been reported as a result of this issue. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic.Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346. |  | 37 | 37 Foreign, 0 Domestic | International Distribution Only: Australia and Canada. | Terminated | December 24, 2011 | 500 | 37 | Software | Device Operation | Replace monitor | Remove or Replace
Z-2212-2010 | 56131 | prosthesis, hip, semi-constrained (metal cemented acetabular component) | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | DYNASTYÂŽ A-CLASSÂŽ POLY LINER, REF DLXP-LD36, 1 EACH, Rx ONLY, STERILE EO, GROUP D, I.D. 36 mm, LIP 15o, LINER STD, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. Product is used in total hip arthroplasty. | DYNASTY ACLASS POLY LINER | 3 | August 12, 2010 | 2010 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life, translations, manufacturing date, CE marking and also listed the incorrect sterilization method. | PRODUCTION CONTROLS: Error in Labeling | On 06/28/2010, Wright Medical Technology sent a letter to consignees. Consignees were instructed to examine their inventory for recalled product and return it to the firm. Consignees should contact customer service at 800-238-7117 for return instructions and replacement inventory. Questions concerning the recall should be directed to Debby Daurer at 800-874-5630. |  | 6 | 6 units | USA distirbution only, in states of WI, WV, IN, IA, and GA | Terminated | May 31, 2011 | 292 | 6 | Not_Computer
Z-2214-2010 | 55364 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal acetabular revision system buttress augment size, 54, 00-4898-001-54, sterile, Zimmer Inc., Parsippany, NJ.Usage: Implantable devices that are used to address bony deficiencies of the pelvic antebulum in revision total hip arthroplasty surgery. | Trabecular Metal acetabular Augments | 2 | August 12, 2010 | 2010 | Zimmer Inc. | The firm has determined that some of the packages could develop a breach of the seal in the sterile barrier pouches. | PRODUCTION CONTROLS: Packaging | The firm sent an "Urgent Device Removal" letter dated 4/6/2010. The consignees were instructed to inventory the products on hand and return the recalled product to the manufacturer. More information is available at 1-800-613-6131. |  | 1972 | 1972 all products | US, Germany, Brazil, Singapore, Canada, Australia, England, India, and Switzerland. | Terminated | October 21, 2011 | 435 | 1972 | Not_Computer
Z-2211-2010 | 55502 | system,network and communication,physiological monitors | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DATEX-OHMEDA NETWORK AND ICENTRAL '05, SALES REVISION 4.3 ( Datek-Ohmeda Network and iCentral (including iCentral Client)) The Datex-Ohmeda Network and iCentral (including iCentral Client) transfers information between networked Datex-Ohmeda devices in the Monitor Network as well as between Monitor Network and networked devices in GE Unity NetworkTM. It also allows information transfer between several iCentrals. Within one Monitor Network it allows a networked device to display, store, print and otherwise process information received from other networked devices. The iCentral maintains the network connections between the Datex-Ohmeda bedside monitors and other networked devices in Monitor Network. Network connections consist of hardwired network cables and/or Wireless LAN (WLAN) connections. Furthermore, it coordinates the transfer of information between devices in the Datex-Ohmeda Network, between the Datex-Ohmeda Network and devices in GE Unity Network as well as between the Datex-Ohmeda Network and Hospital Information Systems (HIS). The iCentral can be used for remote monitor management, storing, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The iCentral Client can be used for remote monitor management, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The Datex-Ohmeda Network will be used for patients in the hospital and it is meant for continuous use. The device is for use by qualified personnel only. | DATEXOHMEDA NETWORK AND ICENTRAL | 2 | August 12, 2010 | 2010 | Ge Healthcare, Llc | The iCentral or iCentral client v5.0.2 with HP xw 4400 and 4600 may spontaneously shutdown, causing monitoring loss at the central monitoring station. Monitoring loss may result in prolonged delay of treatment of potentially lethal arrhythmia. | DESIGN: Software Design | GE Healthcare sent a "Urgent Medical Device Correction letter" dated 6/18/10 on 6/25/10 to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The letter instructs consignees to follow specific detailed instructions for safe use of the device pending provision of a software upgrade by the firm. Consignees will be contacted by a GE Healthcare representative to schedule the software upgrade. | Contact the recalling firm for information | 2962 | 2962 | AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, NE, NH, NH, NY, NC, ND, OH, OK PA, RI, SC, TN, TX VT, VA, WA, WI.ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA , KUWAIT, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND,NICARAGUA, NIGERIA, NORWAY, PHILIPPINES, POLAND, PORTUGAL QATAR, RUSSIAN FEDERATION SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN THAILAND TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN. | Terminated | August 14, 2012 | 733 | 2962 | Software | Device Operation | Software update | Software Update
Z-2219-2010 | 56166 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Hi-Art System, H-0000-0003Usage: The TomoTherapy Hi-Art System is intended to be used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. The megavoltage x-ray radiation is delivered in a rotational, non-rotational, modulated (IMRT), or non modulated (non-IMRT/three dimensional conformal) format in accordance with the physician approved plan | TomoTherapy HIART System | 2 | August 12, 2010 | 2010 | TomoTherapy Incorporated | An issue was identified with the TomoTherapy HI-Art System. In the event a patient or DQA plan has a moved image, roll adjustments applied during registration will be incorrect. The Planning Station Plan Settings and DQA Setup tabs allow for images to be moved during planning. During registration when roll is applied on moved images, the Operator Station incorrectly rolls the image about | DESIGN: Software Design | Consignee was sent a TomoTherapy "Urgent Medical device Correction Field Safety Notice' dated December 01, 2009. The letter described the Issue, Product Affected, Recommended Action and Resolution. |  | 1 | 1 | VA | Terminated | December 17, 2011 | 492 | 1 | Software
Z-2213-2010 | 56189 | accessories, cleaning, for endoscope | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DSD Application Guide (MG04-0001 revision M) for the Pentax Universal Video Scope Hookup (DSD-110-HU0109) only when used with the DSD automated endoscope reprocessor (AER) and the Pentax EG-3630U endoscope. | DSD Application Guide (MG040001 revision M) for Pentax Universal Video Scope Hookup:DSD110HU0109 | 2 | August 12, 2010 | 2010 | Minntech Corp | The DSD hook up Application Guide incorrectly recommends use of the Medivators DSD-110-HU0109 endoscope hookup to connect the Pentax EG-3630U Ultrasound Endoscope to a Medivators DSD endoscope reprocessor or Medivators Scope Buddy Endoscope Flushing Aid. This may result in inadequate disinfection during reprocessing, and subsequent infection risk to the patient. The recall was due to the misprinte | MISBRANDING: Labeling False and Misleading | Consignees were sent a Medivators Reprocessing Systems "Urgent Medical Device Recall" letter date June 17, 2010. The letter was addressed to "Dear Medivators Customers". The letter described the problem and product involved. Recommended consignees to immediately examine the Pentax EG-3630U Ultrasound Endoscope hookups and inform all staff of the notice to ensure that the DSD-11--HU01109 is not used with the Pentax endoscope.Also requested consignees to complete and return the enclosed response form. | Contact the recalling firm for information | 185 | 185 | AZ, AR, CA, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NY, NC, ND, OH, PA, SC, TN, TX, UT, VA, WA, WV, WI, WY. | Terminated | December 17, 2011 | 492 | 185 | Not_Computer
Z-2178-2010 | 56222 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Engage TR Introducer, 6 F - ACT (2.25 mm), 7 cm length, .025" Max Guidewire O.D., Rx, Sterile EO, REF C408524, 100028093. | Engage TR Introducer, 6 F ACT (2.25 mm), 7 cm length, .025" | 1 | August 12, 2010 | 2010 | St. Jude Medical Cardiovascular Division | Potential for a partial or complete separation of the Shaft (sheath) from the hub or a material break in the hub assembly just distal to the hemostasis seal. If either of these were to occur during use, it is likely that fluids would leak around the Introducer hub and strain relief. | DESIGN: Process Design | Consignees were sent "Urgent Medical Device Recall Notice" dated June 24, 2010 from St. Jude Medical . The letter was addressed to Cath Lab Manger/ Risk Manager. The letter described the problem and the product involved. Advised consignees to discontinue use of the product and remove from inventory. Sales Representatives were also sent a St. Jude Medical Memorandum "Field Action Notice" dated June 24, 2010 and a reconciliation form. The memorandum described the problem and product. Advised Sales Rep to visit the affected accounts and to performed the assigned tasks.A St. Jude Medical sales representative will be contacting for removal and replacement of the affected product. For additional information, please contact St. Jude Medical at 651-491-4470. | Contact the recalling firm for information | 5 | 5 OUS | Worldwide Distribution including USA, Italy, Denmark, France, Switzerland, England, Finland, Austria, Spain and Estonia. | Terminated | December 27, 2011 | 502 | 5120 | Not_Computer
Z-1634-2010 | 55077 | humidifier, respiratory gas, (direct patient interface) | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCIÂŽ AQUA+ÂŽ Flex, Hygroscopic Condenser Humidifier, HCH Tidal Volume Range 300-1500mL, ICU/Anesthesia Adult/Child, Catalog number 1570, Distributed by Teleflex Medical. The Aqua+ Flex is a passive humidifier indicated for use on patients under anesthesia and intensive care to effectively warm and humidify the inspired air. | Hudson RCI AQUA Flex, Hygroscopic Condenser Humidifier | 1 | August 13, 2010 | 2010 | Teleflex Medical | The patient end of the 22 cm connector on the flex tube may not fit securely within the endotracheal tube (ET) connector and therefore disconnect from the endotracheal tube. A disconnect in ventilator dependent patients without prompt response to the alarm could lead to hypoxia, organ failure, or cardio-respiratory arrest. | DESIGN: Device Design | Teleflex Medical sent an Urgent Medical Device Recall letter on March 22, 2010 to consignees. They were instructed to: 1) Immediately discontinue use and quarantine any products labeled as AQUA+ FLEX, Catalog number 1570 with the lot numbers provided. 2) Complete and fax the Recall Acknowledgment Form to 866-804-9881. 3) After contact by a Customer Service Representative, who will provide them with a Return Goods Authorization Number, return the product to Teleflex Medical for credit. Letters addressed to distributors instructed them to communicate the recall to any of their customers who have received the product using the customer letter template and acknowledgement form provided by Teleflex Medical. Their customers were to return the product and acknowledgment form to the distributor who would consolidate and return the product to Teleflex Medical. Affected products received by Teleflex Medial will be destroyed upon receipt. For additional instructions or information, contact the firm at 1-866-246-6990. | Contact the recalling firm for information | 3841 | 3841 units | Worldwide distribution in the states of CA, FL, GA, KY, KS, IL, MA, MD, MI, MN, NC, NJ, NM, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI and to the country of Japan. | Terminated | November 28, 2011 | 472 | 3841 | Not_Computer
Z-2220-2010 | 56246 | device, irrigation, ocular surgery | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | CORNEA COAT Hydroxypropylmethylcellulose 1.5ml Syringe. STERILE, Rx ONLY.Manufactured by: Insight Instruments, Inc., 2580 SE Willoughby Blvd. Stuart, FL 34994 USA.REF 60010.000, QTY 14 - This package contains 14 individual packs, each containing 1 sterile, single use ocular irrigation device. For external use only!Used as an ocular irrigation device. | CORNEA COAT | 2 | August 13, 2010 | 2010 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling CORNEA COAT 2.5% Hydroxypropylmethylcellulose solution in 1.5 ml Syringe due to a potential weakness of the peel pouch seals. Products affected are Item Numbers 600100.000 sold individually and 40201.000 thru 10206.000 sold as components of BIOM Pack. | PRODUCTION CONTROLS: Packaging | All consignees were notified of the recall by letter sent via first class mail. Letter contained a Return Response Form. |  | 16128 | 16,128 | Worldwide Distribution -- USA, including Puerto Rico, and Frankfurt, Germany. | Terminated | February 16, 2012 | 552 | 66853 | Not_Computer
Z-2222-2010 | 56256 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | SIEMENS AXIOM Luminos TF and Sireskop SD Solid State x-ray imager (flat panel/digital imager), Fluoroscopic x-ray diagnostic system. | SEIMENS Axiom Luminos TF and Sireskop SD | 2 | August 13, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Fluid may enter the multi-functional handle (Opti Grip) when using the Axiom Luminos TF or Sireskop SD system, causing unintended reaction or movement, increased radiation or useless images. | DESIGN: Device Design | Siemens Healthcare Sector sent an "Urgent Field Safety Notice" to consignees describing the issue and informing them to take specific precautions when performing interventional procedures to avoid fluids penetrating the system. The multi-functional handle (Opti Grip) will be replaced through implementation of Update Instructions AX030/10/S. The modified version of the handle resists fluids from entering the handle and prevents malfunction. |  | 1296 | 1296 | Nationwide distribution: all 50 US states and PR. | Terminated | September 29, 2012 | 778 | 1296 | Not_Computer
Z-2176-2010 | 56290 | apparatus, nitric oxide delivery | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Ikaria INOMAX DS Drug Delivery System, model 10003. Ikaria 6 Route 173, Clinton, NJ 08809 USA.The INOMAX DS Delivery System delivers INOMAX (nitric oxide for inhalation) therapy gas into the inspiratory limb of the patient breathing circuit in a way that provides a constant concentration of nitric oxide (NO), as set by the user, to the patient throughout the inspired breath. It uses a specially designed injector module, which enables tracking of the ventilator waveforms and the delivery of a synchronized and proportional dose of NO. It may be used with most ventilators. The INOMAX DS provides continuous integrated monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. | Ikaria INOMAX DS Drug Delivery System | 1 | August 13, 2010 | 2010 | Ikaria Holdings | The pressure switch in the INOMAX DS drug-delivery system may fail, which may interrupt or delay the administration of INOMAX (nitric oxide) for inhalation to patients. Delay in administration of INOMAX therapy may cause: -worsening of systemic oxygenation indices (i.e., hypoxemia, especially manifested as decreased arterial oxygenation saturation). -hypotension, and/or -increase in pulm | DESIGN: Component Design/Selection | Ikaria sent an Urgent Medical Device Recall letter to customers dated July 21, 2010. This letter is also posted on the Inomax webpage. The letter described the issue, how to recognize it, what to do in the event of a suspected leak, and recommended maintaining back-up cylinders and depressurization of the regulator when not in use. The replacement plan for all INOMAX DS drug-delivery systems was described. Customers can contact Ikaria Customer Care at 1-877-566-9466. | Contact the recalling firm for information | 2507 | 2507 | Worldwide Distribution to include: US Nationwide and Countries of Canada and Malaysia | Terminated | December 27, 2011 | 501 | 2507 | I/O | Treatment/Delivery/Therapy | Replace device | Remove or Replace
Z-2223-2010 | 55804 | cable, transducer and electrode, patient, (including connector) | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | Disposable Surgical Cable, Product Code: FL-601-97 (100), Single use, Rx Only, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005Usage: Temporary pacing of open hearts: used with External Pacing Generators (EPGs). | Safe Connect | 2 | August 13, 2010 | 2010 | Remington Medical Inc. | The Safe Connect component may separate from the cable, causing the cable to become unplugged from the EPG. | DESIGN: Process Design | Consignees were notified on/about 05/19/2010. They were instructed to examine their inventory and quarantine product subject to the recall. If they distributed the product, they were advised to identify their customers and notify them of the recall (via the recall letter). A Return Authorization was provided. More information can be obtained by calling 770.888.8520. | Contact the recalling firm for information | 1488 | 1488 units | AZ, CA, CO, CT, DE, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NM, NY, OH, OK, OR, PA, SC, TX, TN, TX, VA, WA, WI and Canada. | Terminated | November 29, 2010 | 108 | 1488 | Not_Computer
Z-2224-2010 | 54546 | system, dialysate delivery, single patient | FKP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008H Dialysate Delivery Machine, Manufactured by Fresenius USA Inc., Walnut Creek, CA | Fresenius FDS08 Dialysate Delivery Machine | 2 | August 15, 2010 | 2010 | Fresenius Medical Care North America | Power cords may fail if the prongs crack or fail at the plug, and may become a fire/burn/shock hazard or delay/interrupt therapy. | DESIGN: Component Design/Selection | An urgent device recall notification was sent to all consignees. Customers are asked to inspect cords for defects, and are refered to preventive maintenance requirements in the manual. Customers who identify the defective cords will be instructed to discontinue use as soon as possible without jeopardizing patient care. Customers are instructed to monitor cords regularly, and pay attention to machines that are frequently plugged or unplugged. Personnel are instructed to not use the cord to unplug the device from the wall (yanking the plug from the wall) and to refrain from rolling over the cord or plug. | Contact the recalling firm for information | 33 | 33 | Product distributed to 5404 customers throughout the US and the world. | Terminated | February 07, 2011 | 176 | 72142 | Battery
Z-2228-2010 | 56193 | instrument, manual, surgical, general use | MDM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | SpeedPass Disposable Suture Lariat with Nitinol Wire 45 Degree Up, QTY. 1, Biomet Sports Medicine, Warsaw, IN 46581. Sterile. Part number 904008.Used to pass suture through tissue. | Speedpass Suture Lariat | 2 | August 16, 2010 | 2010 | Biomet, Inc. | The nitinol wire may be fractured or may have the potential to fracture making the suture lariat unusable and possibly delaying surgery. | DESIGN: Component Design/Selection | The firm notified all 29 consignees by An Urgent Medical Device Recall Notice letter on 6/18/2010. The letter identified the affected product and explained the reason for recall. Customers are asked to locate and remove the affected products, follow the instructions on the enclosed "FAX Back Response Form," and fax a copy of the form back prior to return of the affected product. Additionally, if product has been further distributed, then hospital personnel must be notified and given the "Dear Biomet Customer" letter. Questions should be directed to Mary Johnson at 800-348-9500 or 574-372-3983. |  | 301 | 301 units total | Worldwide Distribution -- USA, including the states of FL, IL, NY, CO, MD, CA, GA, AL, PA, OH, SD, KY, TX, ID, WA, NC, AK, & MN, and Columbia and The Netherlands. | Terminated | August 06, 2012 | 721 | 301 | Not_Computer
Z-2233-2010 | 56262 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis MP/Artis dMP and Artis zee Multipurpose. Angiographic x-ray system. | AXIOM Artis MP/Artis dMP and Artis zee Multipurpose | 2 | August 16, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The Axiom Artis MP, Artis dMP or Artis zee Multipurpose C-arm gearbox can become damaged. If the gearbox is damaged, the C-arm may tilt away when being repositioned and collide with nearby objects in its range of motion. | OTHER/UNDETERMINED: Pending | Siemens has issued an "Urgent Field Safety Notice: Customer Safety Information to consignees with Update Instructions AX033/10/S. This letter informed customers of the issue, and provides an Addendum to the User Manual. The Update Instructions describe the process for installation of an additional clamping mechanism provided by Siemens into the existing system, which will be activated if the gearbox is damaged. If activated, the clamping machanism will prevent the C-arm from tilting away, stop regular operation, and alert the operator to check the system and repair the gearbox damage. | No consumer action necessary | 131 | 131 units | Nationwide distribution: States of AL, AZ, CA, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OK, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 05, 2012 | 689 | 131 | Not_Computer
Z-2235-2010 | 56225 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | medCOMP(R) MRCTI66841 6.6F Plastic Dignity Low Profile CT Port with Pre-Attached Chronoflex(R) Polyurethane Catheter - 1 Unit. Catalog #MRCTI66841. | medCOMP(R) MRCTI66841 6.6F Plastic Dignity Low Profile CT Port with PreAttached Chronflex(R) Polyur | 2 | August 17, 2010 | 2010 | Medical Components, Inc dba MedComp | Recalled product was packaged with the incorrect port and does not match the label. The kit is labeled as a 6.6F "Dignity" Low Profile CT Port. The kit contains a 6.6F "Pro-Fuse" Low Profile CT Port. | TRAINING: Employee Error | Medcomp sales reps and foreign distributor were notified via email on 7/1/10. Product is to be returned to Medcomp. | Contact the recalling firm for information | 47 | 47 units | Worldwide Distribution -- USA, including states of NC and TX and country of Israel. | Terminated | October 13, 2010 | 57 | 47 | Not_Computer
Z-2234-2010 | 56138 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Accessory Power Pack; a 4 outlet power strip with a custom pole clamp bracket to attach to the IV pole with the infusion pump; Distributed by: Hospira, Inc., 275 N. Field Drive, Lake Forest, IL 60045; part number 39003-04-01, Modo #HSP1700, Model Tripp-Lite PS-415-HGAccessory power pack used for use with medical devices. | Hospira Accessory Power Pack | 2 | August 17, 2010 | 2010 | Hospira Inc. | The Power Pack Accessory is not certified to UL 60601 standard or Subject 1363A for relocatable power taps intended for use with medical equipment. Potential problems could include electrical shock, electrical short or electrical fire. | DESIGN: Device Design | Hospira sent Urgent Device Recall letters dated 7/1/10 to their direct accounts on 7/2/10 via Federal Express 2 day delivery, informing them of the potential for electrical shock, electrical short, or electrical fire, since the power packs were not certified to UL 60601 standard or Subject 1363A of UL standard 1363 for relocatable power taps intended for use with medical equipment. The accounts were instructed to examine their inventory for the affected product and quarantine it immediately, and Stericycle at 1-800-321-7388 to obtain a return kit for returning the affected product. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. They were also requested to complete the enclosed reply form and return it to Hospira by fax at 1-262-577-6921, indicating the amount of product placed on hold for return, if they further distributed the product and if they notified their retail customers. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. |  | 1766 | 1,766 units | Worldwide Distribution -- US and Canada. | Terminated | July 20, 2011 | 337 | 1766 | Battery | Physical Safety Hazards | Return | Remove or Replace
Z-2236-2010 | 55663 | snare, flexible | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SD Snares, Reusable Electrosurgical Snares. Models: SD-9L/U-1, SD-11L/U-1, SD-12L/U-1, SD-13L/U-1, SD-13L/U-1, & SD-16L/U-1. Snare tube model numbers: MAJ-13 and MAJ-15. To be used with Olympus endoscopes to electrosurgically resect tissue within the digestive tract. | Olympus SD Snares | 2 | August 18, 2010 | 2010 | Olympus America Inc. | An investigation of consumer complaints revealed the inability of the snare wire to be retracted into the snare tube after deployment of the snare. | OTHER/UNDETERMINED: Pending | Firm initiated this recall following investigation of customer complaints regarding the inability of the snare wire to be retracted into the snare tube. Customers were notified via certified mail, an Urgent: Medical: Device Recall letter issued on May 4, 2010. The letter identified the affected product and the reason for the recall. Customers were instructed to inspect their inventory and identify the affected product. Customers are to remove from use any identified inventory, quarantine, and discard them. Customers should contact the firm to obtain replacement products. Questions should be directed towards Laura Storms-Tyler at 484-896-5688 or laura.storms-tyler@olympus.com. | Contact the recalling firm for information | 108 | 108 units | Nationwide Distribution -- AR, CA, CO, CT, DC, FL, GA, IN, KS, LA, MD, MO, NC, NY, OH, PA, SC, TN, and TX. | Terminated | May 20, 2013 | 1006 | 108 | Not_Computer
Z-2218-2010 | 54687 | catheter, percutaneous transluminal coronary angioplasty (ptca), cutting/scoring | NWX |  |  | Cardiovascular |  | AngioSculpt Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter, cutting/scoring, Easy Exchange (EX) delivery system, Manufactured by AngioScore, Fremont, CA. Catheter is for use in the treatment of hemodynamically significant coronary artery stenosis, including in-stent restenosis and complex type C lesions, for improving myocardial perfusion. The devices are available in balloon diameters of 2.0-3.5 mm in 0.5 mm increments, and in scoring balloon lengths of 10 to 15 nm. Catheter length is 139 cm and is compatible with 0.014 inch guide wires and 6F guide catheters. The EX catheter is supplied sterile and intended for single use. | AngioSculpt PTCA Scoring Balloon Catheter, manufactured by Angioscore | 1 | August 18, 2010 | 2010 | AngioScore Inc. | Separations of the PTCA catheter shaft at the guide wire exit port have occurred during treatment of coronary artery stenosis. Fragments of the catheter may become lodged in coronary arteries, which resulting in serious patient outcome and require emergency coronary bypass surgery and medical intervention to remove the retained device fragments. |  | AngioScore sent an "Urgent: Medical Device Voluntary Recall" dated December 4, 2009 addressed to "Catheterization Lab Manager" at consignees. The Recall notification letter instructed consignees to inspect their inventory to detect product from the involved lot, and segregate it from their inventory. Consignees were asked to record all affected product and fax return forms to the firm. AngioScore Customer Service will contact consignees and provide instructions for return and replacement of affected product.. Consignees can contact AngioScore at 877-264-4692. | Contact the recalling firm for information | 3870 | 3870 units | Nationwide distribution. | Terminated | June 21, 2011 | 307 | 3870 | Not_Computer
Z-2237-2010 | 49012 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series Computed Tomography (CT) Scanner System, X-Ray System. Part Number: 989605651321. Intended to produce cross-section images of head and body bycomputer reconstruction of X-ray transmission data taken at differentangles. | NeuViz Dual series Computed Tomography (CT) Scanner System | 2 | August 19, 2010 | 2010 | Neuisys, LLC | The reference lines for image generated for a Surview scan may appear in the incorrect position in the "FILM" display mode. | DESIGN: Software Design | On 07/21/2008, NEUISYS Imaging Systems Solutions distributor in Greensboro, NC began sending the URGENT DEVICE CORRECTION letter on behalf of Philips and NeuSoft Medical Systems (PNMS), Co., Ltd in China.The letter informs the consignees that there is potential safety problem in the NeuViz Dual series CT scanners due to the malfunction of "Incorrect Reference Line in the Film". When a "Head first, Couch out" or a "Feet first, Couch in" Surview scan is started then paused midstream, the image generated is what the doctor expects, but the reference lines appear in the incorrect position in "FILM" display mode. The acquired images accurately represent the exact positions planned on the original view. The incorrect/unmatched reference line in film may result in incorrect diagnosis.The consignees are advised not to stop or pause during Surview acquisition while PNMS in China will provide the updated software patch to all affected systems free of charge.The Field Change Order #FCOP-08-CT-NMS-008 SP4 will be implemented by NEUISYS Imaging Systems Solutions in Greensboro, NC., who is a distributor for PNMS. Consignees are advised to e-mail Service Support Department of PNMS at helpdesk@pnms.neusoft.com regarding the recall. |  | 19 | 19 units in the US | Nationwide Distribution | Terminated | August 23, 2010 | 4 | 19 | Software | Display/Image | Software update | Software Update
Z-2238-2010 | 55921 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2000, manufactured by GE Medical Systems, SCS | Innova 2000 Digital Flouroscopic Imaging System | 2 | August 19, 2010 | 2010 | GE Healthcare, LLC | A hardware component (potentiometer and encoder)mismatch in Innova 2000 digital fluoroscopic imaging systems may cause a Source to Image Distance (SID) failure, which will cause the imager to fail to detect the distance between the imager and the patient. This will result in loss of the x-ray image if the recovery process is not performed. The error message to use the Emergency Back-Out Button for | OTHER/UNDETERMINED: Pending | GE Healthcare sent a GE "Urgent Medical Device Correction" dated June 9, 2010letter to consignees. The letter was addressed to Hospital Administrators / Risk Mangers; Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details; Safety Instruction; Product Correction and Contact Information. The letter detailed the method for identifying SID failure and the use of the emergency back-out buttoms to recover the X-ray image. A GE service representative will affix a label to the device on-site to remind operators how to recover from loss of X-ray from SID failure. Consignees may contact GE at 800 437-1171, option 4. |  | 885 | 885 | Worldwide distribution: States of AL, AZ, AR, CA, CO, CT,DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV,NH, NJ, NM, NH,NC, ND, OH, OK OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and countries of VENEZUELA, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, JORDAN, JAPAN, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT , ECUADOR, COLOMBIA, CHINA, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRALIA, and ALGERIA. | Open |  |  | 885 | Hardware | Device Operation |  | Other
Z-2240-2010 | 56258 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | PoleStar FlexiCoil band D/E, revisions A and B, Catalog # 9733063,Usage: Low field MRI device used to gather cranial images. | PoleStar N20 Magnetic Resonance Diagnostic Device | 2 | August 19, 2010 | 2010 | Odin Medical Technologies Ltd. | Electronic box on headstrap may overheat, causing burn to patient. | DESIGN: Device Design | Customers were notified by letter on 07/07/2010, advising them of the overheating problem and alerting them that a service rep will be visiting to inspect and upgrade any affected FlexiCoils. Questions were referred to 800-595-9709. |  | 560 | 560 units | CA, CO, DC, IL, IN, MD, MI, MN, MO, NE, NJ, NV, NY, OH, OR, WA, and WI. Foreign distribution to Austria, Belgium, Canada, China, France, Germany, Israel, Italy, Korea, Netherlands, Poland, Spain, Russia, Switzerland and Turkey. | Terminated | November 09, 2010 | 82 | 560 | Not_Computer
Z-2239-2010 | 55043 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The Starion Instruments Thermal Ligating Shears 35C Model 102-138D and 23C Model numbers 102-137D allow the surgeon to position the distal jaws of the instrument around the region of tissue to be cut/cauterized. The surgeon squeezes the handle and depresses the switch, heating elements activates in the jaws. The heat is conducted to the tissue between the jaws to provide cutting/cauterization. For the simultaneous cutting and cauterization of soft tissue during surgery. Cutting of natural or synthetic, non-metallic, sutures during surgery. | Starion Instruments Thermal Ligating Shears | 2 | August 19, 2010 | 2010 | Starion Instruments | The Starion Instruments Thermal Ligating Shears (STLS) 35C and the Starion Thermal Ligating Shears (STLS) 23C may turn on and remain on without the finger button being depressed. The devices may also work intermittently, incorrectly indicating the device is 'on". | DESIGN: Device Design | Starion distributed a "STLS Device Recall - Important Medical Device Information" letter dated March 9, 2010 to consignees. The letter identified the affected model and lot numbers. Consignees were asked to complete a inventory status form and fax to Starion. the letter included a shipping labe with a pre-assigned return authorization number. Starion will ship replacement after receipt of the inventory status form and returned product.. |  | 228 | 228 units | Worldwide Distribution: USA states of CA, OH, MI, KS and MO and countries of Italy, China, Kuwait, and Germany. | Terminated | November 27, 2010 | 100 | 228 | Not_Computer
Z-2241-2010 | 55153 | chair, dental, without operative unit | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | The magnetic headrest is designed to fit the following dental chairs (THE HEADREST IS THE RECALLED ITEM, NOT THE CHAIRS):The dental chairs are sold under the brand names Marus and DCI Equipment. The affected models are:Marus MaxStar-Agency Model # DC1490;Marus MaxStar-Agency Model # DC1690;Marus NuStar-Agency Model # DC1700;Marus NuStar-Agency Model # DC1702;Marus ProStar-Agency Model # DC1540;Marus ProStar-Agency Model # DC1535;DCI Equipment DC1235 Hydraulic-Agency Model # DC1235; and DCI Equipment DC1335 Electromechanical-Agency Model # DC1335. | Dental chair headrest | 2 | August 20, 2010 | 2010 | Dental Equipment LLC | The magnetic headrest on certain dental chairs may affect the function/programming of some implantable pacemakers or defibrillators if the implanted device is programmed to respond to a magnet. | OTHER/UNDETERMINED: Pending | DCI Equipment sent an "Urgent Field Safety Alert" letter dated March 15, 2010 to consignees advising them of the potential problem. The letter instructed that patients with implantable pacemakers or defibrillators should avoid dental chairs with the magnetic headrest. A warning label was sent with the letter for insertion into the product Use and Care manual for operator warning. Consignees were asked to complete a Field Safety Alert Acknowledgement/Return Form and fax to the firm. Consignees were requested to notify customers.Consignees may contact the firm at 503 537-3602. | Contact the recalling firm for information | 108 | 108 | Worldwide distribution in the USA and Canada. | Terminated | January 04, 2012 | 502 | 108 | Other | Device Operation | Instructions to avoid | Safety Notice/Insructions
Z-2243-2010 | 56291 | for export only - antigen, anti-cardiac, indirect fluorescent antibody test system | NUK |  |  | Clinical Chemistry | Contact ODE | Human Cardiac Specific Fatty Acid-Binding Protein (H-FABP) ELISA TEST KIT Product is an In Vitro Device (IVD) categorized as a Cardiac Marker, manufactured and distributed by BioCheck, Inc., Foster City, CA. Enzyme Immunoassay for the quantitative determination of Human Cardiac Fatty Acid-Binding Protein (H-FABP) concentration in serum, citrate-plasma, EDTA-plasma, and heparin-plasma. | Human Cardiac Specific Fatty AcidBinding Protein (HFABP) ELISA TEST KIT | 2 | August 20, 2010 | 2010 | BioCheck Inc | Products were marketed without 510(k) for diagnostic use. Unapproved device. | PREMARKET APPROVAL: No Marketing Application | BioCheck sent an "Urgent Device Recall" letter dated July 15, 2010 to consignees identifying the products as unapproved and requesting prompt return to the firm. Consignees are asked to sub-recall. The letter requested a reponse form be completed and returned to the firm.Diagnostic claims have been removed from the website. Returned product will be quarantined by the firm pending determination of marketing status and labeling revision. | Contact the recalling firm for information | 187 | 187 kits | Worldwide Distribution - USA, including the states of CA, MI, NJ, and PA and countries of Canada, India, S. Korea, and UAE. | Terminated | December 13, 2010 | 115 | 187 | Not_Computer
Z-2242-2010 | 56285 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Surgi-Tip Transducer Cover, Ref. 610-834, Sterile (20.3 x 244cm) polyethylene telescopically-folded cover with attached (6cm) NeoGuard tip. The latex-free surgi tipped covers are utilized to cover ultrasound transducers and prevent cross-contamination from occurring with equipment and patients. | CIVCO SurgiTip Transducer cover | 2 | August 20, 2010 | 2010 | Civco Medical Instruments Inc | Polyethylene material did not have proper treatment applied during manufacturing. Untreated polyethylene material does not maintain a secure seal with the adhesive tape used to secure the tip to the cover. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CIVCO sent a "Urgent - Medical Device Correctdion" letter date July 19, 2010 to consignees describing the reason for recall and requesting they dispose of the affect lot number of product. A response form was enclosed for completion and return to the firm, indicating the amount of product in inventory and requesting a Return Authorization Number. The form was to be returned via FAX or e-mail. Customers may contact CIVO at 800-445-6741. | Contact the recalling firm for information | 22 | 22 boxes | Nationwide distribution: the states of OH, MD, NE, VA, IA, MI, WA, and WI. | Terminated | April 13, 2012 | 602 | 22 | Not_Computer
Z-2244-2010 | 56200 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee Baseplates, Size 8, Left, Part # 333-01-108, Lot# A1000003. | 3DKnee System | 2 | August 22, 2010 | 2010 | Encore Medical, Lp | One lot of product has an incorrect color label. The color label is a secondary visual indicator of sizing. | PRODUCTION CONTROLS: Labeling Mix-Ups | Firm notified their consignees by letter on June 28, 2010. | Contact the recalling firm for information Contact the recalling firm for information | 18 | 18 | Florida, Missouri, and Italy | Terminated | October 04, 2010 | 43 | 18 | Not_Computer
Z-2261-2010 | 54701 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | REF 70-0077-103, 3mm ARTHRO-KNIFE SHEATHED KNIFE, SERRATED, STRAIGHT, 2-SIDED, STERILE, Rx ONLY. CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA. Intended for use in orthaepedic surgery where a cutting instrument is required. | ARTHROKNIFE | 2 | August 24, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | This recall was due to the possibility that multiple products may have a breach in the poly-Tyvek pouch that could potentially compromise the sterility of the contents. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Each consignee will be notified of the recall by an Urgent Medical Device Recall Notification letter sent by FedEx or an equivalent method. The letter identified the affected product, explained the reason for recall, and health risk. Consignees were instructed to examine their inventory and segregate the affected product for return. Customers are to complete and return the Reply Form and to notify their customers if the product was further distributed. Consignees are instructed to return any remaining product to ConMed Linvatec. Questions should be directed to the Customer Service department at 800-535-8536. | Contact the recalling firm for information | 35451 | 35,451 total | Worldwide Distribution -- United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Republic of Korea, Netherlands, Portugal, Serbia, Singapore, Slovenia, El Salvador, Thailand, Taiwan, South Africa, and throughout the U.S. | Terminated | December 15, 2010 | 113 | 35451 | Not_Computer
Z-2267-2010 | 55156 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Care Assist Bed, Model P1170. | HillRom Care Assist Bed | 2 | August 24, 2010 | 2010 | Hill-Rom, Inc. | Pinch hazard. Potential for shearing action between the side rail arm and the bracket used to secure the side rail to the bed. | DESIGN: Device Design | Hill-Rom sent a customer notification letter dated 4/9/2009. A second notice was sent to all customers on 7/9/2009. The letter included instructions and a Customer Response Form, along with the required parts. Customers were instructed to do the correction on their affected beds and return the completed customer response form within 45 days. | No consumer action necessary | 39971 | 39,971 both products | Worldwide distribution: USA, Afghanistan, Algeria, American Samoa, Argentina, Australia, Bermuda, Brazil, Canada, Chili, China, Colombia, Costa Rica, Cuba, Ecuador, Egypt, Hong Kong, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Kyrgyzstan, Lebanon, Libya, Malaysia, Mexico, Oman, Panama, Paraguay, Philippines, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Syria, Taiwan, Thailand, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | June 06, 2011 | 286 | 39971 | Not_Computer
Z-2263-2010 | 56273 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Philips M1013A IntelliVue G1 Anesthesia Gas Modules with manifolds | Philips M1013A IntelliVue G1 Anesthesia Gas Modules with manifolds | 2 | August 24, 2010 | 2010 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels. | DESIGN: Device Design | Philips Healthcare issued an "Urgent Device Correction" letter dated July 14, 2010 with an addendum to the Instructions for Use including additional instructions for the replacement of the water trap for the Philips M1013A IntelliVue G1, M1019A IntelliVue G5, M1026B/BT Anesthesia Gas Modules. This addendum is being sent along with the Urgent Medical Device Correction notice to all affected customers.Customers will be requested to incorporate the information contained in the Urgent Medical Device Correction notice and the Addendum to the Instructions for Use with the existing labeling (Instructions for Use) provided with the product. Customers can contact Philips at 1-800-722-9377 for further information.. |  | 531 | 531 units | Worldwide Distribution: United States, and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Chile ,China ,Colombia, Czech Republic, Estonia, Finland, France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Latvia, Libyan Arab Jamahiriya, Luxembourg, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Norway, Oman, Pakistan, occupied Palestinian territory, Peru, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | March 09, 2012 | 563 | 4675 | Not_Computer
Z-2280-2010 | 52030 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ABC PROBE REF 160636 5 mm Handswitching Probe, 36 cm QTY: 1 STERILE CONMED CORPORATION 525 French Rd, Utica, NY 13502 ASSEMBLED IN MEXICOABC PROBE REF 160636 5 mm Handswitching Probe, 36 cm QTY: 10 STERILE CONMED CORPORATION 525 FRENCH RD UTICA, NY 13502 ASSEMBLED IN MEXICOUsage: Electrosurgical handpiece for use as monopolar electrosurgical accessory in conjunction with CONMED electrosurgical units equipped with argon beam coagulation. | ABC PROBE, 160636 | 2 | August 25, 2010 | 2010 | ConMed Electrosurgery | The internal electrode can protrude in front of ceramic insulator at the tip of the ABC Handpiece used with Conmed Electrosurgical units. | DESIGN: Process Design | An Urgent Device Recall notification was issued 12/18/2008, which explained the problem/risk associated with the device, the discontinuance of use and the return of the device. If you have any questions, please contact 800-552-0138, ext. 5225. |  | 184 | Dom: 184 cs (1840 units); Int'l: 47 cs (470 units) | Nationwide and International | Terminated | October 22, 2010 | 58 | 5527 | Not_Computer
Z-2270-2010 | 50226 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. The Smart Disclosure product collects patient waveforms, alarms, vital signs and 12-lead reports. A 24-hour waveform acquisition is standard, with an upgrade option to 72 hours. The product provides basic trends in both tabular and graphical format, along with a retrospective ECG analysis. The analysis may be defined by patient, and the data may be presented in multiple formats, including a rhythm, a histogram, and a summary view. Standard reports are available from Smart Disclosure, including saved events, disclosure, trends histogram and summary. | Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. | 2 | August 25, 2010 | 2010 | Spacelabs Healthcare, Incorporated | Potential for device to print out the wrong waveform data for the time indicated in the printout or display. | DESIGN: Device Design | On 07/17/2008, Spacelabs Healthcare began sending the MEDICAL DEVICE CORRECTION letter to their consignees. The letter informs the consignees that there is an issue involving ICS G2 Smart Disclosure (Model 92810) which causes the display or printing of an incorrect waveform for the time indicated in the printout or display.The consignees are advised whether to discontinue to use the device until the firm upgrades the system. If the consignees decide to continue to use the device, then they should discontinue any use of the "Refresh" button. When they print or view waveform reports,t hey should utilize the Saved Event view functionality of the product rather than Arrhythmia, Alarm or Waveform Views. If the printing from the Arrhythmia, Alarm, or Waveform views, the consignees must visually inspect the print out and verify that the waveform contains what you expected to see printed.The consignees could contact Spacelabs Healthcare at 1-800-522-7052 or 1-425-657-7200, x5089 for any questions about the corrective actions of the affected device. The firm's field service engineer will make a schedule a mutually agreeable time with the consignees to update the ICS G2 system at free of charge. |  | 135 | 135 units | Worldwide Distribution -- US, Turkey, Chile, Brazil, Netherlands, Puerto Rico, Switzerland, Australia, Czech Republic, Canada, Italy, Singapore, France, Germany, and Israel. | Terminated | September 29, 2010 | 35 | 135 | Software | Display/Image | Upgrade | Repair
Z-2269-2010 | 54810 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | ESSENTIAL MEDICAL SUPPLY, INC. ENDURANCE HD, BUILT TO LAST, HEAVY DUTY FOUR WHEEL WALKER for users up to 500 pounds. Model W1801B- Blue, UPC 7 54756 91802 6 all serial numbers.W1801R-Red, UPC 7 54756 91803 3 all serial numbers.The Endurance HD is a four wheeled steel frame walker with loop locking wheel locks, a padded seat and an under seat tote for carrying personal articles. Used to assist in ambulation and specifically for obese patients weighing up to 500 pounds.Warranty Card: "...lifetime of the original purchaser...". Usable life of a walker is 5 years but firm will warranty for lifetime of the original purchaser. | ENDURANCE HD | 2 | August 25, 2010 | 2010 | Essential Medical Supply, Inc. | Essential Medical Supply, Inc. is recalling W1801B and W1801R Endurance HD Walkers due to recent reports and tests that show the front caster may detach from the frame. The potential failures include a collapse of the plastic fork or shearing of the stud attaching the caster to the frame. | OTHER/UNDETERMINED: Pending | All Consignees were mailed an Urgent: Product Recall W1801B and W1801R Endurance HD Walkers notification on December 30, 2009. All notifications were sent first class mail in a 9" X 6" envelope. Response cards were included. The letter identified the affected product and the reason for recall. Customers were asked to examine their inventory for affected product and to discontinue further use or distribution of the product. If the product was further distributed, customers should contact their customers, advise them of the recall and have them return the units to them. Customers should follow the instructions provided for return and replacement of units. | No consumer action necessary | 2954 | 2,954 units | Nationwide Distribution | Terminated | February 21, 2012 | 545 | 2954 | Not_Computer
Z-1981-2010 | 55855 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXIMA Biliary Stent System, 8.5Fr-7cm, M00539260, Sterile EO, Boston Scientific, Spencer, IN. The Flexima Biliary Stent System is intended for delivery of the stent to the biliary tract for drainage of the bile ducts, for splinting of a bile duct during healing, or for providing bile duct patency in a stricture or past a stone. | FLEXIMA Biliary Stent System | 2 | August 25, 2010 | 2010 | Boston Scientific Corporation | Flexima Biliary Stent System Devices were packaged with Tyvek tray lids that had incorrect adhesive, resulting in sterile barrier seal defects. The breach of the sterility barrier could lead to the device being non-sterile. | PRODUCTION CONTROLS: Packaging | Boston Scientific sent an "Urgent Medical Device Recall - Immediate Action Required" letter date April 27, 2010 to all consignees. The letter included instructions to immediately discontinue use and segregate the affected product for return, and complete and return by fax a Reply Verification Tracking Form. Customers with affected product were instructed to contact customer service at 866-868-4004 to obtain a Return Authorization (RGA) number, and package and ship the product to Boston Scientific. |  | 1577 | 1577 | Worldwide Distribution: United States and Puerto Rico, and countries of Arab Emirates, Argentina, Austria, Belgium, Canada, China, Columbia, Croatia, El Salvador, Estonia, Finland, France, Great Britain, Germany, Greece, Hungary, Iran, Israel, Italy, Japan, Korea, Latvia, Lebanon, Malaysia, Mexico, the Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Ukraine, and Uruguay. | Terminated | March 23, 2011 | 210 | 13227 | Not_Computer
Z-2274-2010 | 55933 | accessories, catheter | KGZ | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Arrow Percutaneous Sheath Introducer Kit with Integral Hemostasis Valve/Side Port for use with 7.5 Fr. Catheters; Catalog number AJ-09801 | Percutaneous Sheath Introducer Kit with Integral Hemostasis Valve/Side Port for use with 7.5 Fr. Cat | 2 | August 25, 2010 | 2010 | Arrow International Inc | Sterility not assured. The stopcock handle position was inadvertently changed from an open to a closed position, which prevents sterilization of the closed portion of the device. Sterility of the affected product cannot be guaranteed. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | June 7, 2010 Notification letter to customer instructed them to return the product to Arrow. Arrow will destroy the returned product. |  | 3000 | 3000 kits | All affected product was distributed to one foreign consignee: Arrow Japan Ltd. | Terminated | September 29, 2012 | 766 | 3000 | Not_Computer
Z-2271-2010 | 56276 | syringe, balloon inflation | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MedtronicÂŽ EverestÂŽ 20 Disposable Inflation Device, AC2200.Used to inflate/deflate balloon catheters as well as to monitor pressure within the balloon. | Medtronic Everest Disposable Inflation Device | 2 | August 25, 2010 | 2010 | Medtronic, Inc. | Product sterility may be compromised due to breach of package seal integrity. | PRODUCTION CONTROLS: Packaging Process Control | Medtronic issued a Urgent Medical Device Recall Notice to US customer on 7/15/2010 and to foreign accounts advising of the issue and requesting that they segregate recalled product. A Medtronic Sales Rep will visit their facility to account for all affected product and initiate a credit. Foreign accounts will return product to their respective distribution centers. |  | 2875 | 2,875 units | Worldwide Distribution -- USA, including the state of MI and the countries of Austria, Belgium, Czech Republic, Finland, France, Germany, Greece, Ireland, Israel, Italy, Lithuania, Pakistan, Poland, Portugal, Slovakia, Spain, Sweden, Switzerland, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | June 07, 2011 | 286 | 4897 | Not_Computer
Z-2275-2010 | 56141 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE MONITOR B850Usage: The CARESCAPE TM Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. | CARESCAPE MONITOR B850 | 2 | August 25, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of four potential safety issues associated with the CARESCAPE" Monitor B850:1. Loss of user input when using certain displays or the USB remote control.2. Potential for a delay in treatment or missed alarm when the alarm light is used as a primary alarm source and audible alarms are turned down or off.3. The CARESCAPE Monitor B850 will not use user entered hemo | DESIGN: Software Design | Consignees were sent on 7/16/10, a GE "Urgent Medical Device Correction" letter dated July 12, 2010. The letter was addressed to Healthcare Administrator/Risk Manger, Chief of Nursing, Director of Biomedical Engineering. They stated the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. If you have any questions, please contact Technical Support at 1-800-558-7044. |  | 875 | 875 | CA, CO, FL, GA, KY, LA, MA, MI, NE, NJ, NY, NC, OH, OR, PA, TX VA, and WI, Foreign: UNITED KINGDOM, TURKEY, SWITZERLAND, SPAIN , SLOVENIA, SINGAPORE, SAUDI ARABIA, PORTUGAL NETHERLAND, ITALY, GERMANY, FRANCE,FINLAND, DENMARK, CZECH REPUBLIC, BELGIUM, AUSTRIA, and AUSTRALIA. | Terminated | September 18, 2013 | 1120 | 875 | Other | Alarm/Message | Instructions | Safety Notice/Insructions
Z-2287-2010 | 54474 | suture, absorbable, synthetic, polyglycolic acid | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Deknatel BondekÂŽ Plus Polyglycolic Absorbable Surgical Sutures, Bon+VI R 0 HRX27 1N (70 cm), Teleflex Medical. | Bondek Absorbable Sutures | 2 | August 26, 2010 | 2010 | Telefelx Medical | Labeling error; The Bondek Suture product inside the box are of a different catalog number than labeled on the outer box. |  | Teleflex Medical issued consignees a "Urgent Medical Device Recall" letter dated February 1, 2010. Distributors were instructed to forward the letter to their consignees to retrieve relevant product from those locations. The letter described the packaging issue and potential impact. Customers were instructed to immediately discontinue use and quarantine any identified products, complete and fax the Recall Acknowledgement Form to 866-804-9881, in order to receive a Return Goods Authorization (RGA) Number, and then return the product to Teleflex Medical for replacement. Affected products received by Teleflex Medical will be destroyed upon receipt. |  | 17 | 17 sales units (24 packs per sales unit/408 units) | Worldwide Distribution: USA and Germany | Terminated | April 19, 2011 | 236 | 17 | Not_Computer
Z-2288-2010 | 55772 | board, bed | FPS | Bed board. | General Hospital | General Hospital | 510(K) Exempt | ReliaMed Deluxe Portable Adjusting Table. Product is a medical table used on patients for massage and chiropractic treatment. | ReliaMed Deluxe Portable Adjusting Table | 2 | August 26, 2010 | 2010 | MDC Acquisition Co. | Two complaints were received of the vinyl-covered, padded, plywood table surface breaking near the foot end. | DESIGN: Component Design/Selection | "Urgent: Medical Device Recall" letters were sent to consignees from Milliken Medical and Meyer Distributing Company in April 2010 describing the issue. They instructed their customers to immediately stop use of the table. customers were instructed to return the tables to the firm with an enclosed Recall Form for future replacement, or for replacement with a comparable table from another menufacturer at no cost, or for refund of the original cost of the device. Consignees may contact Meyer Distributing Company at 800-472-4221 and Milliken Medical at 800-532-1356. |  | 265 | 265 Tables | Worldwide distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA and WI, and the country of Canada. | Terminated | February 16, 2012 | 539 | 265 | Not_Computer
Z-2285-2010 | 56542 | instrument, biopsy | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Suros ATEC 0912-20 Breast Biopsy and Excision System Handpiece, REF ATEC0912-12, sterile, Hologic, Inc. 6100 Technology Center Drive Indianapolis, IN. The handpiece is inserted into the breast to the site of interest. When the proper location/position is acheived, the handpiece cuts and removes biopsy tissue. | ATEC 091220 Handpiece | 2 | August 26, 2010 | 2010 | Hologic, Inc. | The firm received a customer complaint that the a device labeled as ATEC 0912-12 was actually ATEC 0912-20. | PRODUCTION CONTROLS: Packaging | The firm, Hologic, Inc., sent out "Urgent Medical Device Recall" letters by FedEx on August 11, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to check their inventories and to quarantine any of the recalled product; if the product was distributed further, identify the users and notify them at once of this product recall; segregate all products from this lot and quarantine them in preparation for return to Hologic, and complete and return the enclosed RECALL RESPONSE FORM by fax at 877-574-3255 or mail. Note: Replacement product cannot be sent until Hologic receives your response form.If you have any questions, please contact Hologic directly at 317-344-7630. |  | 990 | 990 | Worldwide distribution: USA and countries including Canada, Belgium, Algeria, Germany, London, Lebanon and Italy. | Terminated | April 13, 2011 | 230 | 990 | Not_Computer
Z-2283-2010 | 55870 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Delta/Gamma X XL Patient Monitor. Portable monitor used to monitor vital signs of neonatal, pediatric, and adult patients for bedside and transport applications. | Infinity Delta/Gamma X XL Portable Monitor | 2 | August 26, 2010 | 2010 | Draeger Medical Systems, Inc. | Monitor keys may become inoperative or activate spontaneously, causing a membrane switch panel malfunction. This may discharge a patient automatically. |  | Draeger Medical issued "Urgent-Medical Device Recall Notification" letters to affected consignees. The Delta/Gamma X XL monitor keyboards will be updated with new material at each affected facility. If the keys on affected monitors become inoperative or activate without user interaction, the customer is instructed to remove the monitor from service and contact Draeger at 978-379-8573. A Technical Service Bulletin (TSB) will be distributed globally, instructing subsidiaries to inform international customers. | No consumer action necessary | 8675 | 8675 units | Worldwide Distribution - United States, including the states of AL, AR, CA, CO, DC, FL, GA, IA, IL, IN, KS, LA, MA, ME, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, WO, and W, and the countries of United Arab Emirates, Angola, Argentina, Austria, Australia, Bosnia, Belgium, Bulgaria, Benin, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Serbia, Czech Republic, Germany, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Coratia, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, South Korea, Kuwait, Kazakhstan, Lebanon,Lithuania, Latvia, Libya, Morocco, Macedonia, Martinique, Mexico, Netherlands, Norway, Nepal, Peru, Poland, Puerto Rico, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovak Republic, Turkmenistan, Turkey, Ukraine, Uruguay, Venezuela, Mayotte and South Africa. The government accounts are Bay Pines VAMC, Bay Pines, FL; Naval Medical Center, Portsmouth, VA; and Walter Reed Army Medical Center, Washington, DC. | Open |  |  | 9475 | I/O
Z-2292-2010 | 56329 | implant, eye sphere | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer | Sterile BioEye Hydroxyapatite Orbital Implant and Conformer | 3 | August 26, 2010 | 2010 | Integrated Orbital Implants Inc | Integrated Orbital Implants, Inc. has initiated a recall of it's Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer due to a labeling issue in which the expiration date of the product was omitted. | PRODUCTION CONTROLS: Error in Labeling | Customers were contacted by telephone and email. | No consumer action necessary | 40 | 40 units | USA (MT, AL, IN, IA, NY, KY, IL, MI, and FL), Canada, Germany, and Turkey | Terminated | August 31, 2010 | 5 | 40 | Not_Computer
Z-2290-2010 | 56032 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VERSYSÂŽ HIP SYSTEMFEMORAL HEAD 12/14 TAPER 32 MM DIA. -3.5 MM NECK LENGTHZIMALOYÂŽ CO-CR-MO ALLOY STERILE QTY-1MANUFACTURED BY ZIMMER MANUFACTURING B.V.,A SUBSIDIARY OF ZIMMER, INC. (WARSAW, IN)The VerSysÂŽ Hip System Femoral Heads are intended for use in total hip arthroplasty and are assembled onto the proximal taper of the hip stem. The femoral heads manufactured from Zimaloy Alloy are intended for mating with Tivanium" Ti-6AL-4V Alloy and Zimaloy Alloy femoral stems equipped with taper necks. These femoral heads are available in two Morse-type taper options, 6-degree and 12/14, each of which mates with the corresponding 6 degree or 12/14 taper of a femoral stem prosthesis. Each taper option is supplied in a variety of diameters and neck lengths for use in total hip arthroplasty. | VerSys Femoral Head | 2 | August 26, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for a package to be labeled as a 28 mm femoralhead but contain a 32 mm femoral head. In addition, there is a potential for the patient record label tobe incorrect. | OTHER/UNDETERMINED: Pending | The firm, Zimmer, Inc., sent an "URGENT: DEVICE REMOVAL" letter dated July 9, 2010, to all consignees. The letter described the issue, clinical implications, affected product, and required actions. The consignees were instructed to carefully review the letter, ensure users of the VerSys Hip System Femoral Head device at their faculty have been informed of this removal and perform a physical count of all affected product at their facility, record the data on the Inventory Return Certification Form and/or the User Facility/HCP Form and return the form via fax at (574) 372-4265 or email: aimee.wood@zimmer.com. Zimmer Inc. will conduct a removal of all remaining (unused) units of the subject lots. For shipping assistance, questions or assistance in notifying your accounts about the device removal, please contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 103 | 103 | Worldwide distribution: USA and countries including Bolivia, Canada, New Zealand, Taiwan, Thailand, Singapore, and India. | Terminated | January 12, 2011 | 139 | 103 | Not_Computer
Z-2289-2010 | 56328 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer ITST" INTERTROCHANTERIC/SUBTROCHANTERIC FIXATION LOCKING BOLT EXTRACTOR - NONSTERILEZimmer, Warsaw, IN 46580, U.S.A. The ITST Intertrochanteric/Subtrochanteric Fixation Locking Bolt Extractor instrument is a reusable surgical instrument used to separate the femoral nail from the locking bolt after nail implantation/fixation in trauma surgeries | Zimmer ITST" INTERTROCHANTERIC/SUBTROCHANTERIC FIXATION LOCKING BOLT EXTRACTOR | 2 | August 26, 2010 | 2010 | Zimmer Inc. | Instruments were manufactured with the incorrect grade of steel, increasing the potential for fracture during use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Zimmer sent "Urgent: Device Removal" letters dated July 8, 2010 to consignees identifying the potential for fracture of the device. Consignees were instructed to inventory affected product and record on the Inventory Return Certification Form, to document users of the affected device on the User Facility/HCP Form, and return both forms by fax or email. Affected product is to be returned to Zimmer with a copy of the Inventory Return Certification Form. Returned product will be destroyed. Each of the affected consignees will be provided with a new instrument, manufactured with the correct 17-4 stainless steel.Consignees can contact Zimmer at 800-613-6131 or 574-372-4463. |  | 20 | 20 individual units | Worldwide Distribution in the countries of US, Canada, Spain, Germany, and Great Britain. | Terminated | February 25, 2011 | 183 | 20 | Not_Computer
Z-2297-2010 | 55877 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Mitek VAPR S90 Suction Electrode (integrated)Product Code: 225370The VAPR Electrodes (end effect, side effect, wedge, angled, hook, flex, side effect short and wedge short) and VAPR Suction Electrodes (S9ÂŽ and S ÂŽ) with Integrated Handpiece with Integrated Handpiece, when used with the DePuy Mitek VAPR Electrosurgical System, are intended for resection, ablation, excision of soft tissue, hemostasis of blood vessels and coagulation of soft tissues in patients requiring arthroscopic surgery of the knee, shoulder, ankle, elbow and wrist. | Mitek | 2 | August 27, 2010 | 2010 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Product does not meet the required packaging specifications for pouch seal integrity which could lead to a breach in sterility |  | The firm, Depuy Mitek, sent an "Urgent Voluntary Product Recall" letter dated June 7, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to not use any of the identified recalled products; to isolate all inventory of the affected product; complete the Business Reply Form, and return the affected products along with the form to Depuy Mitek; ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or fax completed form to 1-508-828-3750. Note the IMPORTANT PRODUCT INFORMATION: Any product that contains a round green sticker has been 100% inspected, meets specifications and is acceptable for use. Please Do Not Return any product that is labeled with the green dots because this product is not affected by the recall.If you have any questions or concerns with regard to this program, please contact your local Depuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682. | Contact the recalling firm for information | 545526 | 545,526 units | Worldwide distribution: USA and countries including Canada, Argentina, Australia, Bahrain, Belgium, Brazil, Chile, China, Columbia, Costa Rica, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhastan, Kuwait, Lebanon, Maylasia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Palestine, Panama, Peru, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Portugal, Sweden, Switzerland, Syria, Thailand ,Turkey, UK, Ukraine, UAE, Venezuela, Vietnam, and Yemen. | Terminated | July 12, 2012 | 685 | 813657 | Not_Computer
Z-2295-2010 | 56365 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK TWO AST-N085 REF 22 225 For In Vitro Diagnostic Use - each test card is packaged into a foil pouch with 20 pouches per box. The firm name on the label is bioMerieux, Inc., Durham, NC and bioMerieux, sa, France.This is an Export Only test kit that is not marketed in the U.S. | VITEK TWO ASTN085 REF 22 225 | 2 | August 27, 2010 | 2010 | Biomerieux Inc | Foil pouch was punctured by drug susceptibility cards during packaging, allowing moisture into pouch, resulting in inaccurate and/or incorrect drug susceptibility results. | PRODUCTION CONTROLS: Packaging | BioMerieux issued an Urgent Product Removal Notice, dated July 29, 2010 to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and destroy any remaining inventory of the affected lot, and to notify their customers who received the product. All consignees were instructed to complete and return the Customer Acknowledgement and Product Replacement Form. These cards will be used to distribute replacement cards. Consignees can contact BioMerieux at 1-314- 731-8667. | Contact the recalling firm for information | 20 | 1297/ 20 card boxes | Worldwide Distribution: Country of Canada only. No product was distributed in the US. | Terminated | November 02, 2010 | 67 | 1297 | Not_Computer
Z-2294-2010 | 56281 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Greatbatch OptiSeal Valved PTFE Peelable Introducer, Model 1000093-001, 1000093-002, 1000093-003, 1000093-004, 1000093-005, 1000093-006, 1000093-007. Sterilized using Ethylene Oxide. Greatbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. The OptiSeal Valved PTFE Peelable Introducer is designed to reduce blood loss and air intake by providing hemostatic sealing at venous pressures. The device is offered in two configurations, one without a sideport and one with a sideport attached to a segment of extension tubing terminating in a 3-way stopcock. | OptiSeal Valved PTFE Peelable Introducer | 2 | August 27, 2010 | 2010 | Greatbatch Medical | The outer sterile barrier (pouch) packaging for the 12 Fr device is not intact. No breaches in the inner sterile barrier (tray/tyvek lid) have been identified. The OptiSeal introducer and accessories sealed in the tray are sterile. Infection may result if a sterile field is contaminated by the outer surface of the inner tray. | PRODUCTION CONTROLS: Packaging | Greatbatch Medical issued an "Urgent Medical Device Recall letter dated July 12, 2010 via email to consignees. The letter described the problem and affected device. Consignees were instructed to immediately discontinue distribution of the product and return the unused units to Greatbatch Medical and to notify their customers to immediately return unused units. For shipping assistance, questions or assistance in notifying accounts, contact Shannon Springer at 763 951-8244 or Kimberly Briggs, 716 759-5702. |  | 216 | 211 5-pack boxes | Worldwide Distribution: United States, including the states of MN and CA. | Terminated | December 17, 2011 | 477 | 216 | Not_Computer
Z-2296-2010 | 56420 | dissector, surgical, general & plastic surgery | GDI | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | V. MuellerÂŽ Peanut Sponges, Catalog number 23275-470, Sterile, For Single Use Only, Size: 3/8, X-Ray Detectable In Holders, Distributed by Cardinal Health, McGaw Park, IL 60085 USA. There are various uses including as gauze/sponge internal, x-ray detectable and as dissectors. | V. Mueller Peanut Sponges | 2 | August 27, 2010 | 2010 | He, Inc | Sterile product packaging contains weak seals, which may result in compromised product sterility. | PRODUCTION CONTROLS: Packaging | Harod Enterprises, Inc. (HE) issued a "Notice of Recall" letter to Cardinal Health/V. Mueller dated March 1, 2010 identifying the issue and affected product. Cardinal Health/V. Mueller will issue a subrecall letter to all their customers (end users) instructing them to discontinue use of the affected lots and return for replacement. and will track all responses and removal of the product from their private label customers. HE will repackage and/or replace all returned product.Consignees can contact Harod Enterprises, Inc. at 1-706-228-5165. |  | 551 | 551 cases of 100 packages per case | Nationwide: including state of IL. | Terminated | November 02, 2011 | 432 | 551 | Not_Computer
Z-2293-2010 | 55612 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartStart MRx Defibrillator/Monitor with Q-CPR Meter option.The HeartStart MRx monitor/defibrillator has an option (Q-CPR Meter) to provide visual and audible feedback to the rescuer on the quality of CPR. The Q-CPR option offers real time measurement and corrective feedback on the rate, depth and complete release of compressions, ventilations rate and lack of CPR activity in accordance with current CPR guidelines. The Q-CPR option is contraindicated as follows: *The Q-CPR option is contraindicated for use on neonatal and pediatric patients (under 8 years of age or weighing less than 25 kg). *The Q-CPR option is not for use when CPR is contraindicated. | HeartStart MRx Defibrillator/Monitor with QCPRTM | 2 | August 27, 2010 | 2010 | Philips Healthcare Inc. | Reduced sensitivity in the force measurement sensor can result in inaccurate feedback to the user on how well he or she is releasing pressure after each chest compression. This can result in incomplete release of chest compression, which will reduce CPR effectiveness. | DESIGN: Device Design | Philips issued an "Urgent Medical Device Recall notification/Field Safety Notice" dated April 2010 to all consignees. The letter described the issue, how to identify affected devices, and required actions. The consignees were instructed to remove affected devices from service. Philips will contact consignees to arrange on-site replacement of affected meters with meters containing revised labeling and calibration software. The new meters will indicate that meter recalibration is required after approximately 50,000 compression cycles. Consignees returning devices to Philips for routine recalibration will receive free replacement. All CPR meters will be replaced with re-designed meters when they become available.For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | Contact the recalling firm for information For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 72 | 72 units | Worldwide distribution in the countries of USA, Canada, Australia, China, Germany, Iceland, India, Libyan Arab Jamahiri, Netherlands, Saudi Arabia, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 19, 2013 | 907 | 72 | I/O | Output/Calculation | Replace | Remove or Replace
Z-2311-2010 | 50214 | dc-defibrillator, low-energy, (including paddles) | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | External Sterilizable Paddles for use with the LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12, and LIFEPAK 20/20e Defibrillator/Monitor. | External Sterilizable Paddles | 2 | August 28, 2010 | 2010 | Physio Control, Inc. | Sterilization efficacy could not be confirmed for sterilization methods on External Sterilizable Paddles and Pediatric Paddle attachments. | MISBRANDING: Labeling False and Misleading | On 11/14/08, Physio Control began sending the Urgent User Information letter via certified mail to their customers. The letter notified the customers about the modification to the cleaning and sterilization instruction of the External Sterilization Defibrillation Paddles and Pediatric Paddle attachments used with LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12, and LIFEPAK 20/20e Defibrillator/Monitors The customers are informed to discontinue use of the gas sterilization methods listed as 12/88 Ethylene Oxide (EO) or 100% Ethylene Oxide (OE). They are notified that they can continue to use the sterilization method such as STERRAD per existing Operating Instructions, but limit sterilization cycles from 100 to 50 cycles.The customers are advised that prior to and after each use, they should examine cables and connector for damage or signs of wear such as loose connections, damaged pins, exposed wires and cable connector corrosion and examine paddles for scratched or pitted surfaces and stop using product if damage has occurred. Within 45 days, the firm will update the Operating Instruction that will include revised Sterilization Guidelines on a website www.physio-control.com/products/product-support. Customers can call Technical Support at 1-800-442-1142, #5 or they can visit the firm's website at www.physio-control-notices.com/extpaddles regarding the recall. |  | 38185 | 38,185 units | Devices were distributed to medical facilities through out the US and WORLD WIDE to Canada, Germany, India, Netherlands, Australia, Hong Kong, Japan, Brazil, Panama, Russia, Korea, Malaysia, Venezuela, and United Arab Emirates(UAE) . | Terminated | September 28, 2010 | 31 | 39874 | Not_Computer
Z-2310-2010 | 56218 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Genzyme Biosurgery Essentials Kit, Item Number 82004 | Genzyme Biosurgery Essentials Kit, Item Number 82004 | 2 | August 28, 2010 | 2010 | Genzyme Corporation | Revised Labeling of Carticel Essentials Kit clarifiesthe non-sterile packaging of the outer clear plastic tray which should not be opened in the sterile field. | DESIGN: Labeling Design | Genzyme notiifed physicians with the Dear Doctor letter on July 13, 2010 via UPS Ground delivery. The letter informed healthcare workers of the recent incidents and advise them to heed the Essentials Kit label and exercise proper handling of the Kit.. A fax return form will be provided for the physician to sign and return to Genzyme to acknowledge receipt of the notification. ContactGenzyme Medical Information at 800-453-6948, option #3 for any additional information or contact your Genzyme sales representative with any questions. |  | 3132 | 3132 kits | Nationwide | Terminated | January 30, 2012 | 520 | 3132 | Not_Computer
Z-2313-2010 | 56284 | slide stainer, automated | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor Heated Reagent Basins; automated slide stainer; Abbott Molecular Inc., Des Plaines, IL 60018; The basins are sold under the the following part numbers:a) list 02J11-013, part number 30-144152 (replacement basins);b) list 02J11-060, part number 30-144100 (as a component of the 100 volt VP 2000 Processor); c) list 02J11-001, part number 30-144101 (as a component of the 117 volt VP 2000 Processor); d) list 02J11-004, part number 30-144102 (as a component of the 230 volt VP 2000 Processor); Designed to automate and standardize FISH slide specimen processing and routine slide staining for the laboratory. | VP 2000 Processor Heated Reagent Basins | 2 | August 31, 2010 | 2010 | Abbott Molecular | The VP2000 Processor Heated Reagent Basins crack along the bottom seam, have non-uniform coating of interior corners, may be warped, and some may have incorrect dimensions. | PRODUCTION CONTROLS: Reprocessing Controls | Abbott sent Urgent Product Recall letters dated 7/20/10 to the VP 2000 Processor customers on same date via FedEx, informing them that the heated reagent basins could crack along the bottom seam, have non-uniform coating of interior corners, warping of the aluminum housing or incorrect dimensions. These deficiencies could result in leakage or the inability to place the basins into the VP 2000 Processor, resulting in possible exposure to chemical reagents, a slippery floor hazard, and/or delay or failure to start a protocol run. An Abbott Molecular customer service representative will call them to coordinate replacement of all of their heated reagent basins before 10/31/10. Customers were to complete and return the Customer Reply form provided. Any questions were directed to their local Abbott Molecular Representative.*** A second recall letter was sent to all customers on 12/21/10 explaining that the VP 2000 heated reagent basins previously inspected and determined to be acceptable could develop a leak, and that they would be contacted by an Abbott Molecular Customer Service Representative or Field Service representative in early 2011 to coordinate replacement of all of their VP 2000 heated reagent basins. ***. | Contact the recalling firm for information | 219 | 219 basins | Worldwide Distribution -- Australia, Belgium, France, Germany, India, Italy, Japan, Pakistan, Poland, Russia, Singapore, South Korea, Switzerland, Taiwan, Thailand, United Kingdom, Uruguay, and the United States. | Terminated | April 19, 2012 | 597 | 219 | Not_Computer
Z-2315-2010 | 56064 | fish (fluorescent in situ hybridization) kit, protein nucleic acid, enterococcus faecalis | OAH | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | E.faecalis/OE PNA FISHIntended Use: E. faecalis/OE PNA FISH" (Cat. No. KT003) is a multicolor, qualitative nucleic acid hybridization assay intended for identification of Enterococcus faecalis and selected other enterococci (OE) from blood cultures. E. faecalis/OE PNA FISH provides rapid identification of E. faecalis and selected other enterococci (OE) on smears made from positive blood cultures and is compatible with all major blood culture media | E.faecalis/OE PNA Fish | 2 | August 31, 2010 | 2010 | AdvanDx, Inc. | Some strains of streptococus anginosus cross-reach with E. faecalis/OE PNA FISH to produce a false positive green signal. | DESIGN: Device Design | The firm, AdvanDx, sent a letter dated June 8, 2010, and email notification on June 11, 2010, to all customers. The letter described the product, problem and action to be taken by customers. AdvanDx notified the customers that there was a change to the package insert under "limitations" to include the statement "some strains of Streptococcus anginosus produce a false green positive fluorescence due to sequence similarities and enclosed a package insert for all E. facealis/OE PNA FISH (KT003) kits. The customers were instructed to replace the Package Insert (revision B) included with any kits in their inventory with the enclosed Package Insert (revision C). Note: All shipments of E. faecalis/OE PNA FISH, as of June 14, 2010, will contain this new package insert.If you have any questions or requests for further information, please contact AdvanDx Technical Support at US (Toll Free): 1-800-376-0009 or email techsupport@advandx.com. |  | 541 | 541 units | Worldwide distribution: USA and countries including Danmark, Finland and Sverige (European). | Terminated | June 06, 2011 | 279 | 541 | Not_Computer
Z-2335-2010 | 53811 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Abbott Spine PathFinder, Cannulated Poly Screwdriver II - MOD, non-sterile, Manufactured by Abbott Spine, Inc., Austin, TX; Ref. 1160-2.Usage: The screwdriver is intended to mate to the hexalobular feature of a PathFinder open polyaxial screw and threads onto the head to hold it in on-axis for implanting. | Abbott Spine PathFinder, Cannulated Poly Screwdriver II MOD | 2 | August 31, 2010 | 2010 | Zimmer Inc. | The screwdriver may fracture, resulting in surgical delays and/or fragments going into the surgical site. | OTHER/UNDETERMINED: Pending | The sales force was provided with new instructions for use on 1/12/09 and again in 05/09. On April 6, 2010 the firm sent out an Urgent Medical Device Correction to it's field representatives for distribution to it's customers. The field representatives were to distribute the forms and provide a list of the customers that recieved the recall notice.For questions or assistance, please contact Zimmer Spine at 800-326-0635. | Contact the recalling firm for information | 691 | 691 | Worldwide. | Terminated | December 22, 2010 | 113 | 691 | Not_Computer
Z-2334-2010 | 53427 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Knee System Primary Tibial Baseplate; Sterile, Made in Ireland.Howmedica Osteonics325 Corporate DriveMahwah, NJ 07430The Triathlon system is comprised of a femoral component, tibial tray, and tibial insert that are intended to be used in total knee arthroplasty (if replacement of the articular surf ace of the patella is required, the Duracon patellar components are compatible with the Triathlon components). The Triathlon PS Total Knee System is intended to be used in situation where the posterior cruciate ligament is absent, or cannot be preserved. The Triathlon CR Total Knee System consists of a femoral component, tibial insert, and all-polyethylene patellar components that are intended to be used with the previously cleared Triathlon Primary Cemented Tibial Tray in primary or revision total knee arthroplasty. The Triathlon All Polyethylene patellare components are intended to be used with the femoral components of the previously released Duracon Total Knee System, as well as the previously release Triathlon PS Femoral component in situation where replacement of the articular surface of the patella is required. The Triathlon CR Total Knee system is intended to accommodate the posterior cruciate ligament (PCL) if it is present. | Triathlon Tibial Baseplate | 2 | August 31, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | A lot for lot mix up occurred between the Triathlon Total Knee System Primary Tibial Baseplate # 4 Cemented, Catalog # 5520-B-400, Lot SH8NT and the #2 Cemented, Catalog # 5520-B-200, Lot SH7RP. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated September 1, 2009, via FedEx to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to examine their inventory and hospital locations to identify the affected product; retrieve and return all affected products via Inter-Org Transfer and ship to Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430; Attn: Regulatory Reporting, and attach the fluorescent orange PRODUCT REMEDIATION sticker and mark the outer box with the words "PRODUCT RECALL; complete and fax the attached Product Accountability Form and spreadsheet to (201) 831-6069 within 5 days of receipt of the notice. If you have any questions, please contact the Manager, Divisional Regulatory Reporting at (201) 831-5970. |  | 6 | Lot Code: SH7RP - 6 in the US. Lot Code SH8NT was not received in the US. | Nationwide distribution: AL, AZ, CA, MT, NJ and TX. | Terminated | August 27, 2012 | 727 | 6 | Not_Computer
Z-2320-2010 | 56331 | prosthesis, esophageal | ESW | Esophageal prosthesis. | General & Plastic Surgery | Gastroenterology/Urology | 510(k) | Boston Scientific Ultraflex Covered Esophageal Stent System - Proximal Release (18/23/12/9)Material/UPN/Catalog Number: M00513850 | Ultraflex Covered Esophageal Stent System | 2 | August 31, 2010 | 2010 | Boston Scientific Corporation | Units have excessive fraying of the deployment suture thread, as a result, these devices may experience a suture break during deployment. | PRODUCTION CONTROLS: Process Control | Boston Scientific issued "Urgent Medical Device Recall - Immediate Action Required" letters dated July 15,2010 to consignees. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Foreign accounts received notification translated into the local language and with changes that are applicable to local regulations.Boston Scientific can be contacted at 1-508-683-4427. |  | 8 | 8 units | Worldwide Distribution: USA, and the countries of: Austria, Germany, Finland, France, UK, Ireland, Iran, Italy, the Netherlands, Portugual, Spain, and Sweden | Terminated | October 14, 2011 | 409 | 118 | Not_Computer
Z-2314-2010 | 56402 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PERI-LOC (R), 3.5 MM ANTEROLATERAL DISTAL TIBIA PLATE, 10 HOLE, RIGHT - 148 MM, REF 71800710, QTY: (1), STERILE R, CE 0086, Smith & Nephew, Inc., Memphis, TN 38116, USA, EC REP: Smith & Nephew Orthopaedics Gmbh, Tuttlingen, Germany. Adult and pediatric patients, as indicated for pelvic, small, and long bone fracture fixation. | PERILOC (R), 3.5 MM ANTEROLATERAL DISTAL TIBIA PLATE | 2 | August 31, 2010 | 2010 | Smith & Nephew Inc | Inner packaging of sterile product was not sealed. Lack of sterility assurance. |  | Smith & Nephew issued an "Urgent:: Medical Device Recall" letter dated July 20, 2010 to consignees (International distributors), identifying the issue and affected product. Consignees were requested to identify and quarantine affected stock, and cease distribution. Consignees were instructed to contact Smith and Nephew for a Return Authorization number. The Return Authorization number and an inventory of affected devices are to be recorded on the Verification Section of the notification letter and faxed to the firm. Affected stock is to be sent to Smith & Nephew. Consignees were requested to notify anyone to whom the product was distributed to also contact the firm for return of the product. Smith & Nephew can be contacted at 1-901-399-5441. | Contact the recalling firm for information | 10 | 10 units | Worldwide Distribution: INTERNATIONAL ONLY: Countries of China, UK, Italy, Australia, and South Africa | Terminated | November 17, 2011 | 443 | 10 | Not_Computer
Z-2339-2010 | 52830 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba Aquilion ONE CT SystemTSX-301 A/2D | Aquilion ONE CT System TSX301 A/2D | 2 | September 01, 2010 | 2010 | Toshiba American Medical Systems Inc | ECG Gated Dynamic Volume Prospective scan may stop if preset eXam Plan is confirmed and CTA On is changed to CTA OFF, due to exposure control PWB. | DESIGN: Software Design | Toshiba America Medical System Inc (TAMS) sent an Urgent: Medical Device Correction dated January 14, 2009 to their consignees. The letter described the issue and provided instructions to test the device to determine if it is an affected device. If the device is affected, contact the Toshiba service representative at 800 521-1968. Consignees were requested to complete and fax a response form to Toshiba. | Contact the recalling firm for information | 16 | 16 | Nationwide Distribution: including the states of AL, CA, FL, GA, KY, MA, MD, MT, NV, NY, OH, PA, and Puerto Rico. | Terminated | September 02, 2011 | 366 | 16 | Software
Z-2342-2010 | 54265 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eye surgery stretcher with crank fowler, model 1079 Stryker Medical 3800 E. Centre Ave Portage MI 49002. Intended Use: A stretcher used for eye surgery. | Stryker Eye Surgery Stretcher 1079 | 2 | September 01, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The fowler of the bed is raised by a crank. The crank consists of a steel screw that moves through a brass nut. Over time the screw can wear the threads on the brass nut. The screw and the nut can become misaligned, causing a loud squealing noise when the screw comes in contact with the drive tube. In addition the fowler can become inoperable over time with excessive wear of the brass nut. T | DESIGN: Device Design | The firm, Stryker Medical, sent an "Urgent Medical Device Correction" letter dated February 26, 2010, to all customers. The letter describes the product, problem and action to be taken by customers. The letter states that Stryker field service tech will contact the firm and perform device corrections within four months. The customers were instructed to locate the recalled stretchers and verify that there is no grinding noise when activating the fowler. If the fowler exhibited a grinding noise, the customers were instructed to remove the stretcher from service and contact the recalling firm at 800-STRYKER, option 3. Note: Do not put the stretcher back into service until it is repaired by our field service representative. The customers were also instructed to return the enclosed post card to confirm receipt of this notice; inform any new users and advise the firm of new locations if any of the listed stretchers were on loan or sold, and advise the firm of the obsolescence by providing the serial numbers if any of the listed stretchers were disposed of. If you have any urgent questions or concerns, contact (269) 324-6884. Normal hours are Monday-Friday 8am-5pm(EST). | Contact the recalling firm for information | 729 | 729 | Worldwide distribution: USA and countries including Australia, Canada, Chile, China, France, Germany, Greece, Spain, Italy, Japan, South Africa, Taiwan, Korea, Singapore, and UK. | Terminated | April 21, 2011 | 232 | 1866 | Not_Computer
Z-2337-2010 | 55699 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 500 Series System | COULTER LH 500 Series System | 3 | September 01, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter identified four (4) issues with the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System:Issue 1:There is a potential for misidentification to occur when the system is configured in languages other than English or Chinese. The characters #,@,[,\,],`,{,¦,}, or ~ are substituted or omitted in other languages. The issu | DESIGN: Software Design | The customer recall notification was initiated when Product Corrrective Actions letters with attached Customer Response form were sent on the week of February 15, 2010 to all customers who purchased the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System. The letters will provided the customers with an explanation of the problem identified and a work around.Customers with questions regarding the Product Corrective Action were instructed to call 800-526-7694 in the UnitedStates or Canada, or contact your local Beckman Coulter Representative. | Contact the recalling firm for information Contact the recalling firm for information | 1373 | 1373 units in US and 147 units in Canada | Throughout the United States and Canada | Terminated | September 10, 2010 | 9 | 1426 | Software
Z-2348-2010 | 56323 | posterior metal/polymer spinal system, fusion | NQP | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | NFlex Stabilization System 150mmThe Synthes N-Hance and NFix II rods are intended to provide immobilization and stabilization of spinal segments as an adjunct to fusion in the treatment of acute and chronic instabilities or deformities of the spine. | NFlex Stabilization System 150mm | 2 | September 01, 2010 | 2010 | Synthes USA (HQ), Inc. | package integrity compromised |  | The firm, SYNTHES, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 9, 2009 and an updated letter dated February 12, 2009, with updated expiration information to customers. The letters described the product, problem and action to be taken by customer. The customers were instructed to call Synthes at 1-800-479-6329 if they have any of the identified devices to obtain a Return Authorization Number; complete the enclosed Verification Section attachment and return this letter with the recalled product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If they do not have the identified product ,complete the Verification Section attachment and return the document to Synthes by Fax 610-719-5120 or scan/email: lewis.lynne@synthes.com.If you have any questions, please call 800-620-7025 x 5454 or 610-719-5454 or contact your Synthes Spine Sales Consultant. | Contact the recalling firm for information |  |  | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TN , TX, UT, VA, WA, WI, and WV and countries including Denmark, Germany, Korea and Switzerland. | Open |  |  |  | Not_Computer
Z-2355-2010 | 56214 | immunochemical, thyroglobulin autoantibody | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Thyroglobulin Antibody II Reagent, Part Number: A32898. The Access Thyroglobulin Antibody II (TgAb) assay intended use is for the quantitative determination of thyroglobulin antibody levels in human serum and plasma using the Access Immunoassay Systems. | Access Thyroglobulin Antibody II Reagent, Part Number: A32898 | 2 | September 02, 2010 | 2010 | Beckman Coulter Inc. | The affected Access Thyroglobulin Antibody II reagent may produce inaccurate results due to a change in the thyroglobulin that was used in the kits. | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an "Urgent Product Corrective Action" letter dated June 29, 2010, with attached Customer Response form to the affected customers. The letter identified the affected product and issue.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers 916271, 917073, and 918433.(2) Review results in the context of the total clinical presentation of the patient, and rerun questionable results using another reagent.(3) Complete and return a response form within 10 days.(4) For replacement contact Beckman Coulter at (US) 1-800-526-3821, (Canada) 1-800-463-7828. Customer Technical Support can be contacted at 1-800-854- 3633. |  | 1071 | 1071 units | Worldwide distribution : United States, including states of CA, CO, FL, GA. IA, IL, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and country of Canada. | Terminated | May 04, 2012 | 610 | 1071 | Not_Computer
Z-2350-2010 | 56309 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 3200 CS and SL, Hematology Analyzer manufactured by Abbott Laboratories, Santa Clara, CA. Waste Reservoir Assembly Part number 8921174802. | CELLDYN 3200 | 2 | September 02, 2010 | 2010 | Abbott Laboratories | Epoxy holding the two parts of the waste reservoir together fails, potentially resulting in biohazardous fluid leaks. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Product Recall letters were sent on 04/16/2010. The letters identified the affected product, explained the reason for recall, the patient impact/user safety, stated the necessary actions, and contact information. Users are to discontinue use until the Waste Chamber(s) have been replaced and follow the instructions in the Hazards section. In addition, appropriate protective equipment needs to be worn along with following their recommended laboratory procedures dealing with spills. Users are to contact their local Abbott Customer Support rep. US customers should be direct their questions to Customer Support. |  | 69 | 69 units total. | Worldwide Distribution -- CA, CT, ID, IN, and MA; and, countries of Australia, Brazil, Canada, China, Colombia, Mexico, Puerto Rico, Singapore, and Uruguay. | Terminated | December 06, 2010 | 95 | 69 | Not_Computer
Z-2340-2010 | 56364 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Long Nail Kit R2.0, Ti RIGHT 011x360 mm x 125 degrees; Sterile, TI Alloy, Implant; Catalog number 3225-0360S.Stryker Trauma GmbH, 24232 Schenjrchen, Germany;distributed in USA by Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ 07430.Intended for temporary fixation, correction, or stabilization of the right femur. | Gamma 3 Long Nail Kit R2.0 | 2 | September 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The nail has no thread for the set screw hindering insertion. | PRODUCTION CONTROLS: Process Control | Urgent Product Recall letters were sent on July 23, 2010 by Federal Express to Stryker branches. A corrected letter was sent on July 26, 2010 to correct an error in the size of the screw. Additionally, letters were sent on July 26, 2010 to Hospital Risk Managers, Chief of Orthopaedics and Surgeons. The letter identified the affected product, explained the issue, discussed the potential hazard, and the risk mitigation. Customers are to examine their inventory, retrieve all affected product and return it; then reconcile all products on the Product Recall Acknowledgement Form and fax it back. Questions should be directed towards Richard Wolyn at 201-972-2100. |  | 12 | 12 units | Nationwide Distribution -- NJ, NY, DE, IN, WI, IL, PA and NM. | Terminated | August 06, 2012 | 704 | 12 | Not_Computer
Z-2354-2010 | 56489 | handpiece, belt and/or gear driven, dental | EFA | Dental handpiece and accessories. | Dental | Dental | 510(k) | Titan 3 5K Low-Speed Motor - Star Titan 3, 5,000 rpm- Lube Free | Titan 3 5K LowSpeed Motor | 3 | September 02, 2010 | 2010 | Dental EZ Stardental Division | The housing of 5K motor is labeled as a 20K motor. | PRODUCTION CONTROLS: Labeling Mix-Ups | DentalEZ group/Star Dental Division sent a "Medical Device Recall" letter to consignees on 8/11/10, informing them that the device's 5K motor was mislabeled as a 20K motor. Customers are to stop using of the device, and return to the manufacturer or to the distributor using supplied shipping labels. Replacement devices will be sent to customers at no charge. Consignees are to complete and return a form indicating whether they have any affected devices, and whether customers have been contacted about this recall. Customers can contact DentalEZ at 866-383-4636 . | Discard product or return to place of purchase for a full refund or exchange | 41 | 41 units | Nationwide Distribution: States of IN and TX | Terminated | December 13, 2010 | 102 | 41 | Not_Computer
Z-2356-2010 | 54625 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker Pain Pump 2 Luer Cap, REF 525-401, Rx, Sterile, Product of Ireland Manufactured for Stryker Instruments Kalamazoo MI | Luer Cap | 2 | September 03, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Defects include, but not limited to: Not sterile, not properly labeled, drugs and swabs may be ineffective, and catheter tips may break in patients. | PRODUCTION CONTROLS: Process Control | Stryker Instruments issued an "Urgent Medical Device Recall Notification" letter to consignee risk managers, directors and material managers on February 2, 2010, identifying the affected devices and product issues. The consignees were instructed to:1) Forward the recall notice and reply form to all affected locations. 2) Review all inventory and quarantine affected product.3) Destroy all affected product using attached instructions.4) Complete and return the Business Reply Form to the firm by fax.Stryker Instruments can be contacted at 1-9800-800-4236 ext 4354. |  | 18826 | 18,826 all products | Nationwide Distribution: United States, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, and UT. | Terminated | March 28, 2011 | 206 | 18826 | Not_Computer
Z-2380-2010 | 56561 | warmer, thermal, infusion fluid | LGZ |  |  | General Hospital | 510(k) | Vital Signs Inc. enFlow Disposable IV Fluid/Blood Warmer Cartridge. Medical point-of-care device designed to warm blood and IV fluids. | enFlow Disposable IV Fluid/Blood Warmer Cartridge | 2 | September 03, 2010 | 2010 | Vital Signs Inc | The enFlow Disposable IV Fluid/Blood Warmer Cartridge's male luer-lock fitting may be oversized and may not fit properly with standard female luer fittings. | PRODUCTION CONTROLS: Equipment Maintenance | GE Healthcare issued an "Urgent Medical Device Correction" letter, dated July 15, 2010, to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and quarantine the affected devices, and contact Vital Signs Inc. Customer Service to arrange for return and replacement of the affected devices. Consignees were instructed to complete and return the Confirmation Form. Consignees can contact Vital Signs Inc. Customer Service at 1-973-956-5300 . |  | 406596 | 385,820 domestic - 20,776 international | Worldwide Distribution; USA (all states), and to the countries of Australia, Ecuador, Finland, India, Israel, Latvia, Lebanon, Malaysia, Mexico, Morocco, New Zealand, Panama, Philippines, Qatar, Saudi Arabia, Singapore, Switzerland, the Netherlands, and Vietnam. | Open |  |  | 406596 | Not_Computer
Z-2384-2010 | 53819 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper- 903 California Newborn Screening Card Ref 110534903 Card are intended for use by healthcare professionals in the collection and in-vitro storage of neonate human blood. The neonate blood is tested to screen the infant for congenital abnormalities of metabolism and other conditions.790 | Whatman Body Fluid Collection Paper 903 | 2 | September 07, 2010 | 2010 | Whatman, Inc. | Neonate newborn screening card demographic and filter portion may be detached | PRODUCTION CONTROLS: Process Control | The firm, Whatman, issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 29, 2009, to the customer. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: inspect the card and ensure the 903 Paper Portion is firmly and completely fixed to the Demographic portio, prior to recording patient demographic details on the card and obtaining a patient sample; discard the card, obtain a new card and inspect as above, if the 903 Paper Portion of the card partially or completely detaches from the Demographic portion of the card; and to please staple the card stub at two points as specified in the illustration, to prevent any opportunity for 903 paper detachment from the demographic portions.If Whatman customers/users have any questions in relation to this customer notification, please contact, either Product Manager at +1 732 885 6570,email: lili.lee@ge.com or Quality Assurance at +1 207 459 7557, email: heidi.kuell@ge.com. |  | 488000 | 488,000 cards | Nationwide distribution: CA | Open |  |  | 488000 | Not_Computer
Z-2383-2010 | 56488 | electrode, electrocardiograph | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | 3M Red Dot Monitoring Electrode, Catalog #2560 & Novaplus Universal Monitoring Electrode #V2560, used for ECG monitoring. | 3M Red Dot Monitoring Electrode, Catalog 2560 &38 & Novaplus Universal Monitoring Electrode V2560 | 2 | September 07, 2010 | 2010 | 3M Company / Medical Division | The sensing element of the 3M Red Dot Monitoring Electrode, Catalog # 2560 and Novaplus Universal Monitoring Electrode #V2560 may have corrosion. In an emergency situation where defibrillation is necessary and electrodes are already in use, the electrodes may malfunction and provide a delayed or no trace response. | PRODUCTION CONTROLS: Process Control | 3M sent an "Urgent Medical Device Recall" letters and a "Distributor Product Recall Form" dated July 29, 2010 to customers and distributors of the 3M and Novaplus electrodes. The letters described the 3M Red Dot and Novaplus Universal products, problems and actions to be taken by customers. Customers were instructed to examine their inventory and quarantine the affected lots pending return to the firm. After completing and returning the enclosed Customer Product Recall Form to 3M, customers will be issued a Return Authorization Number and provided with instructions for return of the affected devices. Distributors were instructed to notify customers to the end user level and have them return affected lots to the distributor for forwarding to 3M, or forward the customer letter to end-users to instruct them on how to contact 3M for recovery of the device. The 3M Customer Helpline can be contacted at: 1-800-228-3957. | Contact the recalling firm for information | 50 | 11,063 cases (20 pouches/case, 50 electrodes/pouch) | Nationwide Distribution: States of CA, FL, IL, KS, KY, MI, MO, NE, NJ, NY, OH, PA, RI, SD, TN, VA, and WI. | Terminated | April 19, 2012 | 590 | 50 | I/O
Z-2382-2010 | 56315 | pump, blood, cardiopulmonary bypass, non-roller type | KFM | Nonroller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Impella 5.0 Circulatory Support SystemCatalog number: 004840The Impella 5.0 System is intended for circulatory support using an extracorporeal bypass control unit, for periods up to 6 hours | Impella 5.0 Circulatory Support System | 2 | September 07, 2010 | 2010 | Abiomed, Inc. | Pump may separate at the motor housing while device is explanted |  | ABIOMED issued the recall notification letter on July 5 , 2010 to the one affected customer (France) and requested that they segregate and return any unused pumps. The pumps in Canada were already used.If you have any questions, contact +49 241 8860-0 or e-mail europe@abiomed.com. | Contact the recalling firm for information | 10 | 10 units | International only: Canada and France | Terminated | September 10, 2010 | 3 | 10 | Not_Computer
Z-2385-2010 | 54460 | ferrozine (colorimetric) iron binding capacity | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Dimension IBCT Flex Reagent Cartridge. The total (IBCT) method used on the Dimension system is an in vitro diagnostic test intended to quantitatively measure Total Iron Binding capacity in human serum and plasma. | Dimension IBCT Flex Reagent Cartridge | 3 | September 08, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Test produces falsely elevated IBCT results and Abnormal Reaction test report messages on heparinized plasma samples. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens issued an "Urgent Field Safety Notice" letter dated December 2009 to users identifying the issue and affected product. Customers were instructed not to use heparinized plasma samples for IBCT testing. IBCT Instructions for Use (IFU) has been updated to reflect that heparinized plasma is no longer an acceptable sample type. Beginning with lot ED0341, the new IFU will be packaged in the carton with the IBCT FlexÂŽ reagent cartridges. Consignees were requested to forward the notification letter to anyone to whom the product was distributed. The Siemens Technical Solutions Center can be contacted at 800-441-9250. | Contact the recalling firm for information |  | unknown - all lots involved | Worldwide Distribution: USA (all states) and the countries of Canada, Switzerland, New Zealand, Australia, Japan, and South Korea. | Terminated | October 25, 2010 | 47 |  | Not_Computer
Z-2386-2010 | 54624 | enzyme immunoassay, n-acetylprocainamide | LAN | Digoxin test system. | Toxicology | Toxicology | 510(k) | ONLINE TDM N-acetylprocainamide, Part number 04642520190 N- acetylprocainamide reagent Roche Diagnostics Corporation Indianapolis, IN. | N acetylprocainamide reagent | 2 | September 08, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The N-acetylprocainamide reagent lot 14456200 shows an incorrect expiration date of 7/31/2011. The correct expiration date for the lot is 1/31/2011. | PRODUCTION CONTROLS: Error in Labeling | Roche issued an "Urgent Medical Device Removal" letter dated February 2010 to consignees identifying the issue and affected product. Customers were requested to complete the attached form and return in to the firm to obtain replacement product. After receipt of replacement reagent, the customer should discard the affected product product. If unaffected lots are not available, customers may use the affected lot until replacement product is received. Use of the affected product before the corrected expiration date of January 31, 2011 should not produce erroneous test results. Roche Diagnostics Technical Support can be contacted at 1-800-428-2336. | Contact the recalling firm for information | 40 | 40 kits | Nationwide Distribution: USA states of AL, CT, NY, PA, and TX. | Terminated | November 01, 2010 | 54 | 40 | Not_Computer
Z-2332-2010 | 56162 | ventilator, emergency, manual (resuscitator) | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | BagEasy manual resuscitation devices are used for patients requiring complete or intermittent ventilatory support. It can provide positive pressure ventilation or oxygen for the spontaneously breathing patient with a face mask or through an artificial airway.A manual emergency ventilator is a device, incorporating a bag and valve, intended to provide emergency respiratory support by means of a face mask or a tube inserted into a patient's airway. | BagEasy Resuscitatior | 1 | September 08, 2010 | 2010 | Westmed Inc | This recall was initiated after Westmed, Inc. became aware of a potential for disconnection at the patient port retention ring assembly of the Bag Easy manual resuscitation device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Westmed, sent a letter dated June 21, 2010, to all customers. The letter describes the product, problem and action to be taken by customer. The customers were instructed to cease distribution immediately, notify any facilities that they may have distributed the product identified of the recall, call Customer Service Manager at 1-800-975-7987 x 1223 and/or email kpeterson@westmedinc.com to receive Return Authorization to return identified product, and complete and return the enclosed Bag Easy Recall Reconciliation Checklist.If you have any question, contact Customer Service at 1-800-975-7987 or 1-520-294-7987. | Contact the recalling firm for information | 22145 | 22145 | Worldwide distribution. | Terminated | March 22, 2011 | 195 | 24879 | Not_Computer
Z-2381-2010 | 56523 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Inc. Symbiq One- and Two- Channel Infuser, List Numbers 16026 and 16027 respectively. | Symbiq Infuser | 1 | September 08, 2010 | 2010 | Hospira Inc | If the administrator set is removed prior to the cassette carriage fully opening without closing the slide/roller clamp, unrestricted flow and delivery may result. | DESIGN: Device Design | Hospitra Urgent Device Correction Notification mailed on 03/26/2010 via federal express. Tracking done by delivery tracking. Distribution of product continues and has not stopped. Additional notification was issued in 09/2010. A software correction is planned. | Contact the recalling firm for information | 28485 | 28485 units | Worldwide distribution: USA, Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | March 30, 2012 | 569 | 28485 | Other | Treatment/Delivery/Therapy | Software update | Software Update
Z-2336-2010 | 55852 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Mayfield Infinity XR2 Base Unit, model A-2079, manufactured by Integra LifeSciences Corporation. The MAYFIELD Infinity XR2 Base Unit (A-2079) is an operating room table accessory intended to be used to support a patient during diagnostic examination and/or surgical procedures where a rigid support between surgical table and headrest, or skull clamp is necessary, positional freedom is required and where x-ray imaging modalities will be used. | Mayfield Infinity XR2 | 2 | September 09, 2010 | 2010 | Integra LifeSciences Corporation | Storage case incorrectly labeled. The labels were for the Infinity XR2 Skull Clamp storage containers but were inadvertently placed on the Infinity XR2 Base Unit storage cases. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Integra Life Sciences Corporation,contacted their customers via telephone and Field Correction Letter on May 19, 2010 to notify their customers of the problem. The firm indicated that their sales representatives will visit the consignees affected, correct the case labeles according to the specific instructions indicated, revise their account tracebility records to reflect the labeling change, then, sign and return the Acknowledgement of Reciept and Return Form for confirmation. For questions regarding the correction, please contact Integra Life Sciences Corporation at (877) 44-1114 or (513) 533-7923. | No consumer action necessary No consumer action necessary | 13 | 13 Units | Product was shipped to the following states: IN, MO, NJ, TX and VA. Product was also sold in Belgium, Canada and New Zealand. | Terminated | October 27, 2011 | 413 | 13 | Not_Computer
Z-2387-2010 | 54743 | calibrator, multi-analyte mixture | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista Chemistry 2 Calibrator, Catalog number KC120.Used to calibrate Phosphate, Salicylate, and Triglycerides methods on the Dimension Vista system. | Dimension Vista Chemistry 2 Calibrator | 3 | September 09, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | There is incorrect barcode information, which can potentially cause the Alert Message "Calibrator Insert Missing" to be displayed. | TRAINING: Employee Error | The recalling firm issued an Urgent field Safety Notice, dated February 2010, to all affected customers informing them of the problem. Customers have been told to run calibrators from cups when calibrating or request a replacement lot. The notification should be forwarded to that this product was further distributed to. Siemens Technical Solutions Center should be contacted at 800-441-9250, if they are any technical questions. |  | 1310 | 1310 cartons | Worldwide Distribution -- USA, Austria, Germany, Netherlands, Italy, France, Denmark, Portugal, Canada, South Korea, Malaysia, Australia, and New Zealand. | Terminated | October 13, 2010 | 34 | 1310 | Not_Computer
Z-2393-2010 | 56333 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare AC-Powered Adjustable Hospital Bed.Model 5890 IVC and 5490 IVC.Manufactured by:Invacare Corp. 2101 East Lake Mary Blvd., Sanford, FL 32773. |  | 2 | September 09, 2010 | 2010 | Invacare Corporation | Invacare Corporation, Florida issued a device correction in October 2006 for the IVC Bed Foot Actuator used on the IVC Bed Series, Models 5890 IVC and 5490 IVC. The actuators were not in the fully extended position and could not be mounted onto the beds. All affected consignees were notified of this correction in October 2006. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Invacare Corporation issued a notice of an Urgent Field Correction to Consignees on 10/18/2006. Consignees were instructed to check the serial number of their units to see if they were potentially affected. If so, the firm provided inspection and adjustment instructions of the foot actuator to prevent a likelihood of injury. After adjustment, the consignees were requested to complete and return the enclosed tracking sheet acknowledging receipt of the recall notice and correction of the bed(s). For questions or assistance regarding this correction, please contact the Customer Relationship Management Group at (800) 333-6900, or via e-mail at CRMGroup@Invacare.com. |  | 731 | 731 | Product was distributed throughout the U.S. | Terminated | September 13, 2010 | 4 | 731 | Not_Computer
Z-2389-2010 | 56424 | alkaline picrate, colorimetry, creatinine | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Systems Modular Creatinine (CREm) Reagent Kit, Part Number: 472525Systems AQUA CAL 1 and 2, is intended for the quantitative determination of creatinine concentration in human serum, plasma or urine.CREm reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 800 System and SYNCHRON Systems AQUA CAL 1 and 2, is intended for the quantitative determination of creatinine concentration in human serum, plasma or urine. | Synchron Systems(s) Creatinine Reagent (CREm) | 2 | September 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports that black / brown particulates were observed in some bottles of the Creatinine Alkaline Buffer (Buffer lot M911427 and M911518) in modular Creatinine (CREm) reagent kit lots M911428 and M911519. Beckman Coulter have had reports of erroneous results when using the lots of Creatinine Alkaline Buffer listed above. Singlecreatinine d |  | The firm, Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 3, 2010, to all customers. The letter describes the product, problem, and action to be taken by customer. The customers were instructed to: (1) Please discontinue use of Creatinine (CREm) reagent kit lots M911428 and M911519 and discard appropriately.(2) Please contact Beckman Coulter at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discussCREm kit M911428 and M911519 replacement options. Outside the United States and Canada, please contactyour local Beckman Coulter Representative.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.(4) Please share this information with their laboratory staff and retain this notification as part of their Quality Systemdocumentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, they were instructed to please contact Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/SupportRequests, or call (800) 854-3633 in the United States and Canada. | No consumer action necessary | 3106 | 3106 units | Worldwide Distribution - USA including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, KS, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY and country of Canada | Terminated | July 02, 2012 | 662 | 3106 | Not_Computer
Z-2390-2010 | 56464 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq One-Channel InfuserLIST No. 16026 Symbiq One-Channel Infuser,List Number Configurations: 16026-04-01/03/04,51/52, 53/54, 79/80, 81/82, 83/84, 87/88 and 89/90,16026-13-25/26, 53/54,Distributed by: Hospira, Inc. Morgan Hill, CA | Symbiq Infuser | 2 | September 09, 2010 | 2010 | Hospira Inc | The clamp may not secure to the IV pole and the pump may slip or fall off. The pump may fall on a person or pull the tubing out of the IV container. | DESIGN: Device Design | Hospira issued an "Urgent: Device Field Correction" Letter dated February 22, 2010 to all consignees, identifying the reason for the recall and affected product. Customers were informed that they will be contacted by a Hospira representative to replace the affected pole clamps. Consignees can contact Hospira Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA (all states and Puerto Rico) and the countries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Singapore. and the UAE. | Terminated | April 01, 2011 | 204 |  | Not_Computer
Z-2388-2010 | 55359 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLUCm used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number: 472500 ( ynchron LX20, LX20 PRO, LXI 725, UniCel DxC 600, UniCel DxC 600 PRO, UniCel DxC 600i, UniCel DxC 800, UniCel DxC 800 PRO, UniCel DxC 880i, UniCel DxC 680i, UniCel DxC 660i, and UniCel DxC860i ), | Synchron LX and UniCel DxC Synchron Clinical Systems | 2 | September 09, 2010 | 2010 | Beckman Coulter Inc. | UniCel DxC and Synchron LX instruments produced inaccurate (low) GLUCm results. False low results could affect or delay diagnosis or treatment. | OTHER/UNDETERMINED: Pending | Beckman Coulter issued an "Urgent: Product Corrective Action" (PCA) letter dated April 2010 to all consignees, identifying the reason for the recall, the affected lots, and the potential for inaccurate GLUCm results due to bubble retention and short samples. Customers were instructed to:1) establish specific protocols for samples from defined patient populations or sample types, which could include running specimens using GLUCm in duplicate or using cartridge glucose.2) To reduce the possibility of inaccurate results, ensure that all elements are properly installed and secured, the sample probe is adequately tightened, and proper sample volume and integrity are maintained. A copy of the PCA letter will also be included in all new instrument ship kits.Consignees were instructed to complete and return an enclosed response form. Consignees can contact Beckman Coulter at 1-973-956-5300 . |  | 84186 | 84,186 units (79,071 in US & 5,115 in Canada) | Worldwide Distribution: USA and Canada | Terminated | December 31, 2013 | 1209 | 84186 | Not_Computer
Z-2392-2010 | 56133 | forceps, general & plastic surgery | GEN | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | BIOMET MICROFIXATION REF: 02-0461 REES INSULATED FORCEPS, 21CM, (8.25") (L), STAINLESS STEEL, NYLON, ***NON-STERILE PRODUCT***. The Rees Insulated Forceps are used during surgery to grasp, dissect and coagulate selected tissue. Distributed by:BIOMET MICROFIXATION, 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218, www.biometmicrofixation.com | Rees Insulated Forceps | 2 | September 09, 2010 | 2010 | Biomet Microfixation, Inc. | The forceps insulation may have cracks or breaks in the insulation that may cause burns to patients when used. | PRODUCTION CONTROLS: Manufacturing Material Removal | Biomet Microfixation issued a letter dated July 2010 to their consignees. Each consignee was ask to discontinue use of the affected product and return the product back to the company for credit. Biomet asked their customers to complete a "Inventory Reconciliation" form and fax the completed form to Biomet Microfixation. If the consignee has further distributed the product, the consignee is asked to contact their customers. Biomet Microfixation can be contacted at 904 741-4400, ext 468. |  | 130 | 130 units | Worldwide Distribution: USA including states of CA, FL, GA, ID, IL, IN, MA, MD, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, WA, and WI, and countries of Norway, Canada, and Mexico. | Terminated | June 21, 2011 | 285 | 130 | Not_Computer
Z-2398-2010 | 55700 | phosphomolybdate (colorimetric), inorganic phosphorus | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Clinical Chemistry Phosphorus assay | AEROSET, ARCHITECTURE C8000 & C16000 | 2 | September 10, 2010 | 2010 | Abbott Laboratories Inc. | Product may give inaccurate results and aspiration errors with reagent configuration, including results outside of linear range & calibration failure. | OTHER/UNDETERMINED: Pending | Abbott Laboratories issued a "Product Recall: Immediate Action Required" letter dated May 4, 2010 to consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to:1) Identify affected product in inventory2) Discontinue use and destroy any remaining inventory of the affected lot: 7D71-213) Complete and return an enclosed Customer Reply Form4) Order alternative kit sizes for Clinical Chemistry Phosphorus: LN 7D71-31 or 7D71-225) For AEROSET Systems, manually update parameters as described in the letter.6) For ARCHITECT cSystems, implement alternative kit sizes without system changes.7) Provide a copy of the letter to other laboratories that have been distributed the affected product.US customers can contact Abbott Customer Service at 1-877-4ABBOTT. Customers outside the U |  | 3136 | 3136 | Worldwide Distribution: US, including Puerto Rico, and countries of Germany, New Zealand, Trinidad/Tobago, Cayman Island, Hong Kong, Chile, Curacao, Jamaica, Australia, and Bahamas. | Terminated | April 05, 2012 | 573 | 3136 | Not_Computer
Z-2395-2010 | 55478 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion PumpsIntended use:deliver insulin therapy | Animas 2020 Insulin Infusion Pumps | 2 | September 10, 2010 | 2010 | Animas Corporation | sensor could be damaged | PRODUCTION CONTROLS: Equipment Maintenance | An Email was sent by the firm, Animas, to US patients and health care providers where email addresses were on file. Animas sent two "URGENT: MEDICAL DEVICE CORRECTION " letters one dated March 18, 2010, to all US customers, HCPs and distributors of One Touch Ping Glucose Management System and the Animas 2020 Insulin Infusion Pump and the other letter to all US patients, HCPs and distributors of Animas 2020 Insulin Infusion Pump, Animas IR 1200 and Animas IR 1250 (1200 plus). Animas also notified all global distributors and affiliates and sent the letter to customers and health care professional in those regions. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow their User Guide instructions; Pay attention to all pump warnings; NEVER rewind the pump motor, load the insulin cartridge, or prime their infusion set while connected, and fill out the enclosed reply card and send it back to Animas as soon as possible.Should you have any questions or concerns, please call the Animas Customer Support Department at 1-866-949-1525 or call toll free number on back of pump at 1-877-937-7867. If you are traveling outside of the U.S., please call 610-644-8990 and press option 1. |  |  |  | Worldwide distribution: USA including states of CA,FL,LA, MA, MI, NY, OH, PA,TX, UT, VA and WA, and countries including Canada, United Kingdom, Finland, Germany, France, Hungary, Puerto Rico, Australia, New Zealand, Israel, Sweden, Czech Republic, Spain, Mexico and Norway. | Terminated | October 27, 2010 | 47 |  | I/O
Z-2349-2010 | 56330 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Picasso/Picasso Lite lasers Intended use: Dentistry. | Picasso/Picasso Lite lasers | 2 | September 10, 2010 | 2010 | AMD Lasers, LLC | Product is noncompliant with the Federal laser product performance standard in certain requirements. The ON/OFF key on some of the Picasso/Picasso Lite lasers could be removed with force while the key was in the ON position. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Letters, dated April 30, 2010, addressed to Dental Professionals were sent identifying the affected product and reason for the recall. The letter asked customers to check their units and, if applicable, AMD Lasers would replace the key assembly on their lasers. The service would be provided free of charge including round trip shipping charges to the corporate office. Customers are to call to schedule the service and obtain a RMA number and shipping instructions. Questions should be directed to 1-866-999-2635 or support@amdlasers.com. | Contact the recalling firm for information | 3000 | 3,000 units | Worldwide Distribution | Terminated | April 21, 2011 | 223 | 3000 | Not_Computer
Z-2400-2010 | 55610 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus CS Acute Care Troponin I TestPak ( Troponin I assay) | Stratus CS Acute Care Troponin I TestPak | 2 | September 11, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | False elevated Stratus CS CTNI results | PRODUCTION CONTROLS: Process Control | Recalling firm issued " Urgent Field Safety Notice" dated April 2010, informing users of the problem and to discuss with laboratory directors regarding the need to review previous test results, conduct patient follow up, and/or repeat testing for tests conducted on or after Nov 1, 2009. | Contact the recalling firm for information Contact the recalling firm for information | 7500 | 7500 cartons | Nationwide Distribution | Terminated | October 25, 2010 | 44 | 7500 | Not_Computer
Z-2399-2010 | 56271 | laser, ophthalmic | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Peregrine 20 GA Endo Ocular Adjustable Laser Probe | Peregrine 20 GA Endo Ocular Adjustable Laser Probe (Smart Flex) | 2 | September 11, 2010 | 2010 | Peregrine Surgical Ltd | Product did not have proper 510(k) clearance | OTHER/UNDETERMINED: Pending | A recall letter is being sent to all affected accounts with instructions to return the recalled product to the recalling firm. | Contact the recalling firm for information | 132 | 132 boxes | Product was distributed to hospitals and eye surgery centers in Florida, Maryland, Missouri, North Carolina, Ohio, Pennsylvania, Texas. | Terminated | September 30, 2010 | 19 | 132 | Not_Computer
Z-2401-2010 | 56268 | tubing, pump, cardiopulmonary bypass | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet HLM Tubing Pack; Bioline Coating | Maquet HLM Tubing Pack | 2 | September 12, 2010 | 2010 | Maquet Inc. | The color of the tape on the arterial and the venous line were placed incorrectly, causing the connection of the arterial and the venous lines to be interchanged. | PRODUCTION CONTROLS: Process Control | Recall notification letters were sent by Federal Express on July 13, 2010 to affected account and also an e-mail to the Maquet field representative who services the account. | Contact the recalling firm for information Contact the recalling firm for information | 5 | 5 units | Distribution: Pennsylvania | Terminated | August 10, 2011 | 332 | 5 | Not_Computer
Z-2114-2010 | 56370 | microwave oven, commercial | RCS |  |  |  |  | Microwave Ovens, Model Number DG-2 | Duke Microwave Ovens | 2 | September 12, 2010 | 2010 | Duke Manufacturing Co. | The oven did not meet Federal Performance Standard for certain requirements. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Letters were sent to Duke Manufacturing Speed Oven Owners. The letters identified the affected product and stated the reason for recall. The letter also states that the firm instituted a repair program on all DG-2 Speed Ovens. The firm will repair of replace all affected products at no charge to the customer. Customer Service personnel will contact customers directly and coordinate the schedule of the required repairs. |  | 62 | 62 units | Nationwide Distribution | Terminated | April 17, 2012 | 583 | 62 | Not_Computer
Z-2402-2010 | 56336 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | CareGuard Mattress Models CG10180 and CG10180A used with IVC and ValueCare Series of manual, semi-electric and electric beds. Mattresses can be identified by looking on the law label tag for the following information: Model # CG10180 or CG10180CA and a BLACK date code with the corresponding month and year of manufacture (i.e. September 2006. October 2006...August 2007). Manufactured by: Invacare Corp., 2101 East lake Mary Blvd. Sanford, FL 32773. Affected bed package model numbers: BED24IVC, BED25IVC, BED24CA or BED25CA, | CareGuard Mattress Models CG10180 and CG10180A used with IVC and ValueCare Series Beds | 2 | September 12, 2010 | 2010 | Invacare Corporation | In October 2007 Invacare Corporation, Florida recalled CareGuard Mattresses, Models CG10180 and CG10180A, associated with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Beds, when used with Invacare bed rails. Affected bed packages are BED24IVC, BED25IVC, BED24CA, and BED25CA. The mattresses when used in combination with the IVC a | DESIGN: Device Design | Consignees were notified of the recall by letter dated October 15, 2007 and entitled "CareGuard Mattresses Provided Since September 2006 In combination with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds, and the Value Care Bed Series Semi Electric Bed, and Full and Half-Length Rails". Depending on the mattress usage the consignees were instructed to either: A: 1) Move the bed rail cross bar from the 5th and 6th spring location on the head section of the bed to the 6th and 7th spring location on the head section of the bed in accordance with the enclosed instructions and document your actions on the enclosed tracking sheet. 2) Return of the Product Tracking Report indicating disposition of the recalled product and acknowledgement of receipt of the recall package. 3) The tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credit. or B) 1) contact Customer Service and request part No. 6022, which is a kit that contains a 2" convoluted foam pad, warning label and installation instructions. Testing has shown that the addition of the foam pad on a new mattress with Invacare half rails / bed combination meets the Guidance requirements." 2) When you have completed placing the label and foam topper on all of your affected mattresses, return the tracking sheet to us and we will credit your account $10 for each mattress you have listed. the tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credi. | Contact the recalling firm for information Contact the recalling firm for information | 19143 | 19143 | Nationwide Distribution | Terminated | September 13, 2010 | 1 | 19143 | Not_Computer
Z-2403-2010 | 56335 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | Mattress Model 5185 and 5185XL. Manufactured by:Invacare Corp., 2101 East Lake Mary Blvd., Sanford, FL 32773.Mattress Model 5185 and 5185XL were packaged in the IVC Bed Models BED2-1633, BED4-1633, BED8-1633, BED9-1633, BED10-1633, BED12-1633, BED38LOW-1633, VCPKGIVC-1633 OR VCPKGIVC2-1633. | Mattress Model 5185 and 5185XL used with IVC and ValueCare Series Beds | 2 | September 12, 2010 | 2010 | Invacare Corporation | In September 2007 Invacare Corporation recalled the innerspring mattresses associated with Invacare's IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Bed. These mattresses were manufactured between June 28, 2007 and July 12, 2007. These mattresses when used in combination with the IVC and ValueCare Series of manual, semi-electric a | DESIGN: Device Design | Consignees of the product were notified by letter dated 09/05/2007. Consignees were instructed to destroy the affected mattresses and to return the completed confirmation of the destruction form to Invacare. | Contact the recalling firm for information Contact the recalling firm for information | 4896 | 4896 | Nationwide Distribution | Terminated | September 13, 2010 | 1 | 4896 | Not_Computer
Z-2404-2010 | 46217 | device, spot-film | IXL | Spot-film device. | Radiology | Radiology | Enforcement Discretion | Diagnost 56/66/76/76 Plus, DSI is a Spot-Film and has a Digital Imaging Spot.Intended use is for the following applications: as a multifunctional/universal system, general R/F, Fluoroscopy, Radiology, and Angiography can be performed along with pediatric examinations and more specialized interventional applications. | Diagnost 56/66/76/76 Plus | 2 | September 13, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential for the tower (holding the serial changer and/or image intensifier) to slide unexpectedly and potentially injure a patient or user in its path. | OTHER/UNDETERMINED: Pending | On 1/4/08, the firm issued the URGENT PRODUCT CORRECTION letter dated 1/3/08 to their consignees via certified mail, return receipt requested. The letter informs consignees the issue of the safety switch and counter weight steel band on the Diagnost 55/66/76Plus, Easy Diagnost, and Easy Diagnost Eleva X-ray systems. However, the Easy Diagnost Eleva X-ray system was distributed and applied only to foreign consignees outside US.If the consignees hear the primary steel band clicking/rattling noise while tilting the table to 90 degrees table or positioning the serial changer, the consignees are instructed to tilt the table base towards the 0 degree tilting position and do not use the tilting drive movement until the mandatory action by a Philips service representative has been performed. The Field Change Order (FCO) would be implemented on 01/23/08 and expected to be completed by July 2008. The firm's representatives will visit sites and will inspect all systems where the switch has been replaced or adjusted for proper operation of the safety sensor and switch. Two steel bands will be inspected and replaced if any abnormity is noted.In the mean time, consignees should contact Philips Call Center at 800-722-9377, #5, #2 and reference the FCO 70600029 with any questions. |  | 28 | 28 units | Nationwide Distribution | Terminated | September 13, 2010 | 0 | 38 | Not_Computer
Z-2420-2010 | 55923 | instrument, surgical, orthopedic, ac-powered motor and accessory/attachment | HWE | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Bur Guard REF 2296-301, Stryker Instruments, Kalamazoo, MI (USA)The Impaction Bur Guard is intended to be used with the Impaction Drill. The drill is used for surgical procedures involving precision drilling and cutting of bone and hard tissue. The bur guard covers the nose of the handpiece and the shaft of the cutting accessory to prevent exposure to heat and moving parts, as well as minimize bur whip and handpiece vibration. | Stryker Bur Guard REF 2296301 | 2 | September 13, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The reason for the field action is that the inner diameter of the Bur Guards may be out of specification. If there is excess material on the inside of the Bur Guard at the location where the guard tapers, it may come into contact with the Nose Cone of the drill. This could result in an injury or burn to the patient and/or user. | DESIGN: Device Design | Stryker Instruments sent an Urgent Medical Device Recall Notification dated June 10, 2010, to their customers. The letter described the product the problem and the action to be taken by the customers. The customers were instructed to:1) immediately review this recall notification.2) forward this letter and the attached Business Reply Form (BRF) to all affected locations. 3) Immediately check and quarantine any affected product found.4) Destroy all affected product by breaking off at least 3 of the tabs on the bur guard and then dispose of the affected product in normal trash.5) Complete and sign the BRF and fax to Stryker Instruments Regulatory Department, 866-521-2762.NOTE: Upon receipt of the BRF, a replacement bur guard will be sent to the customer's account for each affected bur guard destroyed.For questions regarding this recall please contact Stryker Instruments:1-800-800-4236, ext. 3808jennifer.mars@stryker.com. | Contact the recalling firm for information | 10982 | 10982 | Worldwide Distribution: US AND COUNTRIES OF SCANDINAVIA, AUSTRALIA, CANADA, CHILE, CHINA, BRAZIL, FRANCE, GERMANY, GREECE, SOUTH AFRICA, ITALY, JAPAN, MEXICO, SWITZERLAND, PORTUGAL, SINGAPORE and UK. | Terminated | April 21, 2011 | 220 | 10982 | Not_Computer
Z-2407-2010 | 56145 | enzyme immunoassay, tracrolimus | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex Reagent Cartridge, manufactured by Siemens Healthcare Diagnostics Inc., Newark, Delaware. | Dimension Tacrolimus Flex Reagent Cartridge | 2 | September 13, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low patient sample recovery results in all affected lots, with possibility of inappropriate increase in tacrolimus dosage based on false low result. | OTHER/UNDETERMINED: Pending | Siemens issued an "Urgent Field Safety Notice" dated June, 2010 to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the TACR Flex reagent lots DC0306, CA0320, DA0334 and GA0348. The letter contains the following instructions: 1) Immediately discontinue use of the affected lots; 2) Discard all remaining inventory of the affected lots; 3) Contact Siemens Healthcare Diagnostics Technical Solutions Center for replacement product; 4) Complete and return the Field Correction Effectiveness Check by fax to 302 631-8467;4) Siemens is recommending that the content of the letter should be discussed with the laboratory director. If unexpectedly low results are obtained on a patient sample, caution should be exercised before decisions regarding therapeutic dosage adjustments are made. As stated in the package insert, "Confirmation of unexpected or atypical results by an alternative methodology is recommended prior to any adjustments in tacrolimus dosage. Also individual tacrolimus values cannot be used as the sole indicator for making changes in the treatment regimen. Each patient should be thoroughly evaluated clinically before treatment adjustments are made."Siemens Technical Solutions Center may be contacted at 800-441-9250. |  | 4580 | 4580 Cartons | Worldwide Distribution to USA states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, KS, LA, MI, MO, MS, NC, NE, NH, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI, and countries of Austria, Belgium, Canary Islands, Egypt, France, Germany, Great Britain, Hungary, Italy, The Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, United Arab Emirates, Argentina, Brazil, Canada, Columbia, El Salvador, Venezuela, Australia, Hong Kong, India, Japan, Malaysia, New Zealand, South Korea, Singapore, Taiwan and China. | Terminated | December 16, 2010 | 94 | 4580 | Not_Computer
Z-2410-2010 | 55876 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen complete knee solution MIS total knee procedure stemmed tibial component fixed bearing precoat size 3, sterile, REF 00-5950-037-01, Zimmer Inc. Warsaw, IN.This device is indicated for patients with severe knee pain and disability due to: Rheumatoid arthritis, osteoarthritis, traumatic arthritis, polyarthritis. Collage disorders, and/or avascular necrosis of the femoral condyle. Post-traumatic loss of joint configuration, particularly when there is patellofemoral erosion, dysfunction or prior patellectomy. Moderate valgus, varus, or flexion deformities. The salvage of previously failed surgical attempts or for a knee in which satisfactory stability in flexion cannot be obtained at the time of surgery. NexGen TM Tibial Trays may be used with or without bone cement (biological fixation). NexGen MIS Tibial Components and MIS Modular Tibial Plates and Keels are intended for cemented use only. | NEXGEN COMPLETE KNEE SOLUTION MIS TIBIAL COMPONENTS, LOCKING SCREW AND STEM EXTENSIONS | 2 | September 13, 2010 | 2010 | Zimmer Inc. | The firm has received complaints of loosening of the implanted device requiring revision surgery. There have been 114 MDRs filed all reported that the device loosened and the patient required additional surgery to replace the device. | OTHER/UNDETERMINED: Pending | The firm, Zimmer, Inc., sent an "URGENT DEVICE CORRECTION" letter dated April 26, 2010, to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to destroy or disregard all previous versions of the surgical technique, review the attached revised surgical technique to familiarize yourself with the modifications and warnings, and complete and return the Health Care Professional Dcoumentation Form via fax to Zimmer, Inc. at (574) 372-4265. (Note: Zimmer made a modification to the surgical technique and instructions for use. They added a warning to fully cement and pressurize the anterior and posterior surfaces of the tibial component, and to strongly recommend the use of a drop down stem extension. This change was completed and approved on April 23, 2010.) For questions please contact Zimmer at 1-877-946-2761. |  | 68384 | 68,384 all products | Worldwide distribution: USA including states of AZ, IL, IN, KY, MO, NE, NM, NY, OH, TX, VA, WA, and WI; and countries of Austria, Belgium, France, Germany, Italy, Netherlands, Spain, Sweden, Switzerland, UK, Ausralia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, Tawain and Thailand. | Terminated | March 28, 2011 | 196 | 68384 | Not_Computer
Z-2422-2010 | 47477 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost X-ray System is a stationary Diagnostic X-ray using Digital Solid State Imaging Detector. It is identified with Part Number 4512 201 04752 Software Version 1.5.1 and Part Number 4512 201 04753 Software Version 1.5.2. The device is manufactured by Philips Medical Systems DMC GmbH, RontgenstraBe 24, 22335 Hamburg, Germany. | DigitalDiagnost Xray system | 2 | September 14, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | When using DICOM print functionality, the printed images may contain data for the wrong patient. | DESIGN: Software Design | Philips Healthcare issued an "URGENT PRODUCT CORRECTION " letter dated March 10, 2008 to their consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: 1) Validate that image and patient information are correct on hard copy prints.2) Reprint images that have mixed patient information, and re-check that image and patient information are correct.3) Do not use printed images if the information is incorrect.The firm will issue the Field Change Order (FCO) 71200025 and the field service engineers will visit each consignee and conduct a software upgrade to solves the printing problem. The consignees may contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 71200025". |  | 119 | 119 units | Nationwide Distribution: including the states of AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, UT, VT, WA, and WI. | Terminated | September 14, 2010 | 0 | 119 | Software | Output/Calculation | Software update | Software Update
Z-2421-2010 | 56422 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Probe Detector PCB Assembly (also called the Aspiration Bottom Sensor) used on the Cell-Dyn Sapphire Hematology Analyzer, Manufactured by Abbott Laboratories, Santa Clara, CA | CELLDYN Sapphire Hematology Analyzer | 2 | September 14, 2010 | 2010 | Abbott Laboratories | Under certain conditions in closed mode of operations, the Aspiration Probe Detector PCB assembly may fail to operate properly and provide inaccurate results. | OTHER/UNDETERMINED: Pending | Abbott Diagnostics Division of Santa Clara, CA initiated a field correction of the CELL-DYN Sapphire Probe Detector Assembly used on the CELL-DYN Sapphire Hemotology Analyzer, on April 5, 2010. The correction is being addressed through field service representatives and a mandatory Technical Service Bulletin (TSB). The TSB requires inspection, testing, and potential replacement of the affected assembly. The firm will not be issuing a seperate notification. For questions regarding this correction please contact (408) 982-4800, or fax (408) 982-4863. | Contact the recalling firm for information | 926 | 926 units | Product was distributed to 926 retailers throughout the US and to argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong King, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | May 17, 2011 | 245 | 926 | Hardware | Output/Calculation | Replace | Remove or Replace
Z-2424-2010 | 56320 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | NeoMarkers Rabbit Monoclonal anti-Human Cyclin D1 Antibody (Clone SP4), immunochemistry Reagent and Kit, For In Vitro Diagnostic use, manufactured by Lab Vision Products, Fremont, CA, a subsidiary of ThermoFisher Scientific | NeoMarkers Rabbit Monoclonal antiHuman Cyclin D1 Antibody (Clone SP4) | 2 | September 15, 2010 | 2010 | Lab Vision Corporation | Two lots may be contaminated with Ki-67 antibody, which could cause abnormal staining patterns, and produce false positive results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Thermo Fisher Scientific issued an "Advisory Notice" letter dated March 5, 2010 to all domestic customers and international distributors. The letter described the product, problem and action to be taken. Customers were instructed to:1) discard the affected lots, and2) retest specimens using a replacement lot supplied by the firm, and 3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489International distributors were instructed to:1) place the affected lots and replace with a replacement lot supplied by the firm, and 2) provide customers with copies of the customer letter, and3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489For questions, contact Thermo Fisher Scientific at 1-800-828-1628 or 1-510-771-1595. |  | 103 | 103 units | Worldwide Distribution: USA, including the states of CA, SC, IL NY, WI, PA, MN, AL, WA, MA, and KS, and the coutnries of Austria, Australia, Denmark, Finland, France, Germany, Hong Kong, Italy, Japan, Malaysia, the Netherlands, Norway, Pakistan, Poland, Portugal, Russia, the UK, and Turkey. | Terminated | March 21, 2011 | 187 | 103 | Not_Computer
Z-2423-2010 | 55763 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc., Dimension Vista Intelligent Laboratory System with Alpha Pinnacle 10kVA UPS External Battery Pack. Part numbers: 070200001, 070200011, 070200021; Dimension Vista 1500 and 3000T instruments. | Dimension Vista System | 2 | September 15, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Defective batteries used in the Alpha Pinnacle P10 UPS External Battery Pack may overheat and cause a strong sulfuric odor to be released and a corrosive sulfuric acid solution leak to occur. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice in May 2010 to customers with information on what to do if the battery leaks, overheats, or vents an odor. Firm also issued a Support Bulletin to field service personnel to inspect and replace affected units. | Contact the recalling firm for information | 228 | 228 units | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Portugal, Slovenia, Spain, and Switzerland. | Terminated | October 25, 2010 | 40 | 228 | Battery
Z-2419-2010 | 54682 | enzymatic method, creatinine | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista ECREA Flex Reagent CartridgeThe ECREA method is an in vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista System. Creatinine measurements are used in the diagnosis and treatment of renal diseases in monitoring renal dialysis, and as a calculation basis for other urine analytes. | Dimension Vista ECREA Flex Reagent Cartridge | 2 | September 16, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | The ECREA test results may be falsely elevated or depressed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated February 2010 to all affected customers. The letter described the product, the problem, and the action to be taken by the customers. The customers were instructed to immediately discontinue the use of and discard the ECREA Flex reagent cartridge lot 09224AA. Customers are to contact their Siemens Customer Service Center at 800-241-0420 to report their additional no-charge replacement needs, and complete and return the enclosed Field Correction Effectiveness Check Form via Fax at 302-631-8467.For any technical questions regarding this Recall contact Siemens Technical Solutions Center at 800-441-9250. |  | 324 | 324 | Worldwide Distribution - USA, including the states of AZ, CA, FL, IA, IL, MT, NC, NY, OR, PA, TX, and WI and countries of Belgium, Italy, Canada, Germany and the Netherlands. | Terminated | October 13, 2010 | 27 | 324 | Not_Computer
Z-2429-2010 | 49489 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Edge Nuclear Magnetic Resonance Imaging System. "1.5T Edge High-Field Magnetic Resonance Imaging System, Picker International, Inc. World Headquarters 595 Miner Road Cleveland, Ohio 44143, U.S.A.". | Edge 1.5T MRI | 2 | September 17, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Water may collect and freeze in the vent pipe. This can rupture the magnet vessel during a cryogen transfer, and release helium into the MRI suite. | OTHER/UNDETERMINED: Pending | Philips issued an Urgent Correction letter dated 09/19/2008 to their consignees, identifying the reason for the recall and the affected products. The firm issued the Field Change Order (FCO) 78100232. Field service engineers will visit each consignee and conduct an inspection of the magnets venting system of Edge, Vista, Eclipse, Polaris, Infinion, and Panorama 0.6T MRI systems. The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 78100232". |  | 28 | 28 units | Nationwide Distribution: USA., including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY. There is no distribution to foreign consignees. | Terminated | September 17, 2010 | 0 | 374 | Not_Computer
Z-2426-2010 | 56564 | bag, stomal | GDS | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Reliamed Odor Eliminator Drops (8 oz, 4 oz, 1 oz package sizes) | Reliamed | 3 | September 17, 2010 | 2010 | BioMed Laboratories | The product may contain the mold Aspergillus ssp. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Biomed Laboratories, Inc. of Garland, TX is conducting a recall of the Reliamed Odor Eliminator Drops, Securi-T Ostomy Deodorant, and the Safe n' Simple (Sn'S) Ostomy Appliance Deodorant. Lots being recalled were distributed between June 9, 2009 though June 28, 2010. Biomed will work with their 3 consignees to inform sub-distributors and down to any firm that keeps the product in stock to have the product removed from further distribution and returned. If there are any questions or concerns, please contact Biomed Laboratories, Inc. at (972) 282-8008 or email them at biomedlabs.net. | Return product to place of purchase for a full refund Return product to place of purchase for a full refund | 33772 | 33,772 Units (27,148 8oz bottles. 3,492 4oz bottles. 3,132 1oz bottles). | OH, FL, MI | Terminated | January 24, 2011 | 129 | 72259 | Not_Computer
Z-2436-2010 | 56626 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | The Gyrus ACMI Pleatman Sac Tissue Removal System facilitates removal of specimens while minimizing contamination of the abdominal cavity or wound tract. The system contains one (1) Sac and one (1) obturator in a sterile, ready -to -use, single- use procedure tray. | Pleatman Sac Tissue Removal System | 2 | September 19, 2010 | 2010 | Gyrus Acmi, Incorporated | Gyrus Medical, Inc. has issued a Urgent Medical Device Recall for the Gyrus ACMI product 004942-903 - Pleatman SacÂż Tissue Removal System, sold as lot 186540FC with an expiration date of 06/2015. As a result of a manufacturing error in the production of the trays used in this lot, it is possible that a crack could form that would compromise the package integrity, and possibly the sterility of th | PRODUCTION CONTROLS: Packaging Process Control | Gyrus Medical, Inc. of Maple Grove, MN issued an Urgent Medical Device Recall drawing attention to a possible package integrity compromise that could affect the product sterility. USA consignees were contacted via telephone on 8/3/10. The firm directed consignees to inspect their stock to confirm if any of the affected lot of product remains in their possession. Use of remaining affected product should cease immediately. Affected products should be return to Gyrus ACMI. For questions or concerns, consignees should contact the firm at (888) 524-7266. | Contact the recalling firm for information Contact the recalling firm for information | 2 | 2 cases (20 devices) | UT, PA, Canada | Terminated | August 23, 2012 | 704 | 2 | Not_Computer
Z-2435-2010 | 54668 | ventilatory effort recorder | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | The Provent Nasal Cannula is a nasal appliance device placed in both nostrils intended to be used for transmission of respiratory airflow signals between the Provent Professional Sleep Apnea Therapy device and physiologic recorders. | Nasal Cannula Diagnostic Kit | 3 | September 19, 2010 | 2010 | Ventus Medical, Inc. | The catalog number printed on two Nasal Cannula Diagnostic Kits (MRS0231) was different that the catalog number printed on the bulk pack shipper box and the other Nasal Cannula Kits in the same shipper (MRS0208). No devices were used on any patient . | PRODUCTION CONTROLS: Labeling Mix-Ups | Ventus Medical, Inc. of Belmont, CA has notified all customers who received distribution from the affected manufacturing lot (1880783) to immediately return the product in question for replacement. If there are any questions regarding this correction, please contact the firm at (650) 632-4165. | Return product to place of purchase for a full refund Return product to place of purchase for a full refund | 60 | 60 | Product was distributed in the US through a single distribution supplier. | Terminated | March 11, 2011 | 173 | 60 | Not_Computer
Z-2437-2010 | 53044 | enzyme immunoassay, gentamicin | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | SYNCHRON Gentamicin (GEN) ReagentGEN reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems Drug Calibrator 3 Plus, is intended for quantitative determination of Gentamicin concentration in | SYNCHRON Gentamicin (GEN) Reagent | 2 | September 20, 2010 | 2010 | Beckman Coulter Inc. | The reagent produces false low test results. Beckman Coulter has confirmed low recovery of values at the low end of the therapeutic range ( | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter dated February 27, 2009 to customers identifying the affected product and issue prompting the recall. The letter asked them to discontinue use and discard the reagent lot, and contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact their local Beckman Coulter Representative to discuss reagent replacement options. They were also asked to complete and return the enclosed response form within 10 days.Customers were directed to call the Customer Support Center (Hotline) at (800) 854-3633 in the US and Canada if they had questions in regards to the PCA letter. Outside the US and Canada they were told to contact their local Beckman Coulter Representative. |  | 1547 | 1,547 Kits | Worldwide Distribution - USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV, and countries of Canada, Belgium, Denmark, France, Germany, Ireland, Lithuania, Netherlands, Spain, Switzerland, United Kingdom, New Zealand, Singapore, Egypt, South Africa, Taiwan, and United Arab Imarets. | Terminated | July 16, 2012 | 665 | 1547 | Not_Computer
Z-2441-2010 | 55163 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Synchron Clinical SystemsClosed Tube Aliquotter (CTA) for Synchron LXi 725, Part Number: 388982, Domestic; 389695, InternationalThe UniCel Synchron Clinical Systems are fully automated, computer controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | Unicel Synchron Clinical Systems | 2 | September 20, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated due to " A loose Aliquot Probe tube fitting in the Closed Tube Aliquotter (CTA), may result in erroneous Access immunoassay results or Synchron chemistry results".This may result in dilution of the sample with Wash Buffer; patient results can potentially be affected for Access immunoassays or Synchron chemistries. Access Quality Control may also be affected if run throug | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated January 18, 2010, to all affected customers. The letter provided the customer with an explanation of the problem identified, the affected product, and instructions for the customer. The letter requested customers complete the following actions for :(1) CTA Aliquot Probe Inspection and Maintenance:a. Check daily for signs of leaks or damage to the tube fitting by looking and feeling for liquid at the fitting, the end of the probe and the surrounding areas;b. The CTA Aliquot Probe tube fitting should be checked to ensure that it is connected tightly to the Aliquot Probe. Without loosening the tube fitting, ensure it is securely tightened;c. When replacing the CTA Aliquot Probe, ensure that the Aliquot Probe tube fitting is securely tightened, prime the CTA, inspect for leaks, and run Access Quality Control through the CTA;d. If a leaking Aliquot Probe is suspected, review all sample results for both Access and Synchron in light of the total clinical presentation of the patient including: symptoms, clinical history, data from additional tests, and other appropriate information;e. As this issue may result in dilution of the sample with Wash Buffer; it may be appropriate to discard and redraw any affected sample tubes run on the CTA. The letter also informed customers that Beckman Coulter is currently evaluating potential solutions for this issue and requested they share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected products to another laboratory, they were asked to provide a copy of this letter to them.They were also instructed to complete and return the enclosed response form within 10 days. Customers were told to call Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada with any questions. |  | 672 | 672 units for both systems | Worldwide Distribution - USA: including the states of AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 13, 2012 | 693 | 672 | Not_Computer
Z-2438-2010 | 55427 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Integrated Data Management System - Version 7.1, 7.2 and 7.3; ITC Corporation, 8 Olsen Place South, Piscataway, NJ 08854 | Integrated Data Management System Version 7.1, 7.2 and 7.3; | 3 | September 20, 2010 | 2010 | International Technidyne Corporation | A tabulation error in the Integrated Data Management Systems (IDMS) data management software that results in incorrect Electronic Quality Control (EQC) Summary Reports: The Operator and the Device Performance Summary Reports. The IDMS software is not used for the clinical management of individual patients. | DESIGN: Software Design | ITC sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 10, 2010 to Distributors and First Consignee End Users identifying the affected product and affected summary reports, the reason for recall, and the actions to be taken by the distributors and customers. New Errata Instructions were enclosed with each letter and are to be put in the idms user guide for reference. The new instructions are to be followed when running reports for the ESV or TVT tests run by the Hemochron Response. The enclosed Customer Account Tracking Form is to be completed and returned to ITC indicating that the letter has been received and appropriate questions answered as to the use of the software system. Distributors are asked to check inventory for recalled product. If affected product is found, shipment should be stopped and placed on hold. Distributors should notify customers of the correction and provide them with a copy of the correction letter and updated instructions. This should be done by traceable mail. Distributors should return the account tracking form and return affected software disks for replacement.Questions should be directed to ITC Technical Support at 800-631-5945. |  | 84 | 84 total idms software systems; 66 in US, 18 outside of US. | Worldwide distribution: USA including the states of AL, AZ, CA, FL, GA, IL, KS, LA, MD, MI, NC, ND, NJ, NM, NY, OH, PA, TN , TX, UT, VA, WA, and WI, and the countries of Italy, Turkey, Canada, France, Belgium, Germany, UAE, and New Zealand. | Terminated | May 25, 2011 | 247 | 84 | Software | Output/Calculation | Replace software disks | Software Update
Z-2439-2010 | 56390 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TRABECULAR METAL ÂŽ FEMORAL CONE AUGMENT - LEFT, Part Number: 00-5450-012-31.The objective of the Trabecular Metal Femoral Cone Augment is to fill and reconstruct large bone deficiencies and cavitary defects in the distal femur and to provide for a stable platform for an LCCK or Rotating Hinge Knee femoral component. | Zimmer Trabecular Metal" Femoral Cone Augment | 2 | September 20, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for the packages to be mislabeled. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. | PRODUCTION CONTROLS: Labeling Mix-Ups | Zimmer, Inc. sent Urgent: Device Removal letters, dated July 9, 2010, identifying the affected product and stating the issue. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. The letter also discussed the clinical implications along with required actions. Customers are to perform a physical count of all affected product and record the data on the Inventory Return Certification Form included with the letter. In addition, customers should complete and return the Inventory Certification and User Facility/HCP Documentation Forms. Affected product along with the Inventory Return Certification Form should be returned. Zimmer, Inc. should be contacted at 1-800-613-6131 if there are any questions or need for additional information. |  | 50 | 50 total | Worldwide Distribution -- USA, Canada, Singapore, Australia, and Switzerland, Spain, UK, Sweden, and Israel. | Terminated | December 22, 2010 | 93 | 50 | Not_Computer
Z-2446-2010 | 56183 | electrode, needle, diagnostic electromyograph | IKT | Diagnostic electromyograph needle electrode. | Physical Medicine | Neurology | 510(k) | Disposable Detachable Monopolar Needle Electrode for electromyograph model number 110-725, for Physician use, manufactured by Chalgren Enterprises Inc., Gilroy, CA. | Disposable Detachable Monopolar Needle Electrode for EMG | 2 | September 21, 2010 | 2010 | Chalgren Enterprises Inc | Sterility Compromised-- Investigation found that individual packaging seals were separating, which may compromise sterility. | PRODUCTION CONTROLS: Packaging | Chalgren Enterprises, Inc. removed all of the affected inventory from inventory and had all affected inventory retrieved from distribution. |  |  |  | Nationwide Distribution: USA states of CA, IL, NJ, OH and TX. | Terminated | March 24, 2011 | 184 |  | Not_Computer
Z-2445-2010 | 56460 | lenses, soft contact, extended wear | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Primary (inner) Package Label: ACUVUE OASYS BRAND CONTACT LENSES for ASTIGMATISM with HYDRACLEAR Plus (senofilcon A) Johnson & Johnson Vision Care, Inc. Rx Only, STERILE. Lot B006WBS1 -9.00D -1.25 CYL 80 , 2013/05 AXIS. (Correct labeling) Mislabeled Outer (secondary) Label: Lot B006X641D -8.50 CYL -1.25 AXIS 100, 2013/05 DO NO PAY Free Sample Not For Sale Consumer Note: This product has been provided free of charge to your Eye Care professional to aid in proper fitting. The sale of this product is prohibited. You should not have been charged money for this product. Contents: 1 lens (38% H2) in buffered saline with up to 0.005% methyl ether cellulose. | ACUVUE OASYS BRAND CONTACT LENSES for ASTIGMATISM with HYDRACLEAR Plus (senofilcon A) | 3 | September 21, 2010 | 2010 | Vistakon | ACUVUE OASYS Brand Contact Lenses for ASTIGMATISM may be mislabeled on the outer (secondary) packaging carton. The outer (secondary) cartons labeled as Lot B006X641 -8.50D -1.25cyl. 100 Axis may actually contain primary packages from Lot B006WBS1 -9.00D -1.25cyl. 80 Axis. The inner (primary) package labeling is correct. Product was only distributed to Eye Care Professionals for use in proper f | PRODUCTION CONTROLS: Error in Labeling | Vistakon issued an Important Product Notification dated July 28, 2010 to customers identifying the affected product, the labeling issue, and actions to be taken by the customer. Vistakon issued a separate Important Product Notification dated July 28, 2010 to distribuitors identifying the affected product, the labeling issue, and actions to be taken by the distributor. Customers are instructed to check their inventory for affected products and return a business reply card. If replacement -8.50 lenses are needed, the customer can contact Customer Service at 1-800-874-5178.Distributors are instructed to check their inventory for affected products and return them using an enclosed return label. Replacement product will be sent by the firm. Customer Service can be contacted at 1-800-874-5178.Accounts will be instructed to go to their inventory and look for the B006X641, -8.50 lot and open the secondary carton and verify the product inside. If it is the -9.00 product they can keep the product and re-slot it in the proper place within the kit (product is ok, nothing wrong with primary package or lens inside) and order a replacement -8.50 if necessary. Business reply cards will be supplied for accounts to return to third party indicating if they had any mix/mislabeled product. | Contact the recalling firm for information | 2123 | 2,123 lenses | Worldwide Distribution: USA including all states, District of Columbia and Puerto Rico, and the countries of Canada, Bahamas, Honduras, and United Arab Emirates. | Terminated | March 28, 2011 | 188 | 2123 | Not_Computer
Z-2447-2010 | 56327 | patch, pledget and intracardiac, petp, ptfe, polypropylene | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) products. Sterile, Rx only, SynovisÂŽ Surgical Innovations. Intended to be used as a staple line buttress. | PeriStrips Dry Staple Line Reinforcement with Apex Processing (PSD) and PeriStrips Dry with Verita | 2 | September 21, 2010 | 2010 | Synovis Surgical Inovation Div. of | SynovisÂż Surgical Innovations is voluntarily recalling PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple line Reinforcement (PSD-V) products. Internal investigation have shown there is a potential for small defects in the PSD gel tube pouch which may compromise the sterility of the outside of t | PRODUCTION CONTROLS: Packaging | Consignees were sent on 6/30/10 a Recall Notification letter dated June 30, 2010. The letter was addressed to "Risk Manager". The letter described the problem and product involved. They listed both the device and gel lot numbers since it is possible that some hospitals have separated the pouched gel tubes from the PSD/PSDV products. They advised consignees to quarantine the affected lots and to complete and fax the attached "Product Return/Replacement" form to Synovis Surgical Innovations. |  | 5384 | 3,989 gel tubes and 1,395 PSD & PSDV devices. | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL,IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI and countries of AUSTRALIA, GERMANY, GREECE, HONG KONG, NETHERLANDS, PUERTO RICO, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | December 17, 2011 | 452 | 5384 | Not_Computer
Z-2443-2010 | 56121 | pulse-generator, pacemaker, external | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Oscor PACE 203H, Dual-Chamber External Pulse GeneratorThe PACE 203H / Model 3085 external pulse generator / temporary cardiac pacemaker is designed to be used with cardiac pacing lead systems for temporary atrial or ventricular or A-V sequential pacing. The PACE 203H / Model 3085 has applications where such pacing modes are indicated for therapeutic, prophylactic, or diagnostic purposes. | Oscor PACE 203H, DualChamber External Pulse Generator | 2 | September 21, 2010 | 2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e., not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t |  | OSYPKA Medical sent "Urgent Medical Device Recall" letters dated June 30, 2010 and July 7, 2010 to the direct customers. The recall notice informed the customers of the Reason for Recall, the Potential Hazard, the Temporary Fix of the Problem, the Permanent Fix of the Problem and Recommended Actions to be taken by the customer. The customers were instructed to contact Osypka Medical via telephone (858) 454-0021 or e-mail to m.osypka@osypkamed.com to verify that the Model 3085 and/or PACE 203H is subject to this product removal, obtain a return-to-manufacturer-authorization (RMA) number ,and then return the affected devices to Osypka Medical.If you have any questions call (858) 454-0021 or e-mail to mail@osypkamed.com. |  | 205 | 205 units (US Only) | Nationwide Distribution: USA including states of CA & FL | Terminated | August 10, 2012 | 689 | 217 | Hardware | Device Operation | Return | Remove or Replace
Z-2461-2010 | 49594 | coil, magnetic resonance, specialty | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Nuclear Magnetic Resonance Imaging System (MRI). Trace item 9896-030-14351 -1.5T Sense XL Torso coil 12nc on coil: 4535-671-4188x. | Intera 1.5T and 1.5T Achieva MRI systems | 2 | September 22, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The cables of the coil may become too hot and cause burns when placed too close to the patient. While using the 1.5T Sense XL Torso coil, unwanted radiofrequencies (RF) interaction with the patient can occur. Especially in combination with high SAR levels and when the RF cables of the coil are too close to the patient, this RF interaction has resulted in RF burns to patients. | OTHER/UNDETERMINED: Pending | Philips Healthcare issued URGENT-Device Correction Notices dated September 30, 2008 to their consignees. The notice identified the recalled device as 1.5 Intera and 1.5T Achieva MRI systems, described the risk to patients, and the actions to be taken by customers..The firm has issued Field Change Order (FCO) 78100242. Field service representatives will visit each consignee and conduct a hardware upgrade, applying a sleeve over the cable that will ensure proper spacing between the cable and patient, reducing the chance of RF interaction and burns.The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #3, #1 and reference "FCO 78100242 - PATIENT CABLE SLEEVE FOR 1.5T SENSE XL TORSO COIL". |  | 47 | 47 units total (Achieva 1.5T and Intera 1.5T) | Nationwide Distribution: Throughout USA. No distribution to foreign consignees. | Terminated | September 30, 2010 | 8 | 47 | Not_Computer
Z-2464-2010 | 50066 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Endura Mobile X-Ray system with the extended rotation option. | BV Endura Rel 2 | 2 | September 22, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The C-arm rotation brake does not always function properly, and may not hold the C-arm position if the C-arm is not positioned in its maximum rotation and if it is accidentally bumped by medical staff or patients. Unwanted C-arm movement may cause the C-arm to collide with a patient. The brakes function properly if the C-arm is rotated to its maximum positions. | OTHER/UNDETERMINED: Pending | Philips Medical Systems issued an Urgent Product Correction letter dated September 30, 2008 to consignees. The letter identified the affected device, what the problem is and under what circumstances it can occur, and the actions that should be taken by the customer or user in order to prevent risks for patients or users. Philips will contacdt customers to implement Field Corrective Action (FCO) #71800026), which describes a mechanical brake modification which will be installed by Philips. Until the Field Corrective Action has been completed, users should ensure that the C-arm is placed in the maximum extended position, or ensure that the C-arm is not touched during critical procedures if the C-arm is not in a maximum position.The Philips Call Center can be contacted at 800-722-9377. |  | 71 | 71 units | Nationwide Distribution: Throughout the US. | Terminated | January 07, 2011 | 107 | 1420 | Not_Computer
Z-2471-2010 | 50114 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura XPER 10/10 Angiographic X-ray system. This system with Velara Generator can be identified as type number 9890 000 6501 * (number located on the generator cabinet).Intended use for Angiographic X-ray system with and without a solid state X-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | Allura XPER 10/10 | 2 | September 22, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A high voltage reliability issue resulting in the loss of the x-ray generation and this can cause an intermittent stop of fluoro acquisition and exposure during an examination without warning and hence no image on the display. | OTHER/UNDETERMINED: Pending | The firm, Philips, sent "URGENT-Device Correction" letters dated November 10, 2008, to their customers and informed them about the reliability issue in the Allura XPER and Integris X-ray systems with the High Voltage Velera (HV) Generator. This reliability issue can lead to a system error causing intermittent loss of fluro and/or exposure acquisition, potentially resulting in the loss of the X-ray image on the display.The customers are advised if they experience High Voltage Generator reliability issue and the HV Generator stops working generating an O3EA error, a cold restart is required and reboot the system when the X-Ray image disappears.The firm will implement multiple mandatory Field Change Orders (FCOs) to solve the Generator reliability by updating the software for each X-ray system.The customers could contact Philips Care Center at 800-722-9377, #5, #4, #1 for any questions regarding this recall or 425-487-7469. |  | 70 | 70 units in US | Worldwide distribution. | Terminated | August 10, 2011 | 322 | 1681 | Computer
Z-2465-2010 | 56179 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Spider Flex Monopolar Hook (Spider Surgical Instruments), TransEnterix, Inc., Durham, NC | Spider Flex Monopolar Hook | 2 | September 22, 2010 | 2010 | Transenterix, INC | The hook end effecter from a monopolar device may become dislodged and fall into the patient when in use during laproscopic surgery. Retrieval from the patient may involve extended anesthesia time. | OTHER/UNDETERMINED: Pending | TransEnterix issued a "Field Action" letter dated June 22, 2010 to consignees, identifying the affected product and actions to be taken by the customer. Customers were instructed to return the product to TransEnterix, Inc. using the address label and return information provided. The firm will replace the product at no charge. TransEnterix Customer Care can be contacted at 1-888-879-4111. |  | 46 | 46 boxes (6 units per box) | Nationwide Distribution in the states of AZ, FL, GA, NC, NY, OH, OR, SC and TX. | Terminated | January 31, 2011 | 131 | 46 | Not_Computer
Z-2449-2010 | 56093 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | BIPLANE X-ray Imaging Systems included Allura XPER FD10/10, FD20/10, FD20/20, Integris BH5000, and Integris Allura 12-15 Biplane. The BIPLANE system has two C-arms and it allows the operator to image in two planes: Lateral and Frontal. It has a C-arm on the stand and C-arm suspended from the ceiling.The product's brochure is labeled in part: "Philips Healthcare Global Information Center...P.O. Box 1286...5602 BG Eindhoven...The Netherlands". | BIPLANE XRay Imaging Systems with Velara Generator | 2 | September 22, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A lock-up might occur inside the Velara Generator due to intermittent shorts inside the system. The problem can only resolved by a warm or cold system restart. In exceptional cases the power converters of the KVmA unit in the generator are damaged and the system can not be restarted. | DESIGN: Software Design | On 02/04/2010, Philips Healthcare began sending out the URGENT-Device Correction Notices to their consignees. The firm issued two notifications, one for the Allura Integris and Xper FD10C, FD10F, FD10F/C and FD20 Monoplane X-Ray imaging systems with the Velara Generator. The second notification is for the Allura Integris Xper FD10 and FD20 Biplane X-ray imaging systems with the Velara Generator.This notice informs the consignees that a system locks up during a critical interventional case and it might impose a serious injury to patients. This lock up issue could only be resolved by a warm or cold system restart. A warm restart would take about 90 seconds or a cold restart would take about four minutes, and the device could not be used during this time. If the system does not work after the restart, the consignees should contact the Philips Healthcare Customer Service Representative at 1-800-772-9377 #5 (then follow prompts). A service representative will visit each consignee and conduct a software and hardware upgrade. | Contact the recalling firm for information Contact the recalling firm for information | 267 | 267 | Nationwide Distribution | Terminated | September 23, 2011 | 366 | 1737 | Hardware | Device Operation | Software and Hardware upgrade | Software Update
Z-2425-2010 | 56190 | instrument, surgical, non-powered | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Integra NeuroSciences Implants, NeuroBalloon Catheter,Sterile/EO, Rx only, single use, latex free: Catalog Number 7CBD10 | NeuroBalloon | 1 | September 22, 2010 | 2010 | Integra LifeSciences Corp | Certain NeuroBalloon Catheters have the potential to improperly inflate or deflate under certain conditions. | DESIGN: Process Design | All customers and Integra Sales Specialists were notified of the recall on July 2, 2010 by fax, e-mail and letter send by Fed Ex overnight delivery. |  | 338 | 338 units | Nationwide, Canada, Taiwan and to Integra, Australia. | Terminated | May 29, 2012 | 615 | 338 | Not_Computer
Z-2466-2010 | 56303 | tampon, menstrual, unscented | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons | 2 | September 22, 2010 | 2010 | First Quality Hygienic Inc | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons package contained Super Plus Unscented Plastic Applicator Tampons. | PRODUCTION CONTROLS: Labeling Mix-Ups | Consignees were notified via telephone and letter of the need to withdraw product from the market and requested to return all product to manufacturer. | Contact the recalling firm for information | 195 | 195 Cases/2340 Cartons | Canada | Terminated | October 06, 2010 | 14 | 195 | Not_Computer
Z-2460-2010 | 56506 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Heavyweight Gantry Scanner, M/N 728306, S/N: 100016, 100018 -100025The Brilliance iCT is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance iCT Heavyweight Gantry Scanner, M/N 728306, S/N: 100016, 100018 100025 | 2 | September 22, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | The gantry covers of the iCT have a latch at the base of the system which holds the cover in the closed position. During the course of maintenance the CT Field Service Engineer (FSE) will activate this latching mechanism which in turn has a potential to penetrate the insulation on cabling which carries 277 VAC. If the latch does come into contact with the cabling there is a potential that the FS | OTHER/UNDETERMINED: Pending | Philips Healthcare sent an URGENT - Field Safety Notice letter, dated April 14, 2010, to their customers notifying them of the product, the problem, and the action to be taken. The letter asked them to identify the affected product by checking for the serial number tag (s/n 10016  10025), ensure that the electrical power was shut-down before opening or closing the gantry cover to prevent potential of an electrical shock, and that a Field Service Engineer would contact the facility to make arrangements for the implementation of a Field Change Order (FCO) before December to correct the problem. For further information or support concerning this issue, the customer was instructed to contact their Philips representative: Customer Care Center (1-800-722-9377, option 5 or their local Philips Healthcare Office). |  | 9 | 9 units | Worldwide Distribution - USA, including the states: DC, GA, IL, IN, NY and OH; and the countries of Germany and the Netherlands. | Terminated | September 24, 2012 | 733 | 9 | Not_Computer
Z-2450-2010 | 56563 | tubing, noninvasive | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Medi-Vac Non-Conductive Suction Tube with Maxi-Grip Connectors, Inner Diameter: 5 mm (3/16"), Length: 1.8 m (6 ft.); Sterile, for single use only, 50 tubes per case; Cardinal Health, McGaw Park, IL 60085 USA, Made in Mexico; catalog N56A and NEX56A (European label) | MediVac NonConductive Suction Tube | 2 | September 22, 2010 | 2010 | Cardinal Health | The packaging seals may not be intact, resulting in compromised sterility for the listed Medi-Vac Non-Conductive Tubing. | OTHER/UNDETERMINED: Pending | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated August 16, 2010 to consignees, identifying the reason for the recall, the affected lots, and actions to be taken. Cardinal has inspected all packages in their warehouses and removed the affected products. All packages that have been inspected have a 3/4" green dot on the front of each unit and on the case carton label." Cardinal Health end users were instructed to segregate and destroy their affected product and request credit or replacement" Distributor customers were instructed to destroy their product and request credit or replacement" Distributors were instructed to conduct sub-recalls and return their inventory to Cardinal Health" To arrange for credit or replacement of the product destroyed, the accounts were instructed to call customer service at 800-964-5227 or 888-444-5440." All customers were requested to complete the enclosed Recall Acknowledgment Form acknowledging receipt of the letter and the number of tubing units destroyed, and fax it to 847-689-9101.Cardinal Health can be contacted at 800-292-9332.Cardinal sent recall letters dated 8/26/10 to their PreSource kit/module customers via UPS next day air, informing them that the Medi-Vac Non-conductive Suction Tubing listed may have packaging seals that are compromised. The accounts were informed that Cardinal has inspected tubing packages in their warehouses and removed the affected products. All packages that have been inspected will have a 3/4" green dot on the front of each package. Enclosed with the recall letter were recall labels to affix to each affected Presource pack/module in the account's inventory. These labels instruct the user NOT to use the affected tubing component and to remove and discard the component upon opening the kit. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of packs/modules on hand. Distributo. |  |  |  | Worldwide Distribution, including USA. | Terminated | November 28, 2011 | 432 |  | Not_Computer
Z-2470-2010 | 56667 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DRIVER STAR #8 CANNULATED DART-FIRE (TM), COMPRESSION SCREW, SIZE: #8, REF: DSDS-0008, 1 EACH, Rx ONLY, NON-STERILE, EC REP Wright Medical EMEA, Krijgsman 11, 1186 DM Amstelveen, The Netherlands, Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 | DARTFIRE CANNULATED DRIVER STAR 8 | 2 | September 22, 2010 | 2010 | Wright Medical Technology Inc | After autoclaving, the epoxy identifier (colored) band bled and there is concern that it may flake. |  | The firm initiated their recall the product by letter on 08/26/2010. The letters were addressed to distributors and hospital administrators with a fax back response form. Please contact Customer Service at 800-238-7117, if you have any questions. |  | 13 | 13 units | TX, PA, TN, VA, GA, and Spain | Terminated | August 02, 2012 | 680 | 13 | Not_Computer
Z-2467-2010 | 56686 | elevator, wheelchair, portable | ING | Wheelchair elevator. | Physical Medicine | Physical Medicine | 510(k) | Braun wheelchair lift, model NL, Braun Corporation, Winamac, INA wheelchair lift for motor vehicles. | Braun wheelchair lift, model NL, | 2 | September 22, 2010 | 2010 | The Braun Corporation | The wheelchair lifts threshold warning system may not operate and the wheelchair user can fall from the vehicle and be seriously injured. | DESIGN: Device Design | The Braun Corporation issued a Service Bulletin 34273 entitled "Lift Threshold Sensitivity Bulletin 34273 on October 2007. Braun Corporation also sent a "Safety Recall" letter to the Braun lift owner and letters to all customers on or about August 15, 2007. The letters described the product, problem and the action to be taken by the customer. The customers were instructed to notify the National Highway Traffic Safety Administration (NHTSA) and conduct a safety recall if they used recalled lifts as original equipment in a system or vehicle that they built; ensure that any associated customers are notified to discontinue the use of their vehicles until they schedule a repair, and verify the threshold warning system is within the NHTSA specification. The customers were informed to contact the NHTSA with questions by calling the Vehicle Safety Hotline at 1-888-327-4236 or go to http//www.safercar.gov. If they have any problem obtaining the needed repair or wish to schedule service with an approved Braun technician, contact the Braun Corporation Product Support at 1-800-946-7513. | Contact the recalling firm for information | 8356 | 8,356 | Worldwide distribution | Terminated | May 12, 2011 | 232 | 12940 | Not_Computer
Z-2482-2010 | 55360 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECTÂŽ i1000SR" Processing Module System For In Vitro Diagnostic Use, Abbott Laboratories Abbott Park, ILProduct List Number, Part # or Model: 01L86-01In Vitro Diagnostic use only. The Abbott ARCHITECT System is designed to perform automated: Chemistry tests, utilizing photometry and potentiometric technology, and Immunoassay tests, utilizing CMIA (chemiluminescent microparticle assay) detection technology. | ARCHITECT i1000SR" Processing Module | 2 | September 23, 2010 | 2010 | Abbott Laboratories, Inc | During manufacturing, two components of the ARCHITECT i1000SR were reversed which could potentially cause falsely elevated patient results due to a sample-reagent and reagent-sample carryover. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | ABBOTT issued Technical Service Bulletins to Field Service Representatives who then inspected every unit . | Contact the recalling firm for information | 404 | 404 units distributed worldwide (including US) | Worldwide distribution:USA including states of: CO, DC, FL, GA, IL, LA, MA, MN, NY, OH, OK, TX, UT, and WV and countries including: Australia, Brazil, Bogota, Canada, China WFOTC, Columbia, Curacao (NETH ANT), Germany, Hong Kong, Japan, Mexico, Singapore, and Taiwan. | Terminated | April 13, 2012 | 568 | 404 | Not_Computer
Z-2480-2010 | 55911 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, Catalog number 801763, 100/120V System 1 base, Terumo Cardiovascular System, Ann Arbor MI. | Terumo Advanced Perfusion System 1 | 2 | September 23, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The design document states that the oxygen sensor would operate for 300,000 hours but the supplier of the oxygen sensor indicated that it has a 100,000 hour warranty. The oxygen sensor is a consumable part that slowly depletes over time. At the end of its service life, a depleted oxygen sensor results in a gas system calibration failure or a message to the user prompting service for the gas syst | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo issued a URGENT MEDICAL DEVICE CORRECTION letter dated September 10, 2010, identifying the affected product, the issue prompting the correction, and what actions are to be taken by the customer. The firm included the addendum to the operators manual with the recall letter. The addendum recommended replacement of the oxygen sensor every six months instead of every 12 months. The addendum provided additional user instructions in the event of a calibration failure of the sensor. Customers are to complete and return the enclosed Customer Response Form to the firm. Terumo CVS Customer Service can be contacted by 1-800-521-2818 Monday through Friday, 8 AM  6 PM EST. |  | 1340 | 1340 | Worldwide Distribution: USA, including the states of AL,AZ, CA, CO, CT, DC, FL, GA, HI, IA, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WI, and WV, and the countries of Canada, Indonesia, China, Chile, UAE, Japan, Singapore, UK, Malaysia, Taiwan, Belgium, Korea, Phiippines and Thailand. | Terminated | December 20, 2011 | 453 | 2655 | Not_Computer
Z-2478-2010 | 56676 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Anesthesia Machine, OR/Anesthesia, Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969inhalation anesthesia machine for use in operating, induction and recovery rooms | Fabius GS | 2 | September 23, 2010 | 2010 | Draeger Medical Systems, Inc. | potential failures to Electri-Cord Manufacturing AC power cord | OTHER/UNDETERMINED: Pending | Draeger Medical, Inc., sent an Urgent - Medical Device Recall letter dated August 2010 to all customers. The letter described the product, the problem, and actions to be taken by the customer. Customers were instructed to:1) Immediately take the cord out of service if the plug has bent or cracked prongs, burned plastic, or shows signs of excessive wear.2) Monitor the affected cords regularly and be mindful of excessive wear and tear, misuse, or abuse until the affected cord is replaced. 3) Customers were also asked to complete and return the enclosed Customer Power Cord Form via mail: Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969; fax: 1-215-721-5808; or e-mail: kathy.benedetto@draeger.com.4) Disable the cords by cutting off the plug and dispose of the cord in accordance with their facility's disposal procedure when affected power cords are removed.NOTE: The firm will send replacement power cords free of charge. Draeger will only replace power cords originally provided by Draeger (identified with the white flag/part number 4117266.If you have any questions regarding this letter please call 1-800-543-5047 (press 1 at the first prompt, 2 at the second prompt and then 32349#). |  | 1973 | up to 1973 products (GS & Tiro) | Worldwide Distribution: USA and countries of Brazil, Canada, Columbia, Costa Rica, Cuba, Ecuador, Japan, Lebanon, Nicaragua, Mexico, Panama, Philippines, Saudi Arabia, Taiwan, and Venezuela | Terminated | August 23, 2011 | 334 | 1973 | Battery |  | Replace | Remove or Replace
Z-2462-2010 | 56560 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GCX Mountable Downloader-Recharger;(accessory to the i-STAT Portable Clinical Analyzer (model 300).The GCX Mountable Downloader-Recharger (GCX) is a connectivity component that is intended to be mounted to a GCX compatible mount or "arm" and provide a data connection to an authorized manufacturer's patient monitoring systems. | GCX Mountable DownloaderRecharger; | 3 | September 23, 2010 | 2010 | Abbott Point of Care Inc. | Reliability issues associated with broken connection pins within the downloader that mate with the pins from the analyzer. | DESIGN: Device Design | Abbott Point of Care, Inc. sent Urgent Recall Notices, dated August 20, 2010, by Federal Express to two Clarion Health locations.The letter identified the product, the issue, and the actions to be taken by the firm and the customers. 1) An Abbott Point of Care representative will contact the customers to plan the replacement of affected product at their facility.2) Replacement will occur in batches of approximately 50 units at a time.3) Abbott will provide personnel to perform the replacements to help expedite the replacement.4) The customer is required to acknowledge that all units have been replaced by signing the form provided with the Urgent Recall Notice.. If you have any questions regarding this information please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. |  | 277 | 277 | Nationwide Distribution - USA, including the state of Indiana | Terminated | July 25, 2012 | 671 | 277 | Not_Computer
Z-2484-2010 | 56498 | pump, breast, powered | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile double breast personal accessory convenience kit for the Symphony or Lactina Breast pumps; the kit connects to a hospital grade breast pump for the removal, collection, and storage of mother's breast milk; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; article #67116S | Medela Symphony & Lactina Double Pumping Kit (Sterile) | 2 | September 24, 2010 | 2010 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | PRODUCTION CONTROLS: Packaging | Medela sent Important Device Recall letters dated August 6, 2010 to all direct accounts, identifying the affected product, the labeling issue, and actions to be taken by the customer. The accounts were informed that the sterility of the affected products may be compromised due to holes in the packaging. They were instructed to stop use, sale and distribution of the kits and return their stocks to Stericycle using the enclosed packing slip and prepaid UPS Return Service shipping label. The accounts were requested to complete and return the enclosed postage paid reply card, indicating the amount of product being returned. Distributors were requested to notify their customers. Questions about the recall process may be directed to Stericycle at 1-877-247-9970. Information on the recall is posted on Medela's website www.medelabreastfeedingus.com. |  | 2030390 | 2,030,390 kits | Worldwide Distribution: Nationwide throughout USA, including Guam, and to the countries of Barbados, Bermuda, British West Indies, Canada, Panama, Switzerland, Trinidad and Tobago, and Uganda. | Terminated | September 07, 2011 | 348 | 6924825 | Not_Computer
Z-2487-2010 | 56423 | [6 empty] | Ear Candles made from either Paraffin or Beeswax, in plain, herbal scented, or lavender scented varieties, manufactured by Wally's Natural Products, Auburn, CA. | Wally's Ear Candles | 2 | September 24, 2010 | 2010 | Wally's Natural Products | Product is an unapproved medical device. Use as directed may result in accidental damage to the eardrum, leading to injuries up to loss of hearing. | OTHER/UNDETERMINED: Pending | The firm conducted communications via phone, mail and website, stating that they have stopped manufacturing and distributing ear candles, and asks consumers to return for a refund. |  |  | Not Specified. All | Nationwide Distribution | Terminated | March 28, 2011 | 185 |  | Not_Computer
Z-2501-2010 | 54745 | kit, test, pregnancy, hcg, over the counter | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UCG Beta Slide Monoclonal II is a latex agglutination inhibition slide test for the detection of hCG in urine. This test is used for the diagnosis of pregnancy. | UCG BetaSlide Monoclonal II | 3 | September 25, 2010 | 2010 | Inverness Medical Professional Diagnostics | Specific lots no longer meet their stated limit of detection claims of 2000 mIU/mL for the UCG-Slide test and 500 mIU/mL for the UCG-Beta Slide Monoclonal II test. | OTHER/UNDETERMINED: Pending | The recall was initiated on February 12, 2010. Letters were sent by Fed Ex on 2/12/2010. | Contact the recalling firm for information Contact the recalling firm for information | 585 | 585 kits | Distributed Nationwide and to Australia | Terminated | May 14, 2012 | 597 | 1637 | Not_Computer
Z-2503-2010 | 56441 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | GATOR. REF 9263A, 4.2mm STERLING, STERILE Rx Only. | ConMed Linvatec 9263A 4.2mm Sterling, Blade, Gator Meniscus Cutter | 2 | September 25, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling ConMed Linvatec 9263A,4.2mm Sterling Blade, Gator Meniscus Cutter, Lot Number 178026 for breech in the sterile barrier. | PRODUCTION CONTROLS: Packaging Process Control | Recall Notifications will be sent to all direct consignees by FedEx, email, Fax, etc. If the account no longer has the product in their possession, they will be asked to respond as such. A Return Material Authorizations Number will be assigned for the returned product as required. Product returned will be placed into quarantine and will be dispositioned in accordance with ConMed Linvatec procedures. 100% of the consignees will be contacted. The reply form is the mechanism to verify notification effectiveness to all consignees. If a reply is not received within approximately 30 business days of the first notification ConMed Linvatec will make two additional contacts by phone and / or email, as necessary, to insure all consignees have received notification and have taken appropriate action. If a reply is not received following the third contact by email or phone, no further contacts will be made to the consignee. | Discard product or return to place of purchase for a full refund or exchange Discard product or return to place of purchase for a full refund or exchange | 174 | 174 | Products were distributed to ConMed Linvatec consignees, sales representatives and affiliates within the US and Australia. US distribution was to WA, AZ, PA, KS, NM, UT, ID, IL. Products were distributed to ConMed Linvatec Australia. | Terminated | December 14, 2010 | 80 | 174 | Not_Computer
Z-2499-2010 | 56406 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody Thyroid Transcription Factor-1; 1 mL liquid tissue culture supernatant in a 1 mL vial; an in vitro diagnostic reagent for the qualitative identification by light microscopy of thyroid transcription factor-1 molecules in paraffin-embedded tissue by immunohistochemical staining; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom; Catalog No.: NCL-L-TTF-1 | Novocastra Thyroid Transcription Factor1 (liquid) | 2 | September 25, 2010 | 2010 | Leica Microsystems, Inc. | In combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the instructions for use. | DESIGN: Device Design | Leica Microsystems sent Medical Device Field Correction letters dated 7/28/10 to the affected customers via first class mail on the same date, informing them that in combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the Instructions for Use. Currently they state that lung and thyroid tumors may stain, but Leicas investigation further characterized the expression of the TTF-1 antigen and it may also be present in other tumors, e.g. thymomas. Customers who have received the affected lot numbers of the three products were requested to refer to the new Instructions for Use to review the amendments made to the Results Expected via the following links: http://bsd.leica-microsystems.com/pdfs/products/ttf-1-ce.pdf; http://bsd.leica-microsystems.com/pdfs/products/ttf-1-l-ce.pdf; and http://bsd.leica-microsystems.com/pdfs/products/pa0364.pdf; The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. | Contact the recalling firm for information Contact the recalling firm for information | 860 | 860 units Total for all three products | Nationwide and Canada | Terminated | January 29, 2012 | 491 | 860 | Not_Computer
Z-2502-2010 | 56418 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GreenLight HPS fiber optics, Model Number 10-2090, manufactured by AMS (American Medical Systems), San Jose, CA. Potentially Mislabeled as GreenLight PV 10-2079 | GreenLight HPS Fiber Optics, Model Number 102090 | 2 | September 25, 2010 | 2010 | AMS Innovative Center - San Jose | Product may be packed with incorrect external or internal labels and/or missing product insert. | PRODUCTION CONTROLS: Labeling Mix-Ups | Notification issued via letter with instructions and response sheet through certified mail on 01/12/2010. | Contact the recalling firm for information Contact the recalling firm for information | 11668 | 11668 units | Product was distributed to 148 medical facilities throughout the US. and 6 consignees in Great Britain, Canada and Australia. | Terminated | January 25, 2011 | 122 | 11668 | Not_Computer
Z-2495-2010 | 56363 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Femoral Artery Catheterization Set | Femoral Artery Catheterization Set | 2 | September 25, 2010 | 2010 | Arrow International Inc | Incorrectly assembled 18 Ga x 2 1/2 Introducer Needle | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The recalling firm issued an Urgent Medical Device Recall letter dated 7/23/2010 along with a a recall acknowledgement and stock status form. | Contact the recalling firm for information Contact the recalling firm for information | 19 | 19 | The product was shipped to medical centers and hospitals nationwide and to Australia, New Zealand, and Europe. | Terminated | July 05, 2012 | 649 | 113 | Not_Computer
Z-2504-2010 | 56591 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Modular Microplasty Cup inserter 3/8" thread , REF 31-400600 | Modular Microplasty Cup inserter 3/8" thread | 2 | September 25, 2010 | 2010 | Biomet, Inc. | There was a clerical error in order entry which resulted in the recalling firm recieving spring components from a supplier made of "music wire" The recalling frim conducted tests to determine the effects of theis change in materials. The testing showed that the wire would oxidize and there would be early failure of the device. | TRAINING: Employee Error | The firm send URGENT MEDICAL DEVICE RECALL NOTICES to its distributors on 7/28/2010. The notices instructed the distributors to discontinue use of the product and return it to Biomet. | Contact the recalling firm for information Contact the recalling firm for information | 331 | 331 total for both products | US, Australia, Canada, Mexico, Chile, Japan. | Terminated | October 01, 2012 | 737 | 331 | Not_Computer
Z-2582-2010 | 49599 | suture, surgical, absorbable, polydioxanone | NEW | Absorbable polydioxanone surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Product is labeled in part: "CP-SM397A-02...Mono-Dox Violet Monofilament Polydioxanone Synthetic Absorbable Sterile Suture. Do Not Resterilize.... 4/0 (1.5 metric) 30" (75cm) Rx Only...LOT L0535440 EXP 2009-11...4/0 M397A...Mono-Dox CUTTING NFS-2 NEEDLE...Manufactured by: CP Medical, Portland, OR USA".Intended for use in general soft tissue approximation and/or ligation including use in pediatric cardiovascular tissue where growth is expected to occur and ophthalmic surgery. | MonoDox Synthetic Absorbable Suture, Sterile | 2 | September 27, 2010 | 2010 | C P Medical | Seal integrity of inner foil pouch of the Mono-Dox Synthetic Absorbable Suture may be compromised. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | On 09/17/2008, the firm's Sale Representative began using the Telephone Script and calling all their customers.On 09/19/2008, the firm began sending the Recall Notice, dated 9/17/2008, letter to their customers. The letter states that the recalled product is Polydioxanone Suture, Code M397, Lot Number L0535440, Expiration Date: November 2009.The customers are advised to discontinue selling this product and return any existing stock to CP Medical. The firm also requests their consignees to notify their sub-accounts about this recall notice if the product is further distributed.The customers are instructed to provide return all accounted quantity and provide quantify information and location of any product that they are unable to recover for accountability purposes. Customers could contact CP Medical Customer Service Representative at 1-800-950-2763 for return goods authorization number and shipping account number. |  | 1696 | 1684/12 pcs. Boxes (1628/12 pcs. Boxes to veterinary market, 56/12 pcs. Boxes to human suture market) | Worldwide Distribution -- US, Australia, and Taiwan. | Terminated | September 28, 2010 | 1 | 1684 | Not_Computer
Z-2584-2010 | 56379 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | WECK IMA/ENT Blade Electrode ÂŽ (1.9 cm), REF 809335, Rx Only Sterile, Teleflex Medical, RTP, NCThe blade electrode is used in a wide variety of surgical procedures. The most common uses for electrodes with extended insulation is in cardiac bypass surgery (used in the take-down of the internal mammary artery or IMA) and ear, nose and throat surgery. They are sold non-sterile and are used one time and then disposed. | IMA/ENT Blade Electrode | 2 | September 27, 2010 | 2010 | Telefelx Medical | Complaints were received regarding peeling and melting of the insulation of the cautery tip of the IMA/ENT blade electrodes. |  | Teleflex Medical sent an "Urgent Medical Device Recall" letter dated July 30, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to immediately discontinue use, quarantine any affected products and return all of the affected product to Teleflex Medical; forward the letter to their customers to retrieve product from those locations and complete and return the Recall Acknowledgement Form via fax at 1-866-804-9881. A second mailing to non-responding customers on the consignee list will be mailed approximately four to six seeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail, or fax. Note: A record of these contacts will be maintained. Affected products received by Teleflex Medical will be destroyed upon receipt. Records will be maintained accordingly.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 190344 | 190,344 eaches | Worldwide distribution: USA and countries of Canada, Australia and Singapore. | Terminated | April 21, 2011 | 206 | 190344 | Not_Computer
Z-2579-2010 | 56199 | defibrillator, automatic implantable cardioverter, with cardiac resynchronization (crt-d) | NIK |  |  | Cardiovascular |  | Sorin Paradym CRT-D Model 8750 and Sorin Group Paradym sonR CRT-D Model 8770.Made in Italy by Sorin Biomedical CRM S.r.l. Via Crescentino s.n. , 13040 Saluggia (VC) Italy | Paradym Defibrillators | 2 | September 27, 2010 | 2010 | ela Medical Llc | The Paradym CRT 8750 and CRT sonR 8770 device models have a software anomaly, which will cause the device to lose the ability to sense/pace and to deliver further therapy. The software anomaly could only occurs under a specific sequence of events. | DESIGN: Software Design | Sorin Group issued a "Medical Device Correction" letter dated June 7, 2010 to consignees. The letter described the product, the problem, and actions to be taken by customers. A Sorin Group representative will contact customers to update the software. Prior to receiving the software update, customers should program the exercise AV delay to 40 ms, which will trigger a reset in the event of an anomaly. This will restore normal operation after shock delivery. If the anolmaly occurs on an implanted unit, a RV coil or SVC coil continuity test should be started, which will ensure that the device recovers proper operation.Corin Group can be contacted at 1-615-788-1807. | Contact the recalling firm for information | 215 | 215 | Nationwide Distribution in the states of AZ, CA, FL, IA, LA, MA, MS,MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, and VA. | Terminated | March 29, 2012 | 549 | 215 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2580-2010 | 56046 | surgical device, for ablation of cardiac tissue | OCL | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | Cardiovascular | 510(k) | Cardima Surgical Ablation Probe (SAP) Electrosurgical cutting and coagulation device, Model Number 01-242208, manufactured by Cardima Inc, Fremont, CA | Cardima Surgical Ablation Probe (SAP) | 2 | September 27, 2010 | 2010 | Cardima Inc | Sterility compromised-- The sterile pouch in which the product is packaged may be compromised, which could result in patient infection. |  | Cardima issued a Product Recall Notification dated May 18, 2010 to customers, identifying the affected product, the sterility issue, and the actions to be taken by the customer. Custormers are to quarantine and return the affected product to the firm. The firm will replace returned product as soon as it becomes available.Cardima Customer Service can be contacted at 800 354-0102. | Contact the recalling firm for information | 28 | 28 units | Worldwide Distribution to Distributors in USA (Pittsburgh, PA) and the countries of Ireland and Belgium, and customers in the states of CA, OH, MI, MO, AND WA, and the countries of Ireland and the Netherlands. | Terminated | March 23, 2011 | 177 | 28 | Not_Computer
Z-2585-2010 | 54561 | expander, surgical, skin graft | FZW | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | CONCEPT GRAFIX TENDON STRIPPER. REF TS8850. 7.0MM, QTY 1. Rx Only, Non-Sterile, Made In USA. | Grafix Tendon Stripper | 2 | September 28, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | There is a possibility that the device may exhibit Hydrogen Embrittlement in the material and are subject to potential breakage. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Recall Notification Letters were sent to customers on January 29, 2010. Customers were requested to complete and return a Customer Response Form indicating the number of products available for return and number distributed. Each distributor was asked to notifiy their customers if the product was further distributed. | Contact the recalling firm for information Contact the recalling firm for information | 149 | 149 | Distributed throughout U.S. and Internationally to Arab Emirates, Belgium, Canada, China, Columbia, Cyprus, Germany, Denmark, Egypt, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, Republic of Korea, Lithuania, Morocco, Malaysia, Poland, Portugal, Saudi Arabia, El Salvador, Thailand, Taiwan and Viet Nam. Domestic Distribution to CA, FL, GA, IL, IN, KS, MI, MT, NY, OH, PA, TX, and WA. | Terminated | January 12, 2011 | 106 | 149 | Not_Computer
Z-2588-2010 | 55708 | apparatus, suction, single patient use, portable, nonpowered | GCY | Nonpowered, single patient, portable suction apparatus. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Integra Suction Reservoir Kit for use with Integra NeuroSciences Subdural Drainage Catheter;Sterile, Rx only.Manufactured by Integra NeuroSciences Implants S.A.,Sophia Antipolis Cedex, France;distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, NJ 08538;This product is distributed by Integra NeuroSciences under catalog number 910-500 and in Suction Valve Reservoir Kits under catalog number 31223 and 31253. | Integra Suction Reservoir with AntiReflux Valve | 2 | September 28, 2010 | 2010 | Integra LifeSciences Corp. | The sterility of the product may be compromised. Integra LifeSciences Corp. received a recall notification from Saint-Gobain Performance Plastics, France, contract manufacturer of the Integra Suction Reservoir with Anti-Reflux Valve, that the product's sterility may be affected by a sterilization issue discovered by the supplier. |  | Integra sent Urgent Product Recall letters dated May 7, 2010 to customers identifying the affected product, the sterility issue, and actions to be taken by the customer. Customers are instructed to quarantine and return all affected inventory to Integra's distribution center in Sparks, NV, using a Returned Material Authorization Number received from Integra LIfeSciences Customer Service at 800 654-2873. All Customers are requested to return the attached "Recall Acknowledgement. and Return Form" Integra LIfeSciences can be contacted at 609-936-2495. |  | 313 | 313 units | Worldwide Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, WI, and the countries of Australia, Chile, Costa Rica, and Sri Lanka. | Terminated | August 27, 2012 | 699 | 313 | Not_Computer
Z-2609-2010 | 56160 | negative pressure wound therapy powered suction pump | OMP | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Engenex Non Acute Care Suction Pump for use with Engenex Negative Pressure Wound Therapy | Engenex Non Acute Care Suction Cup | 2 | September 28, 2010 | 2010 | Convatec Inc. | Reduced vacuum capacity has been reported with a small number of units caused by electrostatic discharge (ESD) or static electricity. | DESIGN: Component Design/Selection | Recall notification letters were sent to 46 customers, acute care facilities and durable medical device distributors on July 9, 2010, by Federal Express. | Contact the recalling firm for information Contact the recalling firm for information | 1190 | 1190 units | Nationwide distribution. | Terminated | June 07, 2011 | 252 | 1190 | Not_Computer
Z-2587-2010 | 56507 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Zoom Critical Care Bed, Model 2040. Stryker Medical, Kalamazoo, MI. Critical Care Patient Bed. | Stryker Zoom Critical Care Bed | 2 | September 28, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The Fowler mechanism may malfunction, preventing it from being lowered into a flat position. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Stryker Medical sent an Urgent Medical Device Alert letter dated May 14, 2010 to their customers. The letter identified the product, the problem, and the actions customers should take. The letter instructed users to do the following:" Return the postcard included with the letter to confirm receipt of the notification." Locate the recalled beds, verify the Fowler mechanism is functioning properly and does not display uneven motion when raising and lowering, and is not unusually noisy. " If the Fowler mechanism is not operating correctly, they were to call Stryker at 1-800-STRYKER, option 1, Reference RA-2010-014 to obtain service for the bed. " If any beds have been loaned or sold, the customer is to forward a copy of the notice to the new users and advise the firm of the new location." Provide the firm with the serial numbers for any beds that were disposed of." Retain the revised maintenance manual included with the notification.If the customer had any urgent questions or concerns, they were instructed to call the firm at 800-STRYKER, option 1, reference RA-2010-014, Monday-Friday 8am to 5pm (EST). | No consumer action necessary | 4926 | 4,926 | Worldwide Distribution - USA (all states), and including the countries of ARGENTINA, BRAZIL, CHILE, CHINA, CANADA, MEXICO, GREECE, JAPAN, KOREA, PUERTO RICO, SPAIN, POLAND, TAIWAN, SINGAPORE, INDIA, SWITZERLAND, UK and MALAYSIA. | Terminated | June 13, 2011 | 258 | 14060 | Not_Computer
Z-2589-2010 | 56543 | dressing, wound and burn, interactive | MGR |  |  | General & Plastic Surgery |  | Apligraf is supplied as a living, bi-layered skin substitute.Article Number: GS100Indicated for use with standard therapeutic compression for the treatment of non-infected partial and full-thickness skin ulcers due to venous insufficiency of greater than 1 month duration and which have not adequately responded to conventional ulcer therapy. | Organogenesis Apligraf | 2 | September 28, 2010 | 2010 | Organogenesis, Inc. | Product is contaminated with a yeast identified as Pseudozyma antarctica. | PRODUCTION CONTROLS: Process Control | Organogenesis notified accounts by telephone, e-mail or fax on 8/2/010 to remove from use and return recalled units. If unit was applied to patient then monitor for adverse events. The Notification Letter dated August 6, 2010, described the issue and provided instructions for customers. Customers are to place the affected unit on hold for its return to the firm. The firm will contact customers to make arrangements for its return. Replacements have already been sent. If the affected unit has already been applied, the patients should be monitored closely for any potential adverse events. An Organogenesis Medical Affairs rep will be in contact regarding patient follow up. Customers are to complete the attached form and fax it back to the number provided. Questions should be directed to the Medical Information Center at 1-888-432-5232. |  | 32 | 32 units | Nationwide Distribution -- AL, AZ, CA, CO, IL NC, PA, TN, and VA. | Terminated | March 10, 2011 | 163 | 32 | Not_Computer
Z-2608-2010 | 56709 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, 6F, REF 5540, Sterile EO, RX only, Vascular Solutions, Inc, 6464 Sycamore Court, Minneapolis, MN 55369, 110 cm length, | Langston Dual Lumen Catheters | 2 | September 28, 2010 | 2010 | Vascular Solutions, Inc. | Investigation of a recent Device Experience Report has made us aware of a potential problem with our 6F Langston dual lumen pigtail catheter (Model 5540) with the following lot numbers: 548110, 548227, 548364, 548430, 548554 & 548723.Recently, it was reported that pouches containing Langston Pigtail catheter (Model 5540) were missing a seal, posing a risk of product contamination. It is possib | PRODUCTION CONTROLS: Process Control | Consignees were sent on 9/2/10, a Vascular Solutions "Urgent Medical Device Field Action" letter dated September 1, 2010. The letter was addressed to Cath Lab Manager. The letter descried the product and the problem. Provided instructions to examine pouches of opened boxes to see if they have not been sealed. If unsealed, they recommended to remove pouches from the current inventory. For unopened boxes they recommended to remove the boxes from the inventory. Requested consignees to complete the Field Action Customer Inventory Form included with the letter. For any questions, please contact the firm at 763-656-6032. | Contact the recalling firm for information | 1128 | 1,128 boxes ( 5pk) | AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS,MO, NE, NY,NH, NJ, NM, NY,NC, OH, OK, OR, A, PR, SC, TN, TX, UT, VA, WA, WV, WI, and WY.SINGAPORE, MEXICO, CANADA, SPAIN, AUSTRIA, GERMANY, ITALY, ENGLAND, and SWITZERLAND. | Terminated | April 24, 2012 | 574 | 1128 | Not_Computer
Z-2601-2010 | 56522 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 7 MM DISPOSABLE DOWEL HARVEST TUBE WITH PROTECTION SLEEVE, REF 900737, Sterile, BIOMET Sports Medicine, Warsaw, IN 46581.Usage: Bone Coring Device used in ACL and PCL procedures. | Disposable Dowel Harvest Tube | 2 | September 28, 2010 | 2010 | Biomet, Inc. | The affected instruments are discolored and/or have a sticky residue on the end. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | The firm sent out "Urgent Medical Device Recall" notices dated 7/1/2010. Letters were sent to distributors, customers and OR managers. The letters stated that the devices should be located, quarantined and shipped back to the recalling firm. On 7/15/2010, the firm added one lot of one additional product to the recall (product 900788 lot 780470). The firm sent letters to physicians that had used that lot of product.Questions relating to this recall should be directed to 574-371-3755. |  | 35 | 35 | US, Canada, Australia, Chile, Argentina, Netherlands, Costa Rica and Japan. | Terminated | August 20, 2012 | 692 | 701 | Not_Computer
Z-2591-2010 | 56695 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 250 mL PainPump with Luer Lock Tubing Set, sterile, REF 525-350, Stryker Instruments, Kalamazoo, MI. Delivers controlled amounts of medication and narcotics. | PainPump2 | 2 | September 28, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | All of the failures lead to the patient receiving less medication than intended. Certain lots of the pain pump have a high failure rate associated with a cracked port housing which leads to internal or external leaks. Some lots have a high failure rate associated with E2 errors. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm sent our Urgent Medical Device Recall Notices dated 7/12/2010 to Risk Managers and Materials Managers. The recall notice listed specific lots of PainPump2 and PainPump2 BlockAid that were under recall. Consignees were told to check stock for the recalled product, quarantine the product, and contact Stryker to arrange return shipment of the recalled product.On 9/2/2010 the firm sent out a second Urgent Medical Device Recall Notice. This notice expanded the scope of the recall to include all products distributed within the past 12 months. The shelf life of the product is 1 year -so the recall was for all product currently in the supply chain. The letter again directed consignees to locate the recalled products and contact Stryker to arrange return shipment of the recalled product. The firm sent an additional letter dated 9/1/2010 to its customers. This letter stated that Stryker will no longer be offering any of their PainPump2 products. | Contact the recalling firm for information | 54568 | 54,568 all products | Worldwide Distribution -- US, Canada, and Australia. | Terminated | April 21, 2011 | 205 | 54568 | Not_Computer
Z-2607-2010 | 56485 | radioimmunoassay, human growth hormone | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 Immunoradiometric Assay (IRMA) Kit, PartNumber: DSL-6600usage: The DSL6600 IGFBP-3 immunoradiometric assay (IRMA) kit provides materials for the quantitative measurement of IGFBP-3 in serum. This assay is intended for in vitro diagnostic use. | IGFBP3 IRMA Kit | 2 | September 28, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports that IGFBP-3 measurementsusing IGFBP-3 IRMA kit (REF DSL-6600) lots identified above may result in IGFBP-3 levels that are not in agreement with simultaneously measured levels ofIGF-1. The investigation confirmed that above mentioned lots of IGFBP-3 kit provide results elevated by 84% on average. Erroneously elevated IGF |  | The recall communication was initiated with a Product Corrective Action (PCA) letter that was sent on August 09, 2010 to allcustomers who received the IGFBP-3 IRMA kits, lot numbers referenced. The letter will provide the customers with an explanation of the problem identified and an action to be taken.Customers were instructed to:(1) Discontinue using all lots identified above and discard in accordance withlocal regulations.(2) To continue testing for IGFBP-3 levels, find an alternate method.(3) Immediately notify all parties of interest (e.g., laboratory staff, clinicians) ofthe possibility of erroneously elevated IGFBP-3 results for patient samples tested with these lots, and advise all parties that clinical decision making and patient safety may be impacted.(4) Review patient IGFBP-3 results for the lots listed above, and interpret them in the light of additional tests (e.g., IGF-1) and the patient's clinical status and history.Customers were also asked to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days. For any questions regarding this notification, customers were to contact the DSLTechnical Support Center at 1-800-231-7970 option 4 in the United States, or contact your local Beckman Coulter Representative. |  | 63 | 63 | Nationwide in US and Canada | Terminated | August 13, 2012 | 685 | 63 | Not_Computer
Z-2596-2010 | 56366 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MODEL P8400 HANDLE WITH INTEGRATED CABLE, FOR USE WITH MODEL P8400 CONTROLLER, REF 7013-8050, QTY (1), STERILE EO, HEMOSTATIX MEDICAL TECHNOLOGIES, LLC, BARTLETT, TN 38133 USA, CE 0473, EU REP : QFI LTD, STRATFORD, LONDON E15 2ST UKThe Hemostatix Thermal Scalpel is a surgical instrument designed to retain the precise, clean cutting characteristics of the traditional steel scalpel while minimizing blood loss by simultaneously sealing blood vessels as they are cut, with minimum damage to surrounding tissue and virtually no muscle stimulation, using heat thermally conducted to the tissue from an elevated-temperature blade. | MODEL P8400 HANDLE | 2 | September 28, 2010 | 2010 | Hemostatix Medical Technologies LLC | Lack of sterility assurance | OTHER/UNDETERMINED: Pending | On July 27, 2010, the firm, Hemostatix Medical Technologies, LLC, notified their international distributor in Japan of the recall by email. Hemostatix sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 29, 2010, to U.S. customers delivered via FedEx Priority Overnight. All non-responding customers were to be notified by telephone after 2 weeks. The other international customer, Italy, were to be notified by email on July 30, 2010. The letter describes the product, problem and action to be taken by customers. The customers were instructed to immediately check their inventory and remove any affected product found; ensure that satellite facilities/clinics receive the recall notice; complete and return the Recall Response Form via fax to Hemostatix's international fax number at 901-261-0017, and contact the Customer Care Department's recall hotline at 901-261-0012 for further instructions, if they do have affected stock.If you have any questions or concerns, please feel free to contact our Customer Care Department at: 1-800-722-5735 in the USA or 1-901-261-0012 outside the USA. | Contact the recalling firm for information | 40 | 40 units | Worldwide distribution: USA including CA, GA, LA, MA, MD, MO, NC, OH, OR, PA, SC, TX, VA, and WI; and countries of Japan and Italy. | Terminated | November 29, 2010 | 62 | 259 | Not_Computer
Z-2606-2010 | 56322 | pump, blood, cardiopulmonary bypass, non-roller type | KFM | Nonroller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump with X-Coating, packaged within custom ordered Cardiovascular Procedure Kits, P/N 3ZZ164275X, 3ZZ164275, CEOT028, CEOT032; and as single sterile units with product codes CV-164275, 164275 and 164275X. CV-164275 is an item number for product that is sold to foreign affiliate.Usage: The Sarns Centrifugal Pump (with X-Coating) is indicated as an extracorporeal blood pump for use in cardiopulmonary bypass procedures. | Sarns Centrifugal Pump with XCoating | 2 | September 28, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device component may malfunction and require user intervention during surgical bypass procedures. | OTHER/UNDETERMINED: Pending | The recalling firm notified end users, perfusion groups, and distributors by phone, fax, email and mail on 03/15/10 to advise of the required recall actions with instructions for labeling and returning the involved product from the customer level. The firm expanded their recall to include additional lots on 05/17/10. If you have any questions, please contact Customer Service at 800-521-2818. | Contact the recalling firm for information | 28257 | 28257 pumps | The product was sold to medical device distributors and to direct accounts nationwide and to a subsidiary in Japan. | Terminated | November 04, 2011 | 402 | 28257 | Not_Computer
Z-2624-2010 | 50095 | prosthesis, knee, femorotibial, constrained, cemented, metal/polymer | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GMRS Extension Piece, 60 mm; SterileStryker Howmedica Osteonics; Howmedica Internation S. de. R.L. Limerick, Ireland.The Global Modular Replacement System is intended to be used with the components of the Howmedica Osteonics Modular Replacement System , the Howmedica Osteonics Modular Rotating Hinge Knee System and/or the Howmedica Osteonics Kinematic Rotating Hinge Knee System in situations where replacement of extensive bone loss in the femur and/or proximal tibia is required. The Proximal Femoral Module and Distal Femoral Module of the GMR may be used together, or with the components of the MRS, in replacement of the total femur. | Global Modular Replacement System (GMRS) Extension Piece 60 mm | 2 | September 29, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Mislabeled package: Package labeled as GMRS Extension Piece 60 mm, lot ZYCOC1 actually contains a GMRS Extension Piece 70 mm extension piece. | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker issued Important Market WIthdrawal letters dated May 5, 2005 to customers and distributors. The letters identified the affected product, the labeling issue, and actions to be taken by the customer or distributor. Customers were instructed to identify and quarantine affected product, which will be reconciled and returned by the Stryker sales representative. Distributors were instructed to identify affected product, reconcile all product on an attached Product Accountability Form to be faxed to Stryker at 201 831-4825 within two days of receipt of the letter. The affected product was to be returned using the enclosed pre-paid mailing label. Contact the Stryker sales representative to re-order product that is being returned.Stryker can be contacted at 201 831-5825. |  | 5 | 5 units | Nationwide Distribution: USA, including the states of AZ, FL, MD, and PA.G | Terminated | October 01, 2010 | 2 | 5 | Not_Computer
Z-2618-2010 | 54617 | wheelchair, powered | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Jazzy Select with XLR Port Harness | Jazzy Select with XLR Port Harness | 2 | September 29, 2010 | 2010 | Pride Mobility Products Corp | Improper set-up of the wire harness can allow wires to be pulled from their intended position and become caught under the motor mounting bracket located under the powered wheelchair. This can cause the wire connector and/or the wire itself to be pulled or pinched which can lead to a short or open circuit. | OTHER/UNDETERMINED: Pending | The recalling firm issued recall notifications with response cards to their dealers. Included with the notification was fasteners and an example letter for customer notification of the problem. Dealers were instructed to contact the end users to make arrangements with the customer/end user to verify that the devices were properly set up. If the dealers' inspection revealed that additional fasteners were required they dealer would install them. The recalling firm also issued a recall letter to consumers. The consumer recall letter included instructions that if the correction was completed by anyone other than the wheelchair provider, then the consumer should schedule an appointment with the wheelchair provider to ensure the correction was completed properly. |  |  |  | The product was shipped to dealers nationwide who further shipped the product to end users. The product was also shipped to distributors in Canada, Australia and United Kingdom | Open |  |  |  | Not_Computer
Z-2620-2010 | 55188 | radioimmunoassay, progesterone | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems PROGESTERONE, Part Number: 33550 | Access Immunoassay Systems PROGESTERONE | 2 | September 29, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter after the firm confirmed customer reports that the affected lot of Access Cortisol reagent (REF 33600) identified above may produce suppressed patient results. Beckman Coulter also confirmed customer reports that the two lots of Access Progesterone reagent (REF 33550) identified above may produce suppressed patient results and an increased incidence of I | OTHER/UNDETERMINED: Pending | Two (2) Product Corrective Action (PCA) letters with attached Customer Response form were sent on the week of February 15, 2010 to the affected customers. The letters provide the customers with an explanation of the problem identified and instruct the customers to discontinue use of Access Progesterone and Access Cortisol Reagents, lot numbers referenced. For product replacements consignees were instructed :(1) In the United States, please contact Customer Service at 1-800-526-3821; (2) In Canada, please contact Customer Service at 1-800-463-7828;(3) Outside of the United States and Canada, contact your local Beckman Coulter Representative.If consignees need assistance or have any questions regarding this notification, they were instructed to contact TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. | Contact the recalling firm for information | 6412 | 6412 | United States & Canada | Terminated | April 19, 2012 | 568 | 13694 | Not_Computer
Z-2622-2010 | 56110 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Stryker Scope Warmer, disposable. Endoscopic accessory composed of polyvinyl Chloride (PVC) and Sodium Acetate, Model Number 502-360-000, manufactured by Stryker Endoscopy, San Jose, CA. | Stryker Scope Warmer | 2 | September 29, 2010 | 2010 | Stryker Endoscopy | Outer packaging seals have inconsistencies that have resulted in compromised seals and may result in compromised sterility. | PRODUCTION CONTROLS: Process Control | Stryker issued an Urgent Device Removal letter, dated June 21, 2010, to customers which identified the affected product, the sterility issue, and the actions to be taken by the customer. Customers were instructed to identify and quarantine any affected product in their inventory, and contact Stryker Customer Service at 800 624-4422 and request an RMA#. The RMA# is to be written on the outside of the package used to return products, using an enclosed pre-paid shipping label. All cusotmers are to complete and sign an enclosed "acknowledgement of receipt" and fax to (408) 754-8278.Stryker can be contacted at 800-624-4422. |  | 191855 | 191855 units | Worldwide Distribution: USA, Including the states of AK, AL, AR, CA, CO,CT, DE, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NJ, NM, NH, NV, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Brazil, Canada, Chile, China,France, Germany, Iberia, India, Italy, Netherlands, New Zealand, Japan, Korea, Latin America, Polska, Singapore, South Africa, Switzerland, Thailand, and United Kingdom. | Terminated | December 16, 2010 | 78 | 191855 | Not_Computer
Z-2548-2010 | 56548 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | All Med Medical 22ga x 1" Straight Hub NeedleModel #M2215HSReorder #55-2235Intended to be used to inject fluids into, withdraw fluids from inside the vascular space through a subcutaneously implanted port. | All Med Medical 22ga x 1" Straight Hub Needle | 1 | September 29, 2010 | 2010 | Multi-Med, Inc. | Huber needles may core and result in port leakage or emboli being flushed our of the port. | DESIGN: Device Design | Multi-Med issued an Urgent: Medical Device Recall notification on 8/30/10 via fax and followed by hard copy letter. Users were advised of the problem and asked to immediately examine their inventory and quarantine the affected product. They were to also notify their customers of the recall. The firm stated in the letter that they would contact customers to determine the number of product that they had and the amount they expected from the field. The firm will coordinate a plan for the return of the affected product. Questions should be directed to Sue Starkey at 603-357-8733. |  | 494 | 494 units | Worldwide Distribution -- USA and Mexico. | Terminated | January 31, 2013 | 855 | 546873 | Not_Computer
Z-2623-2010 | 56261 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Strattice Reconstructive Tissue Matrix for Stoma Reinforcement.8 cm x 8 cmPackaged in an inner foil pouch inside outer foil pouch tertiary carton.LifeCell Corporation, One Millennium Way, Branchburg, NJ 08876Strattice is intended for use as a foil tissue patch to reinforce soft tissue where weakness exists and for the surgical repair of damaged or ruptured soft tissue membranes. It is used for repair of hernias and/or body wall defects which require the use of reinforcing or bridging material to obtain the desired surgical outcome. It is intended for single patient one time use. | Strattice Reconstructuve Tissue Matrix for Stoma Reinforcement | 3 | September 29, 2010 | 2010 | LifeCell Corporation | Mislabeled -One lot of product labeled as 8 cm x 8 cm were actually cut using a 6 cm x 6 cm template. |  | LifeCell issued Recall Notice letters dated July 12, 2010 to customers, identifying the affected product, the labeling issue, and actions to be taken by the customer. Two versions of the recall notification were sent, one for implanted units and one for those not yet implanted. Both letters stated that the mislabeling has no impact of the safety or effectiveness of the medical device. Replacement units were offered for all affected units. Those with unused units were to contact LifeCell Customer Solutions for return and replacement of the unit at 1-866-423-2433.Those with implanted units may contact LifeCell Customer Solutions at 1-866-423-2433 with any additional questions or concerns. | Contact the recalling firm for information | 32 | 32 units | Nationwide Distribution: USA, including the states of AZ, DC, FL, GA, ID, IL, KY, LA, MO, NC, OH, OK, OR, TX, WA, and WI. | Terminated | October 01, 2010 | 2 | 32 | Not_Computer
Z-2633-2010 | 56505 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Gemstar Pump Set with Convertible Pin, Latex-Free set with Attached Pressure-Activated Anti-Siphon Valve, Pump Yellow Striped Tubing-SL, 95 inch; Rx, sterile, 24 individually packaged sets per case; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list number 20635-01 | Gemstar Pump Set with Convertible Pin | 2 | September 29, 2010 | 2010 | Hospira Inc. | Underdelivery of infusion during clinical use at low rate settings (less than 10 mL/hour). Patient may not receive the intended amount of medication. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hospira issued Urgent Device Recall letters dated August 12, 2010 to Hospira's direct accounts, informing them of the potential for under delivery under low rate settings using the affected lots of pump sets. The accounts were instructed to examine their inventory for the affected product, quarantine it immediately, and complete the attached reply form indicating the amount of product being returned, faxing it to Stericycle at 1-877-523-9110. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. Accounts may contact Hospira Customer Care at 1-877-946-7747 or their Hospira Account Manager to obtain replacement product. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. Customers may contact Stericycle, which is managing this recall, at 877-274-7163 for further information concerning this recall. |  | 71016 | 71,016 sets | Worldwide Distribution: USA, (all states), and the countries of Canada, Costa Rica and Taiwan. | Terminated | June 21, 2012 | 631 | 340728 | Not_Computer
Z-2625-2010 | 56646 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips M3002A IntelliVue X2Model: 865039Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring, recording, and alarming of multiple physiological parameters of adults, pediatrics, and neonates in a hospital environment. ST segment monitoring is restricted to adult patients only. The M3002A and M8102A are used for monitoring, recording and alarming of multiple physiological parameters of adults, pediatrics, and neonates in health care facilities | Philips M3002A IntelliVue X2 | 2 | September 29, 2010 | 2010 | Philips Healthcare Inc. | Speakers on the Philips IntelliVue X2 and MP2 patient monitors may fail, causing absence of an audible alarm and delaying patient treatment. | DESIGN: Process Design | Philips issued a URGENT - MEDICAL DEVICE letter dated August 26, 2010, to each affected customer. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue and gives instructions on actions to take while awaiting their replacement speaker assemblies. The speaker replacement assembly is expected to be available in 8 weeks.Affected products may continue to e used while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP at power-up.If you experience this INOP or there is no sound from your X2/MP2 speaker while it is not connected to another audible alarm annunciation system, remove the device from use and contact your local Philips service representative at 1-800-722-9377. | Contact the recalling firm for information | 19400 | 19,400 units | Worldwide Distribution - USA and the countries of : ARGENTINA AUSTRALIA AUSTRIA BAHRAIN BANGLADESH BELGIUM BRAZIL CHINA CZECH REPUBLIC DENMARK EGYPT ESTONIA FINLAND FRANCE GERMANY GREECE HONG KONG HUNGARY ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KENYA KUWAIT LATVIA LIBYAN ARAB JAMAHIRI LITHUANIA MALAYSIA MAURITIUS MEXICO MOROCCO NAMIBIA NETHERLANDS NEW ZEALAND NORWAY OMAN PAKISTAN PALESTINE PARAGUAY POLAND PORTUGAL QATAR ROMANIA RUSSIA SAUDI ARABIA SINGAPORE SLOVAKIA SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SRI LANKA SWEDEN SWITZERLAND SYRIAN ARAB REPUBLIC TAIWAN THAILAND TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN and VIET NAM | Open |  |  | 22232 | I/O | Alarm/Message | Replace | Remove or Replace
Z-2617-2010 | 56413 | heat-exchanger, cardiopulmonary bypass | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit X-Coated H/L pack with Cardioplegia, Item # 70067-03 | Cardiovascular Procedure KitTerumo Cardiovascular Systems Corporation | 2 | September 29, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid. | PRODUCTION CONTROLS: Equipment Maintenance | Consignees were contacted on June 8, 2010, by phone, fax, email and priority mail to advise distributors and end-users of the potential occluded vent port defect and provide instructions for the replacement product recall. Consignees were mailed follow-up communications by priority overnight mail on June 9, 2010. | Contact the recalling firm for information | 60 | 60 kits | The products were distributed nationwide. | Terminated | February 29, 2012 | 518 | 3979 | Not_Computer
Z-2590-2010 | 55173 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 880i, 860i and 600i Systems; Software Part Number: A84500, A86646; Softwareversion: 4.9.01 The UniCel Synchron Clinical Systems are fully automated, computer-controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent) | Beckman Coulter UniCel DxC 880i, 860i, and 600i Systems | 2 | September 30, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall on their UniCel DxC Clinical Systems (880i, 860i,680i and 660i Systems with serial number | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an URGENT: FIELD SAFETY NOTICE letter, dated March 25, 2010, to all customers. The letter identified the affected product, the problem, and the actions customers take.Customers were instructed to:if the UniCel DxC system serial number was less than 2760 the customer should contact their local representative to arrange for an upgrade to software version 4.9 build 02.Complete and return the enclosed response form within 10 days of the notice.Share the information with their laboratory staff and retain the notification as part of the Quality System documentation. If affected products were forwarded to another laboratory, provide a copy of the notification letter to them.Customers may contact the Customer Support Center (Hotline) at 1-800-854-3633. For questions or comments please contact the firm at (714) 961-3634. | No consumer action necessary | 22 | 22 | Class II Recall - USA, including the states of CA, IN, GA, NC, NY, OH, SC, and WI | Terminated | July 02, 2012 | 641 | 22 | Software | Output/Calculation | Software update | Software Update
Z-2634-2010 | 55680 | test, urine leukocyte | LJX | Leukocyte peroxidase test. | Hematology | Hematology | 510(K) Exempt | Hypoguard, DiaScreen 4pH, Reagent Strips for Urinalysis, 100 strips, Cat. No. D11400. | DiaScreen Reagent Strips | 3 | September 30, 2010 | 2010 | ARKRAY USA INC. | This recall has been initiated due to a labeling issue with the Hypoguard DiaScreen 4-way pH 100 strips. The front and back panel of the bottle carton lists five reagents when it should only list four reagents. The Ketone reagent should not be included on the label. The test strips contained in the bottle and all other labeling are correct and properly align with 4pH part number D11400 listed on | MISBRANDING: Labeling False and Misleading | Distributors were sent a Arkray "URGENT: Voluntary Medical Device RECALL " letter dated March 24, 2010. The letter described the problem and the product being recalled. The letter requested the distributors to immediately examine their inventory , quarantine and return the product subject to the recall. If the distributors had further distributed the recalled product, then they were requested to notify the customers of the recall. Arkray enclosed a customer recall letter that the distributors were to use for notifying their customers.Distributors were also contacted by phone on March 24, 2010, and were informed of the recall issue, the product involved, instructed to quarantine the product and to notify their customers. | Contact the recalling firm for information | 5504 | 5504 | FL, NJ, MI, NV, FL, OH, GA, WI, NY, KY, LA, CA, CT, TN, ILISRAEL | Terminated | March 08, 2012 | 525 | 5504 | Not_Computer
Z-2653-2010 | 56247 | igg, antigen, antiserum, control | DEW | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent, Part Number: 467920 | SYNCHRON System(s) Immunoglobulin A (IgA) Reagent | 2 | September 30, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter Inc. confirmed reports of high recovery of immunoglobulin A (IgA) in proficiency surveys or patient samples using Synchron Ig-A reagent lots M812164 or M902345. Using these reagent lots, some samples deficient in IgA have yielded results in the normal range.Other samples with normal levels of IgA have yielded falsely high results. | OTHER/UNDETERMINED: Pending | A Product Corrective Action (PCA) letter was sent on the week of June 28, 2010 to the affected customers. The letter provided the customers with an explanation of the problem identified and advised customers to discontinue use of the affected lots of the SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent.Customers were instructed to:(1) Discontinue use of Ig-A reagent lots M812164 and M902345. Contact your Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact your local Beckman Coulter Representative to discuss reagent replacement options.(2) You may need to re-evaluate samples for IgA that were tested with either of the listed lots of Ig-A reagent.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.Customers were also instructed to share this information with your laboratory staff and retain this notification as part of yourQuality System documentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them. If they have any questions regarding this Product Corrective Action letter, customers were told to contact our Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Requests, or call (800) 854-3633 in theUnited States and Canada. |  | 852 | 852 units | US Nationwide & Canada | Terminated | July 31, 2012 | 670 | 852 | Not_Computer
Z-2636-2010 | 56220 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Glenoid Head Inserter (RSP Inserter/Impactor), Device Part #804-03-041, All Lots.Product is a Shoulder Instrument used to place the Glenoid Head Trial onto the Baseplate | Glenoid Head Inserter (RSP Inserter/Impactor) | 2 | September 30, 2010 | 2010 | Encore Medical, Lp | Complaints indicate product's distal threaded tip may fracture under certain circumstances. | DESIGN: Device Design | DJO Surgical sent an "Urgent Field Safety Notice" Letter dated July 12, 2010, to all of their consignees. The letter described the product, the problem, and the action consignees should take. Consignees were instructed to: Contact their Customer Service Representative for an R-RPR (RMS) number and return all RSP Glenoid Head Inserters, 804-03-041, under the assigned number to receive credit. Complete and return the Field Safety Notice Response to the firm. Identify and inform any customers who were sold or received the product of the Field Safety Notice Any questions consignees should call the firm at (512) 834-6330. | Contact the recalling firm for information | 104 | 104 | Worldwide Distribution - USA, including the states of AR, AZ, CA, FL, GA, ID, IL, IN, LA, MA, ME, MI, MO, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, WA, and WI; and the countries of Puerto Rico, Germany, and Saudi Arabia. | Terminated | June 30, 2011 | 273 | 104 | Not_Computer
Z-2651-2010 | 56304 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GYRUS ACMI, dissector PlasmaKnife" , REF 7035-3005, Rx only, STERILE R, CE 0344, 161416-CA, MANUFACTURED FOR; GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104, USA, EC REP, GYRUS MEDICAL LIMITED, FORTRAN ROAD, ST. MELLONS, CARDIFF, CF3 0LT, UK | GYRUS ACMI, dissector PlasmaKnife | 2 | September 30, 2010 | 2010 | Gyrus Acmi, Incorporated | Lack of sterility assurance due to compromised package seals. | PRODUCTION CONTROLS: Packaging | The firm initiated their recall to their direct accounts by Priority Mail with delivery confirmation on 08/06/2010. The letter requested return of the product and was extended to the retail level. |  | 4244 | 4244 units | WorldwideIn the United States: AL, AK, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV and the District of Columbia. Foreign: UK, TAIWAN, SOUTH KOREA, GERMANY, CANADA, AUSTRALIA, SPAIN, ITALY, ISRAEL, NORWAY, LITHUANIA, SWEDEN, SWITZERLAND, SAUDI ARABIA, TURKEY, and GREAT BRITAIN | Terminated | January 26, 2012 | 483 | 4244 | Not_Computer
Z-2637-2010 | 56463 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Integra CUSA Ultrasonics NXT Ultra Surgical Aspirator SystemCUSANXT1:Latex Free, Rx only:Integra LifeSciences (Ireland) Limited,County Offaly, Ireland;Distributed by Integra Life Sciences Corporation311 Enterprise DrivePlainsboro, NJ 08536. | CUSA NXT Ultra Surgical Aspirator System | 2 | September 30, 2010 | 2010 | Integra LifeSciences Corp. | A ground fault malfunction at the rectifier retaining screw may cause the aspiration pump in the CUSA NXT Service module to run at maximum speed, and continue to run until the entire system is turned off. This may result in excess aspiration that requires medical intervention. | DESIGN: Component Design/Selection | Integra issued Urgent: Product Recall Notification letters dated July 29, 2010 to all customers on July 29, 2010 identifying the affected products, issue prompting the recall, and actions to be taken by the customer. Letters were also forwarded to all Integra Account Reps.Customers were advised to discontinue use of and quarantine the affected service modules. The Surgical Aspirator may be operated without the Service Module, utilizing the hospitals wall aspiration. Company representatives will contact customers to arrange product return and repair of the affected service modules. Customers are to complete and return an enclosed Recall Acknowledgement and Return Form . The firm can be contacted at 609 936-2485. |  | 26 | 26 in US, 8 in Europe and 1 in Asia (both unit types. | Worldwide Distribution: USA , including the states of AZ, CA, CO, FL, GA, IA, KY, MN, MO, NV, NY, OH, OR, RI, TX, VA, WI, and the countries of Germany, Hungary, Indonesia, Italy, Netherlands, and Switzerland. | Terminated | May 15, 2012 | 593 | 26 | Not_Computer
Z-2635-2010 | 56292 | laparoscope, general & plastic surgery, reprocessed | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EnsealÂŽ TRIO Product Code ETRIO335H, Lot Code: G4TC14, Exp. Date 2012/05 | EnSeal TRIO | 2 | September 30, 2010 | 2010 | Ethicon Endo-Surgery Inc | The weld that connects the electrode to the active rod was separated. This connection is critical to the proper function of these devices. If the weld is separated, the product will not deliver energy to the jaws, resulting in the potential for a cut/transection without a seal. Ethicon Endo-Surgery identified this issue during a routine finished goods quality inspection. |  | Ethicon Endo-Surgery issued an Urgent: Device Recall, Event 2193 letter dated July 19, 2010 identifying the affected device, the issue prompting the recall, and actions to be taken by consignees.Consignees were instructed:Do not use the affected product.Identify and quarantine affected product from inventory.Complete the Business Reply form and fax to Stericycle at 877-597-9580 within three daysTo return product, place it in a box with a copy of the Business Reply form and mail to Stericycle using the pre-printed UPS shipping label. Replacement product should be received by the customer within 60 days of return of affected product. Customers can contact Stericycle about this recall at 877-243-7314 and referencing Event 2193. | Contact the recalling firm for information | 396 | 396 instruments | Nationwide Distribution to the following states: AZ, AL, AZ, CA, CO, CT, FL, GA, IN, KY, MA, MD, MI, MO, MS, NC, NV, NY, OH, PA, TN, TX, WA, and WV. | Terminated | February 14, 2011 | 137 | 396 | Not_Computer
Z-2641-2010 | 56188 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Test Strips, CHOL +HDL+GLU Panel for CardioChek PA brand analyzers, 15 tests 1 MEMo Chip, manufacturer Polymer Technology Systems, Inc., 7736 Zionsville Rd., Indianapolis, IN 46268. | CardioChek PTS Panels CHOLHDLGLU | 2 | September 30, 2010 | 2010 | Polymer Technology Systems, Inc. | The test strips appear to under-recover Cholesterol, HDL and glucose. | OTHER/UNDETERMINED: Pending | Distributors and direct customers were notified by letter on July 2, 2010. They were told to discontinue use and dispose of the affected lot of test strips. The firm also asked customers to complete an acknowledgement and confirmation document . |  | 5805 | 5805 | US accounts in: CA, MA, FL, NY, KY, VA, UT, MN, OH, and MD. Direct foreign accounts in: Belgium, England, Sweden, and Italy. | Terminated | May 11, 2011 | 223 | 5805 | Not_Computer
Z-2652-2010 | 56481 | laser, ophthalmic | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Endo Ocular Directional Laser Probe, Lumenis Inc. Salt Lake City, UT STERILE EO | Endo Ocular Directional Laser Probe | 2 | September 30, 2010 | 2010 | Lumenis, Inc. | Firm did not have proper 510(K) clearance. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Lumenis sent letters to all customers on 7/27/10 instructing them to review their inventory, isolate affected product, complete a recall verification form and return all affected product to them. | No consumer action necessary | 250 | 250 probes (25 boxes, each containing 10 probes) | FL, IL, MD, CO and TX | Terminated | October 05, 2010 | 5 | 250 | Not_Computer
Z-2654-2010 | 49387 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Laparoscopic Ultrasound Transducer (LAP Transducer) is used on the SonoSite MicroMaxx ultrasound system. The LAP Transducer label is labeled in part: "MicroMaxx LAP...Manufactured for SonoSite, Inc., Bothell, WA 98021". | LAP Surgical Transducer | 2 | October 01, 2010 | 2010 | Sonosite, Inc. | Sterilization of the SonoSite LAP surgical Transducer with the Sterrad NX System may result in damage to the transducer. | PRODUCTION CONTROLS: Process Control | On November 2, 2007, SonoSite sent out the Important Information Regarding Sterrad NX Incompatibility letter and the Technical Bulletin 005/07 Rev A to their consignees. The letter advises the consignees that SonoSite has recently found that sterilization of the SonoSite SLT and LAP surgical ultrasound transducers with the Sterrad NX system may result in damage to the transducer.The Technical Bulletin 005/07 Rev B dated November 2, 2007 informs the consignees that the damage has been observed primarily on the circuit board in the transducer connector, and cable and laparoscope damage may also be occurring. All damages have occurred after processing using the Sterrad NX systems in its "Advanced" cycle and the potential damage from the use of "normal" cycle is being evaluated by the firm.Consignees should call the Technical Support at 1-425-951-1330 or 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall. |  | 16 | 16 units | Devices were distributed to medical centers and hospitals in AR, CA, IN, KS, MA, MD, MI, NC, NE, NJ, NV, OH, SC, TN, TX, WI, Washington DC and to FOREIGN COUNTRIES Australia, China, France, Germany, Greece, Hong Kong, India, Iraq, Italy, Korea, Republic of, Libyan Arab Jamahiriya, Mexico, Spain, United Arab Emirates, United Kingdom, and United States. | Terminated | October 05, 2010 | 4 | 54 | Not_Computer
Z-2639-2010 | 54382 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Test Strips, Part Number 020-245-32, lot number 2964512, Manufactured by LifeScan Inc., Milpitas, CA. | One Touch Ultra Test Strips | 2 | October 01, 2010 | 2010 | Lifescan Inc | One Touch Ultra Test Strip products may actually contain a different product, One Touch Select strips. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Notification was issued by the firm during the first week of February. This letter asked consignees to share the information with patients, and are asking users to contact the firm and to examine their stock for the recalled product. |  | 9631 | 9631 cartons, 4 vials per carton | Product was distributed to one wholesale dealer and one distributor and 2 medical facilities in CA, OH and PA, as well as the UK. | Terminated | March 14, 2011 | 164 | 9631 | Not_Computer
Z-2658-2010 | 55196 | radioimmunoassay, free thyroxine | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Free T4, Part Number: 33880 | Access Free T4 Assay | 2 | October 01, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter determined that kit boxes for Access Free T4 reagent (REF 33880), lot number 916293 may contain reagent packs labeled as Access Vitamin B12, (REF 33000), lot 916277. Using an Access Vitamin B12 reagent pack, lot 916277, filled with Access Free T4 reagents may generate incorrect patient results. | OTHER/UNDETERMINED: Pending | A Product Corrective Action (PCA) letter was sent on March 24, 2010 to the affected customers. The letter provides the customers with an explanation of the problem identified and instructs them to discard the FT4 kit, Lot 916293 in their stock if the FT4 packs within the kit packaging are labeled asVitamin B12 Lot 916277.The PCA requested customers:(1) Inspect all Access Vitamin B12 reagent packs that have been removed from their kit boxes, including reagent packs currently loaded on your system and discard any reagent packslabeled as Access Vitamin B12, lot 916277;(2) Inspect your Access Free T4 reagent kit inventory for any reagent packs labeled as Access Vitamin B12, lot 916277 and discard any reagent packs labeled as Access Vitamin B12, lot916277; (3) Review your historical quality control to ensure that assay performance for Access Vitamin B12 lot 916277 has been within your laboratory specifications.In addition customers were instructed to share this information with their laboratory staff and retain this notification as part of your laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that you have received this important notification.If customers need assistance or have any questions regarding the notification, they were instructed to contact Customer Technical Support Center at 1-800-854-3633 in the United States and Canada. | Contact the recalling firm for information | 10826 | 10826 | Nationwide In the US & Canada | Terminated | May 07, 2012 | 584 | 10826 | Not_Computer
Z-2657-2010 | 56565 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | AVI Sterile Impad Rigid Sole Foot Cover, Left FootProduct Code:6067 | Sterile ImPad Rigid Sole Foot Cover | 2 | October 01, 2010 | 2010 | Covidien LP | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any questions please contact, Covidien Customer Service at (800) 962-9888, Option 1. |  | 1436 | 1436 units (359 cases) | Nationwide | Terminated | June 12, 2012 | 620 | 3420 | Not_Computer
Z-2642-2010 | 56593 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Camera Head OTV-S7H-1D-L08E and Camera Head OTV-S7H-1D-F08E (designed to be used with Olympus endoscopic equipment) | Olympus VISERA OTVS7V Camera Head | 2 | October 01, 2010 | 2010 | Olympus America Inc. | Mislabeled. Endoscopic body-type camera heads labeled as cardiac-type. | PRODUCTION CONTROLS: Labeling Mix-Ups | Asked consignees to return camera heads to correct the labeling. |  | 217 | 217 camera heads | nationwide; camera heads were serviced and returned to consignees in March 2003 - September 2007 timeframe | Terminated | September 27, 2012 | 727 | 217 | Not_Computer
z-2640-2010 | 56652 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System, Model UST-2001, Manufacturer Miwatec Co. LTD, Tokyo, Japan. |  | 2 | October 01, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments acquired the Sonopet Ultrasonic Surgical System from Mutoh co. Ltd on 3/29/2010. Stryker determined that the Sonopet consode did not meet the IEC-60601-1 standard. Stryker decided to recall the Sonopet consode. |  | The firm to send URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to customers dated 7/15/2010. The customers were instructed to continue to use the device. Stryker technicians were dispatched to replace a board with a new board. |  | 1161 | 1161 | Wordwide | Terminated | October 19, 2011 | 383 | 1161 | Not_Computer
Z-0004-2011 | 56411 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | AortaScan AMI 9700 The AortaScan AMI 9700 is an ultrasound device that projects ultrasound energy into the mid-abdomen to obtain an image of the abdominal aorta for aortic diameter measurements | AortaScan AMI 9700 | 2 | October 04, 2010 | 2010 | Verathon, Inc. | The AortaScan AMI 9700 devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | DESIGN: Software Design | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to:"Add the enclosed addendum to their current Users Manual."Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 orcservice@verathon.com. |  | 73 | 73 units | Worldwide Distribution - USA and the countries of Australia, Canada, China, Columbia, European Union, United Kingdom, Saudi Arabia, and Singapore. | Terminated | September 08, 2011 | 339 | 684 | Software | Device Operation | Instructions for reset/Software update | Software Update
Z-0012-2011 | 49138 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Essenta DR is a digital multifunctional X-ray systems, suitable for all routine radiographic exams, including specialist areas like trauma or pediatric work, excluding mammography.It is designed for radiographic examination of the standing or seated patient or the recumbent patient in combination with a mobile X-ray table (trolley).The intended use is from direct digital imaging using the built in flat panel detector and in addition for free exposures on radiographic cassettes. | Essenta DR Xray system | 2 | October 06, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | When a patient is selected in the patient list and examination tab pressed with background query running at the same time, the wrong patient may be selected. | OTHER/UNDETERMINED: Pending | Philip Medical System sent an "URGENT-Device Correction" letter dated August 13, 2008 to their customers. The letter describes the recalled product as Essenta DR Patient and Image Data Software (SW) release 1.0. The firm informs the customers that when patient is selected in patient list and examination tab pressed with background query running at the same time, there is a risk that the wrong patient is selected on the Essenta DR Computed Radiography System.Customers are advised to verify the selected patient's information before releasing an exposure while the firm works on upgrading the Software. All affected units will receive a software upgrade to correct the issue.Customers should contact Philips Customer Care Service at 800-772-9377, #5, #4, #1 and reference FCO 712000031 for any questions or support concerning the recall. |  | 23 | 23 units | Nationwide distribution: CA, GA, FL, KS, MN, NY, OH, OR, SC, TN, TX, UT, and WA. | Terminated | October 07, 2010 | 1 | 23 | Software | Output/Calculation | Software update | Software Update
Z-0014-2011 | 50100 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee Distal Femoral Locking Screw for 5 mm Spacer; Product No. 6632-5-600.Made in Ireland;Howmedica Osteonics;325 Corporate Drive,Mahwah, NJ 07430Intended for primary or secondary reconstruction of the distal femur. | Duracon Total Knee Distal Femoral Locking Screw | 2 | October 06, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that a box labeled as Duracon Total Knee Distal Femoral Locking Screws for 5 mm spacer, may actually contain a 10 mm screw. | PRODUCTION CONTROLS: Labeling Mix-Ups | Important Market Withdrawal Letters were sent via Federal Express October 5, 2006. The letters identified the affected product and the reason for recall. Customers were asked to examine their inventory to identify the product; and retrieve all affected product lots and return it their branch or agency warehouse for reconciliation. All product should be reconciled on the Product Accountability Form and the form should be faxed back. Customers are to contact a Customer Service Rep to re-order the product that is being returned. Questions can be directed to Rita Intorrella on 201-831-5825. |  | 7 | 7 units | Nationwide Distribution -- Including Ohio. | Terminated | October 07, 2010 | 1 | 7 | Not_Computer
Z-0011-2011 | 56728 | tubing, pump, cardiopulmonary bypass | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | MAQUET HLM Tubing set with Bioline Coating (BEQ-TOP 21400 with 1/2 x 1/2 HG 0286 Connector; Device Part Number BEQ-TOP 21400 | MAQUET HLM Tubing Set with Bioline Coating with HG 0286 Connector | 2 | October 06, 2010 | 2010 | Maquet Inc. | Tubing packs may contain defective 1/2 inch x 1/2 inch tubing connectors. Leaks at the shank of the connector may result in air aspiration in the venous line. | PRODUCTION CONTROLS: Process Control | Maquet Getigne Group issued Product Recall/Urgent Medical Device Corrective Action letters dated September 7, 2010 to customers, identifying the affected product, the issue prompting the recall, and actions to be taken by customers.Customers are to check their inventory for affected product. Customers with affected product are to contact Laura Olilver at 201-995-8849 for an RMA number and shipment instructions. All firms are to coomplete and return the Corrective Action Response Form.Maquet Customer Service can bre contacted at 800-777-4222. |  | 22 | 22 items | Worldwide Distribution: USA State of IL and country of the Netherlands. | Terminated | September 25, 2012 | 720 | 22 | Not_Computer
Z-0013-2011 | 56026 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Summit Broach Handle, NON-STERILE, MFG: DEPUY ORTHOPAEDICS, INC., PO BOX 988 WARSAW, IN 46581, 1 800 366 8143Catalog Number: 2570-00-000The Summit Broach Handle (instrumentation) is used with the DePuy Summit Tapered Hip System. The Broach Handle is used to prepare the femur for hip system implantation. The broach handle is an impaction/extraction device that holds broaches for creating the cavity in the femur during hip arthroplasty. | Summit Universal Broach Handle | 2 | October 06, 2010 | 2010 | Depuy Orthopaedics, Inc. | Firm discovered that some of their straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring). | PRODUCTION CONTROLS: Process Control | On June 3, 2010, an "Alert: Femoral Broach Handle Repair" went out to the entire sales force/customers via "Sales Mail" and posted to the DePuy Internal website. The Alert stated that "We have discovered that some of our straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring).As you perform routine maintenance on your sets, if you have any of these broaches stamped with the lot numbers listed, it is recommended you have them sent in to be reworked with a new spring device."Send the broach handles to: DePuy OrthopaedicsAttn: Repairs700 Orthopaedic DrWarsaw, IN 46581Please include return shipping informationIf you have any questions regarding this information or need additional information, call 574-372-7333. | Contact the recalling firm for information | 8358 | 8358 | Nationwide distribution: AR, AZ, CA, FL, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, NC, NV, NY, OH, PA, TN, VA, and WI. | Terminated | November 30, 2011 | 420 | 8358 | Not_Computer
Z-0010-2011 | 56599 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Amingo Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland | GE DatexOhmeda Amingo Anesthesia System | 2 | October 06, 2010 | 2010 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety.If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila | TRAINING: Employee Error | GE Healthcare issued an " Urgent Medical Device Correction" letter dated August 30, 2010 to consignees. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager and Director of Biomedical/ Clinical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Instructions.Customers were instructed to:1. Ensure the plug is removed at the completion of the Low P Leak check.2. Review the Preoperative Checkout procedure listed in section 4 of the User Reference Manual with appropriate personnel. Ensure that the Preoperative Checkout procedure is completed in its entirety prior to each case, or as required.All facilities will be contacted by their local GE GE Field Technical Service Team to schedule installation of a software update and a revised test plug for your machines configured with SCGO. Customers are to contact their local GE Healthcare Customer or Technical Support Representative if they need additional information. |  | 131 | 131 | Worldwide Distribution: USA: All States, including DC, except Delaware, and the countries of YEMEN, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SYRIA, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI, ARABIA, RUSSIA, PORTUGAL POLAND, PERU, PANAMA, NORWAY, NETHERLAND, MACEDONIA, LITHUANIA, LEBANON, LATVIA, KUWAIT, KOREA, KENYA, JORDAN, JAPAN, ITALY, IRELAND, INDIA, HUNGARIA, HONG KONG, HONDURAS, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | September 13, 2012 | 708 | 7711 | Computer
Z-0022-2011 | 56707 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The BABIX holder is an accessory used with the DigitalDiagnost X-ray system. It is used to hold the baby either in a secured or immobilized standing position in front of the X-ray detector during the examination. | BABIX holder for DigitalDiagnost Xray system | 2 | October 07, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential of broken strands in the main steel cable supporting the accessory holder used with BuckyDiagnost and DigitalDiagnost X-ray systems. | PRODUCTION CONTROLS: Equipment Maintenance | Philips issued URGENT-Device Correction Notice letters dated September 9, 2010 to their consignees identifying the affected products and action to be taken by customers. The letter states that Babix holder used with Philips BuckyDiagnost and DigitalDiagnost X-Ray System - VS Wall Stand Risk of steel cable fracture in the support holder. The consignees are informed that there is a potential of broken strands at the main steel cable within the holder which cannot be recognized by the user. The firm will issue the Field Change Order (FCO) and the Service Engineer will replace the Babix holder.Consignees can call Philips Call Center at 800-722-9377, #5 (then follow prompts) and reference to FCO 70400038 for the BuckyDiagnost X-ray system and FCO 71200046 for the DigitalDiagnost X-ray system. | Contact the recalling firm for information | 1 | 1 unit | Nationwide Distribution in the states of CA, GA, and TN. | Terminated | July 22, 2011 | 288 | 5 | Not_Computer
Z-0020-2011 | 56536 | catheter, ventricular, general & plastic surgery | GBS | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | HydraGlide Silicone Chest Tube Catheter, 32 Fr StraightProduct Code 14132 | Atrium HydraGlide Silicone Chest Tube Catheter, 32 Fr | 2 | October 07, 2010 | 2010 | Atrium Medical Corporation | Mislabeled: One lot of 32Fr Straight Catheters packaged and labeled as 24Fr Straight Catheters. | PRODUCTION CONTROLS: Packaging Process Control | Atrium Medical isued Product Recall Notice letters dated August 10, 2010 to customers, identifying the affected product and actions to be taken by customers. Customers were instructed to remove the affected products from inventory and return them to Atrium, and complete and fax a recall notice acknowledgement to Atriuim at 603-821-1420. Customers can contact Atrium at 800 528-7486, x5350 (Joe DePaolo) or x5351 (Karen Hall) for further information. |  | 1 | 1 case, 10/case | Nationwide Distribution: USA, including the states of AK, CT, OH, and UT. | Open |  |  | 11 | Not_Computer
Z-0023-2011 | 56663 | needle, aspiration and injection, disposable | GAA | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | BD 30G1/2 PrecisionGlideÂŽ Needle;Catalog/Lot Number: 305106.Needles are intended for use in aspiration and injection of medications. | BD 30G1/2 PrecisionGlide Needle | 2 | October 07, 2010 | 2010 | Becton Dickinson & Company | There have been reports of blocked/clogged needles for the 30 gauge PrecisionGlide Needles |  | BD sent out Product Correction letters to all end users of the product on 8/24/10. Letters were mailed by BD via UPS second day air. The letter stated the issue and said that customers should assess the patency prior to injection. The letter also stated that a BD 30 gauge PrecisionGlide Needle from another lot can be used in place of a clogged needle. If customers require replacement needles or further assistance, they can contact BD at 1-888-237-2762. Finally, the letter states that BD has implemented corrective actions to reduce the potential for further recurrence of the blockage issue. |  | 2599000 | 2,599,000 units | Worldwide Distribution -- USA and Canada. | Terminated | August 25, 2011 | 322 | 2599000 | Not_Computer
Z-0017-2011 | 56630 | programmer, pacemaker | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular |  | Guidant, Vitality HE ICD, Model: T180DDR, Sterile Lot, Guidant Corporation, Cardiac Rhythm Management, 4100 Hamline Avenue North, St. Paul, MN 55112-5798. | Vitality HE Model T180 Implantable Cardioverter Defibrillator (ICD) | 2 | October 07, 2010 | 2010 | Boston Scientific CRM Corp | Exposure to a magnet (typically in a clinic/hospital environment) may cause a magnetic reed switch to become permanently stuck in a closed position in certain Boston Scientific CRT-Ds and ICDs (CONTAK RENEWALÂż 3 / 3 RF / 4 / 4 RF and VITALITYÂż HE implantable defibrillators). This prevents delivery of programmed tachy therapy, and use of the Patient Triggered Monitor feature. This issue may require | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Boston Scientific Issued Urgent Medical Device Correction and Physician Device Advisory Notice letters dated July, 2010 to customers identifying the affected devices, the issue prompting the recall, clinical implications, and actions to be taken by customers.Customers were instructed:1) Physicians should continue routine follow-up sessions, and patients should be reminded to contact their clinic or go to the hospital emergency room immediately if they hear tones/beeps from their device. 2) In a hospital/clinic/surgery setting, if tones are heard upon magnet application but do not cease upon magnet removal, the device should be interrogated with a programmer and checked per normal standard of care. 3) In the United States, use of the LATITUDE remote patient monitoring system may help identify loss of daily measurements and thereby facilitate timely detection of a stuck reed switch. [NOTE 1: A pop-up message and/or LATITUDE alert do not appear for missing Daily Measurements. NOTE 2: Daily Measurements may be unavailable for reasons other than a stuck switch.] 4) Prophylactic explant is NOT recommended.5) If a stuck magnetic switch is confirmed, program the Enable Magnet Use feature to Off, which ensures that programmed therapy will be provided to treat tachyarrhythmias. However, if Enable Magnet Use is programmed Off, a magnet will no longer inhibit tachy therapy, and the Patient Triggered Monitor feature will no longer be available. Do not routinely program Enable Magnet Use to Off in the absence of a confirmed stuck magnetic reed switch because the benefits of magnet use to disable tachy therapy in emergent situations outweigh the probability of patient harm associated with a stuck reed switch.In the US, Boston Scientific Cardiac Rhythm Management can be contacted at 800-227-3422, and in Europe, Boston Scientific Europe S.A. can be contacted at +32 2 416 7222. | Contact the recalling firm for information | 5821 | 5821 | Worldwide Distribution: All States in the continental USA including PR and DC , but excluding VT, the countries of Canada, Mexico, and in Europe.ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRITISH VIRGIN ISLANDS, CANADA, CAYMAN ISLANDS, CHILE, COSTA RICA, CYPRUS, CZECH REPUBLIC, CZECHOSLOVAKIA, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, PORTUGAL , QATAR, SAINT KITTS & NEOIS, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM,VENEZUELA, URUGUAY, ZIMBABWE, IRAN. | Terminated | April 03, 2012 | 544 | 38377 | Not_Computer
Z-0015-2011 | 56219 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Exodus(TM) 8F (2.7 mm)/25 cm Standard Loop Multipurpose Drainage Catheter with Hydrophilic Coating, REF Catalog No. 10002, UPN Product No. H965100201, Rx ONLY, STERILE/EO --- Made in USA: 2301 Centennial Blvd., Jeffersonville, IN 47130 --- Manufactured for: Navilyst Medical, Inc. 26 Forest Street, Marlborough, MA 01752 USA --- Each box unit contains 5 catheters. --- The Exodus Standard Loop Multipurpose Drainage Catheter consists of a radiopaque polyurethane catheter with a locking pigtail distal end and is coated with a hydrophilic coating. Catheter markings are provided in centimeters on the shaft to indicate the distance form the distal tip. The catheter is provided with a trocar, a metal stiffening cannula and a plastic cannula. Multipurpose Drainage Catheters are intended for percutaneous drainage of fluid in the chest, abdomen and pelvis. | Exodus Standard Loop Multipurpose Drainage Catheter | 2 | October 07, 2010 | 2010 | Navilyst Medical, Inc | One box containing five catheters did not complete the sterilization process prior to distribution. | TRAINING: Employee Error | On/about 5/24/10, NMI telephoned the consignee. The recall notification package was sent by NMI to the sole consignee on 6/11/10, followed by a telephone call by NMI Quality staff. On 6/15/10, the recall notification package was sent via FedEx to the sole consignee. Subsequently the consignee was provided a return goods authorization (RGA) number for return to NMI. The letter described the product, problem and action to be taken by the consignee/customer. The consignee/customer was instructed to immediately discontinue use of and segregate recalled product; remove the affected lot from their inventory; complete and return the Reply Verification Tracking Form via fax to the QA Manager; Navilyst Medical, Inc., at 1-518-742-4414; to please call the Navilyst Medical Customer Service at 1-800-833-9973 between 8:30 a.m. and 7:00 p.m. (Eastern Standard Time) to obtain a Return Authorization (RGA) Number, and package/ship the recalled product to Navilyst Medical, Inc, Distribution Center, Attn: QA Returns, 10 Glens Falls Technical Park, Glens Falls, NY 12801. If you have any questions, call Navilyst Medical Customer Service at 1-800-833-9973. Subsequent to the initial notification letter, NMI identified that a typo was made to the lot number on the letter. NMI telephoned the consignee and sent a revised letter via fax on 6/25/10 indicating the correct lot number. |  | 1 | 1 box unit (contains 5 catheters) | Nationwide distribution: Ohio only | Terminated | October 07, 2010 | 0 | 1 | Not_Computer
Z-0018-2011 | 56600 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Emergency and Transport Ventilator; Catalog #2M86300; Drager Medical Inc., 3135 Quarry Road, Telford, PA 18969. | Oxylog 3000 | 2 | October 07, 2010 | 2010 | Draeger Medical Systems, Inc. | Patients received insufficient ventilation. The Oxylog 3000 Instructions are inadequate relating to dead space volume. | PRODUCTION CONTROLS: Error in Labeling | The firm sent Important Safety Notice Letter, dated August 2010, and Instructions for Use supplement to consignees. The letter explained the issue with the affected product. The letter asks customers to attach the Instructions for Use amendment sheet supplied with the letter to their corresponding Instructions for Use. The firm also states that a they will provide a dedicated pediatric breathing circuit for the Oxylog 3000. Questions regarding the letter should be directed to Michael Kelhart at 1-800-543-5047. Questions regarding the use of the device should be directed to DragerService Technical Support at 1-800-543-5047 or your local Drager Sales Representative. |  | 78 | 78 potentially afftected units | Nationwide Distribution | Terminated | November 18, 2010 | 42 | 78 | Not_Computer
Z-0034-2011 | 50215 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The device's marketing brochure states: "IntelliVue M3150 Information Center...Philips Medical Systems".The beige-colored PAS-210 external speakers can be identified by their use of DB15-to-3.5 mm mini-phone jack cable (Agispec-081Y-4 or Agispec-081-6) to connect to the rp5700 PC.Indicated for central monitoring of multiple adult, pediatric, and neonatal patients. | Philips IntelliVue Information Center | 2 | October 08, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential for Internal Sound Amplifier to malfunction which could delay treatment. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | On 11/20/08, Philips began sending the URGENT-Medical Device Correction letter to their consignees via e-mails. On 11/21/08, the firm sent out the recall letter via United State Postal Service.The letter states the affected device as Philips IntelliVue Information Centers, running on the HP rp5700 PC platform equipped with an active speaker solutions using a DB-15 to 3.5 mm cable set.The consignees are advised to ensure that the audible alarm system is clear, free of noise and distortion and adjustable to different sound levels which is explained in p. 5-47 "Adjusting Alarm Tone Volume" of IntelliVue of the Instruction for Use, Release K Manual, supplied with the hardware, Philips Part Number 453564062301.Consignees can call Philips Healthcare Customer Care Service Center at 1-800-722-9377, #3, #1 and reference FCO 99862012. |  | 424 | 424 units in US. | Nationwide Distribution | Terminated | October 12, 2010 | 4 | 424 | Hardware | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0026-2011 | 54527 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SpeedLink 5.5 Transverse Connector Short 39.5-44mm, non-sterile, 1302-55S | SpeedLink 5.5 Transverse Connector | 2 | October 08, 2010 | 2010 | Zimmer Inc. | The firm received complaints that the device was damaged or fractured while the surgeon was bending it prior to implantation. There were also complaints that the cam may dislodge during tightening. The firm added additional instructions for use to reduce the probability of these two types of events. | OTHER/UNDETERMINED: Pending | The firm provided a revised surgical technique document to the Zimmer spine sales force in January, 2009. The firm will provide an Urgent Medical Device Correction notification to speedlink users through the sales force. | Contact the recalling firm for information | 23205 | 23,205 Units for ALL PRODUCTS IN THIS RECALL | The devices are contained in implant cases stocked by Zimmer spine and made available to the Zimmer sales force for surgeries. They are not in general distribution. | Terminated | January 05, 2011 | 89 | 23205 | Not_Computer
Z-0025-2011 | 56767 | device, percutaneous, biopsy | MJG | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Tru-Core II Disposable Automatic Biopsy Instrument and Tru-Core II URO Disposable Automatic Biopsy Instrument. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recalling lot's for the product are as follows:Model Number: 763114100XLot #'s:92661WY5, 93211XXV, 00141YVE, 00541ZKX, 010920HL, 01522192, 92791X5I, 93271Y1W, 00201YYE, 00621ZQ3, 011320LH, 015221BD, 92791X7X, 93411Y9O, 00211Z0C, 00671ZRN, 011720Q, 015521CL, 92861XCT, 93421YBJ, 00251Z1T, 00761ZZP, 012420U0, 016221H2, 93001XLG, 93431YBI, 00351Z86, 0081201Y, 0138211F, 016721KJ, 93081XPC, 93571YJO, 00461ZE2, 0088206B, 0141215R, 93141XST, 00061YOV, 00541ZK0, 009620BO, 0146219M.Model Number: 763114160XLot #'s92661WY7, 00111YPZ, 00671ZRO, 013220Z7, 016221H7, 93131XLS, 00331Z61, 0110201Y, 014125Q, 00882065, 00391Z8Y, 013120YC, 015521CJ.Model Number: 763114200XLot #'s92661WY0, 00751ZYJ, 012420U1Model Number: 763116100XLot #'s92791X80, 93491YG9, 00331Z6M, 0083203O, 012420U3, 01672KL, 92791X81, 93501YGA, 00351Z8D, 009620BL, 013120Y9, 92931XDD, 00121YRT, 00431ZE1, 009920D3, 015221AX, 93001XLN, 00131YUG, 00541ZKV, 011020IU, 015221BE, 93081XPH, 00541ZJY, 00601ZOT, 011320LG, 015521CK, 93291Y43, 00251Z1R, 0081201W, 012020QC, 016221H3.Model Number: 763116160XLot #'s92661WY3, 00882068, 00351Z8A, 00761ZZQ, 012020QB, 016721KK, 92791X89, 00912084, 00471ZG0, 010320FM, 012420TZ, 92861XD4, 01472193, 00541ZK1, 010920HK, 0141215U, 93001XLR, 00151YVC, 00601ZOR, 0110201V, 015521CM, 93291Y40, 00331Z6J, 00701ZUX, 011720OP, 016221HSModel Number: 763116200XLot #'s121KNModel Number: 763116200XLot #'s92671X09, 00471ZG1, 00882067, 010320FN, 00211Z09, 00601ZOP, 009920D0, 011020ITModel Number: 763118100XLot #'s92661WY8, 93491YG7, 00331Z6L, 00621ZQ2, 009920D2, 0138213F, 92791X7Z, 93491YG8, 00351Z88, 00701ZUW, 010320FL, 01472194, 92861XDS, 00061YOT, 00471ZG2, 00761ZZO, 0110201W, 016221H4, 93091XPL, 00131YUF, 00541ZJZ, 0083203P, 012020QE, 93211XXW, 00251Z1S, 00541ZKU, 009620BM, 012420TX.Model Number: 763118160XLot #s92661WY6, 93271Y1V, 00201YYD, 00471ZG3, 012420U2, 016221GZ, 92671WZJ, 93411Y9P, 00251Z1Q, 00671ZRK, 013120YB, 92671WZK, 93421YBK, 00331Z6N, 00701ZUY, 013220Z9, 92861XD0, 00061YOW, 00351Z8E, 00882069, 015221BB, 93001XLM, 00111YQO, 00431ZE3, 009920D4, 015521CO.Model Number: 763118200XLot #s92861XCD, 93571YJN, 00331Z6K, 00671ZRM, 011320LF, 01382134, 93001XL6, 0061YOU, 00351Z89,00701ZUZ, 011720OO, 0141215S, 93001XLJ, 00912082, 00391Z8Z, 009620BQ, 012020QA, 01472195, 93081XPJ, 00111YQM, 00471ZG5, 009920D5, 013120Y8, 016221H1, 93211XXX, 00131YUD, 00481ZG4, 010320FJ, 013220Z6, 93400Y9K, 00201YYC, 00621ZQ1, 010920HJ, 0137211I.Model Number: 763120100XLot #s93001XLV, 00351Z8B, 0088206A, 01472196, 00331Z60, 00601ZOS, 01202QF.Model Number: 763120160XLot #s92661WY1, 00882066, 00111YQP, 00351Z8C, 013120YD, 01472197, 92861XDG, 00912083, 00211Z08, 00621ZQ5, 013220Z8.Model Number: 763120200XLot #s00351Z85, 012020QG, 012420TW.Model Number: 763418200XLot #s92671WZV, 93291Y41, 00351Z87, 00601ZNH, 010320FK, 0141215T, 92671WZW, 00061YOX, 00431ZE0, 0083203N, 0110201S, 015221BC, 92861XDH, 00131YUE, 00491ZIO, 009620BN, 0011720ON, 015521CN, 93081XPR, 00141YVB, 00541ZKT, 009920D1, 0138213E, 016221H6.Model Number: 763418250XLot #s92751X60, 00671ZRL,Model Number: 763418250XLot #s92661WY4, 93271Y1N, 00141YVD, 00601ZOQ, 0110201X, 013220ZA, 92751X6P, 93291Y42, 00211Z0B, 00761ZZR, 012020QD, 0137211G, 92781X5G, 93631YJP, 00221Z13, 0081201X, 012420TY, 01382133, 92861XDN, 00912085, 00431ZDZ, 009620BP, 013120YA, 016221H0. | TruCore II Disposable Automatic Biopsy Instrument and TrueCore II URO Disposable Automatic Biopsy | 2 | October 08, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc., dba Angiotech, Gaineville, Florida is recalling Tru-Core II Disposable Biopsy Devices due to compromised packaging integrity. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recall is for the following Model Numbers: 763114100X, 763114160X, 763114 | OTHER/UNDETERMINED: Pending | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated July 2, 2010 and response card to all customers including distributors. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review their current inventory and segregate the affected lot numbers; arrange for return and replacement of the affected product, and complete and return the Customer Acknowledgment Form via fax to the Quality Assurance department at 1-352-338-0662 or 1-800-333-0440.If you have any questions concerning this notification, contact Shannon Brooks at (352) 338-0440 ext 355 or Katrenia Williams at ext 353. | Contact the recalling firm for information | 11123 | 11,123 boxes (111,230 units) | Worldwide distribution: USA and countries including: Argentina, Australia, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Israel, Japan, New Zealand, Peru, South Africa, , Uruguay, and Venezuela. | Terminated | June 28, 2011 | 263 | 11123 | Not_Computer
Z-0024-2011 | 56645 | detector, bubble, cardiopulmonary bypass | KRL | Cardiopulmonary bypass bubble detector. | Cardiovascular | Cardiovascular | 510(k) | Air bubble detection cable assembly, catalog numbers 78-8067-7375-6 and 149892, Terumo Cardiovascular System, Ann Arbor, MI | Sarns Ultrasonic Air Sensor | 2 | October 08, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Malfunction of the air bubble detection cable assembly causes a false air alarm or a check sensor message during system set-up. In some instances, users were not able to reset the alarm to allow for further air bubble detection. . | OTHER/UNDETERMINED: Pending | The firm sent URGENT MEDICAL DEVICE CORRECTION notices dated September 17, 2010. The notices told consignees that a Terumo representative would visit, inspect the cables and replace the recalled cables. if the customer has a service contract, the visits would occur at the next scheduled preventative maintenance visit. Customers without a service contract will be contacted by Terumo representative to schedule an appointment.Customers can contact Terumo Customer Service at 800-521-2818. | Contact the recalling firm for information | 5000 | 5,000 estimated | Worldwide Distribution: USA states and territories of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Belgium, Canada, Columbia, Chile, Costa Rica, Germany, Dominican Republic, Ecuador, Hong Kong, Indonesia, Japan, Korea, Mexico, Malaysia, Panama, Philippines, Saudi Arabia, Singapore, South Africa, Thailand, Taiwan, UAE and Vietnam. | Terminated | June 29, 2012 | 630 | 5000 | I/O | Alarm/Message | Replace cables | Remove or Replace
Z-0040-2011 | 55602 | accessories,arthroscopic | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith & Nephew Aimer Endofemoral 3mm Offset (Blue)Catalog Number: 72201713ACUFEX Surgical Instruments are intended for use in orthopedic and general surgery. | Smith & NephewAimer Endofemoral | 2 | October 12, 2010 | 2010 | Smith & Nephew, Inc. Endoscopy Division | Endofemoral Aimer tip may separate due to failure of the soldered joint | DESIGN: Component Design/Selection | Smith & Nephew sent a "Medical Device Recall Information" letter dated April 29, 2010, with an attached "Urgent-Product Recall 1st Notification - Urgent" response form via Federal Express on April 29, 2010, to all customers. An e-mail was sent to the contacts in the identified OUS countries alerting them of the recall. ( Note: On June 24, 2010, the firm expanded the recall to include 4 additional catalog numbers 7220l713F, 72201714F, 72201715F, 722017l6F that are refurbished and one additional lot number to catalog Number: 72201714.) The letter described the product, problem and action to be taken by customers. The customers were instructed to immediately check stock for affected product; immediately remove the items from stock and set aside for quarantine; complete and return the attached Urgent-Product Recall 1st Notification - Urgent" response form, even if no product is being returned, via fax at +1-508-261-3636 attention Cindy Burns or email a PDF to Cindy.Burns@smith-nephew.com; call for a return Authorization number at 1-800-343-5717 Option #3, and return product to Smith & Nephew, Inc., Endoscopy Division as soon as possible. If you have any questions or require additional information, please contact Cindy Burns at 1-508-261-3655. |  | 577 | 577 units | Worldwide distribution: USA and countries including: Australia, Austria, Belgium, Brazil, Canada, Costa Rica, Denmark. Dubal, Finland, France. Germany, Greece, Hong Kong, India, Italy, Japan, Malaysia, Netherlands. Norway, Portugal, Puer To Rico, Singapore. South Africa, Spain. Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 575 | 3607 | Not_Computer
Z-0036-2011 | 55802 | device, pasteurization, hot water | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD SYSTEM 520 is a two-tank Medical Device Cleaning and High Level Disinfection Systems (Washer/Pasteurizers). Each unit has a metal serial tag that is riveted to the enclosure of the unit and it is labeled in part: "HLD SYSTEMS 520***FOR SALES.SERVICE.SUPPLIES CALL***Cenorin*** Kent, WA 98032***P:800.426.1042***F: 253.395.2650***www.cenorin.com".Medical Device Cleaning and High Level Disinfection Systems (Washer/Pasteurizers) are intended to use at healthcare provide clinical areas such as sleep labs, central services, sterile processing, respiratory care departments. The system is used to clean and high level disinfecting medical devices. They are designed to automate the pre-washing and rinsing of critical devices prior to sterilization in other devices. | HLD SYSTEM 520 | 2 | October 12, 2010 | 2010 | Cenorin | The water level sensor in the Medical Device Cleaning and High Level Disinfection (Washer/Pasteurizer) is sensitive to externally induced vibration noise, and it produces a false signal that the water level in the tank is full while it is only 75% full. It could present a health risk of biological hazard because the medical devices are not being adequately cleaned or pasteurized with enough wat | DESIGN: Component Design/Selection | On June 7, 2010, firm, CENORIN, sent "URGENT: MEDICAL DEVICE RECALL" Initial Notification letter with Recall Response Form, Series 500 Water Level Check Instructions, and the Water Level Switch Upgrade Kit Instruction (service kit P/N # 082395-00) to their consignees. The firm also sent these documents in Spanish language to their foreign consignees. The letter describes the recall product as HLD Systems Medical Device Cleaning and High Level Disinfection Series 500 Washer/Pasteurizer-Model 540, Model 540HT and Model 520. Customers are advised to call the Customer Service at 1-800-426-1402 or e-mail the firm.The Initial Notification letter instructs consignees to conduct and quarantine the recall products and notify their sub-accounts (customers) if products have been further distributed. While waiting for a replacement kit, the consignees are asked to follow the Series 500 Water Level Check Instructions to perform a test to ensure the water tank is completely filled up the expected water level. The firm also advised the consignees to add a cautionary label on the product to avoid any sharp or bump impact that may create any externally induced vibration noise to the device.The firm is developing a replacement kit that should be ready to send out to their consignees by June 30, 2010. A response card would be also sent out with a new kit, and it is used to verify that the kit has been successfully implemented. If you have any questions, please email jbennett@cenorin.com or call 253-395-2400 or toll free: 1-800-426-1042. | Contact the recalling firm for information | 9 | 9 units | Worldwide distribution:USA and countries including: Costa Rica, Ecuador, Japan, Mexico, and Taiwan. | Terminated | May 07, 2012 | 573 | 93 | Not_Computer
Z-0002-2011 | 56425 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product codes: 2M8153, and 2M8163. | Colleague Triple Volumetric Infusion Pumps | 1 | October 12, 2010 | 2010 | Baxter Healthcare Corp. | The FDA sent a letter to Baxter on April 30, 2010, ordering the company to recall and destroy all models of its Colleague Volumetric Infusion Pumps currently in use in the United States. FDA determined that this action is necessary, as Baxter has failed to adequately correct, within a reasonable timeframe, the deficiencies in the Colleague pumps still in use. This recall is a consolidation of al | DESIGN: Device Design | Baxter Healthcare Corporation sent "Urgent Product Recall" letters dated August 4, 2010 to all Colleague Infusion Pump customers in the U.S., to the attention of the Hospital Administrator, Chief Executive Officer, Chief Operating Officer, Chief Medical Officer and Chief Nursing Officer, referencing Baxter's May 3, 2010 announcement to recall all Colleague pumps from the U.S. market pursuant to an order under its existing consent decree with the FDA. The letters provided the customers with important information regarding the actions they must take prior to November 14, 2010 in order to continue using their Colleague pumps during the transition period. Customers were instructed to complete the attached Certificate of Medical Necessity Form and return it to Baxter as soon as possible, but no later than November 14, 2010, by scanning and e-mailing to Baxter at FCA@Baxter.com. Any questions were directed to their local Baxter Sales Representative or by calling the Center for One Baxter 1-800-422-9837, between 8AM and 5 PM Central Standard Time (CST).*** On July 20, 2011 Baxter sent recall letters to those Colleague customers who had not completed a decision form regarding the status and replacement of their Colleague Pumps. The accounts were asked to inform their Baxter sales representative as soon as possible of their pump decision, and to complete the attached customer reply form and return it to Baxter via fax or scanned e-mail. ***. |  | 81775 | 81,775 pumps | Nationwide, including Puerto Rico | Terminated | November 26, 2012 | 776 | 302872 | Not_Computer
Z-0006-2011 | 56257 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CareFusion Alaris PC Unit Model 8015 (formerly MedleyPC Unit): Alaris PC Units (Model 8015) with logic board part number TC10005522 manufactured and serviced between December20, 2008 and May 6, 2009; and Alaris PC Units (with logic board part number TC10005572 manufactured and serviced between May 7, 2009 and September 8, 2009. For use with Alaris infusion and monitoring systems. | Alaris | 1 | October 13, 2010 | 2010 | CareFusion Corporation | Under certain wireless network conditions a communication error can occur, which freezes the PC Unit screen, which may result in a delay of therapy. A delay of therapy may result in serious injury and/or death. CareFusion identified this potential risk associated with the Alaris PC unit model 8015 only (PC Unit) manufactured between December-2008 and September-2009. | DESIGN: Component Design/Selection | CareFusion sent the customer notification letter to all customers using the affected Alaris PC Unit beginning August 24, 2010. The customer notification letter is addressed to the Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required to confirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, self-addressed mail, fax or email. The recall notice informed the customers of the reason for the notice, the units affected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion will contact their facility through phone or in person within 60 days to initiate implementation of the corrective action. The corrective action will require a hardware update to all affected units. CareFusion does not require that you return your devices. In the interim, if you observe this issue, remove the device from service and contact CareFusion Recall Center immediately.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems should also be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. | Contact the recalling firm for information | 17081 | 17,081 units (17,050 in US) | Nationwide and worldwide to: Canada, Netherlands, New Zealand and Singapore | Terminated | December 12, 2011 | 425 | 17081 | I/O | Display/Image |  | Other
Z-0050-2011 | 56680 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio/ NRG 4in1 Cutting BlockThe Scorpio total knee system is indicated for primary Total knee Arthoplasty (TKA) procedures. The Scorpio Femoral Component is implanted on femur after making bone cuts to accommodate the shape of the implant. The Scorpio instrumentation includes a 4 in 1 Cutting block 98010-00xx) which serves as a guide to allow the surgeon to make the bone e cuts necessary for seating the femoral component (implant) in place. The Scorpio Cutting blocks (8010-00xx) are designed to align with two previously located holes and for seating on the distal surg ace of the femur, and provide guided slots for the sawblades to make bone cuts. | Scorpio/ NRG 4in1 Cutting Block | 2 | October 13, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker s has received reports in which the Cutting Blocks have fractured on impaction during surgery. | DESIGN: Device Design | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated August 30, 2010, to all consignees/customers (branches/agencies, hospital risk management, hospital Chief of Orthopaedics, surgery) including foreign countries. The letter described the product, problem, and action to be taken by customers. The customers were instructed to complete and return the attached PRODUCT RECALL ACKNOWLEDGMENT FORM via fax to 201-831-6069.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance (201) 972-2100. |  | 839 | 839 (total) | Worldwide distribution: USA and countries including: Asia, Canada, Europe, Middle East and Africa | Terminated | August 30, 2012 | 687 | 839 | Not_Computer
Z-0058-2011 | 56694 | ventricular (assisst) bypass | DSQ |  |  | Cardiovascular |  | HeartMate II System Controller, a component of the HeartMate II Left Ventricular Device System (LVAS), distributed both as stand alone devices and as components of the LVAS Implant kit. Manufactured/Distributed by Thoratec Corporation, Pleasanton, CAThe HeartMate II Left Ventricular Assist System (LVAS) consists of an implantable blood pump connected to an eternal system controller by a percutaneous lead. The external controller is powered bye either batteries or a power supply that connects to AC Main power. The device is intended for use as a bridge to transplantation in cardiac transplant candidates at risk of imminent death from non-reversible left ventricular failure. It is also for use in patients with New York Heart Association class IIIB or IV end-stage left ventricular | HeartMate II System Controller; component of the HeartMate II Left Ventricular Device System (LVAS) | 2 | October 13, 2010 | 2010 | Thoratec Corp | Connection failure-- product may malfunction because of bent pins or broken wires in power cord connection |  | Thoratec Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 30, 2010, to all customers. The letter described the affected product, description and symptoms of problem, and immediate action to be taken by the customers. The customers were instructed to:1) contact all ongoing patients using the affected Systems Controllers (serial numbers starting with "EPC") to determine if they are experiencing any unexplained intermittent alarms or have any bent pins in the connectors between the System Controller and the power sources, if so, have them return to the hospital so that their power leads can be examined. If further assistance is needed with the evaluation, please contact Thoratec Technical Service (800-456-1477).2) review the attached addendum to the HeartMate II labeling with all VAD clinicians and ongoing HeartMate II LVAS patients, and3) complete and sign the attached Acknowledgment Form via fax to Thoratec Regulatory Affairs at (952) 847-8571.If you have any questions, please contact 925-847-8600. | Contact the recalling firm for information | 8965 | 8965 units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX , UT, VA, WA, and WI; and countries including: Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Netherlands, Singapore, Slovenia, Sweden, Switzerland, and the UK. | Terminated | April 07, 2011 | 176 | 8965 | Other
Z-0060-2011 | 56651 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular metal Modular Acetabular system 56 mm Porous shell with cluster holes, Catalog 00-6202-056-22, Zimmer Inc., Warsaw, IN.This device is indicated for primary or revision surgery for rehabilitating hips damaged as the result of noninflammatory degenerative joint disease, including osteoarthritis, fused hip, avascular necrosis, traumatic arthritis and fracture of the pelvis. | Zimmer Trabecular metal Modular Acetabular system | 2 | October 14, 2010 | 2010 | Zimmer Inc. | The threads on the device were not completely machined and therefore would not mate with athe corresponding cup positioner/shell inserter. | OTHER/UNDETERMINED: Pending | The firm, Zimmer, made telephone calls using a phone script document entitled "Urgent: Medical Device Recall" Phone Script Call Communication, to the customers that received this product starting on 8/18/2010. The telephone calls were to determine the status of the recalled product. A Zimmer Field Action Strategy notification dated August 9, 2010 was also provided. The notifications/phone script described the product, problem and actions to be taken by the customers. The customers were instructed to remove the 00-6202-056-22, lot 61091653 product from their inventory and return to Zimmer Product Service Department, Attn: Product Service, 1777 West Center Street, Warsaw, IN 45680. Note: If the product had been used, no action was necessary. The returned product will be evaluated to determine how many of the devices had the defect.If you have any questions, please call (574) 371-8852. | Stop use and contact the recalling firm for information. | 24 | 24 | Worldwide distribution:USA including states of: MI, IN, WI, MO, CA, VA, MD, TX, NC, PA, OR, WA, FL, GA, NY, ME, AZ, and UT and country including Canada. | Terminated | February 07, 2011 | 116 | 24 | Not_Computer
Z-0059-2011 | 56589 | pusher, socket | HXO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Closed knot pusher REF 902813, Biomet Sports Medicine, Warsaw, IN.The device is intended to push knots down a limb of suture. | CLOSED KNOT PUSHER | 2 | October 14, 2010 | 2010 | Biomet, Inc. | The firm has received four complaints that the closed knot pusher is missing a radius which contributes to sutures being cut during surgery. |  | Biomet Sports Medicine sent URGENT MEDICAL DEVICE RECALL NOTICES dated 8/2/2010 to its customers and distributors. The notices stated that the use of the device should be discontinued and the products returned to Biomet.Healthcare professionals and consumers were to report serious adverse events (side effects) or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online at www.fda.gov/MedWatch/report.htm, by regular mail to MedWatch, 5600 Fishers Lane, Rockville, MD 20852-9787 or by phone at (800)FDA-1088 or Fax (800)FDA-0178.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 101 | 101 | Worldwide Distribution - USA, including FL, NY, VA, CA, GA, AL, PA, MI, LA, KS, SD, KY, TX, AR, and WV, and the countries of Belgium, Italy, Argentina, Poland, Netherlands, Turkey and Portugal. | Terminated | August 20, 2012 | 676 | 101 | Not_Computer
Z-0071-2011 | 55157 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Continuum Acetabular Systems Shell Inserter Adapter with Rotational Control, 00-8790-003-00, Zimmer Inc., Warsaw, IN.Used to mate the inserter with the shell implant. | Continuum Acetabular Systems Shell Inserter Adapter with Rotational Control | 2 | October 15, 2010 | 2010 | Zimmer Inc. | The firm has received reports of users encountering difficulties in usage leading to surgical delays or minor damage to a threaded surface of the shell implant. | DESIGN: Device Design | The firm sent Urgent: Device Removal letters dated 3/15/2010. The letter identified the affected product, described the issue, discussed the clinical implications, and provided a list of required actions. The consignees were instructed to perform an inventory of the recalled product and report the amount on hand to Zimmer by completing the Inventory Return Certification Form. After Zimmer received the report, they would send out new product to the consignee. After new product was received the consignees were instructed to return the recalled product to Zimmer. The firm expects to have all product replaced by 4/30/2010. If there are any questions or additional info needed, consignees should contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 1125 | 1125 all products | Worldwide Distribution -- US, Singapore, Australia, Portugal, Germany, Switzerland, Spain, Finland, United Kingdom, Israel, Italy and Sweden. | Terminated | January 12, 2011 | 89 | 1125 | Not_Computer
Z-0068-2011 | 56719 | enzyme immunoassay, tracrolimus | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Architect Tacrolimus Reagent Kit; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of tacrolimus in human whole blood; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 1L77-25The Architect Tacrolimus assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of tacrolimus in human whole blood as an aid in the management of liver and kidney allograft patients receiving tacrolimus therapy. | Architect Tacrolimus Reagent Kit | 2 | October 15, 2010 | 2010 | Abbott Laboratories | There are reagent barcode read errors when using reagent lot 86599M500 due to the print quality of the barcode. | OTHER/UNDETERMINED: Pending | Abbott Laboratories sent a "Product Recall Immediate Action Required" letter dated September 7, 2010, to all customers. The letter described product, problem and action to be taken by customers. The customers were instructed to determine if they are currently using and/or have inventory of lot 86599M500; discontinue use of and destroy any remaining inventory; inform and provide copies of letter to other laboratories if appropriate; retain the letter for laboratory records and complete and return the Customer Reply Immediate Action Required form via fax at 1-800-777-0051 or email QAGCO@abbott.com. IMPORTANT NOTE: If you do not currently have replacement material than you may continue to use this reagent lot until your laboratory receives replacement material. Upon receipt of the replacement material, discontinue use of and discard lot 86599M500 and switch to the replacement material. Contact Abbott Customer service for replacement material (1-877-4ABBOTT, Option 1).If you or any of the health care providers you serve have any questions regarding this information, U.S. customers should call Customers Service at 1-877-4ABBOTT. Customers outside of the U.S., please contact your local Customer Service representative. |  | 1995 | 1,995 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Lithuania, Montenegro, Netherlands, New Zealand, Norway, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | May 13, 2011 | 210 | 1995 | Not_Computer
Z-0070-2011 | 56699 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 83" (2.1 m), 150 mL Burette Drip Chamber Filter Valve, 2 Injection Sites, Male Luer Lock Adapter, Large Bore 4-Way Stopcock with Rotating Male Luer Lock Extension Set, 38" (97 cm), Vol. 5.8 mL, Injection Site, Male Luer Lock Adapter, 60 DPM; a sterile Rx IV fluid pathway device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 3C0166The administration of fluids from a container into the patient's vascular system through a vascular access device | Baxter Interlink System Buretrol Solution Set | 2 | October 15, 2010 | 2010 | Baxter Healthcare Corp. | There are leaks in the Drip Chamber Check Valve Subassembly of the Solution Set. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. | TRAINING: Employee Error | Baxter sent an "Urgent Product Recall" letter dated September 2, 2010, via first class mail to the direct account, to the attention of the Director of Materials Management, with a copy sent to the Director of Nursing, on the same date, informing them that leaks were detected in the Drip Chamber Check Valve Subassembly used in the affected lot of Interlink System Buretrol Solution Sets. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. The account was requested to examine their inventory in all locations to determine if they have any of the affected lot and remove the affected product from inventory; call Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 948 | 948 units | Nationwide distribution: Wisconsin only. | Terminated | February 25, 2011 | 133 | 948 | Not_Computer
Z-0069-2011 | 56278 | medical computers and software | LNX |  |  | General Hospital | 510(k) | Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.02.40.1, and 4.0.2.41.1.SCC Soft Computer, 5400 Tech Data Drive, Clearwater, FL 33760. Manufacture Dates:4.0.2.38.1 06/15/07 4.02.39.2 06/29/07, 4.02.40.1 07/19/07, 4.02.41.1 08/20/07. The Result Reporting Interface is a highly versatile interface designed to electronically convey clinical laboratory orders and results for the SCC applications to other recipient information systems. | Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.02.40.1, | 2 | October 15, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.0.2.40.1, and 4.0.2.41.1 on 10/05/2007. A software anomaly in the result reporting interface caused result comments entered after an empty line | DESIGN: Software Design | SCC Soft Computer sent a RISK-TO-HEALTH NOTIFICATION TASK VERBIAGE dated October 5, 2007, through their Task Management System to all affected clients on October 17, 2007. The notice identified the product, the problem and the action clients should take to correct the problem. Clients were to acknowledge receipt of the task notice and to identify if the Workflow or functionality is not used and if they Agree to use alternative working solutions. Each client was asked to grant permission to load the required software to correct the issue.For any questions regarding this recall call (727) 789-0100. | Contact the recalling firm for information | 35 | 35 | Worldwide Distribution - USA, including AL, CO, FL, GA, IN, KS, KY, MA, MO, NC, NH, NJ, NY, OH, PA, TN, and WA and the country of Canada | Terminated | October 19, 2010 | 4 | 35 | Software | Output/Calculation | Software update | Software Update
Z-0066-2011 | 56756 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Control Syringe, 12ml, CAT No: CCS201E Version B, Sterile R, Merit Medical Systems, Inc.Used for hand injections of contrast medium during Coronary or Peripheral Angiography and Angioplasty. They are intended to be used as sterile, disposable, single use syringes. | Coronary Control Syringe | 2 | October 15, 2010 | 2010 | Merit Medical Systems, Inc. | Small holes in the packaging may render the product non-sterile. | TRAINING: Employee Error | Merit Medical sent an Urgent Product Recall Notice dated September 9, 2010, to all sales representatives. The notice identified the product, the problem, and the action to be taken by the sales representatives and their customers. Sales representatives were instructed to contact all customers via telephone and instruct them to immediately discontinue use and quarantine any unused product. Recall Notification Forms were to be completed and signed by the site representative and the customer and faxed to Merit customer service at 801-208-3378. If customers were returning affected products they were to place the completed original Recall Notification Form in the box with the products to be returned, write the RGA# on the shipping box and return by Federal Express 2nd Day air to: RGA DepartmentMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416For any questions regarding this recall call 801-316-4822 or 801-208-4365. |  | 199 | 199 units | Worldwide Distribution - USA, including IL, PA, WI and WV and the countries China, Greece, Hong Kong, Japan, New Zealand and Spain. | Terminated | March 21, 2011 | 157 | 16574 | Not_Computer
Z-0067-2011 | 56275 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Releases 3.17 through 4.1.0.SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760.Manufacture/Distribution Dates:Version 3.17.5 10/14/2003. Version 3.17.6 03/28/2003. Version 3.17.7 07/02/2003. Version 3.17.8 03/03/2004. Version 4.1.0 07/03/2003. Version 4.1.1 10/03/2003. Intended for use in the Pathology Lab for receiving and accessioning specimens, documenting tissue processing, producing material labels, generating final diagnostic reports, and creating billing and management reports for surgical, GYN, NON-GYN, and Autopsy cases. | SoftPath GUI Releases 3.17 through 4.1.0. | 2 | October 15, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for SoftPath GUI Releases 3.17 and 4.1 in March 2007 for clients with high volume output of final reports who use the Batch Print for Final Reports in the Scheduler, there is a potential for some reports not printing. This issue was corrected for the one affected user in October 2003. | DESIGN: Software Design | SCC Soft Computer sent a SAFETY NOTIFICATION TASK VERBIAGE dated January 5, 2004, to one client. The Notice identified the product, the problem, and the action to be taken by the client. SCC recommended the client set the Search Date Range to a small enough range to assure that the number of cases scanned during the job would be less than 10,000.The client was asked to acknowledge receipt of the task and to indicate if the Workflow or functionality was used/not used and Agree to use the alternative working solution. For any questions regarding this recall call (727) 780-0100, ext. 4235. | Contact the recalling firm for information | 84 | 84 | Nationwide - USA, including the states of CA, CO, CT, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, PA, SD, TX, WI, and WV | Terminated | October 19, 2010 | 4 | 84 | Software | Output/Calculation | Software update | Software Update
Z-0075-2011 | 54550 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Pathfinder End Screw Extender Sleeve, REF 1165-10, Rx, non-sterile, sold under label Abbott Spine, Inc. Austin TX (prior to 2008) and Zimmer Spine Austin TX (7/2009 to present) | PathFinder end screw extender sleeve. | 2 | October 18, 2010 | 2010 | Zimmer Inc. | The tip of the PathFinder End Screw Extender Sleeve has fractured during use. In some cases, this instrument fracture led to surgical delay or additional surgical steps necessary to remove the fragments of the broken tip. | DESIGN: Device Design | Zimmer Spine issued an "Urgent Medical DeviceCorrection" letter through the Zimmer Spine sales force to be provided to PathFinder system users in February 2010. the letter identified the affected product, safety issue, and actions to be taken by customers. Customers will receive replacement devices with a design change for the recalled devices, as they are available from the manufacturer. |  | 1471 | 1471 (888 subject to removal) | Worldwide Distribution: USA including states of AZ, CA, CO, FL, GA, LA, MI, MN, MO, NC, NH, NY, OH, OK, PA, TN, TX, WA, and WI, and the countries of Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, France, Germany, Greece, Italy, Japan, Mexico, Thailand, Turkey, UK and Switzerland. | Terminated | February 10, 2012 | 480 | 1471 | Not_Computer
Z-0078-2011 | 55150 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyRA Cuvette Segment 1, P/N 007000-001(In Vitro Diagnostic )EasyRA Cuvette Segments included in the following Kits:EasyRA Cuvettes, REF 10700-4Startup Kit - Chemistry, REF 10990Startup Kit - Chemistry & ISE, REF 10991 | EasyRA Cuvette Segments for EasyRA Clinical Chemistry Analyzer | 2 | October 18, 2010 | 2010 | Medica Corporation | Chemistry analyzers may report Calcium results below the sample's actual concentration. Reported results may be ten to fifteen percent lower than expected sample results.The defect is a shortened path length at specific cuvette positions of those segments manufactured using molding cavity 1. The shortened path length was determined by measuring the cuvette dimensions with a Co-ordinate Measuri | OTHER/UNDETERMINED: Pending | Medica issued an Urgent Product Recall letter dated March 29, 2010 to customers. Users were requested to cease using and discard the affected product, and complete and return a Response Form in order to obtain replacement product.Customers can contact Medica Technical Support at 877-777-5895. |  | 14926 | 14,926 | Worldwide Distribution: USA, and the countries of Bulgaria, Canada, India, Mexico, Panama, Phillipines, Russia, Turkey, New Zealand, South Africa, United Arab Emirates, and Venezuela. | Terminated | June 07, 2011 | 232 | 14926 | Not_Computer
Z-0079-2011 | 56805 | prosthesis, hip, constrained, cemented or uncemented, metal/polymer | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) 22 MM I.D., 52 MM O.D., CONSTRAINED LINER, UHMWPE, CO-CR, TI-6AL-4V, REF 71339152, STERILE EO, QTY 1, Smith & Nephew, Inc., Hip prosthesis acetabular component.Acetabular Liner | R3 Constrained Acetabular Liner | 2 | October 18, 2010 | 2010 | Smith & Nephew Inc | The firm received reports of a femoral head popping out of the liner intraoperatively. |  | Smith & Nephew notified all distributors via e-mail and sent a letter dated August 13, 2010 to all hospitals notifying them of the product, the problem, and the action to be taken. 1) Distributors were to immediately cease distribution and use of the product, examine their stock and the stock of the accounts they supplied if applicable and remove any affected devices from all inventories and place them into quarantine to prevent accidental distribution. 2) Identify customers that received these products and immediately notify them of the recall.3) Carry out a physical count and record the data on the Verification Section included in the letter.4) Contact their inventory specialist for an RA# and record this on the Verification Section.5) Return the Verfication Section even if they do not have the recalled product.6) Return the recalled product to: Smith & Nephew, Inc. Attn: Return Goods 3303 Holmes Road Memphis, TN 38118Smith & Nephew, Inc. sent a second notice via e-mail on 09/03/2010 to non-responding distributors and by Fed Ex to the hospitals 09/21/2010.For any questions regarding this recall customers were to call (901) 399-5441. |  | 118 | 118 units | Worlwide Distribution, USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, WA, and WI and the countries of Australia and Portugal | Terminated | September 21, 2011 | 338 | 716 | Not_Computer
Z-0074-2011 | 56274 | medical computers and software | LNX |  |  | General Hospital | 510(k) | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1.SCC Soft Computer, 5400 Tech Data Drive, Clearwater, FL 33760.Soft Query Results Reporting version 2.2.0.6.3 was released for General Availability on 09/11/03.Soft Query Results Reporting version 2.2.0.7.1 was released for General Availability on 01/08/04. | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 | 2 | October 18, 2010 | 2010 | SCC Soft Computer | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 HIS may display incorrect collection data for a procedure or test result. Clients using HBO Star RR interface to query for results on a canceled test may see the previous occurrence of the test with the current date and time displayed as the collection date and time instead of the correct collection date and time sent by the SCC interf | DESIGN: Software Design | Client Notification Task Verbiage was sent through the Task Management System on December 3, 2003, identifying the software issue, the clinical impact, and corrective action. Customers were informed that the HIS may display incorrect collection data. Customers received the SQRR Service Pack with revised software, which will will not produce inaccurate information if there are no verified results. Customers were to acknowledge receipt of the task notification and indicate whether they would like the corrected software.SCC Soft Computer proprietary Task Management System is used to manage client communication. Clients use this system to request services, receive and provide software updates for any issue.Customers can contact SCC Soft Computer at 727-789-0100. | Contact the recalling firm for information | 5 | 5 | Nationwide Distribution: including states of GA, LA, and NH. | Terminated | October 19, 2010 | 1 | 5 | Software | Display/Image | Software update | Software Update
Z-0073-2011 | 56789 | pouch, colostomy | EZQ | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ActiveLife One-Piece Drainable Pouch with Stomahesive Skin Barrier;1 1/4 in., 32mm;Made in Dominican Republic;ConvaTec.Colostomy pouch | ActiveLife OnePiece Drainable Pouch 32mm | 3 | October 18, 2010 | 2010 | Convatec Inc. | Mislabeled: the market units were incorrectly labeled as # 125339 ActiveLife Drainable Pouch 25 mm; the product inside was actually ActiveLife Pouch Small Tan 32 mm, ICC # 125333. | PRODUCTION CONTROLS: Labeling Mix-Ups | ConvaTec, Inc. sent a Returned Goods Authorization form by e-mail dated September 3, 2010, to one customer. The form identified the product, the problem, and the action the customer should take. The customer was instructed to detach the label and affix it to the outside of the box, copy the RGA # and retain a copy for their records, enclose the completed form and return with the product, For questions call (908) 904-2675. | Contact the recalling firm for information | 1 | 1 case (20 market units) | USA, including IL | Terminated | October 19, 2010 | 1 | 1 | Not_Computer
Z-0076-2011 | 56736 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Sequencer Verify & Record System, for use with radiotherapy treatment machines, Class IIb, manufactured by Impac Medical Systems, Sunnyvale, CA. | Sequencer | 2 | October 18, 2010 | 2010 | Impac Medical Systems Inc | Software issue-- matching criteria can be configured that permits an invalid patient match, resulting in overwriting of patient records with information from a different patient. | DESIGN: Software Design | Elekta Impac Software issued a Important Safety Notice letter dated September 2010 to all customers identifying the software issue and clinical impact. Customer Support representatives will contact each site and modify site configuration or upgrade existing software to the version containing the corrective fix. Customers can contact Impac Software at 800 488-4672. |  | 797 | 797 units | Worldwide Distribution: throughout the US and the countries of Australia, Austria, Belgium, Canada, China, Hong Kong, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Ireland, Israel, Italy, Japan, Jordan, Morocco, Netherlands, New Zealand, Poland, Portugal, Singapore, Spain, Sweden, Switzerland, and UK. | Terminated | June 24, 2011 | 249 | 797 | Software | Output/Calculation | Software update | Software Update
Z-0077-2011 | 56729 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | DDC-6-AA Data Card, an accessory to the Sentry DDU-1 00 semiautomatic external defibrillator (AED).The DDC-6-AA data card is an accessory to the AED and does not have its own intended use. | Sentry semi automatic external defibrillator (AED) | 2 | October 18, 2010 | 2010 | Defibtech, LLC | Data cards may include an incorrect format which may cause the AED to not function correctly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Defibtech Initially notified customers by telephone on 8/27/10. Customers were notified of the product, the problem and the action to be taken. Customers were instructed to locate affected data cards and send them back to Defibtech for evaluation. If the data card was inserted into the AED, then the customer was instructed to check the AED software version. If the software version was at 2.005, the customer was instructed to keep the AED out of service until Defibtech could evaluate the returned data card. (Any other software version indicated that the device was not affected by a faulty data card). If the software version was any other version other than 2.005, then the customer was instructed to keep the AED in service unless the original service inquiry required it to remain out of service.If the data card was not inserted into the AED, then the customer was instructed to return the data card for evaluation.For any questions customers were to contact Customer Service at (866) 333-4241 from 8:30-5pm EST, M-F. | Contact the recalling firm for information | 38 | 38 | USA - Nationwide Distribution and Puerto Rico | Terminated | October 05, 2011 | 352 | 38 | Other |  |  | Other
Z-0088-2011 | 56851 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftMic Versions 4.0.4.7, 4.5.1.3. Version 4.0.4.7 Manufactured 02/03/2010. Version 4.5.1.3 Manufactured 03/18/2010. SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760. | SoftMic Versions 4.0.4.7 and 4.5.1.3 | 2 | October 19, 2010 | 2010 | SCC Soft Computer | Softmic system may delay or omit reporting of clinically significant results including organisms and drug sensitivities. Incomplete culture results may appear complete. SCC Soft Computer is issuing a Correction for SoftMic Versions 4.0.4.7-4.5.1.3 software due to this potential hazard. | DESIGN: Software Design | SCC Soft Computer issued Risk-to-Health notifications to clients on July 19, 2010 via proprietary communication software. Corrective actions will be via programming correction via hot fix (software modification) or such so that Global variable is not used. Clients are asked to acknowledge receipt of this task and update the task with permission to load the required correction. | Contact the recalling firm for information | 35 | 35 | Worldwide Distribution: USA, including the states of CA, CT, DE, FL, IA, KY, MA, MA, MI, MN, NJ, NY, OH, PA, TX, VA, and WI, and the country of Canada. | Terminated | July 16, 2012 | 636 | 35 | Software | Output/Calculation | Software update | Software Update
Z-0090-2011 | 56280 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 3.17.6.3.1.SCC Soft Computer 34350 US Highway 19 N, Palm Harbor, FL 34684.Date of Manufacture:Version 3.17.6.3.1 06/03/2003. | SoftPath GUI version 3.17.6.3.1. | 2 | October 19, 2010 | 2010 | SCC Soft Computer | For clients using the SoftPath GUI version 3.17.6.3.1 Results Reporting interface with parsed text and having the auto correct and auto formatting functions active in MS Word, there is a possibility for some ANSI characters/symbols to be translated incorrectly to ASCII. SCC Soft Computer initiated a Correction in March 2004 for all affected clients. | DESIGN: Software Design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on March 11, 2007. SCC recommended turning off the autocorrect and autoformatting functions in MS Word if custom configuration parses text to an ASCII file (HIS) for the Result Reporting interface. If this function was turned on at the time of creating the canned message in the setup file, edit and re-save all canned messages in SoftPath. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide distribution: One hospital in NY. | Terminated | November 03, 2010 | 15 | 1 | Software | Output/Calculation | Software update | Software Update
Z-0087-2011 | 56502 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray: Anterior Cervical Plate (ACP) System, 510 (k) #K080690. The labeling for the plate component reads in part: 'REF: 22-'**** MAT'L: Titanium per ASTM F-136, Lot: W****, Description: Plate - 40 mm, Qty: 1.' The labeling for the screw component reads in part: 'REF: 22-**** MAT'L: Titanium per ASTM F-136 LOT W**** Description: Screw-Fixed Tapping 4x10mm QTY: 1.Manta Ray Anterior Cervical Plate (ACP) System is intended for temporary stabilization of the cervical spine C2-C7 due to the following indications: DDD (defined as neck pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies); Spondylolisthesis; Trauma (i.e. fracture/dislocation); Tumor; Spinal stenosis; Deformity (i. e. scoliosis, kyphosis and/or lordosis); Pseudoarthrosis; and Failed previous fusion. | Manta Ray ACP System | 2 | October 19, 2010 | 2010 | Theken Spine Llc | The Manta Ray Cervical System was subject to corrective action by the firm after two (2) customer complaints were received alleging that screws were 'backing out' of the device during surgical procedures. | OTHER/UNDETERMINED: Pending | Integra sent a Technical Bulletin dated July 29, 2010 via Fedex to their customers. The bulletin identified the product, the problem, and action to be taken by the user/customers (surgeons) of the two (2) reported instances of screws backing out of the plates. The Bulletin provided instructions outlining the proper surgical techniques for implantation of screws using the Manta Ray ACP System. The bulletin also included an attached acknowledgement form which the user/customers were asked to complete and FAX to (609) 275-5363 or E-mail back to the E-mail address provided. The form verified that each customer: 1) received the technical bulletin; and 2) has read and understood the proper technique for surgical placement of the screws as outlined in the bulletin.For technical support customers were to contact the Customer Service Department at (866) 942-8698. |  | 1081 | 1081- includes totals distributed for both plates & screws | Nationwide Distribution - USA including AL, AZ, CA, CT, FL, GA, HI, IL, ID, IN, KS, LA, MI, MD, MN, MA, MD, MS, NE, NY, NC, NJ, NV, NM, OH, OR, OK, PA, TN, TX, UT, WI, and WA.WA. | Terminated | September 21, 2012 | 703 | 1081 | Not_Computer
Z-0091-2011 | 56808 | ligator, hemorrhoidal | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ten Shooter Saeed Multi-Band Ligator, Rx Only, Wilson Cook Medical 4900 Bethania Station Road, Winston-Salem, NC 27105This device is used to endoscopically ligate esophageal varices at or above the gastroesophageal junction or to ligate internal hemorrhoids. | WilsonCook 10 Shot MultiBand Ligator | 2 | October 19, 2010 | 2010 | Cook Endoscopy | A section of the ligator barrel may separate from the barrels friction fit adapter that attaches to the patent end of the endoscope. This could result in a section of the ligator barrel detaching inside the patients gastrointestinal tract. | OTHER/UNDETERMINED: Pending | COOK Endoscopy sent an URGENT: PRODUCT RECALL dated September 3, 2010, via USPS Certified Mail and by email to all consignees. The letter identified the products, the problem, and the action to be taken by the consignees. All consignees were instructed to quarantine any affected product and return them to Cook Medical. Distributors were instructed to contact their customers and coordinate the return of the affected product to Cook Medical. Within 2 weeks of the initial recall notification, Level B effectiveness checks (40% of the total number of consignees) will be conducted. These customers will receive a recall effectiveness survey. All non-responding consignees will receive a total of 3 recall notifications (2 follow-up requests for a response in addition to the initial recall notification. Monthly status reports regarding this recall will be provided to FDA. All devices returned as part of this recall will be destroyed by the recalling firm upon receipt of authorization to do so by FDA. Forty percent (40%) of total customers affected received a second notice on 9/17/2010 as part of the effectiveness checks. As of 9/21/2010, 38% of all affected customers have responded.For questions contact the Customer Relations Department at CustomerRelationsNA@cookmedical.com or call 1-800-457-4500, press 4 then enter extension 2146. |  | 2011 | 2011 units (total including both 6 & 10 Shooters) | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, WA, WI, WV, and WY; and the countries of Algeria, Australia, Austria, Belgium, Canada, Chile, Cyprus, Denmark, France, Georgia, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Jordan, Korea, Mexico, Malaysia, Morocco, Netherlands, New Zealand, Peru, Poland, Portugal, South Africa, Slovakia, Spain, Switzerland, United Arab Emirates, UK and Venezuela. | Terminated | March 31, 2011 | 163 | 2012 | Not_Computer
Z-0089-2011 | 56277 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions 3.17 and 4.1.SCC Soft Computer 5400 Tech Data Drive, Clearwater, FL 33760.Manufactured: Version 3.17.7.6 10/23/2003 | SoftPath GUI Versions 3.17 and 4.1 | 2 | October 19, 2010 | 2010 | SCC Soft Computer | SoftPath Cases with Word Perfect documents that were not signed out in ASCII are printing in GUI as a converted document from Word Perfect to PDF document instead of the current MS Word Document. This affects clients upgrading from ASCII to SoftPath GUI 3.17.7.6 Version. All affected clients have been notified and corrections have been made. | DESIGN: Software Design | SCC Soft Computer issued Risk-to-Health notifications with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System. beginning on December 30, 2003. Clients were instructed to sign out all Word Perfect documents before upgrading to GUI. The correction for this anomaly is to upgrade to the latest release of version 3.17.7 or 3.17.8. Clients were asked to acknowledge receipt of this Task and indicate whether they would like to take the correction at this time. When the client decides to accept the correction a new task will be created to initiate the patch using the Task Management System. |  | 5 | 5 | Nationwide Distribution to the states of IN, NC, NH, and NJ. | Terminated | November 03, 2010 | 15 | 5 | Software |  | Software update | Software Update
Z-0107-2011 | 55195 | radioimmunoassay, parathyroid hormone | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH Enzyme-Linked Immunosorbent Assay(ELISA), Part Number: DSL-10-8000The DSL-10-8000 ACTIVE I-PTH Enzyme-Linked Immunosorbent Assay (ELISA) Kit provides materials for the quantitative measurement of I-PTH in serum or plasma. This assay is intended for in vitro diagnostic use as an aid in the diagnosis of hyperparathyroidism | DSL108000 Active IPTH ELISA | 2 | October 20, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA assay maynot be achieved for lot numbers: 891198, 990980 and 991448.The kit lot numbers listed above passed the stability testing performed at 6 months, but failed the testing performed at 12 months. Therefore the affected I-PTH ELISA kit lots may produce falsely elevated patient r | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an URGENT: PRODUCT CORRECTIVE ACTION letter dated March 17, 2010 with attached Customer Response form to all customers who received Active I-PTH ELISA kits with the affected lot numbers. The letter identified the product, the problem, and the actions to be taken.Customers were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot numbers 891198, 990980 and 991448;(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if your laboratory uses only the control materials provided with the I-PTH ELISA kit;(3) Abnormally elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinicaldiagnostic tests; (4) I-PTH ELISA kit lot numbers 891198, 990980 and 991448 will no longer be shipped.Alternate lots will be provided for current and future replenishment orders or requests.in addition customers were instructed to share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. Complete and return the enclosed response form within 10 days.For questions regarding this recall contact (800) 231-7970. | No consumer action necessary | 70 | 70 kits | Nationwide Distribution including the states of CA, CT, IA, IL, MI, NY, SC, and TX | Terminated | April 27, 2012 | 555 | 70 | Not_Computer
Z-0099-2011 | 56282 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftWeb version 4.0.4.10.1.SCC Soft Computer, 5400 Tech Data Drive, Clearwater, FL 33760.Manufacture Date: 05/16/2006. | SoftWeb Version 4.0.4.10.1 | 2 | October 20, 2010 | 2010 | SCC Soft Computer | If a case number is re-used in SoftWeb with SoftPath, the SoftPath case is linked to the original MRN in SoftWeb. This causes SoftPath cases being linked to the wrong patient in SoftWeb. Incorrect results on the patient with original MRN and no results on the patient with the second MRN (and the reused case number) are displayed in SoftWeb Results Query. SCC Soft Computer issued a Correction for | DESIGN: Software Design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on June 8, 2006. SCC recommended SoftPath/SoftWeb users NOT re-use SoftPath case numbers when cases are cancelled. The generic option setting "Specimen Registration - Standard Settings - Minimum Time to Reactivate Case Number" should be set to "0" so the user is not prompted to reuse the same case number. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide Distribution: 1 hospital in Alaska. | Terminated | November 03, 2010 | 14 | 1 | Software | Output/Calculation | Software update | Software Update
Z-0098-2011 | 56687 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 blockAid 400 mL Painpump for continuous Nerve Block, REF 575-100, Sterile, Rx only, Stryker Instruments, Kalamazoo, MI. |  | 2 | October 20, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The instructions for use (IFU) has been revised for health care professionals. The indication for use has been changed to remove site specific pain management. The firm has revised the warning to not use the pump around joint spaces. A warning has been added regarding the use of the pump with anticoagulants in epidural applications. A special reminder to users to always evaluate the potential | OTHER/UNDETERMINED: Pending | Stryker Instruments sent Urgent Medical Device Recall Notification letters dated 8/16/2010 to healthcare professionals. The recall letters outlined the changes for the instructions for the use of these products. The letters informed the healthcare professional to review the new IFU, to distribute the new IFU to all healthcare professionals that use the device, and to complete and return the business reply form. The letters included a complete IFU that was to be used with any remaining stock of devices. Customers should dispose of the old version of the IFU. There is no product being returned for this recall. | Contact the recalling firm for information | 11976 | 11,976 all products | US, Canada and Australia. | Terminated | April 21, 2011 | 183 | 11976 | Not_Computer
Z-0103-2011 | 56675 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LIBERTY Automated Peritoneal Dialysis System Line Cycler Set with stay safe connectors (Dual Patient Connector) PIN: 050-87212, manufactured for Fresenius Medical Care, Waltham MA by Erika de Reynoza, Reynosa Tampa, Mexico. | Fresenius LIBERTY 05087212 Automated Peritoneal Dialysis System Line Cycler Set | 2 | October 20, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Cycle set cassettes may leak due to holes/scratches in the film, contaminating dialysate and potentially resulting in peritonitis. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Fresenius Medical Care issued an Urgent FMCNA LIBERTY Cycler Set Recall letter on September 3, 2010 identifying the specific affected part numbers and product lots and requesting that customers identify and discontinue the use of the affected lots immediately. Customers were requested to return a receipt of notification. Customers may arrange return unused product to FMCNA by contacting customer service at 1-800-323-5188. |  | 363265 | 363,265 units | Nationwide Distribution: Throughout USA. | Open |  |  | 370588 | Not_Computer
Z-0108-2011 | 56802 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | 2.0/1.0 4 hole long plate (titanium plate ).Lot 120150.Product labeling reads: (1. ) REF: 01-9205, Lot 999999, " 2.0MM" SYSTEM PLATE 4 HOLE PLATE, STRAIGHT, LONG.(2.) QTY 00001, CE 0086, *** NON-STERILE PRODUCT***, 1.0 MM TITANIUM, BIOMET MICROFIXATION, 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA, www.biometmicrofixation.com.(3.) #01-9205, LOT 999999, QTY 00001, 4 HOLE PLATE, STRAIGHT, LONG, 1.0 MM, CE 0086 | 2.0/1.0 4 hole long plate Model 019205 | 2 | October 20, 2010 | 2010 | Biomet Microfixation, Inc. | Mislabeled size: The 2.0/1.0 4 hole long plate, Lot 120150 is labeled as 01-9205 (2.0/1.0 4 Hole Long Plate), however, the package contains a 01-9204 (2.0/1.0 4 Hole Regular Plate). | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Biomet Microfixation sent URGENT MEDICAL DEVICE RECALL NOTICES dated August 31, 2010 to its customers and distributors, identifying the affected product and actions to be taken by customers. Customers were requested to check their inventory for product and return any affected product to Biomet Microfixation. Customer questions or concerns should be directed to Laura Sabo, VP Global QA/RA at 904-741-4500. |  | 56 | 56 | Worldwide Distribution: USA including states of Hawaii and Kansas, and the countries of Australia, Denmark, Italy, Jordan, and South Africa. | Terminated | November 19, 2010 | 30 | 56 | Not_Computer
Z-0093-2011 | 56279 | neurosurgical paddie | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | Codman 5/16" Round Surgical Patties(20 pouches of 10 patties per box)Product code: 80-1396 | Codman | 2 | October 20, 2010 | 2010 | Codman & Shurtleff, Inc. | Surgical Pattie x-ray stripe (radio-opaque marker) became detached | PRODUCTION CONTROLS: Process Control | Codman notified customers by letter dated 7/16/10 issued July 23 through July 26, 2010 via Fed'x overnight advising users that the x-ray marker on Codman Surgical Patties may become detached. Customers are instructed to return product from affected lots. An error in the original letter was discovered and a corrected letter dated July 28, 2010 issued. If there are any questions, please contact Codman at (800) 225-0460. | Contact the recalling firm for information Contact the recalling firm for information | 673 | 673 boxes (20 pouches of 10 patties per box) | NationwideForeign: Canada, Singapore, Belgium, Mexico, Madrid, Dubai, Tokyo, Brazil, New Zealand, Egypt, South Africa, Singapore, | Terminated | July 09, 2012 | 628 | 5309 | Not_Computer
Z-0102-2011 | 56272 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 4.2.3.4 2008-06-30 (Manuf Date). SoftPath GUI version 4.2.4.1 2008-09-30 (Manuf Date). SoftPath GUI version 4.3.0.3, 2008-07-23 (Manuf Date). SoftPath GUI version 4.3.0.4 2008-09-16 (Manuf Date). SCC Soft Computer, 5400 Tech Data Drive, Clearwater, FL 33760. | SoftPath GUI versions 4.2.3.4, 4.2.4.1, 403.0.3, and 4.3.0.4. | 2 | October 20, 2010 | 2010 | SCC Soft Computer | For SoftPath GUI versions 4.2.3.4, 4.2.4.1, 403.0.3, and 4.3.0.4., the Previous Patient History and some QA Management reports may have missing or incorrect diagnosis text. However, the patient final report, HIS resulting and Query display are all correct. All affected clients have been provided corrected software.. | DESIGN: Software Design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System on January 26, 2009. SCC issued a software upgrade with corrective patch electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software.SCC will contact each client to coordinate the running of the utility and the loading and validation of the correct patch. |  | 52 | 52 | Worldwide Distribution: Throughout USA and Canada. | Terminated | November 03, 2010 | 14 | 52 | Software | Output/Calculation | Software update | Software Update
Z-0106-2011 | 56346 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand TrueBeam and TrueBeam STx, Model Number:TMX- H19, Distributed by and Manufactured by: Varian Medical Systems Inc., Palo Alto, CA | TrueBeam and TrueBeam STx | 2 | October 20, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to a software anomaly, guidance-based couch shift values may not be applied as expected when the operator selects "Apply Shift" and presses Motion Enable buttons. This can result in the actual patient position differing from that position indicated by the imaging application affecting poistion accuracy and precision. | DESIGN: Software Design | Varian Medical Systems issued an Urgent Medical Device Correction letter date dated May 24, 2010 to users identifying the affected device, the issue prompting the correction, and the actions to be taken by the user.Customers were asked to set couch tolerance values to small numbers, and for radiosurgical treatment tolerance values to be set to zero. A Varian representative will contact customers to schedule installation of a software version correcting the issue. US Customers can contact Varian at 1-888-827-4265,. European customers can contact Varian at +41 41 749 8844. |  |  |  | Worldwide Distribution: Two consignees, one in the US and one in Switzerland. | Terminated | July 20, 2011 | 273 |  | Software | Output/Calculation | Software update | Software Update
Z-0101-2011 | 56682 | bottle, collection, vacuum | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | NOVAPLUS (TM), SUCTION CANISTER, 1200cc, RIGID, WITH ELBOW, Cat. No. V71-3001, Rx only, NON-STERILE, Made exclusively for Novation by DeRoyal, 200 DeBusk Lane, Powell, TN 37849Most commonly used during medical medical emergencies and surgery to remove fluid and secretions from a patient. | NOVAPLUS (TM), SUCTION CANISTER | 2 | October 20, 2010 | 2010 | DeRoyal Industries Inc | Suction canisters may crack during use | OTHER/UNDETERMINED: Pending | DeRoyal sent a Recall letter dated July 1, 2010 to all their consignees via UPS Air. The letter identified the product, the problem, and action to be taken by the consignees. Consignees were to complete the Recall Response Form by checking their inventory for DeRoyal Suction Canisters V71-3001, Lot #22622679.Indicate the number of replacement products needed on the Recall Response Form, including the address, reference PO#, etc. to be used for the replacement products and fax the form to DeRoyal at 865-362-3755 no later than July 7, 2010.Recalled products were to be returned to DeRoyal via UPS Ground Attn: Quality Control 300 DeBusk Lane Powell, TN 37849Once the recalled product was received by DeRoyal, the replacement product would be shipped per the instructions on the Recall Response Form.Distributors were to forward the recall notice to their end-users, or provide customer listing so that DeRoyal may contact them.For any questions or assistance regarding this recall call 865-362-1037. | Contact the recalling firm for information |  |  | Nationwide | Terminated | April 10, 2012 | 538 |  | Not_Computer
Z-0035-2011 | 56685 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Excelsior Disposable Syringe W/Normal Saline (0.9% sodium Chloride)General hospital use | Excelsior Disposable Syringe W/Normal Saline | 1 | October 20, 2010 | 2010 | Excelsior Medical Corp | There is the potential for a dimensional issue in 6 ml syringe products, codes E0100-50, 10056-1000, 10056-240, 14056-240, 910056-1000, S5, to cause leakage and.or loss of sterility | DESIGN: Device Design | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 2, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue using syringes from the listed lots and return the attached Customer Reply Form via fax (866) 688-3185 as confirmation that notification had been received and that proper arrangements can be made to have the product returned to the firm. The customers were also instructed to inform their customers, if they have resold products, of the recall and instruct them to return the form, and verify that they have returned all unused syringes from any of the listed lots . NOTE: The customers were informed that they may alternatively have the products destroyed and certified to Excelsior that they have done so. They must include a copy of the certificate of destruction with the attached Customer Reply Form via fax. - Excelsior issued a press release on October 14, 2010. The press release advised consumers of the recall and requested that use of the product be immediately discontinued and returned to the point of purchase.Should you have any questions regarding the recalled product, please contact Excelsior Medical Customer Service at (800) 487-4276 or if you prefer by fax (866) 688-3185. You may also contact us at recall10@excelsiormedical.com. |  | 13734640 | 13,734,640 | Worldwide distribution: USA and countries including Canada and Brazil. | Terminated | August 23, 2012 | 673 | 13734640 | Not_Computer
Z-0100-2011 | 56283 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftLab version 4.01 ASCII, ASCII Base 3.1.3.0-3.1.3.32, 3.1.4-3.1.4.27, 3.1.6-3.1.6.4, 4.0.0.0-4.0.0.9; LAB/MIC Add-on 4.0.1, 4.0.1.0-4.0.1.13, 4.0.2.0 in 2006. SCC Soft Computer, 5400 Tech Data Drive, Clearwater, FL 33760. | SoftLab version 4.01 ASCII Lab/Mic Addon | 2 | October 20, 2010 | 2010 | SCC Soft Computer | For SoftLab version 4.01 ASCII, the qualifying test result does not reach the Call List because the order is locked by either another application, or two Hosparams (HISRR_RBS_sync and Auto_sync_RBS). This creates a lock on the RBS server. The problem was identified in the RBS "C" rule which sends qualifying results to the Call List for SoftLab. All affected clients received the upgrade for Soft | DESIGN: Software Design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System in June 2006. SCC recommended SoftLab version 4.01 ASCII users change program settings to prevent locking of the RBS server. A software upgrade with corrective patch will be delivered electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software. |  | 131 | 131 | Nationwide Distribution: USA in the states of AL, AR, CA, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SD, TN, TX, VA, WA, WI, and WV. | Terminated | July 16, 2012 | 635 | 131 | Software | Device Operation | Software update | Software Update
Z-0110-2011 | 55922 | defibrillator, automatic implantable cardioverter, with cardiac resynchronization (crt-d) | NIK |  |  | Cardiovascular |  | Medtronic Concerto II CRT-D, model D274TRK and Model D294TRK not available in the US. Digital implantable cardioverter defibrillator with cardiac resynchronization therapy (DDE-DDR). Sterilized using ethylene oxide. Medtronic, Inc., Minneapolis, MN 55432 USA. Manufactured in: Juncos, Puerto Rico, USA.The Dual chamber implantable cardioverter defibrillator with cardiac resynchronization therapy (CRT-D) is a multiprogrammable cardiac device that monitors and regulates the patient's heart rate by providing single or dual chamber rate-responsive bradycardia pacing, sequential biventricular pacing, ventricular tachyarrhythmia therapies, and atrial tachyarrhythmia therapies. The device automatically detects ventricular tachyarrhythmias (VT/VF) and provides treatment with defibrillation, cardioversion, and antitachycardia pacing therapies. The device also automatically detects atrial tachyarrhythmias (AT/AF) and provides treatment with cardioversion and antitachycardia pacing therapies. Simultaneous or sequential biventricular pacing is used to provide patients with cardiac resynchronization therapy. The device responds to bradyarrhythmias by providing bradycardia pacing therapies. | Concerto II CRTD | 2 | October 21, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic is informing you of a rare device software issue in certain ICD and CRT-D devices. A software update will be available to correct his issue. There have been 5 reports of this issue occurring with no reported patient injuries or deaths. Medtronic has identified the root cause to be a rare and specific sequence of events that must occur within a few milliseconds of each other. In the u | DESIGN: Software Design | Medtronic, Inc. sent an IMPORTANT MEDICAL DEVICE CORRECTION LETTER, dated May 2010 to physicians beginning May 05, 2010. The letter informed doctors of the issue and informed them that a software update would be available to correct this issue. The software was to be installed on the programmer, with a Medtronic representatives assistance, so that on a patient's next in-clinic interrogation the device would receive the upgrade automatically and their device would be corrected.For any questions regarding this recall physicians were to contact their local Medtronic Representative or Medtronic Technical Services at 800-723-4636. | Contact the recalling firm for information | 18778 | 18,778 | Worldwide Distribution - USA, Puerto Rico, Virgin Islands, and the countries of Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City, Venezuela, and Viet Nam. | Terminated | March 15, 2012 | 511 | 210237 | Software
Z-0121-2011 | 56484 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics Systems QCFIRST Custom RuleThe EasyLink Informatics System is classified as a data calculator and processing module for use with In-vitro diagnostic devices. The EasyLink is included with the Dimension Vista System (K051087), and may be used with other Siemens analyzers, as a communications and connectivity workstation for integration with laboratory information system (LIS) networks. | EasyLink | 2 | October 21, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | the EasyLink QCFIRST custom rule does not operate as intended. | OTHER/UNDETERMINED: Pending | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated July 2010 to all customers that have the EasyLink QCFIRST custom rule installed. This Urgent Field Safety Notice included information on the problem and confirmed that a Siemens representative would contact them to facilitate the removal of the rule from the EasyLink System.For questions contact the Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 82 | 82 | Worldwide Distribution - USA including AZ, CA, FL, GA, IA, IL, IN, KS, KY, LA, MO, MS, NC, NH, NJ, NM, NY, OH, OR, PA, TX, VA, WA, WI, and WY and the country of Canada. | Terminated | March 19, 2013 | 880 | 82 | Not_Computer
Z-0120-2011 | 56701 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂŽ Complete Knee Solution Stemmed TibialComponent Precoat Size 3, REF 00-5980-037-01, Sterile, Zimmer, Warsaw, IN.Intended for mating with a UHMWPE articular surface and Zimaloy Cobalt-Chromium-Molybdenum alloy femoral component and is indicated for cemented use only. | NexGen Complete Knee Solution Stemmed Tibial | 2 | October 21, 2010 | 2010 | Zimmer Inc. | Zimmer had determined that two lots of these tibial implants have incorrectly positioned or missing flange plugs that were not seated in the device upon receipt to the customer. A missing flange plug could lead to a delay of surgery or possibility of wear leading to more surgery. | TRAINING: Employee Error | The firm sent an URGENT MARKET WITHDRAWAL to its sales force and distributors on 1/26/2010. The letter identifies the affected product and explains the reason for the Market Withdrawal. The letter also explains the clinical implications and instructs consignees to locate any unused devices and quarantine them immediately. A physical count should be carried out of all affected devices and recorded on the Inventory Return Certification Form that's included with the letter. A copy of the letter should be faxed to Zimmer, Inc. and the affected product should be returned along with the form to the firm. |  | 61 | 61 units | Worldwide Distribution -- US, Canada, Singapore Australia, and India. | Terminated | October 23, 2010 | 2 | 61 | Not_Computer
Z-0109-2011 | 56762 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) A1c Test Cartridges (Catalog 281-0001EX) and in2it (II) A1c Test Cartridges (Catalog 281-0001), manufactured by Bio-Rad Laboratories Deeside Ltd., Flintshire, UK, distributed by Bio-Rad Laboratories, Hercules, CA.The in2it (I) is for Physician use, the (II) is for prescription home use. The products are medical devices used in affinity chromatography method, intended for the in-vitro quantitative determination of A1c (HbA1C) in capillary blood taken from a finger prick. The test is indicated for monitoring the time averaged blood glucose levels of known diabetics, for professional use as an indicator of overall glycemic control. | in2it (I) and (II) Test Cartridges | 2 | October 21, 2010 | 2010 | Bio-Rad Laboratories Inc | Test Cartridges may leak and have the potential to generate error messages and biased inaccurate results. | DESIGN: Component Design/Selection | BIO-RAD Laboratories Deeside sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated September 7, 2010, via registered mail to all subsidiaries/customers. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to stop using and discard the test cartridges from the listed lots; complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax +44 1244 833401; record the number of affected kits of product, confirm that they have discarded the cartridges and record the method of disposal.If you have any questions, contact +44 1244 288888. | Contact the recalling firm for information | 4056 | 4056 | Worldwide distribution: USA and countries including: Australia, New Zealand, Taiwan, France, the UK, Singapore, Germany, Italy, Thailand, China, Korea, Sweden, Sri Lanka, Maldives, Canada and the Philippines. | Terminated | May 17, 2011 | 208 | 4056 | Not_Computer
Z-0133-2011 | 56508 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 & 3056.Produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Computed Tomography XRay System MX 8000 IDT 16 | 2 | October 22, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken.Customers were instructed to locate and correct the error. Field Service Engineers were to visit each site to update the softwareFor further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377. |  | 16 | 16 units | Product was sold in the following states: CA, DE, IN,OR, MA, NV, TN. Product was also sold in Canada and Australia | Terminated | September 21, 2012 | 700 | 16 | Software | Output/Calculation | Software update | Software Update
Z-0123-2011 | 56627 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. | GE Centricity Laboratory Instrument Interface Software | 2 | October 22, 2010 | 2010 | Ge Healthcare It | Instrument may incorrectly report the result of an antimicrobial sensitivity test. | DESIGN: Software Design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated August 27, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue associated with the use of GE Centricity Laboratory instrument interface to Biomerieux's Vitek analyzer where results may be incorrectly reported. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 22 | 22 units | Worldwide Distribution -- USA, including states of Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee, and Texas and countries of Australia, Canada, England, India, New Zealand, Qatar, and Scotland. | Terminated | July 15, 2011 | 266 | 22 | Software | Output/Calculation | Software update | Software Update
Z-0131-2011 | 56791 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | ARRAY and POLARIS Spinal System Medium Cross ConnectorNon SterileBiomet Spine, Parsippany, NJ. 07054 | ARRAY and POLARIS Spinal System | 2 | October 22, 2010 | 2010 | Ebi, Llc | One of the two connector ends of the medium cross connector may not be fully secured to the rod after tightening. This may not be readily apparent to the operating surgeon, resulting in inadequate torsional stability. | OTHER/UNDETERMINED: Pending | Biomet Spine issued an Urgent Medical Device Recall letter dated September 16, 2010 to distributors and physicians, identifying the affected device and actions to be taken.Customers and distributors were requested to identify, remove, and immediately return the affected product to Biomet, following enclosed Product Return Instructions, and complete and return the Removal Tracking and Verification Form.Customers and distributors can contact Biomet at 1-800-526-2579. |  | 82 | 82 units | Worldwide Distribution: USA and the countries of Australia, Columbia, Great Britain, The Netherlands, New Zealand, and Switzerland. | Terminated | June 01, 2012 | 588 | 82 | Not_Computer
Z-0128-2011 | 56462 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq Two-Channel Infuser,LIST No. 16027 Symbiq Two-Channel Infuser, List Number Configurations: 16027-04-01/03/04, 51/52, 53/54, 79/80,81/82,83/84,87/88 and 89/90 16026-27-89/90, 16027-27-89/90, 16027-13-25/26,53/54, Distributed by: Hospira, Inc. Morgan Hill, CA | Hospira brand Symbiq TwoChannel Infuser | 1 | October 22, 2010 | 2010 | Hospira Inc | Hospira has identified motor encoder failures in the Symbiq pumping mechanism that causes the infuser to cease operation during infusion therapy, resulting in delay or interruption of infusion therapy. | DESIGN: Device Design | Hospira issued a Recall Notification letter dated October 4, 2010, updating the Urgent Device Field Correction letter dated February 22, 2010, ideentifying the affected products, the issue prompting the recall, and actions to be taken by customers.Customers are to examine their inventory to determine if they have any of the affected products, and recommended to remove them from service in critical care areas, and contact Hospira to schedule removal and replacement of the motor. Replace affected devices with a loaner or corrected back-up pump if available.Custoemrs can contacvtd Hospira Global Product Safety and Complaints at 1-800-441-4100, and Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA and the coutnries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | December 03, 2013 | 1138 |  | Not_Computer
Z-0132-2011 | 56610 | flowmeter, blood, cardiovascular | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | PD Access Vascular Access Device. Escalon Vascular Access, Inc., 2440 South 179th Street, New Berlin, WI 53146:Model 78050 18 Gauge GA Doppler Guided Bare Needle Assembly, 2 3/4 in Length, 0.038 in guidewire, Sterile EO, PD Access Vascular Access Device.Model 78060 Extended 18 Gauge Doppler Guided Bare Needle Assembly, 3 1/2 in Length, 0.038 in Guidewire, Sterile EO, PD Access Vascular Access DeviceModel 75010 18 Gauge 2 3/4 in Length, 0.038 in guidewire, Sterile EO, SmartNeedle Vascular Access Devices | PD Access Vascular Access Device | 2 | October 22, 2010 | 2010 | Escalon Medical Corp | The packaging of the vascular access device may potentially become damaged, compromising the sterility of the product. Use of a non-sterile product increases risk of patient infection. This recall involes: Model 78050 18 Gauge PD Access Vascular Access Device, Model 78060 Extended 18 Gauge PD Access Vascular Access Device, Model 75010 18 Gauge SmartNeedle Vascular Access Devices with Expiration Da | DESIGN: Packaging Design/Selection | Escalon Vascular Access issued an "Urgent: Recall" letter dated May 5, 2010 to customers. The letter identified the affected product being recalled, described the problem, and the actions to be taken.. Consignees were recommended to immediately examine their inventory, quarantine the product, notify their customers, and return product affected. They also advised them that if they have further distributed the product, to please identify the customers and notify them at once of the product recall. Replacement product can be ordered from Vascular Solutions, Inc at 888-240-6001. Customer service can be contacted at 800-676-0043. |  | 188375 | 188,375 | Worldwide Distribution: USA, including the states of AL, AR, AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD.MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, MN, NY, NC, ND, OH, OK OR, PA, RI, SC, SD,TN, TX UT, VA, WA, WV, and WI, and the countries of: Australia, Austria, Bahamas, Canada, Colombia, Denmark, England, Germany, Greece, Hong Kong, Italy, the Netherlands, New Zealand, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | March 28, 2013 | 888 | 188375 | Not_Computer
Z-0125-2011 | 49488 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD10/10 Biplane X-ray system.Intended for use in cardiovascular and vascular x-ray imaging applications, including diagnostic, interventional procedures (such as PTCA and stent placement and atherectomies), pacemaker implantations and electrophysiology. It is compatible with specified magnetic navigation systems. | Allura Xper FD10 | 2 | October 22, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential for table and stand not to move due to a software error. | DESIGN: Software Design | Philips Medical Systems North America Co. sent an URGENT - Device Correction, dated September 15, 2010, to their consignees. The letter identified the product, the problem, and the action to be taken. The consignees were informed that there was a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible.The consignees were advised that as geometry software errors happen intermittently, no actions can be performed by the consignees to minimize the risk except for not performing DRA procedures until the firm upgrades the software.For question consignees were to contact the Philips Call Center at 800-722-9377 #5, #4, #1 and reference FCO 72200111 or FCO 72200125. | No consumer action necessary | 26 | 26 units | Devices were distributed to hospitals and medical facilities through out the U.S. | Terminated | November 19, 2010 | 28 | 73 | Software | Device Operation | Software update | Software Update
Z-0122-2011 | 56503 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) and (II) System A1c Test cartridges and in2it Self Test System A1c Test Cartridges, for use with the in2it (I) and (II) self test system analyzer, manufactured by Bio-Rad, Flintshire, UKThe in2it (I) and (II) is an affinity chromatography method and is intended for the in-vitro quantitative determination of A1c in capillary blood taken from a finger prick. The test is indicated for monitoring the time averaged blood Glucose levels of known diabetics, for professional use as an indicator of overall glycemic control. The in2it (I) assay is intended for use in a physicians/doctors' office, the in2it (II) is intended for home use. The in2it Self Test system in intended for testing blood taken from a fingerstick. The test system is for self testing and for in-vitro diagnostic use only (not to be swallowed). The system shows how good glucose control has been over a two to three month period. The test system is the same as the test that has been CE marked for use in a clinical laboratory. The instructions for use have been changed with suggestions from people with diabetes. A patient study was conducted where untrained users ran their own blood tests and produced results which were as accurate as results obtained by trained users. | in2it Self Test System A1c Test Cartridge | 2 | October 22, 2010 | 2010 | Bio-Rad Laboratories Inc | Inaccurate results with high bias-- Inaccurate results may be read by the machine, creating the risk for untreated hypoglycemia at elevated altitudes. | DESIGN: Device Design | BIO-RAD Laboratories Deeside Ltd., sent an URGENT: MEDICAL DEVICE CORRECTION letter, dated July 26, 2010, to their subsidiaries and to all customers identifying the product, the problem, and the action to be taken. Subsidiaries were to ensure that all customers received the letter with the Customer Medical Device Correction Response Form and collect the completed form from every customer. Once forms were received from all customers they were to be faxed to BIO-RAD Deeside Regulatory Affairs at +44 (0) 1244833401. Customers were instructed to run a test and visually inspect the test cartridge to detect any residual clear liquid. If the cartridge did not run properly the test was to be rejected.Customers were to run Quality Control test with each new shipment of test cartridges. Patient's results were not to be reported if the QC failed.Customers were to complete the CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM and fax it to the number provided on the form. For questions customers were to contact their regional BIO-RAD office. |  | 90703 | 90703 | Product was distributed to Bio-Rad's US subsidiary, along with 18 other distribution subsidiaries outside the US. These distributed further to other consignees. POroduct was distributed to Mexico, Korea, Singapore, Italy, South Africa, India, Australia, Spain, Greece, France, Thailand, Germany, New Zealand, China, Hong Kong, Portugal, Czech Republic, United Kingdom, Canada, Russia, Netherlands, Taiwan, Austria, Denmark, Israel, Finland and Belgium as well as to other African and Latin American countries. | Terminated | October 25, 2011 | 368 | 90703 | Not_Computer
Z-0126-2011 | 56730 | calibrators, ethyl alcohol | DNN | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | EndoTool Drug Dose Calculator-Software Version 7.2.1800.3 and 7.2.1825.5 | EndoTool Drug Dose Calculator | 2 | October 22, 2010 | 2010 | Hospira, Inc. | The EndoTool software did not always provide an audible alert when a blood glucose measurement was due. This could result in a patient not receiving an appropriate dose of insulin. | OTHER/UNDETERMINED: Pending | Hospira issued an Urgent Device Field Correction letter dated July 23, 2010 to consignees, providing details of the potential issue and the associated patient risk. Hospira is finalizing software enhancements which will resolve this situation within the next 90 days ,and will contact each consignee when it is available to be downloaded from the Hospira website. The notification letter also included a clinical bulletin, outlining mitigation instructions for clinicians to use until the software enhancement is available. Consignees were requested to complete and return an attached Reply Form by fax.Customers can contact the Hospira EndoTool support line at 1-877-349-4582 for clinical inquiries about this recall. |  | 97 | 97 units total (96 version 7.2.1800.3 and 1 version 7.2.1825.5) | Nationwide Distribution in states of CA, FL, IL, IN, KY, KS, MI, MO, MT, NC, NJ, NM, NV, OH, SC, SD, TN, TX, VA, WA, and WY. | Terminated | February 25, 2014 | 1222 | 97 | Software | Alarm/Message | Software update | Software Update
Z-0140-2011 | 49605 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Pulsera R2.2 Mobile X-ray system. The Type Product Code 718094 can be found under the X-ray trolley cover on the front lower right side of the module or under the stand cover. The device's brochure is labeled in part: "Philips...Mobile C-Arm systems BV Pulsera". | BV Pulsera R2.2 Mobile Xray system | 2 | October 25, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Reliability of the Stand-Trolley cable may result in loss of images on the monitor during an examination. This intermittent problem can lead to System lock-up, where no radiation is possible and images are not visible on the monitor | DESIGN: Component Design/Selection | On 10/04/2008, Philips began sending the URGENT-Device Correction BV Pulsera R2.2 and Endura R2.2 letters to their customers. On 09/03/2008, the firm issued the Field Change Order (FCO) 71800024 and the firm's representative will replace the stand-trolley cable free of charge.Customers should contact Philips Call Center at 1-800-722-9377, #5, #2, and reference "FCO 71800024" with any questions. |  | 1331 | 1331 units | Nationwide distribution. | Terminated | December 07, 2010 | 43 | 1458 | Not_Computer
Z-0135-2011 | 56657 | instrument, coagulation, automated | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base)PN 000280000Bench top, fully automated, random access analyzer designed specifically for in vitro diagnostic clinical use in the hemostasis laboratory for coagulation and/or fibrinolysis testing in the assessment of thrombosis and/or hemostasis. The systems provide results for both direct hemostasis measurements and calculated parameters | Instrumentation Laboratory | 2 | October 25, 2010 | 2010 | Instrumentation Laboratory Co. | Software: New data check, "Multiple Threshold Check", was added to identify and fail abnormal (bimodal) clot curves | DESIGN: Software Design | Instrumentaion Laboratory notified the US Distributor/Canada via BCI Service Modification Notice (#10793) on August 13, 2010. The Mandatory software upgrade to Verison V4.3.0 in the United States and Canada will be conducted by BCI service representatives and tracked through the distributor, Beckman Coulter (BCI). International Affiliates/Dealers:Worldwide Regulatory Notices were e-mailed to IL international affiliates and dealers, notifying them of the mandatory software upgrade, providing issue description and a risk summary and the support. The upgrades were to be handled through the local representatives. |  | 177 | 177 units | Worldwide Distribution, USA and Canada | Terminated | May 23, 2012 | 576 | 1076 | Software
Z-0139-2011 | 56703 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista V-LYTE Standard A, Catalog No K820For the calibration of Na+/K+/Cl- on the Dimension vista System. | Dimension Vista VLYTE Standard A | 2 | October 25, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | High potassium content in the affected lots results in the potential for low potassium recovery for Dilution Check, QC and patient samples when run on the Dimension Vista System. | OTHER/UNDETERMINED: Pending | Siemens sent an Urgent Field Safety Notice (dated "August 2010") to customers identifying the product, the problem and the action to be taken. Customers were instructed on how to continue testing until new lots of the recalled product were available and asked to complete and fax a Field Correction Effectiveness Check to 301-631-8467 to confirm receipt of the Urgent Field Safety Notice. For technical questions or concerns customers were instructed to contact the Siemens Technical Solutions Center at 800-441-9250. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 3706 | 3,706 cartons | USA, Nationwide Distribution | Terminated | October 26, 2010 | 1 | 3706 | Not_Computer
Z-0143-2011 | 56790 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | The label (which is affixed to ScanPoint Communications Cradle) states: "...VERATHON ScanPoint Docking Station...Part No. 0570-0168...Bothell, WA USA..." | Scanpoint Communication Cradles | 3 | October 26, 2010 | 2010 | Verathon, Inc. | Incorrect programming causes QuickPrint to fail to recognize the communication Cradles. As a result, ScanPoint Communications Cradles could not communicate with ScanPoint Online Tool while using the QuickPrint interface. This will affect the transfer of exams and device calibration through QuickPrint interface. | DESIGN: Software Design (Manufacturing Process) | On 09/20/10, Verathon issued an URGENT: Medical Device Accessory Recall Notification letter dated September 20, 2010 to consignees, identifying the affected product, the issue prompting the recall, and actions to be taken by consignees.The consignees are instructed to return any cradles within the serial number range between CC012000-CC012084, and the Part Number 0570-0168.The firm will provide a replacement communication cradle at no cost. Consignees can contact Verathon Customer Care at 425-867-1348 or 800-2331-2313 or send e-mail to cservice@verathon.com. | No consumer action necessary | 70 | 70 units | Worldwide Distribution: U.S. nationwide including states AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, LA, MA, MD, MI, MN, MO, ND, NY, OH, PA, SC, TN, TX, UT, WI, and Washington DC, and the countries of Australia, Canada, France, South Korea, Taiwan, The Netherlands, and United Kingdom. | Terminated | April 01, 2011 | 157 | 70 | Software | Device Operation | Replace | Remove or Replace
Z-0413-2011 | 56580 | solvent, adhesive tape | KOX | Surgical skin degreaser or adhesive tape solvent. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Detachol adhesive remover, 4 fl oz (118 mL), Marketed by Eloquest Healthcare, Inc., a Ferndale Pharma Group company, Manufactured by Ferndale Laboratories, Inc., Ferndale, MI 48220 U.S.A. | Detachol Adhesive Remover | 3 | October 26, 2010 | 2010 | Ferndale Laboratories, Inc. | There is a defect on the neck surface of the bottles, causing bottles to leak. Bottle defect may potentially cause contamination of solvent stored in bottles. | DESIGN: Component Design/Selection | Ferndale Pharma Group issued an Urgent Medical Device Recall letter dated July 7, 2010 to customers identifying the affected product. Customers were instructed to identify and quarantine all affected product and return it to Ferndale Laboratories for a credit. Customers are requested to complete and return an enclosed response form.Customers can contact Ferndale at 800-621-6003. |  | 55000 | 55,000 bottles | Nationwide Distribution: Throughout the US. | Terminated | March 02, 2011 | 127 | 55000 | Not_Computer
Z-0145-2011 | 56893 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Guide, Catalog Number 9732355, to be used with a Medtronic Computer-Assisted Surgery System to help assist with screw placement during spine procedures.To be used with a Medtronic Computer-Assisted Surgery System to help assist with screw placement during spine procedures. | Surgical drill guide | 2 | October 26, 2010 | 2010 | Medtronic Navigation, Inc | Drill guide handle may detach unintentionally during use. | OTHER/UNDETERMINED: Pending | Medtronic Navigation sent a "Product Correction Notification" letter dated September 22, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to set aside and do not use the Vertex Max Drill Guide 9732355, Lot Number 100526 and complete and return the Verification of Field Action Completion form via email: steve.murray@medtronic.com or fax: 720-890-3517. A Medtronic representative will contact the customer, visit their facility, remove the defective product, and provide a replacement Vertex Max Drill Guide that is correctly manufactured.For further information or if you have questions regarding use of the Vertex Max Drill Guide, please contact Medtronic Navigation Technical Services at 1-800-595-9709. | Contact the recalling firm for information | 8 | 8 units | Worldwide distribution: USA including states of DC, IN, PA, and VA and country of Japan. | Terminated | April 13, 2011 | 169 | 8 | Not_Computer
Z-0144-2011 | 56754 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems VB23A-VB23H with Perivision functionality, Model Numbers: 5904433,5904649, 5904466, 5904441, 5904656, 7728350, and 7007755Angiographic X-ray System | Axiom Artis | 2 | October 26, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Sporadic system crash while Perivision is performed. | DESIGN: Software Design | Siemens issued an Urgent Field Safety Notice to all customers. The letter informed customers of the potential issues and provided additional information on the implemented resolution. To resolve the problem, a system update was to be installed on each affected system under Update Instructions AX054/10/S. For questions phone (610) 448-3237. |  | 286 | 286 units | USA - Nationwide Distribution | Terminated | September 29, 2012 | 704 | 286 | Not_Computer
Z-0147-2011 | 56844 | chair, with casters | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Recalled chairs include the following: 2008 Model Numbers: 653N - Care Cliner, 654N - XL Care Cliner, 655N - Drop Arm Care Cliner, 657N - XL Drop Arm Care Cliner, 690N - Elite Care Cliner, 691N - XL Elite Care Cliner, 694N - Dual Swing-Arm Elite Care Cliner, 695N - XL Dual Swing-Arm Elite Care Cliner. 2009 Model Numbers: 6530 - Care Cliner, 6540 - XL Care Cliner, 6550 - Drop Arm Care Cliner, 6570- XL Drop Arm Care Cliner, 6900 - Elite Care Cliner, 6910 - XL Elite Care Cliner, 6940 - Dual Swing-Arm Elite Care Cliner, 6950 - XL Dual Swing-Arm Elite Care Cliner. | Care Cliner | 2 | October 27, 2010 | 2010 | Winco, Inc. | Winco Mfg., LLC, Ocala, Florida is recalling their Care Cliner. Due to malfunctioning axles which become lose from the caster assembly. This chair is a multi position recliner with a steel frame used primarily in clinics, hospitals and treatment centers for recovery, dialysis, oncology, examination, infusion. Recalled chairs include the following: Care Cliner, XL Care Cliner, Drop Arm Care Cliner | DESIGN: Device Design | Winco sent an Urgent Device Recall letter dated October 1, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Replacement parts were sent out to the consignees and customers were asked to dispose of the defective casters. Any consignee not reponding by November 15, 2010, would receive an additional letter. The firm placed a follow up phone call to the non responding consignees after two weeks.Customers were advised to call 800-237-3377 or email customer service at customerservice@wincomfg.com with questions or concerns.For questions or concerns regarding this recall call 352-854-2929 x 110. |  | 18522 | 18,522 casters | Worldwide Distribution - USA (all states except Vermont and Nebraska) and the countries of Belgium, Canada, Chile, Ireland, Kuwait, Netherlands, Saud Arabia, Singapore and United Arab Emirates. | Terminated | May 20, 2011 | 205 | 18522 | Not_Computer
Z-0149-2011 | 56718 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Non-DEHP Y-Type Catheter Extension Set, 5.3" (13.5 cm), 1.0 mL, Male Luer Lock Adapter with Retractable Collar; a sterile Rx device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product 2N1196For the administration of drugs and solutions to a patient's vascular system through a vascular access device. | Baxter NonDEHP YType Catheter Extension Set | 2 | October 27, 2010 | 2010 | Baxter Healthcare Corp. | During packaging testing, Baxter observed tears and pinholes in Chevron-style pouches for certain infusion disposable products. The pinholes and tears could compromise the sterile barrier properties of the pouch. | DESIGN: Packaging Design/Selection | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 9, 2010, to all customers. The letter identified the product, the problem and the action to be taken by the customer. The customers were requested to examine their inventory in all locations to determine if they had any of the affected lots of the affected product codes and to remove the affected product from their inventory; they were to call the Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they could comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 413750 | 413,750 sets | Worldwide Distribution - U.S., Canada, Japan, Australia, New Zealand, and Chile. | Terminated | September 07, 2011 | 315 | 1607774 | Not_Computer
Z-0159-2011 | 55712 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify & Record System, for MOSAIQ, manufactured by Elekta Impac Software, Sunnyvale, CA. The intended use of Sequencer (generically referred to as a verification system) is to assist in the process of process of patient treatment on a radiotherapy treatment machine. | Elekta Impac Software Sequencer | 2 | October 28, 2010 | 2010 | Impac Medical Systems Inc | Mistreatment-- The database conversion utility used during the upgrade errantly changes the start gantry angle to zero. Any affected treatment fields, if not already treated, would start delivering from zero instead of the original prescribed value. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | ELEKTA IMPAC SOFTWARE sent a USER NOTICE dated February 19, 2010, to all customers. The notice identified the product, the problem, and the action to be taken by the customers.Customers were urged to verify the start angle of the gantry against the treatment field definition prior to delivery.Review all treatment plans and charts with conformal arcs. If a treatment plan with an errant gantry start angle was identified, customers were to cease using that plan immediately. A workaround was to be re-imported into their current version of MOSAIQ.IMPAC Support was to follow-up with each site and execute a script to detect and correct all affected treatment plans.For questions, concerns, and requests for upgrade, customers were to email support@impac.com or call their support center. |  | 19 | 19 units. | Worldwide Distribution -- USA, Australia, Belgium, France, Germany, Hong Kong, Ireland, Israel, and Italy | Terminated | March 28, 2011 | 151 | 19 | Software | Treatment/Delivery/Therapy | Upgrade | Repair
Z-0160-2011 | 56148 | computer, diagnostic, pre-programmed, single-function | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Edwards' MHANM8 Pressure Slave Cable, Part Number:693657001The Edwards Lifesciences Analog Slave Cable is used to transfer monitoring output from patient monitors to the analog input ports of the Vigileo Monitor series or compatible Edwards monitoring systems. This cable has instrument specific connectors. | Edwards' MHANM8 Pressure Slave Cable | 2 | October 28, 2010 | 2010 | Edwards Lifesciences, Llc | The recall was initiated because the Edwards' MHANM8 Pressure Slave Cable was wired incorrectly, thus will not work as stated. | DESIGN: Process Design | Edwards Lifesciences contacted customers via telephone on June 22, 2010 informing them of the product, the problem and the action they should take. Customers were asked to locate their cables and send them back overnight using the provided Edwards FedEx account number.Edwards Lifesciences indicated their personnel would follow-up with each customer until all units were returned.For any questions regarding this recall call (949) 250-2007. | Contact the recalling firm for information | 20 | 20 units | Worldwide Distribution - USA including CA, PA, and MI and the countries of Switzerland and Thailand | Terminated | May 24, 2011 | 208 | 20 | Not_Computer
Z-0162-2011 | 56826 | accessories, catheter | KGZ | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | IV Catheter - Optiva IV Catheter. This is found in box labeled 'Graft Jacket Xpress Flowable Soft Tissue Scaffold'; labeled by Wright. | Graft Jacket Xpress Flowable Soft Tissue Scaffold OPTIVA IV catheter component | 2 | October 28, 2010 | 2010 | LifeCell Corporation | Graft Jacket Xpress kits contain catheters that will expire or have already expired prior to the labeled expiration of the human tissue component of the product on the outer box. | PRODUCTION CONTROLS: Labeling Mix-Ups | Lifecell Corporation, Branchburg, NJ initiated a recall to their distributor Wright Medical Technology, Arlington, TN on 9/13/2010. |  | 1507 | 1507 kits | Nationwide Distribution: USA, to distributor in TN | Terminated | August 17, 2012 | 659 | 1507 | Not_Computer
Z-0167-2011 | 56975 | calibrator, dose, radionuclide | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Captus 3000 Thyroid Uptake System with Portable Stand option. Capintec item number 5430-2130, serial numbers: 91001, 91003 through 91010. Manufactured by Capintec, Inc, Pittsburgh, PA. | Captus 3000 Thyroid Uptake System with Portable Stand | 2 | October 28, 2010 | 2010 | Capintec Inc | The spring arm on portable stand may have been over tightened during assembly, which may result in failure. If the spring arm fails, it will not balance the collimator weight, and the arm will drop to the lowest position. | TRAINING: Employee Error | Capintec issued an Alert letter to all customers identifying the affected device, informing them of the problem and what actions are to be taken by customers. Capintec will send a replacement arm for the system, which is to be installed by the customer using included instructions. The firm included a correction acknowledgement form to be completed and returned by the customer.Customers can contact Capintec concerning this recall at 201 825-9500. | Contact the recalling firm for information | 9 | 9 units | Worldwide Distribution: USA in the states of NC and TN, and in the countries of Pakistan, Paraguay, Portugal, Saudi Arabia, and Spain. | Terminated | July 05, 2012 | 616 | 9 | Not_Computer
Z-0161-2011 | 56820 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Licox Brain PMO Probe Kit;Ref IP2.P;Multiparameter Monitoring Oxygen + ICP Channel: Contain Products: REF IP2 Bolt kit for 2 brain probes, Ref CC1.P1 combined Oxygen & Temperature probe;Contains sterile and non-pyrogenic if package is unopened and undamaged.Integra US, 311 Enterprise Drive, Plainsboro, NJ 08536 | Licox Brain PMO Probe Kit; for Cerebral Pressure, Ref IP2.P; | 2 | October 28, 2010 | 2010 | Integra LifeSciences Corp. | Incorrect expiration date - The outer boxes of the IP2P Licox Complete Kit for Cerebral Pressure, Temperature and ICP Monitoring Probes, Lot 250210A are labeled with an incorrect shelf life date (10 months too long). | TRAINING: Employee Error | Integra issued Urgent Product Recall Notification letters dated September 16, 2010 to US customers and foreign consignees, identifying the affected lots and actions to be taken by customers.Affected product is not to be further distributed and should be quarantined pending return to the firm. Customers and Distribors should contact Integra Customer Service for a Returned Material Authorization Number, and then return the product to the Distribution Center. In addition, a Recall Acknowledgement and Return Form is to be completed and returned.Customers can contact Integra at 609 936-2495. |  | 23 | 23kits | Worldwide Distribution: USA, including the states of FL, ID, IL, and NJ, and the countries of Argentina and Costa Rica. | Terminated | April 18, 2012 | 538 | 23 | Not_Computer
Z-0163-2011 | 56852 | system, radiation therapy, radionuclide | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion with EXTEND Frame System, manufactured by Electa Instrument AB, Stockholm, Sweden. | Leksell Gamma Knife Perfexion with EXTEND Frame System | 2 | October 28, 2010 | 2010 | Elekta, Inc. | The anterior fixation bar of the EXTEND mouthpiece frontpiece may separate from the dental impression tray. This may result in reduced treatment accuracy. | OTHER/UNDETERMINED: Pending | Elekta issued an Important Notice: Urgent Field Safety Notice dated September 17, 2010 concerning Field Change Order 10000202003-A342 "Important Notice regarding recall of mouthpieces for EXTEND system, to all users of the EXTEND system for LGK Perfexion. the letter identified the issue and associated potential hazard, and requested the identification and return of all affected mouthpieces, and confirmation of receipt of the letter. Elekta will arrange return and replacement through local service representatives.Customers can contact Elekta concerning this recall at 770 300 9725. |  | 3 | 3 units | Nationwide Distribution: to the states of MO, VA, and WA. | Terminated | November 10, 2010 | 13 | 3 | Not_Computer
Z-0165-2011 | 56908 | applier, hemostatic clip | HBT | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | LeMaitre Vascular AnastoClip GC Vessel Closure System-Large SizeModels:4007-03 Large size 3 Pack4008-03 Large size 10 Pack4007-07 Large size 1 PackManufactured by LeMatire Vascular Inc, Burlington, MA 01803 | LeMaitre Vascular AnastoClip GC Vessel Closure SystemLarge Size Models: 4007 | 3 | October 28, 2010 | 2010 | Lemaitre Vascular, Inc. | Mislabeled size on external packaging: The contents section listed on the Tyvek lid labels of the Medium and Large sizes incorrectly identified the size of the clip as XL, while the device inside was either a Medium or Large size. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | LeMaitre Vascular issued an Urgent AnastoClip GC Device Field Safety Notice letter dated Sepember 28, 2010 identifying the affected devices and actions to be taken by customers. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. The letter provides an option for the notified hospital to return the affected product for replacement if unused or if they simply wish to have it corrected. Customers can contact LeMaitre Customer Service for replacement at 800 628-9470. | Contact the recalling firm for information | 108 | 108 units | Nationwide Distribution: to the states of CA, FL, GA, IL, IN, LA, MD, MS, NJ, OK, OR, and TX. | Terminated | September 15, 2011 | 322 | 108 | Not_Computer
Z-0166-2011 | 56352 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Table, Part No. 7346534, Distributed by and/or Manufactured by Siemens AG, Medical Solutions, Germany | Siemens brand 550 TxT Patient Table | 2 | October 28, 2010 | 2010 | Siemens Medical Solutions USA, Inc | An increase in voltage may cause a short circuit of a power transistor that may cause a smoldering and smoking effect of the electrical wiring and also an unexpected lowering of the Table. Both issues present the potential for injury to a patient or the operator. | OTHER/UNDETERMINED: Pending | Siemens sent a Customer Information Letter on March 31, 2010, to all consignees. The letter identified the product, the problem, and the action that should be taken. Siemens implemented minor modifications to address the issues. No further action was required for the continual operation of the 550TxT Table. For questions regarding this recall call 925-602-8083. |  | 326 | 326 units | Worldwide Distribution including USA | Terminated | March 01, 2012 | 490 | 326 | Hardware | Physical Safety Hazards | Correction | Repair
Z-0170-2011 | 56909 | cardiac ablation percutaneous catheter | LPB |  |  | Cardiovascular |  | Boston Scientific Peristaltic Tubing Kit, Model 2104, UPN M0043521040. For use with Boston Scientific CHILLI Cooled Ablation Catheters and Model 8005 Continuous Operation Pump. Sterilized using Ethylene Oxide. Contents: peristaltic pump fluid delivery extension tube set. Manufactured for: Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, USA.Tubing Set (model 2104) for use with Boston Scientific Cooled Ablation System. | Boston Scientific Peristaltic Tubing Kit, Model 2104 | 2 | November 01, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is conducting a Recall on CircuCoolÂż Tubing Sets used with the Chilli IIÂż Cooled Ablation System because there may be a potential for a sterile barrier breach which could compromise device sterility and potentially the sterile field. Even though there have been no reported complaints. | PRODUCTION CONTROLS: Packaging | Boston Scientific Corporation sent an Urgent Medical Device Recall letter, dated September 27, 2010, to all customers via overnight Federal Express. The letter described the product, the problem, and the action the customer should take. The letter instructed customers to cease distribution of affected devices and remove them from their inventory, segregate in a secure location for return to Boston Scientific. A Reply Verification Tracking Form was to be completed and returned via email or faxed to (510) 624-2475 even if the customer did not have any of the affected product. If affected units were located the customer was to call (510) 624-2563 to obtain a Returned Good Authorization (RGA) Number and return to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: _________________For additional questions call (510) 624-2563. | Contact the recalling firm for information | 4694 | 4031 US , 663 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, OH, PA, RI, SC, TN, TX, VA, WA, and WV and the countries of France, Germany, Great Britain, Hungary, Israel, Italy, Netherlands, Saudi Arabia, Spain, and Taiwan | Terminated | April 17, 2012 | 533 | 4694 | Not_Computer
Z-0169-2011 | 56858 | [6 empty] | Bone Void Filler, sterile orthopedic product, 510 (k) #K053228. The product is delivered in a syringe which is packaged in an aluminum-backed polyethylene pouch and is distributed one (1) pouch per box. The product affected by this recall was distributed by the recalling firm under six (6) different names. They are: Form-Putty; SBvFShape; Valeo BP; X-Putty; TheriGraft Putty; and SynFlow. The product is indicated for use in filling the gaps or voids of osseous defects surgically created or resulting from trauma and are intended for treatment of osseous defects not intrinsic to the stability of the bone structure. The products are intended for use in defects of the skeletal system (i. e. the extremities, spine and pelvis). | Bone Void Filler | 2 | November 01, 2010 | 2010 | Theken Spine Llc | The firm's quality control testing determined that a separation was occurring between the flowable carrier and the B-TriCalcium Phosphate within some of the product syringes. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Theken Spine Llc sent Recall Notification Letters dated August 26, 2010, to all customers. One (1) document, titled an 'Urgent Product Recall Notification' letter was addressed: 'ATTENTION; PRODUCT RECALL COORDINATOR/RISK MANAGER.' The other letter was titled: 'RECALL NOTIFICATION.' The recall letters were accompanied by an attached 'Recall Acknowledgement and Return Form' and were sent out via Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. The letters requested that the customers review their product in inventory and identify any product remaining from the affected lot numbers; part numbers; and/or item names for immediate return of these items to the recalling firm. The letters also asked that the customer complete and return the attached Recall Acknowledgement and Return form according to the instructions provided on the form. In addition, the 'RECALL NOTIFICATION' letter instructed the distributor/customers provide information concerning their sub-account customers to Theken Spine so they could issue a Recall Notification letter to these customers as well. To obtain instructions from the recalling firm concerning product return and issuance of credit, customers were instructed to call (866) 942-8698. For questions regarding this recall call (330) 825-8258. | No consumer action necessary | 2904 | 2904 units (syringes) | Nationwide Distribution - including UT, SD, TX, AL, PA, CA, MI, OH, NJ, NH, MD, ID, FL, LA, NM, TN, IA, and OR. | Terminated | October 27, 2011 | 360 | 2904 | Not_Computer
Z-0171-2011 | 56921 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP and Mobilett XP Hybrid, Mobile radiology deviceSystem, x-ray, mobile | Mobilett XP | 2 | November 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Units not equipped with protective plastic cover over electronics. | DESIGN: Device Design | Siemens Medical Solutions USA, Inc issued a mandatory field update. The letter identified the product, the problem and the action to be taken by the customer. All systems that were not yet equipped with protective cover over the D972 board would be upgraded.For questions regarding this recall call (610) 448-3237. |  | 177 | 177 units | Nationwide Distribution including AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA, WA, WI, WY AND PUERTO RICO. | Terminated | September 29, 2012 | 698 | 177 | Not_Computer
Z-0168-2011 | 56794 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i imaging system Family Software version 3.2, Part Numbers 435-0602.01 to 05, including ChromaFlo and Virtual Histology, manufactured by Volcano Corporation of Rancho Cordova, CA. | Volcano s5/s5i imaging system software version 3.2 | 2 | November 01, 2010 | 2010 | Volcano Corporation | Software defect results in improper image brightness, which may result in introduction of artifacts, which may cause misclassification of results and affect classification of tissue types. This may result in misdiagnosis and inappropriate treatment. | DESIGN: Software Design | Volcano issued a Voluntary Recall of Volcano s5 Family Software version 3.2 letter dated June 30, 2010, identifying the software defect and actions to be taken by Volcano. Field Service Engineers will visit each site and reinstall the previous software version, which will resolve the issue.Customers can contact Volcano at 916 281-2790. |  | 79 | 79 software kits | Worldwide Distribution: USA, and the countries of France, Germany, Japan, and the UK. | Terminated | December 22, 2010 | 51 | 79 | Software | Display/Image | Reinstall previous version | Software Update
Z-0134-2011 | 56902 | instruments, surgical, cardiovascular | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medtronic Octopus Nuvo Tissue Stabilizer, model TSMICS1. The Octopus Nuvo Tissue Stabilizer is a disposable tissue stabilization system consisting of a rigid shaft, detachable tissue stabilizer headlink and a flexible articulating arm. The device is assembled within the thoracic cavity. Sterilized using ethylene oxide, non pyrogenic, disposable, single use only. The product is intended to stabilize and minimize the motion of selected areas of the beating heart during minimally invasive cardiac surgery procedures under direct visualization through a thoracotomy. | Medtronic Octopus Nuvo Tissue Stabilizer | 1 | November 01, 2010 | 2010 | Medtronic Perfusion Systems | Medtronic has received reports regarding the Octopus Nuvo Tissue Stabilizer stating that during use, the collet (a component of the device that stabilizes the head link) fractured, thereby causing immediate separation of the head link from the shaft of the device. The resulting potential hazards are that fragments of the collet could fall into the patient's chest cavity and/or damage the heart ti | DESIGN: Component Design/Selection | Medtronic sent an "Urgent Medical Device Recall Notice" dated September 14, 2010, to Risk Managers at each affected account. The letter described the issue, identified affected lot numbers, asked for immediate discontinued use of the device, quarantine all unused devices and to return unused devices to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428 and requested a recall certificate be faxed back to Medtronic indicated acknowledgement of the Recall Notice and quantity to be returned.If you need additional information, please contact Medtronic Cardiovascular Lifeline Technical Services at 877-526-7890, or your local Medtronic sales representative. | Contact the recalling firm for information | 571 | 571 units | within the US to: AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and Washington D.C. OUS to include: Belgium, Canada, Germany, Italy, Poland, Switzerland, United Kingdom. | Terminated | April 16, 2012 | 532 | 571 | Not_Computer
Z-0288-2011 | 53863 | display, cathode-ray tube, medical | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | The product is a 19" Spacelabs Medical Flat Panel Display, Model 94260-19, manufactured by ELo Touchsystems | Spacelabs Medical Flat Panel Display, Model 9426019 | 2 | November 02, 2010 | 2010 | Spacelabs Healthcare, Incorporated | The Spacelabs Medical Flat Panel Display, Model 94260-19 may fail prematurely and causing the screen to go blank. This inability to view data may result in delayed patient treatment. |  | On 10/23/09, Spacelabs issued URGENT: MEDICAL DEVICE RECALL notification letters dated October 23, 2009 to their consignees.The notification states the recalled product as Spacelabs Medical 19" Flat Panel Display, Model 94260-19. The consignees are informed that there is apossibility that the internal power distribution circuit board will fail prematurely, causing the display to go blank. Audio alarms will continue to function normally.The firm's Field Service Engineer will be contacting the consignees to set up a convenient time to schedule the upgrades of the affected displays. Displays with part number ET1920L-7SKR-1-35S-SPC are not part of this recall and should not be sent in for upgrading.The firm recommends that their consignees brief their staff on the possibility of the display going blank to help prevent an adverse event from happening. The firm will be providing loaners while the affected displays are upgraded at free of charge.If the consignees have any questions about the recall, they could contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support . Consignees outside the United States may call 1-425-657-7200 x 5089. | Contact the recalling firm for information | 2560 | 2560 units (2020 domestic, 540 international) | Worldwide Distribution: USA in the states of AK, AL, AR, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WY, and the following countries: Australia, Austria, Canada, Czech Republic, Egypt, France, Germany, Grand Cayman, Greece, India, Israel, Italy, Kuwait, Lebanon, Mexico, Netherlands, Pakistan, Panama, Poland, Puerto Rico, Saudi Arabia, Singapore, Slovakia , South Africa, Switzerland, Trinidad and Tobago, Turkey, United Kingdom, and Venezuela. | Terminated | November 03, 2010 | 1 | 2560 | Hardware | Display/Image | Replace | Remove or Replace
Z-0308-2011 | 55242 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Panta Support Device PantaÂŽ Arthrodesis Nail Instrumentation Tray One unit per package.The PANTA(R) Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. | Panta Arthrodesis Nail | 2 | November 02, 2010 | 2010 | Integra LifeSciences Corporation | Inability to properly align the calcaneal screw placement on a consistent basis, due to a manufacturing defect of of the targeting frame (Panta Support Device). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | OIntegra issued Urgent Product Recall Notification letters dated 3/1/2010 to customers and distributors to notify them of the problem. Customers were requested to check their inventory to determine if they have an of the affected devices. If so, custo0mers should stop using the product. Custoemrs with the affected product should contact their local Sales Representative for assistance, or contact Integra at 800-654-2873 - 3 to obtain a returned material authorization number, and return the product to their local representative.All customers were requested to complete and return a Recall Acknowledgement and Return Form.. |  | 108 | 108 Units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VA, WA and WI; and countries of: Australia, Austria, Belgium, Czech Republica, Denmark, Finland, France, Germany, Ireland, Italy, Portugal, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | February 16, 2012 | 471 | 108 | Not_Computer
Z-0293-2011 | 56938 | implant, intragastric for morbid obesity | LTI |  |  | Gastroenterology/Urology |  | REALIZEÂŽ Adjustable Gastric Band Pak with Endoscopic Dissector and Calibration Tube Product Code RLZB22DG1, Lot #: G4RJ98, G4T42A, G4T88X, G4TH90, G4TJ61, G4TP6J, G4TR76, G4TV48, G4TX1V | REALIZE Adjustable Gastric Band Pak with Endoscopic Dissector and Calibration Tube | 2 | November 02, 2010 | 2010 | Ethicon Endo-Surgery Inc | The Tubing Strain Relief component of the Injection Port may detach from the Locking Connector, and migrate from its original position on the Locking Connector. This may result in: Tubing Strain Relief migrating on the Tubing; Kinking of the Tubing resulting in difficulty/inability to adjust the fluid volume within the Band; or a free foreign body in the abdominal wall or abdominal (peritoneal) ca |  | Ethicon Endo-Surgery issued an Urgent : Devoced Recall Event 2240 letter dated October 8, 2010 to all consignees identifying the affected devices and actions to be taken.Consignees were instructed:DO NOTt use product from the affected lots, but implanted devices do not need to be removed as part of this recall.Examine inventory and remove the affected product.If they DO have the affected product, complete the Business Reply Form, keeping a copy for records, and place the original form in the box with the product to be returned, affix the pre-printed label, and return to Stericycle.If they DO NOT have the affected product, complete the Business Reply Form, and fax it to Stericycle.A redesigned replacement product will be sent out to consignees within 30-60 days of product return.Ethicon Endo-Surgery can be contacted at 513 337-8208 concerning this recall. | Contact the recalling firm for information | 105 | 105 pieces | Worldwide Distribution: USA including the following states: AK, AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, VT, WA, WI and WV, and the following countries: ARAB EMIRATES, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LIBYA, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, RUSSIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY and VENEZUELA. | Terminated | November 18, 2011 | 381 | 127130 | Not_Computer
Z-0129-2011 | 56735 | accessories, soft lens products | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AMO COMPLETE Multi-Purpose Solution, Formula: 9424x; Product Code: 93504 (360ml), 93505 (2x 360ml duo pack), 93515 (240 ml). Cleans, rinses, disinfects, and stores all soft contact lenses and helps keep your eyes protected from dryness and irritation throughout the day. | COMPLETE MultiPurpose Solution | 2 | November 02, 2010 | 2010 | Abbott Medical Optics Inc (AMO) | The recall was initiated after AMO determined, through internal testing, that a limited number of the flip top caps used during production of these solutions may leak and, although unlikely, the sterility of the product may be compromised. Products that are non-sterile have the potential to cause eye infections, which may be sight threatening. | PRODUCTION CONTROLS: Process Control | AMO prepared and issued on July 28, 2010 both a customer and distributor notification letter as part of their customer communication process. The letter informed the customers of the reason for recall, the products involved, the hazard involved and actions to be taken regarding the affected products.Customer with questions related to the COMPLETE Multi-Purpose Solution and/or Total Care (1) Contact Lens Solution return process, were instructed to contact an AMO Customer Service Representative at 1800 931 663. |  | 34080 | 34,080 units | International Distribution Only -- United Kingdom, Lebanon, Dubai, Bahrain, Russia, Kazakhstan, Ireland, and France. | Terminated | October 13, 2011 | 345 | 34224 | Not_Computer
Z-0309-2011 | 56837 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm LCP Periarticular Proximal Humerus PlatesIndicated for fractures and fracture dislocations, osteotomies, and nonunions of the proximal humerus, particularly for patients with osteopenic bone | 3.5mm LCP Periarticular Proximal Humerus Plates | 2 | November 02, 2010 | 2010 | Synthes USA (HQ), Inc. | screw can pass through plate hole | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Synthes (USA) notified all Synthes Trauma Hospital accounts of record of this recall via an Urgent Medical Device Recall letter dated September 27, 2010. The notification included instructions for hospital accounts to immediately cease use of the product, examine inventory, remove the 3.5 mm LCP Periarticular Proximal Humerus Plates from their shelves and complete and return the attached Verification Section form via mail: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; fax: 610-215-9005 or Scan/email: FieldAction@synthes.com. If they have any of the identified devices, they were to call Synthes at 1-800-479-6329 to obtain Authorization number prior to completing Verification form.If you have any questions, please call 1-800-620-7025 x5403 or contact your Synthes Trauma Sales Consultant. | Contact the recalling firm for information | 2849 | 2849 units | Worldwide distribution: USA and countries including: Bermuda, Korea, Italy, Japan, Georgia, Qatar and Turkey | Open |  |  | 2849 | Not_Computer
Z-0287-2011 | 56324 | permanent pacemaker electrode | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | enpath medical Myopore Sutureless Myocardial Pacing Lead, Bipolar (labeled as Unipolar) Connector, , REF 511212. 54 cm. Distributed by Guidant/CPI. Contents in unopened, undamaged package have been sterilized by ethylene oxide gas. Manufactured by enpath medical, Inc., 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. | enpath Myopore Bipolar 54 cm Epicardial Lead Sutureless Myocardial Pacing Lead | 2 | November 02, 2010 | 2010 | Greatbatch Medical | The Myopore Bipolar Epicardial Leads packaging is labeled as unipolar leads. The lead body itself is correctly labeled as bipolar. This could result in a bipolar lead being implanted with a pulse generator operating in unipolar mode. | TRAINING: Employee Error | Greatbatch Medical issued an Urgent Medical Device Recall letter dated May 13, 2010 to their distributor. The letter described the problem and product involved, and gave the action required which included:1) immediately discontinue distribution of the units and return any unused units to Greatbatch Medical.2) Reconcile medical recordsbetween BSC and Greatbatch Medical to confirm implant data3) Confirm pulse generator type (bipoloar, unipolar, or unknown) used with each serial number listedGreatbatch will work with customers to contact physicians as necessary regarding patient monitoring. Greatbatch Medical can be contacted at 763-951-8312. | Contact the recalling firm for information | 34 | 34 | Worldwide Distribution: to USA and Europe, through distributor in MN. | Terminated | May 10, 2012 | 555 | 34 | Not_Computer
Z-0311-2011 | 55287 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | microTargeting Drive System with Mounted AccessoriesCatalog numbers : 66-ZD-MD-01, MT-DS-01, FC1006 (Medtronic)Product is the FHC DBS Depth Stop Adapter, a component of the microTargeting Drive and the STar Drive. The Drives are intended to be used with commercially available stereotactic systems for neurosurgical procedures which require the accurate positioning of microelectrodes, stimulating electrodes, or other instruments in the brain or nervous system. | microTargeting Drive System | 2 | November 03, 2010 | 2010 | FHC, Inc. | Fixation thumbscrew on the DBS depth stop adapter may be overtightened and damage the implantable lead. | DESIGN: Device Design | FHC issued an "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 29, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. FHC informed the customers that that they had returned to the original design of the fixation screw in the Depth Stop Adapter. FHC sent replacement Depth Stop Adapter with an original fixation screw. The customers were instructed to discard the recalled Depth Stop Adapter (s) and replace with the enclosed corrected version(s) and sign and fax the receipt notification to 207-666-8309 as proof of their receipt of notification. For assistance with the product replacement, contact the FHC's technical support line at 800-326-2905. This line is staffed 24 hours a day.Should you have any further questions regarding this matter, please contact Dorothee Sasse, product manager, or Keri Seitz, recall coordinator at 207-666-8190. |  | 18 | 18 units | Worldwide distribution: USA in the states of AZ, CA, FL, ID, NC, MA, MI, MN, NV, RI, UT, VA, and WA, and countries of Columbia and South Africa. | Open |  |  | 63 | Not_Computer
Z-0316-2011 | 55748 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Sample Prep Assistant III; (BD SPA III), Manufactured by BD Biosciences, San Jose, CA.The intended use for the BD FACS Sample Prep Assistant is to prepare human whole blood samples for flow cytometry analysis. | BD FACS Sample Prep Assistant III (BD SPA III) | 2 | November 03, 2010 | 2010 | BD Biosciences, Systems & Reagents | Due to inaccurate sample dispense, use of product may result in getting incorrect results on samples prepared using the device. Having a piece of the lid partially or fully clogging the Probe during piercing a hole in the lid of tube may cause inaccurate sample dispense, resulting in incorrect test results.The piece of the lid clogging the Probe can be in the processed sample tube. It has the | DESIGN: Device Design | BD Biosciences sent an "Important Product Information" letter dated February 2010 and an "URGENT: Follow-Up to Letter Dated February 2010" letter dated July 2010 to customers identifying the affected device and actions to be taken: Customers were instructed to stop use of the SPA probes and to contact BD for replacement of the affected probes and assistance in the event of probe clogging. BD has arranged replacement of affected probes. BD's followup communication detailed the potenial hazards and reiterated remedial actions in the event of error messages from FACS flow cytometers.Customers with questions about the performance of theirr SPA instrument or the content of the letters may contact BD Customer Support at 1.877.232.8995, prompt 2. |  | 349 | 349 units | Worldwide distribution: USA including OH, NC, AL,TX, FL, CA, NY, and MN and the country of Belgium. | Terminated | May 18, 2011 | 196 | 349 | Not_Computer
Z-0315-2011 | 56499 | prosthesis, hip, hemi-, femoral, metal/polymer, cemented or uncemented | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Centralizer 10 mm | Ortho Development Primaloc Cemented Hip System | 2 | November 03, 2010 | 2010 | Ortho Development Corporation | The hip stem did not fit the distal centralizer because the hole was not drilled with the appropriate taper. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ortho Development sent a Field Notification Memo dated July 28, 2010 to Japan Medical Dynamic Marketing instructing them to reconcile parts that may have been used and return unused parts to Ortho. Customers can contact Ortho Development at 1-801-619-3419. |  | 39 | 39 units | Worldwide Distribution: Japan | Terminated | November 04, 2010 | 1 | 39 | Not_Computer
Z-0317-2011 | 56668 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT cTnl cartridgeAbbott Point of Care,Princeton, NJ 08540 | iSTAT cTnl cartridge | 2 | November 03, 2010 | 2010 | Abbott Point Of Care Inc. | i-Stat cartridges may not fill, cartridge labels or channels may appear discolored, cartridges may have fluid visible in the cartridge channel, and cartridge packaging may be distended. | OTHER/UNDETERMINED: Pending | Abbott Point Of Care Inc.sent notification letters with a customer reply form to all affected customer in January, 2008. All distributors outside the US who were sent shipments from the affected lots were notified of the lot numbers that were affected. Affected product is to be quarantined and returned for replacement. Customers with questions about this recall can contact Abbott Point Of Care at 609 454-9000. |  | 482500 | 482,500 (total number) | Worldwide Distribution: USA and Canada. | Terminated | September 14, 2011 | 315 | 482500 | Not_Computer
Z-0310-2011 | 56825 | negative pressure wound therapy powered suction pump | OMP | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | smith&nephew; EZCARE Negative Pressure Rx only. 100-240 V 12 V 50/60 Hz 90VA T1A 250 VIndicated for patients who would benefit from a suction device particularly as the device may promote healing when used with accessory wound sealing kits. Made in USA by/Fabrique aux Etals Unis par: Smith & Nephew, Inc., Largo, FL 33773 Smith & Nephew Medical Limited Hull HU3 2BN England | Negative Pressure Wound Therapy Powered Suction Pump | 2 | November 03, 2010 | 2010 | Smith & Nephew Inc. | Smith&Nephew; Inc., Advance Wound Management, St Petersburg, FL is recalling their Ezcare Negative Pressure Wound Therapy pumps manufactured by BlueSky Medical Group Inc., between July 17, 2007 and December 13, 2007 and distributed by Smith&Nephew; between July 18, 2007 and January 14, 2010. Affected products codes are 66800187, 66800322, 66800323, 66800324, and 66800325 with Serial Numbers startin | OTHER/UNDETERMINED: Pending | Smith & Nephew Inc. FEDEXed overnight an URGENT DEVICE RECALL letter dated January 27, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were to check their inventory and identify any affected product, fill out and return the response form by both mail and fax, and return the affected product with the prepaid Federal Express (FedEx) Ground label provided by Smith & Nephew. For any questions regarding this recall contact the Customer Care Center at (800) 876-1261. |  | 2717 | 2717 units | Worldwide Distribution - US including AL, AR, AZ, CA, GA, IL, IN, MA, MD, MD, MO, NC, NJ, NY, OH, OR, PA, TN, TX, and UT and the countries of Australia, UK, South Africa, Canada, Puerto Rico, and New Zealand. | Terminated | November 10, 2010 | 7 | 2717 | Not_Computer
Z-0320-2011 | 52231 | accessories, implant, dental, endosseous | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann T-Sleeve 2.8mm, height 6mm, guidedArticle Number: 034.055V46Intended for use as a drilling guide to prepare a site for placement of a dental implant into the dental implant site. | Straumann TSleeve | 3 | November 04, 2010 | 2010 | Straumann Manufacturing Inc. | Incorrect Inner diameter of the T-sleeve | PRODUCTION CONTROLS: Process Control | Strauman USA LLC notified the consignee by telephone on June 4, 2009, and sent an Urgent Device Recall letter dated, June 8, 2009 as a follow-up. The letter described the product, the problem, and the action to be taken by the customer. The customer was asked to return any affected product, in the prepaid preprinted mailer pouch as soon as possible to:T-sleeve Recall, Straumann USA, LLC60 Minuteman RoadAndover, MA, 01810Customers were asked to sign and indicate on the Recall Acknowledgement Form if they would like a replacement product or a credit to their account for the product returned and fax the form to (978) 747-0023.For any questions regarding this recall call (978) 747-2575. | Contact the recalling firm for information | 3 | 3 units | Worldwide Distribution - Canada | Terminated | February 20, 2012 | 473 | 3 | Not_Computer
Z-0321-2011 | 56785 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Single-stage 3.5mm Healing Abutment, 4 mm, REF SYHA40, Rx Only, STERILE R, gamma irradiated, expires 2014-44, CE 0473, BIOHORIZONS, BIRMINGHAM, AL 35244.Endosseous dental implant healing abutment. | Singlestage 3.5mm Healing Abutment, 4mm | 2 | November 04, 2010 | 2010 | BioHorizons Implant Systems Inc | Healing abutment may not fully seat on the implant. | OTHER/UNDETERMINED: Pending | Consignees were notified by an Urgent: Medical Device Recall letter on 09/07/2010. The letter identified the affected product and gave some background regarding the issue. It also stated the possible impact and the cases that were not affected. Consignees were instructed to contact BioHorizons Returns & Complaints Manager Rick Self at 205-986-1250, if they have unused product from the affected lot, for a Return Authorization Number and to receive replacements. The attached Notification Certification form is to be completed and returned via fax. | Contact the recalling firm for information | 396 | 396 | Worldwide Distribution -- USA, including PA, MI, NY, AL, WA, and WV, and countries of Canada and Chile. | Terminated | December 13, 2010 | 39 | 396 | Not_Computer
Z-0305-2011 | 56164 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | B. Braun addEASE Binary Connector, Catalogue #N7990, 200 units (50 units per bag with 4 bags per shipping box) Intended for use in a pharmacy setting to connect a B. Braun 250mL Excel IV solution bag to a 20 mm drug vial for reconstituting or mixing the drug in the vial with the solution in the bag. | B. Braun addEASE Binary Connector | 1 | November 05, 2010 | 2010 | B. Braun Medical Inc | During this insertion, fragmentation of the PAB container stopper may occur resulting in a small amount of visible particulates in the solution. | DESIGN: Device Design | On July 8, 2010 B Braun sent an Urgent - Medical Device Recall letter, dated June 28, 2010, to consignees informing them of the recall and that further distribution or use of the affected product should cease immediately. The attached "Product Removal Acknowledgement" form should be used to record the number of units and full-unopened cases that customers have in their inventory. The completed form should be returned via mail or fax within 2 weeks of receipt. If consignees have any full or partial cases of these affected products, they are to please call the Customer Support Department at (800) 227-2862 to arrange for return. Canadian customers may call (800) 624-2920. International customers may call (610) 691-5400. A Customer Support Representative will provide consignees with instructions for handling the affected product. The firm will arrange for unopened cases to be returned to B. Braun Medical, Inc. for proper disposal. Individual units (out of shipping cases) may be destroyed at the consignees facility.Consignees were informed that B. Braun is currently working to resolve this issue and provide an acceptable alternative to the market. They expect this product to be available in August 2010. In the interim, B. Braun PAB containers can continue to be used safely with a standard syringe and needle in accordance with the Direction for Use. If customers have additional questions regarding with a matter, they are to call B. Braun Customer Support Department at the number listed above. Should consignees identify product where this condition has occurred, consignees were asked to report it promptly to B. Braun Medical Inc. by contacting the B. Braun Clinical and Technical support at 1-800-854-6851. | Contact the recalling firm for information | 33905600 | 33,905,600 units | Worldwide Distribution -- USA and Canada. | Terminated | October 19, 2011 | 348 | 36405600 | Not_Computer
Z-0307-2011 | 56750 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | WalkMed Triton Pole Mount Infusion Pump, Catalog 300000.Indicated for delivering infusions of medications, as well as providing total parenteral nutrition, and patient controlled analgesia. | Volumetric Infusion Pump | 1 | November 05, 2010 | 2010 | Walkmed Infusion LLC | Pump door may be in a near shut position, but unlatched, and the "Door Open" alarm may not sound. | DESIGN: Device Design | All customers were notified by letter on 06/17/2010, describing the affected product and advising them of the possible malfunctioning alarm in the case a user does not follow the instructions as specified in the Triton Operation Manual. The letter also reminded all users to follow the procedures in the IFUs. The firm stated that it was best if customers re-educated all nurses on the proper door closure procedure to help prevent the potential bypass issue from occurring. Attached to the letter were two pages explaining the proper process for closing the pump door. Follow-up phone calls were made to coordinate an upgrade to all affected pumps. Questions or concerns should be directed to Ross Kurz at 720-531-4933. | Discard product or return to place of purchase for a full refund or exchange | 1565 | 1565 pumps | Nationwide Distribution -- AL, CA, FL, and PA. | Terminated | March 14, 2012 | 495 | 1565 | Software | Alarm/Message | Software update | Software Update
Z-0146-2011 | 56833 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- Common/Classification Name: Infusion Pump --Class 2 medical device -- SIGMA LLC, Medina, NY 14103 -- Shipping box label: Infusion Pump SIGMA Model 35700 (with accessories), SIGMA International, Inc. Medina, NY 14103 ** Serial Number: XXXXXX.To be used in a healthcare facility in an IV pole mounted configuration or carried by the user in an ambulatory manner. | Sigma Spectrum VOLUMETRIC INFUSION PUMP | 1 | November 05, 2010 | 2010 | Sigma International General Medical Apparatus, Llc | The infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. | DESIGN: Device Design | Urgent: Medical Device Recall Notification letters and forms (dated 9/14/10) were sent to the consignees via Certified Mail on/about 9/15/10. The letters inform the consignees of the recall of the pumps since these units may fail suddenly causing inaccurate flow conditions during use, ranging from back flow to over-infusion, including free flow. Consignees are instructed to check their inventory to verify whether their serial numbers fall within the range of pumps being recalled. Due to possible bearing failure, SIGMA is taking the precautionary measure of servicing or replacing all affected Spectrum Infusion Pumps and is requiring the return of the recalled devices to its facility. SIGMA has provided the following instructions for users: (1) Do not use the affected Spectrum Infusion Pumps on patient populations where inaccurate flow, ranging from back flow to over-infusion, including free flow, could result in serious adverse health consequences or death. (2) Do not use on neonatal patients. -- If the user facility has no acceptable alternative methods for delivering therapy, the user should weigh the risks and benefits for the continued use of the affected pumps. If you must use these pumps, SIGMA recommends the following actions which may mitigate the risk. (1) for pediatric and adult patients, an add-on Buretrol may limit the amount of an over-infusion if the failure occurs. Always follow the Buretrol manufacturer's instructions. (2) consider using air-eliminating filters. Under free flow conditions pumps cannot stop fluid flow. If air is in the line with fluid, the air will also infuse. (3) Frequently observe that the drip rate (mL/hr) compares to the programmed pump flow rate. Also, make an observation of the container volume to make sure it is correct. If a discrepancy is observed, immediately remove the pump from service and treat the patient as appropriate. There are additional instructions to Asset Management Customers: SIGMA requests that each A. | Contact the recalling firm for information | 17557 | 17,557 Infusion Pumps | Nationwide Distribution | Open |  |  | 17557 | Not_Computer
Z-0323-2011 | 55655 | stimulator, muscle, powered | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | The device is labeled in part: " RS Medical RS-4i Stimulator". The device is sold in a kit that includes a carrying case, connector cables, pads (electrodes), battery charger, patient document, instructional DVD, return mail bag, and the operation manual.This device is used with cables and electrodes that are placed on skin in the treatment area, and allow an electrical micro-current to relieve pain. The device requires a prescription from a physician. It is programmed by an RS medical employee (or Account Manager) to the physician's specific treatment plan for each patient. The patient is given instructions on how to utilize the device (proper electrode placement). Powered MUSCLE Stimulation indications for use including: 1) Relax muscle spasm. 2) Prevention or retardation of disuse atrophy. 3) Maintain or increase range of motion. 4) Increase local blood circulation. 5) Re-educate muscle. 6) Immediate post-surgical stimulation of calf muscles to prevent venous thrombosis. INTERFERENTIAL Stimulation indications for use including: 1) Relieve acute pain. 2) Relieve and manage chronic pain. | RS4i Sequential Stimulator; | 2 | November 08, 2010 | 2010 | International Rehabilitative Sciences, Inc. | The charge current is leaking through two Diodes, which are supposed to block the reverse current, when the device is warm during the main battery charge operation, it then causes the lithium battery to swell around two years of use and it causes the device to be inoperable. | OTHER/UNDETERMINED: Pending | Since June 11, 2008, RS Medical has taken the following actions when the device was sent back through a consumer complaint, or when the device was returned for routine maintenance, or through the rental service. The firm has upgraded the blocking diodes with higher temperature rated diodes, replaced the Lithium coin cell battery, and the software register bit has been changed to denote for subsequent automated testing that the diode upgrade/Lithium coin cell battery replacement has occurred. The firm reported that the device must be returned for the rechargeable NiMH battery replacement around two years of use, and the battery should not be replaced by patients.On 10/29/10, the firm began sending out URGENT NOTICE of VOLUNTARY RECALL letters to their customers. The firm will continue to send out the letters through 2011. The letter states the recall product as RS-4i Sequential Stimulator, Affected Serial Numbers 4i300000 to 4i467119. The customers are informed that "RS-4i Sequential Stimulators, only those within the affected serial number range indicated above, are susceptible to early expiration of the lithium coin cell battery that provides back-up power to the device memory when the stimulator is powered OFF. Customer experience indicates that the 3V lithium coin cell battery can be depleted as early as one third of its rated life. The RS-4i Stimulator will perform as intended until the lithium coin cell battery is fully depleted. When this happens, the Stimulator will turn OFF and fail to turn ON, even if it has been charged by battery charger connection to a standard electrical wall outlet".Customers are advised to make a schedule for a free Return and Repair of a RS-4i Sequential Stimulator and should call the Toll Free Service Line at 877-363-5753 (English) or 877-363-5754 (Spanish) from 6AM to 5PM (Pacific Time) Monday-Friday.If the customers have any other questions or to review additional information before scheduling a device . |  | 297 | 166, 131 units | Nationwide Distribution | Terminated | September 10, 2012 | 672 | 297 | Battery | Device Operation | Replace | Remove or Replace
Z-0328-2011 | 56628 | stretcher, wheeled | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker M1 Ambulance cot, Model 6100, Stryker Medical, Kalamazoo, MI. | Stryker Model 6100 M1 Ambulance Cot | 2 | November 08, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The firm has received 42 complaints regarding the head end lock rod, head end slide tube, brushing, and upper cross tube failures. | DESIGN: Device Design | Stryker Medical sent an Urgent Medical Device Alert letter dated May 15, 2010 to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter stated that there is a possibility for fatigue related failures of certain components of the cot and that a Stryker representative will visit the firm to correct the device. A representative from Stryker field service would contact the customer to perform the device corrections, free of charge.Immediate Action required by the customer:Locate the ambulance cots listed in this notice.Return the enclosed post card to confirm receipt of this notification.If any of the cots listed in the letter were loaned or sold the customer was to forward a copy of the notice to the new users and advise Stryker of their new location.If the customer disposed of any of the affected stretchers and they were no longer in use, customers were to advise Stryker of their obsolescence by providing the serial numbersFor any questions or comments regarding this issue, please call (800) 869-0770, ext. 6902. | Contact the recalling firm for information | 7314 | 7,314 | Worldlwide Distribution - US including AZ, CA, CO, FL, LA, MS, NC, NJ, OH, PA, and TN and the countries of Australia, China, France, Germany, Hong Kong, Italy, Japan, New South Wales, Poland, Portugal, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | February 07, 2012 | 456 | 7314 | Not_Computer
Z-0322-2011 | 56255 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartStart MRx Defibrillator/MonitorModels: M3535A / M3536A / M3536J / M3536MCThe HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips Healthcare HeartStart MRx | 2 | November 08, 2010 | 2010 | Philips Healthcare Inc. | Updated Instructions for Use- HeartStart MRx monitor/defibrillator. The update provides more detailed information on:" Default ECG Lead Behavior" Synchronized Cardioversion using External Paddles | DESIGN: Labeling Design | Philips issued an Urgent Voluntary Medical Device Correction notification/Field Safety Notice and an Instructions for Use Addendum on July 20, 2010, to affected customers. The Notice identified the product, the problem, and the action customers should take. In the USA, the notification sent to all affected customers via UPS and confirmation of delivery would be tracked. Outside the USA, Philips representatives in the affected geographies would manage their mailings locally. Customers were instructed to review the HeartStart MRx Instructions for Use addendum and keep a copy with their equipment - Instructions for Use. Additionally, they were instructed to insure that all users are fully trained on the MRx and, in particular, ensure that they review and understand have reviewed and understand the additional detail provided in the notification.For questions regarding this recall call (800) 722-9377. |  | 41119 | 41,119 units | Worldwide Distribution - US Nationwide and the countries of: CANADAAFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBELGIUMBOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIACHILECHINACOLOMBIACYPRUSCZECH REPUBLICDENMARKECUADOREGYPTESTONIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGHUNGARYICELANDINDIAINDONESIAIRELANDISRAELITALYJAPANJORDANKENYAKUWAITLATVIALEBANONLIBYAN ARAB JAMAHIRILITHUANIALUXEMBOURGMACEDONIAMALAYSIAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIASAUDI ARABIASERBIA&MONTE; NEGROSINGAPORESLOVENIASOUTH AFRICASOUTH KOREASPAINSRI LANKASWEDENSWITZERLANDTAIWANTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMVENEZUELAVIET NAMYEMEN | Terminated | September 05, 2012 | 667 | 41119 | Not_Computer
Z-0332-2011 | 55301 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Excel Care Bariatric Bed, P 610A, Hill-Rom, Batesville, IN. Intended Use: The Excel Care Bariatric Bed is for bariatric patients with varying medical and physical conditions. The bed supports patients between 250 and 1000 pounds. | Excel Care Bariatric Bed | 2 | November 09, 2010 | 2010 | Hill-Rom, Inc. | The firm received complaints regarding the siderails not operating correctly or becoming inoperable. The siderails could experience false latching, unintentional lowering or become inoperable. There have been reports of 5 patient falls and 2 injuries to caregivers. | DESIGN: Device Design | Hill-Rom sent out "Urgent Medical Device Correction" letters dated March 5, 2010. The letters described the product, problem and actions to be taken by the customers. The letters instructed the customers to locate the beds and examine the siderails for any visible damage (cracks, bending, loose parts) and conduct a physical assessment of all four siderails on each bed to determine if the sideral could be placed in all positions. If the siderails were found to be inoperable, the bed was to be taken out of service until a qualified Hill-Rom technician can make any necessary correction. The customers were instructed to contact Hill-Rom to make arrangements and inform them of all beds that they have loaned, sold, no longer use or disposed of at 800-445-3720, option 2. The letter states the firm is evaluating a design change to the siderail.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720, option 2. |  | 1406 | 1,406 beds both models. | Worldwide distribution: USA states including: AK, FL, HI, IN, KS, LA, MS, OK, PA, TN, VT, and WA and countries including: Canada, Mexico, Angola, Australia, Saudi Arabia, Qatar, and United Emirates. | Open |  |  | 1406 | Not_Computer
Z-0329-2011 | 55840 | lamp, infrared, therapeutic heating | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(k) | Titan XL Handpiece Infrared Therapeutic Heating Lamp, Manufactured by Cutera Inc, Brisbane, CAInfrared therapeutic heating lamp, intended to emit energy in the infrared spectrum to provide topical heating for the purpose of elevating tissue temperature for the temporary relief of minor muscle and joint pain and stiffness. | Titan XL | 2 | November 09, 2010 | 2010 | Cutera Inc | Risk of burns-- under certain conditions, devices containing a gold plated end block could produce energy above what is stated on the device setting. | DESIGN: Device Design | Cutera sent Certified letters sent to all consignees on May 8, 2010. The letter identified the product, the problem, and the action the consignees should take. Consigness were instructed to immediately stop using the handpiece and complete the attached Recall Tracker Form. A second letter was sent to customers who received a replacement handpiecce but did not return the affected devices. All affected products were to be exchanged for new or reworked devices.For questions contact customer service at (888) 428-8372, ext 5561. |  | 1344 | 1,344 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OK, OR, PA, OH, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Argentina, Australia, Austria, Barbados, Belgium, Bermuda, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cypress, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Lle de la Reunion, Mexico, New Zealand, Phillipines, Poland, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, UAE, and Ukraine. | Terminated | December 28, 2010 | 49 | 1344 | Not_Computer
Z-0330-2011 | 56800 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | External Compression Device; 1 unit per packageStryker Trauma GmbH;Distributed by Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430Active Apposition/compression is achieved by threading the External Compression Device into the Nail holding Screw. While being threaded into the Nail Holding Screw the tip of the External Compression Device will contact the (shaft screw(cross screw) and will push the screw in distal direction. As the cross screw is screwed into the bone of the patient, the bone fragment will be moved too is distal direction. Finally both sides of the bone fracture will be pressed together - compression is now applied to the fracture side. | External Compression Device | 2 | November 09, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Variations in the outer diameter of the shaft of the External Compression Device and the inner diameter of the Nail Holding Screw may make assembly of these 2 instruments not possible. | OTHER/UNDETERMINED: Pending | Stryker sent Recall notification letters by Federal Express on September 9, 2010 to all Stryker branches/agencies and Risk Managers, Chief of Orthopaedics of affected hospitals and Surgeons on September 10, 2010, that used the product. The letter identified the product, the problem, and the action that customers should take. Customers were instructed to:Examine their inventory and hospital locations to identify the affected product and return it to their branch or agency warehouse for reconciliation.Reconcile all product on the attached Product Recall Acknowledgment Form and fax a copy to (201) 831-6069 within 5 days of receipt of notice.A Stryker representative would contact the customer with details for return of the product.For any questions regarding this recall call (201) 972-2100. |  | 11 | 11 | Nationwide Distribution including AZ, CA, CO, FL, GA, IL, KY, MD, OH, OK, and TX | Terminated | August 21, 2012 | 651 | 33 | Not_Computer
Z-0334-2011 | 57024 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic Intrathecal Catheter, 8731SC with sutureless connector. Method of Sterilization: Ethylene Oxide. Manufacturer Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 The implantable Medtronic Model 8731SC intrathecal catheter is part of an infusion system that stores and delivers parenteral drugs to the intrathecal space. The implanted infusion system components consist of a Medtronic pump and a Model 8731SC catheter. The catheter connects to the pump at the catheter port.To deliver parenteral drugs to the intrathecal space. The needle is used to facilitate the placement of the catheter. The needle is inserted percutaneously into the intrathecal. The needle stylet is then removed to confirm the needle placement through observing backflow of cerebrospinal fluid (CSF) through the needle. The needle stylet prevents tissue coring and also minimizes loss of CSF. The spinal segment tip of the catheter is inserted through the needle until the catheter is advanced to the desired location. Once the catheter position is verified through fluoroscopy, the needle is carefully removed and discarded. | Medtronic Intrathecal Catheter, 8731SC | 2 | November 09, 2010 | 2010 | Medtronic, Inc. - Neuromodulation | Medtronic is recalling specific lots of Model 8731SC intrathecal catheter kits and Model 8598A intrathecal catheter spinal revision kits because there is the potential that the 11.4 cm introducer needles within these kits may exceed USP requirements for bacterial endotoxin. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Medtronica sent an "Urgent: Medical Device Recall" letter, dated September 2010, was sent to medical facilities/customers beginning September 29, 2010. The letter described the issue, identified affected product, requested that unused inventory be segregated and returned to Medtronic Neuromodulation per the enclosed Product Return Instructions form. A Recall Reply Card was also asked to be returned. Additionally an "Urgent: Medical Device Recall" letter was addressed to Healthcare Providers/customers to inform them of the recall of the Catheter and Revision Kits, identified patients by means of registration information, Informed them of the issue with the needle inside the kits, and requested they stop using the specific lots and to monitor any patients for symptoms. If you have any questions or comments, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. | Contact the recalling firm for information | 1323 | 1323 | AL, CA, CO, CT, FL, GA, HI, ID. IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, WY. OUS to include: Argentina, Austria, Belgium, Canada, Chile, Czech Republic, France, Germany, Greece, Hungary, Ireland, Israel, Italy, San Marino, Jordan, Lithuania, Netherlands, Norway, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, United Arab Emirates, United Kingdom. | Terminated | December 10, 2011 | 396 | 1743 | Not_Computer
Z-0339-2011 | 56973 | stent, carotid | NIM |  |  | Cardiovascular |  | RX ACCULINK Carotid Stent System 7-10 X 40mm, Part Number: 1011344-40, Lot Number: 0072061The RX Acculink Carotid Stent System, used in conjunction with Abbott Vascular RX Accunet embolic protection system, is indicated for the treatment of patients at high risk for adverse events from carotid endarterectomy who require carotid revascularization and meet the criteria outlined below. (1) Patients with neurological symptoms and | RX ACCULINK Carotid Stent System | 2 | November 10, 2010 | 2010 | Abbott Vascular | The recall was initiated because Abbott vascular has discovered through internal testing that the affected lots may not meet the firm's quality specifications for catheter shaft tensile strength. If the affected device is used it may result in acute stent deployment difficulties and subsequent intervention. | PRODUCTION CONTROLS: Process Control | Abbott Vascular sent an URGENT DEVICE RECALL letter dated September 10, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken. The Sales Representatives were to personally visit each account to deliver the recall letter and assist in identifying and returning anyunused devices to Abbott Vascular. Also, Abbott Vascular Regulatory Compliance initiated contact with the Risk Management department at each account by phone to verify their mailing address and provide a courtesy copy of the recall letter.The consignees were instructed to work with their local account representative to review their inventory, complete the Recall Effectiveness Check and return identified products to Abbott Vascular.For questions regarding this recall call (951) 914-3324. |  | 42 | 42 units total for all products | Nationwide Distribution including AL, CA, CT, DC, FL, GA, IA, IL, IN, MA, ME, MI, MN, MO, OH, OK PA, TX, VA, and WI | Terminated | November 16, 2010 | 6 | 42 | Not_Computer
Z-0343-2011 | 56976 | fluorometric method, cpk or isoenzymes | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Rotors, catalog number CROTOR, lot number NH42-251-09Intended use: Stratus(R) rotors is intended for the separation of whole blood samples for analysis on the Stratus(R) CS analyzer | Stratus(R) CS Rotors | 2 | November 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low frequency defect in the gel filling of this lot that may result in inadequate separation of cells in centrifugation. | PRODUCTION CONTROLS: Process Control | The recalling firm, SIEMENS, issued an "Urgent Field Safety Notice" dated December 2009 to all affected customers. The notice described the product, problem and the action to be taken by the customers. The customers were instructed to review the situation with their laboratory director to determine the need to evaluate previously reported test results; immediately examine their CROTOR inventory and discard all bags that are from lot NH42-251-09 including rotors that were not stored in their original packaging (bag); contact Siemens Customer Service Center at 800-241-0420 to request additional no-charge replacements (Note: A "no-charge" order has already been placed so that one bag of replacement rotors would be recieved shortly); forward this notice to anyone they may have distributed this product, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK via fax at 302-631-8467.If you have any questions regarding this information, please contact the Siemens Technical Solution Center at 800-405-6473. |  | 330 | 330 bags | Worldwide distribution: USA including states of AK, AZ, AR, CA, FL, GA, HI, IA, IL, IN, KY, LA, MI, MN, MO, MS, MT, NC, ND, NH, NM, NY, OH, TX, UT, WA, WI, and WV and country of Canada | Terminated | November 10, 2010 | 0 | 330 | Not_Computer
Z-0341-2011 | 56751 | apparatus, autotransfusion | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | The box label states: "ATF 40 Fast Start Kit 9108491...STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician".The ATF 40 Fast Start Kit contains one each AT1 Autotransfusion Set, ATS Suction Line, and the ATR40 Autotransfusion Reservoir.ATS is a suction line to aspirate and add anticoagulate to shed blood from the surgical field into an autotransfusion reservoir (ATR). The ATS has a ÂŽ inch. suction port and is connected to a vacuum source.ATR is an autotransfusion reservoir to collect, de-foam, filter, and store blood prior to processing. The ATR reservoirs are offered with either a 120 micron or a 40 micron filter.ATF 40 Autotransfusion set is a disposable set, dedicated component of the Fresenius CATS Continuous Autotransfusion System, which in turn is an Autotransfusion device for the processing of autologous shed blood collected intraoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. | ATF 40 (item 9108491) | 2 | November 10, 2010 | 2010 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode. | OTHER/UNDETERMINED: Pending | Fresenius Kabi sent out a Technical Notification letter dated March 29, 2010, to one distributor located in Ann Arbor, Michigan. The distributor sent the technical notification with a cover letter to each of their customer's that had ordered AT1's, ATF120's and ATF40's over the last year. The letter identified the product, the problem, and the action to be taken. Customers were notified that the autotransfusion prime and / or washing procedure can be continued after alarm blood flow by following Alternative procedures. The Technical Notification informs the distributor about the problem of the AT1 (item #9005101), but also mentions the other two products ATF120 (item #9108501) and ATF40 (item #9108491) which have no known problem.The distributor is advised to contact Fresenius Kabi, LLC for any question regarding the Technical Notification.For questions regarding this recall call (425) 242-2074. |  | 444 | 17, 444 units distributed in US | Worldwide Distribution - USA and the countries of Australia,Austria, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Malaysia, Netherlands, Phillipines, Portugal, Switzerland, Thailand, UK, and Vietnam. | Terminated | June 15, 2011 | 217 | 1384 | Not_Computer
Z-0344-2011 | 56863 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab/Mic ASCII Add-on Versions 4.0.1.14, 4.0.1.15, and 4.0.1.16.SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760.Version 4.0.1.14 Manufactured & Distributed 12/29/2006. Version 4.0.1.15 Manufactured & Distributed 05/02/2007. Version 4.0.1.16 Manufactured & Distributed 08/22/2008. Used in a medical research or clinical laboratory setting by knowledgeable, trained and experienced personnel. | SoftLab/Mic ASCII Addon | 2 | November 10, 2010 | 2010 | SCC Soft Computer | When a test is being verified, the wrong (expired) ranges may be associated to that test result, causing an incorrect flag to be posted for that test result. | DESIGN: Software Design | Clients were notified via the proprietary communication software on 02/16/2010. The Notification Verbiage identified the affected product and informed clients of the potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 31 | 31 | Nationwide Distribution - US and Puerto Rico. | Terminated | July 17, 2012 | 615 | 31 | Software
Z-0236-2011 | 56760 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N134, REF 22 319, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, 69280 Marcy-l'Etoile France. | VITEK 2 Gram Negative Susceptibility Card, ASTN134 | 1 | November 11, 2010 | 2010 | Biomerieux Inc | AMENDED REASON FOR RECALL: The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | bioMerieux issued a "Product Correction Notice" letter dated September 13, 2010, to their entire customer base (regardless of whether they ever ordered the recalled cards). The letters explained the reason for the recall and instructed the account to confirm all results for E. coli using a second method, and for K. pneumoniae, to confirm with a second method only for resistant results. These alternate test confirmations are to continue until further notice. The customers were to complete and return the enclosed Acknowledgement Form via fax at (314) 731-8689 or (919) 620-6350 which indicate that the notice has been read, the instructions provided in the "Actions Required" section were followed, and notes if they have received any reports of illness or injury related to this issue. All future shipments will include an orange multilanguage insert to go into the boxes reminding customers they must perform an alternate test for TZP/E. coli combinations and that customers must also perform an alternate method prior to reporting TZP results when a resistant result is obtained for the TZP/K. pneumoniae. Additionally, this recall letter will accompany shipments of instrument software kits for the new customers.The recalling firm issued a press release on 10/27/10 via the Associated Press notifying the public of the recall.The recalling firm expanded their recall to add 7 more organisms to the list of organism testing that were to be suppressed. As a result the recalling firm issued a letter dated 3/30/11 via Fed Ex on 4/4/11 flagged "URGENT PRODUCT CORRECTION NOTICE (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786." The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Taz. | Contact the recalling firm for information | 10091 | 10,069 boxes as of 3/8/11 | Worldwide Distribution: USA (all states), and the countries of:SUBSIDIARIES - Argentina, Canada, Switzerland, Chile, Colombia, Germany, Spain, Great Britain, Hong Kong, Japan, Korea, Mexico, Poland, Brazil, Sweden, Singapore, and Thailand. DISTRIBUTORS - Dutch Antilles, Australia, Bermuda, Uruguay, Costa Rica, Dominican Republic, Ecuador, France, Guatemala, Honduras, Indonesia, Israel, Italy, Singapore, Malaysia, Nicaragua, The Netherlands, Panama, Peru, Philipines, Paraguay, El Salvador, Venice, and Viet Nam. | Terminated | May 08, 2012 | 544 | 1859979 | Not_Computer
Z-0349-2011 | 51883 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ViewSonic 19" flat panel LCD display Model VG930M which is used as part of the Welch Allyn Acuity Central Monitoring System. The affected serial numbers are those less than QC2082880001.The ViewSonic 19" flat panel LCD display Model VG930M is used as part of the Welch Allyn Acuity Central Monitoring System (Acuity System.) The Acuity System consists of a central monitoring station, Acuity System software, and a collection of other commercially available networked products | 19 ViewSonic LCD VG930M Display | 2 | November 12, 2010 | 2010 | Welch Allyn Protocol, Inc | The 19" Model VG930M ViewSonic flat panel LCD display may have faulty capacitors at certain locations in the power supply section. The faulty capacitors may cause a loss of power to the display resulting in the display screen going dark and loss of both visual and secondary patient alarms. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On May 1, 2009, Welch Allyn Protocol, Inc. sent out recall notification to their consignees in via UPS and US Postal Certified Mail.The letter describes the recalled product as 19" model VGM930M ViewSonic flat panel LCD display with serial numbers less than QC2082880001. For the Welch Allyn service contract consignees/customers, the firm's Field Service Engineer will replace the LCD monitor for the consignees. The consignees could contact Technical Support at 1-800-289-2501.For the Warranty/non-Warranty contract consignees/customers-For the Welch Allyn service contract consignees, Stericycle Return Management Services (RMS) will manage all recall and recovery initiatives. Consignees/customers can call Welch Allyn Customer Care at 1-800-535-6663 for any questions or concern about this recall. |  | 637 | 637 units | Worldwide distribution: USA including states of: AK, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VT, WI, and WV; and countries including: AUSTRALIA, CANADA, ESTONIA, GERMANY, ICELAND, IRELAND, ITALY, JORDAN, MALAYSIA, POLAND, and RUSSIA. | Terminated | July 25, 2011 | 255 | 637 | Hardware | Display/Image | Replace LCD | Remove or Replace
Z-0358-2011 | 53048 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Renegade Fiber Braided Microcatheter, 150/10/2RO, Rx Only; Model Numbers: 18257, 18259; Product is distributed by Boston Scientific Corporation, Natick, MA;Product is manufactured by Boston Scientific Corporation, Cork, Ireland | Renegade Fiber Braided Microcatheter | 2 | November 15, 2010 | 2010 | Boston Scientific Corporation | Sterility Loss-- The sterile barrier pouch seal (top seal) may be breached by the product, resulting in loss of sterility which may result in transmission of disease or infection. | PRODUCTION CONTROLS: Packaging Process Control | Boston Scientific issued an Urgent Medical Device Recall letter dated August 26, 2009 identifying the affected devices, the issue prompting the recall, and actions to be taken by customers.Customers were instructed to determine if they have affected product in their inventory, and if so, immediately quarantine and return to Boston Scientific.Customers can contact Boston Scientific at 510 624-1734. | Contact the recalling firm for information |  |  | Distribution Worldwide, including throughout the US. | Terminated | February 03, 2011 | 80 |  | Not_Computer
Z-0367-2011 | 55839 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit Recall, Catalog Number 150-FT-Q; Kit includes angiographic injector and syringeThe contents of the package are intended to be used in the delivery of contrast media. | MEDRAD Sterile Disposable Syringe Kit Recall | 2 | November 15, 2010 | 2010 | Medrad Inc | Some kits have a hole in the packaging which could compromise the sterility of the components. | PRODUCTION CONTROLS: Packaging | MEDRAD issued an URGENT MEDICAL DEVICE RECALL letter dated May 17, 2010, to all affected customers. The letter identified the product, the problem, and the action the customer should take.Customers were instructed to identify the product with the affected lot numbers.Contact MEDRAD Customer Support at 1-800-633-7231 to make arrangements for the return of the product.Complete and sign the form attached to the letter and fax it back to MEDRAD at 412-406-0942.For any questions regarding this recall call (412) 767-2400, ext 6851. |  | 39448 | 39,448 syringe kits | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of Belgium, China, Columbia, Germany, Denmark, France, Great Britain, Hungary, Ireland, India, Israel, Italy, Mexico, Netherlands, Sweden, and Singapore. | Terminated | November 17, 2010 | 2 | 39448 | Not_Computer
Z-0351-2011 | 56792 | device, vein stabilization | LBJ | Vein stabilizer. | General Hospital | General Hospital | 510(K) Exempt | Baxter V-Link Luer Activated Device with Vitalshield Protective Coating Non-DEHP Catheter Extension Set Kit; 7.6" (19.3 cm), Vol. 1.0 mL, 1 Catheter Extension Set with Luer Activated Device with Silver for IV Access and Male Luer Lock Adapter with Retractable Collar, 1 VITAL-HOLD medium catheter stabilization device, 2 foam tape strips; an Rx sterile convenience kit for IV infusions; 50 units per case; Manufactured by an affiliate of Baxter Healthcare Corp., Deerfield, IL 60015 U.S.A., Made in Costa Rica; catalog 6N8220 | Baxter VLINK Luer Activated Device with Vitalshield NonDEHP Catheter Extension Set Kit | 3 | November 15, 2010 | 2010 | Baxter Healthcare Corp. | The removable paper liner covering the adhesive strip that is used to secure the catheter stabilization device to the patient is pyrogenic. The device is labeled as nonpyrogenic. | DESIGN: Component Design/Selection | Baxter issued Urgent Product Recall letters dated September 17, 2010 to the affected direct accounts. The accounts were informed that the paper lining covering the adhesive strip that is used to secure the catheter stabilization device to the patient was found to be pyrogenic. The accounts were requested to discontinue use and segregate the affected lots of products from inventory and call Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, 7:00 AM to 6:00 PM CT, to arrange for the return of the affected product for credit. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and the amount of each product being returned. If the accounts further distributed the products, they were requested to forward the communication to those accounts. Clinical questions regarding the communication are directed to the Medical Information Services at Baxter at 1-800-933-0303. Any other questions may be directed to the Center for One Baxter at 1-800-422-9837. |  | 73350 | 73,350 units | Nationwide Distribution: Throughout the US. | Terminated | October 04, 2011 | 323 | 172300 | Not_Computer
Z-0363-2011 | 56461 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Transport ProÂŽ Monitor with the CARESCAPE" Patient Data Module. | Solar 8000i with Patient Data Module / Transport Pro Monitor | 2 | November 15, 2010 | 2010 | GE Healthcare, LLC | Transport Pro Monitor stops communication with the CARESCAPE Patient Data Module (PDM) after 414 days of continuous run time. This time will equate to different amounts of real time depending on how much the units is actually in service per day. Transport Pro contains an internal timer that is used to control the software and remind users to perform preventive maintenance. When this internal timer | DESIGN: Software Design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated September 17, 2010 to consignees. The letter was addressed to Hospital Administrator, Head of Biomedical Engineering and Nursing Manager. The letter described the product, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Service representatives will updates all of the affected Transport Pros with PDM that were distributed.Customers may contact GE at 262-548-2731 about this correction. | Contact the recalling firm for information | 3256 | 3256 | Worldwide Distribution: Domestic in the states of AL, AZ, AR, CA,CT, CO, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, ME, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, VT, WA, WI, and WY, and Internation in the countries of BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, CANADA, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY GREECE, INDONESIA, IRELAND, ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MAURITIUS, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD 7 TOBAGO, and the UNITED KINGDOM. | Terminated | May 17, 2012 | 549 | 3256 | Software | Display/Image | Software update | Software Update
Z-0364-2011 | 56865 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SA BASE with SoftLab Version 3.1.4.x. Manufactured 03/2003. SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760.Intends to provide clients with an interface that links all clinical laboratory departments throughout the enterprise. | SA BASE with SoftLab Version 3.1.4.x. | 2 | November 15, 2010 | 2010 | SCC Soft Computer | Functionality issue: For clients using the instrument to generate results on the interface menu, two different result records could be displayed under the same order and sequence numbers. This could cause incorrect results to be populated to a patient record. | DESIGN: Software Design | Clients were notified via the proprietary communication software on 06/01/2010. The communication discussed the risk-to-health notification task, potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 40 | 40 | Worldwide Distribution -- US, including states of CT, FL, IL, KY, GA, LA, IN, MS, NJ, NY, PA, OH, and TN and country of Canada. | Terminated | July 16, 2012 | 609 | 40 | Software | Display/Image | Software update | Software Update
Z-0369-2011 | 56814 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE BrightSpeed Excel/Edge/Elite Select / Elite CT Scanner System, models 2377708-560, 5377708-540, 5377708-560. Composed of a gantry, patient table, operator console, computer, and PDU and includes image acquisition hardware, image acquisition and reconstruction software, associated accessories and connections/interfaces to accessories. Is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes, including Axial, Cine, Helical (Volumetric), Cardiac, and Gated acquisitions. | GE BrightSpeed CT | 2 | November 15, 2010 | 2010 | GE Healthcare, LLC | GE has become aware of a potential set of circumstances that could cause x-ray continuation during an unexpected table stop on certain LightSpeed VCT and BrightSpeed systems. | DESIGN: Software Design | An Urgent Medical Device Correction letter dated September 2, 2010 was sent to consignees beginning September 3, 2010. This letter addressed the LightSpeed VCT systems. Another Urgent Medical Device Correction letter, dated September 24, 2010 was sent September 24, 2010 for the BrightSpeed Systems. The letters described the issue and provided safety instructions for the health care provider to be aware if the table should stop unexpectedly and the x-ray remains on to manually stop the scan until the devices software is upgraded.If you have any questions, please contact the US (800) 437-1171, Japan 0120-055-919, Korea 1544-6119, Australia/New Zealand 800 659 465, China 800-810-8188; for other countries, please contact your local GE Healthcare Services Representative. | Contact the recalling firm for information | 946 | 946 | Worldwide distribution: USA including states of: MI, NM, OH, TX, and Puerto Rico and countries including: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hungary, India, Indonesia, Ireland, Islamic Republic of Iran, Israel, Italy, Jamaica, Japan, Kenya, Lebanon, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Pakistan, Paraguay, Philippines, Poland, Portugal, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | May 15, 2012 | 547 | 1014 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0362-2011 | 56872 | software, blood bank, stand alone products | MMH | Automated blood grouping and antibody test system. | Hematology | Hematology | 510(k) | SoftBank Versions 23.1.2.x, 23.2.0.x, and 25.0.0.x with SoftReports.SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760.23.1:BK040048 23.2:BK080020 25.0:BK090017The SoftBank II application, using SoftScape user interface, is a decision support software device that requires knowledgeable user intervention to document steps and events in a transfusion service. The application provides single and multiple site facilities the ability to manage their transfusion service by integrating patient and unit information. SoftBank uses an SCC proprietary interface called SoftScape that facilitates access to the application. | SoftBank Versions 23.1.2.x, 23.2.0.x, and 25.0.0.x with SoftReports. | 2 | November 15, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a Correction for SoftBank II Versions: 23.1.2.x, 23.2.0.x, and 25.0.0.x software applications. Functionality issue - For clients generating patient history reports using Patient | DESIGN: Software Design | SCC Soft Computer notified Customers on June 22, 2010, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated June 11, 2010. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; to grant permission to load the required hot fix, and to make the manual correction. For the manual correction to the datapath for patient report layouts, your SoftBank Support Specialist will Coordinate those changes and your validation of affected environments.If you have any questions, contact (727) 789-0100. |  | 30 | 30 | Worldwide distribution: U.S. and country of Canada. | Terminated | September 05, 2012 | 660 | 30 | Software | Output/Calculation | Software update | Software Update
Z-0360-2011 | 56873 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftID.NET versions 2.0.0.6 - 2.0.0.19. SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760. SoftID.NET is to be used in a medical or clinical laboratory setting by knowledgeable, trained, and experienced professionals. SoftID.NET supports bedside collection of patient specimens used in diagnostic testing. This product is integrated with Soft Computer's SoftLab, SoftScape and can also interface with all other SCC laboratory systems and Hospital Information Systems (HIS). | SoftID.NET versions 2.0.0.6 2.0.0.19. | 2 | November 15, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer issued a Correction for SoftID.NET versions 2.0.0.6 - 2.0.0.19. When using the Cancel functionality, if the collection list refreshes during the transaction and changes the position of the patient on the list, the wrong patient order can be cancelled. All affected clients have been notified. | DESIGN: Software Design | SCC Soft Computer notified Customers on August 28, 2009, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated August 17, 2009. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; document a request to load the software, if they wish to receive the correction; document that the Alternative Working Solution will be used by the customers facility until the correction has been taken live.If you have any questions, contact (727) 789-0100. | Contact the recalling firm for information | 41 | 41 | Nationwide distribution: AR, CA, FL, IL, IN, KS,LA, MA, MI, MN, MO, NC, ND,NH, NJ, NY, OH, PA, RI, TN, TX, VA and WI. | Terminated | November 16, 2011 | 366 | 41 | Software | Output/Calculation | Upgrade | Repair
Z-0361-2011 | 56940 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ENDOPATHÂŽ ETS-FLEX45 Articulating Endoscopic Linear Cutters, Model # ATS45. Manfuactured by Ethicon Endo-Survery, LLC; 475 Calle C; Guaynabo PR 00969. | ENDOPATH ETSFLEX45 Articulating Endoscopic Linear Cutters | 2 | November 15, 2010 | 2010 | Ethicon Endo-Surgery Inc | The mechanism that connects the articulation joint to the device shaft may cause the jaws of the cutters to remain closed and clamped down on tissue after the device is fired. If the jaws of the device remain clamped down on the tissue, there is a risk the surgery will need to be modified to remove the device. This may lead to an increase in procedure time, tissue manipulation, and a change in t | OTHER/UNDETERMINED: Pending | Ethicon Endo-Surgery issued "Urgent Device Recall Event 1952" notification letters dated January 21, 2009 to their customers. They instructed their custormers to immediately discontinue use of the product. They are to complete and fax back the Business Reply Form. Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle. Replacement product should be received within 30-60 days after return.Customers can contact Stericycle concerning this recall at 1-866-861-6921, reference Event 1952. |  | 4113 | 4,113 UNITS | Worldwide Distribution: USA only, in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI. IA. ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | May 10, 2012 | 542 | 4113 | Not_Computer
Z-0359-2011 | 56664 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Power supply boards on Dash 3000, 4000 and 5000 Monitors.GE Healthcare, 8200 West Tower Avenue, Milwaukee, WI 53223. | GE Healthcare Dash 3000, 4000 and 5000 patient monitors | 2 | November 15, 2010 | 2010 | GE Healthcare, LLC | The power supply board in GE Healthcare Dash 3000, 4000 and 5000 patient monitors may fail due to insufficient solder of a component. Insufficient solder on the power supply board could result in excess EMC emissions, smoking of components and eventual failure of the power supply, tripping of the fuse and immediate shutdown resulting in a loss of monitoring that could delay recognition of a clinic | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 8, 2010, which described the loss of monitoring issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to remove the units from use. Local service representatives will contact firms to schedule installation of replacement power supply boards.GE Healthcare can be contacted at 1-800-558-7044, Option 2, Technical Support, about this correction. | Contact the recalling firm for information | 137 | 137 | Worldwide Distribution: USA in the states of IL, NV, TX, and VA, and the countries of Australia, Brazil, China, Japan, South Korea, and Taiwan. | Terminated | May 29, 2012 | 561 | 137 | Battery | Device Operation | Replace boards | Remove or Replace
Z-0386-2011 | 57052 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sorin OptiPack, Perfusion Tubing Set, Sterile EO, REF 084503000, Manufactured by Sorin Group USA, Inc. 14401 W. 65th Way Arvada, CO 80004Indicated for use in surgical procedures requiring extracorporeal support for periods of up to six hours. | Perfusion Tubing Set | 2 | November 15, 2010 | 2010 | Sorin Group USA, Inc. | Defective product packaging may compromise product sterility. |  | Sorin Group sent an "URGENT: FIELD SAFETY NOTICE" letter dated October 7, 2010, to all customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to check all affected inventory to exam for holes. Any damaged products were to be returned to Sorin. If no damage was found, product could be used. The customers were also instructed to contact Sorin Cardiopulmonary Customer Service at 1-800-650-2623 to return product and receive replacement; they were to share this information with all personnel who may use this product; and to complete and return the attached Customer Response Form via the self-addressed prepaid envelope provided or fax to (303) 467-6502 as soon as possible.For further information, contact 1-800-650-2623. |  | 13 | 13 | Nationwide distribution: CA, FL, IL, IN, KS, MD, MO, NE, NJ, NM, TN, TX, and VA. | Terminated | December 17, 2010 | 32 | 3021 | Not_Computer
Z-0366-2011 | 57001 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" ID/AST Panels, Item # PMIC/ID-107, Catalog number 448607, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663*** | BD Phoenix" Automated Microbiology System ID/AST Panels | 2 | November 15, 2010 | 2010 | Becton Dickinson & Co. | The diagnostic reagent packaging may not be completely sealed, which may cause substrate degradation, and produce incorrect test results, including false positive and false negative results. | PRODUCTION CONTROLS: Packaging Process Control | BD Diagnostic Systems issued an "Urgent Product Recall" letter dated May 2010 to direct accounts, identifying the affected product, potential impact, and actions to be taken by the customer. Accounts were requested to respond by returning the enclosed reply form indicating the number of replacements needed for the affected product to be discarded.For assistance iwth replacements, contact BD Customer Service at 800 675-0908. For other inquiries, contact BD Technical Services at 1 800-638-8663. | Contact the recalling firm for information | 1675 | 1675 units | Nationwide Distribution: To direct accounts in CA, MD, NJ, PA. and WI. | Terminated | November 29, 2010 | 14 | 1675 | Not_Computer
Z-0394-2011 | 55999 | table, radiographic, stationary top | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Rectangular Extension, REF MTIL6600, Reusable non-sterile extension for use with Universal Couchtop. The firm name on the label is CIVCO, Orange City, IA.For use to support and aid in positioning a patient during radiologic and other medical procedures. | CIVCO Rectangular Extension | 2 | November 16, 2010 | 2010 | Med Tec Inc | Under certain conditions, the extensions can become detached from the main body of the couch. | DESIGN: Device Design | The recalling firm, CIVCO, issued an 'Urgent Field Safety Notice" dated June 15, 2010, via overnight mail on June 21, 2010, regarding the first issue for the three extension models. The letter described the product, problem, and actions to be taken by the customers. The customers were instructed to load the patients on the couch from the bottom of the couch and not from the superior end with the extension and do not let patients crawl to the top of the couch with their full upper body weight at the end of the couch; notify others within their organization that may have the extension, and complete and return the enclosed response form via fax: Attention: CIVCO MTL6600 Rectangular Extension at 877-271-2310. The letter informs the customer that they will be receiving a kit to replace the locking tab so that the extension will be less likely to come loose. Once the kit is received, the customers are ask to return the replaced part and sign a confirmation form for their records.CIVCO sent a second "Urgent Field Safety Notice" dated July 20, 2010 via overnight mail on July 21, 2010, and by phone on July 22, 2010, regarding the second issue which covered only MTIL6600 regarding the weight being dropped onto the corner of the rectangular extension. The letter described the product, problem and action to be taken by the customers. The customers were instructed to accept a one-piece Universal Couch Top replacement; sign the attached form stating the they will stop using the rectangular extension and return it to CIVCO or sign the form stating they have been advised of the recall and they will exercise the necessary precautions to prevent the event from occurring by always using the thumb screw lock and never allowing the patient to crawl on the couch top or extension, and to notify others within their organization that have the extension. The customer was informed that their current system would be replaced and sent back to the recalling firm. Please contac. | Contact the recalling firm for information | 22 | 22 extensions | Worldwide distribution: USA including states of: CA, FL, IA OH, and WI; and countires of: United Kingdom, Australia, and Switzerland. | Terminated | April 17, 2012 | 518 | 30 | Not_Computer
Z-0365-2011 | 56867 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI Version 4.0.1.0 - 4.0.1.16, 4.0.2.0 - 4.0.2.10, 4.0.3.0 - 4.0.3.13, and 4.0.4.0 - 4.0.4.5.SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760. Manufacture/Distribution Dates:4.0.1.0 - 08/06/04, 4.0.1.4 - 10/18/04, 4.0.1.5 - 01/18/04, 4.0.1.6 - 12/15/04, 4.0.1.7 - 01/10/05, 4.0.1.8 - 03/29/05, 4.0.1.9 - 05/06/05, 4.0.1.13 - 01/25/06, 4.0.1.14 - 03/26/07 4.0.1.15 - 05/02/07, 4.0.1.16 - 08/22/08, 4.0.2.0 - 04/08/05, 4.0.2.1 - 08/13/08, 4.0.2.2 - 06/06/05, 4.0.2.3 - 09/19/05, 4.0.2.4 - 09/19/05, 4.0.2.5 - 09/19/05, 4.0.2.7 - 02/13/06, 4.0.2.8 - 12/12/06, 4.0.2.9 - 02/20/08, 04.2.10 - 08/13/08, 4.0.3.2 - 05/17/06, 4.0.3.3 - 09/18/06, 4.0.3.4 - 09/27/06, 4.0.3.5 - 11/29/06, 4.0.3.6 - 06/22/07, 4.0.3.7 - 03/28/07, 4.0.3.8 - 06/22/07, 4.0.3.9 - 08/03/07,4.0.3.10 - 08/24/07, 4.0.3.11 - 09/24/07, 4.0.3.12 - 03/07/08, 4.0.3.13 - 06/13/08, 4.0.4.2 - 02/22/08, 4.0.4.3 - 05/22/08, 4.0.4.4 - 07/28/08, 4.0.4.5 - 05/04/09.Intends to provide clients with an interface that links all clinical laboratory departments throughout the enterprise. | SoftLab GUI Version 4.0.1.0 4.0.1.16, 4.0.2.0 4.0.2.10, 4.0.3.0 4.0.3.13, and 4.0.4.0 4.0.4. | 2 | November 16, 2010 | 2010 | SCC Soft Computer | Clients using Patient Maintenance to perform Moves of Stays may send the wrong billing number during an ADT update to SoftWeb, SoftPath or SoftMedia. The wrong billing number can come from a different patient in Order Entry. Only Moves of Stays are affected. | DESIGN: Software Design | Clients were notified via the proprietary communication software on 06/01/2009. This is a MANDATORY CORRECTION. The Risk-To-Health Notification Task gives the background on the issue, the potential hazard, root cause, corrective action, alternative working solution, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. | Contact the recalling firm for information | 49 | 49 | Nationwide Distribution | Terminated | August 07, 2013 | 995 | 49 | Software | Output/Calculation | Software update | Software Update
Z-0392-2011 | 56988 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 4100/4100IQ GE Innova 4100 / 4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector)The Innova systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational imaging procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. Those devices are not intended for mammography applications. | Innova 4100 | 2 | November 16, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of a potential video splitter failure with certain Innova 2000 and 4100 X-ray systems that may impact patient safety.A potential video splitter failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray image display. Should this occur, the systemcannot be recovered without service intervention.GE Healthcare has obse | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent an " Urgent Medical Device Correction" letter dated July 9, 2009, to Hospital Administrators / Risk mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter identified the product, described the Safety Issue, Affected Product Details, Safety Instructions Product correction and Contact Information.Hospitals were instructed to establish emergency workflow procedures should a component stop operating during a patient examination. A GE Healthcare service representative would contact each customer and correct all affected systems.For questions regarding this recall call (262) 548-2731. | Contact the recalling firm for information | 43 | 43 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX UT, VA, WA, WV, and WI, and the countries of UNITED KINGDOM, TUNISIA, TAIWAN, PROVINCE OF CHINA, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN D=FEDERATION, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, ISLAMIC IRAN, INDONESIA, INDIA, GREECE, GERMANY, FRANCE, FINLAND , EGYPT, CHILE, CANADA, BULGARIA, BELGIUM BELARUS, AUSTRALIA. | Terminated | May 17, 2012 | 548 | 408 | Not_Computer
Z-0402-2011 | 56876 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Vortex VX Port, AngioDynamics, Inc., One Horizon Way, Manchester, GA 31816 | Vortex VX Port | 2 | November 17, 2010 | 2010 | Angiodynamics, INC | Labeling Error; The labeling for the specific products incorrectly identified that the packaged device kit contains a Lifeguard Safety Infusion Set. | OTHER/UNDETERMINED: Pending | AngioDynamics issued a Notification of Medical Device Correction letter dated September 27, 2010 to customers. The notification served to make the customer aware of the mislabeling. Customers were instructed not to return the affected products.Customers can contact AngioDynamics at 1 800 772-6446 about this correction. | Contact the recalling firm for information | 6000 | 6000 units (includes all port systems listed) | Nationwide Distribution : Throughout the US. | Terminated | June 21, 2011 | 216 | 6000 | Not_Computer
Z-0407-2011 | 56967 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131--1Q Biplane Cardiovascular Imaging System. The Innova Biplane systems are the first and only digital flat panel biplane systems with a full sized lateral plane to cover lateral anatomy without requiring multiple contrast injections and radiation exposures. The Innova 3131IQ (3131-IQ) and 2121 IQ (2121-IQ) biplane systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational angiography procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic, and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. These devices are not intended for mammography applications. | Innova 2121IQ/31311Q Biplane Cardiovascular Imaging System | 2 | November 17, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue regarding Innova systems with GE Fluoro UPS 20KVA that may impact patient safety.GE Fluoro UPS 20KVA (Uninterrupted Power Supply) is used to maintain fluoroscopy availability for approximately 5 minutes should there be a main power failure. Once power is restored, the equipment power will be returned to the main hospital power.UPS fluoro ba | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter on August 5, 2009 to all Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letter was dated August 4, 2009, and described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For questions regarding this recall call (262) 548-2731. |  | 32 | 32 (2121IQ); 37 (3131IQ) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK ,OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI and the countries of ANGOLA, BELARUS, BULGARIA, CANADA, EGYPT, FRANCE, GERMANY, HUNGARY, IRELAND, MEXICO, ITALY, NIGERIA, NORWAY, POLAND, PORTUGAL, RUSSIAN FEDERATION, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, TUNISIA, TURKEY, and the UNITED KINGDOM | Terminated | January 16, 2014 | 1156 | 578 | Not_Computer
Z-0399-2011 | 57005 | activated whole blood clotting time | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | i-STAT ACT Celite Cartridgesin vitro diagnostic test used to monitor moderate and high-level heparin therapy through analysis of arterial and venous whole blood samples. | iSTAT ACT Celite Cartridges | 2 | November 17, 2010 | 2010 | Abbott Point Of Care Inc. | Certain lots of i-STAT Celite ACT and Kaolin ACT cartridges may be difficult to fill or will not fill when attempting to transfer blood to the sample entry port of the cartridge. | OTHER/UNDETERMINED: Pending | Abbott Point of Care, Inc. sent an URGENT RECALL NOTICE dated September 2010 on October 6, 2010 by Federal Express to all US customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from the identified lots. If customers did not currently have replacement cartridges then they should continue to use the cartridge lot until their facility receives replacement material.Upon receipt of replacement material, customers were to discontinue use of and return the affected cartridges and switch to the replacement material. All unused cartridges will be credited and will be returned.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. | No consumer action necessary | 49025 | 49,025 cartridges. | Worldwide Distribution | Terminated | August 22, 2012 | 644 | 144350 | Not_Computer
Z-0398-2011 | 57029 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Gamma Camera Gearbox brake, a component of the SKYLight Gamma Cameras 3/8" SPECT, Model number 4535 600 66661; and 5/8" SPECT, Model number 4535 602 20381, manufactured by Philips Medical Systems (Cleveland), Cleveland OH.Intended to produce images depicting the anatomical distributions of single photon and positron emitting radioisotopes within the human body for interpretation by medical personnel. | SKYLight Gamma Camera | 2 | November 17, 2010 | 2010 | Philips Medical Systems | Detector arm assembly may slide to hardware limit, potentially causing impact. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Firm will notify consignees via Urgent - Medical Device Correction, dated Nov 3, 2010, via FedEx. The letter identified the affected product, described the problem, the hazard involved, how to identify affected products, actions to be taken, and actions planned by Philips. Customers are to contact their Field Safety Engineer immediately to schedule an inspection of their system. If customers suspect that the identified problem has occurred with their device, they should immediately stop using the system until Philips has confirmed it as safe to use. If customers need further information concerning this issue, they are to contact their local Philips representative. |  | 24 | 24 units | Worldwide Distribution -- USA and the Netherlands. | Terminated | March 21, 2012 | 490 | 24 | Not_Computer
Z-0401-2011 | 57015 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Disposable Angiographic Catheter (Left Heart), manufactured for Millar Instruments, Inc 6001-A Gulf Freeway, Houston, Texas 77023 | Millar MikroTip Disposable Angiographic Catheter (Left Heart) | 2 | November 17, 2010 | 2010 | Millar Instruments, Inc | The sterile packaging was found to be compromised when tested after shipping conditions. This may result in infection or transmission of disease. A test for transportation of sterile package was conducted and resulted in several samples failing. | PRODUCTION CONTROLS: Packaging Process Control | Millar issued a Recall Report letter on October 14, 2010 to consignees identifying the affected product, sterility issue, and actions to be taken. Customers are to return all unused catheters to Millar for replacement. Customers can contact Millar about this recall at 832-667-7131. |  | 12 | 12 Units | Worldwide Distribution: USA, including the states of MA, OH, PA, and D.C., and the countries of Canada and Germany. | Terminated | February 15, 2011 | 90 | 12 | Not_Computer
Z-0412-2011 | 52108 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips EasyVision MM workstation with software R10.2 and R11.1The EasyVision MM Workstation is intended for the manipulation and display of medical images, including mammograms. It can show images from different modalities and interfaces to various image storage and printing devices using DICOM or similar interface standards. Device options enable mammography reading, teleradiology, fast demonstration, 3D, etc. Device accessories enable orthopedic pre-operative planning; angiography/cardiology imaging capabilities; and bone densitometry capabilities. Typical users are trained professionals such as physicians, radiologists, nurses, medical technicians and assistants. | Philips EasyVision MM workstation with software | 2 | November 18, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Review by firm of Philips EasyVision MM workstations found that to assure compliance with 21 CFR 900.12(c)(5) corrected software was needed so that all sites are in compliance. | OTHER/UNDETERMINED: Pending | Philips Medical Systems North America Co. sent an URGENT-Field Safety Notice dated January 30, 2009, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised to ascertain the quality of a hard or soft copy explicity before further distribution to make sure that no part of the tissue is covered by image identification information. Philips will issue the Field Change Order (FCO) 83000129 and release a level on EasyVision MM R10.2 L6 and EasyVision MM R11.1 L3 to resolve the issue.Customers could contact Philips Healthcare Call Center at 1-800-722-9377, #2, #3 and Reference FCO 83000129 for any questions. |  | 11 | 11 units in US | Nationwide Distribution including DC, FL, MI, MO, NM, NY, OH, AND TX | Terminated | November 19, 2010 | 1 | 11 | Software |  | Software update | Software Update
Z-0416-2011 | 56978 | vehicle, motorized 3-wheeled | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Rascal LiteWay 224 manufactured by Giant Manufacturing, Taiwan. | Rascal LiteWay Scooter Model 224 | 2 | November 18, 2010 | 2010 | Electric Mobility Corp | Scooter seat bases were found to have insufficient welding which may lead to the seat base becoming loose and causing the seat to wobble. This may result in complete failure of the seat base, causing the user to fall off the vehicle. | DESIGN: Device Design | The Rascal Company (Electric Mobility Corporation) sent Urgent Medical Device Recall Notice Scooter Seat Base letters dated October 7, 2010 and November 11, 2010 letter to their affected retail customers and affected dealers. The letter instructs the reader that the seat base may become loose, causing the seat to wobble and if not replaced the wobble could lead to a complete failure of the seat base and cause an individual to fall off their vehicle and become injured. The letters also instruct the reader to replace seat bases immediately. Respone forms are attached. Seats with new seat bases will be shipped with return shipping labels to return the original base to the factory. Service representatives will replace the seat base for customers.Customers can contact the Rascal Company about this recall at 1-800-662-4548. |  | 166 | 166 units | Worldwide Distribution: Throughout the US and Canada | Open |  |  | 653 | Not_Computer
Z-0346-2011 | 56874 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | ATM100, LB-AP-M0001Fibrijet, Catalog Number ATM100, Autoloous Thrombin Mixer (Blending Connector), Micromedics, Sterilized by Ethylene Oxide, Rx Only, MicromedicsInc., St Paul, MN 55121-1385. Distributed by Lifebond, P.O. Box 3048 Caesarea Industrial Park 38900, Israel, LP-AP-M0001, LifeSeal SLR Applicator Sterilized by Ethylene Oxide, Manufactured By Micromedics, St Paul, MN 55121-1385 USA. | ATM100, LBAPM0001 Fibrijet, Catalog Number ATM100, Autologous Thrombin Mixer (Blending Connector) | 1 | November 18, 2010 | 2010 | Micromedics, Inc. | Micromedics, Inc is conducting a product recall on a number of medical device products because of weak seals of the sterile pouches potentially resulting in a non-sterile product which may cause transmission of disease or infection. | PRODUCTION CONTROLS: Packaging Process Control | Micromedics issued a Product Recall letter dated March 15, 2010 to customers, describing the product being recalled, the product sterility issue, and actions to be taken. The letter advised customers to contact Customer service to return any affected product to Micromedics for replacement.Customers can contact Micromedics at 1-800-624-5662 or 651-452-1977 ext. 226 concerning this recall. | Contact the recalling firm for information | 620 | 620 (800 pouches) | Nationwide Distribution in the states of CA, CT, DE, FL, GA, ID, IL, IN, MD, MA, MI, MN, NJ, NY, NC, PA, TX, UT, and WI.FRANCE, NETHERLANDS, BELGIUM, UNITED KINGDOM, CANADA, ISRAEL, PERU, SPAIN, IRELAND. | Terminated | June 21, 2011 | 215 | 4880 | Not_Computer
Z-0421-2011 | 57041 | suture, nonabsorbable, synthetic, polyethylene | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Smith & Nephew FAST FIX 360, REVERSE CURVED NEEDLECatalog Number; 72202469.Intended for use as a suture retention device to facilitate percutaneous or endoscopic soft tissue procedures such as shoulder stabilization (Bankart Repair), rotator cuff repair, meniscal repair, and gastrostomy. | Smith & Nephew | 2 | November 18, 2010 | 2010 | Smith & Nephew, Inc. Endoscopy Division | FasT-Fix 360 Needle Delivery System may not allow the T-2 (anchor) to advance for implantation. | DESIGN: Device Design | Smith & Nephew notified customers by an Urgent - Product Recall 1st Notification - Urgent letter sent via Federal Express October 14, 2010. The letter identified the affected product and explained the problem. The letter asked customers to call for a Return Authorization (RA) Number. Customers are to complete the form and return of the affected product along with a copy of the letter to the address provided being sure to reference the RA number. Questions were to be directed to Cindy Burns at 508-261-3655. Consignees out side the US were notified by e-mail. |  | 130 | 130 units | Worldwide Distribution - USA, Dubai, and Great Britain. | Terminated | May 09, 2012 | 538 | 804 | Not_Computer
Z-0418-2011 | 57094 | implant, intragastric for morbid obesity | LTI |  |  | Gastroenterology/Urology |  | LAP-BAND Adjustable Gastric Band System; Product codes:B-2105 (ACCSS PRT II KIT (0-10CC) US),B-2106 (ACCESS PRT II KIT (0-14CC) US),B-2210 (LP-BND SYS, 9.75, STER, US),B-2215 (LAP-BAND 9.75 W/ACC PRT II US),B-2220 (LP-BND SYS, 10.0, STER, US),B-2225 (LAP-BND 10.0,ACCESS PORT II,US),B-2255 (LAP-BAND VG W/ACS PORT II (US),B-2260 (LP-BND AP STANDARD SYS, W ACC PRT II),B-2265 (LP-BND AP LG SYS, W/ACC PRT II);The Lap-Band System is a long-term implantable device intended to induce weight loss in morbidly obese patients by limiting food consumption restrictive and satiating, rather than malabsorptive). The device is surgically implanted, using either a laparoscopic or open procedure, to create a restricted opening (stoma) and a small gastric pouch to limit food consumption and induce early satiety. The main components of the device are the silicone elastomer band, access port. and kink-resistant tubing used to connect the other two components. The inner surface of the silicone band, which is placed around the stomach, is inflatable and connected by the tubing to the access port (a remote injection site. The access port is implanted on, or attached to, the rectus muscle to permit non-surgical, percutaneous adjustments to the band and thus, the stoma diameter, using sterile saline. | LAPBAND Adjustable Gastric Band System | 2 | November 18, 2010 | 2010 | Allergan | In a LAP-BAND System adjustment, failure to inject the needle perpendicular to the access port septum, as described in Directions for Use, may lead to port damage, resulting in subsequent port leakage. Leakage results in deflation of the LAP-BAND to its widest, open position requiring a procedure to replace the port. The patient may have symptoms of reduced satiety and increased appetite. Stalling | DESIGN: Process Design | Allergan sent an URGENT FIELD CORRECTIVE ACTION letter dated September 15, 2010, to all affected customers. The recall communication was initiated on 09/16/2010 when the firm begin forwarding Customer letters with attached current Directions for Use distributed via Federal Express, with tracking verification provided by Federal Express, to the direct consignees.The Customer letters informed consignees of the products affected, description of the Issue, clinical implications and firm recommendations.Customers were instructed to follow the Directions for Use for inserting the needle. Customers were instructed to report any malfunction or adverse event related to the device to Allergan Product Support at 1-800-624-4261 ext. 5972 or 1-805-961-5972 and FDA's Medwatch Program (1-800-FDA-1088 or www.fda.gov/medwatch). For questions, consignees were instructed to contact their Allergan field representative or contact Allergan Product Support at 1-800-624-4261. |  | 152000 | Approx. 152,000 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, AND WY | Terminated | December 22, 2010 | 34 | 152000 | Not_Computer
Z-0411-2011 | 57047 | orthopedic stereotaxic instrument | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | smith & nephew TRIGEN SURESHOT Distal Targeting System, Trauma Interface, V2.0.2, Trauma Launcher V1.01.1Product is intended to be an intraoperative image-guided localization system. It is a computer-assisted orthopedic surgery tool to aid the surgeon with drill positioning for screws during intramedullary nail implantation. | TRIGEN SURESHOT Distal Targeting System | 2 | November 18, 2010 | 2010 | Smith & Nephew Inc | Targeting software has been found to provide erroneous targeting when used with the 13 mm TAN/FAN Left nails. | OTHER/UNDETERMINED: Pending | Smith & Nephew Inc. sent an URGENT MEDICAL DEVICE CORRECTION letter by FedEx on October 15, 2010, to all customers. The letter was dated October 4, 2010, and identified the product, the problem, and the action to be taken by the customers. Customers were instructed to perform a software upgrade and complete the form provided and return by fax to (901) 566-7835. For questions regarding this recall call (901) 399-5441. |  | 445 | 445 units | Worldwide Distribution - USA including AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NM, OH, PA, TN, TX, AND WA and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, South Africa, Spain, Sweden, Switzerland, Turkey, and the UK. | Terminated | February 17, 2011 | 91 | 445 | Software | Device Operation | Software update | Software Update
Z-0414-2011 | 57039 | display, cathode-ray tube, medical | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio (ICIP), Clinical Management SystemSoftware Version D.00 through D.03. | Philips IntelliVue Clinical Information Portfolio | 2 | November 18, 2010 | 2010 | Philips Healthcare Inc. | Software: Scheduling function of the Philips Intellivue Clinical Information Portfolio Releases D.00 through D.03. Edits to a scheduled even frequency order may result in pending interventions not being generated after the first 24 hours. | DESIGN: Software Design | Philips Healthcare issued an Urgent Medical Device Correction letter dated October 15, 2010 to customers. Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel. Contact Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377. | Contact the recalling firm for information | 200 | 200 units | Worldwide Distribution: USA, including states of CA, DC, FL, IL, IN, KY, MA, MD, ME, NJ, NY, OK, RI, TX, VA, VT, and WV; and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and the UNITED KINGDOM. | Open |  |  | 200 | Software | Device Operation | Software update | Software Update
Z-0425-2011 | 56380 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa OpenSpeed, 0.7T MR Systems (HFO3-5). | Signa OpenSpeed | 2 | November 19, 2010 | 2010 | GE Healthcare, LLC | If the lateral table adjustment crank handle breaks off, it has sufficient attractive force to the magnet and may result in patient injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent on 8/23/10 a GE "Urgent Medical Device Correction" letter dated August 20, 2010. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Customers were asked to discontinue use of their MR system and to contact their local service representative, if the lateral table adjustment crank handle on the table becomes loose. | No consumer action necessary | 176 | 176 | Worldwide Distribution -- USA, including states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, LA, KS, KY, LA, MD, MA, MI, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WV, and WI and countries of VENEZUELA, ECUADOR, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, SPAIN, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PHILIPPINE,S OMAN, MEXICO, KUWAIT, JAPAN, ITALY, HUNGARY, GERMANY, EGYPT, DENMARK, CHILE, CHINA, and BRASIL. | Terminated | July 13, 2012 | 602 | 378 | Not_Computer
Z-0424-2011 | 57010 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BARDÂŽ VentrioÂŽ Small Circle Hernia Patch 3.0"Product Code: 0010213Indicated for use in the reconstruction of soft tissue deficiencies such as for the repair of hernias. | Bard Ventrio | 2 | November 19, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Hernia Patch may be a different size than on the product label | PRODUCTION CONTROLS: Packaging Process Control | Davol notified accounts by an Urgent: Medical Device Recall letter on 10/15/10 via overnight delivery. The letter stated that the product maybe a different size than indicated on the product label. It also identified the affected product. Accounts are requested to examine their inventory, remove, and return recalled product. Customers are to notify any of their customers that they may have further distributed the affected product to by including a copy of the recall notice. Customers should contact Davol Customer Service 1-800-556-6275 for instructions on how to return the affected product. Customers are also requested to complete and fax back the attached fax sheet to 1-401-825-8753. If there are any additional questions, customers should contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Service Support Department at 1-800-562-0027. | No consumer action necessary | 26 | 26 units | Nationwide Distribution | Terminated | October 27, 2011 | 342 | 406 | Not_Computer
Z-0432-2011 | 49384 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart FRx Defibrillator, Automated External Defibrillator (AED), catalog number 861304. | Philips HeartStart FRx Defibrillator, Automated External Defibrillator (AED), catalog number 861304. | 3 | November 22, 2010 | 2010 | Philips Medical Systems | The affected device may not deliver appropriate defibrillation therapy. | TRAINING: Employee Error | Philips notified the customer with the affected device by phone on August 8, 2008 . A replacement device was shipped to the customer on August 11, 2008. |  | 1 | 1 unit | Nationwide Distribution: to one account in Arkansas. | Terminated | December 01, 2010 | 9 | 1 | Not_Computer
Z-0437-2011 | 56215 | prosthesis, wrist, carpal trapezium | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | The Ascension PyroHemiSphere (aka NuGrip), Model #30-L NUG-443-30L, Catalog #NUG-443-30L, Lot #10-0731, is a single-use uncemented, one-component prosthesis for the basal thumb joint.The NuGrip is intended to replace the proximal end of the first metacarpal. | Ascension PyroHemiSphere (aka NuGrip), | 2 | November 22, 2010 | 2010 | Ascension Orthopedics, Inc | Two different product lots were found to be incorrectly labeled and placed into the incorrect box. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Ascension Orthopedics notified by phone on June 25, 2010 and sent a letter dated July 22, 2010, to all customers. The customers were instructed to return the devices to the firm.Should you have any questions or concerns please feel free to contact (512) 836-5001 ext 1585. | Contact the recalling firm for information | 9 | 9 | Worldwide distribution: USA including states of: IL, IN, MD, OH, and TN; and countries of : Puerto Rico and Netherlands. | Terminated | June 30, 2011 | 220 | 19 | Not_Computer
Z-0430-2011 | 57098 | implantable pulse generator, pacemaker (non-crt) | LWP |  |  | Cardiovascular |  | Medtronic Dual Chamber Temporary Pacemaker, model 5388. Manufacturer: Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. | Medtronic DualChamber External Temporary Pulse Generators Model 5388 | 2 | November 22, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | A subset of Model 5388 Dual-Chamber External Temporary Pulse Generators worldwide may be unable to power up or may power down unexpectedly. The issue presents itself in one of two ways. First, during startup, and prior to initiating patient therapy, the instrument may power down in 1 - 2 seconds. Second, the instrument may power up correctly, but power down at a later time while in use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic issued an Important Medical Device Correction letters, dated October 2010 and November 2010 to hospital's biomedical engineering department. The letter described the issue, how it presents, and stated that Medtronic will contact the hospitals to return affected product for servicing. The second letter reiterated that Medtronic would notify customers when FDA approves their design upgrade for returning devices for service. In the interim, hospitals should continue to use normally functioning devices. Cusotmers may contact Medtronic Instruments Technical Service at 800-638-1991 concerning this action. | Contact the recalling firm for information | 944 | 5,595 ( 2,696 US, 2,899 OUS) EXPANDED: 944 (852 US, 92 OUS) | Worldwide Distribution: throughout USA and territories of Guam and Puerto Rico and to the countries of: Algeria, Argentina, Australia, Austria, Barbados, Belarus, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, French Guiana, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Malaysia, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, and Yemen NO ADDITIONAL COUNTRIES ARE AFFECTED BY THE EXPANSION. | Terminated | May 29, 2012 | 554 | 6539 | Hardware | Device Operation | Return for servicing | Remove or Replace
Z-0429-2011 | 56017 | test, system, immunoassay, lipoprotein-associated phospholipase a2 | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit; Manufactured and distributed by diaDexus , Inc., South San Francisco, CA; Model Part Number: 90123The diaDexusPLAC Test ELISA Kit is an enzyme Immunoassay for the quantitative determination of Lp-PLA2 (lipoprotein-associated phospholipase A2) in human plasma and serum, to be used in conjunction with clinical evaluation and patient risk assessment as an aid in predicting risk for coronary hear disease, and ischemic stroke associated with atherosclerosis. | PLAC Test ELISA Kit | 2 | November 22, 2010 | 2010 | DiaDexus, Inc | The reason for the recall is that Real Time Stability Test program showed the kit lot failed stability prior to expiration, affecting samples with results | OTHER/UNDETERMINED: Pending | DiaDexus notified all customers by phone, including distributor's customers, beginning on March 15, 2010. Customers were told to re-test samples that were tested from 2/1/2010 through 2/28/2010 which have values above 500 ng/ml with a new lot of product.For questions regarding this recall call (650) 246-6477. | Contact the recalling firm for information | 239 | 239 kits | Nationwide Distribution - US including NJ, OH, KS, IN, FL, CA, UT, OK, PA, WA and MI. | Terminated | January 11, 2011 | 50 | 239 | Not_Computer
Z-0431-2011 | 56416 | radioimmunoassay, testosterones and dihydrotestosterone | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sex Hormone Binding Globulin (SHBG) Reagent for use on Access Immunoassay Systems, Part number(s): A48617The Access SHBG assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of Sex Hormone Binding Globulin levels in human serum and plasma using the Access Immunoassay Systems. | Sex Hormone Binding Blobulin (SHBG) Reagent | 2 | November 22, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the Instruction for Use (IFU) for Access SHBG does not specify the units of measure to be usedwhen calculating the Free Androgen Index (FAI %). Access SHBG (nmol/L) values and Access Testosterone (ng/mL) values are reported in different default units. The testosterone value (when reported in ng/mL) requires conversion to nmol/L before the FAI % can be calculated. |  | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter on the week of July 26, 2010 to the affected customers. The letter was dated July 30, 2010, and identified the product, the problem and provided the appropriate equation to recalculate the FAI% values. The customers were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were also asked to complete and return the enclosed response form within 10 days to confirm that they had received the important notification.The customers were instructed to contact Beckman Coulter Technical Support at 1-800-854-3633 in the United States, or contact their local Beckman Coulter Representative if they had any questions regarding the notification. | Contact the recalling firm for information | 894 | 1,113 Units ( 894 in US & 219 in Canada) | Worldwide Distribution - USA including AL, AR, CA, CO, FL, GA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NV, NY, OH, OK, OR, TN, TX, UT, VA, WA, AND WI and the country of Canada | Terminated | May 04, 2012 | 529 | 1113 | Computer
Z-0434-2011 | 57028 | tape, measuring, rulers and calipers | FTY | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | LeMaitre VascuTape Radiopaque Tape, LeMaitre Stent Guide, 100 stripsCatalog Number: 1102-00Intended to be placed on the patient's skin to assist during imaging procedures by serving as an internal frame of reference. | LeMaitre VascuTape Radiopaque Tape | 3 | November 22, 2010 | 2010 | Lemaitre Vascular, Inc. | Box label (Top Part). Lot and Expiration Date information were switched. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Lemaitre Vascular notified customers by an Urgent: VascuTape Radiopaque Tape Device Field Safety Notice letter on 10/22/10 via USPS with delivery confirmation. The Notice advised users of the incorrect lot and lot number on the outer box and also provided an image of the box and label for ease in identifying the product in their inventory. The letter states that the use and functionality of the device are not affected by the problem; however, if customers feel uncomfortable continuing to use the affected device, they can contact Customer Service for replacement. The firm recommended discarding the outer box of the affected product. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. Accounts can contact customer service at 800-628-9470. |  | 200 | 200 units | Nationwide Distribution -- AL, CA, CT, IL, ME, MI, MS, NC, NJ, NY, PA, PR, TX, and WA. | Terminated | September 15, 2011 | 297 | 1940 | Not_Computer
Z-0443-2011 | 56985 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Cholestech LDX Calibration Verification, Levels 1-4, Catalog Number 11-255.To be used for verifying the reportable range of tests on the Cholestech LDX System. | Cholestech LDX Calibration Verification | 2 | November 22, 2010 | 2010 | Biosite Inc Dba Innovacon Inc. | The recall was initiated because results were being reported as being out of range. Firm is reassigning ranges for HDL cholesterol in product. |  | All customers were notified and given instructions, requirements, and expectations for their role in the Correction activities using the Notice of Correction with attached Returned Verification Forms which were supplied via fax, email or direct mail beginning on October 6, 2010.Customers are to share the information provided in the Notice with their laboratory staff and retain the letter as part of their laboratory Quality System documentation. If any affected product had been further distributed to another laboratory, then a copy of the letter should be provided to them. Customers are to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice.Consignees with questions about the information contained in the notification were instructed to contact:Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121U.S.A.Phone: 877 441 7440FAX: 858 805 8457E-mail: Response.ts@alere.com. | Contact the recalling firm for information | 388 | 388 kits | Worldwide Distribution -- USA, Australia, Canada, Great Britain, India, Italy, Mexico, Norway, Puerto Rico, South Korea, Spain, Switzerland, and Trinidad & Tobago. | Terminated | August 01, 2012 | 618 | 6529 | Not_Computer
Z-0427-2011 | 56913 | system, perfusion, kidney | KDN | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Belzer UW Cold Storage Solution, 1000 mL. Sterile and non-pyrogenic. Formulated for flushing and cold storage of kidney, liver and pancreas organs. Manufactured by: Preservation Solutions, Inc., Elkhorn, WI 53121. Manufactured for: BTL Solutions LLC, a subsidiary of Bridge to Life, Ltd., 1225 Laurel Street, Columbia, SC 29201. | Belzer UW Cold Storage Solution | 3 | November 22, 2010 | 2010 | Preservation Solutions Inc | Mislabeled volume - 2 liter bags of Belzer UW Cold Storage Solution were labeled as 1 liter. | PRODUCTION CONTROLS: Labeling Mix-Ups | Bridge to Life, Ltd, issued a Notification of Product Recall letter dated October 18, 2010 to all direct accounts identifying the affected product and labeling issue. Customers were informed that all product from the affected lot has been replaced by the firm. Customers with questions about this recall can contact Bridge to Life at 803 673-4548.. | Contact the recalling firm for information | 55 | 55 units | Nationwide Distribution - in US to the states of CA, MO, and PA. | Terminated | April 02, 2012 | 497 | 55 | Not_Computer
Z-0448-2011 | 55200 | radioimmunoassay, estradiol | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Estradiol, Part Number: 33540Intended use: The Access Estradiol assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of Estradiol levels in human serum and plasma using the Access Immunoassay systems. | Access Immunoassay Systems Estradiol | 2 | November 23, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that three lots of Access Estradiol reagent (REF 33540) contains an incorrect raw material which may produce incorrect patient results. | OTHER/UNDETERMINED: Pending | A Product Corrective Action (PCA) letter dated March 17, 2010, with attached PCA Customer Response Form was sent to the affected customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to discontinue use of the three Access Estradiol Reagent lots; discard all remaining inventory of affected product; share this information with laboratory staff and retain the notification as part of their laboratory Quality System documentation, and complete and return the enclosed response form within 10 days via fax to 786-639-4000 or mail: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs.The following product replacement contacts were provided:In the United States, please contact Client Services at 1-800-526-3821.In Canada, please contact Customer Service at 1-800-463-7828.Outside the United States and Canada, contact your local Beckman Coulter Representative.Questions and assistance could be obtained through Technical Support at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact consignees were told to contact their local Beckman Coulter Representative. | Contact the recalling firm for information | 20209 | 20,209 units (5475 in US) | Worldwide distribution: USA and Internationally to: Algeria, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Czech Republic, El Salvador, France, Georgia, Germany, Ghana, Greece, India, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, United Kingdom, Venezuela, Vietnam, and Zambia | Terminated | May 07, 2012 | 531 | 20209 | Not_Computer
Z-0444-2011 | 55889 | dental implant surgical tray (kit) | OFY | Endosseous dental implant accessories. | Dental | Dental | Enforcement Discretion | Zimmer Dental, Drill stop kit, REF DSKIT, for use with dental implants. | Zimmer dental | 2 | November 23, 2010 | 2010 | Zimmer Inc. | The instructions for use contain pre-vacuum sterilization instructions that may lead to compromised sterility at the lid-kit interface location. | DESIGN: Labeling Design | Zimmer sent out "Urgent Device Correction" letters dated 5/18/2010. The consignees were provided with new instructions for use. The letter informed users that the current pre-vacuum cycle of 3 to 18 minutes is not enough time to assure sterility between the lid-kit interface and could lead to cross contamination of multiple patients. The users were informed that a pre-vacuum time of 30 minutes was required to assure sterility. |  | 158 | 158 | Distribution Nationwide, to France, and Spain. | Terminated | May 18, 2011 | 176 | 158 | Not_Computer
Z-0445-2011 | 57282 | [6 empty] | Lumen Biomedical Xtract Aspiration Catheter system, intravascular catheter, consisting of One Catheter, Two 30 mL Syringes, One Extension Tube with Stopcock, and One 40 um Cell Strainer. Single-user, 0.014" guidewire compatible intravascular extraction and aspiration catheter, Part number 1450-0001 size 4.2F, and part number 1451-003 size 5.8F, distributed by Volcano Corporation. | Lumen Biomedical Xtract Aspiration Catheter system | 2 | November 23, 2010 | 2010 | Volcano Corporation | Potential for product failure. Packaging/Storage issue results in compromised integrity and potential product failure during use, leading to potential patient injury. | DESIGN: Device Design | Firm initiated recall on 11/03/2010, and sent notifications to its consignees via letter. The firm has requested that consignees provide the notification to all personnel within the organization and letters will be sent directly to known customers on 11/04/2010 by Volcano. Customers have been instructed to cease use and contact customer service to arrange returns, and credit will be issued. | Discard product or return to place of purchase for a full refund or exchange | 4724 | 4724 units | Worldwide distribution: USA, Greece, Egypt, Kuwait, Netherlands, Saudi Arabia, Italy, France, Sweden, Portugal, Great Britain, Austria, Switzerland, Germany, Romania, Poland, and Bulgaria | Terminated | December 07, 2011 | 379 | 4724 | Not_Computer
Z-0446-2011 | 57008 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BARDÂŽ Mesh Pre-shaped (1.8" x 4.0"/4.5cm x 10cm) Product code: 0112700 Intended use: Bard Pre-shaped mesh is a pre-shaped mesh constructed of knitted polypropylene monofilaments. It is indicated for the repair of inguinal hernia defects | BARD Mesh Preshaped | 2 | November 23, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: The labeled BARDÂż Mesh Pre-shaped product code 0112700 (1.8" x 4.0"/4.5cm x 10cm), but the contents are the BARDÂż Mesh Large Pre-shaped product code 0113700 (2.4" x 5.4"/6.0cm x 13.7cm). | PRODUCTION CONTROLS: Process Control | Davol issued an "URGENT MEDICAL DEVICE RECALL" letter on October 15, 2010, via Overnight Delivery to all customers. The letter describes the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quaratine the product; identify and notify customers they may have further distributed the product to of this recall; contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product, and complete and return the attached EFFECTIVENESS CHECK form via fax to 1-401-825-8753. If you have an questions, please contact Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Service Support Department at 1-800-562-0027. | No consumer action necessary | 144 | 144 units | Nationwide distribution | Terminated | October 27, 2011 | 338 | 144 | Not_Computer
Z-0447-2011 | 57048 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-Arm Imaging System, Door Assembly (BI-700-00006) used in Product Catalog Numbers:9732719,9733346,81-700-00027, 81-700-00027R, 81700-00027-100, 81-700-00027-1 OOR, 81-700-00027-120, 81-700-00027-120R, 81700-00027-230, 81-700-00027-230R, 81-700-00027GER, 81-700-00027GERRIntended use: Designed for 2D fluoroscopic and 3D imaging and is intended to be used by a physician benefits from 2D and 3D information of anatomic structures with high x-ray attenuation such as bony and metallic objects. | Medtronic Navigation OArm Imaging System | 2 | November 23, 2010 | 2010 | Medtronic Navigation, Inc. | Safety Notice: Door Cable Periodic Maintenance Schedule and Cable Wear for the O-ARMÂżÂż Imaging System Breakage of the door cable during opening and closing may occur and cause physical harm or delay of therapy | PRODUCTION CONTROLS: Equipment Maintenance | Medtronic Navigation sent all consignees/customers an "Urgent Field Safety Notice" letter dated October 15, 2010. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review and follow the attached information that describes the process to properly detect and proactively manage cable wear; to ensure that the door is completely open when positioning the patient into or removing them from the gantry and when placing the O-ARM Imaging System Tube Drape (BI-900-00003) onto the unit, and report any visual or audible (cracking noise) wear of the cable to Medtronic Representative ASAP. Note: A Medtronic Representative will be contacting the customers within the coming weeks to schedule a time to inspect the current condition of the cable, explain and train the customers on the recommended workflow and inspection procedure and place a permanent label on their O-ARM Imaging System to reinforce these actions. If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or 720-890-3200. | Contact the recalling firm for information | 244 | 244 units | Worldwide distribution: USA including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, and WI; and countries including: Canada, Austria, Australia, Belgium, Brazil, Libya, South Africa, India, Japan, Singapore South Korea, Haiti, Kuwait, Saudi Arabia, Turkey, UAE, Germany, Netherlands, Great Britain, Italy, Poland, Netherlands, Spain,Sweden, and Switzerland. | Terminated | June 04, 2012 | 559 | 244 | Not_Computer
Z-0456-2011 | 55178 | electrode, ion based, enzymatic, creatinine | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CREA Cartridge, Abbott List Number: 06F10-01, 02; IVD. Manufactured by Abbott Point of Care, Inc., Princeton, NJ. | iSTAT CREA Cartridge Blood Gas Panel | 2 | November 24, 2010 | 2010 | Abbott Point of Care Inc. | The i-STAT CREA Cartridges may produce inaccurate (depressed) results at or near the end of the labeled expiry date when stored above 25 degrees for 2 weeks. | OTHER/UNDETERMINED: Pending | Abbott Point of Care issued Urgent Product Recall Notification letters dated March 2010 to all direct customers identifying the affected product and the actions to be taken. Customers were instructed to determine if they had cartirdges that had been stored above 25 degrees Celsius, and if so, contact Abbott for replacement or credit. Abbott Technical Support can be contacted at 800-366-8020 option 1 about this recall. |  | 34525 | 34525 cartridges | Worldwide Distribution: Throughout USA and Canada. | Terminated | September 22, 2011 | 302 | 34525 | Not_Computer
Z-0458-2011 | 55604 | radioimmunoassay, free thyroxine | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Dimension FT4L Flex reagent cartridge( Free Thyroxine) manufactured by Siemens Healthcare Diagnostics, Inc., Newark, DE. | The Dimension FT4L Flex reagent cartridge | 2 | November 24, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Some LOCI cartridges have sub-optimal seal of the lidstock that can cause rupture of the package at high elevations, causing the cartridge to leak. Use of leaking cartridges may result in inaccurate test results. | OTHER/UNDETERMINED: Pending | Siemens issued an "Urgent Field Safety Notice" dated April 2010, informing users of the problem and instructing them to inspect cartridges for leaking fluid. Cartridges with leaking fluid should be reported to the Siemens Technical Solutions Center. The Technical Solutions Center can be contacted at 800 441 9250. |  | 1520 | 1520 | Worldwide Distribution: USA, Europe, Australia, China, Japan, Korea, and Saudi Arabia. | Terminated | March 19, 2013 | 846 | 14037 | Not_Computer
Z-0449-2011 | 56345 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV, Artiste Digital Linear Accelerators in combination with the Flat Panel, Part No. 8139789,To deliver x-ray radiation for therapeutic treatment of cancer.Product is manufactured and distributed by Siemens Medical Solutions, Concord, CA | ARTISTE Linear Accelerator | 2 | November 24, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The robotic lift armature shaft may break and allow the Flat Panel Positioner with the digital flat panel imager to fall into the Flat Panel cradle. If the positioner is over a patient, there is a potential for injury. | OTHER/UNDETERMINED: Pending | Siemens sent an Urgent Medical Device Correction letter to all on January 27, 2010, via phone, Fax, certified mail and/or delivery via Siemens Customer Service Group to all its direct consignees, informing them of the affected products, the problem, and provided instructions on the recall.Consigneees were to be contacted by an ARTISTE Service Representative for an appointment for corrective action. Consignees were instructed to include Customer Safety Advisory Notice in their ARTISTE System Owner Manual.For questions regarding this recall call (925) 246-8407. |  | 106 | 106 | Worldwide Distribution - USA including CA, DE, FL, LA, MA, MI, MO, NC, NE, NJ, OH, PA, TN, UT, WI, AND WV and Egypt, Germany, Japan, Lebanon, Malaysia, New Zealand, Poland, Russia, Spain, Sweden, and Turkey | Terminated | April 12, 2011 | 139 | 106 | Not_Computer
Z-0452-2011 | 56912 | dura substitute | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | DuraGen XS Dural Regeneration Matrix: 4 inches x 5 inches.Non Pyrogenic; Rx only;Integra LifeSciences Corporation, Plainsboro, NJ 08536 USADuraGen XS is indicated as a dural substitute for repair of dura matter. | DuraGen XS" Dural Regeneration Matrix | 2 | November 24, 2010 | 2010 | Integra LifeSciences Corp. | Test results no longer support a 3 year shelf life. | OTHER/UNDETERMINED: Pending | Integra LifeSciences Corporation sent an URGENT Product Recall Notification letter dated September 29, 2010, to all affected customers. Integra NeruoSpecialists were informed of the recall via telephone conference call and recall notification letters sent by e-mail. Letters were also sent to customers by Federal Express with copies of e-mail provided to Interga Neuro Specialists on October 1, 2010.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately quarantine any affected product and contact Integra Customer Service at 1-800-654-2873, select option 5, followed by option 3.Customers were instructed to fill out and return the enclosed "Recall Acknowledgement and Return Form" indicating whether or not they possessed any of the affected product.For any questions related to this recall call 1-800-654-2873 and selection option 1. |  | 571 | 571 | Nationwide Distribution including AK, AL, AZ, CA, CO, DC, DE, FL, GA, HI, IL, IN, LA, MA, MD, ME, MN, MO, MT, NC, ND, NH, NJ, NM, NY, OK, PA, TN, TX, UT, VA, VT, WA, WI, AND WV. | Terminated | April 23, 2012 | 516 | 2766 | Not_Computer
Z-0457-2011 | 56147 | ventilator, continuous, minimal ventilatory support,facility use | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, manufactured by Respironics California Inc. in Carlsbad, CA. | Respironics V60 Ventilator | 2 | November 24, 2010 | 2010 | Respironics California Inc | Respironics California Inc has received customer complaints of an odor or burning smell coming from the V60 Ventilator, and identified a component on the Power Management PCB that exhibited thermal damage. | DESIGN: Device Design | Phililps Healthcare issued an Urgent Medical Device Field Correction letter dated July 2, 2010 concerning the RespironicsV60 Ventilator issue. Customers will be contacted by US Field Service Technicians and international distributors to schedule and perform inspections and replacement of the affected component. The unit may continue to be used pending completion of the field inspection.Customers may contact Respironics US Customer Service about this action at 1 800 345 6443, option #5, then option #1. | Contact the recalling firm for information | 504 | 504 | Worldwide Distribution: Throughout the US and Puerto Rico,, and to the countries of: Australia, Austria, Belgium, Canada, Denmark, France, Germany, Hong Kong, India, Italy, Kenya, Kuwait, Libya, Malaysia, Oman, Netherlands, Norway, the United Kingdom, Saudi Arabia, and Spain. | Terminated | December 01, 2010 | 7 | 504 | Hardware | Physical Safety Hazards | Replace component | Remove or Replace
Z-0428-2011 | 57166 | catheter, thrombus retriever | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Penumbra System Reperfusion Catheter 032, catalog number PSC032, lot number lot F15020, manufactured by Penumbra, Alameda, CA. | Penumbra System Reperfusion Catheter 032 | 1 | November 24, 2010 | 2010 | Penumbra Inc. | Reperfusion catheter may fail at the mid-shaft joint and may break off. If the segment cannot be removed, it may result in blocked blood flow, requiring emergency medical treatment. | PRODUCTION CONTROLS: Process Control | Penumbra issued an Urgent Voluntary Field Removal letter dated October 28, 2010 to consignees, identifying the affected device and actions to be taken by consignees. Customers were instructed to inspect their inventory for the affected product and quarantine it, pending return to Penumbra. Penumbra employees will contact customers to arrange return and replacement of affected units. Customers were also requested to complete and return a product identification form and return form to Penumbra within 5 business days.Customers can contact Penumbra at 1 510 748-3223 concerning this action. | Contact your health professional | 61 | 61 | Worldwide Distribution: 17 consignees in the US, including in the states of CA, FL, HI, ID, IL, KY, MA, MS, NJ, NY, PA, and TX, 1 in Canada and 21 in the EU countries of France, Germany, Hungary, Italy, Norway, Slovenia, Spain, and the United Kingdom. | Terminated | June 01, 2011 | 189 | 61 | Not_Computer
Z-0462-2011 | 57142 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SignaÂŽ Excite" 1.5 T Magnetic Resonance Systems, manufactured by GE Healthcare, P.O. Box 44 Milwaukee, WI 53201. | Signa Excite" 1.5 T Magnetic Resonance Systems, GE Healthcare, P.O. Box 44 Milwaukee, WI 53201. | 2 | November 24, 2010 | 2010 | GE Healthcare, LLC | SignaÂż MR Excite" 1.5T system containing 11.1_M4_0818.a software, associated with 2D Fast Spin Echo based pulse sequences (2D FSE-XL, 2D FRFSE-XL, 2D FSE-IR, 2D T2FLAIR and 2D T1FLAIR) may experience an image flip along the phase encode direction and incorrect annotation, which may result in incorrect localization of anatomy in oblique axial images with the pulse sequences. | DESIGN: Software Design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 21, 2009. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Field Engineers will contact customers to schedule a field modification of the affected products to address this issue.Customers can contact GE Healthcare at 262 521-6100 about this issue. |  | 1464 | 1464 | Worldwide Distribution: Throughout USA except states of AK, VT and WY, and the countries of: Argentina, Australia, Bulgaria, Brazil, China, Canada, Switzerland, Chile, Colombia, Czech Republic, Ecuador, Egypt, Spain, Finland, France, Great Britain, Guatemala, Honduras, Croatia, Hungary, Indonesia, Ireland, India, Iran, Iceland, Italy, Japan, Kenya, South Korea, Kuwait, Lebanon, Luxembourg, Libya, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Pakistan, Poland, Puerto Rico, Portugal, Reunion, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Syria, Thailand, Tunisia, Turkey, Taiwan, Venezuela, and South Africa. | Open |  |  | 1464 | Software | Display/Image | Software update | Software Update
Z-0455-2011 | 56012 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Inc. Eclipse/VARiS Vision, an Eclipse radiotherapy Treatment Planning System (Eclipse TPS); Eclipse Version 7.3.10 (multiple workstations) VARiS Vision System 7. x;Model Numbers: H48/H83. | Varian brand Eclipse/VARiS Vision | 2 | November 24, 2010 | 2010 | Varian Medical Systems Oncology Systems | Incorrect treatment plan or record: A software design deficiency has the potential for corrupted treatment plan records or treatment records in cases where multiple users editing a record simultaneously can result in a corrupted (but accessible) treatment plan record. | DESIGN: Software Design | On May 13, 2010, the firm issued letters to all its direct consignees, informing them of the affected products and providing instructions on the recall.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***. |  | 2417 | 2417 units***12-20-12 AMENDED TO: 49 units*** | Nationwide and Worldwide | Terminated | February 24, 2012 | 457 | 2417 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0481-2011 | 50117 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T2 Ankle Arthrodesis System Nail and Nut Adapter;Stryker Trauma GmbH, Germany;Distributed in USA by Howmedica, Osteonics Corp., 325 Corporate Drive, Mahwah, NJ | T2 Ankle Arthrodesis System Nail and Nut Adapter | 2 | November 26, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The thread of the nut which should fix the nail adaptor to the targeting arm may seize on the threat of the nail adaptor, interrupting surgery. | DESIGN: Process Design | Stryker Orthopedics issued an Important Stock Recovery letter dated August 2, 2006 The customer was instructed to inspect inventory for the affected product and return it to the firm. The customer was asked to complete and return a Product Accountability Form and contact their Stryker representative for replacement product.Stryker can be contacted about this issue at 201 831-5118. |  | 8 | 8 sets | Nationwide Distribution: To one location in New Jersey. | Terminated | November 22, 2013 | 1092 | 8 | Not_Computer
Z-0482-2011 | 55760 | intracranial neurovascular stent | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Neuroform 3 Microdelivery Stent System; Distributed and manufactured by: Boston Scientific Corporation, Fremont, CA | Neuroform 3 Microdelivery Stent System | 2 | November 26, 2010 | 2010 | Boston Scientific Corporation | The product may contain manufacturing aids (mandrels) which should not be present in the final product. If the mandrel remains in place, the catheter cannot be loaded over the required guidewire, so cannot be used. | PRODUCTION CONTROLS: Process Control | Boston Scientific issued an Urgent Customer Notice dated March 1, 2010 to all US customers, with personal delivery by sales representative where applicable outside the US, identifying the affected device. Customers were instructed to return the device for replacement, and to return an acknowledgement form.Customers can contact Deborah Springer at Boston Scientific at 800-845-4598 about this recall. |  | 638 | 638 units | Worldwide Distribution: Throughout the US and to the countries of: Austria, Australia, Belgium, Brazil, Canada, China, Columbia, Czech Republic, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hungary, India, Ireland, Italy, South Korea, Luxemburg, Mexico, the Netherlands, New Zealand, Peru, Russian Federation, Sweden, Slovenia, Taiwan and Ukraine. | Terminated | November 26, 2010 | 0 | 638 | Not_Computer
Z-0479-2011 | 52100 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Product's marketing brochure is labeled in part: "Philips Easy Diagnost Eleva Conventional R/F system...PHILIPS". manufactured by Philips Medical Systems Gmbh, Hamburg, Germany. | Philips Easy Diagnost Eleva Conventional R/F system | 2 | November 26, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Patient and image data can be exchanged when a patient is selected in the patient list and the examination tab is pressed with background query running at the same time. There is a risk that the data may then be linked to the incorrect patient, which may adversely affect diagnosis and treatment. | DESIGN: Software Design | Philips Medical Systems issued URGENT-Device Correction letters dated January 27, 2009 to consignees describing the issue with the Easy Diagnost Eleva GXR RF systems. The consignees were informed of the potential for incorrect diagnosis or treatment if an incorrect image is used for diagnosis. Consignees should contact the Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FC0 70600030" with any communication. |  | 314 | 314 units | Nationside Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, and WV. Devices are not distributed to any foreign consignees. | Terminated | December 30, 2010 | 34 | 314 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0469-2011 | 52040 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The product's marketing brochure states: "1.5T coils and accessories...Philips Achieva and Intera 1.5T MR systems".The Identification of the Synergy Flex-M/ Shoulder Coil 1.5T can be found on the driver box. The 12nc of the affected system is labeled as 4522-131-6656x. | ACHIEVA 1.5T and INTERA 1.5T Nuclear Magnetic Resonance Imaging | 2 | November 26, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Combined use of the Synergy Flex-M / Shoulder Coil 1.5T increases the chance of RF interaction and heating up of the coil. This may result in possible burns of the patient. | OTHER/UNDETERMINED: Pending | Philips issued URGENT-Field Safety Notice letters dated May 1, 2009 and October 14, 2009 to their consignees. The letter informed customers that the firm will replace the "old" Flex M/Shoulder Coils with a new Flex-M coil.Consignees could contact Philips Healthcare Call Center at 800-722-9377, #5, #3, #1 and reference FCO 78100290 for any questions. |  | 389 | 389 units | Nationwide Distribution: in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VA, WA, WI, WV, WY, and Washington DC. There were no foreign consignees. | Terminated | January 06, 2011 | 41 | 389 | Not_Computer
Z-0478-2011 | 56428 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | The software is ARM version 60.80.101.025 (referred to as Nano 1.0.5).The software is used with the SonoSite NanoMaxx ultrasound system. | NanoMaxx ultrasound system | 3 | November 26, 2010 | 2010 | Sonosite, Inc. | Software ARM version 60.80.101.025 (referred to as Nano1.0.5) on the NanoMaxx ultrasound system results in errors when the "print" command is selected. As a result, the displayed MI or TI values may be incorrect. | DESIGN: Software Design | Sonosite issued Urgent Medical Device Correction letters dated August 9, 2010 to consignees. Consignees were advised of the software problem and given directions to verify the software version they are using. They are advised to contact SonoSite at 877-657-8118 to obtain a free software upgrade. | Contact the recalling firm for information | 69 | 69 | Worldwide Distribution: USA and to the countries of Australia, Great Britain, India, Israel, the Netherlands, New Zealand, Norway, and Spain. | Terminated | July 21, 2011 | 237 | 69 | Software | Output/Calculation | Software update | Software Update
Z-0473-2011 | 55682 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartStart FR2+ automated external defibrillator (AED) model: M3841: LAERDAL brand; text prompts only on screen (NO ECG); no configurable manual charge. | HeartStart FR2 automated external defibrillator (AED) | 2 | November 26, 2010 | 2010 | Philips Medical Systems | Failure of the voltage detector, U8 may cause the AED battery to drain more rapidly than normal or make the AED unusable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Medical Systems issued an Urgent Medical Device Recall letter dated May 28, 2010 identifying the affected device and actions to be taken. Customers were instructed to review their inventory for affected devices. If affected devices are found, the customer should contact Philips to arrange for a replacement device. The customer should also run a battery insertion test. If the device passes the test, the AED may remain in service until a replacement is received. For further information, contact Philips at 1 800 263-3342, option 6. |  | 1478 | 1,478 units | Worldwide Distribution: Throughout the USA, and to the following countries: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BERMUDA, BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIETNAM. | Terminated | September 13, 2013 | 1022 | 16573 | Other
Z-0480-2011 | 57023 | device, gas generating | KZJ | Gas-generating device. | Microbiology | Microbiology | 510(K) Exempt | GasPak EZ Large Incubation Container, a component of the GasPak EZ Gas Generating Container System, catalog # 260672, labeled in part ***Becton Dickinson and Company Loveton Circle Sparks, MD 21152 USA*** | GasPak EZ Large Incubation Container | 2 | November 26, 2010 | 2010 | Becton Dickinson & Co. | In vitro diagnostic reagent containers may be defective and could cause incorrect test results in patient samples. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter dated August 2010 to consignees. Consignees were advised of the defect and were requested to discard product for replacement. Distributors were additionally requested to provide distribution lists of end users for direct BD notification. Accounts were requested to return a response form.BD Customer Service can be contacted at 800 675-0908, and BD Technical Services can be contacted at 800-638-8663. |  | 496 | 496 units | Worldwide Distribution: USA and the countries of Australila, Belgium, Canada, India, Mexico, New Zealand. Singapore, and South Korea, | Terminated | January 05, 2011 | 40 | 496 | Not_Computer
Z-0476-2011 | 57077 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1 | Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1 | 2 | November 26, 2010 | 2010 | GE Healthcare, LLC | The measurement values provided by the GE Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems.If not noticed by the caregiver, this could lead to an overestimate of the size of the breast lesion that may impact patient treatment and safety. The measurement values for GE images are not affected. | DESIGN: Software Design | GE Healthcare issued "Urgent Medical Device Correction" letters dated August 28, 2009 or August 31, 2009 to either Seno Advantage Customers (Hospital Administrators/Risk Mangers, Radiology Department Managers), and Image Diagnostic Customers (Hospital Administrators/Risk Managers, Radiology Department Managers). The letters described the Safety Issue, Affected Product Details, safety Instructions, Product correction and Contact Information. GE will linstall new software with the release of Field Modification Instruction 12129.Customers can contact GE Healthcare about this issue at 800-437-1171. | Contact the recalling firm for information | 2514 | 2514 | Worldwide Distribution: USA in the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI,MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY, and to the countries of UNITED KINGDOM, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SOUTH, KOREA, SAUDI ARABIA , PORTUGAL, NEW ZEALAND, NETHERLANDS, KUWAIT, JAPAN, ITALY, GERMANY, FRANCE , CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, VENEZUELA, UNITED ARAB EMIRATES, THAILAND, SWEEDEN, SOUTH AFRICA, SLOVENIA, SINGAPORE, RUSSIA, PHILLIPNES, NORWAY, POLAND, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LUXEMBOURG, LEBANON, JORDAN, JAMAICA, ISRAEL, IRELAND, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUADOPUO, GREECE, GEORGIA, FINLAND, EGYPT, DOMINICAN REPUBLIC, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, BULGARIA, BOTSWANA, and ALGERIA. | Terminated | June 05, 2012 | 557 | 2672 | Software | Output/Calculation | Software update | Software Update
Z-0474-2011 | 57090 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Medtronic Sprint Quattro Secure S, model 6935. Steroid eluting, tripolar, screw-in, ventricular lead with RV defibrillation coil electrode. Contents Sterile. Manufactured in: Villalba, Puerto Rico, USA. Medtronic, Inc., Minneapolis, MN 55432 U.S.A. | Medtronic Sprint Quattro Secure S, model 6935 | 2 | November 26, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Over--retraction of the helix during initial implant or subsequent repositioning may result in the inability to extend the helix. This does not impact acute or chronic performance of successfully implanted leads. | DESIGN: Device Design | Medtronic issued an Important Medical Device Correction letter, dated October 2010, to doctors. The letter contains a Performance Note. The letter describes the issue and the performance note provides information on how to mitigate the issue through replacement with a new lead.Customers can contact Medtronic Technical Services at 800 723-4636. | Contact the recalling firm for information | 154000 | 154,000 total | Worldwide Distribution: Throughout the US, including Guam, Puerto Rico, and US Virgin Islands, and to the countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bangladesh, Barbados, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | February 04, 2012 | 435 | 154000 | Not_Computer
Z-0465-2011 | 57117 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Rectal Tubes, 22 French, 20 Long (51 cm), X-Ray Opaque Rubber with Funnel end, Open Tip, One Eye , Non sterile, 12 Each, C.R. Bard, Inc., Covington, GA 30014 | Rectal Tubes 22 French/20 French | 3 | November 26, 2010 | 2010 | C.R. Bard, Inc., Urological Division | The product is misbranded. The product is labeled as 22 French, however the product is actually 20 French. | TRAINING: Employee Error | Bard sent an Urgent Product Recall Notification letter dated October 28, 2010 to the consignee. They were instructed to return any outstanding inventory of the affected product. They were advised to notify their consignees of the recall and to retrieve relevant product from those locations. Affected products received by Bard Medical will be destroyed.BMD can be contacted at 1 770 784-6969 for questions or assistance. |  | 108 | 108 eaches | Nationwide Distribution: In the state of Ohio. | Terminated | March 31, 2011 | 125 | 108 | Not_Computer
Z-0467-2011 | 57095 | tap, bone | HWX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CSS-059-55-4, Modified 5.5 mm Modified Depth Gauge, Lot 0985021 | Modified 5.5 mm Modified Depth Gauge for MaxTorque Screw System | 2 | November 26, 2010 | 2010 | Orthohelix Surgical Designs Inc | The print for modified CSS depth gauges has inaccurate definition of the etched screw length. The error results in a length reading that is longer (5 mm in the case of the 7.0 and 4mm in the cases of the 5.5 and 4.0) than the actual length of the buried guide wire segment and even longer (10 mm in the case of the 7.0 and 8mm in the cases of the 5.5 and 4.0) than the design intention readings. This | DESIGN: Device Design | OrthoHelix issued an Important Recall Notice letter dated October 15, 2010 to their customers, identifying the affected products. Customers were requested to immediately return the afected devices. OrthoHelix can be contacted at 330 247-1444 concerning this recall. |  |  |  | Nationwide distribution: to the states of GA and MO. | Terminated | September 06, 2012 | 650 |  | Not_Computer
Z-0464-2011 | 57109 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-G Digitizer, manufactured by Agfa-Gevart NV, Mortsel, Belgium. | DXG Digitizer | 2 | November 26, 2010 | 2010 | AGFA Corp. | Packaging was switched between a DX-G Digitizer and a DX-M Digitizer, resulting in the wrong product, which is not sold in the US, being shipped to the consignee. | OTHER/UNDETERMINED: Pending | Agfa halted Installation of the incorrect unit, and the unit was sent to Agfa warehouse in Carlstadt, NJ, where it is being held in quarantine for export. |  | 1 | One unit | Nationside Distribution: to NJ. | Terminated | December 30, 2010 | 34 | 1 | Not_Computer
Z-0463-2011 | 57115 | crown, preformed | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Stainless Steel Primary Molar Crowns, Made in USA by 3M ESPE Dental Products, St Paul, MN. Unitek" Primary Stainless Steel Crowns- Series UL3 and UR3 Product Code 900223, 900213, 902150 | 3M ESPE Stainless Steel Primary Molar Crowns | 3 | November 26, 2010 | 2010 | 3M Company / Medical Division | Unitek Stainless Steel Permanent Crowns with product Codes 902150, 900213, and 900223, is mislabeled.The crown contained within the package is incorrectly labeled; series UL3 crowns are labeled series UR3 and vice versa. The lid of the package correctly identified the product number and lot. This mislabeling could result in placement on the right side of a crown intended for use on the left, and | PRODUCTION CONTROLS: Packaging Process Control | 3M ESPE issued a Device Recall letter dated October 6, 2010 to distributors and dentists. The letter listed the product being recalled including code and lot codes, background information including a picture of the product problem, and the recall procedure. The letter included a template letter to be used by distributors to inform their customers who have received the product. Mislabeled product is to be quarantined and returned to 3M ESPE.3M ESPE can be contacted at 651 733-1110. | Contact the recalling firm for information | 9819 | 9819 | Nationwide Distribution: Throughout USA, including the states of AZ, CA, GA, IL, IN, IA, LA, MA, MN, MO, NV,NJ, NY, NC, PA, TN, TX, WA, and WI.Canada and Dominican Republic | Terminated | February 04, 2012 | 435 | 9819 | Not_Computer
Z-0484-2011 | 56816 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartStart FR2+ AED with Text Display (No ECG), Model M3861, Philips brand, no configurable manual charge.The label of the device states: "HEARTSTART...FR2+ Automated External Defibrillator...Manufactured by: Philips Medical Systems Seattle, WA 98121". | Philips HeartStart FR2 AED with Text Display , Model M3861 | 2 | November 29, 2010 | 2010 | Philips Medical Systems | A flash memory component in the HeartStart FR2+ and Automated External Defibrillator was improperly relabeled by a third party. Therefore the component could not be identified by its source, quality, and date of manufacture. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 10/5/10, Philips began sending out the Voluntary Medical Device Removal notification letter to their consignees.This letter informs the consignees that Philips Healthcare is conducting a voluntary product removal of a limited number of HeartStartFR2+ automated external defibrillators (AEDs), model numbers M3860A, M3861A, M3849A, and M3841A, manufactured by Philips and shipped between March 2007 and August 2010. The owners of affected AEDs are advised to run a battery insertion test. If it passes, the AED should remain in service until its replacement is received. Philips will be replacing affected units.The consignees can visit the firm's website www.philips.com/FR2AEDAction or contact Philips at 1-800-263-3342 for any information or support concerning the issue. |  | 169 | 11, 158 units | Devices were distributed to hospitals, medical facilities, and distributors through out the US and to foreign consignees. Foreign consignees including countries:Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Ecuador, El Salvador, Finland, France, French Guiana, Georgia, Germany, Ghana, Gibraltar, Guatemala, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Japan, Jordan, Kampuchea, Kenya, Korea, Latvia, Lebanon, Libya, Luxembourg, Malaysia, Malta & Gozo, Mexico, Netherland, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Vietnam, and Zimbabwe. | Terminated | September 05, 2013 | 1011 | 16281 | Not_Computer
Z-0489-2011 | 56775 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Vent Needle used on the CELL-DYN Sapphire and CELL-DYN 4000 Hematology Analyzers; CELL-DYN Sapphire Vent Head Assembly, List Number: 08H53-02,CELL-DYN Sapphire Accessory Kit, List Number: 08H03-01,CELL-DYN 4000 Vent Needle, List Number: 02H61-01,CELL-DYN 4000 Accessory Kit, List Number: 01H20-01,CELL-DYN 4000 Vent Needle Upgrade Kit, List Number: 9130576,CELL-DYN 4000 Vent Head Assembly, List Number: 8921187101,CELL-DYN 4000 Sample Processor Assembly, List Number: 8921198201;Product is manufactured and distributed by Abbott Diagnostics Division, Santa Clara, CA | CellDyn 4000 and CellDyn Sapphire Vent Needle head assembly | 2 | November 29, 2010 | 2010 | Abbott Diagnostics Division, 5440 Patrick Henry Drive, Santa Clara, CA | The Vent Needle, which is built into the Vent Head Assembly on the CELL-DYN Sapphire Instrument and which is used directly on the CELL-DYN 4000 Instrument, may be defective, which may result in an undetected short sample, generating incorrect low results, or which may bend, preventing aspiration of the sample. | PRODUCTION CONTROLS: Process Control | Abbott Laboratories issued a Product Recall letter dated July 23, 2010 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were instructed to discard all affected vent needles and replace them with the needle included with the recall letter. Customers may contact customer support at 1-877-422-2688 concerning this recall. |  |  |  | Worldwide Distribution: Thoroughout the US, and in the countries of Argentina, Canada, Germany, Hong Kong, Japan, and Singapore. | Terminated | September 01, 2011 | 276 |  | Not_Computer
Z-0490-2011 | 57002 | kit, identification, enterobacteriaceae | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH Culture Identification Kit for the diagnosis of E. coli and/or P. aeruginosa bacteremiaCatalog Number: KT007 | AdvanDx E. coli/P. aeruginosa PNA FISH Culture Identification Kit | 2 | November 29, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK/P. aeruginosa PNA FISH to produce a false positive red signal. | DESIGN: Device Design | AdvanDx issued a customer letter dated October 6, 2010 with a PDF file of a new package insert and an acknowledgement form to fill out and return to AdvanDx. Customers were instructed to replace the package insert in any existing inventory with the enclosed package insert.Customers can contact AdvanDx technical support about this action at 1-866-376-0009. |  | 260 | 260 units | Worldwide Distribuition: Throughtout the US, and to the countries of Denmark, Italy, Sweden, and UK. | Terminated | June 06, 2011 | 189 | 260 | Not_Computer
Z-0493-2011 | 55119 | instrument, compression | HWN | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Reconstruction System Tenaculum Left Tine Straight 9 inch, Item # 00-1179-021-00, Zimmer Inc., Warsaw, IN.Intended to temporarily reduce pelvic fractures for fixation during reconstructive surgery. | Zimmer Reconstruction System Tenaculum Left Tine Straight 9 inch | 2 | December 01, 2010 | 2010 | Zimmer Inc. | The firm determined that the recalled instruments were manufactured using the wrong grade of material. As manufactured, the instruments are more brittle, increasing the potential for fracture. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Zimmer sent an Urgent Device Recall letter dated February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to monitor all potentially affected patients through regular office consultations. Ensure Zimmer Reconstruction System Tenaculum System Instrument Users at their facility were informed of the urgent device recall.Locate all instruments and quarantine them immediately.Count all affected product in their territory and record the data on theInventory Return Certification Form included with the letter.Fax a copy of the completed Inventory Return Form to:Zimmer, Inc. at (574) 372-4265.Return the recalled product along with the original completed Inventory Return Certification Form to:Zimmer Distribution CenterAttn: Product Service1777 West Center StreetWarsaw, IN 46580For questions customers were instructed to contact their Zimmer Sales Representative.For questions regarding this recall call 1 - (800) 613-6131. |  | 217 | 217 all products | Worldwide Distribution - USA including AL, FL, HI, IA, IL, KS KY, LA, MI, MO, MS, NC, ND, NJ, OH, OK, OR, PA, TN, TX, UT, VA, WA, AND WI and the countries of Australia, Canada, Central or South America, China, Germany, India, Italy, Saudi Arabia, and UK, | Terminated | October 05, 2011 | 308 | 217 | Not_Computer
Z-0502-2011 | 56892 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System, Micro-Infusion Manifold; and Interlink Manifold. An Rx sterile device with 3 injection sites, volume per injection site 0.1 mL, priming volume 0.8 mL, length 8.7 cm. Product code 2N3410 and A2N3410 - JapanBaxter Healthcare corporation, Deerfield, IL 60015 USA; Intended for use with a vascular access device for blood sampling and the administration of multiple drugs and solutions. | Baxter Interlink System MicroInfusion Manifold | 2 | December 01, 2010 | 2010 | Baxter Healthcare Corp. | If the manifolds are connected in series, there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. | OTHER/UNDETERMINED: Pending | Baxter sent an "Urgent Device Correction" letter dated October 20, 2010, via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, informing them that there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. The loosening or disconnection can result in an interruption of therapy, potentially resulting in patient injury and death. The accounts were requested to place the notification where the inventory of affected product is dispensed, so that users are made aware of the procedures to follow: "If these devices are connected in series, and the luer connector is tightened, do not twist the manifolds. This may lead to loosening or disconnection, which may cause a leak. Additionally, Baxter recommends checking all connections before starting and during infusion therapy." The accounts were requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. | Contact the recalling firm for information | 752965 | 752,965 units in US and all lots in Japan | Worldwide Distribution -- United States and Japan. | Terminated | October 16, 2011 | 319 | 752965 | Not_Computer
Z-0491-2011 | 56997 | kit, identification, enterobacteriaceae | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx EK/P. aeruginosa PNA FISH- Culture Identification Kit for identification of E. coli and/or K. pneumoniae and P. aeruginosa bacteremia (In-Vitro Diagnostic)Catalog number: KT008EK/P. aeruginosa PNA FISH is a multicolor, qualitative nucleic acid hybridization assay intended for identification of Escherichia coli and/or Klebsiella pneumoniae and Pseudomonas aeruginosa on smears made from positive blood cultures containing Gram-negative rods. | AdvanDx EK/P. aeruginosa PNA FISH | 2 | December 01, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK /P. aeruginosa PNA FISH to produce a false positive red signal | DESIGN: Device Design | AdvanDX sent an e-mail notification along with a PDF file to all customers on October 8, 2010, followed by written notification dated October 6, 2010. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to replace the package insert(s) included with any kits in their inventory with the enclosed package insert(s). Customers were also instructed to fill out the enclosed Acknowledgement of Receipt Form and return it to AdvanDx to confirm that they received the new package insert. For questions regarding this recall contact AdvanDx Technical Support at 1-888-376-0009. |  | 27 | 27 kits | Worldwide Distribution - USA including IL, MI, NY, and PA and the country of Denmark | Terminated | June 06, 2011 | 187 | 27 | Not_Computer
Z-0504-2011 | 57102 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | VACLOCK Syringe, 20ml, REF: VAC120E, Sterile EO, Single Use, Merit Medical Systems, Inc., South Jordan, Utah 84095. | VACLOCK Syringe | 2 | December 02, 2010 | 2010 | Merit Medical Systems, Inc. | Packaging for syringes may have holes which would compromise the sterility of the package contents. | PRODUCTION CONTROLS: Packaging Process Control | Merit Medical notified product managers of the product recall on October 22, 2010 following a voice mail script. Frequently Asked Questions, Sales Rep Letter, and Recall Notification Forms were then sent.. Sales reps were instructed to contact all of their customers and have them identify and quarantine any unused inventory and either return it to Merit or discard. A response form was to completed by the sales rep and the firm's representative, indicating amount of product quarantined, returned , or destroyed.Also, on 10/22/2010, authorized distributors were contacted by phone, mail, and / or fax. They were also provided a Distributor Letter, FAQs, and Recall Notification Forms instructing them to identify and quarantine any unused inventory and return it to Merit or discard. They were instructed to continue the recall to customers to whom they had shipped the affected Merit product.Additional information is available from Merit at 801-316-4822. |  | 789 | 789 units | Worldwide Distribution: USA including the states of AL, CA, GA, OH, PA, TX, UT, and WA, and to the countires of Japan and UK. | Terminated | January 18, 2011 | 47 | 960 | Not_Computer
Z-0397-2011 | 56048 | reagents,specific,analyte | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD Bioscience CD45 APC-H7 Analyte Specific Reagent, manufactured by BD Biosciences, San Jose, CA.CD45 recognizes members of the T200 family of human leucocyte antigens with molecular mass of 180 to 220 kilodaltons (kDa). | CD45 APCH7 | 2 | December 02, 2010 | 2010 | BD Biosciences, Systems & Reagents | Incorrect expiration date: product with an expiration date of 03/31/2010 was marked with an incorrect expiration date of 06/30/2010. | PRODUCTION CONTROLS: Labeling Mix-Ups | A Product Field Action Notification letter, dated November 2009, was sent to customers who purchased the affected device. The letter identified the affected device and explained the reason for recall. Customers were asked to not use the current CD45 APC H-7 lot number 58633 past 2010-03-31 until they receive their new lot of CD45 APC H-7 at no cost to them. Customers were also asked to complete and return the attached fax form. Questions should be directed to BD Customer Support Center at 1-877-232-8995. | Contact the recalling firm for information | 87 | 87 units | Nationwide Distribution -- CA, WA, NJ, UT, TX and OH. | Terminated | December 03, 2010 | 1 | 87 | Not_Computer
Z-0533-2011 | 57178 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer, NexGen Complete Knee Solution LPS Flex Gender Femoral Component, Size G Left, sterile, REF 00-5764-017-51. | Zimmer NexGen LPS Flex Gender | 2 | December 02, 2010 | 2010 | Zimmer Inc. | These lots of Zimmer NexGen Complete Knee Solution LPS Femoral Component and Zimmer NexGen Complete Knee Solution LPS Flex Gender Femoral Component exhibit a nonconforming internal CAM radius. The condition is intermittent and not all devices have nonconforming geometry. | DESIGN: Process Design | The firm sent URGENT DEVICE CORRECTION AND REMOVAL notices to distributors and sales staff, Risk Managers, and implanting surgeons dated 9/15/2010. Sales staff and risk managers were instructed to return any product on hand. Surgeons were informed of the issue but there was no action to be taken by surgeons. | No consumer action necessary | 158 | 158 all sizes of 5764 | US, Germany, Spain, France, UK, Italy, Sweden, South Africa, Czech Republic, Bulgaria, Poland, Romania, and Austria. | Terminated | July 27, 2011 | 237 | 158 | Not_Computer
Z-0537-2011 | 56818 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XELERIS (1.0, 1.1, 2.0 & 2.1) PROCESSING AND REVIEW WORKSTATIONThe display, processing, archiving, and communication of data acquired by Emission Tomography cameras used in diagnostic radiology, including procedures for planar imaging, whole body imaging, tomographic (SPECT) imaging, positron imaging by coincidence, attenuation correction, and anatomical image registration. | XELERIS (1.0, 1.1, 2.0 & 2.1) PROCESSING AND REVIEW WORKSTATION | 2 | December 03, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate user protocol setup associated with the Thyroid Uptake Index application on certain Xeleris workstations(1.0, 1.1, 2.0, 2.1) that may impact patient safety.Inaccurate protocol setup by the user while using Thyroid Uptake Index application has been identifiedon certain Xeleris workstations. Incorrect protocol setup during clinical data acquisition a | OTHER/UNDETERMINED: Pending | GE Healthcare sent an "Urgent Medical Device Correction" letter dated October 27, 2010, to all consignees. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists and Identifed the product, the problem and the action to be taken by the consignees. Consignees were instructed to replace the original protocol instructions provided with their system with the updated protocol instructions provided with the letter. For any questions consignees were to contact their local GE Healthcare Service Representative or contact the GE Healthcare Call Center:United States or Canada (800) 437-1171Other countries: (262) 896-2890. | Contact the recalling firm for information | 6020 | 6020 | All States in continental USA including DC and PR. Viet Nam, Venezuela, Uruguay, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Sweden, Spain, South Africa, Slovakia, Singapore, Saudi Arabia, Russian Federation, Republic of Korea, Portugal, Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Martinique, Malaysia, Luxembourg, Kuwait, Kenya, Japan, Jamaica, Italy, Israel, Ireland, Indonesia, India, Hungary, Hong Kong, Guadeloupe, Greece, Germany, France, Finland, Ecuador, Dominican Republic, Cyprus, Colombia, China, Canada, Bulgaria, Brazil, Bosnia & Herzegovina, Belgium, Azerbaijan, Austria, Australia, Argentina, Algeria. | Terminated | August 28, 2012 | 634 | 6020 | Not_Computer
Z-0538-2011 | 57111 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta VBH Head FIXThe VBH HeadFIX is intended for positioning and immobilization of the head and neck, stereotactic diagnostic localization and stereotactic radiotherapy of cranial targets. | Elekta VBH Head FIX | 2 | December 03, 2010 | 2010 | Elekta, Inc. | The HeadFIX Baseplate is not screwed down to the adapter to allow to compensate for roll and pitch with the HeadFIX leveling screws of the HeadFIX Baseplate. | OTHER/UNDETERMINED: Pending | Elekta, Inc. sent an Important Notice A343 dated October 14, 2010 to all affected consignees. The Notice identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed constantly monitor the patient before and during the treatment. Check the patient's position after every patient movement (e.g. coughing). Advise all appropriate personnel of the content of the notice place a copy of the Notice In the Important Notice Section of the HeadFix User Manual P10103-108.For questions regarding this recallservice.schwabmuenchen@elekta.com. |  | 34 | 34 units | Worldwide Distribution - USA including AZ, CA, CT, FL, GA, IL, KY, MI, MS, NY, NC, OK, OR, RI, TN, TX, VA, WA and the country of Canada | Terminated | December 28, 2010 | 25 | 34 | Not_Computer
Z-0540-2011 | 56047 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Autoimmune EIA ANA Screening Test Kit, 576 Tests,Model/Catalog Number: 576AN manufactured by Bio-Rad, Hercules, CA. | Autoimmune EIA ANA Screening Test Kit, Model 576AN | 2 | December 06, 2010 | 2010 | Bio-Rad Laboratories Inc | The ANA Screening kit generates a high percentage of false positive results, associated with a low check control reading. | OTHER/UNDETERMINED: Pending | Bio-Rad issued an Urgent Medical Device Recall Notification letter dated March 1, 2010 to customers describing the issue and actions to be taken. Responses will be tracked via response forms. The letter instructed customers that affected products should be identified, quarantined, and discarded, or reworked with shorter expiration dates until newer lots are available.Bio-Rad can be contacted at 510-724-7000 concerning this recall. | Contact the recalling firm for information |  |  | Worldwide Distribution: Throughout the US, and to the countries of: Australia, Canada, France, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and the UK. | Terminated | January 05, 2011 | 30 |  | Not_Computer
Z-0544-2011 | 57069 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva with FD.The Philips MultiDiagnost Eleva intended use is for the following applications: As a diagnostic-imaging device multi-functional/ universal system, general R/F, Fluoroscopy, Radiography and Angiography can be performed along with pediatric examinations and some more specialized interventional applications. | MultiDiagnost Eleva with FD | 2 | December 07, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result, the complete system may shut down and not re-start. | OTHER/UNDETERMINED: Pending | Philips began mailing out an URGENT-Device Correction letter on May 26, 2009. Philips sent out a second URGENT-Device Correction letter dated June 4, 2009, to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to contact their local Philips representative immediately if their system completely shut down and would not restart. A mandatory Field Change Order 72200144 was issued for replacement of the Printed Circuit Board (PCB) in the Allura XPER FD system. A mandatory Field Change Order FCO 70800091 was issued for replacement of the Printed Circuit Board (PCB) in the MultiDiagnost Eleva with FD system. The consignees were to be contacted by their Philips Representative for implementation of the corrective action. The firm's Field Engineer would replace the power tray of the systems.If the consignees needed any further information or support concerning this issue, they were to contact the call Center at 800-722-9377, # 5 and reference FCO 72200144 or FCO 70800091 with any questions. |  | 5 | 5 units | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KY, LA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI. | Terminated | December 09, 2010 | 2 | 50 | Hardware
Z-0542-2011 | 57143 | adaptor, stopcock, manifold, fitting, cardiopulmonary bypass | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(k) | Edwards Protection Cannulae - Monitoring and Infusion Set, Model ISP3147, Sterile EO, Edwards Lifesciences LLC, One Edwards Way, Irvine, CA 92614 | Edwards Protection Cannulae Monitoring and Infusion Set | 2 | December 07, 2010 | 2010 | Edwards Lifesciences Llc | Monitoring and Infusion Set connectors may break during manipulation. | PRODUCTION CONTROLS: Process Control | Edwards Lifesciences issued a Urgent Product Recall letter dated October 13, 2010 to customers identifying the affected devices, requesting the return of unused devices and completion and return of a response form.Further information is available at 800-424-3278. | Contact the recalling firm for information | 322 | 322 sets | Nationwide Distribution: in the states of FL, KS, MO, TX, VA, and PR. | Terminated | May 18, 2011 | 162 | 892 | Not_Computer
Z-0562-2011 | 57241 | generator, pulsatile flow, cardiopulmonary bypass | JOR | Cardiopulmonary bypass pulsatile flow generator. | Cardiovascular | Cardiovascular | PMA | Sarns 8000 Roller pump for Sarns System 8000 Heart Lung Machine console, 115V, Catalog No 16402, Terumo Cardiovascular System Corp Ann Arbor, MI. The system 8000 is indicated for use in extracorporeal circulation of blood for arterial perfusion, reginal perfusion and cardiopulmonary bypass procedures. | Sarns 8000 Roller pump | 2 | December 08, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The system 8000 roller pumps may fail resulting in a persistent pump stop. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm sent "URGENT MEDICAL DEVICE CORRECTION" dated 11/02/2010 to customers. The letter describes the product, problem and corrective actions taken and/or needed to be taken by the customers. The letter states the system responds to a failed inductor by displaying an overspeed message and stopping the roller pump without an audible alert. The pump cannot be restarted. The customers/users are instructed to stop the pump, replace the pump with a backup or if necessary use the handcrank. The letter states that Terumo will replace the roller pump at the next preventative maintenance visit - or sooner- if the customer has a service contract. If the customer does not have a service contract, Terumo will contact the customer within 2 weeks to set up an appointment for the service. Terumo CVS expects all affected inductors to be replaced to be replaced within three months. Call Terumo CVS Customer Service upon completion of the case at 800-521-2818. The customer was also instructed to complete and return the enclosed CUSTOMER RESPONSE FORMPlease note that we are not recommending that you stop using the roller pumps for your Sarns system 8000.For any questions, call toll free 800-262-3304 or phone 734-663-4145. |  | 95 | 95 roller pumps. | Worldwide distribution: USA including states of CA, FL, NC and NY and countries of United Arab Emirates, China, Ecuador, and Panama. | Terminated | October 19, 2011 | 315 | 95 | Not_Computer
Z-0561-2011 | 57064 | absorbent, carbon-dioxide | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Grace SODASORB LF Carbon Dioxide Absorbent, Medical Grade (16%) Soda Lime USP-NF with an ethyl violet indicator which turns from white to violet as absorbent is depleted or upon moisture loss; Manufactured by: W.R. Grace & Co., 6050 West 51st St., Chicago, IL 60638; the product is packaged in pre-paks pre-filled 2.5 lb. canisters, 12 per carton Sodasorb LF is used for absorption of carbon dioxide in low flow anesthesia rebreathing systems and respiratory therapy equipment with oxygen, nitrous oxide, halothane, enflurane, desflurane, sevoflurane and isoflurane. | SODASORB LF PrePaks | 2 | December 08, 2010 | 2010 | Darex Container Products Div of W.R. Grace & Co. | Pre-Pak units of SODASORB CO2 absorbent manufactured prior to August 2010 may have missing or illegible expiration dates and/or batch numbers. | PRODUCTION CONTROLS: Packaging Process Control | Darex/W.R. Grace sent a Customer Notification to their sole direct account of the recall and the account provided Darex/W.R. Grace with lists of their customers who received the recalled products. Darex then sent undated recall letters to those accounts on 10/20-22/10. The accounts were informed of the product, the problem, and the action that should be taken. The accounts were warned not to use Pre-Pak units with missing or illegible expiration dates and/or batch numbers and were requested to examine their inventories of Pre-Pak units, placing all units that had missing or illegible expiration dates and/or batch numbers on hold for destruction and replacement. The accounts were instructed to call W.R. Grace at 1-800-492-7632 to arrange for replacement. The accounts were also requested to complete and return the reply postcard, indicating how much product was destroyed. | Contact the recalling firm for information | 8096 | 8,096 cartons | Nationwide Distribution | Terminated | September 21, 2011 | 287 | 103638 | Not_Computer
Z-0557-2011 | 57176 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Custom Kit, Sterile EO, Rx Only, Catalog No. K09-09815, Manufacturer: Merit Medical Systems, Inc. South Jordan, Utah 84095, USA, | Merit Custom Kits | 2 | December 08, 2010 | 2010 | Merit Medical Systems, Inc. | Defective custom kits may draw in air during use which may compromise patient safety and affect treatment. The valve's fixed luer may be over- or under-tightened, causing leaking at the bond joint. | DESIGN: Process Design | Merit issued an Urgent Product Recall Notice letter dated October 25, 2010 to consignees. Copnsignees were instructed to quarantine the product immediately for return and to notify customers of recall activity and instruct them to return product.. All distributors and customers are to complete and return a Recall Notification Form.For further information about this recall, contact Merit at 800 356-3748. | Contact the recalling firm for information | 1033 | 1033 units | Nationwide Distribution in the states of CA, IL, IN, and UT. | Terminated | March 23, 2012 | 471 | 6943 | Not_Computer
Z-0584-2011 | 56250 | radioimmunoassay, parathyroid hormone | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH ELISA, Part Number: DSL-10-8000The DSL-10-80000 ACTIVE I-PTH Enzyme-Linked Immunosorbent Assay (ELISA) Kit provides materials for the quantitative measurement of I-PTH in serum or plasma. This assay is intended for in vitro diagnostic use as an aid in the diagnosis of hyperparathyroidism. | Active I PTH ELISA | 2 | December 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays may not be achieved for the lot numbers affected by the recall. The affected I-PTH ELISA and the I-PTH IRMA kit lots may produce falsely elevated patient results. Internal stability testing showed an average 16% elevation for I-PTH ELISA lot 99 |  | Beckman Coulter sent an Urgent Product Corrective Action Letter to their consignees dated June 18, 2010. The letter identified the product, the problem, and the action to be taken. Consignees were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372.(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if their laboratory uses only the Control materials provided with the I-PTH ELISA and IRMA kits.(3) Abnormal elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinical diagnostic tests.(4) I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372 will no longer be shipped. Due to the calibrator instability recently encountered with DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays, Beckman Coulter has decided to immediately discontinue these two part numbers.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that they received this important notification.If they have any unused product the were asked to contact the numbers listed below:(1) In the United States, contact Customer Service at 1-800-526-382.1(2) Outside of the United States, consignees were asked to contact their local Beckman Coulter RepresentativeConsignees were told to contact DSL Technical Support Center at 1-800-231-7970 in the United States, or their local Beckman Coulter Representative with any questions.For questions regarding this recall call (714) 961-3634. | Contact the recalling firm for information | 3 | 3 kits | Worldwide Distribution - USA including MD, NY, GA, NJ, CA, MI, TX, PA, IA, and IN. Internationally to Austria, Brazil, China, Colombia, Czech Republic, France, Germany, Hong Kong, India, Japan, Mexico, and the Russian Federation | Terminated | May 07, 2012 | 515 | 199 | Not_Computer
Z-0572-2011 | 56731 | prosthesis, shoulder, non-constrained, metal/polymer cemented | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BIGLIANI/FLATOWÂŽ THE COMPLETE SHOULDER SOLUTION INTRAMEDULLARY REAMER used for implantation of the ZimmerÂŽ Trabecular Metal Humeral and/or ZimmerÂŽ Trabecular Metal Reverse Shoulder Systems: Catalog Number 00-4301-006/018-00. The intramedullary reamer manual surgical instrument is used in the surgical procedure to implant the ZimmerÂŽ Trabecular Metal Humeral and ZimmerÂŽ Trabecular Metal Reverse Shoulder Systems. The reamer is used to open the mid to distal humeral canal to accept the implant. The Zimmer Trabecular Metal Shoulder System consists of conventional and reverse, semi- and nonconstrained (resurfacing) shoulder prostheses for total or hemi-arthroplasty applications. The Zimmer Trabecular Metal Reverse Shoulder System is a treatment option for patients with severe rotator cuff (RC) deficiency. The stem was designed so that its bearing surfaces, as compared to a traditional humeral stem, is such that the humeral component contains a concave ultra-high molecular weight polyethylene (UHMWPE) liner that articulates with a hemispheric metallic glenoid component. These two systems are used for: - the treatment of severe pain or significant disability in degenerative, - rheumatoid, or traumatic disease of the glenohumeral joint; - ununited humeral head fractures of long duration; - irreducible 3-and 4-part proximal humeral fractures; - avascular necrosis of the humeral head, or other difficult clinical management problems where arthrodesis or resectional arthroplasty is not acceptable. is used to open the mid to distal humeral canal to accept the implant. | BIGLIANI/FLATOW THE COMPLETE SHOULDER SOLUTION INTRAMEDULLARY REAMER | 2 | December 09, 2010 | 2010 | Zimmer Inc. | The results of an investigation into complaints revealed that Trabecular Metal Humeral and/or ZimmerÂż Trabecular Metal Reverse Humeral implants engaging in the canal earlier than expected could lead to incomplete seating of the implant, delay in surgery, humeral fracture, or the need for cement in a press-fit application. | DESIGN: Device Design | Zimmer, Inc. sent an URGENT DEVICE CORRECTION letter dated September 10, 2010 to surgeons, risk managers, and distributors. The letter identified the product, the problem, and the action to be taken. Sales Representatives were instructed to:a. provide the Dear Surgeon (Urgent Device Correction) letter and revisedsurgical technique directly to surgeons and risk managers at facilities usingthe device orb. provide a list of facilities and surgeons using the TM Humeral and TMReverse shoulder systems to the Zimmer Recall Coordinator, andc. destroy previous versions of the surgical technique, andd. inform users of when the new surgical technique, instrument case, distalreamers, and distal pilots will be shipped to them based on communicationfrom Zimmer, Inc.Handling of returned product:HCPs were notified to return all of the distal pilots in the scope back toZimmer which were to be be destroyed once the recall has been completed. They willalso be instructed to either return the reamers back to Zimmer or to remove these reamers from the TM Reverse shoulder systems sets and place into theBigliani/Flatow Complete Shoulder Solution set. New surgical techniques will also be introduced and the user will be asked to remove previous versions of the surgical technique from all points of use. This removal will be verified byZimmer Sales Representatives.For questions regarding this recall call (574) 371-8852. |  | 13093 | 13,093 total reamers in distribution. | Worldwide Distribution - USA, Nationwide, Switzerland and Eschbach, Germany, Switzerland, Australia, China, India, Japan, Korea, Malaysia, Singapore, Taiwan, Brazil, Venezuela, El Salvador, Colombia, Canada, and Thailand. | Terminated | September 14, 2012 | 645 | 13093 | Not_Computer
Z-0577-2011 | 57157 | automated urinalysis system | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Teco Diagnostics Uritek TC-101 Urine Analyzer, Catalog Number: TC-101 The Uritek TC-101 (TC-101) urine chemistry analyzer is a semi-automated analyzer for urine chemistry analysis. This device has been manufactured and developed with advanced scientific technology of modern microprocessors, optics and electronics. The apparatus uses advanced testing technology of a cold-light source, integrating sphere, and modularized blocking systematical design which enhances the stability of the optical system. This system also automatically compensates for the influence of ambient and scattered lights through the coordination of software and hardware. The TC-101 can test for 10 parameters including Glucose (GLU), Urobilinogen (URO), Bilirubin (BIL), Ketone (KET), Specific Gravity (SG), Occult Blood (BLD), pH, Protein (PRO), Nitrite (NIT), Leukocytes (LEU) in 60 seconds when used with Urine Reagent Strips (URS) 10. | Teco Diagnostics Uritek TC101 Urine Analyzer | 2 | December 09, 2010 | 2010 | Teco Diagnostics | This recall has been initiated due to the CLIA-101 Urine Analyzer not having 510(k) clearance by the FDA for use in measuring urinary glucose or occult blood with the CLIA waived, Inc. CLIA-10 test strips. Teco Diagnostics has been notified bythe FDA that the CLIA-101 requires 510(k) clearance and is not CLIA Waived for use with the CLIA-10 test strips. Without FDA approval or clearance, Teco Di | OTHER/UNDETERMINED: Pending | Teco Diagnostics initiated an "Urgent: Medical Device Recall" letter dated October 14, 2010, with an attached "Recall Return Response Form" to the affected consignees via US Postal on October 15, 2010. The recall notice informed the consignees of the reason for recall, the products affected and the actions required.Consignees were instructed to complete and return the enclosed Recall Return Response Form as soon as possible via mail, fax, or email. Upon receipt of the Recall Return ResponseForm, Teco Diagnostics would work with the consignees/distributors to process any refunds or reimbursements for the returned product and return shipping. If consignees/distributor had questions, they were to call Teco's Technical Support Department at (800) 222-9880 or (714) 463-1111. |  | 2140 | 2,140 units total for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 03, 2012 | 603 | 2140 | Not_Computer
Z-0587-2011 | 57089 | stopcock, i.v. set | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Three-Way Large Bore Stopcock with Male Luer Slip; a sterile, nonpyrogenic fluid pathway; Baxter Healthcare Corporation, Deerfield, IL 60015; product code 2C6202. Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. | Baxter ThreeWay Large Bore Stopcock | 2 | December 09, 2010 | 2010 | Baxter Healthcare Corp. | Rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. | DESIGN: Device Design | Safety Alert letters dated October 28, 2010 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. The accounts were requested to post the enclosed attachment, which illustrates the correct direction to turn the handle of the stopcock, where the stopcocks/sets are dispensed. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at (847) 270-5457. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 10/27/10, and instructed to notify their customers and ministries of health of the recall.Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  |  |  | Worldwide distribution: USA and countries including Puerto Rico, Guam and American Samoa, and internationally to Canada, Bermuda, Jamaica, Guatemala, Panama, Bahrain, United Kingdom, Hungary, Lebanon, Pakistan, Saudi Arabia and New Zealand. | Terminated | November 01, 2012 | 693 |  | Not_Computer
Z-0565-2011 | 57315 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | MONOJECT Prefill Heparin Lock Flush Syringe10 Unit Heparin, 3 ml in 12 ml syringeItem code 8881580123 | MONOJECT Prefill Heparin Lock Flush Syringe | 2 | December 09, 2010 | 2010 | Covidien LP | The heparin sodium USP used to manufacture was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | COVIDEN sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and in-use stock to determine if they had any units of the affected product and quarantine the product immediately. Distributors were instructed to notify their customers immediately of the URGENT PRODUCT RECALL letter. Customers were to fill out the enclosed response form and return it by fax to (203) 822-6009 or email to adfeedback@coviden.com following the instructions on the form. For questions customers were to contact Customer Service at 1-800-962-9888. |  | 329520 | 329,520 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, LA, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and Bermuda | Open |  |  | 3510660 | Not_Computer
Z-0551-2011 | 57045 | oximeter | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Handheld Pulse Oximetry Model 512 Catalog Number: 9100-00 | Dixtal Handheld Pulse Oximetry | 2 | December 09, 2010 | 2010 | Dixtal Medical, Inc. | Unit may overestimate Sp02 levels in a high humidity environment. | PRODUCTION CONTROLS: Equipment Maintenance | Dixtal Medical notified accounts by email on 9/27/10 and followed by hard copy Urgent Product Recall letter, dated September 27, 2010. The letter identified the affected product and stated the reason for the recall. The letter also requested that customers to cease distribution and/or use of the suspect product and to return product for replacement. The attached Confirmation Notification should be completed and returned. Customers are to contact Customer Service: sales@dixtal.com, Phone 1-203-269-1112, Fax 1-203-269-1760, with questions or to arrange to have a replacement product provided. |  | 4 | 4 units | Worldwide Distribution -- USA, including states of AK, KY, NJ, and NY and countries of Australia, Austria, Bolivia, Canada, Finland, Indonesia, Italy, Poland, Sweden, and Switzerland. | Open |  |  | 39 | Not_Computer
Z-0604-2011 | 54143 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spinal Concepts polyaxial screwdriver II, Manufactured by Spinal Concepts, Austin, TX, REF 781-2. | Spinal Concepts polyaxial screwdriver | 2 | December 10, 2010 | 2010 | Zimmer Inc. | The devices may fracture, resulting in surgical delay or the need for intervention to remove fragments from the surgical site. | DESIGN: Device Design | The affected instruments were removed and replaced in the U.S. Abbott Spine and Spinal Concepts instrument kits between December 2005 and October 2007. User accounts were not advised of the removal Zimmer has confirmed that none of these instruments are currently available for use in the U.S.For further information about this action, contact Dale Miller of Zimmer at 574 372-4962. | No consumer action necessary | 27 | 27 | Worldwide Distribution, including US. | Terminated | June 01, 2011 | 173 | 493 | Not_Computer
Z-0589-2011 | 55977 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | GE Responder 2000 Defibrillator/Monitor, models 2025653, 2026109, and 2026114 that were manufactured between July 25, 2008 and October 2, 2009 by GE Healthcare. | GE Responder 2000 Defibrillator/Monitor | 2 | December 10, 2010 | 2010 | Cardiac Science Corporation | Some units may be non-operational during a rescue, due to a defective PCB interconnect on traces connected to the reed relay. This will result in an inability to analyze the patient's heart rhythm or deliver therapy during a rescue. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cardiac Science issued a Urgent Medical Device Corrective Action letter dated April 28, 2010 requesting GE Healthcare cease distribution of affected devices and to instruct end users of distributed devices to schedule a service call to have the device corrected. Cardiac will send rework kits no later than June 30, 2010. CSC can be contacted at 44 161 926 0011 to arrange return material authorization and replacement of affected devices. | Contact the recalling firm for information | 130 | 130 units | Worldwide Distribution: Throughout the US, and in the countries of Germany, Korea, Peru, Singapore, and Spain. | Terminated | December 08, 2011 | 363 | 130 | Hardware | Treatment/Delivery/Therapy | Replace | Remove or Replace
Z-0594-2011 | 56832 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FUSO Analyzer 300FAbbott Point of Care, Princeton, NJ 08540Model 200. the 37 degree 37 degree C tests for sodium, potassium, ionized calcium, hematocrit, pH, PCO2, PO2 and the i-stat portable analyzer as part of the i-STAT system, are intended for use in the in vitro quantization of arterial, venous or capillary whole blood. | iSTAT Analyzer FUSO Analyzer 300F | 2 | December 10, 2010 | 2010 | Abbott Point Of Care Inc. | i-STAT Analyzers batteries may overheat due to a circuitry failure, resulting in the analyzer becoming uncomfortably hot in the area of the battery component. | OTHER/UNDETERMINED: Pending | Abbott Point of Care issued Urgent Recall Notice letters dated September 20, 2010 to customers. The letter informs users to of the possibility of the unit becoming hot to touch, to remove the unit if this is the case, not to remove batteries unless instructed to do by APOC and to forward notification of the recall to customers if it has been further distributed.Abbott Point of Care can be contacted at 800 366-8020, option 1, about this action.Recall extended in May 2011 to included distribution dates of 9/20/2010-6/6/2011. |  | 2202 | 2202 | Worldwide Distrbution: Throughout USA and the countries of Argentina, Australia, Austria, Belgium.Bolivia, Bosnia-Herz, Canada. Croatia, Chile, Colombia, Costa Rica, Denmark, Ecuador, El Salvador, Finland, Georgian Republic, Greece, Honduras, France, Germany, Italy, Netherlands, New Zealand, Hong Kong, India, Ireland, Israel, Japan, Kingdom of Saudi Arabia, Korea, Kuwait, Mauritius, Mexico, Nigeria, Norway, Panama, Paraguay, Peoples Republic of China, Peru, Poland, Portugal, Qatar, Russia, Singapore, South Africa, Spain, Sweden, UK, Sultanate of Oman, Switzerland, Taiwan, R.O.C., Thailand, Turkey, UAE, Uruguay, and Yemen. | Terminated | January 28, 2013 | 780 | 656504 | Battery
Z-0588-2011 | 55571 | lipase-esterase, enzymatic, photometric, lipase | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension LIP Flex Reagent Cartridge, Catalog# DF55A | Dimension LIP Flex Reagent Cartridge | 2 | December 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Open Well instability: After calibration, QC shifts high (out of range) after the Flex Well is hydrated. Shifts of up to 20% have been observed in both QC and patient samplesas soon as two hours post hydration. | OTHER/UNDETERMINED: Pending | Recalling firm issued "Urgent Field Safety Notice", dated March 2010, informing users of the problem and to transition to the revised Lipasae method. |  | 21171 | 21171 | Nationwide Distribution throughout the US. | Terminated | August 04, 2011 | 237 | 21171 | Not_Computer
Z-0592-2011 | 56864 | drape, surgical | KKX | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Microtek Medical, Inc., Microscope Drape for Zeiss StandardModel/Catalog Number: 70890655Product Description: The Microtek Medical, Inc., Microscope Drape for Zeiss Standard is intended is to protect surgical microscopes from contamination during various procedures.Labeling for the product is as follows: (1.) INNER CARTON LABEL:Distributed by; MICROTEX, MICROSCOPE DRAPE FOR ZEISS STANDARD, 20 x 66 (51cm x 168 cm), REF 70890655, QTY 12, LOT DA102032, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, CE0044, EC, REP, Microtek Medical B.V. Zutphen, The Netherlands, Manufactured in the Dominican Republic of U.S. Materials, STERILE, EO, 2015-07, (01)30748426064206, (17) 150700(10) DA102032, INNER, Rev. C.(2.) PRODUCT LABEL:Distributed by; MICROTEX, MICROSCOPE DRAPE FOR ZEISS STANDARD, 20 x 66 (51cm x 168 cm), REF 70890655, QTY 1, LOT DA102032, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, CE0044, EC, REP, Microtek Medical B.V. Zutphen, The Netherlands, Manufactured in the Dominican Republic of U.S. Materials, STERILE, EO, 2015-07, (01) 00748426064205, (17) 150700(10) DA102032, EACH, Rev. D.(3.) CARTON LABEL:Distributed by; MICROTEX, MICROSCOPE DRAPE FOR ZEISS STANDARD, 20 x 66 (51cm x 168 cm), REF 70890655, QTY 12/CS, LOT DA102032, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, CE0044, EC, REP, Microtek Medical B.V. Zutphen, The Netherlands, Manufactured in the Dominican Republic of U.S. Materials, STERILE, EO, 2015-07, (01)50748426084200, (17) 150700(10) DA102032, INNER, Rev. D. | Microtek Medical, Inc., Microscope Drape for Zeiss Standard | 2 | December 10, 2010 | 2010 | Microtek Medical, Inc. | This product is being recalled due the product is miss-labeled as being sterilized by EO processing; however, the product was processed using Gamma sterilization. | PRODUCTION CONTROLS: Error in Labeling | On 08/23/10 Microtek Medical, Inc., Jacksonville, Florida notified all of their customers via a telephone call and requested immediate return of the product . The firm followed up with each of the customers by sending out an Urgent: Medical Device Recall letter (written notice), dated September 02,2010, to each on 9/8/10. The letter identified the affected product, stated the reason for the recall, and asked that the product be returned immediately. Customers are to complete and return the enclosed response form. Anyone needing further assistance is asked to contact Elizabeth Jolly, V.P. Legal at 1-888-844-0988. | Contact the recalling firm for information | 5 | 5 Cases, 12 drapes per case, 60 Drapes total distrubuted | Nationwide Distribution | Terminated | December 14, 2010 | 4 | 5 | Not_Computer
Z-0601-2011 | 57096 | heparin, vascular access flush | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medefil Heparin I.V. Flush Syringe 1 units/mL; a polypropylene luer locking syringe containing Heparin Lock Flush Solution, USP, a sterile non-pyrogenic, isotonic injection; each mL in a Heparin I.V. Flush Syringe contains 1 units of heparin sodium derived from porcine intestinal mucosa and 9 mg of sodium chloride in water for injection; Medefil, Inc. Glendale Heights, IL 60139. The pre-filled syringes were packed as 60 syringes per dispenser box, 16 boxes per shipping carton in the following syringe configurations: a) NDC# 64253-444-22, catalog # MIH-4422: 2 mL fill in 6 mL syringe; b) NDC# 64253-444-25, catalog # MIH-4425: 5 mL fill in 6 mL syringe | Heparin I.V. Flush Syringe 1 units/mL | 2 | December 10, 2010 | 2010 | Medefil Incorporated | The heparin lock flush solution was manufactured with contaminated Heparin Sodium USP Active Pharmaceutical Ingredient (API). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Medefil faxed/emailed the direct accounts on 11/1/10, and sent the accounts a hard copy of the Urgent Product Recall letter, dated 11/1/2010, on 11/2/10. The accounts were informed of Medefil's recall of 18 lots of Heparin I.V. Flush Syringes produced from the active pharmaceutical ingredient (API) Heparin Sodium, USP, lot 1035-0780, that has been recalled by the manufacturer, Scientific Protein Laboratories, due to the presence of a contaminant, Oversulfated Chondroitin Sulfate. The accounts were instructed to stop use of the 18 lots of Heparin I.V. Flush Syringes immediately, place them under quarantine, and contact Medefil to arrange for the product to be returned. The accounts were requested to complete and return to Medefil the enclosed response form, indicating the quantity on hand of each recalled lot. Any questions were directed to Pradeep Aggarwal, President at 630-682-4600. |  | 339180 | 339,180 syringes | Nationwide Distribution --- including states of Arizona, California, Florida, Illinois, Ohio, Louisiana, Massachusetts, Michigan, Minnesota, New York, Texas, and Virginia. | Terminated | February 13, 2012 | 430 | 2919180 | Not_Computer
Z-0590-2011 | 57011 | bone marrow collection/transfusion kit | LWE |  |  | General Hospital | 510(k) | Orthofix Bone Marrow Aspiration Needle, Catalog number 21-5000, 21-5011, Distributed By: Orthofix Inc, 1720 Bray Central Drive, McKinney, TX 75068-8207. Product includes handle, stylet, cannula, probe, probe guide, and universal cap. Cannula is available as an 8 gauge needle or 11 gauge needle. | Orthofix Bone Marrow Aspiration Needle | 3 | December 10, 2010 | 2010 | Orthofix Inc | Outer cartons for the product may be labeled with incorrect needle gauge. | OTHER/UNDETERMINED: Pending | Orthofix issued an Urgent Product Recall letter dated October 14, 2010 to consignees. Customers were instructed to inspect their inventory, and remove and return affected product to the firm. Replacement needles will be provided. Orthofix account representatives can be contacted at 1 888 298 5700, and Orthofix Regulatory Affairs can be contacted at 1 214 937 2061. | Contact the recalling firm for information | 274 | 274 | Worldwide Distribution - USA, including the states of AL, AR, CA, CT, FL, LA, MI, MO, NC, NJ, NY, PA, SC, and TX, and the country of Spain. | Terminated | February 28, 2011 | 80 | 274 | Not_Computer
Z-0591-2011 | 57216 | lamp, surgical | FTD | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | CHROMOPHARE Surgical Lighting System & CHROMOVIEW Flat Panel Monitor, Mechanical Spring Arms. | Surgical Lighting and Panel Monitor Arms | 2 | December 10, 2010 | 2010 | Berchtold Corp. | A mechanical failure of a welded joint within the spring Arm may result in the rapid drop of a monitor to the mechanical stop. | OTHER/UNDETERMINED: Pending | Berchtold issued a Product Safety Notification letter on November 16, 2009 identifying the affected device and issue prompting the action. Consignees were instructed to:Exercise caution when adjusting the position of the surgical lights or flat panel monitors. While an event is highly unlikely, the likelihood does increase when the spring arm is being repositioned. Identify a specific contact person for their facility to facilitate the communication and field action process with BERCHTOLD. The following action is being planned: To further identify if their units are part of the affected population, determine the supplier serial number identification and specific location for each spring arm in their facility. Report the facility information, supplier spring arm serial numbers and specific location or spring arms to a BERCHTOLD. Detail instructions on serial number identification and the method of reporting to BERCHTOLD will also be provided in a subsequent communication. Once BERCHTOLD has received the serial numbers of their spring arm assembly, customers will be notified whether their particular units are affected and the further required action.Berchtold Corporation can be contacted at 1 800 968 6097 about this action. |  | 7031 | 7031 arms | Worldwide Distribution: Throught the U.S,.and the countries of Canada and Germany | Terminated | July 06, 2012 | 574 | 7031 | Not_Computer
Z-0614-2011 | 51242 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Heartstring II Proximal Seal / VASOVIEW 6 Endoscopic Vessel Harvesting System,Model Number: VH-2000,Product is manufactured and distributed by Maquet Corporation San Jose, CA | Heartstring II Proximal Seal/VASOVIEW 6 Endoscopic Vessel Harvesting System | 2 | December 13, 2010 | 2010 | CTS, Inc. dba Guidant Cardiac Surgery | Sterility may be compromised by a loss of package integrity. | PRODUCTION CONTROLS: Packaging | Maquet issued an Urgent Device Removal letter dated February 5, 2009 to customers, explaining the reason for the recall and requesting the product be returned for replacement. Customers were instructed to examine their inventory and discontinue distribution and use of the affected devices. Customers also were requested to complete and return a Field Action Response Form.Maquet can be contacted at 888-880-2874 concerning this recall. |  |  |  | Distribution: Nationwide and Internationally. | Terminated | January 19, 2011 | 37 |  | Not_Computer
Z-0608-2011 | 55815 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Quest EIA ANA Screen Bulk Kit, model number 96AN-BPU-QUEST. Manufactured by Bio-Rad Laboratories, Inc., Hercules, CA. For labeling and distribution by Quest International Inc.Medical device is used as a qualitative enzyme immunoassay (EIA) intended to screen for the presence of antinuclear antibodies (ANAs) in human serum as an aid in the diagnosis of certain rheumatic diseases. This assay collectively detects, in one well, total ANAs against double stranded DNA (dsDNA, nDNA), histones, SS-A/Ro, SS-B/LA, Sm, Sm/RNP, Scl-70, Jo-1 and centromeric antigens along with sera positive for Immunofluorescent (IFA) HEp-2. | Quest EIA ANA Screen Bulk Kit | 2 | December 13, 2010 | 2010 | Bio-Rad Laboratories Inc | The value for the ANA Cutoff and ANA Positive Controls will decrease overtime resulting in an unusually elevated positivity rate for patient samples tested. |  | Bio-Rad Laboratories, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 24, 2010, to their customer. The letter identified the product, the problem, and the action to be taken. The customer was instructed to:Examine inventory and quarantine affected product. Identify all customers that have received the affected product and immediately notify them. Instruct them to discontinue use of the affected product.Complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM.For questions regarding this recall call (510) 724-7000. |  | 453 | 453 kits | Nationwide Distribution including FL | Terminated | January 05, 2011 | 23 | 453 | Not_Computer
Z-0609-2011 | 55363 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution Stemmed tibial broach impactor, Catalog # 00-5977-011-00, Zimmer Inc., Warsaw, IN. | NexGen Complete Knee Solution Stemmed tibial broach impactor | 2 | December 13, 2010 | 2010 | Zimmer Inc. | The impactor may disassemble or detach during use. The detached piece may not be observed intraoperatively and may be left in the surgical site. Complications may include pain, infection, and prolongation in surgery. These complications may also require revision surgery. | DESIGN: Device Design | Zimmer issued "Urgent Medical Device Safety Alert" letters dated April 16, 2010 to Risk Managers, user facilities, Health Care Professionals and Zimmer Sales force. The users were informed to inspect the device for loosening. The alert also informed users to inspect the device after each use to verify that the device was intact. Customers were instructed to complete and return a User Facility Documentation Form. Zimmer will repair and replace all affected devices.Customers can contact ZImmer, Inc. at 800 613-6131 concerning this recall. |  | 10666 | 10,666 all products. | Worldwide Distribution: Throughout US including Puerto Rico, and the countries of Argentina, Australia, Belgium, Bolivia, Brazil, Canada, Chile, China, Columbia, Denmark, Ecuador, El Salvador, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Korea, Mexico, Netherlands, Nicaragua, Panama, Peru, Portugal, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, and United Kingdom. | Terminated | September 12, 2012 | 639 | 10666 | Not_Computer
Z-0613-2011 | 56386 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad brand Meridian ANA Screen ELISA Test Kit, 96 Test,Catalog No. 4884261, Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA and distributed by Meridian Diagnostics. | BioRad brand Meridian ANA Screen ELISA Test Kit | 2 | December 13, 2010 | 2010 | Bio-Rad Laboratories Inc | False positive test results caused by decrease in the OD values of ANA Cutoff Controls and ANA Positive Controls. The increased positivity rates were associated with the Meridian ANA Screening kit. False positive ANA values could result in misdiagnosis or mistreatment . | OTHER/UNDETERMINED: Pending | Bio-Rad issued an Urgent Medical Device Recall NOtification letter dated February 24, 2010 to their distributor, identifying the affected product and actions to be taken. The distributor was instructed to examine inventory and quarantive affected kit lots, and to identify and notify customers with affected lots to discontinue use and discard the product. Replacement kits will be distributed. Customers were asked to complete and return a recall response form.Bio-Rad can be contacted at 510 724-7000 concerning this recall. |  |  |  | Nationwide Distribution: to Meridian Diagnostics in Ohio. | Terminated | January 10, 2011 | 28 |  | Not_Computer
Z-0607-2011 | 56752 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems, Part Number: 81600 ; Access 2 Immunoassay Systems, Part Number: 81600N;Synchron LXi 725; UniCel DxI 800 Access Immunoassay Systems, Part Number: 973100; UniCel DxI 600 Access Immunoassay Systems, Part Number A30260;UniCel DxC 600i, Part Number: A25656; UniCel DxC 660i, Part Number: A64871; UniCel DxC 860i, Part Number: A64935; UniCel DxC 880i, Part Number: A59102; SYNCHRON Access ClinicalSystemsThe Access/DxI Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitive or qualitative determination of various analyte concentration found in human body fluids. Function of the Peristaltic Pump: to aspirate unbound material from reaction vessels and to move waste from vacuum jar into waste bottle through associated tubing within the Access Immunoassay System. | Access/DxI Immunoassay System | 2 | December 13, 2010 | 2010 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has received reports of erroneous test results being reportedas a result of assay reagent packs being transferred between systems. Erroneous results will be generated as a result of missing or incorrect assay reagents. In these casesboth false positive and false negative results are possible, depending upon the assay configuration. |  | Beckman Coulter, Inc. sent two (2) Product Corrective Action (PCA) letters (Access letter and DxI letter) with attached Customer Response form the week of August 09, 2010 to the affected customers. The letters were dated August 13, 2010, and identified the product, the problem, and the action to be taken by the customers. The letter reminded customers that it is critical that they load reagent packs according to the procedure provided in their Access or Access 2Operators Guide, Instructions for Use, or Help system. A copy of the procedure was included with the letter. Customers were reminded that all operators should be properly trained on these important procedures andunderstand the impact on reported results if these procedures are not followed.Customers were instructed to share this information with their laboratory staff, and retain this notification as part of their laboratory Quality System documentation. Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Supportat 1-800-854-3633, option 1 in the United States or Canada. | Contact the recalling firm for information | 9571 | 9571 units (instruments) | Nationwide Distribution & Canada | Terminated | July 11, 2012 | 576 | 9571 | Not_Computer
Z-0615-2011 | 57214 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Incorrect Label: Endocare 2.4mm VARIABLE ATTACHED CRYOPROBE, REF: CVA 2400, LOT 2751. Correct Label: Endocare 2.4mm CRYOPROBE, REF: CRYO-44F (Sharp Tip). | Endocare 2.4mm VARIABLE ATTACHED CRYOPROBE | 2 | December 13, 2010 | 2010 | Endocare, Division Of Healthtronics | Mislabeled : The label on the outer box and inner pouch of the EndoCare 2.4mm CryoProbe (CRYO-44F) inaccurately reference the CVA2400 (2.4mm Variable Probe aka: V-Probe). |  | Healthronics issued an Urgent Device Recall letter dated November 2, 2010 to consignees. The letter described the product mislabeling, and requested customers to examine their inventory, quarantine all affected product, and complete and return the verfication section of the letter to the firm. A return authorization number and prepaid return label were provide to facilitate return and replacement.Healthronics can be contacted at 877 367-4873 concerning this recall. | Contact the recalling firm for information | 33 | 33 Units | Nationwide Distribution: to the states of FL, MI, PA, TN, and WA. | Open |  |  | 33 | Not_Computer
Z-0611-2011 | 56849 | bed, patient rotation, powered | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoProne Therapy System, Model 209500, Device List Number E232632Product is designed to allow patients to be mechanically rotated to the prone position and provides kinetic therapy to patients. It is intended for the treatment and prevention of pulmonary complications associated with immobility. | RotoProne Therapy System | 2 | December 13, 2010 | 2010 | KCI USA, Inc. | Product complaints indicating that while attempting to resolve an alarm or error condition, operator was unable to return product to the prone position for patient. | OTHER/UNDETERMINED: Pending | KCI USA, Inc. sent an URGENT - Medical Device Safety Alert letter, dated September 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the cusotmer. Customers were instructed to have a practiced back-up plan, such as a plan for the manual proning of patients in the event that clinical circumstances dictate the need for more immediate prone therapy. For questions, customers in the US should call KCI's Technical Support Center at 1-800-275-4524, option 5. For customers outside the US they should contact their local sales representative. Customers should contact their accounts with affected product and advise them of the safety notice. |  | 240 | 240 units | Worldwide Distribution - USA Nationwide, United Kingdom, Germany, and Saudi Arabia | Terminated | March 21, 2011 | 98 | 240 | Not_Computer
Z-0612-2011 | 56126 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad Laboratories Autoimmune EIA ANA Screening kit, 96 test, model 96AN, Distributed and Manufactured by Bio-Rad Laboratories, Inc. | BioRad Laboratories Autoimmune EIA ANA Screening kit, 96 test | 3 | December 13, 2010 | 2010 | Bio-Rad Laboratories Inc | Product mislabeled-- the ANA cutoff Control and Prostiva controls were mislabeled. Use of the incorrect controls will cause run failure - patient results cannot be reported because controls will not pass.. |  | Bio-Rad issued Medical Device Recall Notification letters dated June 24, 2010 to consignees, identifying the affected products and requested that all affected product be disposed of in accordance with local procedures. Response forms are to be completed and returned to the firm. Bio-Rad can be contacted for replacement product and about this recall at 510 741-6114. |  | 37 | 37 units | Nationwide Distribution: including the states of California, Ohio, Massachusetts, and Missouri.Product delivered to 11 consignees only in the US. | Open |  |  | 37 | Not_Computer
Z-0618-2011 | 53182 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Titanium Intraline Anchor, Model Numbers 3910400050, 3910400055, 3910400060, 3910400065, Soft Eyelet Anchor, manufactured by Stryker Endoscopy, Puerto Rico | Stryker Titanium Intraline Anchor | 2 | December 14, 2010 | 2010 | Stryker Endoscopy | Potential breakage -- product may break while inserting it into the bone during surgery. This may prolong surgery, or, if undetected during surgery, result in retained fragments that may cause metal sensitization or immune response, or necessitate further intervention to retrieve the titanium fragments. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Stryker issued an Urgent Device Removal notification letter dated August 21, 2009 to consignees. Consignees were asked to acknowledge receipt and distribute the letter, track reponses, and destroy any returned material and document through weekly summary and records.Stryker can be contacted at 408 754-2000 about this recall. |  | 447 | 447 units | Nationwide Distribution: to one distributor in NJ. | Terminated | December 14, 2010 | 0 | 447 | Not_Computer
Z-0620-2011 | 57370 | display, cathode-ray tube, medical | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Flat Panel Display is a touchscreen display, 19-inch thin film transistor liquid crystal display (TFT-LCD). The PNs are 010-1619-00 and 010-1655-00, Models 94260-19 and 94256. | Spacelabs Medical Flat Panel Display | 2 | December 14, 2010 | 2010 | Spacelabs Healthcare, Llc | The power supply of the Spacelabs Medical Flat Panel Display may fail, causing the display to go blank. | OTHER/UNDETERMINED: Pending | Spacelabs Healthcare sent an URGENT: MEDICAL DEVICE RECALL Spacelabs Medical Flat Panel Power Supply PNs: 010-1619-00 & 010-1655-00 letter dated March 20, 2009 to US consignees. On March 27, 2009, Spacelabs sent the same recall letter with translation option to their international consignees.The letter stated the firm would be sending a replacement power supply. While waiting for the replacement of the power supply, Spacelabs Medical recommended their consignees brief the staff on the possibility of the display going blank to help prevent any adverse event from happening. The consignees were advised to dispose the old power supply per their premise protocol or local regulations. The consignees from US can contact Spacelabs Medical at 1-800-522-7025, #2 for Technical Support , and consignees outside the United States can call 1-425-657-7200 x5089 for any questions about the recall. |  | 2877 | 2877 units in US and 731 units outside US | Worldwide Distribution: USA in the following states: AK, AL, AR, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WY, and including the following countries: AUSTRALIA, AUSTRIA, CANADA, Caribbean, CZECH REPUBLIC, EGYPT, ESTONIA, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, ITALY, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | June 13, 2011 | 181 | 2877 | Battery | Display/Image | Replace power supply | Remove or Replace
Z-0619-2011 | 57304 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Inter-Op Acetabular System Hooded Trial Insert size 32mm/57mm, non-sterile, REF 9363-32-057, Zimmer, Warsaw, IN. | InterOp Acetabular System hooded trial insert | 2 | December 14, 2010 | 2010 | Zimmer Inc. | Mislabeled - Device packages labeled as hooded trial inserts contain standard trial inserts. | TRAINING: Employee Error | Zimmer contacted the two affected accounts directly by phone on November 15, 2010 to inform them of the recall. Zimmer will monitor the return of the affected instruments and account for the outstanding instruments until the removal action is complete. Each of the affected accounts will be provided with a new trial insert, manufactured and packaged with the correct component. The accounts will be asked to return the affected units back to Zimmer where the returned product will be destroyed. Customers can contact Zimmer at 574 371-8852 concerning this recall. |  | 3 | 3 | Worldwide Distribution: US in the states of KS and UT, and the country of Canada. | Terminated | May 23, 2011 | 160 | 3 | Not_Computer
Z-0617-2011 | 57223 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The PercuNav MP008R-B Button Probe is designed to be used with Philips (formerly Traxtal) PercuNav Image Guided Intervention System. Some of the clinical applications of the Button Probe require sterility and others may not. However, it is possible that patient contamination could occur if a non-sterile Button Probe is used in clinical applications where aseptic technique is required. The device's labeled in part: " TRAXTAL PERCUNAV...(1) Button Probe Right hand version, 25mm...Traxtal Inc...49 Spadina Avenue, Suite 310 Toronto, ON Canada M5V 2J1...Made In Canada...www.traxtal.com...STERILE...For use with Traxtal's PercuNav system only. Do not use if package is opened or damaged. Dispose of safely after single use to avoid risk of infection...Rx only...Use immediately after opening unit package. Store at room temperature". | PercuNav MP008RB Button Probe | 2 | December 14, 2010 | 2010 | Philips Healthcare Inc. | The protective cover over the tip of the MP008R-B Button Probe may fall off, causing the probe tip to damage the packaging and compromise the sterility of the probe. | OTHER/UNDETERMINED: Pending | On 11/10/10, Philips issued URGENT - Medical Device Recall notification letters dated November 10, 2010 to their consignees.Consignees are advised to examine their inventory of MP008R-B Button Probes to confirm the probe tip cover is properly placed over the tip of the probe. They should discard any Button Probes that have been confirmed to have the protective probe tip cover removed from the Button Probe tip. Consignees are also reminded to follow the MP008R-B Button Probe and PercuNav Instructions for Use prior to using the device. Do not use the probe if the unit package is damaged, broken or soiled.The firm's representative will contact the consignees to confirm quantities and Lot Codes of MP008R-B Button Probes that the firm may have examined or discarded. The firm estimates replacement probes will be available in the first quarter of 2011.Consignees can contact Philips Customer Service at 800 722-9377 and reference the FCO 79500189. |  | 25 | 25 units | Nationwide Distribution: in the states of FL, IL, MA, NE, NY, SC, and VA. There are no foreign consignees and US government consignees. | Terminated | August 15, 2011 | 244 | 25 | Not_Computer
Z-0621-2011 | 57269 | surgeon's gloves | KGO | Surgeon's glove. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ENCORE (R) Microptic (R) Powder-Free Latex Surgical Gloves, 1 Pair, Size 5ÂŽ, Sterile R, CE 0086, EC REP Ansell (UK) Ltd., Tamworth, Staffordshire, B79 7UL UK, Ansell Lanka (pvt) Ltd., Biyagama Export Processing Zone, Biyagama, Sri LankaIntended use: Surgical Gloves | PowderFree Latex Surgical Gloves | 2 | December 15, 2010 | 2010 | Ansell Healthcare Products LLC | open package seams on sterile gloves | OTHER/UNDETERMINED: Pending | The firm, Ansell, sent a letter dated September 3, 2010 to all customers. The letter described the product, problem and actions taken. An Ansell representatives will contact the customer to remove the gloves bearing the noted lot number from their facility; Ansell will ensure that the gloves removed will be replaced shortly with products from a different manufacturing batch and the recall was extended to the retail (hospital and clinic) level requesting that the product be returned. If you have any questions, contact US (334) 615-2563, international telephone 01.34.24.52.52 and email: infofrance@eu.ansell.com. |  | 4000 | 4000 units | International distribution only: France and Germany | Terminated | November 28, 2011 | 348 | 4000 | Not_Computer
Z-0689-2011 | 55861 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | 3.1 mm Inion CPS, 3.1 x 18 mm screw, REF SCR-1230, sterile, Manufacturer, INION Oy, Finland. | Inion CPS | 2 | December 16, 2010 | 2010 | Stryker Craniomaxillofacial Division | Inadequate Instructions for Use. The stated strength retention is ambiguous and the maximum immersion time of the plates/meshed is not defined. | OTHER/UNDETERMINED: Pending | Stryker issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2010. The notice informed users of the change in the instructions for use (IFU), and included a copy of the revised IFU. The consignees were informed that the affected product does not need to be returned.Stryker can be contacted about this action at 877-534-2464 (option "Regulatory Affairs", reference Product Field Action #2009-445). |  |  |  | Nationwide Distribution: Throughout the US | Terminated | June 06, 2012 | 538 |  | Not_Computer
Z-0624-2011 | 56659 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Cmax Surgical Tables, Model #'s: Cmax110, Cmax 220 manufactured by Steris Corporation in Montgomery, AL | Cmax Surgical Table | 2 | December 16, 2010 | 2010 | Steris Corporation | Cmax Surgical Tables may lose calibration if the table is operated only on battery power and if the table is in a low battery charge condition. This could result in the table moving into an unexpected tilt position when the Level button is depressed on the hand control keypad. | OTHER/UNDETERMINED: Pending | Steris issued Urgent Field Correction Notice letters dated August 26, 2010 to each affected consignee notifying them of the voluntary field correction. A Steris Representative will contact customers to arrange for STERIS to upgrade the software in the table.For further information about this corrective action, contact Steris at 1 800 333-8828. | Contact the recalling firm for information | 1635 | 1,635 tables | Worldwide Distribution: USA in the following states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV, and in the countries of Australia, Canada, China, Columbia, Dominican Republic, India, Israel, Mexico, New Zealand, Saudi Arabia, Singapore, South Korea and Taiwan. | Terminated | September 27, 2012 | 651 | 1635 | Software | Physical Safety Hazards | Software update | Software Update
Z-0627-2011 | 56716 | intracranial neurovascular stent | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Wingspan Stent System, Model Number M003WE0300150, Manufactured by Boston Scientific, Fremont, CAIndicated for use in improving cerebral artery lumen diameter in patients with intracranial atherosclerotic disease, refractory to medical therapy, in intracranial vessels with less than or equal to 50% stenosis that are accessible to the system. | Wingspan Stent System | 2 | December 16, 2010 | 2010 | Boston Scientific Corporation | Products may not meet radial force specifications. | PRODUCTION CONTROLS: Process Control | Recall initiated August 25, 2010. "Urgent Medical Device Recall - Immediate Action Required" notifications, dated 25 August 2010, were delivered by FedEx where possible, and by personal delivery assigned to sales representatives. The letter identified the affected product and the reason for recall. The letter stated that further distribution or use of any remaining product should cease immediately. Customers were asked to read the enclosed instructions and told that their local Sales Rep could answer any questions that they may have regarding the recall. |  | 10 | 10 units | International Distribution Only -- China, Germany, Italy, the Philippines, Portugal and Turkey. | Terminated | December 17, 2010 | 1 | 10 | Not_Computer
Z-0629-2011 | 56757 | calibrators, drug specific | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | ARCHITECT iVancomycin Reagent; an in vitro chemiluminescent microparticle immunoassay (CMIA) for the quantitative measurement of vancomycin in human serum or plasma on the ARCHITECT i System with STAT protocol capability; Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1P30-25The ARCHITECT iVancomycin assay is an in vitro chemiluminescent microparticle immunoassay (CMIA) for the quantitative measurement of vancomycin in human serum or plasma on the ARCHITECT i System with STAT protocol capability. The ARCHITECT i Vancomycin assay is used in the diagnosis and treatment of vancomycin overdose and in monitoring levels of vancomycin to help ensure appropriate therapy. | ARCHITECT iVancomycin Reagent | 2 | December 16, 2010 | 2010 | Abbott Laboratories | The assay may produce falsely elevated vancomycin results. Inaccurate test results may result in inappropriate patient treatment. | OTHER/UNDETERMINED: Pending | Abbott Laboratories issued a Product Correction letters dated September 13 and 15, 2010 to all customers that received ARCHITECT iVancomycin reagents in the last 12 months. In the September 13, 2010 letter, the accounts were informed of the falsely elevated results and were instructed to continue to follow all sample handling instruction in the package insert and to take the following additional actions to mitigate the risk of obtaining an incorrect result:1. After following the primary tube manufacturer's recommended centrifugation instructions transfer serum or plasma for all specimens to a microcentrifuge tube for centrifugation at 150,000-300,000 g-minutes, following centrifugation guidelines in the chart in the letter.2. Prior to testing, transfer the serum or plasma to a sample cup or secondary tube for testing. Care must be taken to transfer only the clarified specimen and not the lipemic material or pellet.The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. On September 15, 2010, Abbott discovered that the calculation for the determination of RCF (xg) given in the September 13, 2010 letter was incorrect. Incorrect Calculation previously provided: Relative Centrifugal Force (RCF) = 0.0000118 x rotational radius in centimeters x RPM2 Correct Calculation within revised letter: Relative Centrifugal Force (RCF) = 0.00001118 x rotational radius in centimeters x RPM2 Abbott stopped distribution of the previous letter and initiated redistribution of the revised letter to all impacted customers, globally, on September 15, 2010. Additionally, all customers which had already received the initial letter were sent a cover letter identifying the error.Customers can contact Abbott at 1-877-422688 about this action. |  | 2411 | 2,411 kits | Worldwide Distribution: Throughout USA, including Puerto Rico, and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Latvia, Malaysia, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Saudi Arabia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 22, 2011 | 188 | 2411 | Not_Computer
Z-0625-2011 | 56957 | wearable automated external defibrillator | MVK |  |  | Cardiovascular |  | Model 4000 Monitor component of the LifeVest Wearable Defibrillator | LifeVest Wearable Defibrillator | 2 | December 16, 2010 | 2010 | Zoll Lifecor Corporation | The response button function may become intermittent. This can result in the device not progressing beyond the response button test during startup, which may prevent a conscious patient from delaying a treatment shock. | OTHER/UNDETERMINED: Pending | Zoll Life Vest contacted patients that are expected to be in the device more than 60 days by letter and phone to arrange replacement monitors. Instructions for the return of the recalled monitor accompany the replacement monitor The firm will also contact the prescribing physician via telephone for long-term use patients to inform the physician of this voluntary recall and to assist in contacting any patients the firm was unable to reach by telephone.Zoll Life Vest can be contacted at 1-800-543-3267 about this recall. | Contact the recalling firm for information | 494 | 494 units | Nationwide Distribution: Throughout the US. including the states of AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, and WV. | Terminated | July 05, 2012 | 567 | 494 | Hardware | Treatment/Delivery/Therapy | Replace monitor | Remove or Replace
Z-0623-2011 | 56679 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Vacuette 2 ml Lithium Heparin, Collection Tube, Manufactured by: Greiner-Bio One North America, Inc., 4238 Capital Drive, Monroe, NC.Venous blood collection tubes. | Vacuette 2 ml Lithium Heparin, Collection Tube | 2 | December 16, 2010 | 2010 | Greiner Bio-One North America, Inc. | Lithium heparin solution was contaminated by EDTA K (potassium) solution due to an operator error and a misinterpretation of a test result. Contaminated tubes will produce inaccurate test results. | OTHER/UNDETERMINED: Pending | Greiner Bio-One instructed customers to stop distribution and to complete an attached Product Disposition form and return it via fax to Greiner Bio-One. They were further instructed to destroy any product they may have. Distributors were instructed provide a list of their customers that have received the affected item via Excel file so that they can be notified of the recall. |  | 18 | 18 cases (total 21,600 pieces) | Nationwide Distribution: USA, including the states of CA, FL, IL, IN, KS, KY, ND, NE, NH, OH, SC, SD and WI. | Terminated | May 15, 2012 | 516 | 18 | Not_Computer
Z-0626-2011 | 57112 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 Inc. 4600 Nathan Lane, Plymouth, MN ProtÂŽgÂŽ " GPS" Self-Expanding Biliary Stent System, Sterile EO, 14 mm 60mm 80 cm, REF SERB65-14-60-80, Rx only.ev3 Inc. 4600 Nathan Lane, Plymouth, MN ProtÂŽgÂŽ " GPS" Self-Expanding Biliary Stent System, Sterile EO, 12 mm 60mm 80 cm, REF SERB65-12-60-80, Rx only. | Protg "GPS" SelfExpanding Biliary Stent System | 2 | December 16, 2010 | 2010 | Ev3, Inc. | Mislabeled size: Specific ProtÂżgÂż GPS lots have product labeling that does not match the actual stent sizes contained in the packages. The affected product outer packaging is labeled as a 12 mm x 60 mm ProtÂżgÂż GPS stent but contains a 14 mm x 60 mm stent. This situation can potentially result in vessel damage. Conversely, a lot was labeled as a 14 mm x 60 mm ProtÂżgÂż GPS stent but actually contai | PRODUCTION CONTROLS: Packaging Process Control | ev3 issued a "Medical Device Recall" letter dated September 29, 2010 to consignees. The letter described the product, Issue Summary, & Required Action. Consignees are to review inventory and remove all affected devices. ev3 representatives will contact customers to arrange for the return and credit of unused product.ev3 can be reached at 800 716-6700 concerning this recall. | Contact the recalling firm for information | 30 | 30 | Nationwide Distribution: including the states of AZ, DE, C, FL, GA, ID, IL, IN,LA, MN, MO, NJ, NY, NC, OH, VA, and WA. | Terminated | March 29, 2012 | 469 | 30 | Not_Computer
Z-0606-2011 | 57104 | aid, surgical, viscoelastic | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic |  | Beaver-Visitec International (Beaver-Visitec) BD 1% OVD Sodium Hyaluronate, 10 mg/ml 1.0 mL. Sterile. RxFor Intraocular UseREF 585304Use as a surgical aid to protect corneal endothelium during cataract extraction procedures, intraocular lens implantation and anterior segment surgery. | BD 1 OVD Sodium Hyaluronate | 2 | December 16, 2010 | 2010 | Becton Dickinson and Company | Finger flange on the BD 1% OVD syringe can disengage during cataract surgery and and may lead to injury. | DESIGN: Component Design/Selection | Beaver Visitec issued an Urgent: Field Safety Notice letter, dated 10/4/10 to direct accounts, identifying the affected product and informing users of the potential for the finger grip to disengage during use. The firm offers solutions as follows:Allow the finger grip to remain attached to the syringe and continue to use the product. Position the open side of the finger grip towards the palm to avoid detachment of the grip from the syringe. Remove the finger grip and continue to use the product. Return the unused product to Beaver-Visitec for reimbursement.Users were requested to complete the Customer Response Form. Questions can be directed to: Inside Sales at 781-906-7952. |  | 72278 | 72278 units | Nationwide Distribution | Terminated | February 16, 2012 | 427 | 72278 | Not_Computer
Z-0622-2011 | 56855 | pack, hot or cold, disposable | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | ThermaCare Menstrual 8 Hour Heat Wraps, 3 count, NDC 0573-3020-02, SKU: 302002, UPC: 3-0573-3020-02-9, Wyeth Consumer Healthcare, P.O. Box 26609, Richmond, VA 23261-6609 | ThermaCare Menstrual 8 Hour Heat Wraps | 2 | December 16, 2010 | 2010 | Pfizer Global | The chemical components contained in the heat wrap may leak directly onto the skin, which could cause skin injury, including burns. | PRODUCTION CONTROLS: Packaging Process Control | Wyeth Consumer Healthcare issued an Urgent Voluntary Medical Device Recall letter dated September 24, 2010 to direct accounts, and issued a press release. The firm is removing the product in question from store shelves and asking consumers who have purchased the product to record the lot number, throw the product away in its entirety without opening the foil pouch and call 1 800 232-3383 for replacement. Direct accounts were instructed to remove the affected product from sale and return it immediately. If they further distributed the product they were instructed to notify their customers of the recall and request return of the products to them credit. Instructions were provided in the event consumers inquired about the recall. A prepaid post card was enclosed to record the quantity on hand to be returned. | Contact the recalling firm for information | 192000 | 192000 cartons (3 wraps per carton) | Nationwide Distribution: Throughout the U.S. and Puerto Rico. | Terminated | September 16, 2011 | 274 | 192000 | Not_Computer
Z-0616-2011 | 57290 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | Fresenius Red Blood Cell Set, Catalog Number 9007601.The product is labeled in part: "RBC 9007601...STERILE...FRESENIUS KABI...Made in the Netherlands".Intended use: The Fresenius RBC red blood cell set (Art./Cat. No. 9007601) is used for depletion or exchange of red blood cells during therapeutic apheresis procedures on the Fresenius AS104 Blood Cell Separator Device. This set is used with a single stage separation chamber. | Fresenius Red Blood Cell Set (RBC) | 1 | December 16, 2010 | 2010 | Fresenius Kabi, LLC | Potential for patient hematocrit dropped to levels lower than expected post-treatment. |  | On October 21, 2010, Fresenius Kabi, LLC in Redmond, WA sent a "Field Safety Corrective Action" Letter dated October 18, 2010 to their US Consignees/Customers. On October 22, 2010, the firm followed up with their customers by telephone call.The letter informed the consignees/customers that the Red Blood Cell (RBC) exchange sets for use on the AS104 Blood Cell Separation Device are being recalled.The consignees/customers were advised to examine their stock and determine if they have any affected products on hand. They were instructed to discontinue distributing, using, dispensing the affected products, and return the product to Fresenius Kabi, LLC, 8635 154th Avenue, NE, Redmond, WA 98052.The consignees/customers were also instructed to notify their sub-account customers if the products were further distributed and to complete the Product Recall Response Form and fax it back to 425-242-2101. TIf you have any further questions concerning the recall, please contact the Quality Assurance Manager at 800-909-3872. |  | 255 | 255 units. | Nationwide distribution: including states of: AL, AR, AZ, CA, CT, GA, KY, MN, NC, NJ, OH, TN, and TX. | Terminated | October 05, 2011 | 293 | 255 | Not_Computer
Z-0711-2011 | 57230 | enzymatic esterase--oxidase, cholesterol | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Genzyme Diagnostics Total Serum Cholesterol Kit, IVDCatalog Number: 225-26Intended use: For the IN VITRO quantitative measurement of total serum cholesterol in serum. | Genzyme Diagnostics Total Serum Cholesterol Kit, | 3 | December 16, 2010 | 2010 | Diagnostic Chemicals, Ltd. | Cholesterol Reagent bottles labels reversed for R1 and R2 | PRODUCTION CONTROLS: Labeling Mix-Ups | Genzyme Diagnostics P.E.I. Inc. sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated November 10, 2010, via Federal Express to customers. The letter described the label error and requested that the reagents not be interchanged with another reagent lot. The firm ask that the customers complete and return the attached IMPORTANT PRODUCT FIELD CORRECTION confirmation of notification via fax at 902-628-6504. Contact a Genzyme Diagnostics Technical Services Representative at 1-800-565-0265, if you require further information, or if you have any questions. | Contact the recalling firm for information | 30 | 30 kits | Worldwide distribution: USA including states of: KY, PA, NV, and WA and country of Canada. | Terminated | September 19, 2011 | 277 | 30 | Not_Computer
Z-0713-2011 | 57220 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Custom Comprehensive Shoulder Bio-Modular Base, 44 mm Humeral Tray standard, sterile, REF CP561510, Biomet Orthopedics, Warsaw, IN Intended for total shoulder replacement in a reverse configuration. | Biomet Compehensive Reverse Shoulder | 2 | December 16, 2010 | 2010 | Biomet, Inc. | The firm received complaints regarding fracturing of the device. |  | The firm sent "URGENT MEDICAL DEVICE RECALL NOTICE" dated 9/13/2010 to their distributors and customers. The letters stated the reason for the recall and what may happen as a result of the fracture of the device. The notices stated that the recalled devices should be immediately located, discontinued from use, and shipped back to Biomet. Customers should follow the instructions on the "FAX Back Response Form." If the affected product was further distributed, hospital personnel must be notified of the recall via the enclosed "Dear Biomet Customer" notice. The firm sent "URGENT MEDICAL DEVICE RECALL NOTICES' to implanting surgeons. These letters informed the surgeons of the recall and instructed them to continue to monitor the patient for fracture of the device.Customers are to confirm receipt of the notice by calling 800-348-9500 ext 3755 or 3983. Questions should be directed to 574-371-3755 or 574-372-3983. | No consumer action necessary | 45 | 45 both products | Nationwide Distribution -- NY, VA, MD, AL, CO, SD, TX, and OK. | Terminated | July 27, 2011 | 223 | 45 | Not_Computer
Z-0628-2011 | 57281 | endoscope and/or accessories | KOG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ASP Automatic Endoscope Reprocessor (AER), Product Code: 20300 & 20301The AER is a machine designed to automatically wash and high-level disinfect flexible, submersible fiberoptic or video endoscopes. | ASP Automatic Endoscope Reprocessor (AER) | 2 | December 16, 2010 | 2010 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) received three customerreports of elevated temperature of the high-level disinfectant solution used in the ASP AER as a result of apossible temperature monitoring system malfunction. This recall communication advised consignees that if high-leveldisinfectants are exposed to elevated temperatures, fumes may occur and resul |  | Advanced Sterilization Products sent an URGENT: PRODUCT CORRECTION letter dated February 19, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall communication was initiated with ASP forwarding a "Urgent: Product Correction" letter with attached customer fax-back response form to the affected customers. Customers were also instructed to complete and sign the enclosed customer fax-back response form indicating that they have disconnected the heater on the ASP AER according to the instructions provided and fax the form to 1-888-345-1435.If customers have any questions or need assistance, they were instructed to call ASP Professional Services at 888-783-7723, option 3 then option 1. |  | 1881 | 1,881 units (1514 in US) | Worldwide Distribution -- United States, Argentina Brazil, Canada, Chile, Colombia, Costa Rica, Ecuador,Hungary, Mexico, Panama, Guatemala, Honduras, Poland, Puerto Rico, Virgin Islands, Singapore, Taiwan,Venezuela, Croatia, and Bulgaria. | Terminated | July 30, 2013 | 957 | 1881 | Not_Computer
Z-0714-2011 | 57236 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, medical charged-particle radiation therapy system, manufactured by Accuray Inc., Sunnyvale, CA.Indicated for treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | CyberKnife Robotic Radiosurgery System | 2 | December 16, 2010 | 2010 | Accuray Inc | If users modify the dose calculation box to cover just the contoured anatomy and not the entire CT volume, dosage may not be properly calculated. | OTHER/UNDETERMINED: Pending | Accuray Inc., sent an "URGENT DEVICE CORRECTION" letter dated October 21, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to increase the size of the dose box to include CT volume, not just the contoured anatomy; recalculate dose at high resolution, and examine the treatment plan to ensure that applicable organs at risk are accounted for and all portions of the patient receive the desired level of doses. If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA), +1-408-716-4700 (non-USA) or email: customersupport@accuray.com. | No consumer action necessary | 219 | 219 units | USA and countries of: Athens, Belgium, Canada, China, Czech Republic, France, Germany, India, Italy, Japan, Korea, Malaysia, Russia, Saudi Arabia, Spain, Switzerland, Taiwan, Thailand, The Netherland, Turkey, Ukraine, United Kingdom, Vietnam, Venezuela. | Terminated | April 01, 2011 | 106 | 219 | Other |  |  | Other
Z-0007-2011 | 54154 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed X/i Computed Tomography operator consoles with the Model Number 2153675 and 2247802.The expected usage of this product is to generate head and whole body CT images of human subjects. | HiSpeed X/i Computed Tomography | 2 | December 17, 2010 | 2010 | GE Medical Systems, LLC | It was discovered that the operator console did not have a label affixed in accordance with 21CFR 1010.2 indicating certification to applicable performance standards. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will bring customer systems into compliance with 21 CFR by having a GE Healthcare field service representative, who will perform a field corrective action, visit each customer site. This will be carried out via a Field Modification Instruction planned to be released in October 2009 and planned to be completed by February 2010. Additionally, all forward production operator consoles are manufactured so as to comply with 21 CFR and to include a manufacturing check for appropriate labeling. |  | 45 | 45 units (13 Nationwide) | Worldwide Distribution -- USA, including states of ME, FL, IA, AL, CT, CA, TX, KS, KY, OH, and WV and countries of Japan, Russian Federation, India, Argentina, Brazil, Italy, Poland, Puerto Rico, Romania, and Saudi Arabia. | Terminated | May 16, 2012 | 516 | 45 | Not_Computer
Z-0717-2011 | 57311 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Luminos dRFIntended use: solid state x-ray imager (flat panel/digital imager) | Axiom Luminos dRF | 2 | December 17, 2010 | 2010 | Siemens Medical Solutions USA, Inc | patient leg can be pinched | OTHER/UNDETERMINED: Pending | SIEMENS sent a Customer Safety Advisory Notice dated October 25, 2010, to affected customers via Update Instructions AX072/10/S. The notice described the issue and the actions to be taken by customers. The customers were instructed to retain the annex together with their operating instructions and bring the hazard to the attention of their employees; inform all users and other persons of this notice; if they sold the device, forward this notice and attachment to the new owners of the device and provide SIEMENS with details of the new owners.If you have any questions, please contact +1-888-826-9702. |  | 24 | 24 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, NC, NJ, and NY. | Terminated | September 29, 2012 | 652 | 24 | Not_Computer
Z-0716-2011 | 57242 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes A1, 20 - Dispensing tips, 29493. SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes A2, 20 - Dispensing tips, 29494.SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes A3, 20 - Dispensing tips, 29495.SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes B1, 20 - Dispensing tips, 29499.Revolution is a flowable light cure hybrid resin restorative. | Kerr Revolution Formula 2 Flowable Light Cure Composite | 2 | December 17, 2010 | 2010 | Patterson Dental Supply, Inc. | Patterson Dental is conducting a recall because they have been advised by the manufacturer of Revolution Formula 2 (Kerr Corporation) that some syringes and tips marked with the brand name are not manufactured and distributed by Kerr and, thus, not authentic product. | OTHER/UNDETERMINED: Pending | Consignees were sent a Patterson Dental "Urgent Device Recall" letter dated October 15, 2010. The letter was addressed to "Valued Customer". The letter described the product and the problem. It also advised the consignees to quarantine and return the product with the specific distribution dates. Consignees were to complete and return the enclosed Recall Return Form along with the affected product. Questions should be directed to Nancy Strand at 651-686-1966 or nancy.strand@pattersondental.com. | Contact the recalling firm for information | 3204 | 3,204 packages | Nationwide Distribution -- Including DC except VT. | Terminated | September 11, 2012 | 634 | 3204 | Not_Computer
Z-0718-2011 | 57312 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP Hybrid/Digital/CRIntended use: mobile x-ray system | Mobilett XP Hybrid/Digital/CR | 2 | December 17, 2010 | 2010 | Siemens Medical Solutions USA, Inc | battery fuse | OTHER/UNDETERMINED: Pending | SIEMENS released a field update dated October 25, 2010, via SP011/10/S to the representatives. The update described issue and corrective action plan. The representatives are to visit each account to upgrade the systems with a new battery fuse and a new holder.If you have any questions regarding this report, please contact Quality Engineer by phone 610-448-3237 or email: meredith.adams@siemens.com and Director, Product complaints & Field Support by phone 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 170 | 170 units | Nationwide distribution: AZ, CA, DC, MD, PA, SC and VA. | Terminated | September 29, 2012 | 652 | 170 | Battery |  | Repair | Repair
Z-0715-2011 | 57277 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra System Control Unit, a component of the StealthStation S7 System. Medtronic Navigation, Louisville, CO 80027.Intended as an aid for precisely locating anatomical structures in either open or percutaneous procedures. | Polaris Spectra System Control Unit | 2 | December 17, 2010 | 2010 | Medtronic Navigation, Inc | Electronic surgical navigation aid may lose tracking ability or stop working. | DESIGN: Component Design/Selection | Customers were notified by a Product Correction Notification letter on 11/05/2010. The letter advised customers of the affected product, the potential problem, and the description of the correction. Customers were told that a Medtronic rep will visit them to install a replacement control unit. Further information is available at 800-595-9709 or 720-890-3200. | No consumer action necessary | 48 | 48 devices | Worldwide Distribution -- US, including states of AL, AZ, CO, FL, GA, KY, MD, MN, NE, PA, TX, WI, and MD and countries of Australia, Belgium, Canada, Czech Republic, Italy, Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Turkey, and UAE. | Terminated | May 03, 2011 | 137 | 48 | Other
Z-0797-2011 | 57219 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital, Vitrea fX and Vitrea Enterprise Suite fXModel number: Vitrea fX 2.1, fX 3.0, and fX 3.1Vitrea Enterprise Suite (VES) fX 1.2 and VES fX 1.3The Vitrea 4D CT Brain Perfusion option is intended for post processing based on dynamic CT images continuously acquired during the infection of contrast, for the visualization of apparent blood flow in brain tissue and pictorial illustration of perfusion-related parameters to aid in the assessment of the type and extent of cerebral perfusion disturbances. | Vital, Vitrea fX and Vitrea Enterprise Suite fX | 2 | December 21, 2010 | 2010 | Vital Images, Inc. | Vital Images Inc., reported that a software defect exists with Vitrea fX 2.1, 3.0 and 3.1 CT Brain Analysis whereby the annotated numeric ROI values on a batched perfusion map are incorrect when created using the auto batching function (from the right click menu) with the auto ROI templates enabled. | DESIGN: Software Design | Vital Images, Inc sent a Vital "Urgent Software Update Alert" letter dated 14 October 2010 to all Consignees/Customers. The letter described the product and problem. They advised, until new software is installed, that the customers can avoid obtaining incorrect values by tuning off ROI templates prior to selecting the batch option from the right click menu in the CT Brain Analysis application; and recommended that the customers re-examine any ROI perfusion values that were previously generated using automatic collage batching with ROI templates.Please direct any questions or comments to Customer Support at support@vitalimages.com or at 1-800-208-3005. Any help needed to install this software update can also be directed to Customer Support. | Contact the recalling firm for information | 182 | 182 | Worldwide distribution: USA including states of: AL, AZ, AR, CA, CO, FL, GA, IA, KY, LA, MD, MA, MI, MN, MO, MT, NY, NJ, ND, OH, PA, PR, TN, TX, WA, WI, and WY, and countries of: CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, and JORDAN.CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, JORDAN. | Terminated | May 15, 2012 | 511 | 182 | Software | Output/Calculation | Software update | Software Update
Z-0794-2011 | 57351 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accessory Cart for The ACHIEVA 3.0T and Panorama 1T Magnetic Resonance Imaging unit.Consists of diagnostic devices that produce cross sectional images, spectroscopy images and/or spectra in any orientation of the internal structure of the whole body. | ACHIEVA 3.0T and Panorama 1T | 2 | December 21, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The Accessory Cart contained magnetic material and is attracted to the magnet when bringing it close to the system. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 05/12/2009, Philips began sending out the URGENT-Field Safety Notice MR Accessory Cart to their consignees. The letter identified the affected product and stated that the drawer rails were made from magnetic material and were not according specification. The hazards involved were also provided along with how to identify the affected product. The firm issued the FCO 78100296 and the firm's Field Service Engineer replaced the magnetic rails with the non-magnetic rails at the consignee's site.The consignees can contact contact the Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and reference FCO 78100296 with any communication. |  | 3 | 3 units | Nationwide Distribution -- IL, MI, and MD. | Terminated | December 28, 2010 | 7 | 3 | Not_Computer
Z-0796-2011 | 57349 | implant, cochlear | MCM |  |  | Ear Nose & Throat |  | Advanced Bionics HiRes 90K Cochlear Implant device (unimplanted units), Model Numbers: CI-1400-02HThe Harmony HiResolution Bionic Ear system is a cochlear implant system designed to provide useful hearing to individuals with severe-to-profound sensorineural hearing loss. | Advanced Bionics HiRes 90K Cochlear Implant device | 2 | December 21, 2010 | 2010 | Advanced Bionics Corporation | The recall action is being taken in response to two confirmed instances where the product experienced a malfunction requiring explantation. These recipients experienced severe pain, overly loud sounds and/or shocking sensations at 8-10 days after initial activation of their device. |  | Advanced Bionics (AB) Personal contacted all centers with known surgeries about the recall by telephone on November 22, 2010, and sent Blastfax to affected surgeons and clinicians.AB forwarding a Press Release on November 23, 2010 to the Business Wire News Trak. The firm also faxed Health Care Provider letters to the affected Health Care Providers and placed a Patient notification letter on the AB website on November 23, 2010. Health Care Provider letters were FedEx'ed November 24, 2010 to the affected Health Care Providers. Patient Letters were sent via FedEx'ed beginning Monday November 29th and 30th to approx. 10,500 patients. A Recall Acknowledgement Form was included for each recipient to sign as proof that the notification has been read.The press release & recall letter informed the patients/consignees of the reason for recall, the product affected and the potential hazard involved. The Health Care Provider letter also requested providers with recipient under their care who reports pain, overly loud sounds, or sudden shock sensations while the implant is powered to follow these steps:1. Advise them to remove the external equipment immediately,2. Schedule an appointment to complete standard troubleshooting steps listed in the AB Sound Processor User Guide,3. If standard troubleshooting does not resolve the issue, please contact Advanced Bionics for additional troubleshooting steps that may assist in determining the presence of recall problem, 4. If your evaluation of the patient's condition suggests that the implant exhibits this problem, device explantation and replacement is advised. In addition, the Health Care Provider letter informed providers that if AB records show that you have any unregistered implants in their possession, a list of serial numbers has been provided for devices that must be returned to Advanced Bionics. If no unregistered implants have been identified, no list will be enclosed. The Patient letter also . |  | 82 | 82 units | Nationwide distribution: AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT and VA. | Terminated | March 23, 2012 | 458 | 414 | Not_Computer
Z-0719-2011 | 57291 | urinometer, electrical | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BardÂŽ CritiCoreÂŽ Monitor, Software Version 2.1, Catalog number 000002N, C.R. Bard, Inc., Covington, GA 30014.Intended for multiple uses as it is an electronic battery-operated monitor used in conjunction with a single-use disposable urine collection container and/or a temperature sensing Foley catheter | Bard CritiCore | 3 | December 21, 2010 | 2010 | C.R. Bard, Inc., Urological Division | Error can occur in the reading of the Present Interval Output (PIO) reading when the device is subjected to bumping or forceful replacement of the fluid container. | OTHER/UNDETERMINED: Pending | Consignees were notified by an Urgent: Product Advisory Letter and Medica Device Recall letter on/about November 18, 2010. The letter identified the affected product, stated the reason for recall, the interim precautions, and actions to be taken. The letter requested acknowledgement of the recall and to return the entire affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. Affected product received by Bard Medical will be reworked with the software upgrade and place back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. For international distribution, the hospital will be informed of this issue through Bard affiliate and Notified Body Bard Europe and managed according to the requirement. If there are questions or assistance is needed, customers can contact BD at 1-770-784-6969. | Contact the recalling firm for information | 234 | 234 units | Worldwide Distribution -- U.S., Netherlands, and Japan. | Terminated | February 02, 2012 | 408 | 234 | Software | Output/Calculation | Software update | Software Update
Z-0507-2011 | 57376 | surveying laser product | REP |  |  |  |  | Aimshot HeatSeeker Model HS3500, contains Class IIIb infared laser illuminator. Amstech Incorporated, Alpharetta, GA 30022.Targeting, detecting heated areas, searches. | Aimshot HeatSeeker | 2 | December 22, 2010 | 2010 | Amstech Incorporated | Product contains infra red laser, which can pose a potential hazard to users who unknowingly shine the beam into people's eyes. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Firm sent a letter dated August 5, 2010 to customers stating that a recall was being conducted on the AimSHOT HeatSeeker and specified model. The letter stated the reason for the recall and that the firm has redesigned the HeatSeeker and removed the infra red laser feature. The letter also asks customers to check their inventory and return the affected products to the firm. The firm states that they will send customers their new model and pay for the shipping. Questions should be directed towards 770-840-8039. | Contact the recalling firm for information |  | All units | Nationwide Distribution | Terminated | November 09, 2011 | 322 |  | Not_Computer
Z-0802-2011 | 57209 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | klinoPORT, AC powered Operating Table | klinoPORT | 2 | December 22, 2010 | 2010 | Kreuzer Gmbh & Co. KG | The motorized adjustable pendant arm may suddenly descend to its lowest position. | OTHER/UNDETERMINED: Pending | Consignees were notified by a Warning Notice on/about 10/26/2010. Medical facilities were instructed to adjust the pendant arm to the lowest height position and disable the motor and to place the accompanying warning label on the pendant. They will, therefore, not be able to adjust the height of the pendant until corrective action is developed and completed. TRUMPF Medical Systems, Inc. will receive corrective action repair kits by 11/12/2010. | Contact the recalling firm for information | 15 | 15 units | Nationwide Distribution -- CT, MA, and MO. | Terminated | March 20, 2012 | 454 | 15 | Not_Computer
Z-0801-2011 | 57210 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Turon Shoulder System Instrumentation, Turon Pegged Glenoid Drill Guide, REF 804-25-024, DJO Surgical | Turon Pegged Glenoid Drill Guide | 2 | December 22, 2010 | 2010 | Encore Medical, Lp | Complaints of fracturing were reported for affected trial products. | DESIGN: Device Design | Encore Medical notified consignees by letter on 11/5/2010. | Contact the recalling firm for information | 34 | 34 units | TX, FL, SC, GA | Terminated | April 05, 2011 | 104 | 578 | Not_Computer
Z-0790-2011 | 56965 | system, balloon, intra-aortic and control | DSP | Intra-aortic balloon and control system | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc.Model: IAB-05830-U | Ultra 8 IAB | 1 | December 23, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | DESIGN: Component Design/Selection | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Open |  |  | 44548 | Not_Computer
Z-0789-2011 | 57335 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 NanoCross 0.14 OTW PTA Dilatation Catheter, sterilized using ethylene oxide gas. REF #'s: AB14W040040150, AB14W030040150, AB14W020040150, AB14W025080150, AB14W020080150, AB14W020120150, AB14W030080150, AB14W035120150, AB14W035040150, AB14W030120150, AB14W025120150, AB14W040080150, AB14W040120150, AB14W040210150, AB14W030150150, AB14W025210150, AB14W020150150, AB14W025040150, AB14W040150150, AB14W035080150, AB14W020210150, AB14W035210150, AB14W025150150, AB14W030210150, AB14W035150150, AB14W040080090, Intended to dilate stenosis in the iliac, femoral, ilio-femoral, popliteal, infra-popliteal, and renal arteries, and for the treatment of obstructive lesions of native or synthetic arteriovenous dialysis fistulae.This device is an over the wire (OTW) 0.014" coaxial lumen catheter with a distally mounted semi-compliant inflatable balloon and an atraumatic, tapered tip. The distal portion of the catheter shaft and the balloon has a hydrophilic coating. The manifold includes a lumen marked "THRU". This is the central lumen of the catheter that terminates at the distal tip. This lumen is used to pass the catheter over a guide wire with a maximum outer diameter of 0.014". The lumen marked "BALLOON" is used to inflate and deflate the dilatation balloon with a solution of contrast medium and saline. The balloon has two radiopaque markers for positioning the balloon relative to the stenosis. The radiopaque marker bands indicate the dilating or working section of the balloon. The 210mm balloon contains two additional radiopaque marker bands that denote the middle of the balloon body. | ev3 NanoCross 0.14 OTW PTA Dilatation Catheter | 1 | December 23, 2010 | 2010 | Ev3, Inc. | ev3 has determined that specific lots of the NanoCross .014 OTW PTA Dilatation catheters may crack or break during use. Cracking or breaking of the catheter shaft can result in the inability to inflate or deflate the balloon, and can result in material separation and potential embolization. | PRODUCTION CONTROLS: Process Control | A Medical Device Recall letter addressed to Risk Managers, dated 11/10/2010, was sent to hospitals beginning 11/12/2010. The letter identified the affected devices, described the issue and requested that affected devices be located and removed from use. ev3 sales representatives will contact customers to arrange return of all unused devices. A credit or replacement will be provided for any unused devices. Questions regarding this issue may be directed towards Customer Service at 800-716-6700. | Contact the recalling firm for information | 4962 | 4,962 (3,858 US; 1,104 OUS) | Worldwide Distribution -- USA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHILE, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KUWAIT, LEBANON, MALTA, NETHERLANDS, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SYRIA, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM, | Terminated | May 17, 2011 | 145 | 4962 | Not_Computer
Z-0816-2011 | 56972 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Vacutainer Push Button collection Set;Non-Pyrogenic, Sterile, Rx only;23 G x 3/4 x 12" (0.6 x 19 mm x 305 mm)Becton Dickinson & Co., Franklin Lakes, NJ, USAIntended use: Blood collection | BD Vacutainer Push Button collection Set | 2 | December 27, 2010 | 2010 | Becton Dickinson & Company | There is the possibility for certain lots of the BD Vacutainer Push Button Blood Collection sets to have separation of the cannula from the hub. | PRODUCTION CONTROLS: Process Control | BD Diagnostics sent "RECALL NOTIFICATION" separate letters October 1 and 5, 2010 and follow up letter dated October 8, 2010, as well as a letter dated December 1, 2010 to all distributors and customers commencing October 5, 2010 by federal express. The letters describe the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory for specific catalog and lot numbers, and return product from the lots following the instructions on the enclosed packing slip (BD will send replacement upon receipt of returned product). If you have any questions on returning and replacing product, please contact BD logistics at 1-866-380-0046. If you have any further detail questions regarding this recall, please contact your BD Sales Consultant or BD Technical Services via email at www.BD.com/Vacutainer/Contact or via telephone at 1-800-631-0174. |  | 4223200 | 4,223,200 | Nationwide Distribution | Terminated | June 06, 2012 | 527 | 4223200 | Not_Computer
Z-0819-2011 | 56894 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Depuy Spine Mountaineer OCT Spinal System Screw Caddy, Non SterileProduct Code: 2883-90-420The MOUNTAINEER Screw caddy is intended to act as a holder in an implant set for implantable screws prior to use in surgery. Further, the caddy provides a means to organize the screws in accordance with the length of the screws. | Depuy Spine Mountaineer OCT Spinal System Screw Caddy Set | 2 | December 27, 2010 | 2010 | DePuy Spine, Inc. | The scale which serves as a secondary means of identifying screw length is incorrectly etched on the screw caddy | DESIGN: Labeling Design | Depuy Spine issued an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" via e-mail on October 2, 2010 to customers. The notification describes the product, problem and action to be taken by customers. The customers were instructed to please distribute this information to appropriate personnel at their facility and return any caddies that they may have to their local distributor for DePuy Spine products. Depuy will return devices to customer promptly (within two business days), after reworking the devices by removing the scale altogether. Once the product is fully corrected, Depuy will provide the customers with a new and updated unit with the corrected scale markings. If you should have any questions or concerns regarding this action, please contact DePuy Spine Customers Service at 800-227-6633. | Contact the recalling firm for information | 237 | 237 units | Nationwide distribution: USA including states of: AR, AZ, CA, FL, HI, IA, IL, KY, LA, MA, MI, MN, MO, NC, NY OH, PA, TN, TX, VA, and WA. | Terminated | April 15, 2011 | 109 | 441 | Not_Computer
Z-0808-2011 | 56846 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Model # 720088-01: AMS AdVance Male Sling System, US; Model # 720088-02: AMS AdVance Male Sling System, INTL; Model # 72404224: Male TO Sling System, INTL.Sterilization Method ETO, American Medical systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA | AdVance Male Sling System | 2 | December 27, 2010 | 2010 | American Medical Systems, Inc. | American Medical Systems (AMS) is recalling all lots of the AdVance male Sling System with an expiration date between July 2012 and August 2013. Internal testing has indicated the potential for degradation of the bioabsorbable tensioning sutures when the product is stored for longer than 15 months after the date of manufacture. | DESIGN: Packaging Design/Selection | Consignees were sent an AMS "Urgent Product Recall" letter dated September 7, 2010. The letter described the problem and the product involved. They requested consignees to return any existing stock product and to return the Product Recall Acknowledgement form enclosed. Customers may request a credit or a full refund or also, arrange for replenishment of AdVance product. Questions regarding the recall should be directed towards AMS Customer Service at 1-800-328-3881. | Contact the recalling firm for information | 7069 | 7,069 | Worldwide Distribution -- All states in the Continental USA including DC and Puerto Rico and countries of ARGENTINA, AUSTRIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, GERMANY, ESTONIA, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, CROATIA, IRELAND, ISRAEL, INDIA, ITALY, REPUBLIC OF KOREA, LEBANON, MOROCCO, MEXICO, MALAYSIA, NETHERLANDS, NORWAY, NEW ZEALAND, PANAMA, POLAND, PORTUGAL, RUSSIAN FEDERATION SAUDI ARABIA, SWEDEN SINGAPORE, SLOVENIA, THAILAND, TURKEY, TAIWAN, VENEZUELA, and SOUTH AFRICA. | Terminated | June 28, 2012 | 549 | 7069 | Not_Computer
Z-0820-2011 | 55976 | prosthesis, testicular | FAF | Testicular prosthesis. | Gastroenterology/Urology | Gastroenterology/Urology |  | Coloplast, Saline Filled Testicular Prosthesis, Sterile R, Manufacturer Coloplast A/S 3050 Denmark.Catalogue Number Description Part Number450-1323 Extra Small 520600450-1325 Small 520610450-1327 Medium 520630450-1329 Large 520650The Coloplast Saline-Filled Testicular Prosthesis is intended for use when cosmetic testicular replacement is indicated i.e., in the case of agenesis or following the surgical removal of a testicle. The Butterfly Needle is used to fill the implant with saline. | Saline Filled Testicular Prosthesis | 2 | December 27, 2010 | 2010 | Coloplast Manufacturing US, LLC | On April 22, 2010, a decision was made by Coloplast to conduct a field correction based on a report that the Saline Filled Testicular Prosthesis is considered mislabeled. This is solely due to a difference in expiration dates between the Terumo Butterfly Needle included in the kit and the testicular implant. The needle has a 3yr expiration date & the Saline Testicularhas a 5yr expiration date. | MISBRANDING: Labeling False and Misleading | The recall will extend to field personnel who have control of all SFTP Systems. The product is ordered by the customer (physician and/or hospital) directly from a field sales representative who ships the device to them or drops it off in person.Coloplast sent all Field Personnel an email on April 30, 2010, with a copy of a letter dated April 30, 2010 and a follow-up letter dated May 1, 2010. The letters described the problem and the product involved in the recall. The letters included a list of Corrections Required. This included the field personnel receiving a stock of sterile Butterfly Needles based on their home consignment inventory to be provided for use with all Saline Testicular Implants; Copies of customer letters to be provided with each replacement needle; Changes in all Saline Testicular orders, and Field personnel to administer the details of the correction and make the direct customer contacts. For questions, contact the Customer Service/Field Inventory Representative at 612-337-7859 or email: usbam@coloplast.com. | Contact the recalling firm for information | 1018 | 1018 | Worldwide distribution: USA including states of AL, AK, AZ, AR, CO, CT, CA, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI,MN, MS, MO, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT and WA; and countries of Australia, Canada, Hong Kong and Latin America (Chile, Columbia, Ecuador and Venezuela). | Terminated | December 17, 2011 | 355 | 1018 | Not_Computer
Z-0815-2011 | 57292 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, & Specialty Head Center Accolade Broach Handle; Stryker Howmedica Osteonics, Non-sterile Instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; | CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, and Specialty Head Center Accolade Broach H | 2 | December 27, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, and Specialty Head Center Accolade Broach Handle may fracture in two pieces along the width of the body under the striking plate. This may extend surgical time under anesthesia, cause trauma to the surgeon or patient, or contaminate the surgical site. |  | Stryker issued a Product Correction Bulletin dated November 2010 to their consignees, alerting surgeons of the need to inspect the instruments before and after each surgical use, and, if a defect is detected, to contact Stryker for return of the product. Stryker can be contacted at 201 831-5288 concerning this issue.In response to their revised strategy, Stryker issued notification letters and product recall acknowledgement forms to their branches and agencies on 2/17/2011 for 2/18/2011 delivery via Fed Ex and to their hospital risk managers on 2/18/2011 for 2/21/2011 delivery via Fed ex. |  | 459 | 459 units | Worldwide Distribution: Throughout the US, the countries of Canada and Japan, and the regions of Asia Pacific, Europe, Middle East, Africa, and Latin America. | Open |  |  | 459 | Not_Computer
Z-0807-2011 | 57274 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ACL ELITE PRO Coagulation Analyzer, PN 00001000311Coagulation testing. | ACL ELITEPRO | 2 | December 27, 2010 | 2010 | Instrumentation Laboratory Co. | Sporadically lowered APTT results were observed with the PT/APTT Test Groups. | DESIGN: Device Design | On 11/11/10, An Urgent Product Notification informing customers to discontinue use of all IL locked PT/APTT-based Test Groups was mailed to U.S. and Canadian customers, instructing them to run in Single Test or Profile Mode until further notice. The letter identified the affected product, the issue, risk and stated a list of mandatory actions. Customers are to immediately acknowledge the receipt of the notification by signing and returning it via email or fax. A Worldwide Regulatory Notice was sent on November 8, 2010 to international affiliates (non-U.S. or Canada), notifying them of the issue, providing a risk summary and the support documentation for the notification. Secondary international dealers were notified Nov 8 and a fax response of notification of the dealers will be tracked by International Laboratories. The customer notification of the secondary dealers will be handled per the local regulatory requirements of the countries where they are located. |  | 379 | 379 in use in USA & Canada, 24 in distribution channel subject to recall | Worldwide Distribution -- USA, Canada, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Thailand, Trinidad & Tobago, Tunisia, Turkey, Turkmenistan, United Arabian Emirates, Uruguay, Venezuela, Vietnam, Austria, Belgium, China, France, Germany, Holland, Italy, Japan, Spain (and Portugal), and UK. | Terminated | August 23, 2013 | 970 | 466 | Not_Computer
Z-0813-2011 | 57348 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Blease Sirius Anaesthesia System is a anaesthesia Gas Machine. The system has the capability to mount two anaesthetic vapourisers on the back bar via a demountable system. The delivery of anaesthetic agent is achieved by the rotation of a control knob which actuates a feature designed "shot bold" mechanism. The purpose of which is to prevent two vapourisers being turned on simultaneously..The product is labeled in part: "BleaseSirius Anesthesia System...SPACELABS Healthcare...Blease Medical Equipment Limited Beech House, Chiltern Court, Asheridge Road, Chesham, Bucks HP5 2PX, England...Spacelabs Healthcare Company Headquarters 5150 220th Ave. SE, Issaquah, WA 98029 PO Box 7018, Issaquah, WA 98027-7018". | Blease Sirius Anaesthesia Systems | 2 | December 27, 2010 | 2010 | Spacelabs Healthcare, Llc | Users may experience difficulty turning the control knob to an active vapour delivery setting, or may not be able to turn on the Left Hand mounted vapouriser on the Blease Sirius Anaesthesia Systems. | OTHER/UNDETERMINED: Pending | Spacelabs Healthcare issued an URGENT MEDICAL DEVICE CORRECTION letter dated June 25, 2009 to the two consignees in the US. Spacelabs Healthcare UK notified consignees outside the US.The consignees were advised to brief their staff about the issue and test the Anesthesia Systems in advance of a procedure. If there is an issue with turning on a vaporizer in the left hand position, simply move the vaporizer to the right hand position.The firm contacted their consignees for a schedule to replace the flowmeter mouldings. The US consignees can contact Spacelabs Healthcare Representatives at 425-657-7200 ext. 5089 about this recall. Foreign consignees can contact Spoacelabs at 44 1494 784422. |  | 5 | 5 units to US consignees and 36 units Internationally. | Worldwide Distribution: US in the states of LA and SD, and internationally by Spacelabs Healthcare, Ltd in the countries of China, Czech Republic, Hong Kong, India, Poland, Saudi Arabia, Singapore, and the United Kingdom. | Terminated | December 28, 2010 | 1 | 5 | Not_Computer
Z-0814-2011 | 57337 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ThomoTherapy Hi-Art SystemIntended to be used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | HiArt System | 2 | December 27, 2010 | 2010 | TomoTherapy Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the Hi-Art System and make it so the user cannot generate a Completion Procedure. | DESIGN: Software Design | Consignees were sent a TomoTherapy "Urgent Field Safety Notice" dated November 8, 2010. The letter addressed to TomoTherapy Customers described the Issue, Product Affected, Recommended Action and Resolution. Customers are to contact the TomoTherapy Customer Interaction Center for further instruction and assistance in generating a completion procedure, if they experience the issue. Questions should be directed towards the TomoTherapy Customer Interaction Center via email or telephone. | No consumer action necessary | 40 | 40 (21US; 19 OUS) | Worldwide Distribution -- USA, including the states of AZ, CO, IL, KS, MI, MO, NV, NY, ND, OH, OK PA, UT, and VA and countries of MALAYSIA, SWITZERLAND, KOREA, ITALY, SPAIN, CANDA, TAIWAN, JAPAN, BELGIUM | Terminated | February 04, 2012 | 404 | 40 | Software
Z-0811-2011 | 57317 | biopsy needle kit | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Expect Endoscopic Ultrasound Aspiration Needle, 22GA (0.72 mm) sterile, UPN M00550011, 5 pack, 428 Technology Drive East, Menomonie, WI.Designed to sample targeted submucosal and extramural gastrointestinal lesions through the accessory channel of a curvilinear echoendoscope. | Expect Endoscopic Ultrasound Aspiration Needle | 2 | December 27, 2010 | 2010 | Boston Scientific Corporation | The firm has received a high number of complaints related to difficulty advancing and/or retracting the needle. | OTHER/UNDETERMINED: Pending | "Urgent Medical Device Recall - Immediate Action Required" notification letters were sent via Federal Express Priority mail on October 20, 2010. The letters identified the affected product and stated the reason for the recall. The notification instructs the account to immediately discontinue use of, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. | Contact the recalling firm for information | 567 | 567 all products. | Worldwide Distribution -- US and Canada. | Terminated | August 03, 2011 | 219 | 567 | Not_Computer
Z-0831-2011 | 51260 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Manual Wheelchairs with Transit System Option,Model: Quickie LXI (EILXI),Product is manufactured and distributed by Sunrise Medical, Inc. Fresno, CA Intended to provide mobility to persons restricted to a sitting position and are unable to transfer from their wheelchairs when traveling in motor vehicles, the wheelchair must serve as a vehicle seat. | Manual Wheelchairs with Transit System Option | 2 | December 28, 2010 | 2010 | Sunrise Medical Inc | The stabilizer bar was inadvertently omitted during the manufacturing process. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | An Urgent: Medical Device Correction letter, dated February 2009, with certified return receipt was sent to customers in the United States; Canadian consignees were notified via e-mail. The letter identified the affected product and the reason for the recall. The letters asked customers to contact their Quickie dealer to make arrangements to have their device wheelchair corrected at no charge and to also verify their wheelchair serial number on the form. Questions should be directed towards their dealer or Sunrise Medical at 800-456-8168. |  | 23 | 23 | Worldwide Distribution -- United States and Canada. | Terminated | January 03, 2011 | 6 | 23 | Not_Computer
Z-0830-2011 | 53183 | bougie, esophageal, and gastrointestinal, gastro-urology | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker Infravision Esophageal Kit, Model Numbers 220180548, 220180540, 220180560, manufactured by Stryker Endoscopy San Jose. | Stryker Infravision Esophageal Kit | 2 | December 28, 2010 | 2010 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter during removal of the e-kit from the body. | DESIGN: Device Design | Stryker issued Urgent Product Recall letters dated August 20, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the affected product to Styrker. |  | 2985 | 2985 units | Worldwide Distribution: USA in the states of: Alabama, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, Nebraska, New Jersey, New Mexico, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and the countries/region of Australia, Canada, Scandinavia, Italy, and Switzerland. | Terminated | January 03, 2011 | 6 | 2985 | Not_Computer
Z-0825-2011 | 54147 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker brand End Cutter Shaver Blade F-Series, 4.0 mm; Model Number -375-747-000, manufactured by Stryker Puerto Rico | End Cutter Shaver Blade FSeries, 4.0 mm | 2 | December 28, 2010 | 2010 | Stryker Endoscopy | Potential breakage of the internal shaft of the End Cutter Shaver Blade; F-Series, 4.0mm. if breakage occurs, the device becomes non-functional and the defect could be observed through the camera/console system. | DESIGN: Device Design | Stryker issued Urgent Product Recall letters dated November 12, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the product to Stryker.Stryker can be contacted concerning this recall at 408 754-2124. |  | 295 | 295 units | Nationwide Distribution includes: Arizona, California, Colorado, Denver, Florida, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Michigan, New York, Ohio, and South Carolina. | Terminated | January 03, 2011 | 6 | 295 | Not_Computer
Z-0823-2011 | 56245 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Steris Reliance EPS Endoscope Processing System , manufactured by Corporation STERIS Canada, Quebec, Canada.Washing and high level disinfection of up to two manually pre-cleaned, immersible, reusable, heat-resistant, semi-critical endoscope devices. | Reliance EPS Endoscope Processing System | 3 | December 28, 2010 | 2010 | Steris Corporation | Reliance EPS had internal water leaks and interrupted cycles when air leaks triggered false Fault 38 alarm that DG/HLD cup has not been replaced. | OTHER/UNDETERMINED: Pending | Steris issued an "Urgent Field Correction Notice" dated July 2, 2010 to each affected US consignee, notifying them of the issue and affected device. Steris Field Service Representatives and Dealers will contact each affected customer to schedule a service visit to install system upgrades. In the event of an interrupted cycle prior to the upgrade, consignees were instructed that the load should be handled in accordance with the Reliance EPS Operators Manual. Steris can be contacted at 440-392-7043. |  | 936 | 936 | Worldwide distribution: USA, including states of AK, AL, AZ, CA, CO, DE, FL, ID, IL, KS, KY, MD, MI, MN, MO, NC, NV, NY, OH, PA, SC, VA, WA, WI, and WY, and countries of Algeria, Australia, Belarus, Canada, France, Germany, Greece, India, Italy, Japan, Kuwait, Lebanon, New Zealand, Russian Federation, Saudia Arabia, Spain, Turkey, United Arab Emirates and United Kingdom. | Terminated | September 27, 2012 | 639 | 936 | Not_Computer
Z-0822-2011 | 56170 | [6 empty] | BD Beaver Slit Blade 3.0mm StraightPart Number: 375521The BD Beaver Slit Blade is used to perform the main incision through which instruments are inserted during cataract surgery. | BD Beaver Slit Blade 3.0mm Straight | 2 | December 28, 2010 | 2010 | Bd Medical Systems | Blades may contain defects at the blade tip or cutting edge. |  | BD Medical notified customers via letter dated 6/24/10 to discontinue use of product and remove from inventory. Customers were instructed to complete the customer response form and fax to BD Medical. BD Medical was to contact customers upon receipt of the customer response form to arrange for replacement.Unused product was to be returned to:BD MedicalOphthalmic SystemsAttn: WT-12582411 Waverly Oaks RoadBuilding 2, Suite 229Watham, Massachusetts 02452-8402For questions regarding this recall call 1-781-906-7952. |  | 524 | 524 pieces | Nationwide Distribution | Terminated | June 29, 2012 | 549 | 524 | Not_Computer
Z-0821-2011 | 56774 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Trilogy System with RapidArc, Model Numbers: H14, H26, H27, H29; Distributed and Manufactured by: Varian Medical Systems Inc., Palo Alto, CA | Varian Trilogy System with RapidArc | 2 | December 28, 2010 | 2010 | Varian Medical Systems Inc. | The system may deliver an inaccurate dose that does not match the treatment plan. As a result of a system configuration problem, the VMAT (Volumetric Modulated Arc Therapy) plans may be delivered as Dynamic Arc therapy plans. If this occurs, the VMAT plan will not be recognized. This issue affects Trilogy machines with the Rapid Arc Treatment delivery option. | DESIGN: Software Design | Varian issued Urgent Medical Device Correction letters dated July 20, 2010 to its direct consignees, informing them of the affected product and providing instructions on the actions to be taken. These instruction include: if the unit does not display "RapidArc" or "VMAT" under Treatment Type during dose delivery, discontinue use of the device and call Varian service.Varian Service Help Desk can be contacted at 1-888-827-4265. |  |  |  | Worldwide Distributiuon: Throughout the US, and the countries of Australia, Canada, China, Japan, and New Zealand, and in the regions of North and Latin America, Europe, and South East Asia. | Terminated | July 18, 2011 | 202 |  | Software
Z-0829-2011 | 57273 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instruments offset reamer bushing - angled, REF 32-488519, Biomet Bridgend, UK.This instrument is used to prepare the tibial cancellous bone for the tibial offset tray. | Offset reamer brushing angled | 2 | December 28, 2010 | 2010 | Biomet, Inc. | The firm has initiated the action following a report that during surgery the Offset Reamer Brushign would not mate with the prepartion mask. After a 5 minute delay in surgery the procedure was completed using another attachment that was on hand. There was no injury to the patient. | OTHER/UNDETERMINED: Pending | The firm, BIOMET, sent a "FIELD SAFETY NOTICES" dated September 10, 2010, addressed to the head of orthopaedic departments, trauma departments, operating departments, sterile service departments, procurement, supplies and risk management. The notice described the product, problem and action to be taken by the customers. The letters instructed the customers to locate and immediately quarantine the recalled device; to complete and return the attached "FAX BACK RESPONSE FORM" via fax to +44 (0) 1656 645454, and the recalled device for credit via mail to Biomet UK Ltd, Waterton Industrial Estate, Bridgend, CF31 3XA. The customers were also instructed to inform all in their organization who are using or ordering the products and send a copy to the organization to which potentially affected products may have been transferred.If you have any questions regarding this communication, please contact the Quality Manager at +44 (0) 1656 655221 or email: david.broderick@biomet.com. |  | 16 | 16 units | International distribution only: Netherlands, Germany, France, Spain, Italy and Denmark | Terminated | January 19, 2011 | 22 | 16 | Not_Computer
Z-0828-2011 | 57286 | extractor | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CLP Instrumentation Removal Hook, distributed by DJO Surgical, manufactured by Encore Medical, Austin, TX. | CLP Instrumentation Removal Hook | 2 | December 28, 2010 | 2010 | Encore Medical, Lp | Revision A and B CLP Extraction Hooks may fracture at the distal end, resulting in surgical delay. | OTHER/UNDETERMINED: Pending | DJO issued an Urgent Recall Notice letter dated November 10, 2010 to consignees identifying the affected devices and requesting return of all affected devices for replacement. Customers are also requested to complete and return an enclosed Recall Response Form.Customers can contact David Harris of DJO at 512 834-6330 with questions about this recall. |  | 27 | 27 Units | Nationwide Distribution to the states of AZ, CA, FL, GA, ID, KS, MO, NY, TX, and WI. | Terminated | April 25, 2011 | 118 | 27 | Not_Computer
Z-0826-2011 | 57289 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 4Z1c matrix array transducer which is used with ACUSON SC2000 ultrasound system.Siemens Medical Solutions USA, Inc.Business Unit Ultrasound1230 Shorebird WayMountain View, CA 94043Intended use: 4Z1 transducer is a special purpose ultrasound transducer designed to help clinicians perform adult echocardiography and pre-operative and intra-operative surgical assessments. Secondary applications include: pediatric, echocardiography, primarily benefiting older and larger pediatric populations, general abdominal and Pinta-operative pediatric abdominal imaging and peripheral vessel imaging. | 4Z1c matrix array transducer | 2 | December 28, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | There is the possibility of a short occurring in the transducer which could cause focal heating on the probe face. | OTHER/UNDETERMINED: Pending | SIEMENS issued a filed correction in the form of a "Customer Safety Advisory Notification" initiated on September 29, 2010 and mailed to customers by Federal Express on the same day. The notification described the product, problem and actions to be taken by the customers. The customers were instructed to examine the 4Z1c transducer prior to use (Note: if transducer does not show a blister on its front face or does not feel unusually when it is safe for use); to immediately discontinue use and contact their local Siemens representative to get issue resolved, if the transducers developed a hot spot that exhibits a visible blister on the patient contact surface and feels hot to the touch during operation, and to pass this notice on to all those within their organization who need to be aware.If you have any questions, please contact (650) 969-9112 or (800) 422-8766. |  | 2196 | 196 units both Transducer and ACUSon SC2000 | Worldwide Distribution: USA and countries including: Australia, Belgium, Brazil, China, Czech Republic, Finland, France, Germany, India, Ireland, Italy, Japan, Kenya, Malaysia, Netherlands, Norway, Portugal, South Korea, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom. | Terminated | October 04, 2011 | 280 | 2196 | Not_Computer
Z-0824-2011 | 57341 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long, Catalog 00-2490-003-50, Zimmer, Warsaw, IN.A reusable surgical instrument used for the insertion of a Lag Screw used in the Zimmer Natural Nail system. | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long | 2 | December 28, 2010 | 2010 | Zimmer Inc. | The firm has determined that two lots of this instrument were inappropriately manufactured by the supplier. The manufacturing process created a step in the lumen which impedes the Lag Screw Retaining Shaft from passing through the lumen without the use of excessive force. | TRAINING: Employee Error | The firm sent Correction and Removal Notification Immediate Response Required letters via email to consignees. The letter identified the affected product and the reason for the recall. The letters also discussed the clinical implications and required actions. Customers are to complete the inventory certification form and user facility documentation form. Each of the affected accounts will be provided with a new instrument, manufactured with the correct lumen diameter and will secure and return the affected instruments per the attached notification, along with the completed inventory certification form. If any further information is needed, customers are to contact Aimee Wood at 574-372-4463. | Contact the recalling firm for information | 34 | 34 | Worldwide Distribution -- USA, including states of MI, NJ, MS, OR, KS, TN, A, UT, and CA and countries of Belgium and Germany, Spain, Italy, Greece, France, UK, and Switzerland. | Terminated | September 16, 2011 | 262 | 34 | Not_Computer
Z-0834-2011 | 56763 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GreenLight MoXy Fiber Optic, part number 0010-2400, for use with the GreenLight XPS Laser System, manufactured by American Medical Systems Inc., San Jose, CAThe XPS Laser system is indicated for surgical incision and excision, vaporization, ablation, hemostasis and coagulation of soft tissue. | GreenLight MoXy Fiber Optic | 2 | December 29, 2010 | 2010 | AMS Innovative Center - San Jose | Product not intended for release was distributed, which may result in old components, not built to specifications, or undergo final testing |  | American Medical Systems sent oral notifications that the product used for patient case evaluations was from the 2 year shelf life study. Products have been returned, as the product was with AMS field personnel.For questions regarding this recall call 847-855-6270. |  | 77 | 77 units | Worldwide Distribution, USA including TX, OH, MO, FL, MA, CA, IL, NY, VA, and the country of Australia. | Terminated | January 03, 2011 | 5 | 77 | Not_Computer
Z-0836-2011 | 56024 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CATreatment Planning System is intended as aid in quickly and precisely planning the radiation treatment of small cranial lesions such as an asteriovenous malformations, pituitar tumors, pincalomas, acoustic neuromas and malignant neoplasms. | FastPlan Versions 5.5 and 5.5.1 | 2 | December 29, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | DESIGN: Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com. |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 422 | 218 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0833-2011 | 57196 | pump, portable, aspiration (manual or powered) | BTA | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Care E Vac 3 Portable Suction; Battery/AC Portable Aspirator; Ohio Medical Corporation, 1111 Lakeside Drive, Gurnee, IL 60031; Model 758000The device is a portable aspirator to help evacuate saliva, mucous, vomit or other aspirant from the mouth and or airway to allow adequate respiration or ventilation of the patient. | Care E Vac 3 Portable Suction | 2 | December 29, 2010 | 2010 | Ohio Medical Corporation | The D-14 diode on the circuit board may fail, resulting in a damaged, overheated circuit board. Also a screw used to attach one of the feet of the unit may damage the battery inside the device if the screw is over-tightened. |  | Ohio Medical sent Urgent Medical Device Correction Notification letters dated November 4, 2010, to all affected customers on November 8, 2010. The letters identified the product, the problems, and the action to be taken by the customer.Customers were instructed to take the affected product out of service, remove diode 14, and replace the foot screw contained therein. Customers were instructed to visit the micro-sites established for these repair procedures for step-by-step instructions on performing the repairs. Accounts with questions after viewing the video demonstration were directed to the Ohio Medical technical assistance line at 224-430-1522. Customers were asked to complete the attached Customer Acknowledgment Form and return the form to the Quality Assurance Department at (fax) 1.847.855.6304, or e-mail to:Complaints@OhioMedical.com, or mail to:Ohio Medical CorproationQuality Assurance1111 Lakeside DriveGurnee, Illinois 60031.For questions regarding this recall call 847-855-6270. | Contact the recalling firm for information | 4183 | 4,183 units | Worldwide Distribution - USA nationwide, Puerto Rico, and the coutries of Canada, Israel, Jordan, Kenya, Kuwait, Saudi Arabia, Taiwan, Thailand and Vietnam. | Terminated | March 16, 2012 | 443 | 4183 | Hardware | Physical Safety Hazards | Repair | Repair
Z-0832-2011 | 57254 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Urethral Warmer System Model number FPRCH3001The Urethral Warming System is designed to circulate a warm solution through a warming catheter to maintain urethral tissue near body temperatures while the surrounding prostate tissue is being frozen. It is intended to warm urethral tissue when performing cryogenic destruction of prostatic tissue with a Galil Medical Cryoablation System. | Urethral Warmer System | 2 | December 29, 2010 | 2010 | Galil Medical | This voluntary recall has been initiated because our supplier (Gaymar Industries, Inc.) notified us that these products contain power cords that may crack or fail at or inside the plug due to a prong design problem. Affected plugs can be identified by a round ground pin and black plastic bridges. Cords with plugs that are not affected have U-shaped ground pins and no plastic bridges. Although Gali |  | Gaymar sent an URGENT: MEDICAL DEVICE RECALL letter dated September 30, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: Customers were instructed to:" Inspect units in their possession to determine if they contain affected power cord plugs or if there were any signs of excessive wear or damage. " Immediately discontinue use and quarantine the affected product. " Contact Customer Service Department at 800-828-7341 to arrange for replacement. " Complete the attached form for each unit and follow the instructions given to return the form by fax or e-mail. " If customers further distributed the product to another person or entity, they were to identify their customers and notify them at once of this product recall by including a copy of the recall notification letter and the attached form. For questions regarding this recall call 800-828-7341 or 716-662-2551. | Contact the recalling firm for information | 4 | 4 | Worldwide Distribution - USA, including CA, GA, TX and the country of Canada. | Terminated | December 10, 2011 | 346 | 4 | Not_Computer
Z-0835-2011 | 57110 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Mammography KeypadRadiologists and Technologist apply hotkeys from a keyboard for ease of use in their clinical and diagnostic activities. The Agfa Mammography Keypad allows users the ability to apply hot keys from a keypad versus a keyboard. | Mammography Keypad | 3 | December 29, 2010 | 2010 | AGFA Corp. | Product did not receive UL Certification | OTHER/UNDETERMINED: Pending | AGFA sent an Urgent Safety Notice letter dated October 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to unplug existing keypad(s)  including USB cable. Plug in loaner Keypad(s). Print out the Fed Ex shipping label that they received via the e-mail notification. Use that label to ship their existing keypad(s) and their corresponding USB cables to the manufacturer. The manufacturer would perform the upgrades and return the keypad(s) along with the USB cables.Upon receipt of the original keypad(s) customers would also receive another e-mail notification with a shipping label so that they may return the loaner keypad(s) to AGFA HealthCare. Customers were to distribute this information to all those who need to be aware of it, and complete the feedback form as soon as possible and return it to AGFA HealthCare.For questions regarding this recall call 838-342-3989.. |  | 26 | 26 units | Worldlwide Distribution - USA including FL, GA, KY, ME, MD, MS, OH, UT and the country of Canada | Terminated | February 01, 2011 | 34 | 26 | Not_Computer
Z-0845-2011 | 57192 | device, neurovascular embolization | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Micrus Endovascular Microcoil Delivery System;The MicrusÂŽ Microcoil Delivery System consists of three components, a Microcoil System, a Connecting Cable, and a Detachment Control Box (DCB). Each component is sold separately. Catalog Numbers: PC4180933, SSR181128, CPL100252, PC4181647, SSR181640, CPL100404, DFS100254, SSR181230, CPL100406, DFS100410, DPL100206, SRC141230, CRC140304, DFS100716, CRC140406, CPL100203, DFS100204, DPL100408, SSR100253, PC4181240, DPL100202, CDF100408, PC4181137, CSP100500, CDF100154, PC4180830, CSP100350, DFS100152, SRC140408, PC4181034, CPL100206, CDF100152, SRC141025, SRC140225, CPL100304, CRC141025, CPL 100254, PC4181447, DPL100252, DFS100154, SRC140303, DFS100720, DPL 100204, PC4181343, CSP180500, DFS1 0031 0, DPL100306, DPL100256, CPL100306, CDF100304, CDF100203, CDF100410, PC4100412, DPL 100254, DPL 100053, PC41 00626, CDF100208, CRC140306, PC4100517, DFS100203, CSP181400, CPL 100152, DFS100210, CDF100720, CPL100204, CSP180200, DFS100204, CDF100510, DFS101025, DFS100408, PC41 00730, SSR181740, CRC140615, CPL100256, CDF100206, DFS100308, CPL100151, CSP100200, CDF100616, CRC140711, SSR181334, CDF100310, SRC140445, DFS100252, CPL100408, CDF100306, DFS100250, DPL 100152, DPL 100201, CDF100258, SRC140306, CDF100156, DPL 100153, DPL100304, SRC140922, SSR100825, DFS100206, DPL 100203, PC4181550, DFS100616, CSP180900, PC4181950, DFS100406, SRC140717, PC4181750, DFS100925, CSP100400, PC4181846, SRC140711, CPL100153, PC4182050, DFS100306, SRC140406, CSP100900, SRC140615, CDF100308, CSP100300, DFS1 0051 0, CDF100406, CDF1 0021 0, DFS100156, CDF100515, DPL100151, CDF100925, SSR181028, CDF101025, SSR181540, CPL 100201, SSR180825, CPL 100202; Manufactured and distributed by Micrus Endovascular Corporation, San Jose, CA | Micrus Endovascular Microcoil Delivery System; | 2 | January 03, 2011 | 2011 | Micrus Endovascular Corp., 821 Fox Lane, San Jose, CA 95131 | The firm has received reports of difficulty connecting to the detachment control cable. Loss of the electrical connection will result in failure to detach the coil for deployment inside an aneurysm. | OTHER/UNDETERMINED: Pending | Micrus Endovascular Corp sent Urgent: Medical Device Recall Letters dated October 21, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine inventory (including consignment) and quarantine the product to prevent it from being used or shipped. If the product was further distributed, customers were to notify their customers by sending a copy of the recall letter and were instructed to work with all customers to ensure that all product was returned and that pre-use protocol was followed.The response form was to be completed and returned indicating the number of devices that had been used (or shipped if distributor) and the number that would be returned. Customers were to call the Micrus Customer Service Department at 1-888-550-4129 from 8:00 am to 5:00 Pacific time for arrange for a Return Authorization number and replacement. Recall letters also stressed the importance of always following the pre-use protocol and inspecting the device immediately prior to use. Additionally, connector functionality should be verified and instructions were provided.Questions or concerns should be directed to Micrus Customer Service at 1-888-4120.Consignees who did not respond to the initial recall notification after 15 days were notified again by letter. A third letter was sent after another 15 days if there was still no response. |  | 6755 | 6755 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, DC, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, France, Germany, Hungary, Ireland, Italy, Japan, Netherlands, Poland, Republic of Korea, Russia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | January 11, 2011 | 8 | 6755 | Not_Computer
Z-0846-2011 | 54638 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The device is the NeuViz Dual series Computed Tomography (CT) Scanner System. The device is identified as NeuViz Dual, PN: 989605651321.Labeling on the device reads in part: "***Neusoft NeuViz Dual MULTI-SLICE CT SCANNER SYSTEM***MANUFACTURED BY PHILIPS AND NEUSOFT MEDICAL SYSTEMS CO., LTD. ADDRESS: NEUSOFT PARK, HUN NAN INDUSTRIAL AREA, SHENYANG 110179, CHINA***".Device is intended to produce cross-section images of head and body by computer reconstruction of x-ray transmission data taken at different angles. | NeuViz Dual Series Computed Tomography Scanner System | 2 | January 04, 2011 | 2011 | Philips And Neusoft Medical Systems Co., Ltd. | Potential for the CT patient table top to become detached from the carrier due to missing washers. | OTHER/UNDETERMINED: Pending | Neusoft Medical Systems Co., China sent an URGENT DEVICE CORRECTION letter dated Feb. 10, 2010 to their distributor in NC. On 02/12/10, The distributor sent the same letter to their consignees.The firm informed this distributor that there was report of the four screws that secure the patient table top to the carrier that did not have the proper washers in place which caused the table top to detach. If the issue would reoccur, there is a potential for the table top to fall and cause patient injury.Customers were instructed to check the patient table top fixing. If it is found that the the assembly is improper customers were instructed to top using the product and contact local Service personnel for repair. Any customer can contact Service Support Department, Neusoft Medical Systems Co, LTD at nms-service@neusoft.com or Service Support Department at helpdesk@neusoft.com for questions about this recall. |  | 26 | 26 devices | Nationwide Distribution including NC | Terminated | April 15, 2011 | 101 | 26 | Not_Computer
Z-0745-2011 | 56400 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed LX/l (2200997-2) | HiSpeed LX/l | 2 | January 04, 2011 | 2011 | GE Healthcare, LLC | A defect with the hardware and software displaying the incorrect dose data. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Important Electronic Product Radiation Warning letter was sent to customers identifying the affected product, explaining the defect and related hazards, giving instructions, and the product correction. Customers are to decrease the CTDI value by a factor of 30% when using 5, 7, 10 mm slice thickness, and by a factor of 20% when using 3 mm slice thickness for more accurate estimate of the CTDI. GE Healthcare will remedy the defect free of charge and provide an updated Operator Manual and configuration update to correct the CTDI display to all affected systems. Questions should be directed to the contact numbers provided in the letter. |  | 102 | 102 total in US | Nationwide Distribution | Open |  |  | 102 | Other
Z-0847-2011 | 56829 | camera, surgical and accessories | KQM | Surgical camera and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | The Power Cord Reel is labeled as follows:REF, VP6547, QTY1, POWER CORD REEL, Rx ONLY.These Video Carts are intended for storage of Endoscopic equipment that is used in single or multiple surgical specialtiesMANUFACTURED FOR: Linvatec, LINVATEC CORPORATION, LARGO, FL 33773 USA. | Linvatec Power Cord Reel Catalog Number: | 2 | January 04, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Endoscopy Division, dba. ConMed Linvatec, Largo, Florida is recalling five products: (1.) Video Cart Premium 120 VAC (VP6501), (2.) Video Cart Multi-Specialty 120 VAC (VP6502), (3.) Video Cart Articulating Arm 120V (VP6504), (4.) Video Cart (VP8500), and (5.) Power Cord Assembly Self Retracting (VP6547). These products are being recalled due to there is a possibility that the Vid | DESIGN: Component Design/Selection | ConMed Linvatec sent URGENT MEDICAL DEVICE CORRECTION NOTIFICATION letters to their customers by FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review their plugs for damage and determine if they were affected. Distributors were asked to contact affected customers and schedule delivery of their replacement plugs within 30 days. If the account no longer has the product in their possession, or if they have transferred the product, they were asked to respond as such via a Reply form.Customers were asked to document delivery of replacement plugs on the Sales Rep Reply Form and mail back or return by fax to 1-888-370-3062.For questions regarding this recall call the Customer Service Department at 1-888-292-0100 or e-mail Custserv1@Invatec.com. | No consumer action necessary | 77 | 77 ea | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Canada, Chile, Colombia, Denmark, Korea, Mexico, Singapore, South Africa, United Arab Emirates, and Venezuela. | Terminated | March 22, 2012 | 443 | 1761 | Not_Computer
Z-0842-2011 | 57271 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioScore, Inc. AngioSculpt PTA Scoring Balloon Catheter; OTW 0.018" Platform, 50 cm, 6.0mm x 20 mm balloon size, Part (REF) number: 2105-6020 | AngioSculpt PTA Scoring Balloon Catheter | 1 | January 04, 2011 | 2011 | AngioScore Inc. | Reports of distal bond failure which has resulted in peeling of the bond and/or detachment of the distal end of the scoring element. May result in vessel injury or embolism. | DESIGN: Component Design/Selection | Medical Device Recall notifications were sent to all domestic consignees on Nov 16, 2010 by overnight mail. Foreign distributors, other than Japan, were sent recall letters by e-mail or international overnight delivery on November 16, 2010. The Japanese distributor received notification and notified their customers on November 10, 2010. |  | 17682 | 17682 units, all sizes and lots | Worldwide Distribution: USA, China, Finland, Italy, Japan, Norway, Saudi Arabia, Spain, Sweden, Switzerland, UAE, United Kingdom | Terminated | October 04, 2011 | 273 | 17682 | Not_Computer
Z-0854-2011 | 57179 | clamp, uterine | HGC | Obstetric-gynecologic general manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Boston Scientific brand Matrix2 2D Standard SR Coils. Intended for embolization of certain intracranial aneurysms and for arterial and venous embolizations in the peripheral vasculature. | Boston Scientific brand Matrix2 2D Standard SR Coils. | 2 | January 05, 2011 | 2011 | Boston Scientific Corporation, Neurovascular, 47900 Bayside Parkway, Fremont, CA 94538 | The product appears to have premature degradation of the PGLA coating on the Matrix Coil. | PRODUCTION CONTROLS: Process Control | US consignees were notified by letter sent by Federal Express on October 19, 2010. Customers out of the United States were notified by Federal Express if possible and by personal delivery by locally assigned sales representatives as needed. |  | 77364 | 77,364 total units | Worldwide distribution. USA, Argentina, Austria, Australia, Belgium, Bahrain, Brazil, Canada, Switzerland, Chile, China, Colombia, Serbia, Czech Republic, Germany, Denmark, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Ireland, India, Israel, Italy, Jordan, Japan, South Korea, Lithuania, Latvia, Macau, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Taiwan, Turkey, Ukraine, Vietnam, and Yemen. | Terminated | January 27, 2012 | 387 | 77364 | Not_Computer
Z-0852-2011 | 57352 | stent, ureteral | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical, Universa Soft Ureteral Stent and positioner, 6 Fr x 24 cm, REF USH-624-R, sterile, Cook Urological, Inc., Spencer IN. Used for temporary internal drainage from the ureteropelvic junction to the bladder. | Universa Soft Ureteral Stent | 2 | January 05, 2011 | 2011 | Vance Products Incorporated | The firm received a customer complaint that a box labeled as Universa Firm Stents but the box actually contained Universa Soft Stents. The inner product label is correct but the outer box is wrong. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm sent URGENT: RECALL NOTICES dated 10/18/2010 to its direct accounts. The notice identified the affected product and instructed firms to pull the recalled product from stock and isolate it. Customers are to complete the form and e-mail or fax the form back to Cook. Product should be returned to the recalling firm using the Return Goods Authorization number. Cook Urological will credit customer accounts upon receipt of the returned product. | No consumer action necessary | 25 | 25 units | Worldwide Distribution -- USA, including states of WI, TX, NY, SC, AZ, TN, FL, AR, CA, AL, and WA and country of Canada. | Terminated | July 27, 2011 | 203 | 25 | Not_Computer
Z-0864-2011 | 55541 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Pioneer Plus Catheter PPlus 120 dual lumen device, inserted through a 6F introducer sheath and placed percutaneously into a peripheral vessel. The device tracks to its intended site in the vasculature over a standard length, commercially available 0.014" (0.36 mm) Rapid Exchange (RX) tracking guide wire. It utilizes an extendable, hollow Nitinol guide tip (needle) to facilitate the redirection and placement of a 0.014 (0.36 mm) Over the Wire (OTW) guide wire into peripheral vessels. Specification Developer: Medtronic, Minneapolis, MN. Manufactured by: Accellent, Laconia NHThe Pioneer Plus Catheter is intended to facilitate placement and positioning of catheters within the peripheral vasculature. The Pioneer Plus Catheter also provides an intra-luminal cross-sectional ultrasound image of the area of interest to facilitate placement of guide wires beyond stenotic lesions (e.g., sub-total, total or chronic total occlusions) prior to additional intervention (i.e. PTCA, stent, etc.). The Pioneer Plus Catheter is not indicated for use in the coronary or cerebral vasculature. | Pioneer Plus Catheter PPlus 120 | 2 | January 06, 2011 | 2011 | Medtronic CardioVascular | Needles may fail to retract and thus cause injury. |  | Medtronic sent an Urgent Medical Device Safety Information letter to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Sales Representatives were to contact each account and complete an acknowledgement form indicating the Physicians had been notified. For questions regarding this recall customers were to contact their local representative. |  | 1080 | 1,080 units | Worldwide Distribution - USA, Germany, Netherlands, Slovakia, and Spain | Terminated | January 06, 2011 | 0 | 1080 | Not_Computer
Z-0865-2011 | 57457 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Groshong 3 Fr. Single-Lumen PICC Intermediate Tray with Safety Introducer, Reorder Number 7717300, Bard Access Systems, Inc., Salt Lake City, UT 84116.The Groshong is a family of single and dual -lumen peripherally placed mid-line and central catheters made from specially formulated and processed medical grade materials, in a tray with accessories for reliable long (greater than 30 days) or short term (less than 30 days) vascular access. | Groshong PICC | 2 | January 06, 2011 | 2011 | Bard Access Systems | Prepackaged sterile latex gloves beyond their expiration date were packaged in surgical kits. | TRAINING: Employee Error | BARD Access Systems sent an URGENT: GROSHON* 3 FR SINGLE-LUMEN PICC INTERMEDIATE TRAY WITH SAFETY INTRODUCER RECALL letter dated December 10, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the product subject to recall, discontinue use, and cease further distribution of any the affected product. If the product was further distributed, those customers were to be notified immediately with a copy of the recall notification letter. Any remaining product was to be returned to Bard. For information regarding this recall call 800-290-1689. | Contact the recalling firm for information | 195 | 195 units | Nationwide Distribution including AL, AZ, CA, GA, IA, MN, MO, NE, NJ, NY, OH, OR, SD, and WI. | Terminated | May 04, 2011 | 118 | 195 | Not_Computer
Z-0883-2011 | 57416 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Occluder module, catalog 803480, Terumo Cardiovascular System Corp, Ann Arbor, MI.The occluder module is used with the Occluder head to provide a computer controlled tube clamping mechanism to regulate the flow in the venous line. | Terumo Advanced Perfusion System 1 | 2 | January 07, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm identified that a component within the System 1 ABD and Occluder Modules that may malfunction. The Air Bubble Detector Module (ABD) is used with an air sensor to detect gross air bubbles in the perfusion circuit. The Occluder module is used with the Occluder head to provide a computer controlled tube clamping mechanism to regulate the flow in the venous line.The firm identified a fai |  | Terumo Cardiovascular Systems Corporation issued an URGENT MEDICAL DEVICE CORRECTION letter dated 12/2/2010 to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager. The letter identified the product, the problem, and the action to be taken. The letter stated the firm will replace all modules in the affected population, and that a representative would contact the consignee regarding the replacement schedule.Customers were to review the Medical Device Correction.Assure that all user were aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call 1-800-521-2818. | Contact the recalling firm for information | 221 | 221 devices | Worldwide Distribution - US, Austrailia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Guatamala, Hong Kong, Indonesia, Japan, Korea, Malaysia, Mexico, Phillipines, Saudi Arabia, Singapore, Taiwan, Thailand, United Arab Emirates, and Vietnam | Terminated | April 17, 2012 | 466 | 729 | Not_Computer
Z-0867-2011 | 54651 | system, digital image communications, radiological | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Stryker Digital Capture Pro 2 device (240-050-800, CD version) and (240-050-810, DVD version) designed for digital documentation of medical procedures. Manufactured by Stryker Endoscopy, San Jose, CA. | Stryker Digital Capture Pro 2 | 2 | January 07, 2011 | 2011 | Stryker Endoscopy | Device will not able to assign the correct time/date stamp for saved filed after December 31, 2009. On Jan 1, 2010, the device system date for all SDC Pro 2 units will re-start from 1 Oct 2001 and begin incrementing from that point onwards. The system date will be incorrect and users will not be able to correct this issue. Problem will not affect how pictures or videos are captured using SDC Pro | DESIGN: Software Design | Stryker Endocoscopy sent an URGENT: Device Correction letter dated December 11, 2009, to the OR Supervisor/Risk Manager and all Stryker Sales Representatives. The letter identified the product, the problem, and the action to be taken by the Sales Representatives. Sales Representatives were to identify affected accounts. Go to customer accounts and install the upgrade for each affected unit. Upon completion of the upgrade, the sales rep was to sign the acknowledgment form indicating that all of their affected accounts received the appropriate software fix and mail to Stryker Endoscopy or fax both sides of the postcard to 408-754-8378 or scan in both sides of the postcard and email to sdcpro2recall@stryker.com.For questions regarding this recall call 408-754-2738 or 408-754-2000. |  | 3070 | 3070 units | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL IN, IA, KS, KY, LA, MA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Africa, Australia, Brazil, Cnanda, China, France, Germany, Hong Kong, Iberia, Italy, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Puerto Rico, Scandinavia, Singapore, Switzerland, Thailand, and the UK | Terminated | January 08, 2011 | 1 | 3070 | Software | Output/Calculation | Software update | Software Update
Z-0885-2011 | 54685 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Guider Softip TM XF Guide CatheterCatheter is intended to facilitate the placement of interventional devices into the neurovascular system. | Guide Catheter | 2 | January 07, 2011 | 2011 | Boston Scientific Corporation | Complaint about product quality: degradation of polymer portion of the distal section of the catheter. | MISBRANDING: Labeling False and Misleading | Boston Scientific sent an Urgent Medical Device Recall and Correction letter dated December 18, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and fill out the Customer Instructions and REPLY VERIFICATION TRACKING FORM indicating the status of the affected product and send a copy of the form by fax to 510-624-1600 or e-mail to Joe.Nesci@beci.com. | Contact the recalling firm for information | 94296 | 94,296 units | Worldwide. US Domestic including US Federal Government, and foreign countries. | Terminated | December 19, 2011 | 346 | 94296 | Not_Computer
Z-0871-2011 | 56388 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Synchro2 0.14" Guide Wire, 300 cm, Soft Pre Shaped;Boston Scientific Corporation: Made in 2405 West Orton Circle, West Valley City, UT 84119Product is distributed by Boston Scientific Corporation, Fremont, CA The Synchro guidewire series is intended for general intravascular use, including neurovascular and peripheral vasculatures. It can be used to selectively introduce and position catheters and other interventional devices within the peripheral and neurovasculature. This device should be used only by physicians trained in percutaneous intravascular techniques and procedures. | Boston Scientific Synchro 2 Guidewire | 2 | January 07, 2011 | 2011 | Boston Scientific Corporation | The Guidewire has the potential to flake off polytetrafluroethylene (PTFE) particulate, which has the potential to embolize and result in infarction when introduced into a patient's circulatory system. | PRODUCTION CONTROLS: Packaging Process Control | Boston Scientific Corporation sent an Urgent Medical Device Recall letter dated June 7, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of and segregate the affected product. Complete and return the Reply Verification Tracking Form even if they did not have any affected product. Contact the Customer Service Call Center to obtain a Returned Goods Authorization (RGA) Number at (800) 811-3211, Option #1. E-mail or Fax the completed and signed Reply Verification Tracking Form to email Bob.Ware@bsci.com or the Field Action Fax (510) 624-2475.PackageShip and return the affected product along with the original Reply Verification Tracking Form by Federal Express to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: ____________For questions regarding this recall call (510) 624-2563. |  | 49 | 36, 049 total units - all sizes | Worldwide Distribution - USA including CA, CO, IN, MD, MI, OR, TX, and VA and the countries of Australia, Austria, Belgium, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Great Britian, Greece, Honduras, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Korea, Sweden, Switzerland, Thailand, Tunisia, Turkey, and Ukraine. | Terminated | January 10, 2011 | 3 | 36049 | Not_Computer
Z-0866-2011 | 57418 | catheter, irrigation | GBX | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Merit One Step Centesis Drainage CatheterIntended for the percutaneous drainage of fluids | Merit One Step Centesis Drainage Catheter | 2 | January 07, 2011 | 2011 | Merit Medical Systems, Inc. | The catheter's valved hub may fail resulting in an inability to drain fluid. | DESIGN: Device Design | Merit Medical Systems, Inc. sent a Product Safety Notice letter dated November 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to provide a copy of the Safety Notice form to their applicable customers. Advise them to immediately stop using the affected product and to discard it.Customers were asked to complete and sign the Safety Notice forms and fax them to 1-801-208-3378 and send the original Safety Notice Form with both signatures via Federal Express 2nd day air to:Customer ServiceMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416entirety and faxed them For questions regarding this recall call 801-316-4929 or Customer Service at 801-208-4365. |  | 20863 | 20,863 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, AND WV and the countries of Canada, Finland, Germany, Netherlands, Singapore, Sweden, Switzerland, and the UK | Terminated | March 23, 2012 | 441 | 20863 | Not_Computer
Z-0884-2011 | 56986 | cable, transducer and electrode, patient, (including connector) | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 System Interface Cable, 34-pin to 10-pin, Red, 10' - Part Number: D-1286-18; Catalog Number CR3410CTThis product provides a signal interface between a Biosense Webster electrophysiology catheter and the appropriate equipment. This cable may be re-used subject to established cleaning and sterilization restrictions | CARTO 3 System Interface Cable | 2 | January 07, 2011 | 2011 | Biosense Webster, Inc. | When connecting this cable to the mapping system, loss of Body Surface ECG signals may occur. This may result in inaccurate locations appearing for catheters displayed 3D on the CARTO monitor. | DESIGN: Device Design | Biosense Webster issued an Urgent Field Safety Notice letter dated October 5, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to return the affected cable, accompanied by the attached Return Authorization form.For questions regarding this recall call Biosense Webster at 909 839-8874. | Contact the recalling firm for information | 1692 | 1692 | Worldwide Distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, MA, MD, MI, MN, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA, WI, and WV, and the countries of Belgium, China, Israel, Japan, Singapore, and Thailand. | Terminated | May 07, 2012 | 486 | 1692 | Not_Computer
Z-0886-2011 | 56347 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Varian brand C-series Clinac, Trilogy and Novalis Tx, Software Versions 6.X and 7.X, Model Numbers: H14, H27, H29, HCX, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Trilogy Radiotherapy Delivery System is a radiation therapy accelerator intended to deliver megavoltage x-ray treatments for conventional radiotherapy (three dimensional conformal radiotherapy and intensity modulated radiotherapy) and stereotactic radiosurgery and radiotherapy. Stereotactic treatments are intended for therapy of lesions, e.g., arteriovenous malformations, primary tumors and metastases. Stereotactic treatments may be intracranial or extra cranial and consist of single-session or fractionated delivery. | Varian brand Cseries Clinac, Trilogy and Novalis Tx, Software Versions 6.X and 7.X, | 2 | January 08, 2011 | 2011 | Varian Medical Systems Oncology Systems | The Auto goto or Auto setup functions in C-Series version 7 software ignore the couch angle and exceed motion zone (unless a tolerance is defined by the user); it has the potential for collision with the patient on the couch. | OTHER/UNDETERMINED: Pending | Varian sent a Customer Technical Bulletin letter to all affected users on May 27, 2010, and was also posted to Varian's customer support site: http://www.MyVarian.com. The letter identified the product, the problem, and provided instructions on the recall. Customers were instructed to return the proof of notification (via completed "Proof of Notification Form or appropriate Postal system certified receipt of delivery to Varian. For questions regarding this recall call 650-483-3153. |  |  |  | Worldwide Distribution - USA including CA, TX, IL, GA, NJ, and NV and the countries of Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland, and the UK | Terminated | December 23, 2011 | 349 |  | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0888-2011 | 51217 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Floorstand. An accessory of the Optical Guidance Platform, Model number HZ5 for use with a charged particle accelerator (such as the Clinac 600C). | Optical Guidance Platform | 2 | January 10, 2011 | 2011 | Varian Medical Systems Oncology Systems | Potential patient impact -- limit switches for Floorstand may be inadequate for some models and allow for potential impact. | DESIGN: Device Design | Varian sent an urgent device correction notification letter 01/26/2009 to all affected users, with a description of the problem and user corrective action steps via certified mail. The letter will also be distributed to Varian Sales, Marketing and Service organizations. | Contact the recalling firm for information | 3 | 3 units | Product was distributed to 3 consignees in Florida and Wisconsin. | Terminated | February 24, 2012 | 410 | 3 | Not_Computer
Z-0891-2011 | 54146 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console, versions: 8.0.12, 8.0.23 SP2, 8.1 SP2, 8.1.1.3, 8.1.1.4, 8.1.2, 8.1.3, 8.1.4, 8.3, 8.6, 8.6.1, 8.6.2, 8.8.0; Model Number: H51; Distributed and manufactured by: Varian Medical Systems Inc., Palo Alto, CA Intended use: The 4D Integrated Treatment Console function is designed to assist the operator of radiation therapy device in providing accurate treatment setups for each patient by monitoring setup parameters and preventing the radiation therapy device from commencing irradiation when any parameter is out of conformance with the treatment plan. | 4D Integrated Treatment Console | 2 | January 10, 2011 | 2011 | Varian Medical Systems Oncology Systems | Software anomaly may result in a design deficiency causing Plan Edit Settings to malfunction under specific circumstances, where the user selects to use a Cone Beam CT set up field. | DESIGN: Device Design | The firm, Varian Medical Systems, sent a Product Notification Letter/Field Safety Letter entitled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2009, to all affected customers. The letter described the product, problem, actions to be taken by the customers and Varian. The customers were instructed to use caution when creating CBCT imaging field and limit use of the 'select all' button when acquiring treatment parameters; review departmental policies and procedures regarding the acquisition of any gantry angle at the 4D ITC, and advise the appropriate personnel of this letter. The firm, Varian, recommended that only vertical, longitudinal and lateral couch parameters are acquired when creating a CBCT field at the 4D ITC; notified all affected customers of this document, and will correct this issue in a future release of the 4D ITC software. If owners of $D ITC, version 8.0.12 and higher, have any questions regarding this product notification letter, please contact Varian Medical Systems Help Desk at 1-888-VARIANS (888-827-4265) for USA and Canada, +41 41 749 8844 or support-emea@varian.com for Europe;Emails for North American: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com and Latin American: soporte.al@varian.com, and Internets for Oncology Systems- www.myvarian.com and Varian Medical Systems- www.varian.com. | Contact the recalling firm for information | 3629 | 3629 | Worldwide distribution: USA and countries including: Australia, Asia, China, Japan, Latin American, New Zealand , North America, South East Asia. | Terminated | July 20, 2011 | 191 | 3629 | Software | Device Operation | Software update | Software Update
Z-0890-2011 | 56326 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System radiation therapy device, manufactured by Accuray Inc., Sunnyvale, CA.Medical device indicated for treatment planning and image guided stereotactic radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | CyberKnife Treatment Planning System | 2 | January 10, 2011 | 2011 | Accuray Inc | Latches responsible for securing the cover to the device may come lose if not properly secured and the cover may unexpectedly drop off. A design change is planned. |  | Accuray sent an "URGENT DEVICE CORRECTION" letter dated July 6, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to assure the LINAC cover is secure following reattachment and adhere to and follow the instructions included in the letter. Note: A new design is being developed, and will be distributed to user base once available.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 207 | 207 units | Worldwide distribution: US including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV, and countries including: Japan, China, Vietnam, Taiwan/China, South Korea, Hong Kong/China, Thailand and Malaysia. | Open |  |  | 207 | Not_Computer
Z-0887-2011 | 57338 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TiTLE 2 Polyaxial Spinal System, Counter Torque Assembly Tool, Part ET1132-01, Zimmer, Warsaw, IN.The Zimmer Title 2 counter torque assembly tool is a reuseable, non- implantable surgical tool that is provided as part of a the device system for use as a secondary option for compression and distraction of the construct assembly and initial tightening of set screws. The instrument is designed to provide equal torque in the opposite direction to the torque that is being applied to the cap screw during tightening. | Zimmer TiTLE2 Polyaxial Spinal System | 2 | January 10, 2011 | 2011 | Zimmer Inc. | There is a chance that the ball plunger could fall out during surgery and fall into the surgical site. There could be a delay in surgery to remove the ball plunger. If the problem was undetected, the ball plunger could remain in the patient causing pain and potentially necessitating subsequent surgery for removal. | OTHER/UNDETERMINED: Pending | Zimmer sent an URGENT MEDICAL DEVICE RECALL LETTER dated October 21, 2010, to all affected customers. The letter identified the product, the problem, the action to be taken by the customer.Customers were instructed to:Inspect their inventory, stop using the device, and place it in quarantine. Contact Zimmer Customer Service at 1-800-774-6368 to obtain a return authorization number. Complete the Inventory Return Certification form and fax a copy to (952) 837-6958, or scan and email to qualitycompliance@zimmer.com.Return the recalled product along with the original completed Return Certification Form to:Zimmer Spine, Inc.Attn: RA Receiving5301 Riata Park Ct.Austin, Texas 78727For questions regarding this recall call Zimmer Spine Customer Service at 1-866-774-6368. . |  | 68 | 68 | Nationwide Distribution including AL, AZ, CA, CO, FL, MI, MN, NC, PA, OH, OK, TN, TX, WA, and WI. | Terminated | October 12, 2011 | 275 | 68 | Not_Computer
Z-0843-2011 | 57293 | set, tubing, blood, with and without anti-regurgitation valve | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical CombiSet BVM Hemodialysis Blood Tubing SetP/N: 03-2695-9 | Fresenius Medical CombiSet BVM Hemodialysis Blood Tubing Set | 1 | January 10, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Hemodialysis Blood tubing Sets may develop kinking of the arterial line. |  | Fresenius Medical issued recall notification dated 11/19/10, "URGENT RECALL FMCNA CombisetÂŽ BVM Hemodialysis Blood Tubing Sets (Part Numbers 03-2695-9 and 03-27957)", by certified mail with return receipt. Customers have been instructed to check their stock immediately to determine if they have any affected lots on hand. If customers have the affected lots, they have been instructed to discontinue use immediately and place the product in a secure segregated area for return to Fresenius Medical Care North America. Customers have been instructed to contact Fresenius Medical Care Customer Service Team Representative for instructions on how to return the recalled product. |  | 24 | 190,080 units (7920 cases x 24 units) | Distribution Nationwide and Canada | Terminated | November 03, 2012 | 663 | 56 | Not_Computer
Z-0889-2011 | 57414 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | TandemHeart Transseptal Cannulas-EF set (THTC-EF) - Dilator assemblyThe TandemHeart Transseptal Cannula Set-EF is intended for transseptal catheterization of the left atrium via the femoral vein for the purpose of providing a means for temporary (six hours or less) left ventricular bypass when connected to the TandemHeart extracorporeal blood pump which returns blood to the patient via the femoral artery or other appropriate site. | TandemHeart Transseptal CannulasEF set (THTCEF) Dilator assembly | 2 | January 10, 2011 | 2011 | Cardiac Assist, Inc | Component (Dilator of TandemHeart Transseptal Cannula Set-EF) does not meet specifications | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CardiacAssist, Inc. sent a Voluntary Recall letter dated November 12, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed to examine inventory for the affected product and remove from inventory and identify it as a recalled product by applying the enclosed "CAUTION: DO NOT USE" sticker(s) and quarantine to prevent usage.Consignees were to complete the form included with the letter and fax a copy of the form to (412) 963-9739 to coordinate the return and replacement of any affected product. For questions regarding this recall call (412) 963-7770, ext. 271. |  | 164 | 164 | Nationwide distribution including AZ, CA, CO, CT, DC, FL, GA, IL, LA, MA, MD, MI, MN, NC, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | December 01, 2011 | 325 | 164 | Not_Computer
Z-0837-2011 | 57099 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook 400ES Safety Infusion System, Model Number 621-400ES. For use with B. Braun Medical Inc. Horizon Pump IV Sets | B. Braun Outlook | 1 | January 10, 2011 | 2011 | B Braun Medical, Inc | Possible pump stop with audible alarm, but no visual error indication that the pump is not delivering the medication. Under certain conditions in a network environment, Management Processor may become non-responsive, ceasing normal operation. | DESIGN: Component Design/Selection | B. Braun Medical Inc. notified consignee by letter on August 24, 2010. |  | 2107 | 2,107 Pumps (582 Pumps to Winchester) | Virginia and New York | Open |  |  | 2107 | Software | Alarm/Message | Notification | Safety Notice/Insructions
Z-0853-2011 | 53047 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, model number 025007, a subsystem of the Cyberknife Robotic Radiosurgery System, manufactured by Accuray Incorporated, Sunnyvale, CA.Intended for use in the support and positioning of a patient during Radiosurgery and radiotherapy procedures and other medical procedures when precise positioning is required. | RoboCouch Patient Support System, a subsystem of the Cyberknife Robotic Radiosurgery System | 2 | January 11, 2011 | 2011 | Accuray Inc | Mechanical Failure-- Patient Support System's table top may unexpectedly descend several inches. | DESIGN: Device Design | Consignees were notified verbally about the issue when an inspection was scheduled. |  | 26 | 26 | Worldwide Distribution -- USA, including states of CA, FL, VA, AL, CT, MA, NC, TN, PA, RI, and NE and countries of Turkey, France, UK, Saudi Arabia, and Switzerland. | Terminated | January 12, 2011 | 1 | 26 | Not_Computer
Z-0892-2011 | 56350 | [6 empty] | Varian brand C-Series Ximatron Radiotherapy Treatment Stimulator,Model H72, Product is manufactured and distributed by Varian Oncology Systems, Palo Alto, CA The product is indicated for use in radiation therapy simulation, using a fluoroscopic and/or radiographic x-ray system for visualizing the volume to be exposed during radiation therapy. | Varian brand CSeries Ximatron Radiotherapy Treatment Stimulator, | 2 | January 11, 2011 | 2011 | Varian Medical Systems Oncology Systems | There is a potential for failure of a lead screw drive system or image intensifier lead screw drive assembly, if normal maintenance inspections are not followed correctly. | OTHER/UNDETERMINED: Pending | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail receipt on January 19, 2010, to all direct consignees. The letter identified the product, the problem, and the action the consignee should take.Consignees were instructed to have maintenance personnel perform the recommended measurements as described in the Varian Customer Technical Bulletin.Ensure that all measurements made as part of the procedure are recorded in the 'wear table' that is part of the procedure.Previous wear measurements should be compared to new measurements and any large increase in wear should be investigated immediately.Pay special attention to the "Recommendations based on wear measurements" table.Ensure that new wear measurements comply with the operating parameters described in the table. For any questions consignees are to contact Varian Medical Systems Customer Support Services at one of the locations described in the CTB.For questions regarding this recall call (650) 424-5731. |  |  |  | Worldwide Distribution including USA, Australia, New Zealand, Europe, Southeast Asia, China, Japan, Latin America, and Canada. | Terminated | February 24, 2012 | 409 |  | Not_Computer
Z-0894-2011 | 57463 | phlebotomy blood collection kit | OFM | Impedance phlebograph. | Cardiovascular | Cardiovascular | Enforcement Discretion | Blood Center Kit w/Chloraprep One-Step, Sterile/Disposable/Single Patient Use, Reorder No. Q921CL, Manufactured for : Blood Center of Wisconsin, Milwaukee, WI. Port VAD Kit w/Chloraprep One-Step, Sterile/Disposable/Single Patient Use, Reorder No. 3540CL, Manufactured for : McCormick Medical Distribution, Edmonds, WA 98020.The kits are used by blood centers to draw blood. | Monoject Prefill Advanced | 2 | January 11, 2011 | 2011 | Continental Medical Labs,Inc | Continental Medical labs, Inc, is recalling various lots of Blood Center Kit and Port VAD Kits because Covidien reported that testing could not pick up an Over-Sulfated Chondroitin Sulfate (OSCS) contaminant in the purified heparin sodium USP. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to their distributors. Continental Medical Labs,Inc telephoned on 12/06/2010 and later sent a letter dated December 7, 2010, with an enclosed copy of the recall notice. The letters were addressed to McCormick Medical Distribution, Inc and to Medico Mart, Inc. The letter described the problem and product affected. Advised consignees to inform all their customers to return the product as well as the distribution information as soon as possible so CML could notify Covidien for replacements.For questions regarding this recall call (262) 723-6715, ext. 313. | No consumer action necessary | 30 | 30 cases, 30 kits per case | Nationwide Distribution including WA and WI | Terminated | July 05, 2012 | 541 | 30 | Not_Computer
Z-0899-2011 | 51401 | catheter, ventricular, general & plastic surgery | GBS | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Concentric 7F Balloon Guide Catheter (percutaneous catheter per 21 CFR 870.1250), REF 90072, manufactured by Concentric Medical, Inc., Mountain View, CAProduct is indicated for use in facilitating the insertion and guidance of an intravascular catheter into a selected blood vessel in the peripheral and neurovascular systems. The balloon provides temporary vascular occlusion during these and other angiographic procedures. Not intended for use to deliver embolic agents or to perform angioplasty. | Concentric 7F Balloon Guide Catheter | 2 | January 12, 2011 | 2011 | Concentric Medical Inc | Product malfunction-- hub may leak at connection to the shaft due to insufficient adhesive curing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Concentric Medical products sent an URGENT VOLUNTARY MEDICAL DEVICE REMOVAL letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and identify any affected product.Quarantine the product and call Concentric Medical Customer Service at 877-471-0076 (USA toll-free) or 650-938-2100 to receive instructions for returning the devices.For questions regarding this recall call 877-471-0076. |  | 36 | 36 units affected. | Worldwide Distribution - USA including AL, FL, GA, IL, MN, NY, OH, OR, PA, TX, VA, WA, and WV and one consignee in Turkey and one in Switzerland | Terminated | January 13, 2011 | 1 | 36 | Not_Computer
Z-0897-2011 | 51274 | table, powered | INQ | Powered table. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precedence SPECT/CT System, Precedence 16 3/8, Upper Patient Pallet, A single photon nuclear medicine and image from an X-ray computed tomography system,Model Number: 4535 602 50851, Catalog Number: 882350,Product is manufactured by Philips Medical Systems (Cleveland), Inc. Cleveland, OH Intended use: A single photon nuclear medicine and image from an X-ray computed tomography system. | Precedence SPECT/CT System Upper Patient Pallet | 2 | January 12, 2011 | 2011 | Phillips Nuclear Medicine | A problem has been identified with some upper patient pallets on Precedence SPECT/CT System, which may reduce its load bearing capacity and may result in its cracking. If this were to happens, could pose a risk to patient. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, Philips Healthcare, sent an "Urgent - Field Safety Notice" dated January 26, 2009 to all consignees/customers. The notice described the product, problem and action to be taken by customers and by Philips. The customers were instructed to reference their system's System Operation and System information Manuals regarding loading and unloading of patient; monitoring the patient during a study ,and careful monitoring during unsupported (cantilevered) table translation. Philips has ordered replacement parts and they will have a Philip Healthcare Field Service Engineer contact and visit the customers site within two months to replace the affected upper patient pallet.If you need any further information or support concerning this issue, please contact your local Philips representative at: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: and Oncology or your local Philips Healthcare office. |  | 2 | 2 units | International distribution only: Mexico and Saudi Arabia. | Terminated | January 13, 2011 | 1 | 2 | Not_Computer
Z-0896-2011 | 54718 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound SystemDiagnostic Ultrasound system | Ultrasound System | 2 | January 12, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Software safety guard measures on the device malfunctioned to prevent the temperature from going over its limits when the device is used. | DESIGN: Software Design | Field notification in the form of a Customer Safety Advisory Notice was initiated on 09/30/2009 by the firm, and the notice was mailed to all affected customersSiemens sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter on September 30, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconnect and reconnect the V5Ms transducer between uses until their local Customer Service Engineer installs the software update to their system. For any questions customers were to contact their local service support person.For questions regarding this recall call 650-694-5398. |  | 452 | 452 | Worldwide Distribution - United States (nationwide) and the countries of Australia, Austria, Azerbaijan, Brazil, Belarus, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Guatemala, Croatia, India, Indonesia, Italy, Jordan, Japan, Latvia, Malaysia, Mexico, Moldova, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey, Ukraine, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Terminated | June 18, 2012 | 523 | 452 | Software | Physical Safety Hazards | Instructions on reconnecting the transducer | Safety Notice/Insructions
Z-0898-2011 | 57218 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow-Howes (TM) Multi Lumen Central Venous Catheterization Set with Blue FlexTip(R) Catheter. Catalog #: CV-15703.Permits venous access to the central circulation. | ArrowHowes (TM) Multi Lumen Central Venous Catheterization Set with Blue FlexTip(R) Catheter | 2 | January 12, 2011 | 2011 | Arrow International Inc | One lot of product was inadvertently mislabeled with a product lidstock identifying the product as a triple-lumen catheter with product number CV-15703, but a double-lumen catheter with part number CV-16702 was actually packaged in the kit. | PRODUCTION CONTROLS: Labeling Mix-Ups | The recalling firm issued Urgent Medical Recall letters dated 11/9/10 to inform their customers of the problem and the need to return any unused product. Customers are to check their stock for the affected product, contact Arrow Customer Service for a Return Authorization Number, and complete the enclosed Recall Acknowledgement & Stock Status Form. For customers with no affected product in stock, the form should still be completed and also faxed to Arrow Customer Service. Affected product should be returned along with the original completed form to Arrow International, Inc. Arrow will replace the affected product once it has been returned. Additional information or clarification can be obtained from Arrow Customer Service at 1-919-361-3964. |  | 1150 | 1150 units | International Distribution Only -- Argentina and Columbia. | Terminated | September 29, 2012 | 626 | 1150 | Not_Computer
Z-0895-2011 | 57462 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Disc Space DistractorLabel states - Specialty Disc Space Distractor, Mfg by: Stryker Spine 2 Pearl Court Allendale, NJ 07401-1677Used to assist in spinal fusions. It is to be placed between the vertebral bodies in the specified disc space; the t-handle can then be rotated clockwise expanding the distal paddle assembly to a specified height. Once the determined height has been achieved the t-handle can be rotated counter-clockwise retracting the paddles and then removed from the defect site. instrument is not used to sustain or support life. it is intended to be used as a manual surgical instrument during inter body disk replacement to distract the vertebrae to aid in removal of disc space, cartilage, bone and soft tissue. | IBD Disc Space Distractor | 2 | January 12, 2011 | 2011 | Stryker Spine | Stryker Spine has received reports of the Disc Space Distractors failing during surgical procedures. |  | Stryker Spine sent an URGENT PRODUCT RECALL letter dated December 7, 2010, via Fedex to all affected customers.. The letter identified the product, the problem, and the action they should take. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile IBD Disc Space Distractors with the catalog and lot numbers referenced. Complete the Customer Response Form and distribution sheet and fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance,Stryker Spine 6 Pearl CourtAllendale, New Jersey 07401For questions regarding this recall call 201-760-8206. | Contact the recalling firm for information | 21 | 21 units | Nationwide Distribution including AZ, FL, MA, NY, PA, and TX | Terminated | November 14, 2012 | 672 | 21 | Not_Computer
Z-0900-2011 | 57436 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology ECG ManagementThis is a software application designed for use in various hospital departments to import, display, store, analyze, distribute, and process ECG procedures from resting ECG devices. It is also intended to provide analysis or reanalysis or testings ECGs and to provide preliminary data for editing and confirmation by an over-reading physician. Horizon Cardiology ECG Management may provide a serial comparison of ECG Data to facilitate the review of the patient's current ECG with previous ECG's of the same patient. | Horizon Cardiology ECG Management | 2 | January 13, 2011 | 2011 | McKesson Provider Technologies - Medical Imaging Group | When importing ECG files from Mortara ECG carts using DICOM interface, the Horizon Cardiology ECG Management System does not display the amplitudes of the waveform properly. | DESIGN: Device Design | McKesson Medical sent an advisory notice letter dated June 16, 2010, to all affected customers. Advisory notice ADV01646 was distributed to 13 affected clinical sites from August 30, 2010 to September 28, 2010. The letters identified the product that is affected, the circumstances under which patient safety issue may occur, the risk, and the immediate actions that need to be taken along with a temporary fix. Customers were instructed to use the ECG cart printout to verify results of procedures acquired by Mortara ECG carts. Users should be aware of the problem when referrring to past ECG procedures and when performing a visual comparison with ECG procedures acquired by Mortara ECG carts.Customers should disable the Glasgow Serial Comparison for ECG procedures.McKesson Support would contact the customer to coordinate the deployment of the software update on all the affected applications and Horizon Cardiology client applications. For any questions regarding this recall, customers should call the Customer Support department at 1-866-777-0202 and speak to their Support Manager. |  | 16 | 16 units | Nationwide Distribution including AZ, CA, FL, GA, MA, PA, UT, and WI. | Terminated | August 17, 2012 | 582 | 16 | Software | Display/Image | Software update | Software Update
Z-0904-2011 | 52716 | microscope, surgical | EPT | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | VISU 200 on S8 Floor Stand, Manufactured by Carl Zeiss Surgical GmbH, GermanyMedical device: Surgical microscope used during ophthalmic surgeries, which allows real time visualization of the surgical area. Product Quantity Distributed (Int) 2540 units | VISU 200 | 2 | January 14, 2011 | 2011 | Carl Zeiss Meditec Inc | Potential Retina Damage-- An error in the user manual may result in misapplied filter and injury during treatment. | PRODUCTION CONTROLS: Labeling Mix-Ups | Carl Zeiss Meditec Inc sent a SAFETY ALERT letter dated May 18, 2009, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded to use the correct retinal protection filter (GG475) when using the OPMI VISU 200 system during ophthalmic surgery.Verify the position of the retinal protection filter and inform their staff of the customizable feature. Verify that the correct filter was in place. Customers were asked to complete the acknowledgement form indicating receipt of the alert. Sign and return the form to:Carl Zeiss Meditec in Dublin, California via the enclosed self-addressed, pre-paid Federal Express envelope.For any questions customers should call (925) 557-4422. |  | 2540 | 2540 units | Worldwide Distribution including USA (nationwide) and the UK | Terminated | February 09, 2011 | 26 | 2540 | Not_Computer
Z-0901-2011 | 57377 | ventilator, high frequency | LSZ |  |  | Anesthesiology |  | One Cat. Life Pulse High Frequency Ventilator, Catalog Number: 203A, Manufactured by: Bunnell Incorporated Salt Lake City, Utah, Rx Only. | Life Pulse High Frequency Ventilator | 2 | January 14, 2011 | 2011 | Bunnell, Inc. | Product software coding error may compromise patient safety. | DESIGN: Software Design | Customers were notified by an "Urgent: Medical Device Recall or Correction" letter starting 10/22/10. The letter identified the product and described the issues. Customers were given the option of either returning their ventilator(s) for rework or to receive installation/re-test instructions with an upgrade kit to rework themselves. The letter states that all Life Pulse ventilators in clinical use must be upgraded and there is no charge for the software upgrade. | Contact the recalling firm for information | 660 | 660 total | Worldwide Distribution -- USA, Australia, Canada, Czech Republic, Korea, Malaysia, Trinidad & Tobago, and Uruguay. | Terminated | April 17, 2012 | 459 | 660 | Software | Physical Safety Hazards | Return/Update | Remove or Replace
Z-0903-2011 | 57021 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | PROCEED Surgical Mesh; Product Code PCDG1; :Lot BMG618; Ethicon, Inc. a Johnson & Johnson Company.Used for the repair of hernias and other fascial deficiencies that require the addition of a reinforcing or bridging material to obtain the desired surgical result. | PROCEED Surgical Mesh | 2 | January 14, 2011 | 2011 | Ethicon, Inc. | There is a potential for delamination in one lot of Proceed surgical mesh. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ethicon sent Urgent Voluntary Product Recall letters via Fed Ex on October 18, 2010. The letters described the issue with the affected product and instructed customers not to use it. They are to immediately remove the recalled product from their inventory and return it according to the instructions provided. A reply card is attached and must be filled out and returned to Stericycle whether or not the customer has product in their possession. Separate correspondence has been sent to end customers stating that if they have purchased the recalled product through a distributor, to return the product directly to Stericycle. Replacement product will be provided at no additional charge for all recalled product. |  | 87 | 87 products | Worldwide Distribution | Open |  |  | 87 | Not_Computer
Z-0908-2011 | 51550 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife System Robotic Radiosurgery System, (Standard Treatment Couch) Synchrony Respiratory Tracking System, a Radiation Therapy Device, Model Numbers: 026429, 028185, Product is manufactured and distributed by Accuray, Inc., Sunnyvale, CAA Radiation therapy device. | CyberKnife System Robotic Radiosurgery System | 2 | January 18, 2011 | 2011 | Accuray Inc | The product has the potential for the extension coupling of the Synchrony extension mounting hardware to become loose over time. The assembly weighs approximately 17 pounds; an unexpected contact with users or patients may result in serious injury if the boom descends unexpectedly. | DESIGN: Device Design | Accuracy Inc., sent an URGENT ADVISORY NOTIFICATION letter, dated March 16, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their system to see if it was affected. Customers were instructed to contact Accuracy Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 or e-mail customersupport@accuracy.com to schedule service. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) and the countries of Turkey, Russia, France, Italy, the Netherlands, the UK, Spain, Greece, Switzerland, Germany, India, Taiwan, China, Hong Kong, Korea, Thailand, Malaysia and Japan | Terminated | January 19, 2011 | 1 |  | Not_Computer
Z-0910-2011 | 53778 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series Clinac, Trilogy and Novalis Tx used for SRS treatments, Model Numbers H14, H18, H27, H29, HCX, manufactured by Varian Medical Systems, Palo Alto, CA. Device is a radiation Therapy accelerator intended to deliver megavoltage x-ray treatments for conventional radiotherapy (three dimensional conformal radiotherapy and intensity modulated radiotherapy) and stereotactic radiosurgery and radiotherapy. Stereotactic treatments are inteded for therapy of lesions such as arteriovenous malformations, primary tumors and metastatses. Stereotactic treatments may be intracranial or extracranial and consist of single-session or fractionated delivery. | CSeries Clinac, Trilogy and Novalis Tx | 2 | January 18, 2011 | 2011 | Varian Medical Systems Oncology Systems | Product may deliver radiation treatment to areas larger than intended,to healthy tissue. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 8, 2009, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Varian recommended to consignees that test runs be performed prior to actual patient treatment to ensure that the entire system is behaving as intended and that appropriate quality assurance procedures should be followed prior tothe delivery of any treatment plan involving radiation therapy. Operators should monitor the patient and the treatment console continuously during treatment. For questions call USA and Canada - 1-888-827-4265, Europe +41 41 749 8844. |  | 5208 | 5208 units | Worldwide Distribution | Terminated | February 24, 2012 | 402 | 5208 | Not_Computer
Z-0912-2011 | 56624 | tourniquet, pneumatic | KCY | Pneumatic tourniquet. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker Instruments, SmartPump Tourniquet System, Dual Channel Tourniquet Pump, REF 5920-011-000, Rx only.Intended to be used to temporarily occlude or decrease blood flow in a patient extremity during a surgical procedure of that extremity. | SmartPump Tourniquet Pump | 2 | January 18, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | This recall is related to the Electi-cord recall. There is a risk when excessive force is applied to the plugs, a fracture can develop on the prongs inside the molded section of the plug. Over time the fracture may lead to arcing inside of the molded sction of the plug and a build up of heat which can cause charring, electric shock or fire. To this date there have been zero complaint for this fa | DESIGN: Component Design/Selection | Stryker Instruments sent an Urgent - Medical Device Recall Notification letter dated April 13, 2010, to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Stryker notified customers again on May 24, 2010, and again on September 14, 2010, via email, fax, or telephone.Customers were instructed to locate the units and visually inspect each affected cord. Remove the affected cord and dispose of it using their healthcare facilities disposal procedure. If the cord did not show signs of damage, the unit could remain in service until a replacement cord was available. Customers were to return the enclosed business reply form to confirm receipt of the notification letter and identify how many affected units were currently in their inventory. Stryker instruments would send out replacement cords.If the customer loaned or sold any of the units listed in the notification, the customer was to forward a copy of the notice to the new users and advise Stryker of the new user's name and their new location. If products were disposed of and were no longer in use, the customer was to advise Stryker of their obsolescence by providing them with their serial numbers. For questions regarding this recall call 800-800-4236, ext 3808 or ext 3584 Monday -Friday 8am-5pm (EST). |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Portugal, Canada, China, Taiwan, France, Germany, Greece, Spain, India, Italy, Korea, Switzerland, Singapore, and England. | Terminated | May 16, 2011 | 118 |  | Not_Computer
Z-0920-2011 | 57535 | gown, surgical | FYA | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Ultra Surgical Gown, Large with Polypropylene Fabric and Towel, Reference number 95111-35, Single use only, Sterile, Made in Honduras. | Kimberly Clark Ultra Surgical Gown | 2 | January 18, 2011 | 2011 | Kimberly-Clark Corporation | Product was not properly sterilized. |  | Kimberly-Clark contacted all ten (10) Distributors with a Distributor Recall Letter (sent FedEx Priority overnight) on 12/23/2010, as well as via telephone. This letter instructs the Distributor to extend the Recall to end-user level by notifying each of their customers of the Recall via the provided Customer Recall Letter. Both the Distributor Letter and Customer Letter include Response Sheets that allows for both Distributor and Customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The Distributor is instructed to immediately cease further shipment of the product and to physically quarantine the product, and fax the completed Response Sheet to our firm'sRecall Coordinator within 5 business days of receipt of the letter. The Customer is instructed to immediately cease further use of the gowns and physically quarantine the product, and fax the completed Customer Product Response Sheet to Kimberly Clark's Recall Coordinator within 5 business days of receipt of the letter. Upon receipt of the Response Sheets, the firm will provide instructions for the return of the impacted gowns for delivery to a quarantine location at our warehouse. |  | 200 | 200 cases (32 units per case) | Nationwide Distribution -- CA, CT, LA, NJ, NM, RI, CO and TX. | Terminated | July 27, 2011 | 190 | 240 | Not_Computer
Z-0911-2011 | 56772 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Client software versions with build numbers from 7.3.10 through 8.6.15 (i.e. excluding 8.6.17); Model Number: H48; Radiation Treatment Planning System | Varian Eclipse | 2 | January 18, 2011 | 2011 | Varian Medical Systems, Palo Alto, CA | Possible incorrect dose calculation: An anomaly in certain Eclipse versions where, under certain specific sequences of user actions, the dose distribution may fail to be re-calculated even though changes have been made which affect dose. | DESIGN: Software Design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected users on August 6, 2010. Letters were also sent to Varian Sales, Marketing and Service organizations to inform them of the recall. | Contact the recalling firm for information | 7019 | 7019 | Worldwide Distribution. | Terminated | July 21, 2011 | 184 | 7019 | Software | Device Operation | Software update | Software Update
Z-0907-2011 | 56773 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR Digital Linear Accelerators configured with the 160 leaf Multi-Leaf Collimator (MLC) and MEVATRON and PRIMUS Digital Linear Accelerators configured with the High Performance Defining Collimator (HPD); Product is manufactured by Kemnath, Germany-Siemens AG OCS and Siemens AG, Erlangen, Germany-Siemens AG OCS and distributed by Siemens Medical Solutions, Oncology Care Systems, Concord, CA;Intended to deliver x-ray radiation for therapeutic treatment of cancer.ARTISTE with 160 MLC,Part No.8139789,ONCOR AG with 160 MLC,Part No.5863472ONCOR Impression Plus with 160 MLC,Part No.5857912,ONCOR Expression with 160 MLC,Part No.7360204,ONCOR Expression with 160 MLC,Part No.7360717,PRIMUS HI with HPD,Part No. 4504200,MEVATRON M2 / PRIMUS Mid-Energy / HPD,Part No.1940035,MEVATRON 6323-2 / HPD,Part No.19244500,MEVATRON K2 / HPD,Part No.1940753,MEVATRON PRIMART / HPD,Part No.5500371,MEVATRON 77 / HPD,Part No.5659503,MEVATRON M 7400 / 7440 (MEX) / HPD,Part No.5672977,MEVATRON M 6700 (MX) / HPD,Part No.5693908,MEVATRON M 6730 / HPD,Part No.5694005,MEVATRON M 7140 / HPD,Part No.5694104,MEVATRON M 7400 / HPD,Part No.5694153,MEVATRON M 7445 / HPD (MEX),Part No.5694203,MEVATRON K 7467 / HPD,Part No.5694252,MEVATRON K 7767 / HPD,Part No.5694302,MEVATRON K 8067 / HPD,Part No.5694401,MEVATRON M 6300 (MX) / HPD,Part No.8317000,MEVATRON M 6730/6740 (MEX) / HPD,Part No.8319758,MEVATRON MD (MD, MDX) / HPD,Part No.8319808,MEVATRON KD / HPD,Part No.8319857,MEVATRON MDX / HPD,Part No.8496200,MEVATRON KD-2 / HPD,Part No.8515520,MEVATRON M3 6300 / HPD,Part No.9401316,MEVATRON M2 6700 / HPD,Part No.9401407,MEVATRON M2 6740 / HPD,Part No.9401506,MEVATRON KDS / HPD,Part No.941522,MEVATRON MD-2 / HPD,Part No.9401654,MEVATRON MDX  2 / HPD,Part No.9401746,MEVATRON KDS-2 / HPD,Part No.9411588,MEVATRON KD-2 / HPD,Part No.9822685,MEVATRON KDS-2 / HPD,Part No.9822693 | ARTISTE, ONCOR, MEVATRON, and PRIMUS Digital Linear Accelerators | 2 | January 18, 2011 | 2011 | Siemens Medical Solutions, Oncology Care Systems, 4040 Nelson Ave., Concord, CA 94510 | The potential safety issue with the 160MLC or the HPD collimator of mechanical noises and/or sticking of the collimator when rotating the collimator. The potential for mistreatment to a patient should the collimator not be aligned correctly if the bearings are damaged. |  | Urgent: Medical Device Correction Safety Advisory Notice letters were sent by certified mail commencing August 3, 2010. The letter described the issue and when it occurs, the measures users could take, and how Siemens would address the issue. It is recommended that customers include a check of the field alignment at 90 and 270 degrees gantry angle in the daily check prior to start of patient treatments. A copy of the Safety Advisory Notice should be included in the Owner's Manual and questions should be directed towards the local Siemens Service Engineer or Application Specialist. |  | 1167 | 1167 | Worldwide Distribution. | Terminated | January 31, 2012 | 378 | 1167 | Not_Computer
Z-0906-2011 | 56871 | carrier, sponge, endoscopic | FGS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | InfraVision IR Illuminator Console, model number 220-180-521, manufactured by Stryker Endoscopy San Jose, for use with the Stryker InfraVision system.Intended use:Medical device for human use, during surgical procedures to transilluminate anatomical structures during surgical procedures with infrared light. It can be used in either open surgical interventions or during laparoscopic procedures. infrared light which is not visible to the human eye is detected by all of Stryker's infrared sensitive cameras. The InfraVision console outputs the infrared light via disposable fibers hooked to the console that are contained in either the E-kit or the U-kit. | InfraVision IR Illuminator Console | 2 | January 18, 2011 | 2011 | Stryker Endoscopy | Product may not adequately visualize IR light emitted from U- and E- kits due to a manufacturing error. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Stryker Endoscopy, sent an "Urgent: Device Correction" letter dated September 28, 2010, t o all consignees/customers via a Fed Ex Mailer that included notification letter, acknowledgment receipts and updated instructions for use. International Stryker sites received product advisory notification. Stryker reps at those sites are responsible for notifying customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to discard any old IFUs from the InfraVision Console and ensure that they only use InfraVision compatible scopes; read the newly updated IFU provided with the letter; and complete and sign the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in or email to infravision@stryker.com. Note: Not returning the "acknowledgement of receipt" card will result in additional notices.Please contact your local Stryker sales representative to obtain information on scopes that are specific to the InfraVision system.If you have any questions, please contact us at 1-800-624-4422 or email at infravision@stryker.com. |  | 861 | 861 units. | Worldwide distribution: USA and countries including: Australia, Africa, Brazil, Canada, China, France, Greece, India, Italy, Korea, Latin America Mexico, New Zealand, Portugal, Romania, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | February 09, 2011 | 22 | 861 | Not_Computer
Z-0919-2011 | 54229 | system, pressure measurement, intermittent | JFC | Intermittent pressure measurement system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Product labeling of the X3 Display reads in part:"XSENSOR Technology Corporation***X3 DISPLAY*** REF: XBOM-002 PID: CTDPDA00 S*** R***60601-1 61010-1***Calgary, AB, CANADA***Made in Canada".Product labeling of the X3 PRO reads in part:"XSENSOR Technology Corporation***X3 PRO***REF:XBOM-004 PID: CMUHUB00 S***R***60601-1 61010-1***Calgary, AB, CANADA***Made in Canada".Pressure mapping and surface assessment for seats, seat backs, and beds. XSENSOR pressure imaging systems, also known as pressure mapping systems, are a valuable tool in helping clinicians prevent and manage pressure ulcers. Clinicians, nurses, and rehabilitation specialists use XSENSOR pressure imaging systems to help prevent pressure ulcers through improved surface selection and patient positioning on wheelchairs and hospital beds. | Pressure Imaging System | 2 | January 18, 2011 | 2011 | Xsensor Technology Corporation | The plastic enclosures for the X3 Display and X3 PRO Pressure Mapping System has a UL94 HB flammability rating while the medical equipment specification UL 60601-1 requires V-2 rated plastic for electronic enclosures. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | On July 6, 2009, Xsensor sent an "URGENT: MEDICAL DEVICE RECALL" notifications to their consignees/customers.The consignees/customers were informed that the incorrect material of the plastic enclosures was used. The incorrect plastic had a UL94 HB flammability rating while the medical equipment specification UL 60601-1 required V-2 rated plastic for electronic enclosures. The firm provided prepaid shipping box to consignees/customers with instructions for returning the components. The components were refurbished and returned to consignees for free of charge.For further information please contact XSENSORs Customer Support at 1-866-927-5222. |  | 33 | 33 units of X3 Display and 2 units of X3 Pro | Devices were distributed to one distributor, medical facilities, and rehab centers through out the US, including following states:AL, AR, AZ, CA, CO, FL, GA, IA, IL, KS, ME, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, and WI. | Terminated | January 31, 2011 | 13 | 33 | Not_Computer
Z-0918-2011 | 57068 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips OmniDiagnost Eleva with Digital Image (DI) Rel.2 system.The Philips OmniDiagnost Eleva is a diagnostic-imaging device intended for radiographic, fluoroscopic, angiographic, and interventional applications. | Philips OmniDiagnost Eleva with Digital Image (DI) Rel.2 system | 2 | January 18, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | Images from two different patients may be combined in one patient folder in the Digital Image (DI) or may be combined as one (printed) image. It is possible that an image of a previous patient registered in the DI will be overwritten with an image of a later patient. | OTHER/UNDETERMINED: Pending | On March 6, 2009, Philips sent an "URGENT-Device Correction Easy Diagnost Eleva, MultiDiagnost Eleva, and OmniDiagnost Eleva with DI Rel. 2 systems notification" to their consignees/customers.The consignees/customers are advised that before diagnosing images, always check the stored (Digital Image, Viewform or PACS) or printed images to ensure that these represent the data and images captured from the examination of the respective patient. Do not use images if they do not represent the patient that was examined. Consignees will be contacted by Philips to schedule a software upgrade of their units.The consignees can contact Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FCO 70800070" with any questions about this recall. |  | 13 | 13 units | Nationwide distribution: USA including states of: AZ, CA, CO, CT, DE, FL, GA, HI, ID, LA, MA, ME, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, and WY. | Terminated | June 16, 2011 | 149 | 95 | Software
Z-0923-2011 | 50074 | alpha-1 microglobulin, antigen, antiserum, control | MGA | Alpha-globulin immunological test system. | Immunology | Immunology | 510(K) Exempt | K-ASSAY Alpha-1 Microglobulin Reagent Kit. Kit contains: 1 vial Reagent 1, Buffer 50 mL, and 1 vial reagent 2, Antibody Reagent 50 mL. The product is packaged in a kit box along with with the package insert. The package insert is labeled in part: "K-ASSAY Alpha-1 Microglobulin...Cat. No. KAI-056...For the Quantitative Determination of human Alpha-1 Microglobulin in urine, serum, and plasma...KAMIYA BIOMEDICAL COMPANY...12779 Gateway Drive...Seattle, WA 98168 USA...". | KASSAY Alpha1 Microglobulin Reagent | 3 | January 19, 2011 | 2011 | Kamiya Biomedical Company, LLC | Standardization of the reagent performance data did not match the standardization used to assign the calibrator values. | TRAINING: Employee Error | Kamiya Biomedical Company, LLC sent an urgent medical Device Recall letter dated September 17, 2008, and the corrected package insert via e-mail. . The letter identified the product, the problem, and the action the customer should take. Customers were instructed to replace their previous package inserts with the corrected version and asked to sign and FAX the letter back once that had been completed. As one customer is also a distributor, the recalling firm told the this distributor's representative to advise their single customer of this correction and replace the package insert, and return the signed letter back to the firm. On September 22, 2008, both customers were contacted by phone and reminded to fax the signed letters once completed.For questions regarding this recall call 1-800-222-0342. | Contact the recalling firm for information | 5 | 5 kits | Kits were distributed to one laboratory in MN and to one distributor in PA who further distributed product to one customer. | Terminated | January 20, 2011 | 1 | 5 | Not_Computer
Z-0922-2011 | 56356 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Stryker brand Cart with Isolation Transformer, Model Numbers: 240-099-001, 240-099-001K, 240-099-011, 240-099-011K, 240-099-020, 240-099-020K, 240-099-050, 240-099-072.240-099-001,240-099-001K,240-099-011,240-099-011K,240-099-020,240-099-020K,240-099-050,240-099-072,Product is manufactured and distributed by Stryker EndoscopySan Jose, CAThe product is indicated for use in the carts include an isolation transfer and a power cord for connecting the transformer to an electrical outlet as means of providing power to the cart for the equipment it houses. | Electrical Power Cord | 2 | January 19, 2011 | 2011 | Stryker Endoscopy | The electrical cord power plug design may allow the prongs to develop a fracture inside or outside the plug, if the plug is misused. | DESIGN: Device Design | Stryker Endoscopy sent an URGENT: Device Recall letter dated April 16, 2010, to Stryker Sales reps and all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Locate the affected units.Visually inspect each cord for damage.If the cord has bent or cracked prongs the unit should be taken out of service immediately. The customer was instructed to call a Regulatory Affairs Recall Representative at (408) 754-2738 for further instruction. If the cord did not show signs of damage, it could remain in service.An acknowledgement form was included to document confirmation of receipt of the letter. Upon completion of the power cord inspection, customers were to sign the acknowledgement form indicating that all of their units were inspected. Customers were asked to mail the form to Stryker Endoscopy or fax to (408) 754-8378 or scan the card and email to PowerCord@stryker.com.For questions regarding this recall call the Stryker RA/QA Department at (408) 754-2738 or (800) 435-0220, Reference RA 2010-034. |  | 7371 | 7371 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MA, ME, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, MY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Canada, Chile, China, India, Korea, Mexico, Singapore, Taiwan and Latin America. | Terminated | January 19, 2011 | 0 | 7371 | Battery
Z-0925-2011 | 57573 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DPOC-15, Cholangioscopy Access Balloon, G54951, 15 mm, Rx only, Cook Medical, Winston-Salem, NC 27105This device is intended to guide both optical and accessory devices for diagnostic and therapeutic applications during endoscopic procedures in the pancreaticobiliary system including the hepatic ducts. Also, used for endoscopic removal of biliary stones. | Cholangioscopy Access Balloon | 2 | January 19, 2011 | 2011 | Cook Endoscopy | Product defect; There is a risk of gastrointestinal perforation due to balloon deflation difficulty or air embolism resulting in poor patient outcomes including death. | OTHER/UNDETERMINED: Pending | The firm, Cook Medical, sent an "URGENT: PRODUCT RECALL" letter dated January 7, 2011, to all customers. The letter described the product, problem and the actions to be taken by the customers. The customers were instructed to quarantine any affected product and return to Cook Medical. If they are distributors, they were instructed to contact their customers, direct them to quarantine any affected product and coordinate the return of the affected product to Cook Medical. The customers were also instructed to complete and return the bottom of the letter and send to Cook Medical as well as the PRODUCT RECALL EFFECTIVENESS SURVEY via fax at 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. Note: All non-responding customers will receive a total of three (3) recall notifications ( two (2) follow up requests for a response in addition to the initial recall notification). Should you have any questions, please contact the Customer Relations Department at Cook Medical for more information. Please use CustomerRelationsNA@cookmedical.com or call toll free at 1-800-457-4500, press 4 then enter extension 2146. |  | 103 | 103 units (total for both units under recall) | Worldwide distribution: USA including states of: AL, CA, FL, IL, IN, LA, MA, NC, NY, OH, PA, SC, TN, TX, WA, and WI; and countries of: Belgium, Canada, Germany, Hong Kong, and the Netherlands.. | Terminated | August 04, 2011 | 197 | 104 | Not_Computer
Z-0928-2011 | 56931 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Fixed Femoral A/P Sizing Guide; non-sterile instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; A subsidiary of Stryker CorporationThe A/P sizer is designed to set the desired external rotation and to provide adjustment of the anterior/posterior position when needed. | Scorpio Fixed Femoral A/P Sizing Guide | 2 | January 19, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that there is a potential for the shoulder screw on the Scorpio Sizing Guide (the "Sizing Guide") to disassociate from the main body of the instrument. | DESIGN: Component Design/Selection | Stryker Orthopaedics sent an 'Urgent Update Product Recall' letter dated December 30, 2010, along with Product Acknowledgement forms to all customers (US and international). The letter described the product, problem and action to be taken by the customers. The customers were asked to complete and return the PRODUCT RECALL ACKNOWLEDGMENT FORM within five (5) days via fax at 201-831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team, at (201) 831-5028. |  | 6 | six units | Worldwide distribution: USA and countries including: Australia, Brazil, China, Columbia, France, Greece, India, Italy, Korea, Poland, Singapore, South Africa, and United Kingdom. | Terminated | July 31, 2013 | 924 | 243 | Not_Computer
Z-0924-2011 | 56859 | table, surgical with orthopedic accessories, ac-powered | JEA | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Mizuho ProFx 6850 Operating Room Table, with orthopedic accessories, AC Powered, manufactured by Mizuho OSI, Union City, CA. | ProFx 6850 | 2 | January 20, 2011 | 2011 | Mizuho Orthopedic Systems Inc | Device component spar gear box locks up and may result in failure of push pull tube, which may result in release of breaking and impact. |  | All customers will be contacted to arrange for removal and shipping arrangements. Verification will be done by the Repair and Verification instructions for the leg spars, via registered letters.*******UPDATED INFO---The firm sent another notification letter---dated 12/22/10---which states in part: "****The upgrade involves the correction of your 6850 PROfxÂŽ spars (6850-2100) at your facility by a Qualified Technician. The spars shall be upgraded to remove the 6850-2123, Drive Bracket, Spar Push Rod and replace it with the 6850-2170, Drive Bracket, Spar Push Rod. Upon completion of the upgrade, the 6850-2123, Drive Bracket, Spar Push rod shall be returned to Mizuho OSI for processing. *******. | Contact the recalling firm for information | 17 | 17 units | Worldwide Distribution -- USA & Germany. | Terminated | May 17, 2011 | 117 | 17 | Not_Computer
Z-0929-2011 | 56403 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CG8+ Cartridge;pH, pCO2, PO2, NA, K, iCa, Glu, Hct; IVDAbbott Point of CareThe test for PCO2, as part of the i-STAT System, is intended for use in the vitro quantification of carbon dioxide partial pressure in arterial, venous or capillary whole blood. PCO2 along with pH is used to assess acid-base balance. | iSTAT CG8 Cartridge | 2 | January 20, 2011 | 2011 | Abbott Point of Care Inc. | Two cartridge lots were identified that may generate falsely elevated PCO2 results. | OTHER/UNDETERMINED: Pending | Abbott Point of Care Inc. sent an Urgent Recall Notice letter dated July 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the cartridges from the identified lots and follow the instructions on the enclosed reply card to receive replacement cartridges. If cartridges were distributed from the lot numbers to another institution, a copy of the recall letter should be provided to them.For any questions regarding this recall contact Abbott-Point of Care Technical Support at 800-366-8020, Option 1. | Contact the recalling firm for information | 55350 | 55350 total cartridges | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of Oman and Turkey. | Terminated | December 05, 2011 | 319 | 55350 | Not_Computer
Z-0936-2011 | 57572 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex(R) PSI SetIntended for use in the hospital catheterization laboratory for the percutaneous introduction of various devices into veins and/or arteries in a variety of diagnostic and therapeutic procedures | Arrow Super ArrowFlex(R) Percutaneous Sheath Introducer Set | 2 | January 21, 2011 | 2011 | Arrow International Inc | Packaging may have been compromised, and sterility of product cannot be guaranteed. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | The firm, Teleflex Medical thru its subsidiary Arrow International, Inc., sent an "Urgent Field Safety Notice" letter dated December 20, 2010, with Field Acknowledgement form to all customers. The letter described the product, problem and action to be taken by the customers. The letter instructed the customers to check their stock for affected product, cease use and distribution of product and quarantine all products from affected lot immediately; contact their local Teleflex Medical Customer Service/Sales Rep (or Distributor) for a return authorization number prior to filling out the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM; and complete and return the form immediately by email. If you do not have the product from the affected lot, indicate it on the form and fax it to your local Teleflex Medical Customer Service/Sales Rep (or Distributor). The Teleflex Medical Customer Service/Sales Rep (or Distributor) will coordinate the product return with the customer.If you have any questions, please contact the Regulatory Affairs Project Manager at 610-378-0131 ext 3318 or email: Jody.Cadd@teleflexmedical.com. |  | 320 | 320 units | International distribution only: countries include: Austria, Belgium, Czech Republic, Germany, Greece, Italy, and Luxembourg. | Open |  |  | 320 | Not_Computer
Z-0931-2011 | 57515 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161 | Colleague CX Volumetric Infusion Pump | 2 | January 21, 2011 | 2011 | Baxter Healthcare Corp. | While the infusion pump was being serviced at the Baxter Service Center, it was inadvertently routed to the packaging area before the final inspection. | TRAINING: Employee Error | Baxter telephoned the one customer on December 1, 2010, informing them that that an outgoing Quality Control Evaluation was not performed on their Colleague pump serial number 15030799CC during the last service. Since this was not conducted on this pump, the device may not function as designed and could lead to an interruption of therapy. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.For questions regarding this recall call (800) 437-5176. | Contact the recalling firm for information | 1 | 1 pump | Texas | Terminated | May 13, 2011 | 112 | 1 | Not_Computer
Z-0933-2011 | 57334 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter FloGard 6201 Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product code 2M8063For delivery of intravenous solutions, drug solutions, enteral feedings and blood. | Baxter FloGard 6201 Volumetric Infusion Pump | 2 | January 21, 2011 | 2011 | Baxter Healthcare Corp. | The FloGard Infusion Pump was released from service without having the Current Leakage test performed. This could lead to the patient/user to receive an electrical shock. | PRODUCTION CONTROLS: Process Control | The firm, Baxter, telephoned the one customer on November 17, 2010, informing them that electrical safety testing was not performed on their FloGard pump serial number 11091330FA during the last service. Since the test was not conducted on the pump, the device may not function as designed and could lead to a hazard of electrical shock. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter advised the customer that a Baxter Field Service Engineer would be sent to their facility to perform electrical safety testing or alternatively, the customer may choose to send the pump to Baxter for testing.If you have any questions regarding this information, please contact Vice President, Quality, GIS at 847-270-5507. |  | 1 | 1 pump | Nationwide distribution: WA only | Terminated | March 16, 2011 | 54 | 1 | Hardware | Physical Safety Hazards | Replace/Return | Remove or Replace
Z-0932-2011 | 57527 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC);All versions up to and including 8.8.17Varian Medial Systems, Palo Alto, CADesigned to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | 4D Integrated Treatment Console (4DITC) | 2 | January 21, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Multiple isocenter plans will load on the 4DITC and are available for imaging and treatment, but the 4DITC software does not recognize that mulitplex isocenter is present. | DESIGN: Software Design | Urgent Medical Device Correction/Field Safety notification letters were distributed to all affected users, and to Varian Sales, Marketing and Service Organizations on December 14, 2010 by certified mail. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters described the issue and recommended user actions along with Varian's actions. Users are to use caution when acquiring actual couch positions for any plan. In addition, each plan should be reviewed for quality assurance purposes in accordance to set departmental or organizational policies. The firm updated the procedure and enclosed a copy with the letter. It was also posted to the customer support site. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 4188 | 4188 | Worldwide Distribution. | Terminated | July 22, 2011 | 182 | 4188 | Software | Treatment/Delivery/Therapy | Procedure update | Other
Z-0934-2011 | 57612 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Lubricating Jelly, Sterile, Non-staining, Made in the USA, packaged into the following size containers 5 gram tube, 2 ounce and 4 ounce tubes, 4 ounce bottle, 3 and 5 gram packetsIntended use: A patient lubricant is a device intended for medical purposes that is used to lubricate a body orifice to facilitate entry of a diagnostic or therapeutic device. Brand names: 1. Allegiance, Distributed by: Cardinal Health, McGaw Park, IL 60085-6787, USA, Cat. LJT2. Net Wt. 2 oz., Cat. LJF3, Net Wt. 3g, Cat. LJT4, Net Wt. 4 oz., Cat. LJT5, Net Wt. 5 g; 2. Select medical products, manufactured for PSS World Medical, Inc., 4345 Southpoint Blvd., Jacksonville, FL 32216, Reorder #137, 3 g/packet, Reorder #136, 4 oz.; 3. Novaplus, Manufactured for Novation, Inc., Irving, TX 75039, Manufactured by: Triad Group, Inc., 700 West North Shore Drive, Hartland, WI 53029, Cat. No. V10-8344, Net Wt. 3 grams, Cat. No. V10-8919, Net Wt. 4 oz., Cat. No. V10-8917, Net Wt. 2 oz.; 4. Triad Group, Inc., 700 West North Shore Drive, Hartland, WI 53029, Cat. No. 10-8917, Net Wt. 2 oz, Cat. No. 10-8946, Net Wt. 5 g, Triad Plus Cat. No. 11-8344, 3 g., Cat. No. 11-8472, 5 grams, Cat. No. 10-8919, Net Wt. 4 oz, Cat. No. 10-8500, Net Wt. 4 oz; 5. IMCO, Manufactured for: Independent Medical Co-op, Inc., Daytona Beach, FL 32114. Reorder No. 8919-IMC Net Wt. 4 oz.; 6. McKesson Medi-Pak Performance, Marketed by McKesson Corporation, McKesson Surgical, Richmond, VA 23228. Reorder No. 66-8919, Net Wt. 4 oz.; 7. Henry Schein, Distributed by: Henry Schein Inc., Melville, NY 11747. 104-9637, Net Wt. 4 fl. oz. | Sterile Lubricating Jelly | 2 | January 21, 2011 | 2011 | H & P Industries, Inc. dba Triad Group | Triad Group is recalling all lots of Sterile Lubricating Jelly manufactured by Triad Group. This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products. | PRODUCTION CONTROLS: Process Control | The firm, Triad Group, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 22, 2010, to all customers via US Postal Service certified. The letter described the issue, identified affected product as that which contained a lot number beginning with the digits 7, 8, 9, or 0, and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; notify their customers at once if the recalled product was further distributed, and complete and return the enclosed Recall Acknowledgement form via fax at 262-538-2947 or mail along with the recalled product to: Recall Coordinator, Triad Group, 700 West North Shore Drive, Hartland, WI 53029.If you have any questions, please call Triad Group Customer Service Monday thru Friday, between the hours of 8:30 a.m. and 4:00 p.m. Central Time at 262-538-2900 ext 2761. | Contact the recalling firm for information | 453787 | 453,787 cases | Worldwide distribution: USA including states of: CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MI, MO, MS, MY, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI and Puerto Rico; and countries including: Australia, Belgium, Canada, China, Lebanon, New Zealand, Philippines, Northern Marianna Islands, Saudi Arabia, Singapore, Taiwan, and United Arab Emirates. | Terminated | June 11, 2012 | 507 | 453787 | Not_Computer
Z-0935-2011 | 57516 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX and CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8161 and 2M9161.Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. | Colleague CX and CXE Infusion Pumps | 2 | January 21, 2011 | 2011 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | TRAINING: Employee Error | Baxter telephoned the two customers on December 1, 2010, informing them that that a post repair test was not performed correctly on their Colleague pumps (serial numbers 12070599CC and 17086323CP) during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.*** Baxter expanded the recall to include twelve additional pumps on June 22, 2011 by telephone calls to the affected accounts. The recall strategy remains the same as the December 1, 2010 instructions. ****. |  | 14 | 14 pumps | Nationwide Distribution -- Arizona, Indiana, *** California, Iowa, New Jersey, Ohio, Virginia and West Virginia***. | Terminated | January 20, 2012 | 364 | 14 | Not_Computer
Z-0938-2011 | 50791 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Crosser 14S Catheter, Catalog/REF CRU14S, Lot number 5053 manufactured by Flowcardia Inc, Sunnyvale, CAIntended use: Catheter for use in surgery. | Crosser 14S Catheter | 2 | January 24, 2011 | 2011 | Flowcardia Inc. | Mislabeled expiration date: Some units were mislabeled with a three year shelf life, but the product is validated only for a two year shelf life. | TRAINING: Employee Error | The firm, FlowCardia, Inc., sent an "URGENT! Medical Device Recall" letter dated January 9, 2009 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to return the product to FlowCardia, at the firms expense, for replacement and Confirm that this product has been used by completing and returning the attached fax document via fax to 408-616-0053. The customers were ask to please make an attempt to contact the customer service representatives, Operations specialist at 888-456-7286 or Vice President, Clinical, Quality & Regulatory Affairs at 408-617-0352 x302, within the next 30 days. Note: FlowCardia Sales, Clinical and Customer support will send notification via certified mail or traceable courier and track response.If you have any questions or need any additional information email (dmichaels@flowcardia.com) or telephone at 408-617-0352 x302. Alternatively you may contact customer service at 888-456-7286. |  | 35 | 35 units | Nationwide distribution: CA, NJ, LA, TX, GA, MI and CO. | Terminated | January 25, 2011 | 1 | 35 | Not_Computer
Z-0937-2011 | 56804 | pacemaker, cardiac, external transcutaneous (non-invasive) | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn PIC 50 Portable Intensive Care System. Model 971081 series.Designed for use in and out of the hospital. | Welch Allyn PIC 50 Portable Intensive Care System | 2 | January 24, 2011 | 2011 | Welch Allyn Protocol, Inc | Potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The loss of the ability to view the patient's Electrocardiogram (ECG) on the screen could cause the user to delay or fail to administer defibrillating shocks needed to resuscitate a patient. | OTHER/UNDETERMINED: Pending | On 9/17/10, Welch Allyn began sending out the Important Product Alert letter to all their direct consignees. On 9/24/2010, the firm sent out the letters to their distributors.The Product Alert letter informs the consignees of a low potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail.The consignees are advised to follow the instruction in the PIC 50 user's manual requiring routine shift equipment checks daily or at each shift change to minimize the likelihood of the malfunction from occurring. The checks ensure that the PIC 50 and its component (including the screen display) are fully operation and undamaged.The letter further states "our analysis indicates that routinely doing such checks will likely lead to discovery of this malfunction as it occurs, and thus timely maintenance and repair".The distributors are advised to communicate this Product Alert to their customers if the devices were further distributed. For customers in US and Canada, distributors are advised to provide the electronic Excel file with customer's name, address, serial number and part number to WA2221@Stericycle.com. For customers in other countries, they are advised to provide information to PIC50Product Alert@welchallyn.comConsignees can call Welch Allyn at 800-289-2500 for any questions during normal business hours. |  | 3606 | 2002 units to US consignees and 1604 units to foreign consignees | Worldwide Distribution -- US, including states: AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and countries ofAustralia, Austria, Belgium, Canada, China, Columbia, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Jordan, Kuwait, Kyrgyzstan, Latvia, Lebanon, Liechtenstein, Lithuania, Malaysia, Mexico, Oman, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and Yemen. | Terminated | July 25, 2011 | 182 | 3606 | Other | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0947-2011 | 57470 | lavage, jet | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | The Stryker InterPulse system is a single patient use, sterile, disposable, pulsed lavage device. The system consists of an irrigation handpiece, irrigation tubing, irrigation tips, suction tips, suction tubing, and splash shields. The components of the InterPulse Irrigation System are used for wound debridement, soft tissue debridement and cleansing of medical, clinical and surgical sites. This includes cleansing of bone in surgical procedures, hydrodebridement of chronic wounds, debridement of loose skin from burns and cleansing of trauma wounds. | InterPulse Handpiece Set With Suction Tubing | 2 | January 25, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | There is potential for a breach in sterility in the blister packaging. This can lead to the use of non-sterile product which, if contaminated, could lead to a patient infection. | PRODUCTION CONTROLS: Packaging Process Control | Stryker sent a Medical Device Recall Notification letter dated October 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and dispose of all affected product and to complete the business reply form and return to Stryker to receive replacement interpulse products at no charge.For questions regarding this recall call (269) 323-7700. | Contact your health professional | 490 | 490 | Worldwide Distribution - USA (including all 50 states) and the countries of Brazil, Canada, France, Germany, Japan, Netherlands, South Africa, South Korea, Spain, Sweden, Switzerland and the UK. | Terminated | June 06, 2011 | 132 | 3187 | Not_Computer
Z-0955-2011 | 57380 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Newdeal (Newdeal Compression Plate) Model #UNI-CP 339900 Sets, Lot Numbers: E71H, E7V1, E86R, EELE, EGTCThe Newdeal Compression Plates are intended for fixation of bone fractures or for bone reconstruction. | UniCP Compression Plate System | 2 | January 25, 2011 | 2011 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has decided to initiate a voluntary recall of the Uni-Comp Kits due to inaccurate information relative to the recommended sterilization parameters described in the products Instructions for Use (IFU). The sterilization parameters noted in the Uni-Comp IFU were incorrect. It is noted that the sterilization recommendations in the UNI-Comp IFU in which the Exposure | OTHER/UNDETERMINED: Pending | The firm, Integra, contacted the Sales Representatives on November 8, 2010 via telephone conference calls. Integra also sent a "URGENT: PRODUCT RECALL NOTIFICATION" letter dated November 9, 2010 all customers via FedEx overnight delivery. The letter described the product, problem and the action to be taken by the customers and the firm. The customers were instructed to immediately examine their inventory, QUARANTINE THE PRODUCT, and complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via mail: Integra LifeSciences Corporation, Attn: Sandy Yates, 311 Enterprise Drive, Plainsboro, NJ 08536, fax to (609)-275-5363 or Scan and email to sandra.yates@integralife.com. The firm will have a Integra LifeSciences Corporation Sales Representative contact the customer to assist them in retrieving and returning their Uni-CP Compression Plate System to Integra. Integra will also ship replacement Uni-CP Compression Plate (Uni-CPII and Large Uni-Clip) Systems back to the customers as soon as possible at no charge.Should you have any additional questions, contact the Director of Marketing, Hindfoot-Extremity Reconstruction or your regional managers at 513-533-7929, 609-275-0500 or toll free at 800-654-2873. | Contact the recalling firm for information | 67 | Sixty seven (67) Kits | Nationwide distribution: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA and WI. | Terminated | November 18, 2011 | 297 | 67 | Not_Computer
Z-0954-2011 | 57498 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Covidien, LigaSure, Blunt Tip Laparoscopic Sealer/Divider, 5 mm-37 cm, REF LF1537, Sterile EO, Covidien, Valleylab, a division of Tyco Healthcare Group LP, 5920 Longbow Drive, Boulder, CO 80301Intended use: LigaSure 5 mm devices are electrosurgical instruments designed to permanently seal and divide vessels, tissue bundles, and lymphatics. | LigaSure Blunt Tip Laparoscopic Sealer/Divider | 2 | January 25, 2011 | 2011 | Covidien LP | During surgical procedures, the device may energize when clamping without pushing the activation button, which could result in unintended energy delivery to tissue. |  | The firm, Covidien, sent a "URGENT: PRODUCT RECALL" letter dated December 6, 2010, to the customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to remove product from inventory, immediately quarantine the product and return any remaining stock of the lots listed to: Covidien Energy-based devices, Attention: R82, Building 4/Receiving, 5920 Longbow Drive, Boulder, CO 80301. The customers were also instructed to complete and return the Customer Response Form immediately via mail (in the envelope provided) or fax to (303) 516-8307.To return any identified product for credit or replacement, contact Customer Service at 1-800-722-8772 Option 1.Should you have any medical questions concerning this recall, please contact your local sales representative or the Covidien Clinical information Hotline at (800) 255-8522 (option 1). | Contact the recalling firm for information | 25009 | 25,009 units | Worldwide distribution: USA and countries including: Australia, Canada, France, Germany, Japan, Puerto Rico, Singapore, and Uruguay. | Terminated | March 27, 2012 | 427 | 25009 | Not_Computer
Z-0939-2011 | 57630 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Linemaster Wireless Footswitch, VASER System, REF Vaser 2.0, Sound Surgical Technologies LLC, 357 South McCaslin Blvd., Suite 100, Louisville, CO 80027 | Linemaster Wireless Footswitch | 2 | January 25, 2011 | 2011 | Sound Surgical Technologies, LLC | Surgical equipment, activated by a wireless foot switch, may remain powered on when switch is no longer depressed. | PRODUCTION CONTROLS: Process Control | Customers were notified by e-mail with an Advisory Notice on 12/21/2010. The Notice identified the affected product and described the issue. The Notice also explained what the potential hazard involved along with a mechanism to eliminate the potential hazard. Sound Surgical would arrange a service call to implement the correction at the firm's expense. If customers experience this potential failure, they are to contact Sound Surgical Technologies at 1-888-677-8765 or 1-888-471-4777. |  | 27 | 27 units | Nationwide Distribution -- AL, AR, AZ, CA, CO, DC, FL, GA, IA, IL, IN, MA, MI, MO, NJ, NY, PA, and TX . | Terminated | February 08, 2011 | 14 | 27 | I/O
Z-0966-2011 | 57708 | catheter, urological | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological, Inc. Koyle Diaper Stent 6.0 Fr, Sterile, GPN G17667, REF KDSCF-060050. | Koyle Diaper Stent. | 2 | January 25, 2011 | 2011 | Vance Products Inc. dba Cook Urological Inc. & Cook OB/GYN | The firm received a complaint that 15 units of the Koyle Diaper Stent were received in packages that had not been sealed. Each stent should be sealed in a poly-Tyvek pouch. | PRODUCTION CONTROLS: Packaging Process Control | The firm sent an URGENT RECALL NOTICE to one consignee on 12/13/2010 (the other consignee returned all their product as a complaint). The letter informed the consignee of the recall and asked the customer to return the recalled product. | Contact the recalling firm for information | 25 | 25 stents | CA, CT | Terminated | January 31, 2011 | 6 | 25 | Not_Computer
Z-0967-2011 | 57188 | [6 empty] | StaXxÂŽ XD Expandable Cartridge Convex Expandable Implant 25mm (L) x 9mm (W) x 7mm (H)Catalog number: 11-1525The StaXxÂŽ XD System is a vertebral body replacement device intended for use in the thoracolumbar spine (TI-L5) to replace and restore height of a collapsed, damaged, or unstable vertebral body or portion thereof, due to tumor or trauma (i.e., fracture). The system is to be placed bilaterally and used with autograft or allograft and supplemental spinal fixation. The supplemental fixation system that is intended to be used with the StaXxÂŽ XD System is the CapSureÂŽ PS Spine System. | Spine Wave | 2 | January 26, 2011 | 2011 | Spine Wave Inc | Package Insert Incorrect | PRODUCTION CONTROLS: Packaging Process Control | Spine Wave notified each affected customer on October 22,2010, via telephone of the product, the issue and action they should take. Customers were asked to segregate the product from other products and immediately return to Spine Wave.For questions regarding this recall call 203-944-9494. |  | 12 | 12 units | Nationwide Distribution including Delaware, Massachusetts, Maryland, and New York | Terminated | March 15, 2011 | 48 | 12 | Not_Computer
Z-0959-2011 | 57609 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Mast Roller Pump System 85, Part No.: 28-95-80, Sorin Group Deutschland Gmbh, Lindberghstrasse 25, Munchen, Germany, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004.Used to perform, control and monitor extracorporeal blood circulation during an operation for up to six hours. | S5 Mast Roller Pump System 85 | 2 | January 26, 2011 | 2011 | Sorin Group USA, Inc. | Touch screen may become unresponsive, inhibiting user input. |  | All affected US customers and distributors were notified by certified mail on Nov 12, 2010 via a Field Correction letter. They were told that they would be contacted by Sorin to arrange an appointment to inspect and replace affected touch screens on site. They were told that they could safely continue using their S5 systems in accordance with the Operator's Manual and previously supplied instructions until the inspection and necessary replacement had been completed. A customer response form was also included for completion and return to the firm. Customers can contact the firm if they have any questions. | Contact the recalling firm for information | 11 | 11 units | Worldwide Distribution. | Terminated | July 13, 2012 | 534 | 6671 | I/O
Z-0968-2011 | 57231 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Sternal Saw II Sterile Blades for the Sarns Sternal Saw System, Catalog number 5589, Terumo Cardiovascular System Corp, Ann Arbor, MI.The sternal saw blades are indicated for use in medial sternotomoties. The saw blades are used wit the Sarns Sternal Saw System. | Sternal Saw II Sterile Blades | 2 | January 26, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm discovered that some outer box labels contain incorrect information regarding the method of sterilization. The label incorrectly indicates the product was ethylene oxide sterlized but the product actually was sterilzed with gamma radiation. | DESIGN: Labeling Design | Terumo Cardiovascular Systems Corporation firm sent out an 'URGENT MEDICAL DEVICE CORRECTION: SAFETY ADVISORY" letter dated October 20, 2010, to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Safety Advisory, assure that all users were aware of the notice, and confirm receipt of the Safety Advisory by faxing the attached Customer Response Form to the fax number on the form. For questions regarding this recall call 1-800-521-2818. |  | 4265 | 42,650 units, 4265 boxes of 10 each | US, Indonesia, UAE, Hong Kong, Japan, Singapore, Malaysia, Taiwan, Belgium, Uraguay, Venezuela, Columbia, Mexico, Guam, Vietnam and Thailand. | Terminated | December 14, 2011 | 322 | 4265 | Not_Computer
Z-0973-2011 | 53198 | cardiac ablation percutaneous catheter | LPB |  |  | Cardiovascular |  | RF Ablation System Foot Switch, used with the EPT-1000 RF Generator, Foot Switch Model Number UPN M0048400, and in kits with UPNs M0044810 and M0044811.Intended use: The EPT-1000xp Cardiac Ablation Controllers and Foot Switch are indicated for use in cardiac ablation procedures with standard and high power BSC cardiac ablation Catheters. | RF Ablation System Foot Switch, used with the EPT1000 RF Generator | 2 | January 27, 2011 | 2011 | Boston Scientific Corp | Product malfunction causes inadvertent delivery of energy. | DESIGN: Component Design/Selection | The firm, Boston Scientific, sent an "URGENT: Medical Device Recall" letter dated January 14, 2009, with attached recall instructions to all customers. The letter described the product, problem, and the action to be taken by the customers. The customers were instructed to complete and return the Reply Verification Tracking Form and Obtain an RGA via fax to 408-957-6204; and return the recalled product using BSC Federal Express Number 1-809-4713-2 via second day delivery to Boston Scientific Corporation, Distribution Center, 500 Commander Shea Blvd., Quincy, MA 02171. Boston Scientific Corporation sales representatives will contact each customer.If you have any questions, please contact 408-935-4952 or email: wareb@bsci.com. |  | 1044 | 1044 units | Worldwide distribution | Terminated | January 27, 2011 | 0 | 1044 | Not_Computer
Z-0974-2011 | 54350 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LT -D Size 3, 10-0002, Oropharyngeal Airway. King Systems Corporation 15011 Herriman Boulevard, Noblesville, IN 46060.Airway management by providing a patent airway to allow patient ventilation. | KING LT D Size 3, 100002 Oropharyngeal Airway | 2 | January 27, 2011 | 2011 | King Systems Corp. | King Systems was issued a warning letter from the Center for Devices and Radiological Health, Office of Compliance on October 28, 2009. The warning letter asserts that Kings Systems KLT(S)-D airway exceeds cleared indications in that it is promoted for "airway management in patients over 4 feet in height, for controlled or spontaneous ventilation, difficult and emergent airway cases, and ambulat | MISBRANDING: Labeling False and Misleading | On October 29, 2009, the firm, Kings System, sent a "Urgent - Product Correction" letter to the customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the promotional material from the King System website which includes literature, and Instructions for Use. The firm instructed all Sales force and distributors to cease dissemination of promotional materials, literature and communications.Note: The firm revised indications, instructions for use, product labels and labeling so it is aligned with the 510Ks. The customers were also instructed to replace current instructions for use with revised instructions for use and destroy the current IFU; return Kings LT(S)-D oropharyngeal airway product inventory to King Systems; and complete and return the attached Instructions for Use Destruction Signature Form via fax to King Systems at 317-776-5197.For further information, please call 317-776-6823 ext 257 or email: LTSD@kingsystem.com. |  | 25000 | 25,000 all products | Worldwide distribution: USA and countries including: Canada, Costa Rica, and Panama. | Terminated | August 16, 2012 | 567 | 25000 | Not_Computer
Z-0987-2011 | 57501 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Physiological Monitoring, Patient MonitorM31250 Information Center Local DatabaseReleases A, B, E.01, G, H, J, K, L, & M of Philips IntelliVue Information Center Software used in conjunction with PAS-210 Kit Remote Active Speakers.The intended use of the Information Center Software is to display physiologic waves, parameters and trends, formal data for strip recordings and printed reports and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. The Information Center Software provides for the retrospective review of alarms, physiologic waves and parameters from its database. An additional intended use is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Physiological Monitoring, Patient Monitor | 2 | January 27, 2011 | 2011 | Philips Healthcare Inc. | Some PAS-201 active speaker assemblies in use with the Philips IntelliVue Information Center (PIIC) have experienced intermittent audio or loss of audio. Speaker failures may delay recognition of an alarm condition, if operators do not recognize that there is no alarm sound from the monitor and the monitor is not connected to another audible alarm annunciation system. | DESIGN: Device Design | Philips notified all affected customers with an Urgent Medical Device Correction notification letter dated December 17, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to observe their system for any of the described symptoms. If they experience intermittent audio or there is no sound from their PAS-210 speaker while it is not connected to another audible alarm annunciation system, customers were instructed to remove the speaker from use and contact their local Philips service representative. Customers were advised not to rely exclusively on the audible alarm system for patient monitoring, the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.ACTIONS PLANNED BY PHILIPSPhilips will replace affected speaker assemblies free of charge.A Philips Healthcare representative would contact the customer when the replacement was available. Field Service Engineers will remove and replace defective speakers to correct the problem. The defective speakers will be returned to vendor (Beach Wire) to be discarded. This correction will be done free of charge for customers. | Contact the recalling firm for information | 1071 | 1071 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, MT, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA WI, WV, WY, and Canada | Open |  |  | 1071 | I/O | Alarm/Message | Replace components | Remove or Replace
Z-0972-2011 | 57051 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Bio Med Devices Disposable Pediatric Breathing Circuit, One SetCatalog Number: 4040LC5 foot /15mm corrugated EVA hoseIntended use: Breathing circuit is a pediatric disposable tubing assembly for use with Lifecare ventilators | Bio Med Devices | 2 | January 27, 2011 | 2011 | Biomed Devices | Breathing circuit may disconnect between the adapter adjacent tubing cuff | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | The firm,Bio Med Devices, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 18, 2010 via fax to the customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to immediately examine their stock for the affected lots, if found, discontinue distributing and using the product and promptly return via UPS #025106 to Bio-Med Devices 61 Soundview Rd, Guilford, CT 06437, Attention: 4040LC Recall Return. The customers were also instructed to immediately contact their accounts, if the product was further distributed , advise them of the recall situation and have their outstanding product returned, and to complete and return the RETURN FAX-IMMEDIATE RESPONSE REQUIRED form via fax to 203-458-0440. Note: Bio-Med informed the customers that they will credit the return product; however, they only plan to replace the product in an emergency situation, because they have very few of the 3/8" x 15mm adapters left and are looking to discontinue the 4040LC due to the discontinuance of this product by the manufacturer.If you have any questions, please contact (203) 458-0202 x 217. | Contact the recalling firm for information | 15 | 15 cases (20/cs) | Nationwide distribution: OH, WI. | Open |  |  | 16 | Hardware |  | Return | Remove or Replace
Z-0970-2011 | 57703 | [6 empty] | RLSense, Restless Legs Sensor, Proto6, Rev-J FirmwareRLSense is a pair of disposable adhesive patches applied to the dorsum of the feet. Each patch contains a microprocessor, memory, movement sensor, and a small coin-type battery. The devices record periodic movements of the feet and legs during the night. | RL Sense | 2 | January 27, 2011 | 2011 | Sleep Modus | Device had no 510(k) clearance prior to patient use. | OTHER/UNDETERMINED: Pending | Sleep Modus, Inc. notified consignees by e-mail on 9/15/2010 and 10/7/2010. The e-mail identified the product, the problem, and the action to be taken. Customers were instructed to let them know if they experienced any of the issues identified in the e-mail. For questions call 214-405-4055. |  | 110 | 110 devices (2 per kit) | Nationwide Distribution including Texas | Terminated | January 31, 2011 | 4 | 110 | Not_Computer
Z-0991-2011 | 57522 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | EndoPlege Sinus Catheter Kit, REF EP, 9 Fr x 18.9", Contains Phthalates, Sterile EO, Sterile R, Edwards Lifesciences LLC, Irvine, CA 92614.The EndoPlege coronary sinus catheter is a three-lumen, 9 Fr catheter with an elastomeric balloon near its tip for occluding the coronary sinus in order to retrograde perfuse the coronary circulation. An 11 Fr catheter introducer sheath, which is compatible with the EndoPlege coronary sinus catheter contamination guard, is packaged with these catheters | EndoPlege | 2 | January 28, 2011 | 2011 | Edwards Lifesciences Llc | Defective product may result in the silicone valve becoming dislodged from the introducer which can result in injury. | DESIGN: Component Design/Selection | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated December 3, 2010 to customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to check their inventory and return the unused devices; complete and return the INTROCSC Introducer Recall Response form to the attention of Customer Service, Edwards LifeSciences, One Edwards Way, Irvine, CA at 949-250-3489. If you have any questions, please contact Edwards Customer Service at 800-424-3278 from 8:00am - 4:30pm Pacific Time. | No consumer action necessary | 3427 | 3,427 kits | Nationwide Distribution | Terminated | October 27, 2011 | 272 | 3517 | Not_Computer
Z-0941-2011 | 57540 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Prelude Short Sheath Introducer, 7F-SMT, 4 cm, REF No: PSS-7F-4-038MT, Sterile EO, Merit Medical, South Jordan, Utah. Intended to provide access and facilitate the percutaneous introduction of various devices into veins and/or arteries while maintaining hemostasis for a variety of diagnostic and therapeutic procedures. Can also provide access to a native or synthetic graft used for hemodialysis. | Catheter introducer | 1 | January 28, 2011 | 2011 | Merit Medical Systems, Inc. | Introducer tips may detach and embolize during use, or the tip material may elicit a thrombogenic response. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Merit Medical Systems, Inc., issued an "Urgent Product Recall Notice" dated December 17, 2010, and contacted all customers (sales reps and authorized distributors) via email, phone and/or fax. The notice described the product, problem and action to be taken by the customers (Sales reps/Distributors). The customers were instructed to contact all their customers to identify, quarantine and return all unused inventory and extend the recall to any of their customers who may have received the affected products. The customers were instructed to complete and return the Manufacturer RECALL NOTIFICATION and Product Retrieval Form via fax to 1-801-208-3378 and send original to: the Manager, Customer Service, Merit Medical Systems, Inc., 1600 West Merit Parkway, South Jordan, UT 84095-2416.Questions should be directed to the Senior Product Manager at 801-208-4344 or the Manager, Customer Service at 801-208-4365. |  | 378 | 378 units | Nationwide Distribution | Terminated | July 11, 2011 | 164 | 378 | Not_Computer
Z-1001-2011 | 56354 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C Series ClinicÂŽs (Includes Trilogy, and Novalis Tx), Model Numbers: H14, H26, H27, H29, Distributed by and/or Manufactured by: Varian Medical Systems Inc., Palo Alto, CA The system consists of two major components, a photon, electron, and diagnostic kV X-ray radiation beam-producing component that is installed in a radiation-shielded vault and a control console area located outside the treatment room. Intended use: The Trilogy Mx" System is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. 7/9/2013 - UPDATE: Determined that the correction only applied to C3 users, not C1 and C2 users. | Varian brand C Series Clinics | 2 | January 31, 2011 | 2011 | Varian Medical Systems Oncology Systems | A circuit design anomaly whereby a faulty balance potentiometer can lead to an undetected electron beam asymmetry. | DESIGN: Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 1, 2009 via Federal Express to consignees/customers. The letter described the product, problem and actions to be taken by the customers and the firm. The customers were instructed to perform a verification of field symmetry on electron beam energies to determine the threshold and report any field symmetry issues to their local Varian service office for attention. The firm, Varian, is developing new functionality to detect the symptoms of this component failure condition under all conditions of electron beam operations and asserting a beam interlock if it is detected. Note: Varian indicated until the solution can be developed, they will conduct an audit of installed Clinacs to identify those the suspect "blue" electron beam balance potentiometers installed, and will execute a program of prophylactic replacement of these potentiometers.If you require further clarification, please contact your local Varian Customer Support District or Regional Manager at phone: USA and Canada-1-888-VARIANS (888-827-4265); Europe- +41 41 749-8844 or email support-emea@varian.com; North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdes@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems - www.myvarian.com; Varian Medical Systems -www.varian.com. | Contact the recalling firm for information | 4411 | 4411 | USA and countries including: Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bolivia, Brazil, Canada, Chile,, China, Columbia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France,French Polynesia, Germany, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Lebanan, Luxembourg, Maca, Macedonia, Malaysia, Martinique, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan,, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea,, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay.. | Terminated | March 14, 2012 | 408 | 4411 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-0990-2011 | 56963 | scalpel, one-piece | GDX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Southmedic/Personna Safety Scalpel with Personna Plus Blades, Reorder No. 73-0615, Size #15, General Surgical Use, Sterile, unless package is damaged or opened., Southmedic Inc., 50 Alliance Blvd., Barrie, ON., Canada L4M 5K3 www.southmedic.com, Made in Canada. | Southmedic/Personna Safety Scalpel | 2 | January 31, 2011 | 2011 | Southmedic, Inc. | Scalpels labeled as sterile have the potential to be unsealed and therefore not sterile. | PRODUCTION CONTROLS: Packaging Process Control | On 9/3/2010, the firm notified consignees about the recall by facsimile letter. The letter instructed consignees to return scalpels for credit or replacement. The firm enlarged the scope of the recall on 11/15/2010. An additional 927 cases were recalled. This includes 11 additional lots. The firm sent out URGENT MEDICAL DEVICE RECALL letters dated 11/15/2010 to its direct accounts. The letter identified that affected product and the reason for the recall. The consignees were instructed to check their inventory, complete and return the attached Voluntary Recall Notification Confirmation Form, return the affected product to the manufacturer. Product return question should be directed towards Marc Pauze at 800-463-7146 ext 299 and regulatory questions should be directed towards Tish Anger at 800-463-7146 ext 342. |  | 1330 | 1330 cases all products | Worldwide Distribution -- USA, including states of TX, MO, NJ, LA, IL, NY, KY, PA, VA, WA, CA, MA, MD, IN, NV, TN, NC, AZ, IA, WI, UT, NE, GA, FL, MN, ME, OH, SD, and KS and countries of Australia and Canada. | Terminated | May 18, 2011 | 107 | 1330 | Not_Computer
Z-1005-2011 | 53780 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 10032746 Acuson Antares PE Product Version 5.0, manufactured by Siemens Medical Solutions, Mountain View, CA.Intended for Abdominal, Intraoperative, Small Parts, Transcranial, OB/Gyn, Cardiac, Transesophageal, Pelvic, Neonatal/Adult Cephalic, Vascular, Intravascular, Musculoskeletal, Superficial Musculoskeletal, Great Vessel, and Peripheral Vascular applications. | Acuson Antares PE | 2 | February 01, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | ECG waveform does not accurately align with 2D, M-Mode or Spectral Images. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm's Customer Safety Advisory notice was mailed to all affected customers. The letter describes when the malfunction occurs and the potential risk. It also recommended that users collect corresponding evidence with spectral waveforms that would confirm or deny findings based on ECG, in order to avoid potential risks for misdiagnosis. The firm planned to exchange the faulty hardware component in the system. Customers would be contacted to schedule for the hardware exchange. Questions regarding this issue should be directed to the local service support person. | Contact the recalling firm for information | 67 | 67 units total | Worldwide Distribution | Terminated | March 29, 2011 | 56 | 67 | Hardware | Output/Calculation | Replace component | Remove or Replace
Z-1008-2011 | 54535 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System, with MultiPlan Treatment Planning System Software version 3.5Medical charged-particle radiation therapy system, intended for treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | CyberKnife Treatment Planning System | 2 | February 01, 2011 | 2011 | Accuray Inc | If electron density values are left empty, the calculation of radiation dose in a patient will be modeled as air-like density material rather than correct density. A plan may be created and saved, thus creating the risk of mistreatment. | TRAINING: Employee Error | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated December 29, 2009, to all customers. The letter described the product, problem, and action to be taken by the customers. The customers were instructed to verify correctness of the CT number to relative electron density and Mass Density Calibration curves. Note: A patch is being developed to address this issue. Accuray Customer Support will contact the customers regarding further resolution.If you have any questions or concerns regarding this issue, please contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non USA) or by email: customersupport@accuray.com. |  | 71 | 71 units are subject to this action | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Italy, France, the Netherlands, the UK, India, Spain, Greece, Saudi Arabia, Switzerland, Germany and China. | Terminated | February 11, 2011 | 10 | 71 | Software | Output/Calculation | Software update | Software Update
Z-1002-2011 | 56357 | bougie, esophageal, and gastrointestinal, gastro-urology | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker Endoscopy, Infravision Esophageal Kit, Model Numbers: 220180548, 220180540, 220180560 | Stryker Infravision | 2 | February 01, 2011 | 2011 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter upon removal of the e-kit from the body. | DESIGN: Device Design | On 5/11/10, Stryker Endoscopy initiated the recall notification letters explaining the reason for the recall and requesting the product be returned. |  | 4125 | 4125 | Nationwide and 3 foreign consignees (Canada, Singapore and Switzerland). | Terminated | February 01, 2011 | 0 | 4125 | Not_Computer
Z-1003-2011 | 56349 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C-Series Clinacs and Acuity, Exact Couch Top, Model Numbers: H69, Product is manufactured and distributed by: Varian Medical Systems Inc., Palo Alto, CA | Varian brand CSeries Clinacs and Acuity, Exact Couch Top | 2 | February 01, 2011 | 2011 | Varian Medical Systems Oncology Systems | The product has been shipped since 2005 with the wrong label regarding max weight to be used on accessory rails. | PRODUCTION CONTROLS: Labeling Mix-Ups | Varian issued via customer technical bulletin on April 19, 2010, to Varian Sales, Marketing and Service organization. Notification. The bulletin identified the product, the problem, and the action to be taken by the customer. Customers were asked to fill out the attached form and send it to their local Varial Medical Systems Field Service Office.For questions regarding this recall call 650-483-3153. |  | 13 | 13 | Worldwide Distribution - Japan only | Terminated | December 23, 2011 | 325 | 13 | Not_Computer
Z-0940-2011 | 56868 | clamp, vascular | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Maquet Heartstring II Proximal Seal System (HS-1045) manufactured by Maquet Cardiovascular, Wayne, NJIntended use: medical device for use by cardiac surgeons during CABG procedures to maintain hemostasis and to facilitate the completion of a proximal anastomosis without application of an aortic clamp. | Marquet Heartstring II Proximal Seal System | 1 | February 01, 2011 | 2011 | Maquet Cardiovascular, LLC | Deployment tube may detach during use due to insufficient adhesive, which may result in product failure to operate. | DESIGN: Process Design | The firm, MAQUET Getinge Group, sent an "URGENT DEVICE REMOVAL IMMEDIATE ACTION REQUIRED" letter dated May 5, 2010, to all customers via Federal Express. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to immediately examine their stocks to determine if they have any of the affected product, if found, discontinue dispensing (distributing) the lot and complete the Field Action Response Form; immediately contact accounts if they acted as a distributor, and acknowledge receipt of this notification whether or not they have any affected products by completing the Field Action Response Form via fax to (U.S.): 1-888-899-2874.If you have any additional questions, please contact your local MAQUET Cardiovascular sales representative or our Customer Service at 1-888-880-2874. | Contact the recalling firm for information | 1630 | 1630 | Worldwide distribution: USA and countries including: Japan, Mexico, Germany, Switzerland, Israel | Terminated | June 21, 2011 | 140 | 1630 | Not_Computer
Z-1004-2011 | 56759 | system, imaging, optical coherence tomography (oct) | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General & Plastic Surgery | 510(k) | Optovue iVue optical coherence tomography system, computer controlled opthalmic imaging system, manufactured by Optovue, Fremont, CA Computer controlled opthalmic imaging system. Device scans the patient's eye using low coherence interferometer to measure the reflectivity of the retinal tissue. The cross sectional retinal tissue structure is composed of a sequence of A-scans. It has a traditional patient and instrument interface. The computer has a graphic user interface for acquiring and analyzing the image. The iVue offers three scans: Retina, Glaucoma and Cornea. For the cornea scan, a lens must be attached to the front of the device for proper scanning. This is the Cornea Anterior Module (CAM). | Optovue iVue | 2 | February 01, 2011 | 2011 | Optovue Inc. | Use of the CAM attachment may result in a positioning anomaly and mistreatment in the eyes of the patient. | DESIGN: Software Design (Manufacturing Process) | Optovue Inc. sent a Field Modification/Correction Notice dated July 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that they could continue to use the iVue until the correction has been made. Customers were instructed to review any existing iVue patient data to ensure that this situation and an improper designation of eye assignment has not occurred. Upon completion of the updates, customers were to sign the attached form indicating they acknowledge that their iVue has been updated and they understand that they must review existing data. For any questions regarding this recall call 1-866-344-8948 or by mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 22 | 22 devices in the US. | Nationwide | Terminated | February 08, 2011 | 7 | 22 | Not_Computer
Z-1009-2011 | 57576 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | AutoSuture SURGIPRO Polyproylene Mesh Clear 3" x 5" SterileREF SPM 35 Surgipro Flat Sheet mesh is a non-absorbable, polypropylene surgical mesh exhibiting high burst strength and tensile strength. Surgipro comes in multiple designs, including multi-filament and open weave. All designs are knitted in such a fashion as to provide bi-directional elasticity, while allowing the mesh to be cut to any shape without unraveling.Mesh is intended to assist in the repair and/or reinforcement of hernia and other fascial defects requiring the additional support of a non-absorbable mesh during and after wound healing. | autosuture | 2 | February 01, 2011 | 2011 | Covidien LP | Sterility compromised due to breach in sterile barrier | PRODUCTION CONTROLS: Packaging Process Control | Covidien notified customers via letter dated November 15, 2010. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to examine inventory and return the affected lot to Field Returns Department, 195 McDermott Road, North Haven, CT 06473. For questions regarding this recall call 203-492-8091. |  | 132 | 132 units | Worldwide Distribution - USA including CA, CT, FL, HI, IL, IN, KS, MA, MD, MI, MN, NC, NJ, NY, OH, PA, SC, TX, and WI and the countries of Canada and Italy | Terminated | January 10, 2013 | 709 | 132 | Not_Computer
Z-1010-2011 | 57595 | powered laser surgical instrument with microbeam\fractional output | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Fraxel re:storeÂŽ Dual Laser System----a laser surgical instrument for use in general and plastic surgery and in dermatology. Model Numbers: MC-SYS-SR1500-D-US; MC-SYS-SR1500-D-US-LOANER; MC-SYS-SR1500-D-UPG-US; MC-SYS-SR 1500-0-1 NTL; MC-SYS-SR1500-D-I-LOANER; MC-SYS-SR1500-D-UPG-INTL; Product is manufactured and distributed by Solta Medical, Inc., Hayward, CA The Fraxel re:storeÂŽ Dual Laser System is indicated for use in dermatological procedures requiring the coagulation of soft tissue. | Fraxel re:store Dual Laser System | 2 | February 01, 2011 | 2011 | Solta Medical, Inc. | The product has the potential to cause inadvertent laser firing which may lead to patient/operator injury if the problem recurs. | DESIGN: Software Design | The firm, Solta Medical, issued three "DEVICE CORRECTION" letters, one dated December 15, 2010, the second one date December 17, 2010 and the third one dated December 20, 2010, to all affected customers. The letters described the product, problem and action to be taken by the customers. The customers were instructed to follow the instructions carefully, sign and date the bottom of the instruction form upon completion of the update and promptly return the signed form and USB stick in the enclosed, return-addressed envelope (Product Support, Solta Medical, 25881 Industrial Blvd., Hayward, CA 94545) Note: Solta Medical sales personnel will assist with the software update process so, if your local representative has already updated your system, you do not need to do anything further.If you have any questions, please call Solta Product Support at (510) 259-7291. |  | 352 | 352 units | Worldwide Distribution: USA and countries including: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Egypt, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Kuwait, Mexico, Netherlands, New Zealand, Pakistan, Philippines, Poland, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, United Kingdom, and Vietnam | Terminated | November 14, 2011 | 286 | 352 | Software | Physical Safety Hazards | Software update | Software Update
Z-1017-2011 | 57661 | kit, i.v. start | LRS |  |  | General & Plastic Surgery | Enforcement Discretion | Port Dressing Tray, Reorder DT12885, Sterile, Centurion Medical Products Corp. Williamston, MI. Intended use: Vascular access port kit | Port Dressing Tray | 2 | February 01, 2011 | 2011 | Centurion Medical Products | The firm is conducting a recall of kits that contain the MONOJECT PREFILL ADVANCED flush syringes containing heparin. The syringes are under recall by their manufacturer Coviden because of testing that showed "that the crude heparin used to produce a specific batch of purified heparin sodium USP contained low level of Over-Sulfated Chondroitin Sulfate (OSCS)". | OTHER/UNDETERMINED: Pending | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated November 12, 2010, to affected customers via Certified Mail, Return Receipt Requested. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and in-use stock to determine if they have any units of the affected product, if found, immediately remove from inventory and quarantine; if they are a distributor they were ask to notify their customers of this letter; and complete and return the PRODUCT RECALL FORM whether or not they have the affected product via fax to: 203-822-6009 or scan and email to: sdfeedback@covidien.com. Note: The returned product will be collected and counted prior to destruction or return to the vendor.Should you have any medical questions concerning this recall, please call 508-261-6393.For Customer Service questions regarding the recall, please use one of the following methods: phone 1-800-962-9888 option 1 (8am to 6:30pm Eastern); email: sdfeedback@covidien.com or fax: 302-822-6009. | Contact the recalling firm for information | 2250 | 2250 kits. | Nationwide Distribution: USA including states of: OH, WI, IL, MO and TX. | Terminated | February 16, 2011 | 15 | 4290 | Not_Computer
Z-1012-2011 | 57259 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Arrow WH-02220 - Epidural Catheterization Kit with FlexTip Plus(R) CatheterIntended use: Permits acccess to epidural space, Epidural Anesthesia Kit | Arrow WH02220 Epidural Catheterization Kit with FlexTip Plus(R) Catheter | 2 | February 01, 2011 | 2011 | Arrow International Inc | Sterility may be compromised. | TRAINING: Employee Error | The firm, Arrow International, issued an "URGENT MEDICAL DEVICE RECALL" letter dated November 11, 2010 to their consignees/customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory for the affected products, cease use and distribution, quarantine all affected product immediately, return any affected product freight collect along with the completed form via mail to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, Attn: Asheboro Quality Assurance Manager, and complete and return the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM via fax to: 800-343-2935.If you require additional information or clarification regarding this matter, please contact your sales representative, or Arrow's Customer Service Department at 866-396-2111. | Contact the recalling firm for information | 2320 | 2320 | Nationwide Distribution: USA states including IL and LA. | Terminated | July 05, 2012 | 520 | 2320 | Not_Computer
Z-1022-2011 | 53014 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | PEAK brand PlasmaBlade EXT, Tissue Dissection Device,Catalogue Number: PS215-040;Product is manufactured and distributed by PEAK Surgical, Palo Alto, CAThe PEAK Surgery System is indicated for the cutting and coagulation of soft tissue during General, Plastic and Reconstructive (including but not limited to skin incisions and development of skin flaps), ENT, Gynecologic, Orthopedic, Arthroscopic, Spinal and Neurological procedures. | PEAK | 2 | February 02, 2011 | 2011 | Peak Surgical | A customer reported that a patient suffered a burn while using the device. |  | Peak Surgical notified afftected customers by phone on May 19, 2009. They were notified of the product, the problem, and the action to be taken. Peak Surgical sales representatives removed PlasmaBlade EXT devices from consignees affected by the recall. | Contact the recalling firm for information | 493 | 493 units | Worldwide Distribution - United States including CA, CO, FL, IL, MD, MN, NV, NY, OH, OK, TN, TX, UT, and WAand the countries of Turkey and Ireland | Terminated | April 21, 2011 | 78 | 493 | Not_Computer
Z-1023-2011 | 54658 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation therapy device. Medical charged particle radiation therapy system with software version. Model Numbers: Multi Plan Treatment Planning System, Software version 3.0, 3.1, 3.5.0, and 3.5.1.For the treatment planning and image guided radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | CyberKnife Robotic Radiosurgery System Device | 2 | February 02, 2011 | 2011 | Accuray Inc | When a user modified the align center of an existing plan during the treatment planning process the effective depths of each beam in the plan were not updated by the system. These depths were incorrect and led to an incorrect dose calculation. Software patch /update to address this issue. Implementation at affected sites will initiate, once the patch is available. | DESIGN: Software Design | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated January 18, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to load the plan and recalculate the dose to reflect the current state of beam, compare the reloaded plan and the recalculated plan to verify maximum dose and prescription values, and re-prescribe and save to overwrite to existing plan if necessary. The firm is developing a patch/software update to address the issue and a Accuray Customer Support person will contact the customer regarding further resolutions.If you have any questions or concerns regarding this issue, contact Accuray Customer Support at +1-877-668-8667(USA) or +1-408-716-4700 (non USA) or email: customersupport@accuray.com. | Contact the recalling firm for information | 99 | 99 units | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Russia, Italy, France, Belgium, Netherlands the UK, Spain, Athens, Saudi Arabia, Switzerland, the Ukraine, Germany, Korea and Hong Kong. | Terminated | February 11, 2011 | 9 | 99 | Software | Output/Calculation | Software update | Software Update
Z-1021-2011 | 56923 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Connect 4.0 & 4.1, radiological image processing softwareA Medical diagnostic software system intended to process, analyze, review, and distribute multi-dimensional digital images acquired from a variety of imaging devices including: CT MR, CR/DR/DX, SC, US, NM, PET, XA and RF, etc. Vital connect has the following intended use: Vessel Probe is intended for viewing the anatomy and pathology of a patient's coronary arteries. Clinicians select any artery to view the following anatomical references: the highlighted vessel in 3D, two rotate-able curved MPR vessel views displayed at angles orthogonal to each other, and cross sections of the vessel. | Vital Connect | 2 | February 02, 2011 | 2011 | Vital Images, Inc. | There is a potential alignment error of fused PET/CT images in the Vitrea Core component of Vitrea Enterprise Suite 1.2 and 1.3. The error occurs during panning of the fused PET/CT images. Because the PET data has a larger voxel size than the CT data, it is possible to pan the CT image in more frequent, smaller steps than the PET image, so that the fused PET image does not move until the CT data | DESIGN: Software Design | The firm, Vital Images, Inc,., sent a "Urgent Software Alert" letter dated August 23, 2010 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to not use the PET/CT fusion functionality until the software updates are provided and to re-examine any past results. If you have any questions or comments, contact Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 178 | 178 | Worldwide Distribution: USA including states of: AK, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, PA, SC, TN, TX VA, WV, WI, and WY; and countries of: AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, and ITALY.AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, ITALY. | Terminated | November 06, 2012 | 643 | 178 | Software | Display/Image | Software update | Software Update
Z-1026-2011 | 57555 | accessories, solution, ultrasonic cleaners for lenses | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Up&Up; Multipurpose Solution, no-rub formula, 12 FL OZ (355 mL), Sterile, Distributed by Target Corporation, Minneapolis, MN 55403Daily cleaning, removing protein deposits, rinsing, chemical (not heat) disinfection, and storage of soft (hydrophilic) contact lenses. | Up&Up; Multipurpose Solution | 2 | February 02, 2011 | 2011 | Bausch & Lomb Inc | There is a potential for some packages to not meet sterility requirements. | OTHER/UNDETERMINED: Pending | On December 8, 2010, Target was visited by FDA/PHI-DO and was informed on the violation. The Target customers that received lot GH0078 were requested to return all affected product directly to the Bausch & Lomb Greenville, SC Distribution Center. Bausch and Lomb sent an Urgent Voluntary Medical Device Recall to all first consignees who were shipped the affected lot GH0066 of Renu Fresh Multipurpose Solution on/about December 22, 2010. Since Target customers had already been notified previously regarding lot GH0078, this lot number was not included in the subsequent recall notification. All consignees were requested to review their inventory, and return any of the affected product to Stericycle Pharmaceutical Services, 2670 Executive Drive, Indianapolis, IN 46241. Stericycle is conducting this recall under the control of Bausch & Lomb.For questions abut the recall process contact Stericycle, Inc. at 1-888-345-8316.For questions regarding this recall call 1-800-828-6974. |  | 12 | 67,296 12 oz Bottles (33,648 twin packs) | Worldwide Distribution - USA including AL, CA, CT, FL, GA, HI, IA, IN, LA, MA, MI, NE, NJ, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WI and the countries Taiwan and Brazil | Terminated | September 04, 2012 | 580 | 14 | Not_Computer
Z-1024-2011 | 57464 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Laryngoscope Handle, Reusable, Rx, Product Code 4558GSP or 4559LED, Vital Signs Colorado, Inc., 11039 E. Lansing Circle, Englewood, CO 80112.Used for tracheal tube intubations during anesthesia, intensive care, and emergency medical airway management. | Greenlight II laryngeal scope handle and cartridge | 2 | February 02, 2011 | 2011 | Vital Signs Colorado Inc. | Battery which powers the laryngeal scope light may dislodge resulting in loss of illumination. | DESIGN: Device Design | Healthcare Systems sent an URGENT MEDICAL DEVICE CORRECTION letter dated December 20, 2010, to all affected consignees advising them of the possible problem and giving them interim safety instructions until new battery cartridges were available. Customers were instructed to complete and return the Confirmation Form which was included with the letter.For information regarding this recall call 973-956-5300. |  | 2202 | 2202 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, MA, MN, MO, NH, NJ, NY, OH, OK, SC, TN, TX, WA, and WV and the countries of Australia, Belgium, Canada, Dominican Republic, Ecuador, Germany, Hong Kong, Latvia, New Zealand, Taiwan, Turkey, and the UK | Terminated | July 01, 2011 | 149 | 2202 | Not_Computer
Z-1020-2011 | 56392 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT Level 2 Control Value Assignment SheetsHuman Consumption | iSTAT Level 2 Control | 2 | February 02, 2011 | 2011 | Abbott Point of Care Inc. | Three i-STAT Level 2 Control Value Assignment Sheets and the i-STAT CLEW A20 Software System Manual Update CD contain the incorrect range for pCO2, BUN and Glucose for the i-STAT EC8+, 6+, EC4+, E3 and G cartridge types. | OTHER/UNDETERMINED: Pending | Abbott Point of Care Inc., sent an URGENT RECALL NOTICE to all affected customers on 8/3/2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to replace all copies of the value sheets with the ones provided in the letter. ASLO, if they have a disc with the Value sheets on it, it is to be replaced with the updated disc and the original discarded. If the control materials, discs or value sheets have been forwarded to another facility, a copy of the letter was to be sent to them. The Business Reply card is to be completed and returned to Abbott. Questions or comments are to be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1. |  | 8293 | 8293 Value sheets worldwide | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of Argentina, Austria, Belgium, Bolivia, Brazil, Chile, China, Columbia, El Salvador, France, Georgia Republic, Germany, Hong Kong, Israel, Italy, Kuwait, Mexico, Netherlands, New Zealand, Panama, Paraquay, Peru, Poland, South Africa, Sweden, Switzerland, Taiwan, Turkey, UAE, UK, Uruquay, and Yemen. | Terminated | September 28, 2011 | 238 | 8293 | Other |  | Manual update | Other
Z-1011-2011 | 57275 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Blood Analyzer Part Number: 393-090. Intended for use by trained technologists, nurses, physicians and therapists and for use in a laboratory environment, near patient or point-of-care setting. Also, Intended for in vitro testing of samples of heparinized whole blood. | ABL90 FLEX Blood Analyzer | 2 | February 02, 2011 | 2011 | Radiometer America Inc | Due to a software error, a patient mix up can occur on the ABL90. If a sample is pre-registered and the sample "fails" the next result will inherit the patient data from the previous result. | DESIGN: Software Design | On 11/16/10 the firm sent Medical Device Field Correction letters to their customers. The letter states that all systems will be upgraded and until each system is upgraded, they instructed their customer to "always query for patient ID prior to introducing the sample." They also instruct each customer to complete the "Fax Return Form" acknowledging receipt of the letter. If customers have questions regarding the letter or the upgrade, they should contact Technical Support at 1-800-736-0600 Opt 4. |  | 432 | 432 units | Worldwide Distribution -- USA, including states of MI, NJ, & PA and countries Australia, Austria, Belgium, Denmark, Finland, France, Germany, Indonesia, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland, United Arab Emirates & United Kingdom. | Terminated | May 09, 2012 | 462 | 432 | Software | Output/Calculation | Software update | Software Update
Z-1030-2011 | 57570 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment TableThe Precise Treatment Table is intended as a universal patient treatment support and positioning table for radiation therapy and simulation use. It is for use with existing Elekta Linear Accelerators and Simulators and adaptable to third party radiotherapy products. | Precise Table | 2 | February 03, 2011 | 2011 | Elekta, Inc. | Possible error in the table longitudinal position. | OTHER/UNDETERMINED: Pending | Elekta sent an updated IMPORTANT NOTICE letter dated December 24, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Elekta recommended that all users, and Eleckta field service engineers do regular QA checks to compare the movement of the treatment table in a given axis with the digital readout. Elekta Service representatives would attach the screws with Loctite and calibrate the axis on these sites. For questions regarding this recall call 770-670-2548. |  | 597 | 597 units | Nationwide Distribution and Puerto Rico | Terminated | August 09, 2012 | 553 | 597 | Not_Computer
Z-1031-2011 | 57571 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Presice Cryosurgical System, model FPRCH2047, and the following system is not approved in the US: CryoHit-P Cryoblation System, model FPRCH2058. The system is a software controlled cryoblation system designed to destroy targeted tissue using a minimally invasive percutaneous approach. Freezing and thawing of the tissue is achieved through the use of the Joule-Thompson effect, using a high-pressure gas distributed through closed-end needles placed into the patient;s target organ.Presice is intended for cryogenic destruction of tissue during surgical procedures. Presice is indicated for use as a cryosurgical tool in the fields of general surgery, dermatology, neurology (including cryoanalgesia), thoracic surgery, ENT, gynecology, oncology, proctology, and urology. The system is designed to destroy tissue by the application of extreme cold temperatures including prostate and kidney tissue, liver metastases, tumors, skin lesions, and warts. Presice has the following specific indications: " Urology (ablation of prostate tissue in cases of prostate cancer and Benign Prostate Hyperplasia "BPH"); " Oncology (ablation of cancerous or malignant tissue and benign tumors, and palliative intervention); " Dermatology (ablation or freezing of skin cancers and other cutaneous disorders. Destruction of warts or lesions, angiomas, sebaceous hyperplasia, basal cell tumors of the eyelid or canthus area, ulcerated basal cell tumors, dermatofibromas small hemanglomas, mucocele cysts, multiple warts, plantar warts, actinic and seborrheic keratoses, cavernous hemanglomas, perianal condylomata, and palliation of tumors of the skin.); " Gynecology (ablation of malignant neoplasia or benign dysplasia of the female genitalia); " General surgery (palliation of tumors of the rectum, hemorrhoids, anal fissures, pilonidal cysts, and recurrent cancerous lesions, ablation of breast fibroadenoma); * ENT (Palliation of tumors of the oral cavity and ablation of leukoplakia of the mouth); " Thoracic surgery (ablation of arrhythmic cardiac tissue cancerous lesions); * Proctology (ablation of benign or malignant growths of the anus or rectum, and hemorrhoids). The Presice system may be used with an ultrasound device to provide real-time visualization of the cryosurgical procedure. | Presice Cryosurgiclal system | 2 | February 03, 2011 | 2011 | Galil Medical, Inc. | This voluntary correction has been initiated due to the discovery of a damaged optocoupler on one of our Presice Systems during periodic preventive maintenance. The optocoupler sends a signal across the locking mechanism in the Presice and Seednet MRI Systems, enabling the software to recognize if the channel on the system manifold is locked or unlocked. If an optocoupler is damaged or malfunction | DESIGN: Device Design | Galil Medical sent an "Urgent Medical Device Correction" letter dated December, 30, 2010, to all affected consignees. The letter described the product, the problem involved, and the action to be taken by the consignee. Consignees were instructed to make sure the locking mechanism on the Precise and Seednet MRI Systems were in the "locked" position when a needle is inserted into the manifold. Gently pull on the black plug to ensure needle is secure in locking mechanism. Apply the enclosed alert sticker on the bottom of the distribution panel and the bottom of the system monitor of the Precise System and on the side of the mobile unit itnerface panel and the top of the system monitor of the Seednet MRI System. Consignees were asked to take a picture of the alert sticker after it has been applied to the appropriate location on their Presice or Seednet MRI System(s) and e-mail to yoni.kadosh@galilmedical.com. Consignees were instructed to complete and return the enclosed form as soon as possible.For questions regarding this recall call Customer Service at (877) 639-2796.problem involved. | Contact the recalling firm for information | 63 | 63 | Worldwide Distribution - USA including CA, CT, FL, GA, MA, MD, MN, MO, NC, NY, OH, PA, SC, TN, TX, VA, and WA and the countries of CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, and JAPAN.CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, JAPAN. | Terminated | March 27, 2012 | 418 | 85 | Computer
Z-1033-2011 | 57650 | pump, infusion, pca | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂŽ-Solis Ambulatory Infusion Pump, Model Number: 2100 and 2110, Reorder Number 21-2101-00, 21-2101-0100-50, 21-2101-0100-51, 21-2101-0200-02, 21-2101-0200-03, 21-2111-0100-51, 21-2111-0200-02, and 21-2102-51. | CADDSolis | 2 | February 03, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of the Cassette Lock mechanism on certain CADDÂż-Solis Pumps becoming stuck in the locked position, making it nearly impossible to remove the medication cassette reservoirs (or administration set). The purpose of the Cassette Lock is to secure the medication cassette reservoir (or administration set) to the pump. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE CORRECTION NOTICE" dated January 13, 2011 to all consignees/customers. The notice described the product, problem, and action to be taken by the customers. The customers were instructed to have their Pumps undergo a correction which consist of removal and reattachment of the Cassette Latch handle using a new adhesive process. The customers were also instructed to remove the Pump from service if unable to remove cassette (or administration set); to inspect their inventory and identify Pumps; prepare a plan for correction of these Pumps; complete and return the attached CONFIRMATION FORM via fax to 1-800-237-8033 or by email to CADDDSolis_CassetteLock@smiths-medical.com, and if they have distributed the affected product to other persons or facilities, promptly forward this notice to the recipients. If you should have any questions regarding this information, please contact Smith Medical at 1-866-216-8806. | Contact the recalling firm for information | 2903 | 2903 (2539 US; 364 OUS) | Worldwide Distribution: USA including states of: AK, AZ, C, CO, DC, FL, GA, ID IL, MA, MI, MN, MO, MT, NE, NH, NM, NY, NC, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WI, and WY; and countries including: Australia, Austria, Belgium, Canada, France, Germany, Great Britain, Ireland, Italy, Luxembourg, New Zealand, Qatar, Saudi Arabia, Switzerland and The Netherlands. | Terminated | April 17, 2012 | 439 | 2903 | Not_Computer
Z-1037-2011 | 57721 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Irrigation Kit, Small, Latex, REF: 6113, Apogee Medical, LLC, 90 Weathers Street, Youngsville, NC 27596.Intended use: Kit components to be used during Intermittent Urinary Catheterizations. | Irrigation Kit | 2 | February 03, 2011 | 2011 | Apogee Medical | This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products of Sterile Lubricating Jelly manufactured by Triad Group. | OTHER/UNDETERMINED: Pending | The firm, TRIAD Group, sent a "Urgent Medical Device Recall" letter dated January 18, 2010 to all consignees/customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to review there inventory, cease distribution of the affected lots and to complete and return the enclosed Product Recall Response Form via fax to 1-919-570-9611 or e-mail to Apogee Medical. If they distributed the product, they were instructed to notify their customers of the recall; to include a copy of the recall notification, and to instruct them to return the recalled product to the wholesaler who would then return the product to Apogee Medical. Non-responders to the recall notice will be contacted on a weekly basis.Should you have any questions regarding this notification, please call the Regulatory Affairs/Quality Assurance Manager at (919) 435-5409. |  | 1 | see line 1. | Nationwide Distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NC, NJ, NY, OH, OK, PA, TN, TX, VA and WA. | Terminated | February 28, 2012 | 390 | 4454544 | Not_Computer
Z-1028-2011 | 57504 | monitor, breathing frequency | BZQ | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | ThermiSense model 5700-A5The ThermiSense Model 5700-A5 is a nasal cannula that detects thermal variation during breathing cycles and transmits those variations to a remote Alice-5 Polysomnograph via electrical connectors to monitors an adult patient's breathing rate in a sleep laboratory study | ThermiSense model 5700A5 | 2 | February 03, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | The incorrect length and connector were utilized in lot 110810 of the ThermiSense Alice 5, making the product unusable as it will not interface with the sleep labs equipment. | TRAINING: Employee Error | Salter Laboratories, Inc. sent an URGENT PRODUCT RECALL LETTER dated December 20, 2010, to all affected customers. The letter identfied the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for this product and lot number. Contact Salter to arrange for the return of the product. Complete the attached reply form, confirming receipt of the letter, and fax it back to Salter Labs at 1-800-628-4690.For questions regarding this recall call 661-854-3166. |  | 41 | 41 units | Nationwide Distribution including CA, CO, DE, ID, IN, KY, MD, MO, and OH, | Terminated | June 06, 2011 | 123 | 41 | Not_Computer
Z-1029-2011 | 55087 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER for Mosaiq interfaced to Electa SL75-5 machines.Sequencer Verify and Record System is to assist in the process of patient treatment on a radiotherapy treatment machine with software control. | SEQUENCER for Mosaiq | 2 | February 04, 2011 | 2011 | Impac Medical Systems Inc | Computer interface is defected. Computer based verification system for recording patient treatment on a radiotherapy treatment machine fails to record treated field. | DESIGN: Software Design | A User Notice was sent by the firm to its consignees. It identified the affected product and the reason for the recall. It also discussed the clinical impact and interim mitigations. The firm states in the letter that a fix for the problem and been developed and is available. Customers are to contact IMPAC Support for the patch to the SL75-5 interface. If customers have questions, concerns, or requests, they can email support@impac.com or one of the numbers listed in the letter for the applicable country. The letter should be distributed to all users affected by the issue at the customer's organization. |  | 40 | 40 | Worldwide Distribution -- US, Canada, Germany, and France. | Terminated | April 06, 2011 | 61 | 40 | Software | Device Operation | Software update | Software Update
Z-0993-2011 | 57507 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arstasls One Access System; Models: AAD100, AAD101, Part Numbers: FG-02279, FG-03010; The system consists of the following components: (1) Device Sheath/Anchor, (2) Shaft and (3) Handle with control features. Product is manufactured and distributed by Arstasis Inc., San Carlos, CAThe Arstasis One Access System is a sterile, single use system intended to provide access for the percutaneous introduction of devices into the peripheral vasculature and to promote hemostasis at the arteriotomy site as an adjunct to manual compression. The System is Indicated for use in patients undergoing diagnostic femoral artery catheterization procedures using SF or 6F Introducer sheaths. | Arstasls One Access System | 1 | February 04, 2011 | 2011 | Arstasis, Inc | Device has the potential to fracture and sheath separation during use. | DESIGN: Device Design | The firm, Arstasis, initiated a market withdrawal and sent notification via letters dated October 19, 2010, to customers informing the customers that Arstasis is implementing a product improvement and will replace all existing inventory that the customer has on their shelf with the improved product. The customers were instructed to work closely with their local Arstasis territory manager to ensure that their product replacement is handled satisfactorily, and complete and return the PRODUCT REPLACEMENT and PRODUCT RETURN forms to Arstasis Customer Service via fax at 1-877-688-8417 or email to customerservice@arstasis.com.If you have any questions or concerns, contact Customer Service toll free at 1-877-594-4545. |  | 330 | 330 units | Worldwide Distribution | Terminated | February 16, 2011 | 12 | 330 | Not_Computer
Z-1042-2011 | 57549 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Tube, Catalog Number 9733140, a component of the Navigated Vertex Max Drill Guide Kit, Catalog Number 9733370. Medtronic, Inc., Louisville, CO.Used in conjunction with the StealthStation navigation systems to assist in screw placement during spine surgery. | Vertex Max Drill Tube | 2 | February 04, 2011 | 2011 | Medtronic Navigation, Inc | Incorrectly manufactured drill tube could cause incorrect screw placement during spinal surgery. | DESIGN: Component Design/Selection | Consignees were notified by a Product Correction Notification letter on 12/15/2010 and given the affected product and a description of the problem. The letter stated that a Medtronic representative would contact customers to visit their facility and replace the affected guide with one that has been correctly manufactured. Additional information is available at 800-595-9709. | Contact the recalling firm for information | 15 | 15 devices | Worldwide Distribution -- USA, including states of GA, IL, IN, MN, MO, OH, OR, PA, and TX and countries of Denmark, Finland, Singapore, Spain, and UK. | Terminated | May 04, 2011 | 89 | 15 | Not_Computer
Z-1019-2011 | 57558 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva. The marketing brochure is labeled in parts: "Philips Medical Systems is part of Royal Philips Electronics... www.medical.philips.com... Philips Medical Systems Nederland B.V".Provides one or more capabilities relating to the acceptance, transfer, display, storage, and digital processing of medical images. | PCR Eleva | 2 | February 04, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | There is a risk of selecting the wrong patient, when patient is selected in the patient list and the examination tab pressed with background query are running at the same time. | DESIGN: Software Design | On 1/30/2009, Philips began sending out URGENT-Device Correction PCR Eleva notifications to their consignees.The consignees were informed of the hazards as when a patient is selected in patient list and examination tab pressed with background query running at the same time, there would be a risk that the wrong patient is selected.The firm issued the Field Corrective Order (FCO) 73200029 to their representative to correct the problem. The corrective action would be a software upgrade that resolve the issue.All consignees can contact Philips Call Center at 800-722-9377, #5, #2 and reference "FCO 73200029" with any questions. | Contact the recalling firm for information | 686 | 686 units | Nationwide Distribution -- USA, including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WI, WV, and WY and to LANDSTUHL ARMY BASE, APO in United Arab Emirates. | Terminated | August 25, 2011 | 202 | 686 | Software | Output/Calculation | Software update | Software Update
Z-1046-2011 | 56448 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnifit M/S PSL Porous Coated Acetabular Shell 40MM;Stryker Orthopaedics,Uses in primary and revision total hip arthroplasty which are indicated for painful, disabling joint disease resulting from degenerative arthritis, rheumatoid arthritis, post traumatic arthritis or late stage avascular necrosis. | Omnifit M/S PSL Porous Coated Acetabular Shell | 2 | February 07, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics determined that specific lots of the Omnifit Acetabular Shells may have a screw hole wall thickness that falls above the design specification. | DESIGN: Device Design | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated August 2, 2010, via Federal Express to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory to identify the affected product; retrieve all affected product and return it their branch or agency warehouse for reconciliation; reconcile all product(s) on the attached Product Recall Acknowledgement Form and fax to (201) 831-6069, and return the affected product via Inter-Org Transfer and ship to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  |  |  | Worldwide distribution: USA including states of: AL, AR, CA, FL, IN, ME, MI, MS, NC, NJ, OK, OR, WA , and WI; and countries including: China, Hong Kong and Korea. | Terminated | November 09, 2012 | 641 |  | Not_Computer
Z-0945-2011 | 57193 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery RX (16 slice) PET/CT System; Discovery RX (16 Slice) PET/CT Systems with technical Reference Manual numbers 2351785-xxx and 2390935-xxx distributed from November 2001 to August 2010.Head and whole body X-ray computed tomography applications. | GE Healthcare Discovery RX (16 slice) PET/CT System | 2 | February 07, 2011 | 2011 | GE Healthcare, LLC | Although the system x-ray output is accurate for the technique factors given, and the dose display is also accurate, certain CTDI adjustment factors used for the manual calculation from the Technical Reference Manual (TRM) are incorrect. CTDI100 Aperture Adjustment Factors for 4x0.625 mode contain incorrect values. This issue may lead to inaccurate CTDI manual calculation. The results of the diffe | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be provided together with the TRM addendum informing the customers of the issue and related hazards. The letter also identified the affected product and gave instructions, which include referring to the TRM addendum, along with the product correction. This correction will be executed via a Field Modification Instruction IFMI 254281 released in November, 2010 with an estimated completion date of May, 2011. Questions or concerns should be directed to the call numbers listed or a local GE Healthcare Service Representative for countries not listed. | Contact the recalling firm for information | 1450 | 1450 units total in US | Worldwide Distribution | Terminated | July 13, 2012 | 522 | 1450 | Not_Computer
Z-1064-2011 | 53054 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Linear Medical Accelerator ARTISTE MV systems, when equipped with a syngo RT Therapist, Therapist Express Basic, syngo RT Therapist Assist or Therapist Express Assist workspace with software version NET VA10A for RTT 4.1.122 or NET V10B for RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA.Intended use: medical charged-particle radiation therapy system | Linear Medical Accelerator ARTISTE MV system | 2 | February 08, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential for Mistreatment-- when treatment plans intended for non-ARTISTE machinese are loaded, if the plans contain unspecified beam parameters, attributes or unknown block codes, the therapy system may remove those parameters. | DESIGN: Software Design | The firm,Siemens, sent two Customer Advisory letters (TH004/09/S and TH008/09/S) contained in the Update Instructions to all consignees/customers with active devices. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow the updated instructions. Siemens service representative will sign a document noted as CORRECTED upon completion of the software installation.If you have any questions, call 1-925-602-8157. |  | 151 | 151 units | Worldwide distribution: USA and countries including: Germany, France, Poland, Norway, Australia, Malaysia, Spain, New Zealand, the UK, India, Republic of Korea, Canada, Ireland, Italy, Egypt, Belgium, Philippines, Saudi Arabia, Netherlands, Hungary, Norway and South Africa. | Terminated | April 12, 2011 | 63 | 151 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1063-2011 | 57531 | needle, aspiration and injection, disposable | GAA | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Inrad sample master aspiration biopsy needle with hiliter sharp bevel needle, 22 ga x 2 3/8'' (6 cm), model 41222, Inrad, Inc. Kentwood, MI.The device is intended for use during percutaneous biopsy and drainage procedures to obtain tissues or fluid specimens. | Inrad sample master aspiration biopsy needle | 2 | February 08, 2011 | 2011 | Inrad, Inc | This lot of product has the wrong expiration date on the carton label. The pouches contained the correct expiration date. |  | The firm, Inrad Inc., sent an "URGENT FIELD CORRECTION" notice dated December 20, 2010 to its customers. The notice described the product, problem and actions to be taken by the customers. The customers were instructed to immediately examine their inventory for the affected product. If the recalled product was in stock, the customers were told to removed the affected product from the carton and dispose of carton; to disregard carton label expiration date and reference the expiration date appearing on the pouch label only, and to complete, sign and return the attached "Carton Expiration Date" INFORMATION FORM via fax to 616-301-779 or mail to Inrad Inc., ATTN: Director RA/QA, 4375 Donker Court SE, Kentwood, MI 49512.If you have any additional questions about this action, please call 800-558-4647 extension 102. | Contact the recalling firm for information | 250 | 250 | Nationwide Distribution: USA including states of: FL, GA, KY, NH, NJ, NM, PA , SC, WA, and WI. | Terminated | February 09, 2011 | 1 | 250 | Not_Computer
Z-1045-2011 | 57610 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Linear Accelerator:Varian Medical Systems, Palo Alto, CA 94304Indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Varian High Energy Linear Accelerator | 2 | February 08, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified whereby the bolts used to fasten the counterweight to the gantry may not, in some cases, be fully tightened to the required torque specification. | PRODUCTION CONTROLS: Process Control | An Urgent Medical Device Correction letter was sent to all affected users with a description of the problem and user corrective action steps. Letters were sent on 12/10/2010. Users are to immediately investigate any unusual noises coming from the affected product and report them to whoever is responsible for servicing the device. For customers that do have a service contract with Varian, Varian service will perform an inspection of the counterweight bolts as part of the regular Periodic Maintenance Inspection procedure. Customers are to advice the appropriate personnel of the contents of the letter. For further clarification, customers can contact their local Varian Customer Support Services Manager. | No consumer action necessary | 1882 | 1882 | Worldwide Distribution. | Terminated | August 02, 2011 | 175 | 1882 | Not_Computer
Z-1000-2011 | 57524 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Diabetes Care, ReliOn Ultima Blood Glucose Test ; | Abbott Diabetes Care, ReliOn Ultima | 1 | February 09, 2011 | 2011 | Abbott Diabetes Care, Inc. | Certain lots of Precision Xtra, Precision Xceed Pro, MediSense Optium, Optium, OptiumEZ and ReliOn Ultima Blood Glucose Test Strips may generate erroneously low blood glucose results. | PRODUCTION CONTROLS: Process Control | On 12/22/10, Abbott Diabetes Care notified customers via Press Release and letters explaining the reason for the recall and requesting the products be returned. |  | 359 | 359 total lots, 1,000,000 strips per lot | Products were released for distribution in US and Puerto Rico. | Terminated | October 25, 2011 | 258 | 359 | Not_Computer
Z-1065-2011 | 57333 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Bond Polymer Refine Detection Kit; an in vitro diagnostic immunohistochemistry reagent kit; Catalog No. DS9800;Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom.Intended for immunohistochemistry on paraffin-embedded tissue sections. | Bond Polymer Refine Detection | 2 | February 09, 2011 | 2011 | Leica Microsystems, Inc. | Users of the Bond Polymer Refine Detection System, DS9800, may experience instances of "weak staining" due to the DAB Part B component, which may give a false negative result. | OTHER/UNDETERMINED: Pending | Leica Microsystems sent Medical Device Field Correction letters dated 11/16/10 to the affected users of the Bond Polymer Refine Detection System, DS9800 via first class mail on the same date, informing them that they may experience instances of "weak staining" due to the DAB Part B component in the affected lots, which may give a false negative result. The users were requested to test their DS9800 System prior to use to rule out that the DAB Part B components have been affected, using the enclosed testing protocol. If their system is affected, the users were instructed not to use it and contact Leica Technical Assistance Center for a replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 653 | 653 units | Worldwide Distribution -- USA and Canada. | Terminated | October 17, 2011 | 250 | 653 | Not_Computer
Z-1066-2011 | 57321 | nadh oxidation/nad reduction, alt/sgpt | CKA | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abaxis brand Piccolo Comprehensive Metabolic Panel; Used with the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer; Catalog/Part Number: 400-0028; Product is manufactured and distributed by Abaxis, Incorporated, Union City, CA The Piccolo Comprehensive Metabolic Reagent Disc, used with the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer, is intended to be used for the in vitro quantitative determination of alanine aminotransferase (ALT), albumin, alkaline phosphatase (ALP), aspartate aminotransferase (AST), calcium, chloride, creatinine, glucose, potassium, sodium, total bilirubin, total carbon dioxide, total protein, and blood urea nitrogen (BUN) in heparinized whole blood, heparinized plasma or serum. The patented selfÂŽcontained reagent rotors are bar coded, contain liquid diluents and dry reagents and use centrifugal and capillary technology to provide the most commonly requested chemistry panel results in approximately 12 minutes. The Piccolo Comprehensive Metabolic Panel, used with the Piccolo Blood Chemistry Analyzer is a diagnostic system that aids in disease diagnosis or monitoring disease treatment. | Piccolo Comprehensive Metabolic Panel | 2 | February 09, 2011 | 2011 | Abaxis, Incorporated | Abaxis has determined that during the manufacture of product lot 0396BC5 may report inaccurate values of Alanine Aminotransferase (ALT) when used in the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer. | OTHER/UNDETERMINED: Pending | The firm, Abaxis Inc., sent an "URGENT PRODUCT RECALL NOTICE" dated November 19, 2010, to all customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory of the lot for boxes #0665 or higher; if found, to discontinue use; contact Abaxis for immediate replacement; complete and return the enclosed PRODUCT RECALL RESPONSE form to Abaxis Technical Support via fax at 1-877-349-2087, and ensure that their laboratory staff is informed of this recall notice.If you have any questions regarding this Product Recall Notice, please contact Abaxis Medical Technical Support at 800-822-2947, or at medtechsupport@abaxis.com. |  | 797 | 797 units | Worldwide distribution: USA and countries of: Kowloon and Germany. | Terminated | August 23, 2011 | 195 | 797 | Not_Computer
Z-1067-2011 | 57676 | [6 empty] | Fujinon Endoscope Series;Model # EPX-4400 and EPX-4400HD.Digital processing of images from endoscope. | EPX4400, EPX4400HDVideo Processor | 2 | February 09, 2011 | 2011 | Fujinon Inc | The FICE feature installed in video endoscopic processors is being disabled since this feature should have been reviewed under the premarket notification/510k process. | PREMARKET APPROVAL: No Marketing Application | Fujinon sent an Urgent Medical Device Corrective Action Letter, dated 1/7/2011, via overnight mail to all consignees. The letter advises users that the FICE feature should no longer be used and that a Fujinon Field Service Engineer will be coming to physically disable the FICE feature on the processor. Instructions labeled as Attachment B are provided to assist in ceasing use of the FICE function. Attachment A is provided for frequently asked questions. Customers are to complete the Response/Acknowledgement Form and either fax it or email back to the firm. If there are additional questions, customers can call 800-385-4666 or email products.fice.@fujinon.com. | Contact the recalling firm for information | 783 | 783 units | Nationwide Distribution | Terminated | June 04, 2012 | 481 | 783 | Not_Computer
Z-1081-2011 | 56391 | kit, i.v. start | LRS |  |  | General & Plastic Surgery | Enforcement Discretion | Centurion Medical Products PICC Line Kit CV12880, Sterile. | Centurion Medical Products | 2 | February 09, 2011 | 2011 | Centurion Medical Products | The kits contain heparin lock flush solution USP code 504505 that is under recall by APP Pharmaceuticals. APP Pharmaceuticals is recalling the heparin because of incomplete documentation associated with test results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Centurion Medical Products sent out Urgent Product Recall Notices on 6/3/2010. A second recall notice was sent to nonresponding consignees on 6/21/2010. The letter instructed the consignees to quarantine the recalled product and contact Centurion so that the recalled product can be retrieved. | Contact the recalling firm for information | 60 | 60 | Nationwide distribution. | Terminated | February 09, 2011 | 0 | 4160 | Not_Computer
Z-1094-2011 | 57202 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald, CELL-DYN Emerald CPU Board Hematology System; List Number: 09H39-01; Field Replaceable Unit (FRU) Numbers: 8701686901, 8701686902;Product is distributed byAbbott Diagnostics Division, Santa Clara. CA and manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, ILIntended use: The CELL-DYN Emerald is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | CELLDYN Emerald | 2 | February 10, 2011 | 2011 | Abbott Laboratories | There is the potential for the instrument to not power up due to the flash memory issue. This will result in an error message displayed is "No Memory Available" and has the potential for delay in generating patient results. | OTHER/UNDETERMINED: Pending | The firm, Abbott Laboratories, sent a "Product Correction" letter dated October 15, 2010 to all customers. The letter described the product, problem and action to be taken by the customer. The customers were instructed to install a new printer driver in the software location using the provided instructions and to complete and return the Customer Reply Form via fax 1-800-777-0051 or email to QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT; customers outside the U.S., please contact your local hematology customer support representative. |  |  |  | Worldwide distribution: USA including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, O, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI; and countries including: Argentina, Australia, Brazil, Canada, Dominican Republic, Germany, Singapore, South Korea, St. Eustalius & St. Martin, Trinidad and Tobago. | Terminated | June 07, 2011 | 117 |  | Hardware | Device Operation | Install a new printer driver | Other
Z-1061-2011 | 57651 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic SynchroMedimplantable drug pump Refill Kits (Model No: 8551, 8555, 8561, 8562, 8564, 8565, 8566 The Model 8551, 8555, 8561, 8562, 8564, 8565, and 8566 Refill Kits are intended for use in refilling Medtronic implantable infusion pumps with the exception of Medtronic MiniMed infusion pumps | Medtronic SynchroMed implantable drug pump | 1 | February 10, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | This letter provides important reminders concerning the potential for a pocket fill during a SynchroMedÂż II or SynchroMed EL implantable drug pump refill procedure, and important patient management recommendations that will be added to our product labeling. A pocket fill is the inadvertent injection of all or some of the prescribed drug into the patients subcutaneous tissue, which includes the pum | OTHER/UNDETERMINED: Pending | The firm, Medtronic, sent an "Urgent Medical Device Correction Cover " letter dated January 2011 to all customers. Enclosed with each letter, an "Urgent Medical Device Correction Important Clinical Information about Pocket Fills" letter dated January 2011 as well as a Clinician Refill Reference Card, Refill Appointment Cards, Physician Reply Card and a Self Addressed Stamped Envelope. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to always check the needle placement within the pump septum throughout the procedure to help perform a successful refill; be aware that the pump refill difficulty may increase due to a variety of factors (listed); monitor the patient closely for signs and symptoms of overdose according to the provided instructions. Note: Medtronic is in the process of supplementing the labeling related to improper injection with the information included in this communication.Patients with questions are encouraged to talk with their physician or contact Medtronic Patient Services at 1-800-510-6735, Monday - Friday, 8 a.m. to 5 p.m. CDT. Physicians with medical questions related to this issue or Medtronic therapies should contact Medtronic at 1-800-328-0810, Monday - Friday, 8 a.m. to 5 p.m. CDT. UPDATE - June 2013 - Medtronic Field Representative began notifying Healthcare Professionals June 3,2013,via visit to deliver and discuss the Pump Refill Procedure Safety Update letter. The letter informs clinicians that SynchroMed Implantable Infusion System labeling and the Clinician Refill Reference Card have been updated. The labeling and reference card were updated to provide important reminders concerning the potential for a pocket fill during a SynchroMed II or SynchroMed EL implantable drug pump refill procedure. The Medtronic Rep will complete a Physician Notification Confirmation form documenting the delivery of the information. A copy of the recall notification letter a. |  | 91369 | 91,369 total for all devices. | Worldwide distribution: USA including DC and PR and countries including: AUSTRALIA, NEW ZEALAND, HONG KONG, TAIWAN, CHINA, SINGAPORE, MALAYSIA, INDONESIA, SOUTH KOREA, THAILAND, ARUBA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, FAROE ISLANDS, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JORDAN, KUWAIT, LEBANON, LITHUANIA, LUXEMBOURG, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTI LES, NORWAY, PANAMA, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, SAN MARINO, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VATICAN CITY. | Open |  |  | 91369 | Not_Computer
Z-1222-2011 | 56860 | cutter, surgical | FZT | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Maquet HEARTSTRING Aortic Cutter 4.3 mm, AC-3043, Lot numbers 9081071 and 9082071, Manufactured by Maquet Cardiovascular, Wayne, New Jersey.Intended for use by cardiac surgeons during CABG procedures to create an opening in the wall of a vessel. | HEARTSTRING Aortic Cutter 4.3 mm | 2 | February 11, 2011 | 2011 | Maquet Cardiovascular | Pouch integrity may be compromised, resulting in loss of sterility. | PRODUCTION CONTROLS: Process Control | Urgent Device Removal letters were sent out on June 10, 2020 by Federal Express. The letter identified the affected product, described the reason for the recall, and asked for all affected product to be returned. The firm asked customers to examine their inventory and to discontinue distributing the affected lots. In addition, customers are to complete and return the Field Action Response form. If the affected product was further distributed, those customers also need to be notified of the recall and return their affected product to their distributor. All affected returned products may be exchanged with other unaffected HEARTSTRING Aortic Cutter 4.3 MM devices. Customers are to contact their local MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874 to arrange for an exchange and if they have any questions. | Contact the recalling firm for information | 395 | 395 units | Worldwide Distribution -- USA, including states of IA, KY, IL, WI, TX and NC and countries of Japan, Germany and Hong Kong. | Terminated | June 21, 2011 | 130 | 395 | Not_Computer
Z-1220-2011 | 57733 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bed, P1900, Hill-Rom, Batesville, IN. TheTotalCare Bed System is intended to porvide patient support ideally suited to be used in health care environments. The TotalCare Bed System may be used in a variety of settings including but not limited to, acute care, including critical care, step down/progressive care medical/surgical, high acuity sub-acute care, post anesthesia care unit (PACU) and sections of the emergency department (ED). | TotalCare AC powered adjustable hospital bed | 2 | February 11, 2011 | 2011 | Hill-Rom, Inc. | The firm received one report that a patient was weighed on the bed at 53 Kg when the patient actually weighed 73 Kg. The patient expired the next day and it is inconclusive as to whether the scale was a contributing factor. The firm investigated and found that the software for the scale system did not clear the last patient weight from the history of the scale system. | DESIGN: Software Design | The firm, Hill-Rom, sent a "URGENT MEDICAL DEVICE CORRECTION" letter dated January 19, 2011, along with a software upgrade (USB stick/flash drive) with instructions, to all affected customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to requests that facility service personnel immediately upgrade their software; complete and return the Customer Response Form via fax to 812-931-2410 or mail to: Hill-Rom Company, 1069 State Route 46 E, Batesville, IN 47006, Attn: MOD Coordinator J12 and forward a copy of this notice to any other facility personnel. Note: The customer can request that Hill-Rom technicians complete the software upgrade.If you have any questions concerning this Urgent Medical Device Correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 905 | 905 beds. | Worldwide distribution: USA and countries including: AUSTRALIA, BRAZIL, CANADA, CHINA, EGYPT, GERMANY, GREECE, ITALY, JAPAN, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, OMAN, POLAND, QATAR, SAUDI ARABIA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 26, 2012 | 593 | 945 | Software | Output/Calculation | Software update | Software Update
Z-1093-2011 | 57405 | instrument, coagulation | KQG | Coagulation instrument. | Hematology | Hematology | 510(k) | Hemochron Whole Blood Coagulation Systems Activated Clotting Time (ACT); Package label states -International Technidyne Corporation Response Celite ACT Tube HRFTCA510; Lot number and expiration date are listed. Each box contains 95 ACT tubes tests.Used for heparin anticoagulation monitoring during cardiopulmonary bypass surgery, percutaneous transluminal coronary angioplasty (PTCA), interventional radiology, extracorporeal membrane oxygenation (ECMO), hemofiltration, hemodialysis and critical care. | Hemochron Whole Blood Coagulation Systems Activated Clotting Time (ACT) | 2 | February 11, 2011 | 2011 | International Technidyne Corp. | Some Activated Clotting Time (ACT) Tubes may contain a component that does not meet its specification, making it possible for a failure to detect a tube or the presence of a clot. | PRODUCTION CONTROLS: Process Control | ITC notified their distributors and end user customers with an Urgent Medical Device Recall letter dated November 29, 2010 through UPS overnight mail. The letter identified the affected product and explained the reason for the recall. Distributors and customers should check their inventory for any affected product. Distributors should immediately discontinue shipment of the affected product and place the product on hold. Then regardless of whether they have any remaining inventory, they should complete the enclosed Distributor Account Tracking Form. The firm will contact them to arrange for any affected product to be returned and will credit them for any unused, returned product. Customers are to stop using the affected product and remove it from their inventory. They should also complete the attached Customer Account Tracking Form and return it to the firm. The firm will contact them to arrange for the return of any affected inventory and will credit the customers for any unused, returned product. If customers believe that they have experienced a problem with the product, they should take any additional actions regarding their records or patients. Questions should be directed to ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International) or e-mail techsupport@itcmed.com. | Contact the recalling firm for information | 265 | 265 boxes (95 ACT Tubes per box) | Worldwide Distribution | Terminated | August 22, 2012 | 558 | 265 | Not_Computer
Z-1224-2011 | 57597 | pump, breast, powered | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pump in Style Advanced (PNSA)Breast Pump; a daily use, double pump, electric powered breast pump; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; the affected breast pumps were sold in the following configurations: a) article 57027 - Kit PNSA Shoulder Bag; b) article 57027T - Kit PNSA Shoulder Bag Tagged; c) article 57036T - Kit PNSA The Metro Bag Tagged; d) article 57043 - PNSA without kit; e) article 57062 - Kit PNSA Backpack; f) article 57062-03 - Kit PNSA Backpack CA/3; g) article 57062T - Kit PNSA Backpack Tagged;h) 57102 - Motor Replacement PNSA Backpack 57062; i) 57103 - Motor Replacement PNSA Shoulder Bag 57027Intended use: A powered breast pump for the removal, collection and storage of mother's breast milk. | Pump in Style Advanced Breast Pump | 3 | February 11, 2011 | 2011 | Medela Inc | The performance of the pumps may be compromised due to the potentiometer used on the circuit board. Affected pumps may not turn on, may turn off on its own, may reduce or lose suction during the pumping process, or may make a clicking noise. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | This firm, Medela, sent an "Important Device Recall" letter dated January 11, 2011 to all of direct accounts via UPS or US Mail. The accounts were informed that a component on the control board, after a period of usage, may cause the motor to slow down or stop, resulting one or more of the following symptoms: the pump will not turn on; the pump turns off on its own; the pump reduces or loses suction during the pumping process or the pump is making a clicking noise. If the users have a pump that exhibits any of these symptoms, they were requested to visit Medela's website at www.medelareturns.com to expedite shipment of a replacement unit. Questions about the recall process were directed to Medela Customer Service at 1-800-435-8316 or e-mail at medelareturns@medela.com.Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelareturns.com on January 11, 2011. | Contact the recalling firm for information | 11154 | 11,154 units | Nationwide, including Puerto Rico | Terminated | May 21, 2012 | 465 | 11154 | Hardware | Device Operation | Replace device | Remove or Replace
Z-1219-2011 | 57730 | heart-valve, allograft | MIE |  |  | Cardiovascular | Contact ODE | Pulmonary Valve & Conduit SGUsed in heart surgery | SynerGraft CryoValve | 2 | February 11, 2011 | 2011 | CryoLife, Inc. | A pre-processing culture performed on a companion allograft (Aortic Valve & Conduct SG) tested positive for Staphylococcus aureus. |  | The firm, Cryolife, sent an "Urgent-Human Tissue Recall" letter dated January 7, 2011 to the consignee/customer. The letter describe the product, problem and action to be taken by the customer. The customer was instructed to determine if any additional patient surveillance or intervention was necessary, and to complete and return the enclosed postcard affirming that the notification was received.If you have any questions, please contact Cryolife's Medical Director at 1-678-290-4403. |  | 1 | One | Nationwide distribution: USA state of: MI | Terminated | November 01, 2011 | 263 | 1 | Not_Computer
Z-1095-2011 | 57660 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Femoral Distal Fixation Peg; Catalog Number 5575-X-00; Manufactured by Stryker Howmedica Osteonics Corp., Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 USA A Subsidiary of Stryker Corp. One single unit to a package.Triathlon Distal Femoral Fixation Pegs are made available separate from the Triathlon Posteriorly Stabilized Femoral Component. These distal femoral fixation pegs are an accessory to the Triathlon Femoral Component and are optional for use. These pegs are designed to provide rotational stability and to aid the surgeon in the placement of the femoral component on the prepared distal femur. The Triathlon Distal Femoral Fixation Pegs are fabricated from wrought cobalt alloy. The distal portion of the fixation peg is thread to allow the peg to be assembled to the distal threaded hole in the Triathlon femoral component. The proximal portion of the peg is recessed to allow the use of a wrench to assemble the peg into the femoral complement. | Triathlon Femoral Distal Fixation Peg | 2 | February 11, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the labels on the Triathlon Distal Femoral Fixation Pegs contain an incorrect expiration date of January 2015 (2015-1) when the actual expiration date should read June 2014 (2014-6). | EXPIRATION DATING: Incorrect or No Expiration Date | The firm, Stryker, sent a "UPDATED (Corrected Lot Code) URGENT PRODUCT RECALL" letter dated January 6, 2011, via Fed Ex ,with return receipt to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to fax the attached Product Recall Acknowledgement form to (201) 831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product, and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team at (201) 831-5028. |  | 15 | 15 units | Nationwide distribution. | Terminated | October 03, 2012 | 600 | 15 | Not_Computer
Z-1232-2011 | 57505 | drills, burrs, trephines & accessories (simple, powered) | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 5190 TPS Series Burs. Powered simple cranial drills, burrs, trephines, and their accessories. These burs are intended for use during various surgical procedures to cut hard tissue or bone and soft tissue. More specifically these burs are used for shaping the bone (femur) during orthopedic surgeries. | Stryker 7.5mm x 7.9mm XX Coarse Diamond Drum, Super Long | 2 | February 14, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | Product was packaged with an incorrect packaging material and therefore has a potential to be non-sterile which, if contaminated, could lead to a patient infection. | PRODUCTION CONTROLS: Packaging | The firm, Stryker, sent a "MEDICAL DEVICE RECALL NOTIFICATION" letter dated December 2, 2010 to all customers. The letter described the product, problem and the action to be taken by the customers. The customer were asked to complete and fax in the business reply form indicating they have received the notification and to indicate whether or not they have affected product on hand, and send back any affected product to Regulatory Affairs at Stryker Instruments. Stryker will then send a replacement product to the customer.Upon receipt of the completed business reply form, the customer will be contacted to arrange return of any affected product they have in their inventory. Once the affected product is received, a replacement 5190 TPS Series Bur will be sent to the customer.For questions regarding this recall, please contact Stryker Instruments at 1-800-800-4236 ext. 3584 or kathleen.kelly.smith@stryker.com. |  |  |  | Worldwide distribution: USA including states of: CA, FL, LA, MI, MT, NE, NH, OH, PA, SD, TN, TX, VT, and WA; and countries including: Sweden and Australia. | Terminated | April 13, 2011 | 58 |  | Not_Computer
Z-1231-2011 | 57550 | instrument, shunt system implantation | GYK | Shunt system implantation instrument. | Neurology | Neurology | 510(k) | Integra Peritoneal Introducer Sheaths, 61 cm;Integra NeuroSciences/REF 901-124; Manufacturer Integra NeuroSciences Implants S.A., 2905 Route Des Dolines, 06921 SOPHIA ANTIPOLIS Cedex, France; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, NJ 08536, USADesigned to be used with a Peritoneal Introducer of appropriate length during the placement of a peritoneal drainage catheter in the treatment of hydrocephalus.Designed to be used with a Peritoneal Introducer of appropriate length during the placement of a peritoneal drainage catheter in the treatment of hydrocephalus. | Integra Peritoneal Introducer Sheaths | 2 | February 14, 2011 | 2011 | Integra LifeSciences Corp. | Peritoneal Introducer Sheaths were incorrectly labeled as 61 cm instead of 46 cm. | PRODUCTION CONTROLS: Labeling Mix-Ups | Integra Life Sciences sent an Urgent Product Recall letter dated December 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine any of the affected product. Contact Integra Customer Service at (800) 654-2873, select option 2 to arrange for product returns. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating the current status of quantities of affected product in their inventory. For questions call (609) 936-2237. | Contact the recalling firm for information | 95 | 95 units | Nationwide Distribution including CO, FL, GA, MO, NC, NJ, PA, TX, UT, and WI. | Terminated | May 07, 2012 | 448 | 95 | Not_Computer
Z-1225-2011 | 57679 | adaptor, stopcock, manifold, fitting, cardiopulmonary bypass | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products, Y-Glide (Adaptor, Stopcock, Manifold, Fitting, Cardiopulmonary Bypass) for use during vascular interventional procedures. Part No. FTB-021-00, Catalog No. FTB-021-00. | YGlide (Adaptor, Stopcock, Manifold, Fitting, Cardiopulmonary Bypass) | 2 | February 14, 2011 | 2011 | Thomas Medical Products Inc | Compromised sterility/weak or open seals. | PRODUCTION CONTROLS: Packaging Process Control | GE Healthcare Thomas medical products Recall notification dated 16 December 2010 (GEHC Ref # 35003) was sent by letter via UPS Next Day Air and/or USPS certified mail. |  | 28679 | 28,679 for all products | Distribution Nationwide USA and Germany. | Terminated | August 08, 2012 | 541 | 28679 | Not_Computer
Z-1230-2011 | 57735 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Bipolar cord for use with the VirtuoSaph Endoscopic Vein Harvesting System, MCBICORD1, Terumo Cardiovascular System, Ann Arbor, MI.The Terumo Bipolar cord MCBICORD1 is a flexible, reusable power cable that is indicated for use with the Terumo VirtuoSaph Endoscopic Vein Harvesting System and a compatible generator during saphenous vein harvesting procedures. | VIRTUOSAPH (TM) ENDOSCOPIC VEIN HARVESTING DISPOSABLE SYSTEM | 2 | February 14, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Bipolar cord product specification indicates the product has a service life of 50 uses. However the IFU does not specify an end of service life. The firm has recieved 46 complaints on the Bipolar cord. The firm investigated and found that the units involved in the complaints were often over 12 months old. The firm believes the units involved in the complaints were beyond their product life | OTHER/UNDETERMINED: Pending | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE CORRECTION letter dated January 10, 2011, to Operating room managers. The letter informed the customers of the Bipolar cord product life of 50 uses. The letter states that the IFU for the bipolar cord is being updated to include the information regarding the limit of 50 uses and sterlization cycles. The letter recommends that user keep a sterile back up bi-polar cord in a convienent place during endoscopic vein harvesting procedures. The firm provided a tracking sheet in each package to help users track the number of sterlization cycles for the cord. Customers were instructed to review the Medical Device Correction notice, assure that all users were aware of the notice, and confirm receipt of the communicaiton by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call Terumo CVS Customer Service at 1-800-521-2818. | No consumer action necessary | 2782 | 2782 cords | Worldwide Distribution - United States including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, , NY, MO, MS, MT, NC, NE, NH, NJ, NM, OH, OK, OR, PA, SC, TN, TX, VA, WI, AND WV and the countries of Chile, Hong Kong, Malaysia, United Arab Emirates, Belgium, Japan, Korea, The Phillipines, Singapore, Taiwan, Thailand, Germany, and United Kingdom. | Terminated | July 13, 2011 | 149 | 2782 | Not_Computer
Z-1223-2011 | 57607 | system, radiation therapy, radionuclide | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan, Model 5.34.Intended to be used for planning the dosimetry of treatments in stereotactic radiation therapy. | Leksell GammaPlan | 2 | February 14, 2011 | 2011 | Elekta, Inc. | Investigation found that if a user accidentally selects the wrong image/tube position images can be displayed flipped. | OTHER/UNDETERMINED: Pending | Product bulletin Leksell GammaPlan - "Leksell GammaPlan version 5.34" (dated October 2005) was sent to affected Customers as notification of the problem. Field Change Order FCO-DCO-05-0006 (dated November 22, 2005) and Technical Notes Leksell GammaPlan Leksell SurgiPlan Important Notice - "Potential issue using flipped angiographic images" (dated October 2005) was sent to Elekta Service Engineers as notification of the problem. The letters described the issue and the FCO asked the Service Engineers to confirm their receipt of the documentation. The Technical Notes also discussed image visualization requirements and listed recommendations. | Contact the recalling firm for information | 272 | 272 units | Worldwide Distribution | Terminated | March 02, 2011 | 16 | 272 | Other |  |  | Other
Z-1229-2011 | 56058 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | AMS brand GreenLight PV Laser System, Model-P/N: 0010-9230;Product is manufactured and distributed by AMS Solutions,San Jose, CAThe GreenLight Surgical Laser System and accessories are intended for use in endoscopic (cystscopic) KTP 532 nm laser resection of the prostate for the treatment of benign prostatic hyperplasia (BPH). The entire laser unit and controls are contained in a single console that is connected to external electricity and water sources. Fibers are sterile, single-use, disposable fiber optic delivery systems, designed to deliver 532nm laser energy, the fiber is connected to the laser console and used to cut, coagulate and vaporize prostate tissue. | AMS GreenLight PV Surgical Laser System | 2 | February 14, 2011 | 2011 | AMS Innovative Center - San Jose | A failure mode whereby the fiber attachment to the GreenLight PV laser system could inadvertently be broken. If fiber were to bend sharply at point of fiber connection, it may break and cause a burn. | OTHER/UNDETERMINED: Pending | AMS Innovative Center - San Jose issued Important Product Update Information letters via express mail on February 8, 2010, and September 15, 2010 to all its direct consignees. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to attach the Fiber Guard to their GreenLight PV Laser System, and to complete the Response Form in the enclosed, self-addressed, postage paid envelope to AMS. For any questions regarding this recall call the Customer Care Department at 1-800-356-7600 or send an e-mail to CustomerCare@ammd.com. | Contact the recalling firm for information | 242000 | ~242,000 units | Worldwide Distribution including US | Terminated | February 14, 2011 | 0 | 242000 | Not_Computer
Z-1258-2011 | 51399 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Monitoring System Test Strips and Freestyle Lite Blood Glucose Monitoring System Test Strips, for use with FreeStyle and FreeStyle Lite Test Strips respectively, 100 count, manufactured by Abbott Diabetes Care, Alameda, CA.Product is used with FreeStyle and FreeStyle Lite Blood Glucose Meters to measure glucose (sugar) in whole blood. The test strips are used for testing outside the body (in vitro diagnostic use). The FreeStyle Blood Glucose Monitoring Systems are intended for use in the home and in professional settings to monitor blood glucose levels. | FreeStyle and FreeStyle Lite Blood Glucose Monitoring Test Strips | 2 | February 15, 2011 | 2011 | Abbott Diabetes Care, Inc. | Product dispositioned for destruction was potentially diverted to end users. | OTHER/UNDETERMINED: Pending | The firm, Abbott Diabetes Care (ADC), will be posting notification on its Abbott Diabetes Care website. Also, the firm issued an "IMPORTANT: MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2009, to customers ( registered users of Free Style, Free Style Flash, FreeStyle Freedom or Free Style Navigator). For foreign consignees/customers, the firm has issued communication worldwide. The letter described the product, problem and action to be taken by the customers. The customers were instructed to please check the lot number on the test strip carton or vial for the affected product number; if found, discontinue use and call the ADC Customer Care line at: 1-800-310-1500 and the ADC representative will provide information on how to return the test strips and arrange to have replacement test strips sent.If you have any questions, please call ADC Customer Care line at 1-800-310-1500. |  | 1860 | 1860 units | The firm, ADC, quarantined the eight lots and disposed them to scrap. These eight lots were not released from the firm ADC's control; therefore, the firm was unable to determine the number of consignees. | Terminated | February 15, 2011 | 0 | 1860 | Not_Computer
Z-1259-2011 | 54746 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II Turbo Link Hemoglobin Testing System, including Hemoglobin A, Program Reorder Pack 1600 Tests, Program Buffer A, Buffer B, Cartridge Set, Analytical Cartridge, Guard Cartridge, Calibrator, Accessory Box, Instruction Manual and Quick GuideIntended use: Measurement of Hemoglobin A1c in human whole blood. | BioRad Variant II Turbo Link Hemoglobin Testing System | 2 | February 15, 2011 | 2011 | Bio-Rad Laboratories Inc | The results from the blood device had the occurrence of some ramping baselines on their chromatograms. The ramping baseline can affect quantitation of the blood results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Bio-Rad Laboratories, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 5, 2010, to all customers and letter dated February 16, 2010 to subsidiaries via Fed-Ex or by fax. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review there chromatograms form the lots listed and compare them to what documented; follow the noted instructions; destroy reorder packs from the affected lots if a ramping baseline issue is observed, and complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, attention: Bio-Rad CSD Regulatory Affairs Department. If you have any questions, please contact your regional Bio-Rad Office or (510)724-7000. |  | 27 | 27 | Worldwide distribution: USA including states of: OH, MN, SC and VA; and country including: Japan. | Terminated | March 03, 2011 | 16 | 27 | Not_Computer
Z-1249-2011 | 56182 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Radiotherapy delivery system (C-Series Clinac, Trilogy and Novalis Tx), model numbers H14, H17, H18, H23, H26, H29, HCX | CSeries Clinac, Trilogy and Novalis Tx | 2 | February 15, 2011 | 2011 | Varian Medical Systems Oncology Systems | Position errors may result in misdelivery of radiation. | DESIGN: Device Design | Varian sent Urgent Medical Device Correction letters on June 23, 2010 by Federal Express to all customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Field Safety notification including customer distribution lists, was e-mailed to Varian international Customer Support Managers. Managers will translate to local language if necessary, notify sites by appropriate methods. |  | 5741 | 5741 units | Worldwide distribution. | Terminated | May 03, 2011 | 77 | 5741 | Not_Computer
Z-1248-2011 | 57485 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex Dental Systems, Gendex eXpert DC Intraoral X-Ray System; a dental x-ray system; Model 110-0205G1, catalog number EXPERTDC75NA | Gendex eXpert DC Intraoral XRay System | 2 | February 15, 2011 | 2011 | Gendex Dental Systems | The intraoral x-ray systems were installed improperly with incorrect screws which allow the unit to pull away from the wall when the arm is fully extended. This condition may harm the patient. | DESIGN: Component Design/Selection | Gendex sent a Urgent Medical Device Field Correction letter dated December 17, 2010 to the consignee, informing them that the x-ray units were installed with improper sized lag bolts and requesting that they contact Gendex Customer Service at 1-800-323-8029 to complete the field correction. |  | 5 | 5 units | California | Terminated | February 15, 2011 | 0 | 5 | Not_Computer
Z-1246-2011 | 57587 | tube, orthodontic | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-327 (NF DB UR 6 Convertible .018-14T 14DO HK-Each);ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-328, NF DB UL 6 Convertible .018-14T 14DO HK-Each;ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-418, NF DB UL/LR 6&7 Mini Single .022 0T 0O;ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-427, NF DB UR 6 Convertible .022-14T 14DO HK-Each;ClassOne Orthodontics Label ID Reference Number 351-428, NF DB UL 6 Convertible .022-14T 14DO HK-Each;ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-459, NF DB LR 6 Convertible .022-25T 4DO HK-Each;ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-460, NF DB LL 6 Convertible .022-25T 4DO HK-Each. | ClassOne Orthodontics Buccal SS Tubes | 2 | February 15, 2011 | 2011 | Ortho Organizers Inc | Products were mis-labeled as nickel free. |  | An Urgent: Product Recall Notice letter, dated November 1, 2010 was sent to consignees at the distribution level who received recalled products. Consignees were instructed to do the following:1) Carry out a physical count of current inventory and record this data on the business reply card provided with this notice. Return the enclosed Reply Card with the amount remaining inventory to Ortho Organizers, Inc., Carlsbad, CA 92008 within 10 days of receipt of this notification.2) Immediately cease distribution/use of this product and remove from your shelves.3) Notify all of your retail customers with the Urgent: Product Recall Notice (Retail Level) letter who purchased the affected item numbers from you that the affected items are stainless steel and contain nickel.4) Contact your sales representative to return any unused items. Return any unused product within 30 days. Please clearly mark the return carton Recall Material Enclosed. Ortho Organizers shall accept any product covered by this recall and credit your account for all product and shipping costs. |  | 1300 | 1300 | Worldwide Distribution -- USA, including states of CA, CO, IL, LA, MI, MN, PA & TX and countries of CA (Canada), BE (Belgium), NO (Norway), and CO (Columbia). | Terminated | April 27, 2011 | 71 | 1300 | Not_Computer
Z-1254-2011 | 57592 | bed, flotation therapy, powered | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂŽ MC3 (Model number: M4000) Low Airloss Therapy UnitsFor use to prevent pressure sore development and wound deterioration in patients who have significant risk of developing these problems and who generally also have one or more to the following conditions: immobility, incontinence, poor nutrition, diminished level of consciousness, reduced subcutaneous tissue, or multi-system failure. | Flexicair MC3 | 2 | February 15, 2011 | 2011 | Hill-Rom Manufacturing, Inc. | Due to product age and parts availability the units are being withdrawn to reduce a potential health risk to patients when bed rails are not used. | OTHER/UNDETERMINED: Pending | Hill-Rom Manufacturing, Inc. sent an Urgent Medical Device Withdrawal letter dated December 15, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The FLEXICAR MC3 AND FLEXICAIR 13000 were no longer available for rental as of November 15, 2010. Customers in Michigan were required to switch to another product by December 31, 2010. All other customers were required to switch to another product by March 4, 2011. Customers advised to continue the use of the product until arrangements could be made to replace the product. were notified by letter on/about 12/06/2010. The letter indicated the units will no longer be available for rental as of 11/15/2010. Patients in Michigan currently using the products will be required to switch to another product by 12/31/2010. All other patients using the products will be required to switch to another product by 03/04/2010. They were advised that a Hill-Rom representative will be contacted them to discuss details of the product withdrawal and possible alternative surfaces for their patients.For questions regarding this recall call 800-445-3720. |  | 45 | 42 Flexicair MC3 units | Nationwide Distribution including CA, FL, GA, IL, MI , MS, NJ, NY, OH and PA | Terminated | November 01, 2011 | 259 | 46 | Not_Computer
Z-1256-2011 | 57626 | mask, oxygen | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Oxygen Masks- Labeling as follows:MADA, Inc.Adult Medium Concentration Oxygen Mask with Non-crush 7 Coiled Tubing, Connector Both Ends No 1327; RESPIREXAerosol Mask Adult Elongated Manufactured for PSS World Medical Inc. RESPIREXNebulizer Kits w/Aerosol MaskAdult7 ft. (2.1m) Aerosol TubingManufactured for PSS World Medical Inc.Unomedical, Inc., Hospilak REF 100-E Aerosol Mask Adult Unomedical, Inc., Hospilak Ref. No. 101-E, Adult Aerosol Mask w/Aerosol Connector, 15 mm; Unomedical, Inc., Hospilak Ref. No. 102-E, Medium Concentration Mask, Adult 7 (2.1m) Tubing Unomedical, Inc., Hospilak Ref. No. 103-E, Medium Concentration Mask, Adult Unomedical, Inc., Hospilak REF 104-E High Concentration Rebreather Oxygen Mask, Adult w/7 (2.1)m) Oxygen Tubing, Unomedical, Inc., Hospilak Ref. No. 105-E, High Concentration Rebreather Oxygen Mask, Adult, No Tubing Unomedical, Inc., Hospilak Ref. No. 106-E, Oxygen Mask, Adult High Concentration Non-Rebreather 7 (2.1m) Oxygen Tubing Unomedical, Inc., Hospilak Ref. No. 108-E, Oxygen Mask, Adult High Concentration Non-Rebreather 7 (2.1m) Oxygen Tubing Unomedical, Inc., Hospilak Ref. No. 118-E, Aerosol Mask Elongated, Adult 6 (1.8 m) Aerosol Tubing Unomedical, Inc., Hospilak Ref. No. 120-E, Venturi Mask Kit, Adult All Diluters No-Crush MADA INC. Aerosol Mask, Adult Size, Elongated MADA, Inc. 100% Latex Free, 1 unit per package/50 units per case Disposable Single Patient UseMADA INC. Aerosol Mask, Pediatric Size, Elongated No 1327 100% Latex Free MADA, Inc. 100% Latex Free, 1 unit per package/50 units per case Disposable Single Patient UseMADA INC. Up- Mist Medication Nebulizer With 7 ft. Oxygen Tubing and Elongated Aerosol Mask No. 1310MA 100% Latex Free 1 unit per package/50 units per case Disposable Single Patient UseMADA INC Up-Mist Medication Nebulizer with 7 ft. Oxygen Tubing and Pediatric Aerosol Mask No. 1310 NMP 100% Latex Free 1 unit per package/50 units per case Disposable Single Patient UseUnomedical, Inc., Hospilak REF No. 131-E, Venturi Mask Kit, Adult with 31% diluter only, MADA, Inc. Child/Pediatric Medium Concentration Oxygen Mask with Non-Crush 7 Coiled Tubing, Connector Both Ends, NO 1327C, Unomedical, Inc., Hospilak REF No. 182-E, 3-in-1 Emergency Mask, Elongated Adult, 7 (2.1m) Oxygen Tubing, Unomedical, Inc., Hospilak REF No. 200-E, Aerosol Mask, Adult Small, No Tubing, Unomedical, Inc., Hospilak REF No. 202-E, Medium Concentration Oxygen Mask, Adult with 7 (2.1m) Oxygen Tubing, Unomedical, Inc., Hospilak REF No. 220-E, Aerosol Mask, Pediatric, No Tubing, Unomedical, Inc., Hospilak REF No, 221-E, Aerosol Mask Elongated, Pediatric 5 (1.5m) Aerosol Tubing Unomedical, Inc., Hospilak REF No. 222-E, Medium Concentration Oxygen Mask, Pediatric with 7 (2.1m) Oxygen Tubing, Unomedical, Inc., Hospilak REF No. 223-E, Medium Concentration Oxygen Mask, Pediatric No Tubing, Unomedical, Inc., Hospilak REF No. 228-E, Up-Mist Nebulizer with 7 (2.1m) Oxygen Tubing, Pediatric Aerosol Mask Elongated Unomedical, Inc., Hospilak REF No. 245-E, Adult Trach Mask Unomedical, Inc., Hospilak REF No. 246-E, Pediatric Trach Mask Unomedical, Inc., Hospilak REF No. 260-E, Aerosol Mask, Infant No Tubing, Unomedical, Inc., Hospilak REF No. 263-E, Medium Concentration Mask Infant, No Tubing, Unomedical, Inc., Hospilak REF No. 282-E, 3 in 1 Emergency Mask, Elongated Pediatric, 7 (2.1m) Oxygen Tubing TRI-ANIM Adult Aerosol Mask Elongated adult aerosol mask with elastic strap and adjustable noseclip Part Number 301-170 50/case eValueMed Tri-anim Pediatric Aerosol Mask Elongated pediatric aerosol mask with elastic strap and adjustable noseclip Part Number 301-171 50/case eValueMed Tri-anim Adult Oxygen Mask Elongated adult medium mask with elastic strap and adjustable noseclip Part Number 301-180 eVal | MADA, Hospilak, respirex, Trianim, McKesson, optimist, portex, medichoice | 2 | February 15, 2011 | 2011 | Convatec Inc. | Face tents and Oxygen masks may have a sticky glue-like substance on the inside or outside of the mask consisting of dioctyl phthalate, epoxidized soybean oil and glycerol monoleate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Unomedical, Inc issued recall letters with attached return response forms dated January 7, 2011 via UPS. The letters request the customer stop use and quarantine product. The recall response form is to be completed and returned to Unomedical.The letters/return response form were also sent out in Spanish.Another recall letter was sent dated February 16, 2011 to capture additional lots that werent included in the initial letter. |  | 1113200 | 1113200 pcs. | Worldwide distribution, including USA, Puerto Rico, Canada, Mexico, Colombia, El Salvador, Paraguay, UK, Austria, Belgium, Cyprus, Czech Republic, Denmark, Estonia, Faroe Islands, Finland, France, Germany, Hungary, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Spain, Sweden and Switzerland, Israel, Jordan, Lebanon, Libya, Morocco, Saudi Arabia, UnitedArab Emirates, Russia, Australia, New Zealand,Japan and Vietnam. | Terminated | February 27, 2013 | 743 | 2692300 | Not_Computer
Z-1068-2011 | 57580 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ENDOPATH ETS-FLEX45 Articulating Endoscopic Linear Cutters, Product Codes: ATS45; LONG45A; 6TB45; ATG45; ATB45; ATW45; ATS45NK; and NAW45. Intended for transection, resection, and/or creation of anastomoses. | ENDOPATH ETSFLEX45 Articulating Endoscopic Linear Cutters | 2 | February 15, 2011 | 2011 | Ethicon Endo-Surgery Inc | The firm initiated this second subsequent recall because the affected devices can remain closed and clamped on tissue during surgical procedures due to the fact that the mechanism which connects the articulation joint to the device shaft is insufficient. | OTHER/UNDETERMINED: Pending | Ethicon Endo-Surgery issued "Urgent Expanded Device Recall Event 1981" notification letters dated April 7, 2009, to their customers. They provided a list of all Affected Product Codes of the ETS FLEX45 Articulating Endoscopic Linear Cutters (including the Product Code: ATS45 from the 1st recall) and Affected Flex Tray Codes subject to recall. The notification document reminds the Operating Room and Materials Management customers of the problem which necessitated the previous recall (Z-0361-2011). It goes on to instruct their customers to immediately discontinue use of the affected products. They are to complete and fax back the Business Reply Form, Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle, Replacement product should be received within 30-60 days after return, Customers can contact Stericycle concerning this recall at 1-888-365-4680, reference Event 1981. |  | 29403 | 29,403 units | Nationwide Distribution -- USA including the following states, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MO, MS, MT, NC, NE, NH, NJ, MN, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, and Puerto Rico. | Terminated | February 15, 2013 | 731 | 29538 | Not_Computer
Z-1257-2011 | 57499 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-301A AQUILION ONE (MPDCT0297EA) - Whole Body X-Ray CT Scanner | Toshiba AQUILION ONE | 2 | February 15, 2011 | 2011 | Toshiba American Medical Systems Inc | Incorrect Dose Calculation. The software that performs the pre-scan dose calculation uses an incorrect parameter for an exam Plan using a tube voltage of 80kV or 100kV. This results in a pre-scan expected dose that is approximately 10% lower than what the displayed value should be when the scanner executes the exam Plan. | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) issued notification to consignees 9/7/2010, via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. | Contact the recalling firm for information | 44 | 44 units distributed between 09/2007 - 07/2010 | Nationwide distribution. | Terminated | April 20, 2012 | 430 | 44 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1247-2011 | 57665 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Mantis Tap Sleeve,Mfg by: Stryker Spine, 2 Pearl Court, Allendale, NJ 07401-1677. packaged in the following products: Specialty Mantis Modular Tap Set - Catalog #IS2162MAN; Custom Tap Sleeve - Catalog #IS830MAN; Custom Tap Sleeve, Mantis - Catalog #IS2271MAN; and Custom Mantis Modular Tap Set - Catalog #IS2080MANThe Mantis Tap Sleeves with Markers was designed to be used in a manner identical to the predicate Standard Mantis Tap Sleeve 48281315. It is to be used in conjunction with Standard Mantis Cannulated Modular Taps: 1) to protect soft tissue from the cutting edge of the tap 2) to indicate the depth at which the Tap has been inserted. | Mantis Tap Sleeve | 2 | February 15, 2011 | 2011 | Stryker Spine | Stryker Spine has received reports that after sterilization but prior to surgery, cracks were noticed at the tantalum marker / PEI tube interface. This can likely be attributed to differences in the rate of thermal expansion between the markers and the sleeve during the sterilization cycle. | OTHER/UNDETERMINED: Pending | Stryker Spine sent an Urgent Product Recall letter dated January 07, 2011, to Stryker Spine sales branches and agencies where affected product was distributed. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile any Mantis Tap Sleeves with the catalog and lot numbers referenced in the letter. Complete the Customer Response Form with distribution sheet provided to assist with product reconcilation. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Customers were to retrieve and return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance, Stryker Spine 6 Pearl Court, Allendale, New Jersey 07401. For questions call 201-760-8180. |  | 25 | 25 units | Nationwide Distribution including MA, MN, and PA. | Open |  |  | 25 | Not_Computer
Z-1251-2011 | 57604 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Elekta Leksell Stereotactic System, used for localization (spatial reference) for cranial surgery using X-ray or CT and MRI Image data. | Elekta Leksell Stereotactic System | 2 | February 15, 2011 | 2011 | Elekta, Inc. | New protocols for MR sequences may result in higher RF energies deposited during MR scanning, generating heat in uninsulated fixation posts. | DESIGN: Labeling Design | Elekta sent Important Notice Affected Product: Fixation Posts, A331, dated June 22, 2009. The Important Notice will be added to the Customer User manual upon receipt. |  | 379 | 379 units | Nationwide | Terminated | February 17, 2011 | 2 | 379 | Not_Computer
Z-1252-2011 | 57731 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVIIntended use: Radiation therapy treatment of malignant neoplastic diseases, as determined by a licensed medical practitioner. | Medical linear accelerator (with patient imaging) | 2 | February 15, 2011 | 2011 | Elekta, Inc. | Flexmap calibration did not cover the full 360 degrees of gantry rotation (0 to 355 degrees) and that this caused an isocenter displacement with 2D PlanarView images required at 180 to -180. | OTHER/UNDETERMINED: Pending | The firm, Elekta Limited, sent an "Important Notice A344" dated January 5, 2011 to all consignees/customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to follow any instructions or recommendations covered in the notice. The notice provides suggestions on how to check and prevent this situation from occurring; it instructs the User to file the document in the Important Notice section of the appropriate User Manual; to complete a FlexMap Check each time they acquire a new FlexMap and amend their local procedures to include the FlexMap Check and complete and return the Important Notice Confirmation of Receipt via Fax to: Field Support Administration, Elekta Limited, Linac House, Fleming Way Crawley, West Sussex, RH10 9RR UK +44 (0)1293 654401.If you have any questions, contact Elekta Limited at +44 (0)1293 654200 or www.elekta.com. |  | 223 | 223 units | Nationwide distribution. | Terminated | April 07, 2011 | 51 | 223 | Other |  |  | Other
Z-1250-2011 | 57159 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Xeleris Processing and Review Workstation 1.0, 1.1, 2.0, 3.0The display, processing, archiving, and communication of data acquired by Emission Tomography cameras used in diagnostic radiology, including procedures for planar imaging, whole body imaging, tomographic (SPECT) imaging, positron imaging by coincidence, attenuation correction, and anatomical image registration. 3X The system is intended for use by Nuclear Medicine (NM) or Radiology practitioners and referring physicians for display, processing , archiving, printing, reporting and networking of NM data including planar scans (Static, Whole Body, Dynamic, Multi-Gated) and tomnographic scans (SPECT, Gated SPECT, dedicated PET or Camera-Based-PET) acquired by gammna cameras or PET scanners. The system can run on dedicated workstation or in a server-client configuration. The NM or PET data can be coupled with registered and/or fused CT or MR scans and with physiological signals, in order to: depict, localize, and/or quantify the distribution of radionuclide tracers and anatomical structures in scanned body tissue for clinical diagnostic purposes.. | Xeleris Processing and Review Workstation 1.0, 1.1, 2.0, 3.0 | 2 | February 15, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of safety issues relevant for the Xeleris Workstation 1.0, 1.1, 2.0, 2.1, 3.0:1. Measurements taken on the CT Images on the CT system are different from the Xeleris Workstation2. Inaccurate PET SUV values may be calculated when data is retrieved from PACS after being processed in the Xeleris.3. Wrong marking of Early and Late result series on display in Bra | DESIGN: Software Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 12, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to please ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare call center in US or Canada at 800-437-1171 and other countries at 262-896-2890. | Contact the recalling firm for information | 6046 | 6046 | Worldwide distribution: USA including DC, PR, and Guam; and countries including: AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA, AZERBAIJAN, BELGIUM, BOSNIA & HERZEGOVIA, BULGARIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, LUXEMBOURG, LEBANON, MALY ASIA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, PORTUGAL POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SPAIN, SOUTH AFRICA, SOUTH KOREA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, and VIET NAM. | Terminated | October 16, 2012 | 609 | 6046 | Software | Output/Calculation | Software update | Software Update
Z-1245-2011 | 57379 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare IT, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. The Centricity Laboratory software is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. It integrates with other Hospital Information Systems (HIS) through HL7 interface by order entering and report processing. The centricity is intended to interface with various lab devices through standard HL7 interfaces to record, report and distribute lab results. | GE Centricity Laboratory | 2 | February 15, 2011 | 2011 | Ge Healthcare It | The GE Centricity Laboratory System Software has safety issues with the "Telephone List" function, the "Anatomic Pathology" application, and the Instrument Interface IO module that may impact patient safety. | DESIGN: Software Design | GE Healthcare, sent "Urgent Medical Device Correction" letters dated December 12, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the product, the problem, and the action to be taken by the customer. The letters provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. | Contact the recalling firm for information | 26 | 26 units | Worldwide Distribution -- USA, including AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, NC, NH, NJ, NY, OR, TN, VA, and WI, , and internationally to the countries of Australia, England, Hong Kong, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 150 | 26 | Software | Physical Safety Hazards | Software update | Software Update
Z-1334-2011 | 50094 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | S T-Bone Plate;M T-Compression Plate;Manufactured and Distributed by Stryker Lebinger GmbH & Co.Freiburg, Germany.The Profyle System is intended for use in internal fixation of the bones of hand and wrist. Examples of these procedures may include but are not limited to replantation, lag screws techniques, joint fusion, corrective osteotomies, and the treatment of fractures | S TBone Plate; M TCompression Plate; | 2 | February 16, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | A package labeled as a S T-Bone Plate, Catalog Number 57-05260, Lot Code G4M00F1576, may actually contain a M T-Compression Plate, Catalog Number 57-10250, Lot Code G4N00F1584 and vice versa. | DESIGN: Process Design | Stryker Orthopedics sent an Important Market Withdrawal letter dated May 26, 2006, by Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory, including their hospital consignment location, to identify the product.If the affected product remained in its original packaging, customers were instructed to retrieve all of the affected product lots and return it to their branch or agency warehouse for reconcillation.Reconcile all product on the attached Product Accountability Form and fax a copy of the form to 201-831-6069 within 5 days of receipt of the notice.Customers were asked to return the affected product via Inter-Org Transfer and ship to Product Surveillance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, USA 07430-2104. Customers were also instructed to contact their Stryker Orthopaedics Customer Service representataive to re-order the product that they were returning.For questions regarding this recall call (201) 831-5825. |  | 21 | 21 | Nationwide Distribution | Open |  |  | 21 | Not_Computer
Z-1320-2011 | 57295 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | WalkMed Infusion Triton 90" 20 drop IV Administration Set with Needlefree Y-siteItem # T090-000VIntended use: Intravascular Administration Set | WalkMed Infusion Triton 90 | 2 | February 16, 2011 | 2011 | Churchill Medical Systems, Inc. | Routine testing conducted to challenge the success of sterilization was not fully completed for sterile products. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | The firm, Churchill Medical/Vygon, sent an "URGENT MEDICAL DEVICE RECALL" letter dated November 16, 2010, via overnight mail to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their stock for affected product; immediately cease use and distribution and quaratine the affected product; contact Vygon's Customer Service Department at 1-800-473-5414 for a return Merchandise Authorization number; complete and return the Recall Acknowledgement and Inventory Return Form via fax to 215-390-2019 and mail return product with form to: Vygon, 103A Park Drive, Montgomeryville, PA 18936, and if the customer is a distributor, they were to provide their customers with a copy of this notification and attached form. Note: An additional follow up communication will be provided to the customers after reconciliation of the returned inventory and all inventory will be credited or replaced.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email at customerservice@vygonus.com. |  | 12150 | 12150 units | Nationwide Distribution: USA including states of: CO, IA, OH, MA, MD, MO, NY, FL, SC, FL, and VA. | Terminated | May 01, 2012 | 440 | 61251 | Not_Computer
Z-1341-2011 | 57454 | activated whole blood clotting time | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Celite ACT CartridgeCartridge is useful for monitoring patients receiving heparin for treatment of pulmonary embolism or venous thrombosis, and for monitoring anticoagulation therapy in patients undergoing medical procedures such as catheterization, cardiac surgery, surgery organ transplantation and dialysis. | Celite ACT Cartridge | 2 | February 16, 2011 | 2011 | Abbott Point Of Care Inc. | Celite ACT cartridges have increased numbers of values outside the accepted allowable error. | DESIGN: Device Design | On behalf of Abbott, Genco Pharmaceutical Services sent an "URGENT RECALL NOTICE" dated November 2010, on December 3, 2010 via Fedex to all consignees/customers. Abbott sent out notification packages to foreign consignees/customers on December 3, 2010. The notice described the product, problem and action to be taken by customers. The firm recommended that customers use cartridges within 12 hours of removal from refrigerated storage based on the instructions provided with the product. The customers were instructed to complete and return the Business Reply cards to acknowledge receipt of the communication and provide the letter to any customers that have recieved the product. If you have any additional questions, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care Representative.On February 27, 2012 Abbott sent an Important Product Information letter via Fed Ex to inform customers that i-Stat Celite ACT cartridges manufactured from the beginning of 2012 can now be stored for two weeks at room temperature as indicated on the cartridge box. Foreign consignees were notified on 2/28/2012. Users are instructed to complete business reply cards to acknowledge receipt of the communication and provide the letter to any customers that have received the product. Any additional questions can be answered by contacting 8003668020 option 1. |  | 1164275 | 1164275 cartridges | Worldwide distribution. | Terminated | July 16, 2013 | 881 | 1164275 | Not_Computer
Z-1308-2011 | 57318 | appliance, fixation, nail/blade/plate combination, single component | KTW | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 02.122.610s, 110mm Spiral Blade, 14mm, Right-Sterile - Synthes Spiral Blade Component of the Synthes Modular Blade Plate SystemFor fixation of fractures and osteotomies of the proximal and distal femur in adolescents and adults. | Synthes Spiral Blade Component of the Synthes Modular Blade Plate System | 2 | February 16, 2011 | 2011 | Synthes USA (HQ), Inc. | Potential for out of specification star grind teeth. | OTHER/UNDETERMINED: Pending | Synthes initiated their recall on October 18, 2010, notifying affected hospital accounts via consultant communication. Affected Synthes Trauma Sales Consultants were notified via email (receipt requested). A copy of the Urgent Medical Device Recall letter dated October 20, 2010, was mailed via UPS second day air on October 20, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected devices and call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. A replacement order would be taken at that time.Customers were also instructed to complete the Verification Section at the end of the letter by checking the appropriate box indicating the affected product had been located or indicating in the appropriate box that no affected product was located. Also, customers were to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the recalled product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.For questions regarding this recall call 1-800-620-7025, x 6883. |  |  |  | Nationwide Distribution including AL, CA, GA, IA, IL, IN, MA, MI, MS, NC, NJ, NV, OH, PA, SC, SD, TN, and WV. | Open |  |  |  | Not_Computer
Z-1335-2011 | 57792 | marker, ocular | HMR | Ophthalmic surgical marker. | Ophthalmic | Ophthalmic | 510(K) Exempt | Oasis Corneal Marking Pad, REF 0441, Rx only, Sterile. For use in ophthalmic surgery. | Oasis Corneal Ink Stamp Sterile | 2 | February 16, 2011 | 2011 | Aspen Surgical Products, Inc. | Fragments from the pad may disengage and stick to the instrument. The condition could cause fragments to come into contact with the intact cornea prior to surgery. | DESIGN: Component Design/Selection | Aspen Surgical sent URGENT MEDICAL DEVICE RECALL NOTICES dated 12/23/2010 by Certified Mail. The letters instructed the users to examine their inventory and place the recalled products in quarantine. Consignees were instructed to return the product and complete the response form. Distributors were asked to notify their customers or requesting that Aspen notify the customers.The firm does not intend to continue manufacturing or distributing the product. |  | 77850 | 77,850 | Worldwide distribution including USA Canada, England, Sweden, Saudi Arabia, England, Hong Kong, Belgium and Israel. | Terminated | October 07, 2011 | 233 | 77850 | Not_Computer
Z-1365-2011 | 57544 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ConMed System 2450 Electrosurgical Generator, REF 60-2450-230, 230V-240V, ConMed Electrosurgery, Centennial, CO.A general-purpose electrosurgical generator used in conjunction with an electrosurgical accessory handpiece for delivery of high (HF) electrosurgical current through an accessory electrode for cutting and coagulation at the operative site. | Electrosurgical generator. | 2 | February 17, 2011 | 2011 | ConMed Electrosurgery | Electrosurgical generator may allow possibility of electrical shock and/or burns to user or patient. | DESIGN: Device Design | ConMed Electrosurgery sent an Urgent Device Recall letter dted December 27, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify anyone who may have received the affected product.. They were to review their inventory and return a Business Reply Form via fax to 303-699-9854 and to contact ConMed Electrosurgery Technical Services at 800-552-0138 to schedule a time to have their equipment upgraded. |  | 442 | 442 units | Worldwide Distribution, including USA (nationwide) and the countries of Australia, Belgium, Bolivia, Canada, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Honduras, Hong Kong, Hungary, India, Iraq, Israel, Italy, Jamaica, Japan, Jordan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, Vietnam. | Terminated | August 01, 2011 | 165 | 2323 | Other
Z-1348-2011 | 57287 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | CENTRAL VENOUS CATHETERIZATION KIT/Product SKU's (Codes) AK-04650-E and AK-04660.Central venous catheterization kit | CENTRAL VENOUS CATHETERIZATION KIT | 2 | February 17, 2011 | 2011 | Arrow International Inc | Sterility may be compromised. | OTHER/UNDETERMINED: Pending | Arrow International, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated October 29, 2010, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customer. Customers were instructed to quarantine the product.If they had no affected product in stock, check the appropriate box on the Recall Acknowledgement & Stock Status Form and fax the form to Arrow's Customer Service fax number 800-343-2935.If they had affected stock, immediately quarantine the product and contact Arrow Customer Service at 800-523-8446 for a Return Authorization Number. If customers wished to return product for replacement, they were to consult their sales representative or Arrow Customer Service for information on product availability, substitutability, and lead times. Complete the enclosed Recall Acknowledgement & Stock Status Form and immediately fax the compoleted form to Arrow, Fax number 800-343-2935, Attn: Customer Service.Return any affected product freight collect, along with the original completed Recall Acknowledgement & Stock Status Form to:Arrow International, Inc.312 Commerce PlaceAheboro, NC 27203Attn: Asheboro Quality Assurance Manager.For any questions regarding this recall call 800-523-8446. |  | 40260 | 40,260 | Worldwide Distribution -- United States - all 50 States, and the countries of Canada, South America, China, South Africa, Australia, Puerto Rico, Philippines, Singapore, South Korea, New Zealand, Mexico, Japan, India, Indonesia, Bahamas, Taiwan, Thailand, Finland, France, Cyprus, Czech Republic, Belgium, Germany, Ireland, Italy, Latvia, Netherlands, Poland, Slovakia, Spain, Bahamas, Costa Rico, Dominican Republic, Ecuador, Guatamala and the United Kingdom. | Terminated | September 29, 2012 | 590 | 963380 | Not_Computer
Z-1342-2011 | 56387 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, Model Number/Catalog No.: 270-2417, which contains the update kit CD-ROM, Model/Catalog No.: 270-2423, Distributed and Manufactured by: Bio-Rad Laboratories, Inc., Hercules, CA.Intended for the percent determination of hemoglobin A1c in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). | VARIANT II TURBO Hemoglobin A1c Program | 3 | February 17, 2011 | 2011 | Bio-Rad Laboratories Inc | The firm received customer complaints associated with the Update Kit CD-ROM not being able to generate a valid calibration after being uploaded, resulting in patient results not being generated. |  | Medical Device Correction letter dated March 24, 2010 was sent to customers. The letter identified the affected product and explained the reason for recall. Customers were asked to destroy Lot No. AA00190 and replace it with the Lot No. AA00838 CD-ROM that is enclosed with the letter. Customers are to complete and fax back the Medical Device Correction Response Form at the number provided. |  | 48 | 48 units | Worldwide Distribution -- USA and Italy. | Terminated | March 04, 2011 | 15 | 48 | Not_Computer
Z-1370-2011 | 57850 | instruments, surgical, cardiovascular | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | MitraXs" Self Adjusting Atrial Retractor, Sterile EO, Models (REF): A0030/1; A0031/1; A0032/1; A0033/1, Cardio Life Research s.a., Green Hill, Hoge Wei, 16, B-1930 Zaventem, Belgium.Indicted for the retraction of the atrium walls during mitral or tricuspid cardiac surgery. | MitraXs" Self Adjusting Atrial Retractor | 2 | February 17, 2011 | 2011 | St. Jude Medical Cardiovascular Division | St. Jude Medical Cardiovascular Division is conducting a voluntary recall of all batches and model numbers of the MitraXs Self Adjusting Atrial Retractor. St. Jude Medical has been distributing this product since their acquisition of CLR in late 2008.They have determined that some distributed batches have very small burrs on the polycarbonate material at the distal edge. As these burrs may | PRODUCTION CONTROLS: Process Control | The firm, St. Jude Medical, sent an "Urgent Medical Device Recall Notice" dated January 26, 201,1 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue use of the product (MitraXs Self Adjusting Atrial Retractor batches) and remove from inventory. St. Jude began production of a new batch of units that will bear the labeling of St. Jude Medical. Shipment of these units started on January 10, 2011. Note: A St. Jude Medical Sales Representative will contact the customer to facilitate removal and reconciliation of the affected product.If you have any questions or concerns, please contact your St. Jude Medical Sales Representative or 651-490-4470 and/or 612-310-7936. | Contact the recalling firm for information | 450 | 450 | Worldwide distribution: USA including states of: NJ, OH, MI, and MN; and countries including: BELGIUM, FRANCE, GERMANY, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 31, 2012 | 469 | 450 | Not_Computer
Z-1374-2011 | 56394 | excimer laser system | LZS |  |  | Ophthalmic |  | AMO brand WaveScan WaveFront System, OPM software versions 3.67 and 3.67,Part Number: 0070-1531, Catalog Numbers: 0070-1664,0070-1649,0070-1654: Version 3.670070-1721: Version 3.671 0070-1662: Version 3.67, US version;0070-1668; IVersion 3.67 International versionProduct is manufactured and distributed by AMO Manufacturing USA, LLC, Milpitas, CAThe WaveScan WaveFront System is a class III accessory device used in conjunction with the VISXÂŽ STAR S4 and STAR S4 IRTM Excimer Laser Systems, which are indicated for wavefront-guided laser assisted in situ keratomileusis (LASIK) eye surgery in certain patient populations. The WaveScan WaveFront System is a diagnostic instrument intended for the automated measurement, analysis, and recording of refractive errors of the eye. The WaveScan System software uses the diagnostic data of the eye's refractive errors and wavefront aberrations to calculate the target treatment shape and then generates the commands for the VISX STAR laser to create the target shape on the cornea. | AMO brand WaveScan WaveFront System | 2 | February 18, 2011 | 2011 | AMO Manufacturing USA, LLC | The failure to comply with Minimum System Requirements, increased system errors could result, and a miscreated treatment table (incorrect LASIK refractive treatment) could result. | DESIGN: Software Design | Abbott Medical Optics, Inc. sent an Urgent Medical Device Field Correction letter to all affected customers starting July 15, 2010. The letter identified the product, the problem, and the action to be taken by the customer.Customers were asked to take the following actions:Surgeons were instructed to complete the facsimile attachment to acknowledge receipt of the notification letter and fax to 714-247-4530.For those using the affected product a 3.68 OPM disc must be requested by e-mailing a request to OPM@pacful.com.For any questions please contact AMO Technical Service Representatiave at 1-877-266-4543. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britian, Greece, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Netherland, New Zealand, Peru, Philippines, Portugal, Russia, Russia Federation, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Taiwan, ThailandChina, India, South Korea, Singapore, Taiwan, Thailand, Turkey, United Arab Emirates, Venezuela, and Vietnam. | Terminated | June 22, 2012 | 490 |  | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1373-2011 | 57502 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part number 110-0234G1 (domestic) and 110-0234G3 (international)The Orthoralix 8500 DDE Panoramic Dental X-Ray System is used as an extraoral source of x-rays for imaging of the dento-maxillofacial area. | Orthoralix 8500 DDE Panoramic Dental XRay System | 2 | February 18, 2011 | 2011 | Gendex Dental Systems | The Automatic Exposure Control (AEC) feature located on the control panel of the Orthoralix 8500 Panoramic Dental X-Ray System was functional, but was subsequently disabled in units manufactured starting 5/17/10. North American units manufactured between 12/4/09 and 5/16/10 had functioning AEC buttons, but their manuals stated they were disabled. Units intended for International distribution bet | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Gendex sent a recall letter to their accounts on December 21, 2010, stating that the AEC button on the control panel was a non-functional feature that becomes illuminated upon activation and that Gendex would send them a kit to correct this issue and inactivate the button. A second letter was sent to the accounts on January 31, 2011, explaining that a corrective kit was no longer necessary, letting them know if the AEC feature was active or inactive for their system, and providing them with the appropriate user manual addendum for their system.Customers were asked to complete and return the required Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com.For questions regarding this recall call 800-323-8029. |  | 574 | 574 systems | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WA, and WI and the countries of Australia, Canada, Germany, New Zealand, Taiwan, and Vietnam. | Terminated | November 09, 2011 | 264 | 574 | Not_Computer
Z-1260-2011 | 57567 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device Product Code STRAP25.Intended for fixation of prosthetic material to soft tissues in various minimally invasive and open surgical procedures, such as hernia repairs. | Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device | 2 | February 18, 2011 | 2011 | Ethicon, Inc. | The Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device may have the potential for the sterile barrier to be compromised due to a defect in the packaging. |  | Ethicon sent Urgent Voluntary Product Recall letters dated, December 28, 2010, to their customers on December 29, 2010 via UPS. The letter identified the affected product and described the reason for recall. Customers are to discontinue use and sampling immediately. They are to return the product using the enclosed prepaid shipping label along with the enclosed business reply card regardless of whether they have any affected product or not. Customers should contact 1-877-551-7147 or speak to their manager if they have questions regarding this recall. |  | 4110 | 4,110 | Worldwide Distribution -- USA and Canada. | Terminated | November 07, 2012 | 628 | 4110 | Not_Computer
Z-1372-2011 | 57736 | electrode, electrodcardiograph, multi-function | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Zoll Medical (Part # 16381, HeartSync, Ann Arbor, MI.The Adult Radiolucent Multifunction Tin Electrodes are indicated for use in external pacing, defibrillation and monitoring applications as a non-sterile, disposable device for single patient use only. The electrodes provide the conductive interface between the defibrillator and/or the external transcutaneous cardiac pacemaker and the patients skin. When a patient requires defibrillation, cardioversion or external pacing, these electrodes will be applied to the patient and connected to the respective low power mono-phasic or bi-phasic defibrillator (360 joules maximum) or external pacemaker. | Medstorm Adult Radiolucent Electrode | 2 | February 18, 2011 | 2011 | Heart Sync, Inc | The seal on the electrode pouch may be open, exposing the electrode to the outside environment. The package does have a caution statement that states, " Do not open package until immediately prior to using electrodes. and Electrodes may dry out when removed from packaging and exposed to air. | PRODUCTION CONTROLS: Packaging | Heart Sync, Inc sent an URGENT MEDICAL DEVICE RECALL letter dated December 23, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Examine their inventory and quarantine product that was subject to the recall that was not sealed properly. Contact any customers if they further distributed the product and notify them at once of this product recall. Contact the firm if they had any unused product. The firm would issue a UPS Call Tag to have the product returned to them and the firm would send replacement product as soon as they were notified.Complete and return the Recall Return Response Form to the firm in the attached prepaid envelope.For any questions regarding this recall call 734-213-5530. | No consumer action necessary | 300 | 300 | Worldwide Distribution - USA including AK, AZ, CA, FL, IL, KS, LA, MA, ND, NM, OH, TX, WA, WI and WV; and Canada | Terminated | July 13, 2011 | 145 | 300 | Not_Computer
Z-1378-2011 | 57666 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | InfusO.R. Infusion Pump, an Rx syringe infusion pump; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A.; product codes 2L3100, 2L3100R, 2L3100U, 6461500 and 6461500R. These product codes includes refurbished units and InfusO.R. pumps previously sold under the Bard label. Intended for the controlled rate delivery of small volume parental fluids as prescribed by a physician. | InfusO.R. | 3 | February 18, 2011 | 2011 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the InfusOR will not operate. The reason being that the InfusOR (2L3100) Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | OTHER/UNDETERMINED: Pending | Safety Alert letters dated January 11, 2011 were mailed via first class mail to the InfusOR direct accounts, to the attention of the Director of Biomedical Engineering, the Director of Anesthesia and the Director of Materials Management, on the same date, informing them that over-the-counter Energizer C size alkaline batteries should not be used with the InfusO.R. Pumps (Product codes 2L3100, 2L3100R, 2L3100R, 2L3100U, 2L3100N, 6461500 and 6461500R). That particular battery incorporates a feature that may inhibit electrical contact and affect the ability of the pump to power-up. The accounts were requested to check their inventory of InfusO.R. Pumps for the presence of over-the-counter Energizer C size alkaline batteries. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST.*** Baxter expanded the scope of their action to include the Mini-Infuser System and sent Urgent Device Correction letters dated 3/4/11 via First Class mail to all InfusOR and Mini-Infuser customers who had purchased either of the affected product lines. The letters informed the accounts that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The accounts were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future refer. |  | 39725 | 39,725 pumps | Nationwide Distribution | Terminated | August 27, 2012 | 556 | 39725 | Battery | Device Operation | Instructions | Safety Notice/Insructions
Z-1376-2011 | 57270 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Discovery MR450 1.5T, The Discovery MR450 is a whole body magnetic resonance scanner designed to support high resolution, high signal-to-noise ratio, and short scan times. It is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images, spectroscopic images, parametric maps, and/or spectra, dynamic images of the structures and/or functions of the entire body, including, but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. Depending on the region of interest being imaged, contrast agents may be used; The images produced by the Discovery MR450 reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and/or spectra when interpreted by a trained physician yield information that may assist in diagnosis. | GE Healthcare Discovery MR450 1.5T | 2 | February 18, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that during certain service procedures, there is a potential for electric shock associated with the gradient cooling manifold hoses of your MR750, MR450, and MR450w MR systems that may impact service or maintenance personnels safety.The gradient cooling manifolds used on the affected systems may contain electrically conductive hoses, which become live when the grad | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 16, 2011, to all affected customers. The letter was addressed to Hospital Administrator / Risk Manger, Radiology Department Manager and Maintenance / Biomedical Engineering Department. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product correction and Contact Information.Customers were instructed to follow appropriate lockout/tagout (LOTO) procedures.Customers were instructed to contact GE Healthcare engineering at 843-687-7213 or 843-697-3009 for instructions on how to proceed. For questions concerning the notiifcaton customers were instructed to contact their local Support Engineer. | No consumer action necessary | 11 | 11 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, LA, MD, MA, MI, MO, NV, NJ, NY, NC, OR, PA, SC, TN, TX, WA, WI. and the countries of AUSTRALIA, BRAZIL, CANADA, EHYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, ITALY, JAPAN, KUWAIT, QATAR, REPUBLIC OF KOREA, SWITZERLAND, TAIWAN, and the UNITED KINGDOM.UNITED KINGDOM, TAIWAN, SWITZERLAND, REPUBLIC OF KOREA, QATAR, KUWAIT, JAPAN, ITALY, INDIA, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, CANADA, BRAZIL, AUSTRALIA. | Terminated | February 04, 2012 | 351 | 104 | Not_Computer
Z-1392-2011 | 53691 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5M TEE Transducer, when used with an ACUSON S2000 ultrasound system that has cardiac options installed and is at software version VA10, VA10A, VA10B, VA10C, VA15, VA15A, VA15B, VA16 or VA16A, manufactured by Siemens Medical Solutions, Mountain View, CA.Multipurpose mobile software controlled diagnostic ultrasound system and on-screen display for thermal and mechanical indices related to bio-effect mechanisms. Its function is to acquire primary or secondary ultrasound echo data. The V5Ms is a multi frequency, multiplane, phased sector array transducer. It can be introduced into the esophagus to obtain images of the heart structure. | V5M TEE Transducer | 2 | February 22, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | If the transducer is not disconnected from the system after an automatic shutdown due to thermal limits being exceeded, then the next time it is used, the thermal safety guards do not take effect. Normally, the thermal limit before automatic image freeze (and therefore power shut off) occurs at 43 degrees Centigrade. | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notification letter on August 28, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their local Customer Service Engineer would install a software update as soon as it becomes available. Customers were instructed to pass the Customer Safety Advisory Notification letter to all who need to be aware of these issues until the corrective action is completed. For questions regarding this recall call 650-969-9112 or 800-422-8766. | Contact the recalling firm for information | 97 | 97 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, China, Czech Republic, Denmark, Germany, Great Britian, Hong Kong, Indonesia, Italy, Japan, Norway, Poland, Portugal, Russian Federation, Singapore, South Korea, Taiwan, and Turkmenistan. | Terminated | March 29, 2011 | 35 | 97 | Software | Physical Safety Hazards | Software update | Software Update
Z-1390-2011 | 56670 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The GMD UNIVERSAL URINARY INCONTINENCE SLING (Code 1010) is supplied sterile and is intended for single patient use only.The GMD Universal Sling is a sterile, single use device for the treatment of female stress urinary incontinence. It is comprised of a polypropylene sheath with a disposable universal low density polyethylene sleeve at each end of attachment of the sling to GMD's single use or reusable trocars. The universal sleeve has three trocar insertion points, the distal and proximal trocar insertion points are for inside-out, bottom-up approaches and a sleeve end trocar insertion point is for outside-in, top-down approaches. The design allows the physician to utilize the device for placement of a polypropylene mesh at the urethra or bladder neck. | GMD UNIVERSAL URINARY INCONTINENCE SLING | 2 | February 22, 2011 | 2011 | Generic Medical Devices | Protective sheath sticks to itself and sleeve breakage has resulted in complaints and MDRs. | DESIGN: Component Design/Selection | The firm, GMD, issued an "URGENT MEDICAL DEVICE RECALL" letter dated February 14, 2011 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed not to use any GMD Model 1010 Universal Slings remaining in their inventory; return any Model 1010 slings in their possession to the Distributors or GMD; notify their customers to whom they shipped Model 1010s of this voluntary recall and request them to return any Model 1010 slings in their possession to them; and return all Model 1010 inventories to GMD using the attached Recalled Model 1010 Return Form. Note: Full credit in the amount of the returned products purchase price will be issued for all sterile, unexpired product returned to GMD.Contact the GMD Product Performance Senior Manager, at phone: (253) 853 3512 or email: 1010recall@gmd-us.com, if you have any questions/issues related to this recall. |  | 1711 | approximately 1711 | Worldwide distribution: USA and countries including: Australia, Argentina, Brazil, Canada, South Africa, and Scandinavia. | Terminated | September 24, 2012 | 580 | 1711 | Not_Computer
Z-1393-2011 | 56367 | sterilization wrap containers, trays, cassettes & other accessories | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332.Intended to be used to contain medical devices that are being terminally sterilized in STERRAD Sterilization Systems. | Tyvek SelfSeal Pouch with STERRAD Chemical Indicator | 2 | February 22, 2011 | 2011 | Advanced Sterilization Products | The recall was initiated after Advanced Sterilization Products (ASP) discovered a non-conforming lot. The misaligned perforations may have caused the seal to be incomplete on some of the pouches, creating a potential breach for pouches being used to store sterilized equipment on shelves. | PRODUCTION CONTROLS: Process Control | Advanced Sterilization Products (ASP) released Urgent: Product Recall notifications to its consignees identifying the affected product and explaining the reason for recall. The letter also discussed the impact of using this affected product and return instructions. The firm asked customers to examine their inventory, discontinue use of Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332 and to return all unused affected products to Stericycle. Customers are to complete the Business Reply Card and return it to Stericycle. | No consumer action necessary | 470 | 470 | Worldwide Distribution -- USA, Argentina, Australia, Canada, Chile, Czech Republic, Hong Kong, India, Japan, Panama, Peru and Puerto Rico. | Terminated | March 14, 2012 | 386 | 470 | Not_Computer
Z-1395-2011 | 57715 | antigens, all groups, streptococcus spp. | GTY | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | SP Brand Strep A Cassette; Part Number: B1077-26; Lot #: STA0010037The SP Brand Strep A Rapid Test Device is a rapid chromatographic immunoassay for the qualitative detection of Strep A antigen from throat swab specimens to aid in the diagnosis of Group A Streptococcal infection. This test is indicated for professional and point of care use only. | SP Brand Strep A Cassette | 2 | February 22, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated after the Investigation of customer-reported complaints, testing of market retention and returned product, It was determined that this lot of product hasdemonstrated an elevated Incidence of non-specific binding, potentially causing false positive results. A falsepositive test result could lead to the administration of antibiotics that are unnecessary. This exposes the |  | The firm, Alere San Diego, sent a "URGENT MEDICAL DEVICE RECALL" letter dated January 17, 2011. The letter described the product, problem and action to be taken by customers. The customers were instructed to: (1) Immediately discontinue use of the affected lot (if applicable) and inventory remaining full or partial kits at their site (if applicable);(2) Destroy any remaining inventory of this lot and document destruction per their site requirements (if applicable);(3) Complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of this notice and to receive replacement of their remaining inventory of this lot (if applicable). The customers were ask to please return the form even if they have no remaining inventory or do not have any involved product in order to confirm receipt of this notification; and (4) to please ensure that all users of the device have received a copy of this Urgent Medical Device Recall.Customers with questions about the information contained in the notice were instructed to contact Technical Services at 888-246.7483, option 2 or by e-mail attechservices@alere.com. |  | 814 | 814 kits/24,420 devices | Nationwide distribution | Terminated | June 26, 2012 | 490 | 814 | Not_Computer
Z-1386-2011 | 57790 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AUTOMIX 3+3 Compounder with Accusource Monitoring System, product 2M8287; automated nutrition compounders; Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield, IL 60015The Automix Compounder Systems are electromechanical positive displacement fluid compounding devices used in the admixture of large volume parenteral nutrition formulations via gravimetric weighing. | AUTOMIX 33/AS | 2 | February 22, 2011 | 2011 | Baxter Healthcare Corp. | The AUTOMIX compounder may not respond correctly to the key pressed by the operator using the keypad. The incorrect keypad response issue can occur both in manual data entry mode (STD Mode) and when entering order numbers to download from host software (MAN ID Mode). It is possible that an incorrect volume and/or an incorrect solution could be admixed into the final bag. | OTHER/UNDETERMINED: Pending | Baxter Healthcare Corp. sent Urgent Device Correction letters dated January 27, 2011, via first class mail to all affected customers. informing personnel that use Automix compounders to exercise care when entering data, and verify the accuracy of the keypad response. The accounts were requested to convey this information to all users of the Automix compounders, emphasizing the need to verify data accuracy each time the keypad is used. The accounts were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667.After determining the root cause of the incorrect keypad responses, Baxter issued a follow-up correction letter on 4/27/11 outlining the recommended work practices to reduce fluid penetration into the Automix compounder keypad/control module. |  | 950 | 950 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Terminated | February 16, 2012 | 359 | 5160 | Computer
Z-1387-2011 | 57700 | oximeter | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Masimo HemoglobinPronto-7 with Rainbow 40 Technology foreasy-to-use spot-check testing of hemoglobin (SpHb),SpO2, pulse rate, and perfusion indexPronto-7 rainbow 4D Reusable Sensors are an accessory to the Masimo Rainbow SET Pronto-7 Pulse CO-Oximeter. The Masimo Rainbow SET Pronto-7 Pulse CO-Oximeter (instrument) with Masimo Rainbow SET Technology is developed to simultaneously and non-invasively measure functional arterial oxygen saturation (SpO2), pulse rate (PR), total hemoglobin (SpHb) and perfusion index (PI). The instrument has a single on/off button and an interactive touchscreen for administering tests along with user selectable options. | Pulse Oximeter | 2 | February 22, 2011 | 2011 | Masimo Corporation | Masimo Corporation is initiating a voluntary recall on the Pronto-7 rainbow 4D Resuable Sensor because when performing an SpHb measurement, the sensor may incorrectly read actual finger temperature. As a result, Pronto-7 Check Pulse CO-Oximeter may provide an incorrect SpHb measurement. | DESIGN: Process Design | Masimo sent an Urgent Voluntary Medical Device Recall letter dated December 21, 2010, to all affected customers via FexEx on December 23, 2010. The Urgent Voluntary Recall Notice provides instructions on the actions that the customers and field personnel must take to correct the issue.Customers were instructed to choose any of the three options below. Option 1 would ensure accurate results are achieved when using the Pronto-7 rainbow 4D Reusable Sensor with the Pronto-7 Spot Check Pulse CO-Oximeter. (1) Option 1: Prior to performing an SpHb measurement, wait at least seven (7) minutes with the patient's finger inserted inside the Pronto-7 rainbow 4D Reusable Sensor. The seven minute wait will help ensure that the actual finger temperature and the sensor temperature are close to equalized.Masimo is working to quickly develop a revision to the Pronto-7 rainbow 4D Reusable Sensor that will address this issue and plan to have it available in early 2011. We will contact you when the new revision of the sensor becomes available to replace the current sensor;(2) Option 2: Return only the Pronto-7 rainbow 4D Reusable Sensor and request a loaner Pronto device with rainbow DCI sensor (part numbers 9167 and 2646) to be utilized until the new version of the sensor is available for use with the Pronto-7. Please let us know if you wish to try the Pronto or see the product brochure for Pronto. We will contact you when the new revision of the sensor becomes available;(3) Option 3: In addition, the Pronto-7 rainbow 4D Reusable Sensor and the Pronto-7 Spot Check Pulse CO-Oximeter may be returned for a refund.Customers were also instructed to complete the Tracking/Verification Form provided with the recall notice and fax it to Masimo at: 1-949-297-7499. | Contact the recalling firm for information | 497 | 497 | Worldwide Distribution - Nationwide in the US and worldwide to Argentina, Austria, Bahrain, Belgium, Benin, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebaron, Libyan Arab, Jamahiriya, Malaysia, Malta, Mauritius, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Reunion, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United kingdom | Terminated | March 26, 2013 | 763 | 497 | Not_Computer
Z-1389-2011 | 57245 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-ROM Distal Femoral Cutting block Guide SZ MED, REF 216311004, non-sterile, Depuy International Leeds, England.The product is used to make cuts in the femur to prepare the femur for total knee replacement surgery. | SROM distal femoral cutting block guide | 2 | February 22, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm was notified that the Hinge Block Femoral Box Guide cut being off alignment when using the S-ROM Distal Femoral Cutting Block Guide SZ SML/X-SML pin holes to locate. The firm then found that the same issue was found int the MED block size. | DESIGN: Device Design | The firm, DePuy, sent an "URGENT INFORMATION - DEVICE CORRECTION" email and letter dated October 19, 2010 to all customers. The notices described the product, problem and actions to be taken by the customers. The notices stated that: Depuy is issuing a device correction; however, the product may still be used; the devices do not need to be returned; the distributors were to provide a letter to clinicians that were using the device; the surgical technique will be revised in October 2010; the product will undergo a design change, and the recalled product will be traded out in December 2010 and January 2011.If you have any additional questions, please contact DePuy's Scientific Information office at 1-888-554-2482.On 3/4/2011, The firm sent out a revised URGENT DEVICE RECALL asking that all affected cutting guides be taken out of service immediately. Devices should be returned to DePuy and new cutting guides will be sent out within 2 weeks. | Contact the recalling firm for information | 219 | 219 | Worldwide distribution: USA including states of: AZ, AR, CA, CO, FL, GA, IA, ID,IL, IN, KY, KS, LA, MA, MD, MI, MN, NC, NV, PA, TN, TX, VA, WA, and WI; and countries including: Canada and Ireland. Australia, Austria, Canada, Denmark, Ireland, Israel Italy, New Zealand, Russia, Sweden, Switzerland, Spain, and UK. | Terminated | December 03, 2012 | 650 | 448 | Not_Computer
Z-1394-2011 | 57082 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet Cardiopulmonary AG, MAQUET RPM 20: Roller Pump for Heart Lung Machine HL20. | Maquet Heart Lung Machine HL20 | 2 | February 22, 2011 | 2011 | Maquet Cardiovascular Us Sales, Llc | A wrong ball bearing was installed in the pump heads of product. This may make the pump head wear out and cease function. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Maquet hand delivered a letter to the consignee on 10/4/2010. |  | 4 | 4 units | One hospital in New York City | Terminated | August 11, 2011 | 170 | 4 | Not_Computer
Z-1391-2011 | 57805 | apparatus, autotransfusion | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT Cardiovascular - autotransfusion SystemCatalog Number: 02050-US | Haemonetics cardioPAT | 2 | February 22, 2011 | 2011 | Haemonetics Corporation | Update Instructions for Use for the CardioPAT&reg; Manual for Leak Detection. User may mis-interpret the air-leak rate displayed by the cardioPAT device when making a determination about chest-drain removal, which could result in the patient experiencing a pneumothorax. | PRODUCTION CONTROLS: Process Control | Haemonetics issued a Field Correction Notice on 1/25/11 . with the Air Leak Detection addendum and index be inserted into the appropriate document, applied to all cardioPAT@ devices on site. The serial number for each cardioPAT device on site should be recorded on the attached acknowledgementand the completed acknowledgement should be returned to your Haemonetics Sales Consultant. A Haemonetics Sales Consultant will train on the content of these items and insert them into the Operator Manuals. |  | 309 | 309 units | Worldwide distribution, including USA, Australia, Belgium, Canada, Switzerland, China, Germany, Finland, France, United Kingdom, Hong Kong, Israel, Italy, Japan, Netherlands, Russian Federation, and Sweden. | Terminated | January 14, 2013 | 692 | 309 | Not_Computer
Z-1401-2011 | 57781 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM VerioNuclear magnetic resonance imaging | MAGNETOM Verio | 2 | February 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc | On some systems the gradient cable connections did not meet firm's specifications after installation | OTHER/UNDETERMINED: Pending | Siemens Medical Solutions USA, Inc sent customers a letter on January 10, 2011 notifying them of an Update Instruction and informing them that recalling firm would preventatively check the gradient cable connections of all systems. For questions regarding this recall call 610-219-6300. | Contact the recalling firm for information | 92 | 92 | Nationwide Distribution including AK, AZ, CA, CT, FL, HI, IL, IN, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NV, OH, OK, OR, PA, RI, TN, TX, VA, WI, and WV | Terminated | September 25, 2012 | 580 | 92 | Not_Computer
Z-1421-2011 | 57775 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM SIRESKOP SD, SIRESKOP SX, AXIOM LUMINOSSystem, X-Ray, Fluoroscopic, image-intensified; solid state x-ray imager (flat panel/digital imager) | AXIOM SIRESKOP SD, SIRESKOP SX, AXIOM LUMINOS | 2 | February 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm became aware of a potential malfunction and possible hazard to patients in the event the tilt buttons on the OPTI Grip become stuck under the housing | OTHER/UNDETERMINED: Pending | Siemens Medical Solutions USA, Inc issued a Customer Safety Advisory Notice via Update Instruction XP005/10/S to affected customers. This letter informs customers of the potential malfunction and of possible danger for patients, users or other persons. A future Update Instruction will be released for the replacement of the tilt buttons on all OPTI Grip handles. Customers were instructed to forward the safety notice to any new owners of this device and to provide Siemens Medical with names of the new customers. For any questions regarding this recall call 610-448-3237. |  | 1323 | 1323 | Nationwide Distribution including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | July 16, 2013 | 874 | 1323 | Not_Computer
Z-1157-2011 | 57663 | catheter,intravascular,therapeutic,short-term less than 30 days | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Spectrum Central Venous Tray, Minocycline/Rifampin Impregnated Double Lumen Polyurethane, 4.0 Fr/ 8cm, C-SDLMY-401J-ABRM-IHI-FST, G48047, Sterile, Cook Medical, Cook Incorporated, Bloomington, IN. | Spectrum Central Venous Tray, Minocycline/Rifampin Impregnated Double Lumen Polyurethane | 1 | February 23, 2011 | 2011 | Cook, Inc. | Cook, Inc. is recalling Single, Double, Triple, and Five-Lumen Central Venous Catheter Trays AND Single and Double Lumen PICC Peripherally Inserted Central Venous Catheter Trays. These kits and trays contain 0.9% Sodium Chloride Injection USP 5 mL fill in 6 mL flush syringes that have been recalled by Excelsior Medical. During a routine test for syringe sterility, a possible integrity failure wa | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cook, Inc. sent URGENT DEVICE RECALLS letters to its customers dated 10/4/2010. The letters asked customers to examine their inventory and notify Cook for product return. The letter included a response form. |  | 51473 | 51,473 total kits/trays | Worldwide distribution: USA, Chile, Colombia, Panama, Barbados, Canada, Guatemala, and Saudi Arabia. | Terminated | June 21, 2011 | 118 | 51473 | Not_Computer
Z-1400-2011 | 57818 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra Radionics HRAIM Intubation Head Ring AssemblyRef: HRAIM Head Rings serve as the general stereotactic treatment platform. Head Rings are used to provide a reference frame for instrumentation used for precise spatial localization and treatment of physiologic targets for stereotactic neurosurgical procedures such as craniotomies, biopsies, functional neurosurgery, and radiation therapy. Head Rings are delivered to the user non-sterile, and are reusable. | Integra Radionics | 2 | February 23, 2011 | 2011 | Integra LifeSciences Corp. | Overall length of the intubation hoop in the HRAIM Intubation Head Ring Assembly is too long and will not allow a device to to attach | PRODUCTION CONTROLS: Manufacturing Material Removal | Integra sent a Medical Device Field Corrective Action letter dated January 12, 2011, to all affected customers via Federal Express. Integra Sales Specialists were provided with the replacement part and written instructions on how to replace the affected intubation hoops from customer inventory identified in the traceability report. Upon completion of the field correction each sales specialist was instructed to return the affected hoop along with a completed copy the Recall Acknowledgement and Return Form. If additional information is required contact me at (609) 936-2485 or james.tillman@integraJife.com. |  | 8 | 8 units | Worldwide Distribution - USA, Canada, and Australia | Terminated | April 04, 2011 | 40 | 8 | Not_Computer
Z-1399-2011 | 57431 | hepatitis viral b dna detection | MKT |  |  | Microbiology |  | Abbott RealTime HBV Assay, List 2N40, in vitro diagnostic; Abbott Molecular Inc., Des Plaines, IL 60018 USAThe Abbott RealTime HBV is an in vitro polymerase chain reaction (PCR) assay for the quantitation of Hepatitis B Virus (HBV) DNA in human plasma or serum from HBV-infected individuals. The Abbott RealTime HBV assay is intended for use in conjunction with clinical presentation and other laboratory markers as an indicator of disease prognosis and for use as an aid in assessing viral response to antiviral treatment as measured by changes in plasma or serum HBV DNA levels. This assay is not intended for use as a screening test for HBV or as a diagnostic test to confirm the presence of HBV infection. | Abbott RealTime HBV Assay | 3 | February 23, 2011 | 2011 | Abbott Molecular | Internal testing identified the potential of under filled and over filled enzyme reagent vials which may result in obtaining less than the total number of tests per kit as described on the package insert. | OTHER/UNDETERMINED: Pending | Abbott Molecular sent a Field Advisory Notice letter to all RealTime HIV-1, CT/NG and HBV assay customers on December 20, 2010, informing them that internal testing identified the potential of under filled and over filled enzyme vials which may result in obtaining less than the total number of tests per kit as described in the package insert. The customers were told that specimens that meet acceptance criteria, per the product labeling are not affected by the situation and are considered valid results, and that the issue has no impact on results for those runs that generate valid results. Any questions were directed to Abbott Technical Support at 1-800-553-7042, Option 2. |  |  |  | Nationwide Distribution | Terminated | November 19, 2012 | 635 |  | Not_Computer
Z-1396-2011 | 57861 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Comfort Sling Plus Model 300, Part number 35300305, Size Medium, LIKO, Sweden. Liko a Hill-Rom Company.Enables a comfortable sitting position and adapts to the patient without the need for individual adjustments when transferring a patient with a lift. | Comfort Sling Plus Model 300 | 2 | February 23, 2011 | 2011 | Hill-Rom, Inc. | Slings were improperly assembled and the loops that are used to support the sling when hung from a lift were too long, potentially putting the patient being lifted at risk of falling out. |  | The firm sent an URGENT MEDICAL DEVICE RECALL notice dated 02/09/2011 to Facility Risk Manager/Facility Administrators. The letter Identified the affected product, explained the recall situation, and asked firms to check their facility for the affected serial numbers. Consignees were asked to return the recalled products, and a replacement sling would be sent to them. Customers can contact Hill-Rom Technical Support at 800-445-3720, if they have any questions concerning this matter. |  | 10 | 10 units | Worldwide Distribution -- US, Switzerland, Germany, Denmark, Finland, France, UK, Italy, Netherlands, Sweden, and Canada. | Terminated | October 07, 2011 | 226 | 10 | Not_Computer
Z-1424-2011 | 57084 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | 1) Signa Lx ( K962061Signa Horizon Cx)The Signa Horizon Cx System is indicated for use as a diagnostic imaging device to produce transverse, saggital, coronal and oblique images of the internal structures of the head or body. The images produced by the signa Horizon Cx System reflect the spatial distribution of protons (hydrogen nuclei) exhibiting magnetic resonance. The NMR properties that determine the image appearance are proton density, spin-lattice relaxation time (T1), spin-spin relaxation time (T2) and flow. When interpreted by a trained physician, these images provide information that can be useful in determining a diagnosis. | Signa Lx | 2 | February 24, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the gradient cables being inadvertently swapped while servicing the gradients on some GE Healthcare MR Systems that may impact patient safety.If the gradient cables are inadvertently swapped while servicing the gradients, and the Service Engineerdoes not perform the geometry check as required in the service procedure, images m | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On November 9, 2010, the firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated November 5, 2010 to all Consignees/Customers. The letter described the product, problem and action to be taken by the firm. GE instructed the customers to contact their local GEHC support Engineer, if they have any questions about the image orientation, or the procedures for geometry checks on thier system. Additionally, a GE Healthcare representative will install a device, which will reduce the potiential for the gradient cables to be swapped.Contact your local GEHC Support Engineer if you have any questions concerning this notification or phone: 262-548-2731. | Contact the recalling firm for information | 3105 | 3105 | Worldwide distribution: USA including DC and PR; and countries including: VIET NAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, FORMER YUGOSLAV REPUBLIC OF MACEDONIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, REPUBLIC OF KOREA, PORTUGAL POLAND, PHILIPPINES, PERU,PARAGUAY, PANAMA, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MAURITIUS, MARTINIQUE,MALTA, MALAYSIA, LUXEMBOURG,LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, KENYA, KAZAKHSTAN, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUYANA, GUATEMALA, GUADELOUPE, GREECE, GERMANY, GEORGIA, FRENCH POLYNESIA, FRANCE, FINLAND, EL SALVADOR, EGYPT, ECUADOR, DOMINICAN REPUBLIC, DENMARK, CECH REPUBLIC, CYPRUS, CROATIA, COLOMBIA, CHINA, CHILE, CANDA, BULGARIA, BRAZIL, BOSNIA AND HERZEGOVINA, BOLIVIA, BELGIUM, BELARUS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | July 03, 2012 | 495 | 8937 | Not_Computer
Z-1434-2011 | 57821 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cardiovascular Review StationThe IMPAX Cardiovascular suite is a cardiovascular information system, providing image archiving, image display and modality/study specific structured reporting. | CRS | 2 | February 24, 2011 | 2011 | AGFA Corp. | Echocardiography volume measurements performed within CRS were not reflected correctly in associated Results Management (RM) reports. | OTHER/UNDETERMINED: Pending | The firm, AGFA Healthcare, sent a "Safety Notice" dated January 27, 2011 to all customers. The notice described the product, problem and action to be taken by firm and customer. AGFA informed the customers that a software upgrade was made between December 2008 and February 2010 and instructed the customers to distribute this information to all individuals within their facility and to complete and return the attached URGENT SAFETY NOTICE FEEDBACK FORM via fax to 864-421-1664 or email to debbie.norris@agfa.com. Note: Service Bulletin DD+DIS100.09E, was issued March 2009 and an upgrade from Software Versions 2.12.05 through 2.14.03.SU1 to Software Version CRS 2.14.03.SU2 was required for the impacted sites. Most identified consignees/customers required the software upgrade, but some consignees/customers had already been upgraded prior to the event and did not require additional upgrades.If you have any questions about this matter, please contact the Post Market/QARA North America via email at debbie.norris@agfa.com or call at 864-421-1754. | Contact the recalling firm for information | 107 | 107 | Worldwide distribution: USA and country of: Canada | Terminated | March 29, 2011 | 33 | 107 | Software | Output/Calculation | Software update | Software Update
Z-1422-2011 | 57704 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV with Artiste Digital Linear Accelerators in combination with the Flat Panel Positioner; Part No. 8139789;Intended use: To deliver x-ray radiation for therapeutic treatment of cancer.Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Concord, CA; and Siemens Medical Solutions AG, Roentgenstrasse, Kemnath, Germany, and Siemens Medical Solutions AG, Henkestrasse, Erlangen, Germany | ARTISTE MV with Artiste Digital Linear Accelerators | 2 | February 24, 2011 | 2011 | Siemens Medical Solutions USA, Inc., Concord, CA | The firm become aware of a potential issue with ARTISTE system that may result in patient injury should the positioner belt break, and the cam disc (secondary safety mechanism) is not installed and if the flat panel is over the patient. | OTHER/UNDETERMINED: Pending | The firm, Siemens, issued a "Customer Information" letter on December 3, 2010 via mail to all its direct consignees/customers. The letter described the product, problem and actions taken by the firm to the customer. Siemens arranged for an inspection of their customers systems to make sure the safety mechanism is fully functional. Siemens instructed the customers to include the Customer Information letter in their ARTISTE System Owner manual chapter 'Safety Advisory Letters' .If you any questions, please call 925-246-8407. |  | 20 | 20 units | Worldwide distribution. | Terminated | June 30, 2011 | 126 | 20 | Not_Computer
Z-1423-2011 | 57629 | generator, oxygen, portable | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | HomeFill Oxygen Compressor IOH200 Homefill, The serial numbers are as follows: 10JF034081, 10JF034082, 10JF034083, 10JF034084, 10JF034085, 10JF034086, 10JF034087, 10JF034088, 10JF034089, 10JF034090, 10JF034091, 10JF034092, 10JF034093, 10JF034094, 10JF034095, 10JF034096, 10JF034097, 10JF034098, 10JF034099, 10JF034100, 10JF034101, 10JF034102, 10JF034103, 10JF034104, 10JF034105, 10JF034106, 10JF034107, 10JF034108, 10JF034109, 10JF034111, 10JF034112, 10JF034113, 10JF034114, 10JF034115, 10JF034116, 10JF034117, 10JF034119, 10JF034120, 10JF034121, 10JF034122, 10JF034123, 10JF034124, 10JF034125, 10JF034126, 10JF034127, 10JF034128, 10JF034129, 10JF034130, 10JF034131, 10JF034132, 10JF034133, 10JF034134, 10JF034135, 10JF034136, 10JF034177, 10JF034178, 10JF034179, 10JF034180, 10JF034181, 10JF034182, 10JF034183, & 10JF034184.The IOH 200 Homefill is an electromechanical, prescription device designed for use in the home. Its intended function and use is to provide supplemental oxygen to patients and to supply pressurized oxygen to fill cylinders for the patients personal ambulatory use. | HomeFill Oxygen Compressor IOH200 Homefill | 2 | February 24, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that inadequate clearance between a motor ground clip and the motor windings may result in an increased risk for shock to service personnel. It also may result in contact with the motor winding, resulting in a shorted winding; this condition has the potential to result in motor failure/shutdown or to cause the motor | OTHER/UNDETERMINED: Pending | The firm, INVACARE Corp., sent an "URGENT PRODUCT RECALL: IOH 200 Homefill Recall" letter dated January 10, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to "immediately examine their inventory and quarantine product subject to recall. In addition, if they may have further distributed this product, please identify their customers and notify them at once of this product recall." The customers were also instructed to complete and return an enclosed response form. Note: All product is to be return to the recalling firm.If you have any questions, contact the Manager, Regulatory Affairs at (440) 329-6356. | Contact the recalling firm for information | 62 | 62 Compressors | Nationwide distribution: USA including states of: FL, GA, KY, LA, MO, NC, NE, NJ and SC. | Terminated | November 21, 2011 | 270 | 62 | Not_Computer
Z-1402-2011 | 57487 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Navigation System - iNtellect Cranial Software, Part number 6000-650-000.It is part of the product series of the Stryker Navigation System. The system comprises software for pre-operative surgical planning and supports computer assisted surgery based on a wireless optical tracking localization device for the use in navigated cranial surgery. It includes automatic multi-modality correlation. The system provides an easy-to-use work flow concept for image importation, FESS, Craniotomy, Biopsy, etc. Additional, automatic segmentation of brain surface, tumor, vessels, etc. are supported. | STRYKER NAVIGATION SYSTEM iNtellect Cranial Software | 2 | February 25, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | When users imports more than one image set and changes the viewing direction for both images, the orientation of every other image series might be stored incorrectly, as a mirror image. This may lead to insufficient or inaccurate correlation results or incorrectly shown orientation of every second image series during navigation. | DESIGN: Software Design | Urgent: Medical Device Recall Notification letters were sent on 12/3/2010, and software upgrades were to be performed in the field by the end of 2010. Specific Customer Instructions: 1) Locate the units listed in the notification.2) Distribute this notification to all departments affected.3) Return the enclosed business reply form to confirm receipt of this notification and identify how many affected units are currently in your inventory. 4) Once the new software is available, Stryker will dispatch a service tech in order to update the software. Note: Software upgrades for all applications will be available by the end of the year. The upgrades will be scheduled as we receive your business reply form.Note: If you currently have a Time License or Pay Per Use card, a new one will be issued to you at the time of the update.5) If you have loaned or sold any of the units listed in this notification, please, forward a copy of this notice to the new users and advise Stryker of the new users name and their new location.6) If you have disposed of any of the affected units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.Questions regarding this recall should be directed to Stryker Field Action Team at 800-800-4236, ext 3808, 3584, or 4354 between the hours of 8 a.m. - 5 p.m. Monday - Friday. | Contact your health professional | 1155 | 1,155 total | Worldwide Distribution -- USA, Canada, Mexico, Australia, Argentina, Brazil, United Kingdom, Japan, Poland, S. Korea, Hong Kong, Taiwan, India, and New Zealand. | Terminated | June 01, 2011 | 96 | 1155 | Software | Display/Image | Software update | Software Update
Z-1439-2011 | 57829 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Curlin Infusion Administration Set, REF 340-4111, Sterile EO, Zevex, Inc., 4314 Zevex Park Lane, Salt Lake City, UT 84123.The Curlin Ambulatory Infusion Pump System with disposable PVC Administration Tubing Sets and user accessories, provide a means for the volumetric delivery of fluids used in Parenteral, Enteral, Epidural, Subcutaneous and Intravenous applications. | Infusion Administration Set | 2 | February 25, 2011 | 2011 | Zevex International, Inc. | Defect in tubing may trigger false Air-In-Line warning. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, MOOG, sent an "Expansion of Moog/Curlin Ambulatory Pump Administration Set Recall" letter dated May 2, 2011 to its customers. This letter informed the customers that this recall has been expanded to included specific lot numbers of Curlin model ambulatory infusion administration sets. The customers were instructed to follow the outlined process in the letter: remove suspect product code from inventory, contact MOOG Customer Service at 1-800-970-2337, prompt #7, contact your distributor to arrange replacement product, request a shipping return label as per the distributor's process and ship back to their location and follow the temporary solution outlined in the letter. Moog Customer Service will provide a call tag for convenient product return.If you have any questions contact MOOG Customer Service at 1-800-970-2337. The firm, MOOG, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 24, 2011 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to remove suspect product code from inventory; contact their distributor to arrange replacement product, if purchased from distributor and request shipping return label per distributors process and ship back to their location or contact MOOG Customer Service at 800-970-2337, prompt #7, if purchased directly from MOOG and MOOG Customer Service will provide a call tag for convenient product return; follow the temporary solution provided until replacements are administered; and complete and return the MEDICAL DEVICE CORRECTION REPLY FORM in the self addressed envelope provided.If you have any questions, contact MOOG Customer Service at 800-970-2337. |  | 240 | 17,120 units (US), 240 units (INTL) | Worldwide Distribution: USA including states of: AR, CA, FL, IA, KS, MS, NE, NY, OH, OK, TN, TX, and UT; and countries including: Australia, Croatia, France, Germany, Kuwait, New Zealand, Qatar, Saudi Arabia, South Africa, United Arab Emirates, and UK. | Terminated | September 27, 2011 | 214 | 629950 | Not_Computer
Z-1435-2011 | 57820 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Results Manager/Results Manager Administration ToolThe IMPAX Cardiovascular suite is a cardiovascular information system providing image archiving and image display and modality study-specific structured reporting. The suite is comprised of modular software that runs on industry standard personal computers and servers, utilizing Windows 2000/2003/XP operating systems. Image data storage/display utilizes DICO | RM, RMAT | 2 | February 25, 2011 | 2011 | AGFA Corp. | User could modify a measurement's units and that change would not be reflected in echocardiography formulas used within the Adult echocardiography RM. | OTHER/UNDETERMINED: Pending | The firm, Agfa Healthcare, sent individual letters dating from April 15, 2009 to August 26, 2009, to the customers. The letters described the product, problem and actions to be taken. Agfa provided the affected customer with a complete list of their impacted patients and the measurements. Agfa also scheduled unit corrections with their customers and received acknowledgement of the letters at the time of scheduling the unit corrections. Service Bulletin DIS026.09E was issued January 2009 to inform users that Agfa will no longer allow customers to request unit changes for measurements that have formulas. Note: Technical Support ran diagnostic scripts for each customer and documented in GSC cases for each consignee/customer.If you have any questions, contact Agfa Healthcare at (949) 830-4959. |  | 6 | 6 devices | Nationwide distribution: USA including states of: IL, MT, NY, OR and SC. | Terminated | October 13, 2011 | 230 | 6 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0893-2011 | 54768 | upc reader (bar code reader) | RFZ |  |  |  |  | Hemochron Signature Elite Whole Blood Microcoagulation (IVD) System.The system is intended to be used only with test cuvettes that are available from ITC. | Hemochron Signature Elite Whole Blood Microcoagulation System | 2 | March 01, 2011 | 2011 | International Technidyne Corp. | There is a mismatch between the class of the laser barcode scanner and the labeling of the Hemochron Signature Elite: some of the devices were inadvertently equipped with class II laser barcode scanner instead of the intended class I barcode scanner. | PRODUCTION CONTROLS: Labeling Mix-Ups | Urgent - Medical Device Recall - Hemochron Signature Elite letters, dated February 24, 2010, by Federal Express to all direct customers. The letter informed customers of the issue and stated that the problem does not negatively affect the performance of the product and to continue using the product as intended. In addition, customers are to notify the appropriate personnel about the notice, complete and return the attached form, and an ITC representative will contact customers to discuss arrangements to re-label their devices. Customers can contact ITC Technical Support at 800-579-2255 (US) or 732-548-5700 ext 4700 (International) if they have any questions regarding this issue and notice. Customers can also email sigelitescan@itcmed.com with questions. |  | 1640 | 1640 units (1303 US, 337 foreign) | Worldwide Distribution -- United Arab Emirates, Belgium, Canada, Switzerland, Germany, Denmark, France, United Kingdom, Ireland, Iceland, Italy, Kuwait, Liechtenstein, Netherlands, Oman, Qatar, Saudi Arabia, Sweden, and Singapore. | Terminated | September 25, 2012 | 574 | 1640 | Not_Computer
Z-1457-2011 | 57402 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | QuickVue One-Step hCG Urine, Quidel Catalog Number: 20109, Lot Number: 706842 | QuickVue OneStep hCG Urine | 2 | March 01, 2011 | 2011 | Quidel Corporation | The recall was initiated after Quidel Corporation determined that a small number of QuickVue One-Step hCG Urine kits have been incorrectly packaged with QuickVue H. pylori gll cassettes inside the kit boxes. These foil pouches and testcassettes are clearly labeled as H. pylori. Nevertheless, there is a risk that customers may overlook this labeling. |  | Quidel Corporation initiated an "Urgent Medical Device Recall" letter on November 23, 2010, with an attached "Distributor Recall Notification Form" to the affected distributor consignees notifying them of this issue and requesting return of all kits from the affected lot. Quidel is further requesting that distributor consignees notify end-users of the issue and request return of kit boxes containing the H. pylori pouched cassettes. The recall notices sent to the distributors were sent via US Postal, certified, with return receipt.The consignees were instructed to:(1) Locate kits with the affected lot number and immediately remove the kits from your existinginventory and place them into a quarantine location, (2) Contact Quidel Corporation for instructions on how to return the kits and obtain replacement product.Quidel Technical Support(800) 874-1517, Monday-Friday, 7am - 5pm PST,(3) Fax the enclosed Inventory Form to Quidel at (858) 431-3522. Indicate the total number of kits currently on hand with lot number 706842 or if you do not have any kits of this lot numberremaining in inventory. |  | 3624 | 3624 kits = 90,600 tests total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, HI, IL, KS, KY, MA, MD, MN, NC, NJ, NM, NV, OK, OR, PA, TN, TX, UT, VA, and WA and Canada. | Terminated | June 26, 2012 | 483 | 3624 | Not_Computer
Z-1458-2011 | 57561 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The Panorama HFO system has the Traceable Item Identification (12NC):9896 030 17741-SW R 2.5.1 (shipped as R 2.5.1, upgraded to R 2.5.3 via FCO).9896 030 19641-SW R 2.5.39896 030 19861-SW R 2.6.19896 030 19862-SW R 2.6.19896 030 19863-SW R 2.6.3 | Panorama HFO | 2 | March 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | In the ExamCard MobiFlex the Patient Position is "Prone" which must be "Supine". The result is mislabeling of left and right side of the patient. If this is not detected by the operator there is a risk of wrong treatment or the need of additional X-RAY during treatment. | OTHER/UNDETERMINED: Pending | Philips began sendng out the URGENT-Device Correction Notice letter dated February 5, 2010, ro all affected consignees on February 10, 2010. The consignees were advised to change the patient position from "Prone" immediately to "Supine" and store the corrected ExamCard in the hospital database if they use the ExamCard, MobiFlex dual Injection.Consignees were instructed to call Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and also reference "FCO 78100318" with any communication. | Contact the recalling firm for information | 82 | 82 units | Nationwide Distribution including the following states: AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV.AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV. There was no foreign consignees. | Terminated | August 10, 2011 | 162 | 82 | Not_Computer
Z-1472-2011 | 57514 | control material, blood circulating epithelial cancer cell | NRS | Hematology quality control mixture. | Hematology | Hematology | 510(k) | Veridex, Cell Search Circulating Tumor Cell Control Kit. Product code 7900003 | Cell Search CTC Control Kit | 3 | March 01, 2011 | 2011 | Veridex, LLC | There is a discrepancy between the assigned range of values on the high control label and the range listed on the Cell control Assigned Value Label for the Cell Search Circulating Tumor Cell Control Kit. | PRODUCTION CONTROLS: Process Control | Veridex notified all affected consignees on November 11, 2010 by Federal Express overnight letter and prodivded with the informaiton regarding the label discrepancy. Foreign consignees were also notified on November 11, 2010 of the issue and instructed to notify their affected consignees. Consignees are asked to return the confirmation of receipt form by 11/29/2010. | Contact the recalling firm for information | 122 | 122 US, 28 International | Worldwide distribution, including: USA, Canada, France, UK, Germany, Italy, China, Spain & Portugal. | Open |  |  | 122 | Not_Computer
Z-1459-2011 | 57560 | hepatitis viral b dna detection | MKT |  |  | Microbiology | PMA | COBAS AmpliPrep/COBAS TaqMan HBV v2.0 US-IVD;COBAS AmpliPrep/COBAS TaqMan HBV v2.0 CE-IVD | COBAS Ampliprep/COBAS TaqMan HBV v2.0 USIVD COBAS Ampliprep/COBAS TaqMan HBV v2.0 CEIVD | 2 | March 01, 2011 | 2011 | Roche Molecular Systems, Inc. | There is an issue regarding liquid level detection of the CS1 Magnetic Glass Particles (MGP) Reagent Cassette of the COBAS AmpliPrep /COBAS TaqMan HBV v2.0 Test on the COBAS AmpliPrep instrument. | DESIGN: Software Design (Manufacturing Process) | Roche sent Urgent Medical Device Correction Letters to their consignees on December 28, 2010 via Fed Ex. The letters advise the customers to verify the correct configuration of the CSI reagent cassette, remind customers to follow the package insert instructions for including all required controls for each batch and advise that review of previously generated results may be warranted. A UMDC Faxback form was attached for the consignee to complete.For questions regarding this recall call 908-253-7569. |  | 22212 | 93 kits to US; 22,119 kits to foreign | Worldwide Distribution - USA including CA, CT, IN, MI, MN, NC, NE, NY, OR, TX, WA, and WI and the countries of Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mauritania, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudi Arabia, Singapore, Slovakia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, and the United Kingdom | Terminated | August 27, 2012 | 545 | 22212 | Not_Computer
Z-0286-2011 | 56922 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel/Edge/Elite Select Computed Tomography X-Ray System.For head, whole body X-ray Computed Tomography applications. | BrightSpeed Excel/Edge/Elite Select Computed Tomography XRay System | 2 | March 01, 2011 | 2011 | GE Healthcare, LLC | Failure to properly document the CTDI in the Technical Reference or User Manual. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Planned action(s) to repair defect or to bring product into compliance: GE Healthcare will bring the Technical Reference Manual into compliance by correcting the noncompliance in forward production and the installed base units. This will be accomplished by delivering an "Important Electronic Product Radiation Warning" letter and a TRM addendum which contains the corrected information to customers. This will be carried out via a Field Modification Instruction IFMI 229251 planned to be released in October, 2010 with an estimated completion date of April, 2011. The letter identifies the affected product and describes the defect and related hazards. It also asks customers to refer to the TRM addendum for the scan time accuracy, the helical scan exposure time, and the CTDI adjustment factors for 4 and 8 slices. Contact information is also provided. |  | 103 | 103 units | Nationwide Distribution | Terminated | December 05, 2013 | 1010 | 103 | Not_Computer
Z-1495-2011 | 53690 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA.EIA and IFA processing system that consists of a PhD fluid handling station (s) linked to a network computer via ethernet hub. The computer provides worklist generation, data management, data reduction and microplate reader control functions. | PhD System EIA/IFAversion 2.1A software | 2 | March 02, 2011 | 2011 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.Customers who run competitive immunoassays must revert to software version 2.0B.Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 510-741-3954. |  | 262 | 262 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, FL, IL, MA, MD, MN, MO, MS, MT, NC, NY, OH, PA, SD, TN, TX, UT, and VT and the countries of Australia, New Zealand, Singapore, China, France, Sweden, and Canada | Terminated | March 11, 2011 | 9 | 262 | Software | Output/Calculation | Revert to a previous version | Software Update
Z-1479-2011 | 56385 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA ANA Screening Kit, 96 Tests,Catalog No. 96AN-BPU-TRINITY. Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA and distributed by Trinity Biotech, USA.Intended to screen for the presence of antinuclear antibodies (ANAs) in human serum as an aid in the diagnosis of certain rheumatic diseases. | Autoimmune EIA ANA Screening Kit | 2 | March 02, 2011 | 2011 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the ANA Screening kit due to a decline in stability. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An Urgent: Medical Device Recall Notification, dated March 1, 2010, was provided to the firm's consignee, identifying the affected product and stating the reason for the recall. The letter also asked the customer to examine their inventory and quarantine the affected kits. If the product had been further distributed, those customers must be identified, notified, and asked to discontinue use of the affected product. Customers were to discard all affected product and complete and return the enclosed Medical Device Recall Response Form. The letter also stated that replacement kits would be available in mid-March. |  | 2109 | 2,109 bulk packs | Nationwide Distribution -- New York, US. | Terminated | March 11, 2011 | 9 | 2109 | Not_Computer
Z-1477-2011 | 57427 | test, epithelial ovarian tumor-associated antigen (ca125) | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 125 II Assay packaged in 100 and 500 test kits as follows:Product code: 09427226 (128532) 100 test kit Product code: 01678114 (128533) 500 Test kit | ADVIA Centaur CA 125 II | 2 | March 02, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc | Positive bias (40-55%) with patient results greater than the upper limit of normal and less than the upper limit of normal | PRODUCTION CONTROLS: Process Control | Siemens Healthcare Diagnostics on 12 /14/10 issued Urgent Device Recall notice via Federal Express to accounts. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Device Recall notice directs customers to immediately discontinue the use of ADVIA Centaur CA 125 II assay kits that contain CA 125 II reagent lot 144. It also recommends that patient results obtained with this reagent lot be reviewed. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory action. Technical questions, please contact the Technical Solutions Center at 1-877-229-3711. |  | 2960 | 2960 units | Worldwide distribution, including USA, Canada, Argentina, Austria, Belgium, China, Denmark, France, Germany, Great Britain, Greece, India, Italy, Japan, Korea, Lavitia, Lithuania, Malyasia, Mexico, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Romania, Saudi Arabia, Switzerland, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Czech Republic, Taiwan, Turkey, Uganda, and Venezuela | Terminated | April 12, 2013 | 772 | 2960 | Not_Computer
Z-1478-2011 | 55293 | antinuclear antibody, antigen, control | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | DiaSorin ETI-ANA Screen Kit, 96 Test, Model Number N0093, manufactured by Bio-Rad Laboratories. | DiaSorin ETIANA Screen Kit, 96 Test | 2 | March 02, 2011 | 2011 | Bio-Rad Laboratories Inc | Inaccurate results-- Consumer complaints of increased positivity rates with the ANA screening kit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Recall was initiated on 3/1/2010. DiaSorin S.R.I. (the sole consignee) was alerted via email using the firm's Medical Device Recall Notification. Consignee informed of the recall and instructed to dispose of the affected product according to local laws. | Contact the recalling firm for information | 999 | 999 kits | Foreign Only:One distributor, DiaSorin S.R.I. in Italy.No wholesalers.No retailers.No domestic distribution, No Federal Government sales/distribution centers. | Terminated | March 11, 2011 | 9 | 999 | Not_Computer
Z-1473-2011 | 57083 | [6 empty] | GE ADVANTAGE SIM MD versions 5.x, 6.x, 7.x up to and including version 7.7.0Advantage Sim MD is used to prepare geometric and anatomical data relating to a proposed external beam radiotherapy treatment prior to dosimetry planning. Anatomical volumes can be defined automatically or manually in three dimensions using a set of CT images acquired with the patient in the proposed treatment position. Definition of the anatomical volumes may be assisted by additional CT, MR, PET or SPECT studies that have been co-registered with the planning CT scan. Additionally, CT & PET data from a respiratory tracked examination may be used to allow the user define the target or treatment volume over a defined range of the respiratory cycle. The geometric parameters of a proposed treatment field are selected to allow non dosimetric, interactive optimization of field coverage. Defined anatomical structures and geometric treatments fields are displayed on transverse images, on reformatted sagittal, coronal or oblique images, on 3 D views created from the images, or on a beam eye's view display with or without the display of defined structures with or without the display of digitally reconstructed radiograph | AdvantageSim MD | 2 | March 02, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that inconsistent DRR (Digitally Reconstructed Radiographic) visualization may occur when the blended DRR feature or the Isocenter Move mode feature is used in combination with specified other functions of the AdvantageSim application on Advantage Workstation. These inconsistencies may affect patient safety. Other AdvantageSim functions are not affected. | DESIGN: Software Design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 23, 2010. The letter was addressed to Hospital Administrators / Risk mangers, Oncology Department Managers / Radiation Oncologists, & Radiation Physicists / Dosimetrists / Therapists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For any questions or concerns regarding this recall, please call one of thefollowing phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. | Contact the recalling firm for information | 1531 | 1531 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MI, MN, MS,MO, NT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, PAKISTAN, NOWAY, NIGERIA, NETHERLANDS, MEXICO, MALAYSIA, LITHUANIA, LATVIA, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, ECUADOR, DOMINICA REPUBLIC, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BRAZIL BELGIUM, AUSTRIA, AUSTRALIA. | Terminated | August 03, 2012 | 520 | 1531 | Other |  |  | Other
Z-1476-2011 | 57692 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning SystemTo plan multiple beam radiation therapy treatments. | PrecisePLAN Treatment Planning System | 2 | March 02, 2011 | 2011 | Elekta, Inc. | PrecisePLAN generates digitally Reconstructed Radiographs (DRR;s) with a shift in the superior direction. | OTHER/UNDETERMINED: Pending | Elekta sent an Important Notice A308 letter dated January 28, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Elekta representative directly if they had any questions. . |  | 277 | 277 units | Worldwide Distribution - USA and the countries of Bulgaria, Switzerland, Czech Republic, Germany, Spain, France, Hungary, Ireland, Italy, Liechtenstein, Netherlands, Poland, Portugal, Sweden, Slovak Republic, Australia, China, Hong Kong, and Japan, | Terminated | April 21, 2011 | 50 | 277 | Not_Computer
Z-1474-2011 | 57419 | motor, drill, electric | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Aesculap MicroSpeed Uni Control Box, Part #GD670Intended for high speed cutting, sawing, drilling and manipulation of soft tissue and bone in the fields of Spine, ENT, Neuro and Maxillo facial Surgery. | MicroSpeed Uni Motor System | 2 | March 02, 2011 | 2011 | Aesculap Inc | Potential safety issue in that the control box can misinterpret the information sent from the low speed motor if the control magnet inside the motor changes its posthion. | DESIGN: Software Design | AESCULAP sent an IMPORTANT CORRECTION NOTICE letter dated November 22, 2010, to all affected consignees. The letter identified the product, the problem, and the action the consignee should take. Consignees were instructed to not use the motors for thread cutting and removal of screws.If any unusual reactions of any cutting tools were observed consignees should check to ensure that the motor is not running in the opposite direction.If any error message is displayed, consigness were instructed to not use the motor.Consignees were also instructed to check both directions of running before use and do not use a motor which produces an error message in one direction. Aesculap would be providing all customers with a software update to the MicroSpeed Uni control box either by the unit being returned to Aesculap for update or by firm representative doing the upgrade on site at their facility.For questions regarding this recall call 800-258-1946 x5254 or 800-258-1946 x5267. | Contact the recalling firm for information | 91 | 91 units | Worldwide Distribution - USA including CA, CO, DE, FL, GA, IN, LA, MO, NC, NJ, OK, PA, TN, TX, and VA and the country of Canada. | Terminated | July 05, 2012 | 491 | 91 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1475-2011 | 57100 | spinal vertebral body replacement device | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Vu a-POD Intervertebral Body Fusion Device. Packaged as surgical kits containing all implants and instruments required for a procedure. Also packaged as individual instruments and implants in poly bags distributed as replenishment for surgical kits. Theken Spine, LLC, Akron, OH 44306.Indicated for use as an adjunct for fusion in patients with degenerative disc disease (DDD) at one or two contiguous levels from L2 to S1. | Vu aPOD Intervertebral Body Fusion Device | 2 | March 02, 2011 | 2011 | Integra LifeSciences Corp. | The IFU provided with the device contains incorrect information in the Indications for Use Section, stating the following: The Theken Spine Vu a"POD Intervertebral Body Fusion Device is intended for use as a stand alone implant when used with the SpinPlate provided in the set. | TRAINING: Employee Error | On 10/19/2010 the firm sent "URGENT : FIELD CORRECTION ACTION Letters to their customers. The letters identified the affected product and explained the reason for the recall. The letter also stated that the issue is not believed to have a health impact on patients. Physicians are asked to continue to exercise their clinical judgment with any periodic patient follow up. Attached to the letters were IFUs containing the correct information. Customers are to complete and return the enclosed Recall Acknowledgement Form per the instructions provided. Questions can be directed to Suzanne B. Wojcik, Sr. Dir Regulatory Affairs and Quality Assurance, at 330-835-8528. |  | 2163 | 2,163 pieces | Nationwide Distribution -- AL, AZ, CA, CT, ID, KS, MO, NH, NV, NY, OH, OK, TX & UT. | Terminated | September 19, 2012 | 567 | 2163 | Not_Computer
Z-1497-2011 | 57884 | tape and bandage, adhesive | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General & Plastic Surgery |  | High Viscosity Dermabond ProPen Topical Skin Adhesive 2-Octo Cyanoacrylate, Sterile, Rx only, Ethicon, Inc. | Dermabond Topical Skin Adhesive | 2 | March 02, 2011 | 2011 | Ethicon, Inc. | Some units of Dermabond Topical Skin Adhesive and Dermabond ProPen Topical Skin Adhesive contain discolored product. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Ethicon sent their recall letters with business reply cards attached on February 2, 2011 via UPS overnight mail. Customers are informed of the issue and asked to return the attached business reply card. |  | 725722 | 725,722 each (total) | Worldwide distribution, including USA, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Czech Repub, Chile, China, Denmark, Dominican Repub, Egypt, Finland, France, French Guiana, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, ireland, Israel, italy, Japan, Korea, Lithuania, Luxembourg, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Palestine, Puerto Rico, Reunion, Russia, Senegal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Ukraine, and Venezuela. | Terminated | November 07, 2012 | 616 | 1451494 | Not_Computer
Z-1494-2011 | 54620 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Male Driver 4.0mm (catalog # MDr140L) and Male Driver 3.2mm (catalog # MDr132L), an instrument used for the insertion of the Fassier-Duval Telescopic IM System intramedullary nail, for bone fixation. | FassierDuval Telescopic IM System | 2 | March 02, 2011 | 2011 | Pega Medical Inc. | Possible breakage. Structural weakness has been detected on the Male Driver 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. | PRODUCTION CONTROLS: Process Control | On 01/29/2010, Pega Medical began contacting their consignees by e-mail and telephone.Pega Medical followed up with their consignees by sending out the replacement components along with the explanation letter. The subject on the letter is Notice for 3.2 mm and 4.0 mm Male Drivers of the Fassier-Duval IM Telescopic System.The consignees were informed that there could be a structural weakness in some of the Male Driver instruments. This problem has been detected on the 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. Consignees can contact Pega Medical at 450-688-5144 x 242. |  | 4 | 4 units of MDr132L (3.2 mm) and 3 units of MDr140 (4.0 mm) distributed in US. | Components were distributed to four hospitals in the US in the state of GA, OK, TN, and UT. | Terminated | March 26, 2011 | 24 | 4 | Not_Computer
Z-1493-2011 | 57827 | electrode, electrocardiograph | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Neotrode Wire Attached, Catalog Number: 1731-003; Neotrode II Wire Attached, Catalog Number: 1741-003; Softrace Neonatal Wire Attached, Catalog Number: 2321-003; Softrace Pediatric Wire Attached, Catalog Number: 2331-003; Softrace Limb Band, Catalog Number: 2310-003; TRU-LINK ECG Electrode (Neonatal Pre-Wired), Catalog Number: 685-0037-00; for use with electrocardiographic monitoring equipment to detect neonatal/pediatric heart action voltages. | Neotrode, Softrade, TruLink | 2 | March 02, 2011 | 2011 | ConMed Corporation | Poor signal or loss of signal. Certain recalled devices have a loose wire attachment that may lead to intermittent or poor signal quality. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ConMed Corporation distributed an Urgent Device Recall notification letter on Tru-Link ECG Electrode (Neonatal Pre-wired) dated September 23, 2010. ConMed Corp. has advised to cease the use of the devices immediately and conducted a recall to the user level. Direct consignees were asked to contact sub-accounts if recalled devices were further distributed and notified them of such recall. Recalled devices were to return to ConMed Corp. at 525 French Road, Utica, NY 13502 with completed Attachment II to indicate the amount returned. Customers with no recalled devices were asked to fill out Attachment II and faxed to 315-624-3225. Questions can be directed to Ms. Patricia Cotter, ConMed Recall Coordinator at 315-624-3237 or fax to 315-624-3225 or email to neorecall@conmed.com. |  | 9009000 | 9,009,000 Pediatric/Neonatal ECG Electrodes with Attached Wire | Worldwide distribution, including USA, Dubai, Ecuador, Egypt, Germany, Holland, Indonesia, Israel, Italy, Jamaica, Kenya, Korea, Lithuania, Malaysia, Mexico, New Zealand, Norway, Qatar, Saudi Arabia, Singapore, South Africa, Sweden, Switzerland, US Virgin Islands, Canada and Belgium. | Open |  |  | 9009000 | Not_Computer
Z-1379-2011 | 57566 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | XenMatrix Surgical Graft Sterile, Porcine Dermal Matrix 10 x 15 cm RectangleCatalog #: 1161015 | XenMatrix Surgical Graft | 1 | March 02, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Davol cannot guarantee that all Xenomatic product released from July to October 2010 meets FDA guidelines for endotoxin limits; and therefore, the grafts are potentially contaminated with bacteria. | PRODUCTION CONTROLS: Process Control | Davol issued recall notification to accounts by letter dated 1/06/11 requesting recalled units be returned. contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. A fax sheet attached to the letter to be completed and fax to 1-401-825-8753. |  | 1456 | 1456 units | Nationwide | Terminated | October 31, 2013 | 974 | 5227 | Not_Computer
Z-1518-2011 | 57552 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | LICOX CMP Tissue Oxygen Monitor; Catalog numbers REF AC31. Manufacturer GMS Gesellschaft fur medizinische, Sondentechnik mbH, Dorfstrasse 2, 24247 Mielkendorf, Germany; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, New Jersey 08536 USA.Used for continuous determination of oxygen partial pressure in cerebrospinal fluid and brain tissue. | LICOX CMP Tissue Oxygen Monitor | 2 | March 03, 2011 | 2011 | Integra LifeSciences Corp. | Certain LICOX CMP Tissue Oxygen Monitors could allow for underestimation of temperature values. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra Life Sciences notified their Integra Neurospecialist via email and their customers via Fed Ex overnight mail on December 17, 2010 with an Urgent: Product Recall letter. The letter identified the affected product and the reason for the recall. The letter states that an Integra Neurospecialist will contact customers to arrange for the return of their Licox monitor to the Service Center for recalibration. A loaner will be provided until the recalibration is complete. The attached "Recall Acknowledgement and Return Form" must be completed and returned to Integra. If there are any additional questions, customers should contact Gary Rabinovich at 609-936-6864. |  | 26 | 26 units | Worldwide Distribution | Terminated | November 05, 2012 | 613 | 26 | Not_Computer
Z-1522-2011 | 56660 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, ApexPro Telemetry System The ApexPro Telemetry System is intended for use under the direct supervision of a licensed healthcare practitioner. The system is designed to acquire and monitor physiological data for ambulating patients within a defined coverage area. The system processes this physiological data to detect various ECG arrhythmia events and select physiological parameter limit violations. The ApexPro Telemetry System is intended to be installed in the hospital or clinical environment in order to provide clinicians with patient physiological data, while allowing for patient mobility. These systems are typically deployed in sub acute care areas in hospitals or clinical sites where patient mobility can enhance the effectiveness of the medical procedures administered. The physiological parameters monitored include ECG, noninvasive blood pressure, non-invasive temperature and SpO2. The ApexPro Telemetry System is intended to provide ECG data via Ethernet to the computer platform for processing. The ApexPro Telemetry System is also intended to provide physiologic data over the Unity Network to clinical information systems for display. | ApexPro Telemetry System | 2 | March 03, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential for loss of patient monitoring when ApexPro Telemetry System is used in conjunction with a mirrored view CARESCAPE CIC Pro in combination monitoring mode (combo mode). The safety issue includes the potential for a delay in annunciation of audible alarms when used in combo mode, and may result in thedelay of treatment of a cardiac event. | DESIGN: Software Design | GE sent an "Urgent Medical Device Correction" letter dated December 29, 2010. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Product Correction, Safety Instruction and Contact Information.Customers were instructed to contact Technical Support or their local Service Representative with any questions regarding this medical device.For any questions regarding this recall call 262-573-4122. | Contact the recalling firm for information | 2346 | 2346 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT,NE, NJ, NM, NY, NC, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY. and the UNITED KINGDOM, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH KOREA, RUSSIA, NORWAY, NEW ZEALAND, NETHERLANDS, MALAYSIA, JAPAN, ITALY, ISRAEL, IRELAND, GERMANY, FRANCE, FINLAND, DENMARK, CUBA, CHINA, CANADA, AUSTRIA, AUSTRALIA, POLAND. | Terminated | May 29, 2012 | 453 | 2346 | Other | Alarm/Message |  | Other
Z-1521-2011 | 57847 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Vari Lase WireFiber Laser Fiber, RED 7137, Rx ONLY, Sterile EO, Vascular Solutions Inc., 6464 Sycamore Court, Minneapolis, MN 55369The VARI-LASE WireFiber is indicated for the treatment of varicose veins and varicosities associated with superficial reflux in the Great Saphenous Vein and for treatment of incompetence and reflux of superficial veins in the lower extremity. | VariLase WireFiber | 2 | March 03, 2011 | 2011 | Vascular Solutions, Inc. | Investigations of a recent Device Experience Report has made us aware of a potential problem with our Vari-Lase WireFiber, Model 7137, manufactured with the following lot numbers: 5500898 and 550896.Recently it was reported that the endovenous laser fiber cage and guidewire tip of the WireFiber became separated. It is possible that separation could occur during patient procedure and has the p | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Vascular Solutions sent an "Urgent Medical Device Recall" letter dated January 10, 2011. The letter described the issue and the product affected. Advised consignees to immediately check their inventory , place affected product in a secure location, and to complete and return the Customer Inventory Form.For questions regarding this recall call 763-656-4210. | Contact the recalling firm for information | 41 | 41 | Nationwide Distribution - USA including CT, MA, MN, NY, VA | Terminated | May 03, 2012 | 427 | 41 | Not_Computer
Z-1498-2011 | 57303 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida; System, Diagnostic Ultrasound | SSH880CV Aplio Artida; System, Diagnostic Ultrasound | 2 | March 03, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the SSH-880CV Aplio Artida Diagnostic Ultrasound System due to software issues. | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter to consignees via USPS return receipt mail. The letter contained a return reply form that is to be faxed to TAMS at 877-349-3054 as proof of notification.Customers were instructed to share the information with all users and reviewing radiologists as well as clinical engineering or biomedical group at their facility.For questions regarding this recall call 800-421-1968 or 800-521-1968. |  | 40 | 40 units | Nationwide Distribution to CA, GA, IA, IL KY, LA, MA, MN, MT, OH, OR, PA, TX, WI, and Puerto Rico. | Terminated | May 02, 2012 | 426 | 40 | Software |  | Instructions | Safety Notice/Insructions
Z-1520-2011 | 57841 | cannula, manipulator/injector, uterine | LKF |  |  | Obstetrics/Gynecology | 510(k) | VCareÂŽ Vaginal-Cervical Ahluwalia's Retractor-Elevator devices. The VCare IFU is packaged one per carton of eight VCare Pouched Devices. The Instructions for Use is an 8.5" x 12" 12-page booklet and is packed loose in the carton. Only the carton will have to be opened to replace the old IFU; there is no need to open the sterile packaging of the individual VCare pouches. The ConMed VCAREÂŽ Retractor/Elevator is indicated for manipulation of the uterus and injection of fluids or gases during laparoscopic procedures such as Laparoscopic Assisted Vaginal Hysterectomy (LAVH), Total Laparoscopic Hysterectomy (TLH), minilap, laparoscopic tubal occlusion, or diagnostic laparoscopy and also maintains pneumoperitoneum by sealing the vagina once colpotomy is performed. | VCare, Standard/Small/Large Cervical Cup | 3 | March 03, 2011 | 2011 | ConMed Corporation | Product manufactured after 4/11/2008, but before 8/01/2009, contains the earlier Instructions For Use, IFU (PN 17416, Rev. D) which does not contain the revisions on the addition of Warning Notes 10 and 11 and the expansion of Warning Note 6 to the VCareÂż Instructions for Use. The reason for the field correction is to provide a copy of the revised IFU to all customers who may have unexpired VcareÂż | DESIGN: Device Design | ConMed Corporation sent a communication to provide its direct consignees with important supplemental information concerning the Instructions for Use (IFU) for the VCAREÂŽ Vaginal-Cervical Ahluwalia's Retractor-Elevator uterine manipulators. The direct consignees were instructed to place a copy of the enclosed Instructions for Use, Part No. 14071 (Rev. A), with any inventory in your possession with an expiration date prior to August 2011 (2011-08) and replace any other Instructions for Use in thier possession. A copy of the enclosed Instructions for Use is already packaged with all devices bearing later expiration dates. Any inventory in their consignees possession with an expiration date prior to January 2011 (2011-01) is expired and should not be used. If the devices were further distributed to other facilities or distributors, or their consignees are a distributor, they were instructed to forward this information as appropriate to any known users of these devices.For questions regarding this recall call 315-624-3225. --- The IFU Warning changes are as follows: A new warning statement # 6 asks the user to verify the presence of all the components of the device upon removal of the VCAREÂŽ device from the patient. Although the prior IFU contained warnings about verifying that all components of the device were properly removed following use during surgery, the new statement expands the earlier, briefer statement. 6. Upon removing VCARE, the surgeon should visually inspect the VCARE device, and the patient, to make sure that the entire VCARE device was properly removed and that no components or fragments of these components were retained in the patient. There are 5 parts/components to the VCARE Cervical Elevator Retractor. These are: 1) the balloon; 2) the forward "cervical" cup; 3) the back or vaginal cup; 4) the locking assembly with thumb screw; 5) the metal shaft and handle with balloon inflation valve. --- A new warning statement # 10 further addresses. |  | 31028 | 31,028 VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices | Worldwide Distribution - USA including NY, TX, CA, IA, AZ, PA, IN, and KS and the countries of Austria, Jordan, Turkey, China, Portugal, Chile, Belgium, Ecuador, Brazil, Romania, Japan, Malaysia, India Switzerland, Colombia, Netherlands, Spain, Saudi Arabia, France, and Canada. | Terminated | August 30, 2012 | 546 | 31028 | Not_Computer
Z-1501-2011 | 57702 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Applied Medical's LAP CHOLE KIT 5 GK; Model Number: GK105The device is used to close a clip in order to ligate a tubular structure or vessel in laparoscopic and general surgical procedures. | Applied Medical's LAP CHOLE KIT 5 GK; Model Number: GK105 | 2 | March 03, 2011 | 2011 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of specific lot numbers of the Direct Drive Laparoscopic Clip Applier, model number CA090, due to a potential defect in the jaws, possibly resulting in jaw breakage. | DESIGN: Process Design | Applied Medical sent an URGENT: MEDICAL DEVICE RECALL letter dated September 10, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and complete the attached Voluntary Recall Notification Confirmation Form and fax it to 949-713-8832. Customers were asked to return all affected product.Secondary consignees were notified by Applied Medical on September 16, 2010. For any questions regarding this recall call 949-713-8041. | Contact the recalling firm for information | 2843 | 2,843 units for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VT, WA, and WI and the countries of Australia, Europe, Japan, and Kuwait. | Terminated | March 09, 2011 | 6 | 2843 | Not_Computer
Z-1562-2011 | 54551 | ninhydrin and l-leucyl-l-alanine (fluorimetric), phenylalanine | JNB | Phenylalanine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Quantase Neonatal Elution Buffer, a component of Quantase Neonatal Accessory Kits (#1, 2 and 5), Model number 532-5035, included in 532-5036, 532-5037 and 532-5040, Manufactured by Bio-Rad Laboratories, Hercules, CA | Quantase Neonatal Elution Buffer | 2 | March 04, 2011 | 2011 | Bio-Rad Laboratories Inc | Pressure contained within the Elution Buffer vial may expel the rubber stopper, causing the potential for injury. | OTHER/UNDETERMINED: Pending | Direct accounts were notified by phone, followed by fax (domestic). Urgent Device Recall notification letters were sent out on January 26, 2010 to subsidiary along with a Bio-Rad customer form to be send to each customer. The letters identified the affected product and explained the reason for the recall. Customers were asked to examine their inventory and quarantine the affected lots listed. In addition, they are to complete and return the enclosed response form according to the instructions provided. Customers are contact their regional Bio-Rad office for Elution Buffer replacement. | Contact the recalling firm for information | 714 | 714 | Worldwide Distribution -- USA, including states of MI and MN and Bio-Rad facilities in Italy, France, Czech Republic, Canada, Argentina, Panama, United Kingdom, Hong Kong, China, Canada, Singapore, India, South Korea, Brazil, and Sweden. | Terminated | March 17, 2011 | 13 | 714 | Not_Computer
Z-1523-2011 | 56344 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Connected to MOSAIQ, Part No. 8168754, Distributed by Siemens Medical Solutions Oncology Care System, Concord CA Medical charged-particle radiation therapy system. The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Conn | 2 | March 04, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A software problem that has the potential for data loss during the transfer of treatment records to the OIS, which may not be recorded in MOSAIQ, and subsequently may lead to mistreatment. | OTHER/UNDETERMINED: Pending | Siemens sent Urgent Medical Device Correction, Safety Advisory Notices to all affected customers on June 29, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their System Owner Manual, Chapter, "Safety Advisory Letter.For questions customers were instructed to contact their Siemens Service Engineer or Application Specialist.For questions regarding this recall call 925-602-8157. |  |  |  | Worldwide Distribution - USA including NC,OH, and the countries of Belgium, Sweden, Germany, Lebanon, Japan, Australia, and Iran. | Terminated | March 09, 2012 | 371 |  | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-1559-2011 | 57500 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Wall Mounting Plate, an accessory to the Gendex eXpert DC (REF 110-0205G1) and 765DC (REF 110-0155G1)Intraoral X-Ray Systems; Gendex Dental Systems, Des Plaines, IL USAThe wall mounting plate is used to install the Master Control unit of the X-ray system to two wooden wall studs spaced at 16", or to a wall supported by steel studs. | Wall Mounting Plate | 2 | March 04, 2011 | 2011 | Gendex Dental Systems | The X-Ray unit may break loose from the mounting plate and fall off the wall due to defects in the mounting plate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Gendex, sent an "URGENT MEDICAL DEVICE FIELD CORRECTION" letter dated December 20, 2010 to their customers. A revised notification letter dated January 20, 2011, was sent to the customers providing further clarification on instruction to correct the problem, as well as a template for end user notification. The letters described the product, problem and actions to be taken. The customers were instructed to contact their customers, inform them of the issue and replace the wall-mounting plate. The customers were also instructed to complete and return the attached acknowledgement form via fax to 215-997-5665 Attn: Regulatory Affairs or email to: regulatory.affairs@gendex.com and contact Gendex Customer Service at 1-800-323-8029.If you have any questions or concerns, contact the Regulatory Affairs Associate at 847-364-3290 or regulatory.affairs@gendex.com. |  | 34 | 34 wall plates | Worldwide distribution: USA and countries including: Australia, Canada and Germany. | Terminated | August 14, 2012 | 529 | 34 | Not_Computer
Z-1539-2011 | 57605 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cylindrical Bullet Nose Reamer 14.0 mm x 250 mm, REF 475816, Biomet Orthopedics 66 East Bell Drive Warsaw IN.The product is used to prepare the femoral canal for stem implantation. | Cylindrical Reamer | 2 | March 04, 2011 | 2011 | Biomet, Inc. | The firm was notified by their supplier that during manufacture of this instrument, the 12 degree cutting edge was not created. Testing found that if the cutting edge is missing, the reamer will not perform as intended. |  | The firm, BIOMET, sent an" URGENT MEDICAL DEVICE RECALL NOTICE" dated December 23, 2010 to its customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to immediately locate, discontinue use of product and return the product to Biomet. Biomet distributors were told to notify hospital personnel of the recall if they had distributed the product to hospital customers. The Customers/Distributors were instructed to confirm receipt of the notice by calling 800-348-9500 extension 3755 and 3983 and to complete and return the Response Form included with the recall notice via fax to 574-372-1683.Questions related to this notice should be directed to 574-371-3755 or 574-372-3983, Monday through Friday, 8am to 5pm. |  | 51 | 51 all products | Worldwide distribution: USA including states of: FL, OH, PA, IN, CA, and NJ; and countries of Argentina, Canada, Mexico, Netherlands and Spain. | Terminated | October 01, 2012 | 577 | 51 | Not_Computer
Z-1548-2011 | 57563 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK12 Defibrillator/Monitor.The LIFEPAK12 Defibrillator/Monitor can be used in AED mode, manual mode or advisory mode. In AED mode, the LIFEPAK 12 is intended for use by personnel who are authorized by a physician/medical director. In manual mode, the operator determines whether or not a defibrillation shock is needed by looking at the ECG rhythm display. Advisory mode can be useful during monitoring; CPSS (Continuous Patient Surveillance System) runs in the background and alerts the operator when a shockable rhythm is detected. | LIFEPAK 12 Defibrillator/MonitorAC Power Adapter | 2 | March 04, 2011 | 2011 | Physio Control, Inc. | L1FEPAK 12 Operating Instructions that describethe process for disconnecting the device from AC (wall) power are not always being followed. When theinstructions are not followed and the Power ON button is pushed within 2 seconds after disconnectingthe device from AC power, the service indicator will illuminate and the device may fail to charge the defibrillator and/or initiate pacing. | PRODUCTION CONTROLS: Environmental Control | Physio Control, Inc. began mailing the URGENT-MEDICAL DEVICE INFORMATION letter to their consignees on October 1, 2010. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were advised to follow the current LIFEPAK 12 Operating Instructions-Section 7.5-AC Power Adapter Operation and wait at least 2 seconds to push the Power ON button after disconnecting the device from AC power. If the instructions are not followed and the Power ON button is pushed within 2 seconds after disconnecting the device from AC power and the service indicator illuminates, it's important to be aware that turning the device off, then back on again, will reset the device and restore full functionality.The complete copy of the current LIFEPAK 12 Operating Instructions can be downloaded from the company's website at www.physio-control.com/products/ defibrillators, Select LIFEPAK 12, Select Documents.The consignees can call the firm's Technical Support Team at 1-800-442-1142, option 5, between 6:00 A.M. to 4:00 P.M. (Pacific), Monday - Friday. |  | 180 | 39, 180 units (17168 units internationally and 22012 units in US) | Worldwide Distribution - USA and the countries of ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE.Devices were distributed to the following countries:ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE. | Terminated | April 05, 2012 | 398 | 39180 | Battery | Device Operation | Advice | Safety Notice/Insructions
Z-1556-2011 | 57693 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning SystemIntended use: To plan multiple beam radiation therapy treatments. | PrecisePLAN Treatment Planning System | 2 | March 04, 2011 | 2011 | Elekta, Inc. | When using the 'Beam Edit" function to modify a bean in an unlocked Fraction Group, there are two conditions that may cause the user to unitentionally modify a locked fraction group. | OTHER/UNDETERMINED: Pending | The firm, Elekta, sent an "Important Notice A312" titled "Unintentional Modification of a locked fraction group in PrecisePLAN" dated February 5, 2008 to all affected customers. The notice described the product, problem and action to be taken. The customers were instructed to use care when using the "Make Opposed" function and when changing the Fraction Group of a beam. The customers were also advised to review the final plan and any locked fraction groups to make sure the planned beam arrangement is what was intended. Note: A Field Change Order 540 073 "All PrecisePLAN 2.15 users" was issued on September 30, 2008 to all customers.If you have any questions, call 770-670-2548 or +44(0)1293 654400. |  | 282 | 282 units | Worldwide distribution: USA and country of: Mexico. | Terminated | April 07, 2011 | 34 | 282 | Not_Computer
Z-1549-2011 | 57589 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Precedence Imaging System; 16 Slice 5/8, 16 slice 3/8;Philips Medical SystemsAn imaging system combining the acquisition of single photon nuclear medicine images and images from an x-ray computed tomography system. | Precedence | 2 | March 04, 2011 | 2011 | Philips Medical Systems | Philips Healthcare Nuclear Medicine determined that the detector arm assembly could be compromised. | OTHER/UNDETERMINED: Pending | Urgent - Medical Device Correction notifications were sent to all domestic consignees on December 28, 2010 via Federal Express. The letter identified the affected products, the reason for the recall, the hazard involved, and actions to be taken by the customer/user. Customers have been instructed to immediately stop using the system and contact a Philips service representative for an immediate inspection if one of the listed reasons apply. Customers were also instructed to follow the instructions in the system's Instructions For Use to avoid a collision or force on the detector. The firm's service will contact customers to arrange a time for Philips to conduct an inspection and to conduct the appropriate field safety correction. The letter should be placed in customers' Instructions For use. Notices to international consignees will be sent as soon as translation is complete. If customers need further information or support concerning this issue, they should contact their local Philips representative. | Contact the recalling firm for information | 540 | 540 units, all varieties | Worldwide Distribution | Terminated | April 02, 2012 | 395 | 540 | Not_Computer
Z-1555-2011 | 57688 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro 4.2, 5.0, 5.1, 6.0, 6.2A user interface of rthe SL/SLi Series Linear Accelerators that are used for radiation therapy treatments of malignant neoplastic diseases, as determined by a licensed medical practitioner. | Desktop Pro 4.2, 5.0, 5.1, 6.0, 6.2 | 2 | March 04, 2011 | 2011 | Elekta, Inc. | Under certain conditions it has been found that movement of the collimators can be induced during multi segment beams such as IMRT and Omniwedge which will not be detected by the Desktop Pro if high values are sent within the iCom tolerance table. | OTHER/UNDETERMINED: Pending | The firm, Elekta, sent an "Important Notice A302" dated August 23, 2007 titled "Update to iCom Tables for Desktop Pro when used with Third Party R&V; Systems",to all affected customers. On September 28, 2007, Elekta issued a second Important Notice A304 which reiterated the same information as was in Important Notice A302. The notices described the product, problem and actions to be taken. This notice provides a workaround on how to avoid the issue. In order to avoid this issues documented, Elekta recommended that the customers set the tolerance for collimators in the iCom tolerance table to be 3.0 degrees-no less than 3.0 and to use the Linac Record, if configured, as a means of checking the final collimator position. Elekta noted that the issue will be resolved in a future release of Desktop Pro.If you any questions related to this issue, please contact your local Elekta, third party R&V; system representative or Elekta Limited at +44 (0)1293 654200. |  | 229 | 229 units | Worldwide distribution: USA including Puerto Rico and countries including: Canada and Mexico. | Terminated | April 07, 2011 | 34 | 229 | Not_Computer
Z-1525-2011 | 57854 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TI ANCHOR 2.5MM W / 2-0 SUTURE, G-FORCE (R) SUTURE ANCHOR, Tissue Anchor, REF 86TA-2520, Rx ONLY, Sterile EO, Wright, Arlington, TN USAUsed in the repair of tendons and ligaments | TI ANCHOR 2.5MM W / 20 SUTURE, GFORCE (R) SUTURE ANCHOR, | 2 | March 04, 2011 | 2011 | Wright Medical Technology Inc | Driver may break during use. | OTHER/UNDETERMINED: Pending | Wright Medical Technology Inc initiated their recall of the product by telephone to their distributors on January 11, 2011. The firm followed with letters sent on February 1, 2011, to distributors, hospital administrators, and surgeons. Distributors were asked to immediately complete the attached response form, confirming receipt of the notice, and return a copy to Wright Medical by fax to 901-867-4788. For questions they were instructed to call 800-874-5630.For questions regarding this recall call 901-867-4788. | Contact the recalling firm for information | 45 | 45 units | Nationwide Distribution including FL, NC, OH, IL, IN | Terminated | July 11, 2011 | 129 | 55 | Not_Computer
Z-1526-2011 | 57851 | [6 empty] | Fresenius 2008T Hemodialysis MachinePart Number 190573, 2008T Hemodialysis System, OLCIDIASAFE PLUSPart Number 190713, 2008T Hemodialysis System with CDXPart Number 190766, 2008T Hemodialysis System with biBagThe Fresenius 2008T is indicated for acute and chronic dialysis therapy. | Fresenius 2008T Hemodialysis Machine | 2 | March 04, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Cursor on the dialysis screen may disappear and not respond to user input | DESIGN: Device Design | Fresenius Medical Care contacted customers by mailing a Customer Notification letter "Urgent correction FMNCA Model 2008T Hemodialysis Machine" via Certified-Mail, Return Receipt requested on January 19, 2011. The Technical Services department began contacting consignees by telephone on January 13, 2011, to schedule an appointment for installing the electronic board on the 2008T machines at the customers' facilities. For questions call Technical services 1-800-227-2572. |  | 1268 | 1268 units | Worldwide Distribution - USA and Canada | Terminated | July 08, 2013 | 857 | 1268 | Other |  |  | Other
Z-1557-2011 | 58008 | accessories, soft lens products | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AquaRefresh, Multi-Purpose Solution, 2 fl.oz/60mL, Made in the Netherlands, Distributed by Marietta Vision, 397 N. Sessions Street, Marietta, GA 30060.Contact lens solutions used to clean, rinse, disinfect, rewet and store all soft contact lenses | AquaRefresh | 3 | March 04, 2011 | 2011 | Marietta Vision | Product is misbranded; Product labeled as, Made in the Netherlands, was actually manufactured in Mexico. | OTHER/UNDETERMINED: Pending | John Patterson, President of Marietta Vision called each optical retailer on February 21, 2011, and told them of the misbranding issue in regards to country of origin. In addition, an Urgent Product Recall letter was sent to each account on February 23, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their AquaRefresh inventory for lot number KD005, and discontinue distribution of the samples immediately and to return any outstanding stock to Marietta Vision for exchange. Marietta Vision would return the misbranded product back to the Mexican factory at their expense to have the product packaging corrected. If they elect not to do so, Marietta Vision would have the product disposed of.Customers were instructed to fax the enclosed form back to Marietta Vision at 678-797-1099.For questions regarding this recall call 866-300-6257. | Contact the recalling firm for information | 502 | 500/2 fl.oz. Bottles | Nationwide Distribution including AR, AZ, CA and TX. | Terminated | November 03, 2011 | 244 | 500 | Not_Computer
Z-1561-2011 | 57974 | patch, pledget and intracardiac, petp, ptfe, polypropylene | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Pulmonary Hemi-Artery SGUsed in heart surgery | SynerGraft CryoPatch Pulmonary HemiArtery SG | 2 | March 04, 2011 | 2011 | CryoLife, Inc. | Staphylococcus aureus (not methicillin resistant) identified in companion tissue. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CryoLife Inc. issued a letter to the physician on February 17, 2011, advising him of the results of the testing performed on a companion cardiac allograft. The letter did not recommend explant of the allograft or special antimicrobial treatment of the patient unless medically necessary.The physician was asked to complete and return the enclosed, stamped, self-addressed postcard affirming the receipt of the notification.For any questions the physician was advised to call 770-419-3355. | Contact the recalling firm for information | 1 | One | Nationwide Distribution including AZ & NV | Terminated | April 24, 2012 | 417 | 1 | Not_Computer
Z-1554-2011 | 57988 | powder-free guayle rubber examination glove | OIG | Patient examination glove. | General Hospital | General Hospital | 510(k) | Medline Industries, Inc. "Cedar" brand Vinyl Economy Exam Glove, Powder-Free, Large Size, Catalog Number: MSV403P, Lot Number: JCU04958569 | Medline Industries, Inc. "Cedar" brand Vinyl Economy Exam Glove, PowderFree, Large Size | 2 | March 04, 2011 | 2011 | Medline Industries Inc | The recall was initiated because the affected lot, lot number JCU004958569, of gloves did not meet the requirements for the water-leak test. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | MEDLINE INDUSTRIES, INC. sent a RECALL NOTICE IMMEDIATE ACTION REQUIRED letter by US Mail on February 10, 2011, to all affected cusomters. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to:1. Immediately quarantine ALL inventory of the affected lot for return to Medline. 2. Complete the attached recall verification form. Please complete this form EVEN IF you have none of the affected lot in stock.3. Fax or email the completed form to Medline. The email address is Recalls@Medline.com or faxed the form to 866-767-1290. Customers with questions regarding the recall were instructed to contact Medline at 1-866-359-1704 or via e-mail at Recalls@Medline.com. | Contact the recalling firm for information | 696 | 696 cases (696,000 units) | Nationwide Distribution including CA, HI, and NV. | Terminated | March 19, 2012 | 381 | 2796 | Not_Computer
Z-1558-2011 | 57696 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra Position Sensor Unit, a component of the StealthStation S7 System, Medtronic Navigation, Louisville, CO 80027.The Polaris Spectra Position Sensor Unit (PSU) is the component of the StealthStation S7 System which detects optical markers on instruments and reference frames, determines their spatial positions, and continuously reports this information to the system computer. The computer uses this spatial information, in conjunction with information regarding the geometry of the instrument currently in use, to determine where the tip of the instrument is located on the patient anatomy. | Surgical Navigation System | 2 | March 04, 2011 | 2011 | Medtronic Navigation, Inc | Position Sensor Unit may generate tracking errors or may stop tracking immediately. | DESIGN: Component Design/Selection | The firm, Medtronic, sent a "Product Correction Notification" letter dated December 22, 2010, to all US Consignees/Customers on December 23, 2010 and foreign Consignees/Customers on December 27, 2010. The letter described the product, problem and actions to be taken. The customer were instructed to discontinue use of the impacted navigation system(s) until a replacement camera (PSU) is installed. A Medtronic Navigation Representative will visit each facility to remove the navigation camera and install a replacement. If you have any questions, contact your Medtronic Navigation Representative or Medtronic Navigation Technical Support services at 1-800-595-9709 or +1-720-890-3200. |  | 11 | 11 (US), 8 (Intnl) | Worldwide distribution: USA including states of: CA, CO, GA, SC, TN, and WI; and countries including: Australia, Canada, Chile, Russia, Saudi Arabia, and South Africa. | Terminated | March 11, 2011 | 7 | 19 | I/O
Z-0905-2011 | 58077 | generator, high-voltage, x-ray, diagnostic | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | BuckyDiagnost general purpose x-ray system with Optimus Generator.General Purpose Diagnostic Radiography. | BuckyDiagnost with Optimus Generator | 2 | March 04, 2011 | 2011 | Philips Healthcare Inc. | Affected systems were determined not to prevent run-on of a selected technique factor if the technique selector and the exposure control were activated simultaneously. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | A Customer Information BuckyDiagnost Diagnostic X-Ray System letter, dated January 6, 2011, was sent to customers. The letter identifies the affected product, the reason for the recall, actions that are to be taken by customers, and actions that are planned by the firm. Customers are advised to make sure that the desired exposure values are displayed and that they have released the technique button BEFORE they start the exposure by pressing the exposure switch whenever manually changing the exposure parameters and in order to avoid the problem. Philips will issue a Field Change Order (FCO) to resolve the issue by updating the system software free of charge. Philips will contact customers to implement the FCO. Further information or support concerning this issue can be obtained by contacting your local Philips representative. The Field Service Engineers began software implementation the week of March 21, 2011. | Contact the recalling firm for information | 1810 | 1810 units in the US | Worldwide Distribution | Open |  |  | 1810 | Software | Device Operation | Software update | Software Update
Z-1568-2011 | 56120 | lens, spectacle, non-custom (prescription) | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Polycarbonate prescription single vision lenses for use in sport goggles.To be worn by a customer in a sport goggle frame to provide corrective refraction in accordance with the customer's prescription. Per the customer's request, some of the devices may also provide protection from bright sunlight. | Polycarbonate prescription single vision lenses for use in sport goggles. | 2 | March 07, 2011 | 2011 | Luxottica Optical Manufacturing | The polycarbonate prescription lenses may not meet the frame manufacturer's minimum lens thickness specification of 2.0 millimeters. LNAD made the decision to remove these sport goggles from the market because of the potential risk that the lenses could break or disengage from the frame in the event of an impact. | OTHER/UNDETERMINED: Pending | Luxottica Optical Manufacturing sent notification letters on June 30, 2010, to their customers informing them of the problem. Customers were instructed to complete the enclosed Sport Goggle Return Form, and to indicate whether they would like a replacement pair of goggles or a refund.Place the completed Sport Goggle Return Form and their sport goggles (without the sport goggle case) in the envelope provided and return it to the firm via the United States Postal Service. Customers were instructed not to return to the retail store where purchased. If customers were unable to return their sport goggles because they were lost or discarded, customers should complete the appropriate information on the Sport Goggle Return Form, being sure to indicate whether they would like a replacement pair of goggles or a refund, and then sign the form. Place the completed Sport Goggle Return Form in the envelope provided and return it to the firm via the United States Postal Service. For any questions, call 1-877-488-2407. They instructed their customers to return the product for replacement or refund. |  | 2916 | 2,916 individual lenses (1,458 pair of sport goggles) | Worldwide Distribution - USA including the following states: AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI & WV; and the country of Canada. | Terminated | September 19, 2012 | 562 | 2916 | Not_Computer
Z-1592-2011 | 56994 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, dental implant planning system. Software Version 7.0 including updates 7.0.2 and 7.0.2.10 | Straumann coDiagnostiX | 2 | March 07, 2011 | 2011 | Straumann Manufacturing Inc. | Software: Surgical Error may occur when the implant length is changed after positioning the surgical sleeve. | DESIGN: Software Design | Straumann notified accounts of this correction by letter datedSeptember 30, 2010 and delivered via Federal Express Delivery Service. The upgrade is available on-line by logging into the coDiagnostiX system and enabling the automatic upgrade to occur. | Contact the recalling firm for information | 20 | 20 systems | Nationwide | Terminated | September 19, 2011 | 196 | 20 | Software | Device Operation | Software update | Software Update
Z-1594-2011 | 57915 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | ARCOS Modular Revision System Torque Limiting T-Handle 55 inch- pounds, REF 31-301850, Biomet Orthopedics Warsaw, IN.The torque limiting T-handle is used with attachments and hex drivers for preparation of the femur and for locking screw insertion at specified torque limit in Arcos hip procedures. | Arcos Modular Revision System Torque Limiting THandle | 2 | March 07, 2011 | 2011 | Biomet, Inc. | The firm received two complaints of the tabs on the Torque Limiting T-Handle fracturing during surgery. The firm performed internal testing and found that the tabs may fracture when used in the "OFF" or "locked" setting for hand-reaming the femoral canal. | DESIGN: Device Design | Biomet, Inc. made telephone calls to distributor accounts on January 28, 2011, with specific instructions that the Arcos Torque Limiting T-Handle was not to be used in the "OFF" position. Distributors were required to use the Arcos Torque Limiting T-Handle on the "55" setting for insertion of locking screws and troch bolt. Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated February 21, 2011, to their distributors. The letter stated that the firm was currently manufacturing replacement handles which would be distributed to the field shortly. Consignees were advised to immediately discontinue use of the T-Handle in the "OFF" position.Send back the current T-Handle as soon as they receive a replacement.Carefully follow the instructions on the enclosed "FAX Back Response Form."Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for shipment.If the product has been further distributed, consignees must notify hospital personnel of this action, via the enclosed "Dear Bioment Customer" notice. This letter must be give to hospital personnel responsible for receiving recall notices. Consignees are charged with the location and return of these products. Consignees were instructed to confirm receipt of this notice by calling, 800-348-9500, ext. 3755 or 3983. For any questions call 574-371-3755 or 574-372-3983. | No consumer action necessary | 40 | 40 units | Worldwide Distribution - USA including the states of CA, NY, FL, VA, UT, IL, MI, KY, OH, MA, PA, SD, MO, TX, NV, Puerto Rico, and the countries of Australia, New Zealand, Canada. | Terminated | July 27, 2011 | 142 | 40 | Not_Computer
Z-1572-2011 | 57732 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information SystemIntended Use : Electronic Patient Records Management | MOSAIQ Oncology Information System | 2 | March 07, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records. | DESIGN: Software Design (Manufacturing Process) | Sent an Urgent Medical Device Correction Safety Advisory Notice to customers on 1/11/11. Letters will be delivered by Siemens Customer Service Group or sent by certified mail. Letter identified the affected product and stated there is a potential risk due to a problem identified by the manufacturer. Customers are asked to refer to the attached Safety Notice from Elekta IMPAC Software. A Siemens Service Engineer will contact customers site to arrange for a patient matching criteria audit to determine if you have a problem. Customers should include the Safety Advisory Notice including the Safety Notice from Elekta IMPAC Software in their product related documentation until further notice. If customers have any questions, they should contact their responsible SIEMENS Service Engineer or Application Specialist. | Contact the recalling firm for information | 10 | 10 units | International Only Distribution -- Germany and Japan | Terminated | April 19, 2012 | 409 | 10 | Other | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-1593-2011 | 57519 | prosthesis, vascular graft, of 6mm and greater diameter | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Finesse Ultra-Thin Knitted Cardiovascular PatchIndicated for cardiac and vascular patch grafting. The fabric is recommended for use in patients requiring systemic heparinization prior to, or during, surgery. | Hemashield Finesse UltraThin Knitted Cardiovascular Patch | 2 | March 07, 2011 | 2011 | Maquet Cardiovascular, LLC | One unit of Hemashield Finesse Ultra-thin Knitted Cardiovascular Patch which did not meet the internal release criteria for water permeability was distributed. |  | The firm, Maquet, contacted their sole consignee/customer via telephone informing the customer of the recall and instructing the customer to put the unit to the side. On December 8, 2010 a sales representative visited the site to ensure the unit had been put to the side and to ensure the unit was shipped back to Maquet. If you have any questions, call the Manager, Regulator Affairs at 973-709-7487. |  | 1 | one unit | Nationwide distribution: SC only. | Terminated | September 12, 2012 | 555 | 1 | Not_Computer
Z-1591-2011 | 57432 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Anchor Tissue Retrieval System; ETO sterilized, single use disposable device; port size 15 mm, 1550 mL capacity; 3 units per box; Manufacturer: Anchor Products Co., 52 Official Rd, Addison, IL 60101 U.S.A.; Ref TRS175SBThe device is a single use, sterile disposable pouch used with a dedicated introducer for the capture and removal of tissue or organ from the body cavity during laparoscopic surgery. | Anchor Tissue Retrieval System TRS175SB | 2 | March 07, 2011 | 2011 | Anchor Products Company, Inc. | The metal component within the unit can break and cause a malfunction. Use of the device may represent a potential health hazard if the metal component is not removed after breakage. |  | The firm, Anchor, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 8, 2010 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any remaining units to their distributor for replacement; and complete and return the attached ANCHOR RECALL RETURN RESPONSE FORM via fax to (314) 961-4535 or mail to: Progressive Medical, Inc., Attn: Marketing Manager, 11085 Gravois Industrial Court, St. Louis, MO 63128, even if they have no inventory. A representative of Anchor Products then visited each of the direct accounts to help them implement their subrecall of the devices from their customers.For additional information please contact: Vice President, Anchor Products Company at 1-800-543-9124 or 1-630-543-9124; and the distributor at 1-800-969-6331 or 1-314-961-4535. | No consumer action necessary | 2256 | 2,256 units | Worldwide distribution: USA including state of: Missouri; and countries including: Australia, Belgium, Canada, France, Greece, Ireland, Japan, Spain and the United Kingdom. | Terminated | January 20, 2012 | 319 | 91041 | Not_Computer
Z-1580-2011 | 57396 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PolyAx Locking Plate System, POLYAX FEMORAL PLATE - RIGHT, 15 hole, Sterile, REF 8141-30-115, Depuy Orthopaedics, Inc., Warsaw, IN. | PolyAx Locking Plate System | 2 | March 07, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm determined that the 8.0 mm hole threads in the PolyAx Distal Femoral Locking Plate were incorrectly manufactured, which prevents the mating screw from locking in place. The screw should lock with the plate, instead the screw passes through the plate. |  | Urgent Information - Recall Notice letters, dated November 1, 2010 were sent to customers. The letters described the reason for recall and identified the product affected along with the affected units. The letters also provided clinical guidance and stated that the problem had been corrected and that DePuy was in the process of replenishing its inventory. Hospitals were to determine if any of the recalled components they received were still on hand, and return them immediately to their DePuy Representative or directly to DePuy for credit. | Contact the recalling firm for information | 246 | 246 all products | Nationwide Distribution -- ME, MA, FL, VA, CO, TX, WY, AZ, MO, WI, GA, PA, AL, NY, OH, NJ, NC, and CA. | Terminated | June 27, 2012 | 478 | 246 | Not_Computer
Z-1569-2011 | 57314 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | MONOJECT Blood Collection Tubes containing heparin13x100 sodium heparin trace elementItem Code 8881307014 | MONOJECT Blood Collection Tubes containing heparin | 2 | March 07, 2011 | 2011 | Covidien LP | The heparin sodium USP used to manufacture blood collection tubes was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the product, and return it to the firm. Distributors were instructed to notify their customers. All customers were required to respond to the notice regardless of their remaining inventory. Responses will be tracked for follow-up and effectiveness checks. For Customers Service questions regarding this recall call 1-800-962-9888 option 1. |  | 18100 | 18,100 | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, KS, KY, MA, MI, MN, MO, MS, NC, ND, NE, NJ, NM, MV, MY, OH, OK, PA, SC, SD, TN, TX, WA, and WI and the country of Canada | Open |  |  | 176100 | Not_Computer
Z-1589-2011 | 57965 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Biomet INNERVUE DISPOSABLE SCOPE KIT 1.2MM WITH CANNULA SET, PRODUCT NUMBER 922141, STERILE | InnerVue Diagnostic Scope System | 2 | March 07, 2011 | 2011 | Biomet, Inc. | The cannula could not be connected to the scope and scopes disengage during use. The outer hub of Innervue scopes may become detached from the inner hub of the scopes during assembly of the scope to the Innervue hand piece. This condition may prevent the Innervue scope from being securely attached to the Innervue hand piece. | DESIGN: Process Design | On 2/10/2011, Biomet issued an "URGENT MEDICAL DEVICE NOTICE" to their Direct Accounts instructing customers to identify and discontinue use of the recalled products and return them to Biomet. Response forms were provided. |  | 264 | 264 | Distribution USA nationwide and to the Netherlands | Terminated | August 27, 2013 | 904 | 607 | Not_Computer
Z-1566-2011 | 57969 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 100/120V System 1 Base, REF 801763, Terumo Cardiovascular Systems Corp., Ann Arbor, MI.The Terumo Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | March 07, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) has received 19 reports of a circuit breaker for the Terumo Advanced Perfusion System 1 tripping and causing the system to switch to battery back-up power. In two cases, the incident was caused when the circuit breaker came into incidental contact with the base of an IV pole; in others there was no known cause. Terumo CVS' investigation has determined that | DESIGN: Component Design/Selection | Terumo sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 14, 2011, addressed to Chief of Perfusion, Department of Cardiovascular Surgery or Director of Operationd Room Services. The letter explains the problem with the circuit breakers and hazards of the problem. The letter states that a service technician will visit the facility to install a guard. The service will be performed at the next scheduled preventative maintenance visit for those consignees with a service contract. For customers without a service contract, Terumo CVS will contact the consignees to schedule an appointment for the correction.Customers were instructed to review the Medical Device Correction notice.Assure that all users are aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions or concerns customers were instructed to contact Terumo CVS Customer Service at 1-800-521-2818. | Contact the recalling firm for information | 728 | 728 devices. | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, ME, MI, MO, MT, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and the countries ofCANADA, COLUMBIA, JAPAN, RUSSIA, EGYPT, TURKEY, PAKISTAN, KUWAIT, COSTA RICA, TAIWAN, SAUDI ARABIA, INDIA, SOUTH KOREA, JORDAN, THE PHILLIPINES, INDONESIA, HONDURAS, DOMINICAN REPUBLIC, REPUBLIC OF GEORGIA, HONG KONG, UNITED ARAB EMIRATES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, GERMANY, BELGIUM, MEXICO, THAILAND, VEITNAM, SOUTH AFRICA | Terminated | March 20, 2012 | 379 | 1113 | Hardware | Device Operation | Install a guard | Other
Z-1577-2011 | 57885 | instrument, biopsy | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needle, Soft Tissue Biopsy Needle, NAC-1825M, 18ga (1.3mm) x 25cm (10), Remington Medical, Inc. | Remington Medical, Inc. Automatic Cutting Needle | 2 | March 07, 2011 | 2011 | Remington Medical Inc. | There is a potential for the needle hubs to break during use. | DESIGN: Device Design | Remington Medical consignees were notified via facsimile, email and certified mail starting 02/07/2011. Distributors are requested to communicate this recall to their customers. If a customer does not respond to the recall notice within one week of the recall notification, the customer will be contacted by telephone. If this is unsuccessful, the customer will be sent a second Recall Notice letter and Customer Inventory Reconcile Fax by certified mail by 28 February 2011. If there is no response at that point, RMI will consider the customer's inventory depleted. All return product for this recall will follow RMIs Returns Goods Authorization procedure (SOP7-200). All product returns and customer responses will be documented. Upon completion of the recall, the status will be updated and submitted for the FDA's recall closure process. | Contact the recalling firm for information | 8161 | 8161 units | U.S. Nationwide | Terminated | August 12, 2011 | 158 | 46878 | Not_Computer
Z-1565-2011 | 57944 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | VersaCare P500 Microclimate Management Surface Upgrade Kits P3218A and P3219A Hill-rom Batesville, IN. The TotalCareÂŽ Bed Systems is intended to provide a patient support ideally suited to be used in health care environments. The TotalCareÂŽ Bed System may be used in a variety of settings including, but not limited to, acute care, including critical care, step down/progressive care, medical/surgical, high acuity sub-acute care, post anesthesia care unit (PACU), and sections of the emergency department (ED). The TotalCareÂŽ Bed System is capable of being used with a broad patient population as determined appropriate by the caregiver or institution | Total Care bed | 2 | March 07, 2011 | 2011 | Hill-Rom, Inc. | Hill-Rom has become aware of a potential issue with some TotalCareÂż and VersaCareÂż beds equipped with the integrated P500 Microclimate Management(MCM) therapy surface, and with the VersaCareÂż P500 MCM surface upgrade. Because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCMblower may experi | OTHER/UNDETERMINED: Pending | Hill-Rom sent an "IMPORTANT MEDICAL DEVICE CORRECTION" notice dated February 21, 2011, to Facility Risk Manager/Facility Administrators. The letter states that because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCM blower may experience the following: delayed start, slow running, or failure to operate.This issue does not pose a hazard to patients as long as normal patient monitoring is in effect in relation to managing skin temperature and moisture. Customers were instructed to continue to use the product as they normally would. The mattress system would provide an alarm in the event the MCM blower does not operate properly. A Hill-Rom representative would be contacting their facility to make arrangements for correcting the affected MCM blowers. If customers had any questions concerning this request or the procedure as outlined, they were instructed to contact Hill-Rom Technical Support at 800-445-3720.For questions regarding this recall call 812-934-7777. |  | 740 | 740 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, DE, FL, HI, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WI, and the countries of ARUBA, AUSTRIA, CANADA, FRANCE, GERMANY, LEBANON, PORTUGAL and SAUDI ARABIA | Terminated | August 06, 2012 | 518 | 2465 | Not_Computer
Z-1605-2011 | 57611 | device, spot-film | IXL | Spot-film device. | Radiology | Radiology | Enforcement Discretion | Philips BV Pulsera, Rel. 2.3, 718095.Mobile C-Arm X-Ray systems offering Radiographic and Fluoroscopic techniques in a wide variety of applications. | Philips BV Pulsera | 2 | March 08, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | The fixing strap of the belt holding the CBX Laser to the Image Intensifier may be detached during a procedure when used with the BV Libra and the BV Pulsera/Endura Systems. | DESIGN: Device Design | On 01/06/11, Philips sent an URGENT Device Correction Notice, dated 01/05/2011, to their consignees. The consignees were informed that the fixing strap of the Laser Aiming Device Image Intensifier may inadvertently detach when used with BV Libra, BV Endura Rel. 2, and BV Pulsera Rel 2.3 Mobile X-Ray systems. Specifically the draw latch of the fixing strap holding the CBX Laser to the Image Intensified can be accidentally released during a procedure when used with the above Mobile X-Ray systems. While waiting for Philips to implement their corrective action, the consignees are advised to fix the closed latch body of the draw latch to the metal band using a Tie Wrap. Without securing the latch with a Tie Wrap, the device should not be used.The consignees can contact Philips Call Center at 1-800-722-9377, select 5, and reference FCO 71800035 with any questions about this recall. |  | 214 | 214 units | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Washington, DC. | Terminated | February 28, 2012 | 357 | 227 | Not_Computer
Z-1606-2011 | 56427 | cancer monitoring test system, soluble mesothelin-related peptides, epithelioid/biphasic mesothelioma | OAW |  |  | Immunology | HDE - Humanitarian Device Exemption | MESOMARK Humanitarian Device. Authorized by Federal Law for use as an aid in monitoring of patients diagnosed with mesothelioma. Each kit contains 96 tests.MESOMARK is an enzyme linked immunosorbent assay (ELISA) for the quantitative measurement of Soluble mesothelin Related Peptides (SMRP) in human serum. Measurement of SMRP may aid in the monitoring of patients diagnosed with epithelioid or biphasic mesothelioma. MESOMARK values must be interpreted in conjunction with all other available clinical and laboratory data. | MESOMARK | 3 | March 08, 2011 | 2011 | Fujirebio Diagnostics, Inc. | presence of material attached to the bottom of some wells in places visible to the naked eye, which caused an overestimation of the sample. | OTHER/UNDETERMINED: Pending | Notification dated 7/7/10 was sent to the customer via e-mail and FedEx Priority Overnight.FUJIREBIO Diagnostics, Inc. sent an URGENT PRODUCT CORRECTION NOTIIFICATION letter dated July 7, 2010, to all affected customers via e-mail and FedEx Priority Overnight. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately discontinue using MESOMARK kit. Lot 57K04510 and discard any remaining inventory.Complete and return the attached Confirmation of Receipt form upon receipt of this notification, no later than July 30, 2010.Notify your ordering physician or health care provider as appropriate.For questions regarding this recall call 877-861-7246 or 610-240-3800. | Contact the recalling firm for information | 42 | 42 kits | Nationwide Distribution including UT | Terminated | July 05, 2012 | 485 | 42 | Not_Computer
Z-1596-2011 | 57892 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Power cord use on Symbiq One-Channel infuser andSymbiq Two-Channel infuser. Power cord part number: HSP3306-E07.Used on List No. 16026 - Symbiq One Channel infuser;List No. 16027 - Symbiq Two-Channel infuser | Symbiq | 2 | March 08, 2011 | 2011 | Hospira Inc | Complaints of broken, bent or missing prongs, charring, sparks, visible smoke, burnt smell have been reported on Symbiq AC Power cords. | DESIGN: Component Design/Selection | Hospira Recall notification letters were sent by Federal Express on November 1, 2010. |  | 32376 | 32,376 power cords | Nationwide distribution and to Australia, Canada, and Malaysia. | Terminated | November 07, 2013 | 975 | 32376 | Not_Computer
Z-1598-2011 | 57697 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Plum XL Micro/Macro Infusion Pump, Single Channel, with a 0.1-999 mL/hr flow rate range; list number 11846 | Plum XL Micro/Macro Infusion Pump | 2 | March 08, 2011 | 2011 | Hospira Inc. | Hospira has received reports of the Plum XL, XLM and XLD Infusion Pumps with no audible alarm at the low audio level setting. | DESIGN: Labeling Design | Hospira sent Urgent Device Recall letters dated February 14, 2011 to their customers on the same date, informing them that Hospira had received complaints of the Plum XL, XLM and XLD infusion pumps in which the audible alarm failed. If the audible alarm fails and the user does not notice the visual alert, the user may not be aware of the change in pump status such as air-in-line or occlusion. This may result in a delay or interruption of therapy which may result in serious injury and/or death. Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. They were also instructed to ensure that all personnel in their facility are following Hospira's Technical Service Manual for routine decontamination of the pumps and cleaning of the buzzers during preventative maintenance, including inspection of the buzzers for leg lifting. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance.The accounts were also requested to complete the attached reply form and return it to Hospira via fax at 1-866-382-4228. | Contact the recalling firm for information | 23385 | 23,385 pumps | Worldwide distribution, including United States, the U.S. Virgin Islands, Puerto Rico, Argentina, Australia, Bahamas, Bahrain, Barbados, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, EI Salvador, Finland, France, Greece, Guatemala, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Luxembourg, Mexico, New Zealand, Oman, Panama, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Venezuela. | Open |  |  | 205689 | Not_Computer
Z-1607-2011 | 57759 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex Free NeoGuard Transducer Cover, Ref 676-117, Sterile (2.6 x 30cm) cover, packaged 24/box. The firm name on the label is CIVCO, Kalona, IA. | CIVCO Latex Free NeoGuard Transducer Cover | 2 | March 08, 2011 | 2011 | Civco Medical Instruments Inc | Non-sterile transducer covers were incorrectly labeled as sterile and contained incorrect temperature storage information. | DESIGN: Labeling Design | The recalling firm notified their customers via phone and e-mail letter entitled Urgent - Medical Device Correction, dated 1/21/11, explaining the reason for recall. The letter stated that the firm will replace the affected product and requested that customers dispose of the product from the affected lots. A response form was enclosed to report final disposition of the product and should be returned to the firm. If there are any questions, customers can call their customer service representative at 800-445-6741 between 8 am and 5 pm CST. |  | 55 | 31/24-cover boxes | Nationwide Distribution -- WA, CA, and MO. | Terminated | July 09, 2013 | 854 | 31 | Not_Computer
Z-1471-2011 | 57767 | kit, wound dressing | MCY | Elastic bandage. | General Hospital | General & Plastic Surgery | Enforcement Discretion | Dressing Change Tray, DT5880A, Sterile, Centurion Medical Products Corp., Williamston, MI.These kits are used for dressing changes. | Dressing Change Tray | 1 | March 08, 2011 | 2011 | Centurion Medical Products Corporation | The firm is conducting a subrecall of kits containing Triad Groups alcohol prep pads. Triad initiated a recall of prep pads due to a concern about potential contamination of the products with an objectional orgainism, Bacillus cereus. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Centurion Medical Products, sent an "URGENT PRODUCT RECALL NOTICE" dated January 12, 2011 to affected customers via Certified Mail, Return Receipt Requested. The notice described the product, problem and action to be taken. The customers were instructed to identify and remove all inventory of the recalled kits in their possession; forward a copy of this notice to any customers that this product was further distributed to, and complete and return the attached Accountability Record form via fax to 517-546-3356 or email to: lcarpenter@centurionmp.com. The notice states that a Centurion representative will contact them regarding retrieval of the product from the facility.Should your require additional information, please contact the Director of Quality Assurance & Regulatory Affairs at (517) 546-5400 ext. 1135. |  | 68056 | 68056 all kits | Nationwide distribution: USA including state of: CT, ID, NM, MA, MO, MT, VA and SC. | Terminated | June 21, 2011 | 105 | 68056 | Not_Computer
Z-1595-2011 | 57819 | catheter, intravascular occluding, temporary | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | AMPLATZER Sizing Balloon II, REF 9-SB-034, Sterile EO, AGA Medical Corporation, 5050 Nathan Lane North, Plymouth, MN 55442. The AMPLATZER Sizing Balloon II is intended for use in those patients with cardiovascular defects wherein accurate measurement of the defect is important to select the appropriately sized occluder device. | Sizing Balloon II | 2 | March 08, 2011 | 2011 | AGA Medical Corporation | AGA Medical is initiating a voluntary recall of a specific lot of mislabeled AMPLATZER Sizing Balloon II. AGA Medical received 2 reports from customers who received sizing balloons that were labeled as 34 mm; however, the package contained 24 mm devices. The proper sizing (24 mm) was printed on the device hub visible through the product package. | PRODUCTION CONTROLS: Labeling Mix-Ups | AGA Medical Corporation sent an Urgent Medical Device Voluntary Recall letter, dated December 17, 2010, via e-mail to all affected customers. A hard copy was sent via Fed Ex to customers on December 20, 2010. The letter identified the product, described the issue, and the action that should be taken by the customer. Customers were requested to cease distribution and use of the devices, check inventory for unused devices, and complete and return a response form to facilitate return of the affected devices to AGA.For questions regarding this recall call 1-763-531-3175. | Contact the recalling firm for information | 54 | 54 | Worldwide Distribution - USA including AZ, CA, KY, PA, SC, TX and OUS to include: Argentina, Germany, Greece, Netherlands, Peru, Poland, Portugal, Slovakia, United Kingdom | Terminated | March 29, 2012 | 387 | 54 | Not_Computer
Z-1609-2011 | 57846 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx Defibrillator Monitor (M3535A/M3536A) with software (SW) versions F.01.00 or F.01.01 and with all of the following hardware options EtCO2, IP and Temp.Used for the termination of ventricular tachycardia and ventricular fibrillation. | Phiips Medical Systems | 2 | March 08, 2011 | 2011 | Philips Healthcare Inc. | Software: EtCO2 and Ventilation Rate values may be labeled incorrectly. | DESIGN: Software Design | Philips Healthcare representatives issued an Urgent Medical Device letter on 2/7/11. Customers are asked to follow the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notification. In this section, customers are given the following instructions: Philips software (SW) version F.01.03 (for M3535A and M3536A) and R.01.00 (for M3536A, English only) corrects this problem.If you have affected product a Philips Service Engineer will contact you to install software version F.01.03 or R.01.00. Until your software is updated your HeartStart MRx monitor/defibrillator can remain in service. If you are using the device in one of the use modes listed under problem description, make sure that operators are aware that even though the EtCO2 and ventilation rate values can display interchangeably (unexpectedly or randomly switch) under either the ventilation rate label or the EtCO2 label, the values themselves can be identified by their color (for example if the ventilation rate label is white, the ventilation value is displayed as white). Customers are to contact their local Philips representative at 1-800-722-9377, if they need further information or support regarding this issue. |  | 2382 | 2382 units | Worldwide -- USA, Canada, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA,FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KENYA, LATVIA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, REUNION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open |  |  | 2382 | Software | Output/Calculation | Software update | Software Update
Z-1610-2011 | 57936 | collector, ostomy | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative/Surgical System (57mm), 2 1/4 inchThis kit is comprised of a pouch, wafer, and related accessories (low pressure adaptor, measuring guide, insert and other packaging components); required for application to a patient immediately after ostomy surgery. | Natura Durahesive Convatec Moldable Techonology PostOperative/Surgical System (57mm) | 3 | March 08, 2011 | 2011 | Convatec Inc. | Natura Durahesive ConvaTec Moldable Technology Post Operative/Surgical System (57mm) units may contain pouches with 45 mm Natura flanges, therefore the pouch, low pressure adapter and barrier would not fit together. | TRAINING: Employee Error | Convatec sent a Medical Device Recall letter dated February 1, 2011, via UPS with return response forms attached to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop distributing and quarantine the recalled lot (0100814).Perform a count of recalled product currently in inventory. Complete the enclosed Recall Response Form and return it to the address on the response form. Return the attached Recall Response Form even if no recalled product is in inventory. Ensure that their account number is correctly identified on the attached Recall Response Form. To return the recalled product call the toll free number 1-800-582-6514 to receive an RGA number and instructions on how and where to return the recalled product. If the product has been distributed to wholesalers then customers should forward the letter to them and ask that they follow the Required Actions and return the Recall Response Form to the address listed on the form. Product must be labeled for shipment to: UPS Supply Chain Solutions1920 Outer LoopDock Door #123Louisville, KY 40219For questions regarding this recall call 908-904-2149. | Discard product or return to place of purchase for a full refund or exchange | 77 | 77 market units | Nationwide Distribution including AZ, IL, KS, MA, MI, NC, OK, OR, PA, SC, and TX. | Terminated | January 11, 2012 | 309 | 77 | Not_Computer
Z-1611-2011 | 58045 | [6 empty] | MMP 200 Articulating Booms, Catalog Number 0682-000-604SX, Product Number 0682-001-060C. Used for operating room equipment. | MMP 200 | 2 | March 08, 2011 | 2011 | Stryker Communications Corp | Utilization of incorrect screws resulted in inadequate thread engagement reducing the load bearing capability of the product. | PRODUCTION CONTROLS: Process Control | Stryker notified consignees verbally on January 28, 2011 and by letter dated February 11, 2011. | Contact the recalling firm for information | 2 | 2 | MA, CO | Terminated | March 30, 2011 | 22 | 2 | Not_Computer
Z-1608-2011 | 57211 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | ORTHOFIX Firebird Spinal Fixation System, Pedicle Screw Spinal System, Reduction Case, Part #68-0111, Lot #16518-LC18.Intended for use in conjunction with Multi-Axial Reduction Screw Assemblies. During assembly of a construct, the Anti-Splay Cap locks onto the Screw Body to prevent the Screw Body tabs from splaying (spreading apart). The Anti-Splay Cap is subsequently removed from the Screw Body after assembly of the construct.The Anti-Splay Cap is a stainless steel machined instrument. The instrument is used within the Firebird Reduction Kit (Domestic, 44-9050) and the ICON Reduction Kit (International, 68-0010). Product Code: NKB-Orthosis, Spinal Pedicle Fixation, For Degenerative Disc Disease. Class II. | Orthofix Pedicle Screw Spinal System | 2 | March 08, 2011 | 2011 | Orthofix Inc | There is a possibility that certain Anti-Splay Caps may not mate properly during assembly of the construct to the Screw Body with Multi-Axial Reduction Screw Assemblies. | OTHER/UNDETERMINED: Pending | The firm, Orthofix, sent an "URGENT: PRODUCT RECALL" letter dated October 22, 2010 to all customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to: 1) identify and remove the affected product 2) complete the enclosed Tracking and Verification Form via fax to: Orthofix Inc. at 214-937-2764, and 3) contact their Account Services Representative at 1-888-298-5700 and obtain a return authorization number to return product to Orthofix, Inc.The letter also states that Orthofix Inc., will provide replacement Anti-Splay Caps via priority delivery.If you have any questions regarding the removal and return of this product to Orthofix Inc., please contact your Account Service Representative at 1-888-298-5700. Should you have any other questions or concerns, please contact the Director of Regulatory Affairs at 214-937-2061 or by email: DarlaChew@Orthofix.com. |  | 166 | 166 Units | Nationwide distribution: USA including states of: CA, CO, FL, LA , MD, MI, MO, NJ, NY, OH, TN, VA, and WA. | Terminated | April 25, 2011 | 48 | 166 | Not_Computer
Z-1619-2011 | 56746 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLU Reagent (GLU cartridge assay) used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number 442640; SYNCHRON and UniCel Systems and Part Numbers:Synchron LX20, Part Number: 466200,Synchron LX20 PRO, Part Number: 476100,Synchron LXI 725, Part Number: 476501,UniCel DxC 600, Part Number: A10405,UniCel 600 PRO, Part Number: A11810,UniCel DxC 600i, Part Number: A27318,UniCel DxC 800, Part Number: A11816,UniCel DxC 800 PRO, Part Number: A11812,UniCel DxC 880i, Part Number: A59102,UniCel DxC 660i, Part Number: A64871,UniCel DxC 680i, Part Number: A64903,UniCel DxC860i, Part Number: A64935(1) The Synchron LX and UniCel DxC Synchron Clinical Systems are fully automated, computer controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (Sample type is chemistry dependent). (2) GLUCm reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems AQUA CAL 1 and 2, is intended for the quantitative determination of Glucose concentration in human serum, plasma, urine or cerebrospinal fluid (CSF). | GLU Reagent | 2 | March 09, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Cartridge Glucose (GLU) on UniCel DxC and Synchron LX instruments may give falsely high results for moderately hemolysed samples.The Impact:(1) Neonatal samples drawn by lancets are more likely to be hemolysed to a degree sufficient tocause interference with glucose estimation and are a special at-risk population,(2) Other hemolysed samples may be similar | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 10, 2010 with an enclosed fax-back Customer Response Form to the customers. The letter provided the customers with an explanation of the problem identified, the impact and an action to be taken. The customers were instructed to run samples using another method/technology and to complete and return the enclosed FAX BACK RESPONSE FORM via fax to (714) 961-4232 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., PO Box 8000, Brea, CA 92822.Firms with questions regarding this Product Corrective Action letter, were instructed to call our Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada http://www.beckmancoulter.com/customersupport/support/default.asp. |  | 81871 | 81,871 units | Worldwide distribution: USA and country: Canada. | Terminated | October 17, 2013 | 953 | 81871 | Not_Computer
Z-1613-2011 | 57831 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | STA-Flex SUBTALAR SCREW, 8 mm x 14 mm, REF 982362, Sterile, Biomet Sports Medicine, 56 East Bell Drive, Warsaw, IN.Used to reduce excessive rearfoot pronation by blocking forward and downward displacement of the Talus, resulting in improved rearfoot alignment and foot function. | STAFlex SUBTALAR SCREW | 2 | March 09, 2011 | 2011 | Biomet, Inc. | The firm has initiated this recall following internal fatigue testing that shows possible initiation of fatigue cracking in implants under load. |  | The firm sent URGENT MEDICAL DEVICE RECALL NOTICES dated 1/21/2011 to distributors, implanting physicians, and medical facilities. The letter to the distributors and medical facilities states the recalled product should be immediately located and returned to Biomet. In addition, the "FAX Back Response Form" should be completed and returned to the number provided. If product has been further distributed, distributors are to contact hospital personnel via the enclosed "Dear Biomet Customer" notice. The letter to the implanting physicians informs the physician of the recall and lists the patients that were implanted with the device. The physician letter quotes from the surgical technique "the implant should also be removed if any postoperative complications occur". The letter to the implanting physicians does not recommend any action be taken with the devices that are implanted. The notice just makes the surgeons aware of the recall. Distributors and Physicians are to confirm receipt of the letter by contacting 800-348-9500, ext 3755 or 3756. Also, questions related to this matter should be directed to 574-371-3755 or 574-371-3756. |  | 254 | 254 all products | Nationwide Distribution. | Terminated | November 10, 2011 | 246 | 254 | Not_Computer
Z-1621-2011 | 56637 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT G3+ Cartridge: PCO2Abbott Point of Care,Princeton, NJ 08540The test for PCO2, as part of the i-STAT System, is intended for use in the in vitro quantification of carbon dioxide partial pressure in arterial, venous or capillary whole blood. PCO2 along with pH is used to assess acid-base balance. | iSTAT G3 Cartridge | 2 | March 10, 2011 | 2011 | Abbott Point Of Care Inc. | Certain lots of G3+ cartridges may be difficult to fill. | OTHER/UNDETERMINED: Pending | The firm, Abbott Point of Care, sent a "Important notice" letter dated September 25, 2006 to all customers. The letter described the product, problem and actions to be taken. The customers were to destroy and discard the i-STAT G3+ Cartridges and complete and return the customer reply form. Abbott sent replacements cartridges to their customers.If you or any of the healthcare providers you serve have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1. | Contact the recalling firm for information | 125 | 12,675 in US, 125 ROW | Worldwide distribution: USA and country of : Canada. | Terminated | February 27, 2012 | 354 | 12800 | Not_Computer
Z-1492-2011 | 58015 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK FlexLink Plus, subcutaneous infusion set. Part Number 05511089001. An infusion set for the subcutaneous infusion of insulin administered with micro dosage insulin pumps.Made in Germany Manufactured for and Distributed byRoche Insulin Delivery Systems Inc. Fishers, IN 46037, USAToll Free: 1-800-280-7801Roche Diabetes Care AG3401 Burgdorf, Switzerland | ACCUCHEK FlexLink Plus | 1 | March 10, 2011 | 2011 | Roche Insulin Delivery Systems Inc. | There is the potential to kink or bend the cannula when inserting the ACCU-CHEK FlexLink Plus infusion set. This can lead to under delivery and elevation of blood glucose levels. | DESIGN: Device Design | Roche Insulin Delivery Systems sent urgent Product Recall letters, dated 2/21/2011. The letter identified the affected product and described the issue regarding kinked or bent cannulas. The firm requests that customers discontinue use of the ACCU-CHEK FlexLink Plus infusion set and return the unused stock. Customers are to contact their health care provider to determine if changes to their pump therapy are needed and how to temporarily continue insulin pump therapy without the ACCU-CHEK FlexLink Plus infusion set. The attached Customer Reply Card should be completed and return via mail or fax per the instructions on the card. The letter also lists some additional information pertaining to signs and symptoms. Questions regarding this recall should be directed to ACCU-CHEK Customer Care at 1-800-688-4578, 24 hours a day, 7 days a week. | No consumer action necessary | 1454918 | 1,454,918 million worldwide | Worldwide Distribution -- US and Canada. | Terminated | July 27, 2011 | 139 | 1454918 | Not_Computer
Z-1438-2011 | 55432 | dna-reagents, mycobacterium spp. | LQF | Mycobacterium tuberculosisimmunofluorescent reagents. | Microbiology | Microbiology | 510(k) | Mycobacterium Tuberculosis Complex Culture Identification Test, 2860. | AccuProbe MYCOBACTERIUM TUBERCULOSIS COMPLEX CULTURE IDENTIFICATION TEST | 1 | March 10, 2011 | 2011 | Gen-Probe Inc | This recall has been initiated following the identification of certain assay tubes that do not contain probe reagent. The affected kits may contain a combination of tubes with correct probe amount, partial probe amount and some probe tubes that are empty. Possible false negative results. | PRODUCTION CONTROLS: Process Control | Gen-Probe issued a Recall Notice to customers recommending that they discontinue the use of the identified batch (lot) numbers listed. For customers that have used these batches and obtained negative results, they recommend that the referring physician be notified about the possibility of a false negative result.With the Recall Notice, consignees received a Customer Response Form to complete. Gen-Probe requested that they segregate the affected batches listed on the Customer Response Form, then reconcile, and destroy all partial and/or full kits left in inventory, being sure to document the amount destroyed on the Customer Response Form. Gen-Probe should be updated via phone, fax, and email information. Customers should then sign and return the completed form to Gen-Probe via fax to 858-410-8250, or scanned and emailed to technicalsupport@genprobe.com.The Technical Support team should be contacted for replacement kits. Technical Support was also available for additional information at 888-484-4747 or technicalsupport@gen-probe.com. |  | 1293 | 1293 kits | Worldwide Distribution -- USA, including Puerto Rico and countries of Australia, Guatemala, Malaysia, Bahamas, Hong Kong, Mali, Singapore, Bermuda, India, Mexico, South Africa, Canada, Israel, Netherlands, South Korea, Finland, Kuwait, Panama, Thailand, France, Latvia, and Philippines. | Terminated | June 08, 2011 | 90 | 2659 | Not_Computer
Z-1625-2011 | 54384 | mixer, cement, for clinical use | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic brand KyphX HV-R Bone Cement with Kyphon mixer, a polymethyl bone cement kit consisting of a quantity of powder (20 g) and a volume of liquid (10 ml/9 g) mixed in a Kyphon mixer for up to two minutes. Model Number: CO1B; Manufactured by Elmdown LTD, London, England, for Medtronic Spine LLC, Sunnyvale, CA.Product is a medical device used during balloon kyphoplasty to cement structures in fractured vertebral bodies. | KyphX HVR Bone Cement | 2 | March 11, 2011 | 2011 | Medtronic Spine LLC, formerly Kyphon Inc | Product may not reach proper viscosity (dough state) and inconsistency or lack of homogeneity during cement preparation and during delivery period. | OTHER/UNDETERMINED: Pending | The firm, Medtronic Spine LLC, sent an "URGENT DEVICE RECALL NOTICE" dated November 9, 2009 to the customers. The notice described the product, problem and actions to be taken. The customers were instructed to return the affected lots to Medtronic Spine; to not use any of the affected product that they may have received (Medtronic will replace any affected product in their inventory), the customer should contact their Medtronic Spine sales representative or customer service at [866-959-7466 or for geography] for assistance with replacement product, and complete and return the URGENT DEVICE RECALL NOTICE: Action Required form to Medtronic Spine LLC via fax at (408) 543-6192 or mail to: Medtronic Spine LLC, Attn: Recall Coordinator, 1221 Crossman Avenue, Sunnyvale, CA 94089. If you have any questions about this notification, please do not hesitate to contact Sr. Director, Regulatory Affairs via email: pamela.n.segale@medtronic.com or phone at (408) 548-5235. |  |  |  | USA including states of: AL, AR, AZ, CA, CO, DC, FL, GA, IA, IN, KY, MA, MD, MO, MT, NC, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries including: Argentina, Austria, Belgium, Brazil, Canada, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Ivory Coast, Jordan, Kuwait, Lebanon, Luxembourg, Lybia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Caledonia, Norway, Panama, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Terminated | March 11, 2011 | 0 |  | Not_Computer
Z-1627-2011 | 57506 | system, test, carcinoembryonic antigen | DHX | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | CEA ELISA Kit.BioCheck, Foster City, CA 9440.Quantitative determination of AFP concentration in human serum. | CEA ELISA Kit | 2 | March 11, 2011 | 2011 | BioCheck Inc | Product labels indicate that the products are for IVD use and they do not have an approved 510(k). | PREMARKET APPROVAL: No Marketing Application | The firm, BioCheck, sent an "URGENT: DEVICE RECALL" letter dated December 22, 2010 to all customers. The letter described the product, problem and actions to be taken by customers. The customer were instructed to examine their inventory for the affected lots, if found, discontinue distributing the lots and promptly return via Fedex to BioCheck at 323 Vintage Park Drive, Foster City, CA 94404, Attn: RETURNED GOODS; immediately contact their accounts, if they further distributed the product, and inform them of the Recall, and complete and return the response form via email at: info@biocheck.com or fax to (650-573-1969) immediately.If you have any questions, please call (650) -573-1968 anytime during office hours (8:30am to 5:30pm Pacific time). | Contact the recalling firm for information | 25000 | estimate 25,000 kits, all varieties | Nationwide distribution. | Terminated | March 25, 2011 | 14 | 25000 | Not_Computer
Z-1623-2011 | 54468 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 4000 Vent Head Spring Assembly, a component of the CELL-DYN 4000 Automated Hematology Analyzer, List Number 07H95-01, manufactured by Abbott Diagnostics Division, Santa Clara, CAMedical device component. The CELL-DYN 4000 System is a multiparameter automated hematology analyzer designed for in vitro diagnostic use in clinical laboratories. The vent head spring assembly is used in the auto-sampling mechanism on the CELL-DYN 4000. | CELLDYN 4000 Automated Hematology Analyzer | 2 | March 11, 2011 | 2011 | Abbott Laboratories | A component used in assembly is out of specification, resulting in failure to retract, eject specimens, exposure below the safety zone, piercing of specimen tubes, or error message to display. | OTHER/UNDETERMINED: Pending | The firm, Abbott Laboratories, sent an "Urgent Field Safety Notice Product Recall" letter dated December 28, 2009 to the Customers via Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to destroy any CELL-DYN 4000 Vent Head Spring Assemblies that they have on hand and have not installed; (No action is necessary for assemblies that are currently installed on the CELL-DYN 4000 Analyzer); complete and return the Customer Reply form (even if they do not have the product) via fax to 1-800-777-0051 or email: QAGCO@abbott.comProduce Recall; indicate the number of Vent Head Spring Assemblies destroyed. the number identified will be replaced with Vent Head Spring Assemblies that meet specifications, and to please keep this communication with their CELL-DYN 4000 System Operator's Manual. If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S., please contact your local hematology customer support representative. | Contact the recalling firm for information | 157 | 157 assemblies | Worldwide distribution: USA including states of: CA, CO, FL, GA, KY, NC, NJ, NY, OR, VA, and WV; and countries including: Argentina, Australia, Canada, Germany and Japan. | Terminated | March 11, 2011 | 0 | 157 | Not_Computer
Z-1647-2011 | 56044 | anti-dna antibody (enzyme-labeled), antigen, control | LRM | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA Anti-dsDNA Test Kit, 96 Tests,Catalog No. 96DS, Distributed by and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CAThe Bio-Rad Autoimmune EIA Anti-dsDNA Test is a qualitative enzyme immunoassay (EIA) intended to screen for the presence of dsDNA antibodies in human serum as an aid in the diagnosis of systemic lupus erythematosus (SLE). For In Vitro Diagnostic Use. | EIA AntidsDNA Test Kit | 2 | March 14, 2011 | 2011 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the Anti-dsDNA test kits. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 2, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and quarantine the affected kit lots; discard all affected product in accordance with all local, state and federal disposal procedures; complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM via fax to: (510) 741-3954, attn: Bio-Rad CSD Regulatory Affairs Department, and to please re-evaluate all positive patient samples and retest if necessary. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  |  |  | Worldwide distribution: USA and countries including: Canada, Carribean, Chile, France, Germany, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and United Kingdom. | Terminated | March 15, 2011 | 1 |  | Not_Computer
Z-1651-2011 | 57534 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | MagNA Pure LC 2.0 Instrument, catalog number 05197686001, Roche Diagnostics, Indianapolis, IN MagNA Pure LC 2.0 instrument is a robotic instrument for automated isolation of nucleic acids (DNA, total RNA , total viral nucleic acids) from different kinds of biological research sample material (whole blood, serum, blood cells, culture cells, tissue, bacteria, fungi) using the specially designed MagNA Pure LC kits, for the purpose of life science research only and automated filling of different kinds of PCR reaction vessels (LightCycler Capillaries, 96 well PCR plates, PCR tube strips) with PCR reaction mixes and template nucleic acid, or other reaction tubes for making of dilution series, reaction mixes. | MagNA Pure LC 2.0 Instrument | 2 | March 14, 2011 | 2011 | Roche Diagnostics Operations, Inc. | If the bar code is printed from the Batch Results Screen, a wrong batch ID may be printed, while the correct batch ID is still assigned to the batch in the data repository of the MPLC 2.0 software. This may result in sample mis identification. If printed from the ordering screens, the bar code is printed with the correct batch ID. | DESIGN: Software Design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 04, 2010 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to print batch ID barcode labels from the ordering or stage set up screens only; print Cooling Block barcode labels from the Post Elution Edit screen only; complete and return the attached UMDC Faxback Form for MagNA Pure LC 2.0 Instrument [UMDC 10-228] via fax to 1-877-551-7152 and file this correction notice for future reference. The firm plans to issue a software update in first quarter 2011. If you have questions about the information contained in this UMDC, for Clinical customers, please contact Roche Molecular Diagnostics Technical Service at 1-800-526-1247; for Research customers, please contact Roche Applied Sciences Technical Service at 1-800-262-4911. |  | 119 | 119 devices | Worldwide distribution: AL, AZ, CA, CO, DC, FL, GA, HI, IL, KY, LA, MA, MD, ME, MN, MO, MT, NC, ND NE, NH, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, VT, WI and Puerto Rico; and country of Europe. | Terminated | August 30, 2011 | 169 | 119 | Software
Z-1649-2011 | 57914 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-ART System, Model # H-0000-0003 Intended to be used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or sterotactic radiosurgery to tumors or the targeted tissues. | HiART System | 2 | March 14, 2011 | 2011 | TomoTherapy Incorporated | It was determined that the Treatment Planning Station (TPS) can potentially under dose. | DESIGN: Software Design | Consignees were sent a TomoTherapy "Field Safety Notice Medical Device Correction" letter dated January 26, 2011. The letter was addressed to "TomoTherapy Customer". The letter described the Planning Issue, Related DQA Consideration, Graphical Presentation of Results, Product Affected, Recommended Action and Resolution. The letter suggests to use the information presented in order to avoid creating plans with an increased potential for discrepancies in planned vs. delivered dose. Specifically include measurements for small, off-axis targets that could be susceptible to the effects of this issue. DQA dose measurements should be performed at each unique full fraction dose that will be used for patient treatment, without using DQA scaling.If Customers have any questions they are to contact TomoTherapy Interaction Center by email or phone. |  | 326 | 326 | Worldwide Distribution -- US, including States of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NJ, NM, NY ,NC, ND,OH, OK OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of SPAIN, GERMANY, FRANCE, INDIA, KOREA, SWEDEN, ITALY, MALAYSIA, SWITZERLAND, CANADA, UNITED KINGDOM, SINGAPORE, TURKEY, TAIWAN, AUSTRALIA, JAPAN, BELGIUM, CHINA, MEXICO, NETHERLANDS, POLAND, SAUDI ARABIA, and UNITED ARAB EMIRATES. | Open |  |  | 326 | Software
Z-1645-2011 | 57724 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+3 Triple Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: a) list 12348 - software version 10.3; b) list 12618 - software version 11.3; c) list 20678 - Hospira MedNet softwarePlum A+3 Triple Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: a) list 12348 - software version 10.3; b) list 12618 - software version 11.3; c) list 20678 - Hospira MedNet software | Plum A3 Triple Channel Infusion Pumps | 2 | March 14, 2011 | 2011 | Hospira Inc. | Hospira has received customer reports of the Plum A+ infusion pumps with no audible alarm conditions, which have been associated with failure of the piezoelectric assembly due to component quality issues. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Hospira, sent "URGENT DEVICE RECALL" letters dated February 14, 2011 to their customers on the same date. The letter described the product, problem and action to be taken by the customers. The Customer/Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-888-216-7330. | Contact the recalling firm for information | 23020 | 23,020 pumps | Worldwide distribution: United States including the U.S. Virgin Islands and Puerto Rico, and countries including: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Vietnam. | Open |  |  | 167595 | Not_Computer
Z-1658-2011 | 54136 | enzyme immunoassay, fetal fibronectin | LKV |  |  | Clinical Chemistry | PMA | Rapid fFN for the TLi IQ system and specimen collection kit in the following configurations: 0797Specimen Collection Kit (8 pack); manufactured by Hologic, Sunnyvale, CA.Medical device for use in assessment of preterm birth risk. Enzyme Immunoassay, Fetal Fibronectin and Device for Specimen Collection (code LIO). The Rapid fFN test is approved for the indication to aid to rapidly assess the risk of preterm delivery in | Rapid fFN for the TLi IQ system (Branded as FullTerm) | 2 | March 16, 2011 | 2011 | Cytyc Prenatal Products/Hologic Inc. | Product name change requires conversion to a different type of PMA and implies that the product could be used to assess whether a patient's pregnancy is full term, or to predict imminent delivery in women who are at full term gestation. | PREMARKET APPROVAL: No Marketing Application | The firm, Hologic, sent a "Important-Product Recall/Correction" letter dated December 18, 2009 to its customers. The letter described the product, problem and the actions to be taken. The customers were instructed to 1) replace the directional inserts for the specimen collection kit and health care providers' brochure with provided instructions, 2) over-label specimen collection kit boxes with provided stickers and individual kits, 3) permanently destroy any marketing and educational materials with the "full term" name and 4) complete and return the CUSTOMER RECALL RESPONSE FORM via fax at 408-744-1905 or scan and email to recallfullterm@hologic.com.If you have any questions about this notification, please contact the Hotline (800-442-9892, Option 7# or 508-263-8510). |  |  |  | Worldwide distribution. | Terminated | March 16, 2011 | 0 |  | Not_Computer
Z-1681-2011 | 56332 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SCORPIO CR TIBIAL TRIAL, non sterile, made in the USAHowmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430.The Tibial trial are utilized during surgery to assess the overall component fit, ligament stability, and joint range of motion prior to final implantation of the insert to the tray. | SCORPIO CR TIBIAL TRIAL | 2 | March 16, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is the potential for the Scorpio T72 tibial Insert Trials to be oversized by a maximum of .024" (0.6mm). | DESIGN: Process Design | Stryker Howmedica Osteonics Corp.sent Urgent Product Correction letters Correction letters dated July 23, 2010, via Federal Express to all affected customers. were sent to all Stryker Branches on July 23, 2010 via Federal Express. Asia Pacific, Canada, Europe, Middle East, Africa, Japan and Latin America sites were sent letters on July 26, 2010 by Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the attached Product Correction Acknowledgement Form by fax as soon as possible to 201-831-6069.For any questions customers were to call 201-972-2100 or 201-831-5028. |  | 159095 | 159,095 units, all types. | Worldwide Distribution - USA including OH, TX, MS, and CO and the countries of Australia, Hong Kong, Taiwan, India, Korea, Japan, Chile, Brazil, Columbia, Venezuela, , Panama, Santo Domingo, Mexico, Argentina, Costa Rica, Guatemala, Peru,Sweden, Germany, Netherlands, France, Spain, Italy, Romania, South Africa, United Kingdom. and Canada. | Open |  |  | 159095 | Not_Computer
Z-1665-2011 | 56407 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE SystemAQUILION ONE is a multislice CT system that supports whole body scanning. The system supports simultaneous data acquisitions of 320 slices using the Selectable Slice-thickness Multi-row Detector. | TSX301A: AQUILION ONE | 2 | March 16, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has found that communication may be cut off in the Aquilion ONE/ Premium Systems, owing to the failure of a part used in the console. | DESIGN: Software Design | The firm, Toshiba America Medical Systems (TAMS), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 22, 2010, to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that a new relay will be installed to prevent the occurrence of the problem within the next 90 days and that their Toshiba Service Representative will contact them to install the part to their Aquilion ONE/Premium System. The customers were instructed to contact their Toshiba Service Representative, if the problem occurs before the corrective action has been taken; to share the information with all users and review radiologists as well as clinical engineering or Biomedical group at their facility, and to complete and return the attached CUSTOMER REPLY FORM via fax to 877-349-3054 or email to raffairs@tams.com.If you have any questions regarding this letter, please feel free to contact the Director, Regulator Affairs at (800) 421-1968 or your local Toshiba Representative at (800) 521-1968. |  | 43 | 43 units | Nationwide distribution: USA including states of: AL, AR, AZ, CA, CO, FL, GA, IA, KY, LA, MA, MD, MI, MN, MT, ND, NJ, NV, NY, OH, PA, PR, TX, VA, WI, and WY. | Terminated | April 05, 2011 | 20 | 43 | Software
Z-1675-2011 | 56415 | table, radiographic, tilting | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Toshiba DUA-450F Kalare Radiographic Tilting Table. | Toshiba DUA450F Kalare Table | 2 | March 16, 2011 | 2011 | Toshiba American Medical Systems Inc | Top table lateral movement on their DUA-450F Kalare Table cannot be performed, or the table top lateral movement is performed at an extremely low speed. As a result, positioning of the patient cannot be performed by tabletop movement. | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) issued Urgent Medical Device Correction notification dated July 8, 2010, to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. | Contact the recalling firm for information | 238 | 238 systems | US Nationwide | Terminated | April 20, 2012 | 401 | 238 | Software
Z-1674-2011 | 57199 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Bone Screw, Self Tapping, Cannulated. BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Multi-Axial Screw, Self Tapping, Cannulated. The Multi-Axial Screw is a pedicle screw which is inserted into the pedicle with the use of the Multi-Axial Screw Driver. The component is used within the Firebird Spinal Fixation System. The screw is cannulated and available modular (77-7630, pre-assembled with screw head) or non-modular (77-8630, without screw head). Product Code: NKB-Orthosis, Spinal Pedicle Fixation, For Degenerative Disc Disease. Class II.Intended for use when implanted over a guide wire for additional accuracy with reference to the pedicle. | Blackstone Medical Inc. Firebird Spinal Fixation System | 2 | March 16, 2011 | 2011 | Orthofix Inc | There is a possibility that certain Multi-Axial Screws have spherical heads measuring less than the allowable tolerance. | OTHER/UNDETERMINED: Pending | Orthofix sent an URGENT: PRODUCT RECALL letter dated October 22, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their accounts immediately and advise them of the recall situation and have them return any affected inventory.For any questions customers were instructed to contact their Account Services Representative at 1-888-298-5700.For questions regarding this recall call 214-937-2061. |  | 27 | 27 (19 of 77-7630) (8 of 77-8630) | Nationwide Distribution including AZ, NY, TX, MS, LA, FL, NC, and AR. | Terminated | May 31, 2011 | 76 | 27 | Not_Computer
Z-1679-2011 | 57513 | colorimetry, acetaminophen | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Acetaminophen Test System for use on the Roche/Hitachi, part number 03255379190, Roche Diagnostics Corporation, Indianapolis, IN. For the quantitative determinations of toxic levels of acetaminophen in human serum or plasma on Roche automated clinical analyzers. | Roche Acetaminophen Roche/Hitachi | 2 | March 16, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The current product labeling for Acetaminophen does not provide specific interference information for bilirubin, hemoglobin or lipemia. Additional testing has been performed to quantify the impact of these interferences at different concentrations and different levels of acetaminophen. Low serum concentrations of acetaminophen were included in the study. | OTHER/UNDETERMINED: Pending | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 07, 2010 to its customers. The letter described the product, problem, and action to be taken. The customers were instructed to be aware that endogenous interferents (i.e., bilirubin, hemolysis, and/or lipemia) may produce a falsely elevated value for samples containing concentrations less than 50 ug/ml of acetaminophen; complete and return the attached UMDC Faxback Form for Acetaminophen Assay via fax to 1-817-868-5343, and file this letter for future reference. The firm ask the customers to please consult with the physician or pathologist at their facility to determine specific clinical implications for their patients. The firm is also updating the product labeling to include information regarding the interference of icteric, hemolytic or lipemic samples.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have any questions about the information contained in this Urgent Medical Device Correction (UMDC). | Contact the recalling firm for information | 8668 | 8,668 units | Nationwide distribution. | Terminated | March 13, 2012 | 363 | 8668 | Not_Computer
Z-1669-2011 | 57690 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs of the Computed Tomography X-Ray Systems under the following Brands and Model #s: Brilliance 6, M/Ns 3001-3708, 30001-30131; Brilliance 10, M/Ns 4001-4067, 40001-40035 ; Brilliance 16, M/Ns 5001-6180, 50002-50139 ; Brilliance 16P, M/Ns 6007-6163, 60001-60029 ; Brilliance 40, M/Ns 9009-9272, 29001-29153, 90028-92034; Brilliance 64, M/Ns 9501-10300, 95003-95697; Brilliance Big Bore, M/Ns 7001-7520; Brilliance iCT, M/Ns 100003-100169 & Brilliance iCT SP, M/Ns 200004-200034.The Brilliance Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Vertical Brake Hubs of the Computed Tomography XRay Systems | 2 | March 16, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems has decided to recall the patient support vertical brake hub in all its Brilliance systems due to the collapse in the vertical direction. The short key connecting the vertical brake hub to the motor shaft can disengage due to the lack of thread locker on set screws. | OTHER/UNDETERMINED: Pending | Philips Medical Systems (Cleveland) Inc sent Field Safety Notice to all affected customers on December 10, 2010. The Safety Notices identified the product, the problem, and the action to be taken by the customer. Customers were instructed to monitor their units and if the vertical brake hub malfunctioned to immediately contact the recalling firm. The letter also stated that a Field Service Engineer would visit to perform a brake hub rework on the units. For further information or support customers were to contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompt(s).For questions regarding this recall call 440-483-7672. |  | 5700 | 5700 units/vertical brake hubs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO,CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT,NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV & WY, and the countries of Afghanistan; Algeria; Angola; Argentina; Australia; Austria; Azerbaijan; Bahamas; Bahrain; Bangladesh; Belgium; Bolivia; Bosnia & Herzegovina; Brazil; Brunei Darussalam; Bulgaria; Cambodia; Canada; Chile; China; Columbia; Costa Rica; Cote d'Ivoire; Croatia; Cyprus; Czech Republic; Denmark; Dominican Republic; Ecuador; Egypt; El Salvador; England; Equatorial Guinea; Finland; France; French Polynesia; Germany; Greece; Hong Kong; Hungary; Iceland; India; Indonesia; Iran; Iraq; Ireland; Israel; Italy; Jamaica; Japan; Jordan; Kazakhstan; Kenya; Kuwait; Latvia; Lebanon; Libyan Arab Jamahiriya; Lithuania; Luxembourg; Madagascar; Malaysia; Malta; Martinique; Mexico; Monaco; Mongolia; Morocco; Mozambique; Namibia; Nepal; Netherlands; Netherlands; Antilles; New Zealand; Nicaragua; Nigeria; Norway; Oman; Pakistan; Panama; Paraguay; Peru; Philippines; Poland; Portugal; Romania; Russia; Saudi Arabia; Serbia & Montenegro; Singapore; Slovakia; Slovenia; South Africa; South Korea; Spain; Sweden; Switzerland; Syrian Arab Republic; Taiwan; Tanzania; Thailand; Tunisia; Turkey; Ukraine; Unidentified KM; United Arab Emirates; United Kingdom; Uzbekistan; Venezuela; Viet am & Yemen. | Terminated | August 21, 2013 | 889 | 5700 | Not_Computer
Z-1667-2011 | 57880 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Innovate Esmark Bandages, a component of various Presource Custom Sterile Packs/Presource PBDS Modules; Packaged by Cardinal Health Medical Products and services, McGaw Park, IL 60085 U.S.A.; There is no label on the Esmark bandage in the Presource pack/module, Some are placed in Tyvek pouches and some are just loose in the pack, but neither have any labeling on them.The Esmark bandage is used to support and compress a part of the patient's body. Presource Packs/Modules are manufactured, according to customer requirements, for use during surgical procedures. | Innovate Esmark Bandages | 2 | March 16, 2011 | 2011 | Cardinal Health | Esmark bandages placed in some Presource kits/modules may not have been adequately sterilized. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 28, 2011 to all affected consignees/customers via UPS next day air. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and return the packs/modules to Cardinal Health for destruction; contact the appropriate Customer Service group listed in the letter for return instructions, and complete and return the enclosed Acknowledgment Form via fax to 847-689-9101, regardless of whether or not they have inventory. Should you have any questions, relating to this recall, please feel free to contact Cardinal Health Quality Systems, 800-292-9332. |  | 4.395 | 4.395 modules | Nationwide distribution. | Terminated | September 04, 2012 | 538 | 4.395 | Not_Computer
Z-1677-2011 | 57926 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OSTEOMATRIX bone graft substitute. OMG10. Osteomatrix is provided sterile for singlepatient use. | Osteomatrix | 2 | March 16, 2011 | 2011 | Biostructures, LLC | The OsteoMatrix Synthetic Bone Graft Composite Instructions for Use document, P4G10A/Rev. 01 had language that was not consistent with the cleared indications. | OTHER/UNDETERMINED: Pending | An email was sent on 06/15/2010 to the firm's distributors that had products with lot #0809D110 or lot #0307I208 in their possesion. The email was in regards to the distributors' OMG10 granule inventory. The distributors need to return all the recalled granule inventory to be swapped out for new inventory. | Contact the recalling firm for information | 248 | 248 units | Products were distributed to 6 consignees that were located in Michigan, Ohio, Maryland, California, Texas, and Utah. | Terminated | March 17, 2011 | 1 | 248 | Not_Computer
Z-1654-2011 | 57777 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper Systems with Velara Generator, x-ray system. FCO72200184. | Allura Xper Systems with Velara Generator | 2 | March 16, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | A wire in the connector of the Generator Grid Switch Supervisor box might short-circuit between 12V power and ground, and it will cause a System breakdown. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 10/15/2010, Philips began sending out the URGENT-Device Correction Notice Failure Causing the Generator to Stop Working notification letter to their consignees.The consignees were informed that a short circuit may occur due to a faulty wire connection in the Supervisor Box Connector in the Generator, and it will cause the generator stops working. When the system fails during a critical interventional case, this may cause risk to patients (no Fluoro available).The systems involved in this recall are Allura Xper with the Velara Generator, Allura Integris with OMCP Generator, MultiDiagnost Eleva, and OmniDiagnost Eleva Imaging system.Consignees can contact Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 70800114, 70800115, 72200184, and 72200185. |  | 1556 | 1556 units distributed in US (1452-3107=1556) | Devices were distributed through out the US. | Terminated | February 23, 2012 | 344 | 2739 | Hardware | Device Operation | Notification | Safety Notice/Insructions
Z-1676-2011 | 57932 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Remote Cardiology Review Station, a cardiovascular information system. Model Number CV7.8 | CRS Remote | 2 | March 16, 2011 | 2011 | AGFA Corp. | A discrepancy in the validation testing which resulted in the product not performing as intended. | DESIGN: Device Design | AGFA sent an "URGENT SAFETY NOTICE" letter dated 2/10/11 via FED-EX to the customer. The letter included an acknowledgment which was to be FAX-Back indicating that the information was received and understood. Agfa has discussed with the customer how the de-installation of their CRS Remote product - Software Version 7.8_HL would be performed. | Contact the recalling firm for information | 1 | One | TX | Terminated | September 15, 2011 | 183 | 1 | Other | Device Operation | Software update | Software Update
Z-1664-2011 | 57967 | compressor, cardiac, external | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | LUCAS 2 Chest Compression System. The device's Instruction for Use is labeled in part: "LUCAS 2 Chest Compression System***jolife***www.lucas-cpr.com***a product by JOLIFE***MANUFACTURER, MAIN OFFICE JOLIFE AB Scheelevagen 17 SE-223 70 LUND Sweden***www.jolife.com***".Designed to uninterrupted chest compressions at a consistent rate and depth. | LUCAS 2 Chest Compression System | 3 | March 16, 2011 | 2011 | Jolife AB | LUCAS 2 Chest Compression System may have malfunctions resulting from the failure of a hood probe on the device that connects to the battery. The device may stop in fail safe mode and the LEDs will start flickering or the device will switch off. | DESIGN: Component Design/Selection | Medical Device Correction letter, dated February 2011, was sent to customers and states that Jolife AB with the servicing support assistant of Physio-Control is conducting a voluntary upgrade of the LUCAS 2 Chest Compression System with the Serial number range of 30090016 through 30101932.Consignees were advised to keep LUCAS 2 in service and continue to use it in accordance with the Instruction for Use, Section 3.8-Safety Precautions. WARNING-MALFUNCTION. If there are interruptions, or the compressions are not sufficient or something unusual occurs during operation: Push ON/OFF for 1 second to stop LUCAS and remove the device. Start Manual chest compression.Physio-Control service representative will contact the consignees to schedule a service call to replace the affected battery connector board in the hood of the device. The new connector board will include a new probe design in which current cannot affect the spring.Consignees can contact Physio-Control Technical Support at 1-800-442-1142, #5 between 6 am-4 pm (Pacific Time), Monday-Friday or visit the website www.physio-control-notices.com/LUCAS2. If customers no longer have the affected device or the device location has changed they should contact Physio-Control Technical Support. | Contact the recalling firm for information | 1843 | 1843 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, China, Croatia, Czech Republic, Denmark, France, Germany, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Terminated | April 24, 2012 | 405 | 1843 | Other
Z-1673-2011 | 57747 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ovation Hip Stem, REF 201-0104, Offset Stem Inserter Neutral, Non-Sterile, Ortho Development, Draper, Utah 84020.The Ovation Hip System is intended for use in total hip arthroplasty. The hip stem inserter is an instrument used to seat the hip stem in the femur. There is a notch in the top of the hip stem, and the hip stem inserter has a corresponding tip at the end of the instrument. The tip of the inserter is designed to fit into the notch of the hip stem, as a guide. Once this is inserted, the surgeon will tap on the opposite end of the inserter with a hammer to seat the press-fitted stem assembly. | Ovation Hip Stem, REF 2010104, Offset Stem Inserter Neutral, NonSterile, Ortho Development | 2 | March 16, 2011 | 2011 | Ortho Development Corporation | Component of total hip arthroplasty may break during surgical procedure. | DESIGN: Process Design | Ortho Development Corporation notified Field representatives of the issue and conditions that could lead to failure by e-mail on December 14, 2010. Redesigned product is being produced for replacement.For questions regarding this recall call 801-619-3450. | Contact the recalling firm for information | 33 | 33 units | Nationwide Distribution including AZ, CA, CO, FL, MI, NV, OR, and UT. | Terminated | November 04, 2011 | 233 | 80 | Not_Computer
Z-1666-2011 | 58053 | tubing, pump, cardiopulmonary bypass | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Labeling on box reads: MAQUET BEQ-HG 0284 3/8" X 3/8" Connectors, Lot 70041098 but the product packaged inside was MAQUET BEQ-HG 1113 3/16" x 3/16" Connectors, Lot 70041096. The HG 0284 and HG 1113 connectors are packaged sterile as single components in a container holding 20 units.The connectors are designed to be used for performing connections between tubes in extracorporeal circulation during cardiopulmonary bypass procedures lasting 6 hours or less. BEQ-HG Connectors are to be used with MAQUET HLM Tubing sets. | MAQUET Bioline Connectors | 3 | March 16, 2011 | 2011 | Maquet Cardiovascular, LLC | Tubing connectors incorrectly packaged and labeled. | PRODUCTION CONTROLS: Process Control | The firm, Maquet, sent a "PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION" letter dated February 22, 2011 via Federal Express to its customer. The letter described the product, problem and the action to be taken. The customer was instructed to examine their inventory immediately to identify and remove the affected connectors and complete and return the attached RECALL- Device Corrective Action Response Form via fax to Director of Regulatory Affairs at (973) 709-7016 or email: whitney.torning@maquet.com. If you have any additional questions, please contact your local MAQUET Cardiovascular Sales Representative or customer service at 1-800-777-4222. |  | 40 | 40 BEQ-HG 0284 3/8" x 3/8" Lot 70041098 Connectors | Nationwide distribution: USA state of CA. | Terminated | September 13, 2012 | 547 | 40 | Not_Computer
Z-1701-2011 | 51955 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott CELL-DYN 3700 Hematology Analyzer CS (Closed Sampler), , and CELL-DYN 3700 SL (Sample Loader), Manufactured by Abbott Diagnostics Division, Santa Clara, CaMedical Device, hematology analyzer for in vitro diagnostic use in clinical Laboratories. | Abbott CELLDYN 3700 Hematology Analyzer CS/SL | 2 | March 17, 2011 | 2011 | Abbott Laboratories | FIRE HAZARD-- Some power supply modules may have incorrect fuses installed, which may result in fires and smoking. | DESIGN: Component Design/Selection | Abbott Laboratories sent Product Correction letters, dated March 22, 2010, to its customers. The letter listed the additional products that are now included in the recall and gave an explanation of the reason for recall. The letter also discussed the safety impact and provided a list of necessary actions for customers. Customers are to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3000 Analyzer. If the fuse is correct, then the steps provided should be followed. If the fuse is incorrect, then customers are to replace the fuse with the correct fuse and follow the instructions provided. Customers are to contact their local Customer Support Representative in the event that they do not have their replacement fuse that was provided with the Accessory Kit. Customers can refer to the Product Information Letter dated April 20, 2010 for additional labeling to assist with the correct fuse replacement. The Customer Reply form should be completed and the letter should be kept with the Operator's Manual. Questions should be directed to Customer Support at 1-877-4ABBOTT.The firm, Abbott Laboratories, sent a "Product Correction" letter dated March 27, 2009 to its customers. The letter included a Customer Reply form and CELL-DYN accessory kits to inform new customers of the necessity of replacing fuses correctly. The letter described the product, problem and actions to be taken. The customers were instructed to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3700 System and follow the instructions provided, and complete and return the Customer Reply form via fax to 1-800-777-0051 or e-mail QAGCO@abbott.com, even if they no longer have the instrument. The firm intends to service items in the field. In April 2010, additional labeling was sent to customers providing additional information when replacing fuses.. | Contact the recalling firm for information | 2961 | 2961 units | Worldwide Distribution -- USA, including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of Argentina, Australia, Bahamas, Barbados, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Curacao, Dominica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Haiti, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand Trinidad and Tobago, Turks & Caicos, Uruguay, and Venezuela. | Terminated | December 21, 2011 | 279 | 10466 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-1698-2011 | 54665 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The Waste Line Assembly and the Waste Outlet Tubing and the Waste Bottle Cable are components that are used on one or more CELL-DYN Systems. Usage: Accessories to human waste container systems | CELLDYN | 2 | March 17, 2011 | 2011 | Abbott Laboratories | The use of the device assemblies has been extended to a 6 month useful life. | OTHER/UNDETERMINED: Pending | The firm, Abbott Laboratories, sent a "Product Correction" letter dated November 30, 2009 to all consignees/customers. The letter described the product, problem and actions to be taken. The letter stated that Abbott recommended that the customers change their Waste Line Assembly and Waste Outlet Tubing every six months and included a tag in the letter to record installation and replacement dates. Additionally, the customers were instructed to keep this communication with their CELL-DYN System Operator's Manual and complete and return the Customer Reply form via fax to 1-800-777-0051 or email QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S. please contact your local hematology customer support representative. |  | 874 | 874 | Worldwide distribution: AL, AR, AZ, CA. CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, MT , NC, NE, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries including: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Dominican Republic, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Panama, Peru, Puerto Rico, Singapore, South Korea, Thailand, Trinidad and Tobago, and Uruguay. | Terminated | March 18, 2011 | 1 | 874 | Not_Computer
Z-1668-2011 | 56968 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Diluent Reagent, 10 mL, List Number: 09H48-02, Product is manufactured and distributed byAbbott Diagnostics Division, Santa Clara. CADesigned for use on the CELL-DYN Emerald System, which is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | CELLDYN Emerald Diluent Reagent | 2 | March 17, 2011 | 2011 | Abbott Laboratories | The product bar code label, when scanned, indicates that the volume is 10 mL rather than the correct volume of 10,000 mL. Analytical results are not impacted, although, there is a potential for delay in generating results. | PRODUCTION CONTROLS: Error in Labeling | Product Recall notification letters were sent on 9/30/2010. The letter identified the affected product, an explanation of the problem, patient impact, and necessary actions. Consignees are to contact their local hematology customer support representative to arrange for replacement of the CELL-DYN Emerald Diluent Lot. Once the replacement lot has been received, customers should destroy the affected lot according to their laboratory procedures. A copy of the letter should be retained for their records and the Customer Reply Form should be completed and returned via fax or e-mail. US customers should contact Customer Support at 1-877-4ABBOTT, if they have questions. |  | 322 | 322 units | Worldwide Distribution -- USA, including states of PA, NY, CA, NC, IL, TX, KS, UT, VA, FL, AR, TN, MN, NE, MI, and DE and the country of Germany. | Terminated | April 07, 2011 | 21 | 322 | Not_Computer
Z-1700-2011 | 58075 | splint, nasal | EPP | External nasal splint. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Recall of Alcohol Pads contained in the following kits distributed by Medtronic: ***REF 1529100***Thermasplint Kit, Small***X5***. ***REF 1529110***Thermasplint Kit, Medium***X5***. ***REF 1529120***Thermasplint Kit, Large***X5***. RECALLED PRODUCT - Front:***Allegiance Alcohol Prep Pad***Antiseptic, Medium, 2-ply, Saturated with 70% isopropyl alcohol, For External Use Only*** Cat. 40000-090***NDC 63517-300-10***Qty. 1 ea***Distributed by: Cardinal Health McGraw Park, IL 60085-6787 USA, Made in USA, Rev. D 12/04***Cardinal Health***. RECALLED PRODUCT - Back: ***Drug Facts***Active ingredient Isopropyl Alcohol, 70% v/v....Purpose Antiseptic***Use For preparation of skin prior to injection***Warnings For external use only. Flammable, keep away from fire or flame. Do not use - with electrocautery procedures - in the eyes. Stop use if irritation or redness develop. If the condition persists for more than 72 hours, consult a doctor. Keep out of reach of children. If swallowed, get medical help or contact a Poison Control Center right away.*** Directions Wipe injection site vigorously and discard.***Inactive ingredient purified water.*** | Nasal Splint | 2 | March 17, 2011 | 2011 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. Jacksonville, FL issued a recall on Feb 21, 2011 for all lots of the alcohol pad packaged with the Medtronic's External Nasal Splints REF 1528116 (sm), 1528126 (med), 1528136 (lg) and the Thermasplints REF 1529100 (sm), 1529110 (med), and 1529120 (lg) distributed from Jan 2008 through Aug 2010. The alcohol supplier, Cardinal Health, issued a recall on the alcohol pads on Jan |  | Medtronic Xomed, Inc. sent an URGENT PRODUCT SAFETY ADVISORY (PSA) letter dated February 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately discontinue use and dispose of the alcohol prep pad co-packaged in their External Nasal Splint or Thermasplint kits. If there has been further distribution of this item, the Urgent Product Safety Advisory letter should be forwarded to the appropriate recipient. For any questions regarding this recall call 904-279-7532. | No consumer action necessary | 8385 | 8,385 units (1,677 boxes) | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Singapore, Canada, Belgium, Croatia, Cyprus, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Latvia, Martinique, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, Taiway, Korea, Puerto Rico, and Costa Rica. | Terminated | April 03, 2012 | 383 | 34690 | Not_Computer
Z-1692-2011 | 57734 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Sabina II EE Lift, Liko, Sweeden. The Sabina and Capella sit-to-stand lifts are especially designed for people who have difficulty in standing on their own. The lifts are intended to be used with patients who are able to bear weight on their legs and actively participate in the standing exercise. | Liko Sabina II EE Lift | 2 | March 17, 2011 | 2011 | Hill-Rom, Inc. | The firm has received reports of injuries related to the Sabina Sit-in-Stand lifts. Of six injury reports, two were injuries in patints during use and four injuries were in non-patients when they tripped over the lift device and fell on the unit. | DESIGN: Device Design | The firm, Liko, issued two "URGENT FIELD SAFETY NOTICES" one dated January 27, 2011 and a follow-up notice dated February 9, 2011 to its consignees/customers. The notices described the product, problem and actions to be taken. The customers were instructed to keep the device in a separate unoccupied area when not in use; use extra care and support when lifting patients with the device; not allow unauthorized persons, especially children, to play around or operate the lift; examine the sling hooks daily to make sure they are secure, and to forward a copy of this notice to any other facility personnel they deem appropriate. The notices state the firm is investigating a design change to mitigate the risk of injuries with the device.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. | Contact the recalling firm for information |  |  | Worldwide distribution: USA including: AL, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK ,OR, PA, RI, SC, SD, TN ,TX, UT, VA, WA, WI, and WV; and country of: Canada. | Open |  |  |  | Not_Computer
Z-1043-2011 | 57763 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM ( SENO ESSENTIAL)The expected usage of this product is for the same clinical applications as traditional mammographic film/screen systems. It generates mammographic images which can be used for screening and diagnosis of breast cancer. | GE SENOGRAPHE ESSENTIAL | 2 | March 17, 2011 | 2011 | GE Healthcare, LLC | It was noticed that 3 units received in Japan were missing the rating plate on the X-ray tube cover. There is a potential that other X-ray tube covers are missing the rating plate. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter was issued to customers identifying the affected product and describing the defect along with the related hazards. The letter states that there are no actions that are required to be taken on the part of the customer. A GE Healthcare Service Representative will apply the missing rating as necessary and free of charge. Questions or concerns regarding the letter should be directed to 1-800-437-1171 in the US. |  | 112 | 112 total | Worldwide Distribution | Terminated | May 16, 2012 | 426 | 112 | Not_Computer
Z-1690-2011 | 57900 | prosthesis, hip, cement restrictor | JDK | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DePuy, a Johnson & Johnson company, CEMENT RESTRICTOR, Size 5, Sterile, DePuy Orthopaedics. This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis | DePuy CEMENT RESTRICTOR, Size 5 | 2 | March 17, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised, the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field, including the implant. Non-sterile product and/or contaminat | TRAINING: Employee Error | DePuy Orthopaedics sent an URGENT INFORMATION  RECALL NOTICE dated January 27, 2011, to DePuy Area Directors, FMDs, Distributors, and Office Managers. The notice identified the product, the problem, and the action to be taken. All territories were required to ensure that all recalled products that remain in customer accounts and field inventory are returned immediately to Warsaw to the attention of Returns. Hospitals that purchased and potentially implanted the now recalled product should be provided appropriate notification of the recall and should complete and return a signed Reconciliation Form(s) within seven business days. If the hospitals have any inventory on hand that is subject to the recall, distributors should remove it or help the hospitals return it for credit. All Reconciliation Forms should be faxed to 574-372-7567.For questions regarding this recall call 574-372-7333. | Contact the recalling firm for information | 20 | 20 | Nationwide distribution including the states of MA, PA, FL, LA, MN, WA, WI, CO, GA, NY, NC and SC and the countries of Canada, Colombia, Ireland and Argentina. | Terminated | April 23, 2013 | 768 | 140 | Not_Computer
Z-1705-2011 | 57970 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LT-D, SIZE 3, QUANTITY 1 EACH, Distributed by North American Rescue, Greer, South Carolina, 26950-4791, Products with a Mission, Manufactured by King Systems Corporation, ÂŽ15011 Herriman Boulevard, Noblesville, IN 46060, 317-776-6823, www.kingsystems.comSize 3: Small adult disposable oropharyngeal airway. This Oropharyngeal Airway is intended for airway management by providing a patent airway to allow patient ventilation. The sterile lubricant is used in Assembly Kits of the KLTD and the KLTSD products. | King Systems KLTD Oropharyngeal Airway | 2 | March 17, 2011 | 2011 | King Systems Corp. | King Systems has been notified by Triad Group, our supplier of sterile lubricating jelly, that a recall has been initiated for this product because the product may be "inadequately sterilized" and use of this product may result in "patient infection". Because this particular sterile lubricating jelly is packaged in the KING LT-D Kit and KING LTS-D Kit, King Systems is sending a Notice of the Reca | PRODUCTION CONTROLS: Process Control | The firm, King Systems, sent an "URGENT PRODUCT RECALL" letter dated January 14, 2011 to its customers. King Systems sent a follow-up "URGENT PRODUCT RECALL" letter dated March 3, 2011 which "SUPERSEDES" the January 14, 2011 recall letter with attached instructions to its customers. The letters described the product, problem and actions to be taken by the customers. The customers were instructed to notify all their purchasers of this recall; immediately cease sales and use of the affected products; identify all inventory on hand or returned to them by their customers; REMOVE Triad Group lubricant from King LT-D and LTS-D kits according to King instructions; quarantine the Triad lubricant, re-seal the LT-D and LTS-D kits and destroy Triad lubricant or return to Triad according to King instructions, and complete, sign and return Recall Acknowledgement Form via fax to: 317-776-6827 or email to: triadluberecall@kingsystems.com. King Systems will provide replacement lubricant packs to all customers. Customers with questions regarding product return, replacement, hospital notification or concerning the information requested above, please contact Customer Service at 1-800-642-5464. If you have any regulatory questions about the recall, please contact Director of Quality Assurance and Regulatory Affairs at 1-800-642-5464. | Contact the recalling firm for information | 577426 | 577,426 airway kits total under recall | Worldwide distribution: USA including: CA,CO, FL, ID, IN, KS, NC, NH, NY, OH, PA, PR, SC, TN, and TX; and Country of: Canada. | Terminated | February 20, 2013 | 706 | 577426 | Not_Computer
Z-1704-2011 | 55617 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Stryker Stiletto Electrosurgical Probe (E-Probe); Model number: 0270-070-550, 0270-070-551, 0270-070-552, 0270-070-553, 0270-070-554, 0270-070-555; Manufactured by Stryker Endoscopy Puerto Rico. | Stryker Stiletto Electrosurgical Probe (EProbe) | 2 | March 17, 2011 | 2011 | Stryker Endoscopy | The heat shrink, PTFE E-Probe Tip were out of specification (OOS) in the longitudinal change (+5% -20%). | DESIGN: Device Design | All Sales reps were notified of the recall on February 25, 2010. Urgent Device Removal advisory notices were also sent to all consignees.Device Removal letters identified the specific part numbers and lot numbers involved. Letters also list the potential risk to patients as a result of the issue. Customers are asked to inspect their inventory for part and lot numbers and quarantine any affected product. If recalled product is found, customers are to contact Stryker Customer Service at 800-624-4422 and request an RMS to send back recalled units. New units will be provided at no charge. The acknowledgement form is to be completed and returned. Questions or concerns should be directed to 408-754-2000 or e-mail at eprobe@stryker.com. |  | 45 | 31, 014 units | Worldwide Distribution -- US, Australia, Canada, Germany, Greece, Iberia, Korea, Switzerland, and Newbury. | Terminated | March 18, 2011 | 1 | 31014 | Not_Computer
Z-1685-2011 | 56399 | radioimmunoassay, estrone | CGF | Estrone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estrone Radioimmunoassay (RIA) Kit, DSL-8700.Intended for in vitro diagnostic use. | Estrone Radioimmunoassay (RIA) Kit | 2 | March 17, 2011 | 2011 | Beckman Coulter Inc. | Beckman Coulter has confirmed that eight Estrone RIA kit (Ref DSL 8700) lots are inaccurately calibrated to the firm's internal reference calibrator. Lots 992637, 100115A, 100205A, 100205RA, and 100226A have now expired and should not be in active use any more. | OTHER/UNDETERMINED: Pending | Beckman Coulter released an Urgent: Product Corrective Action letter dated July 9, 2010 identifying the affected product and informing consignees of the reason for recall, asking them to take the following actions:*For lots 992637, 100115A, 100205A, 100205RA, and 100226A, review historical external quality control results relevant to the patients' results obtained with these lots to ensure the assay performance was within laboratory specifications.*For lots 100319A, 100409A, 100409RA, review patents' results obtained with the results of external quality control and interpret them in the light of additional applicable tests, patient's clinical status and clinical history of patient.Beckman Coulter has finalized the process of re-adjusting Estrone calibrators' values assignment to the new internal reference calibrator starting with lot 100403A.If consignees need products for additional tests they were told to follow replacement instructions below:*In the United States, please contact Customer Service at 1-800-526-3821*Outside the United States, contact your local Beckman Coulter Representative.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They were to also complete and return the enclosed response form within 10 days so the firm can be assured that consignees received the notification.Consignees were instructed to forward any other questions regarding the notification to the DSL Technical Support Center at 1-800-231-7970 in the United States, or contact their local Beckman Coulter Representative. | Contact the recalling firm for information | 917 | 917 units | Worldwide Distribution -- US, including states of CA, FL, GA, IL, IN, KS, MI, NC, NY, OH, OR, TX, VA, WA, and WI and country of Canada. | Terminated | May 04, 2012 | 414 | 917 | Not_Computer
Z-1733-2011 | 55146 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, IL The CELL-DYN Emerald is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | CELLDYN Emerald Hematology Analyzer System | 2 | March 18, 2011 | 2011 | Abbott Laboratories | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance. | DESIGN: Software Design (Manufacturing Process) | The firm, Abbott, sent two "Product Correction" letters one dated March 23,2010 and a follow-up letter dated July 1, 2010, via express mail to all its consignees/customers. The letters described the product, problem and actions to be taken. The customers were instructed to perform the following steps, noted in the letter, to print and verify their gain setting. Abbott informed the customers that they would receive a mandatory software update and were instructed to install the CELL-DYN Emerald according to the attached instructions, and to complete and return the Customer Reply form via fax to 1-800-777-0051 or email to QAGCO@abbott.com.Should you have any questions or would like to discuss in more detail, contact Director Quality Assurance and Regulatory Affairs Abbott Hematology at 408-567-3403. |  | 891 | 891 units | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, and WV; and countries including: Australia, Brazil, Canada, Columbia, Dominican Republic, Germany, Puerto Rico, Singapore, South Korea, St. Eustalius & St. Martin, and Turks & Caicos. | Terminated | March 18, 2011 | 0 | 891 | Other |  | Software update | Software Update
Z-1722-2011 | 55875 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e. Product part numbers: 3202487-xxx, 3202488-000, U3202487-000, 3202488-015, and U3202488-000.The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e are acute cardiac care devices used by authorized healthcare providers in hospital and clinic settings. It is commonly found on "code" or "crash" cars in hospital hallways, plugged into AC power. In an emergency the user retrieves the cart, unplugging the monitor, and uses the device with the door closed for easy 1-2-3 operation. When the code team arrives the user will open the door and have access to all functions, including ECG monitoring. In addition to functioning as a semi-automatic defibrillator the LIFEPAK 20/20e provides options for: ECG monitoring, manual mode defibrillation; Non-invasive pacing; synchronized cardioversion; pulse oximetry (SpO2); hard paddles (adult and pediatric) and internal paddle accessories; printing; and backup batter power. The units are biphasic waveform devices | defibrillator/monitor | 2 | March 18, 2011 | 2011 | Physio Control, Inc. | A component failure on the AC power supply assembly results in "NO AC POWER" operation.A failure of AC power can result in a delay of defibrillation therapy if backup battery power is depleted. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | PhysioControl, Inc. began sending an Urgent Medical Device Correction letter via certified receipt to all affected customers describing the potential issue. Customers were advised to keep the defibrillators in service and to follow the daily Operator's Checklist. All power supplies will be updated.For questions customers were instructed to contact Technical Support at 1-800-442-1142.For questions regarding this recall call 425-867-4000. |  | 53550 | 53550 units | Worldwide distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MI, MD, MI, NC, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY and the coutnries of ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Open |  |  | 53550 | Battery | Treatment/Delivery/Therapy | Instructions on cheklists/Power supply update | Safety Notice/Insructions
Z-1727-2011 | 56962 | filler, recombinant human bone morphogenetic protein, collagen scaffold, osteoinduction - hde | OJZ |  |  | Orthopedic | HDE - Humanitarian Device Exemption | Stryker Biotech OP-1 Putty (OP-1, BMP-7, bone morphogenetic protein)"Catalog numbers: 300-50, 300-55OP-1 Putty is authorized by Federal law for use as an alternative to autograft in compromised patients requiring revision posterolateral (intertransverse) lumbar spinal fusion, for whom autologous bone and bone marrow harvest are not feasible or are not expected to promote fusion. Examples of compromising factors include osteoporosis, smoking and diabetes. | Stryker Biotech | 2 | March 18, 2011 | 2011 | Stryker Biotech | A change to the Preparation for Use sectionwas identified where 2.5cc was replaced with 2cc-3cc volume and was not approved | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, Stryker Biotech, sent a "Medical Device Correction" letter on October 15, 2010 via return receipt letter to all customers. The letter contains labeling with the approved Preparation for Use Section of 2.5 cc volume and included the revised copy of the Package Insert. Additionally, the letter described the product, problem and actions to be taken. The customers were instructed to follow the new instructions provided and to keep the new instructions in mind when using the OP-1 Putty.If you have any questions concerning this communication, please contact 508-416-5326 or Stryker Biotech's Pharmacovigilance department at PVG@stryker.com. | Contact the recalling firm for information | 2861 | 2861 units | Nationwide distribution. | Terminated | March 20, 2012 | 368 | 2861 | Not_Computer
Z-1732-2011 | 57615 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute Care Troponin I TestPak, Troponin I AssayThe Stratus(R) CS Acute CareTM Troponin I method (CTNI) is an in vitro diagnostic test for the measurement of cardiac Troponin I in heparinized plasma. Cardiac Troponin I measurements can be used as an aid in the diagnosis of acute myocardial infarction (AMI). Cardiac Troponin I can also be used as an aid in the risk stratification of patients with acute coronary syndromes (ACS) with respect to their relative risk of mortality. | Stratus(R) CS Acute Care Troponin I TestPak, Troponin I Assay | 2 | March 18, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Low frequency of non-repeatable falsely elevated CTNI results without an associated error message. | OTHER/UNDETERMINED: Pending | The firm, SIEMENS, sent an "Urgent Field Safety Notice" letter dated December 2010 to all Stratus(R) CS Instrument customers. The letter described the product, problem and actions to be taken. The customers were instructed to repeat testing of samples with CTNI results above 0.07 ng/mL; discuss content of this letter with their laboratory director regarding the need to review test results reported from CTNI TestPak lots 230200002 or greater; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to Technical Solutions Center at 302-631-8467, and forward this notification to anyone to whom they may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center at 800-405-6473. |  | 11635 | 11635 | Worldwide distribution: AL, AK, AZ, AR, CA, CO, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY.Product was distributed to foreign accounts in Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh,Belgium, Bosnia & Herzegovina, Canada, Canary Islands, Czech Republic, Cyprus, Croatia, Denmark, Finland, France, Germany, Great Britain, Greece, Guadeloupe, Indonesia, India, Israel, Ireland, Italy, Japan, Libya, Malaysia, Morocco, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Arab Emirates. | Terminated | March 19, 2013 | 732 | 11635 | Not_Computer
Z-1724-2011 | 57674 | toothbrush, powered | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | W Vital Health Pulsating Clean Rechargeable Toothbrush; Distributed by: Walgreen Co., 200 Wilmot Rd., Deerfield, IL 60015-4616, Made in China; The toothbrush was sold with a single head and with three heads: a) Rechargeable Power Toothbrush - Model Number 565222, UPC 0 49022 35396 1; b) Rechargeable Power Toothbrush with 3 Replacement Heads: Model Number WIC 472641, UPC 0 49022 49612 5;Intended to remove adherent plaque and food debris from the teeth to reduce tooth decay. | Walgreens Vital Health Rechargeable Toothbrush | 2 | March 18, 2011 | 2011 | Brushpoint Innovations, Inc. | The power toothbrush may make a loud noise and the battery cap becomes separated from the main body of the toothbrush. The battery cap can become a projectile that could cause potential injury if it comes in contact with the body. | DESIGN: Device Design | The firm, BrushPoint Innovations, sent a email dated January 7, 2011 notifying their customers of the recall. The email described the product, problem and actions to be taken. The customers were instructed to pull and quarantine the toothbrushes due to potential product quality concerns, and hold their stocks until January 14, 2011, call in the number of toothbrushes quarantined, and then return them to the Walgreen distribution centers. A stop sale was placed on the affected item numbers in their cash register scanners.If you have any questions, call 905-833-5122 or 905-944-1000. |  | 8100 | 8,100 units | Nationwide distribution. | Terminated | December 16, 2012 | 639 | 8100 | Not_Computer
Z-1731-2011 | 57757 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter MAXM Hematology Analyzer and MAXM Hematology Analyzer with Autoloader.Hematology Analyzers; Part Number: 6604569, 6604687, 6705995, 6705996, 6705997, 6705998, 6705999, and 6706000.Quantitative, automated, differential cell counters for in vitro diagnostic use. | Coulter MAXM Hematology Analyzer and MAXM Hematology Analyzer with Autoloader | 2 | March 18, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the Coulter MAXM, MAXM AL, HmX and HmX AL Analyzers omit the tilde (~) character when thetilde is used as part of the Sample 10 within a barcode label scanned by the primary mode barcode reader. There is a potential for Specimen or Patient misidentification to occur. |  | The recall communication was initiated on 12/14/2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTER HmX Hematology Analyzer, COULTER HmX Hematology Analyzer with Autoloader, COULTER MAXM Analyzer, COULTER MAXM Hematology Analyzer with Autoloader. The letter provides the customers with an explanation of the problem identified and an action to be taken.Customers were instructed not to use the tilde (~) character in Specimen or Patient identifiers. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 4288 | 4288 units total | Worldwide Distribution -- US, Algeria, Andorra, Antigua and Barbuda, Australia, Austria, Bahrain, Bangladesh, Bermuda, Bosnia and Herzegovina, Botswana, Brunei Darussalam, Bulgaria, Burkina Faso, Burundi, Canada, Cayman Islands, China, Congo, Cote d'Ivoire, Croatia, Czech Republic, Djibouti, Egypt, Finland, France, French Guiana, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Italy, Japan, Jordan, Republic of Korea, Kuwait, Latvia, Lebanon, Libyan Arab, Jamahiriya,Lithuania, Macao, Malawi, Mexico, Mongolia, Morocco, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Oman, Pakistan, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia,Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sudan, Swaziland, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, U.S. Virgin Islands, Yemen, and Zambia. | Open |  |  | 4288 | Not_Computer
Z-1726-2011 | 57778 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFA Software; Model/Catalog Number: 426-0246; The Bio-Rad PhD System is an EIA and IFA processing system that consists of a PhD fluid handling station(s) linked to a network computer via an Ethernet hub. The network computer provides worklist generation, data management, data reduction and microplate reader control functions. Product is manufactured by Bio-Rad Laboratories, Inc., Espace European De L Enterprise Strasbourg Schiltigheim, FRANCE | PhD System EIA/IFA Software | 2 | March 18, 2011 | 2011 | Bio-Rad Laboratories, Inc. | Firm's Correction involves advising customers who have implemented a Laboratory Information System (LIS) that they need to ensure quantitative results have passed all validation criteria prior to reporting patient results. When a patient result fails a Validation Rule, the qualitative result will indicate " INVALID". | OTHER/UNDETERMINED: Pending | The firm, Bio-Rad, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 20, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the PhD Worklist Results to verify that all validations have passed and that the quantitative results are valid; to translate (if necessary) the customer letter, add local contacts information, distribute to all active PhD System customers in their region and provide tracking documentation to CSD-RA for all customers who received the customer letter.If you have any questions, contact the Bio-Rad CSD Regulatory Affairs Department at (510) 741-4618. | No consumer action necessary | 650 | 650 units | Worldwide distribution: USA and countries including:Canada, China, Australia, New Zealand, Czech Republic, France, Hong Kong, India, Italy, Korea, Singapore, Sweden, Taiwan and United Kingdom. | Terminated | August 17, 2011 | 152 | 650 | Not_Computer
Z-1041-2011 | 57784 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM (SENO ESSENTIAL)The Senographe Essential system is intended to be used in the same clinical applications as traditional mammographic film/screen systems. It generates digital mammographic images which can be used for screening and diagnosis of breast cancer. | GE SENOGRAPHE ESSENTIAL | 2 | March 18, 2011 | 2011 | GE Healthcare, LLC | During installation it was discovered that the duplicate of the 21 CFR 1010 Certification label for the X-ray Generator Cabinet was missing from the installation kits that allow installation of the Senographe Essential in a van (mobile unit). | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | The firm plans to repair the defect or to bring product into compliance. The field action is being implemented to inspect the affected units and if found to be missing the certification label, a certification label will be applied. This will be carried out via a Field Modification Instruction planned to be released in March, 2011 with an estimated completion date of July, 2011. |  | 14 | 14 | Nationwide Distribution | Terminated | May 16, 2012 | 425 | 14 | Not_Computer
Z-1737-2011 | 56639 | test, system, immunoassay, lipoprotein-associated phospholipase a2 | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit, Enzyme Immunoassay for Quantitative Determination of Lp-PLA2 in human plasma or serum, part number 90123, manufactured by diaDexus Inc, South San Francisco. | PLAC Test ELISA Kit | 2 | March 21, 2011 | 2011 | DiaDexus, Inc | Product fails stability tests prior to expiration date. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A phone script has been prepared to contact customers. Customers were informed of the affected product and its reason for the recall. If customers didn't use the product after May 11, 2010, the action is considered complete. If the product was used after May 11, 2010, a follow up letter will be sent to the customer and they will be asked to re-test any samples beyond May 11, 2010 with an alternate lot of reagent. |  | 482 | 482 units | Worldwide Distribution -- US, including states of UT, OH, CA, NJ, NC, FL, OK, MA, NY, TX, NJ, WA, and MI and countries of Germany and Israel. | Terminated | March 21, 2011 | 0 | 482 | Not_Computer
Z-1736-2011 | 57443 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Total BhCG (BhCG) Assay Part Number 33500The Access Total BhCG assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of total BhCG levels in human serum and plasma using the Access Immunoassay Systems. | Access Total BhCG (BhCG) Assay | 2 | March 21, 2011 | 2011 | Beckman Coulter Inc. | Beckman Coulter has confirmed that non-reproducible falsely elevated test results may occur when using the Access Total BhCG assay. These elevated results are often attributed to pre-analytical factors and are particularly noted at the low end of the analytical measuring range.The role of preanalytical factors in laboratory testing has been described in a variety of published literature. |  | Beckman Coulter sent an Urgent Product Corrective Action letter dated November 11, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:." Evaluate current laboratory protocols to ensure the necessary procedures are in place for identifying and mitigating falsely elevated test results." Follow the sample handling instructions provided in the Access Total ÂŽhCG Instruction for Use (IFU)." Review the enclosed pre-analytical sample handling information bulletins:1. The Role of Preanalytical Factors in Chemistry and Immunoassay Testing (P/N A34880A).2. Sample Handling Makes a Difference (P/N A28226A)." At the discretion of the laboratory director, any total ÂŽhCG result that is questioned due to lack of clinical correlation should be repeated.NOTE: Due to the clinical utility of the Access Total ÂŽhCG assay as a biomarker of pregnancy, Beckman Coulter feels it is unnecessary to require a review of all previously obtained Access Total ÂŽhCG test results. The Access Total ÂŽhCG assay is intended as a marker for pregnancy at the time of testing. A discrepant or questionable result would have been brought to the attention of the laboratory. The pregnancy status of a patient at this point, post testing, would be known to the patient and the physician." Please share this information with your laboratory staff and your phlebotomy director and/or phlebotomy staff and retain this notification as part of your laboratory Quality System documentation." Complete and return the enclosed response form within ten days so they can be assured that this important notification has been received.If you have any questions regarding this notification, please contact the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call1-800-854-3633 in the United States and Canada. Outside the United States and Canada . |  | 158919 | 158,919 units | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Antigua and Barbuda, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Congo, Croatia, Czech Republic, Djibouti, Egypt, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Moldova, Monaco, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Zambia, Zimbabwe. | Open |  |  | 158919 | Not_Computer
Z-1738-2011 | 57809 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thrombelastograph Hemostasis Analyzer System; a multipurpose system for in vitro coagulation studies; Haemonetics, Haemoscope Division, Niles, IL 60714; list number 07-022a multipurpose system for in vitro coagulation studies | TEG 5000 | 2 | March 21, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemonetics has received a report of smoke from the unit with a power supply failure. | OTHER/UNDETERMINED: Pending | Haemoscope sent letters to all of their U.S. accounts on January 24, 2011, via certified mail. The letters listed guidelines regarding proper use of the TEG 5000 System Power Supply to eliminate any potential risks such as electrical shorting or power supply damage. The information, which was printed in the letter, will also be available in an updated TEG User Manual soon to be released. A Caution label with installation instructions was included with the letter, to be applied to the back of the TEG unit. The accounts were requested to complete the enclosed acknowledgement form and fax it back to Haemoscope at 847-588-0455. Any questions were directed to 1-800-438-2834. |  | 1155 | 1,155 units | Nationwide Distribution including CA, FL, MD, PA, OH, SC, TN, TX, and VA | Terminated | March 12, 2012 | 357 | 1155 | Battery | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-1735-2011 | 57232 | microtome, cryostat | IDP | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific HM 560 Cold Disinfection Cryostat, Microm International GmbH, Waldorf, Germany Tissue processing equipment used to prepare human tissue specimens for diagnostic histological examination by processing specimens through the various stages of decalcifying, infiltrating, sectioning, and mounting on microscope slides. | Cryostat Microtome, Thermo Scientific Microm HM 550 Cold Disinfection Cryostat | 2 | March 21, 2011 | 2011 | Thermo Fisher Scientific | Through internal quality review, firm determined the need to conduct a more comprehensive system validation to support the claim of complete bacterial, fungal and viral disinfection efficacy. | DESIGN: Device Design | ThermoFisher sent Advisory Notice letters beginning on September 28, 2010, to all affected customers. The letters stated that the firm needed to conduct more comprehensive system validation data on the Cold-Disinfection sytem option available on the HM550-D cryostat to support the claim of complete bacterial, fungal and viral disinfection efficacy. The letter recommended that the Cold Disinfection option on the HM550-D not be used as the sole disinfection method on the HM550-D until the validation work is completed. The firm provided addtional safety protocols pertaining to the cryostat disinfection. Customers were instructed to complete and fax the Confirmation of Receipt Form to Technical Support, COLD-D Advisory Notice at (269) 372-2436. | Contact the recalling firm for information | 100 | 100 units | Worldwide Distribution - USA including the states of LA, OH, FL, MD, PA, TX, MO, WI, KS, AL, CA, VA, MS, IL, NE, MA, and the countries of Australia, Belgium, China, Czech Republic, Denmark, Estonia, Finnland, France, Germany, Hong Kong, India, Italy, Japan, Netherlands, Norway, Poland, Portugal, Russia, Spain, Sweden, Taiwan, and the UK. | Terminated | April 02, 2012 | 378 | 100 | Not_Computer
Z-1739-2011 | 57366 | enzyme immunoassay, gentamicin | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Gentamicin (GENT) Flex(R) reagent cartridge, K4012This method is an in vitro diagnostic test for the quantitative measurement of gentamicin, an aminoglycoside antibiotic, in human serum and plasma on the Dimension Vista System. Gentamicin measurements may be used in the diagnosis and treatment of gentamicin overdose and in monitoring levels of gentamicin to ensure appropriate therapy. | Dimension Vista(R) Gentamicin (GENT) Flex(R) reagent cartridge, K4012 | 2 | March 21, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | False elevation of Dimension Vista(R) Gentamicin results for patients and QC when prior reagent aspirated by the reagent probe is any of the following plasma protein methods: ASL, B2MIC, B2MU, CRP,CYSC, HCYS, hcCRP, IGE, IGG3, IGG4, RF, and STFR resulting in inappropriate patient treatment. | OTHER/UNDETERMINED: Pending | Siemens Healthcare Diagnostics, Inc. sent an Urgent Field Safety Notice letter dated November, 2010, to all affected customers. Customers were requested to run GENT on a Dimension Vista(R) system that does not run the listed plasma protein methods. Instructions were provided for disabling methods on the instrument to prevent inadvertently testing Gent along with plasma protein methods.Customers were instructed to :Forward the notification to anyone to whom they may have distributed the product. Complete and return the Field Correction Effectiveness Check Form to Siemens Healthcare Diagnostics at (302) 631-8467.For further information contact Siemens Technical Solutions Center at 1-800-441-9250 or their local Siemens representative. |  |  |  | Nationwide Distribution including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | March 19, 2013 | 729 |  | Not_Computer
Z-1368-2011 | 57106 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura XPER FD 10 and Field Extension for XPER Cardiovascular System and Allura XPER FD 20 and Allura CV20.Angiographic X-ray systems with a solid state X-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | Philips Healthcare Allura XPER FD10 | 2 | March 22, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | Adjustment of the micro-switches in the footswitch used on the Philips cardiovascular X-ray systems was not implemented. If it is programmed for exposure, it is possible that after pressing the pedal on the footswitch it does not release itself. The patient may be exposed to X-ray longer than intended. |  | On 11/5/10, Philips began sending the Electronic Product Radiation Warning notification letter, dated November 4, 2010, to their consignees.The consignees were informed that the middle pedal of the foot switch of the Allura FD10, FD20, CV20 systems, and the Allura Field Extension series does not not always become inactive (sticks) when pedal is released. The function of the middle pedal on the foot switch is freely programmable by the customer. If it is programmed for exposure and if the pedal does not release by itself, the patient will be exposed to X-ray longer than intended.The consignees were advised to check if their footswitch is possibly affected. Footswitch identification 452270000141 is noted. If the pedal does not release by itself press the red emergency stop button on the system.Consignees can call Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 72200189. | Contact the recalling firm for information | 4 | 4 units | Worldwide Distribution -- US, including states of CA, IL, MA, NC, and OC. | Terminated | July 21, 2011 | 121 | 11 | Not_Computer
Z-1743-2011 | 56397 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Chemistry RF Rheumatoid Factor Part Number 475136RF reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems CAL 5 Plus, is intended for quantitative determination of Rheumatoid Factor concentration in human serum or plasma. | SYNCHRON System(s) Rheumatoid Factor (RF) Reagent | 2 | March 22, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter testing indicated that Rheumatoid Factor (RF) reagent lot M907325 will not remain stable to the labeled expiration date. | OTHER/UNDETERMINED: Pending | The firm Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated June 25, 2010 to the consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions:*Discontinue use and discard appropriately Rheumatoid Factor (RF) reagent lot M907325*Contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discuss reagent replacement options.*Complete and return the FAX BACK RESPONSE FORM via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., Brea, California 92821, Mail Code E2.SE.08Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If consignees/customers forwarded the affected products listed above to another laboratory, then they need to also provide a copy of the Product Corrective Action letter to the laboratory.If they have any questions regarding this Product Corrective Action letter, consignees were told to call Beckman Coulter Customer Support Center (Hotline) at (800) 854-633 in the United States and Canada. | No consumer action necessary | 58 | 58 units | Worldwide distribution: USA including states of: AZ, CA, CO, IA, PA, NH, NV, NY, and WA; and countries of: Belgium, China, Czech Republic, France, Germany, Hong Kong, India, Italy, Japan, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Russian Federation, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | July 02, 2012 | 468 | 58 | Not_Computer
Z-1622-2011 | 57947 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F250NRe Hemodialyzer, Single Use, Catalog Number: 0500325E, Advanced Fresenius Polysulfone.Intended for use as an artificial kidney system to the treatment of patients with renal failure, fluid overload, or toxemic conditions by performing such therapies as hemodialysis, hemofiltration, hemoconcentration and hemodiafiltration. | Optiflux F250NRe Hemodialyzer | 2 | March 22, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | There is an increased risk of an internal blood leak. | PRODUCTION CONTROLS: Process Control | FMCNA notified consignees on 2/16/11 by telephone and follow-up with an Urgent FMCNA Optiflux F250NRe Recall letter via Certified Mail, return Receipt requested. The firm states in the letter that there is an increased risk of an internal blood leak affecting a small percentage of the two lots of dialyzers. They are asking customers to examine their inventory, to discontinue use, and segregate the affected product in a secure place. Customers are to contact their FMCNA Customer Service Team at 1-800-323-5188 for instructions on how to return product. Used dialyzers should not be returned. If there are any questions, Customer Service should be contacted. |  | 38364 | 38,364 units | Worldwide Distribution -- USA and Canada. | Terminated | February 12, 2013 | 693 | 38364 | Not_Computer
Z-1643-2011 | 57916 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | AFFIXUS Lag Screw Drill, REF 2112-01-303, Non-Sterile, DePuy Orthopaedics, Inc., Warsaw, IN; DePuy International Leeds, England. | AFFIXUS Lag Screw Drill | 2 | March 22, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm recieved a complaint that the cannula for the drill is not the same size at both ends. The drill is grabbing the guidewire and advancing and not allowing to finish reaming, resulting in the possibility of the guide wire progressing through and penetrating the acetabulum. | OTHER/UNDETERMINED: Pending | On 1/12/2011 DePuy contacted their distributors by telephone and directed the distributors to cancel all pending surgeries and discontinue use of the sets until the investigation is completed. On 1/14/2011 DePuy again contacted all distributors by telephone and directed all distributors to return the instruments to DePuy for inspection as part of the investigation. | Contact the recalling firm for information | 296 | 296 units | Worldwide Distribution -- US, including the states of AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, NC, NV, NY, OH, PA, TN, TX, VA, WA, WI; and country of Switzerland. | Terminated | June 27, 2012 | 463 | 296 | Not_Computer
Z-1744-2011 | 57937 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | ETI-AB-AUK PLUS kits,REF P001931, DiaSorin S.p.A., 13040 Saluggia (VC), Italy, Distributed by: DiaSorin Inc., 1951 Northwestern Avenue, P.O. Box 285, Stillwater, MN 55082-0285In vitro enzyme immunoassay (EIA) intended for use in the qualitative detection of antibodies to hepatitis B surface antigen (anti-HBs) in human serum or plasma. | ETIABAUK PLUS | 2 | March 22, 2011 | 2011 | Diasorin Inc. | Through internal testing, DiaSorin has become aware that the kit lots identified (ETI-AB-AUK PLUS kits, catalog number P001931, lot numbers 0650270A and/or 0650270A/1) may not consistently exhibit stability to defined expiration dating, resulting in decreased binding and lower absorbance. Internal testing has confirmed that when defined assay validation criteria pass, analytical performance is co | OTHER/UNDETERMINED: Pending | The firm, DiaSorin, sent a "Customer Notification Letter" dated 19 August, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to destroy all remaining inventory and complete and return the receipt of notification and destruction of product form, found in the bottom portion of the letter, via fax to: (651) 351-5669, ATTN: TECHNICAL SERVICES or mail to: ATTN: TECHNICAL SERVICES, 1951 Northwestern Ave., Stillwater, Minnesota 55082-0285.If you have any questions or concerns, please contact DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. | Contact the recalling firm for information | 710 | 710 | Worldwide distribution: USA including states of: AL, AR, CA, CT, DE, DC, FL, GA, IL, IA, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK OR, PA, TX, VT, VA, WA,and WI; and country of: SOUTH KOREA. | Terminated | August 16, 2012 | 513 | 710 | Not_Computer
Z-1723-2011 | 57960 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R4.5Radiation therapy treatment of malignant neoplastic diseases, as determined by a licensed medical practitioner. | Elekta Synergy XVI R4.5 | 2 | March 22, 2011 | 2011 | Elekta, Inc. | If another patient is selected on the imaging system (XVI) release 4.5) during the transmission of end of treatment data from Desktop Pro R7.01 to R&V; system, the XVI information is prioritized and the end of treatment data is never received by the R&V; system. | OTHER/UNDETERMINED: Pending | Important Notice A341 titled "Mandatory Upgrade Desktop Pro R7.01 Service Pack and Cancelled Notices" dated January 26, 2011, is in distribution to all affected customers. Field Change Order 20000501011, "Service Pack 2 upgrade to Desktop Pro R7.01 Service Pack 1", dated October 25, 2010 will be utilized by Elekta Service personnel to upgrade the Desktop Pro to include Service Pack 2. The Important Notice letter instructed customers to file the letter in the Important Notice section of the appropriate User Manual. Before doing so, users are to ensure that Desktop Pro R7.01 SP2 is installed on their treatment control system. If it is not installed, customers are to contact their local Elekta representative. Once the installation of Service Pack 2 has taken place, the notices listed in Table 1 of the letter will no longer be applicable and should be removed from the User Manual. Customers should direct their questions to their local Elekta representative. | Contact the recalling firm for information | 13 | 13 units | Class 2 Recall -- Nationwide Distribution -- Including CA, MI, MO, ND, OH, OR, RI, VT, and WI. | Terminated | May 24, 2011 | 63 | 13 | Software | Device Operation | Software update | Software Update
Z-1745-2011 | 52904 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Siemens Medical Solutions USA, Inc., ACUSON AcuNav ultrasound catheters when using SwiftLink catheter connectors; Model Number: 10348560. | ACUSON AcuNav ultrasound catheters | 2 | March 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Screen Image Quality. An assembly issue on SwiftLink connectors may cause reduced/degraded image quality when using ACUSON AcuNav ultrasound catheter transducers. | PRODUCTION CONTROLS: Process Control | Siemens Field correction was initiated on 3/20/2009. A device correction letter was distributed to all affected users, with a description of the problem and its corrective actions. Siemens service personnel will make a site visit. |  | 21 | 21 units | Distribution in US and Internationally. | Terminated | April 01, 2011 | 9 | 21 | I/O
Z-1746-2011 | 57824 | dressing, wound and burn, interactive | MGR |  |  | General & Plastic Surgery |  | Organogenesis Apligraf Interactive burn and wound dressing, supplied as a living, bi-layered skin substitute | Organogenesis Apligraf | 2 | March 23, 2011 | 2011 | Organogenesis, Inc. | Product sterility compromised | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Organogenesis notified accounts on Sunday, January 30, 2011, via fax and phone. The Company will collect written confirmation from the affected customers that they have been notified of the recall. Sales representatives for the affected customers have been directed to visit the customers as soon as possible Monday morning to assist in the recall and prevent additional applications of the product. Apligraf Medical Affairs personnel will provide clinical support to treating physicians, and update them as new information on microbial contamination and antibiotic susceptibility becomes available. Units not applied have been requested to be returned to Organogenesis. |  | 50 | 50 units | nationwide. | Terminated | July 28, 2011 | 127 | 50 | Not_Computer
Z-1748-2011 | 58137 | system, imaging, gastrointestinal, wireless, capsule | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Given Imaging Limited, PillCam Express Delivery Device, For use with PillCam SB, endoscopic imaging device. Part numbers FGS-0367 (3-Pak) or FGS-0366 (1-Pak). | PillCam Express Delivery Device | 2 | March 23, 2011 | 2011 | Given Imaging Ltd. | The capsule holder can break away from the catheter during endoscopic procedures, necessitating retrieval. | DESIGN: Device Design | Given Imaging, Inc. consignees were notified via FedEx letter on/about 03/08/2011. They instructed to return all affected product. A Customer Response form with return instructions was included to be completed and returned to Given Imaging. |  | 389 | 389 | U.S. Nationwide, Canada, Germany Israel and Australia. | Terminated | December 02, 2011 | 254 | 389 | Not_Computer
Z-1747-2011 | 58138 | dressing, wound and burn, interactive | MGR |  |  | General & Plastic Surgery |  | Organogenesis Apligraf - Interactive burn and wound dressing, supplied as a living, bilayered skin substitute | Organogenesis Apligraf | 2 | March 23, 2011 | 2011 | Organogenesis, Inc. | Product sterility may be compromised | PRODUCTION CONTROLS: Packaging Process Control | Organogenesis contacted the accounts by telephone and fax on 2/24/11 to advise of the contamination and identification fo the organism for appropriate patient follow-up. |  | 7 | 7 units | Nationwide. | Terminated | July 28, 2011 | 127 | 7 | Not_Computer
Z-1813-2011 | 57177 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | ASR UNI FEMORAL IMPL SIZE 57, DEPUY ASR XL ACETABULAR CUP SYSTEM. The DePuy ASR Acetabular Cup System is comprised of a one-piece metal acetabular cup, a unipolar femoral head, and a taper sleeve adapter. The acetabular component is designed as a cobalt-chrome molybdenum (CoCrMo) alloy one-piece cup with Porocoat porous coating and is available in outer diameter sizes 44mm through 62mm in two-millimeter increments. The outer surface of the cup has a porous coating with the addition of a hydroxyapatite (H-A) coating. There are no separate liner components to this system, as the liners are integral to the one-piece acetabular cups.Prescription Use, Sterile.The device is part of a system for use in total hip replacement in which the acetabular component articulates with a femoral component. The DePuy ASR XL Acetabular Cup System is compatible with DePuy ASRT femoral components. The device is indicated for use in the following conditions, where there is evidence of sufficient sound bone to seat and support the components: 1. A severely painful and/or disabled joint from osteoarthritis, traumatic arthritis, rheumatoid arthritis, or congenital hip dysplasia. 2. Avascular necrosis of the femoral head. 3. Acute traumatic fracture of the femoral head and neck. 4. Failed previous hip surgery, including internal fixation, arthrodesis and hemiarthroplasty. 5. Certain cases of ankylosis. Porous-coated DePuy ASR MXL Acetabular Cups are indicated for cementless application. | DEPUY ASR XL ACETABULAR CUP SYSTEM (UNI FEMORAL) | 2 | March 24, 2011 | 2011 | Depuy Orthopaedics, Inc. | As part of our ongoing post-market surveillance of all products, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports.DePuy Orthopaedics issued a Field Safety Notice in March 2010 after receiving new data from the | DESIGN: Device Design | DePuy Orthopaedics issued a Field Safety Notice in March 2010 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the Field Safety Notice with their organization and any organization where the ASR products may have been transferred. Customer were instructed not to implant the ASR devices. Patients who received the ASR System should be informed of this recall and instructed to return for a follow up visit.Patients with radiographic changes indicative of product failure should be addressed according to normal procedures.All other patients should be followed according to the April 22, 2010, and May 25, 2010 UK Medicines and Healthcare products Regulatory Agency (MHRA) Device Alerts. Follow up for all patients implanted with ASR acetabular cups at least annually for five years postoperatively. Beyond five years, follow up in accordance with locally agreed protocols.For patients who are symptomatic or implanted with a cup angle greater than 45 degrees, particulary where a small component has been implanted.:" Consider measuring cobalt and chromium ion levels in whle blood and/or performing cross sectional imaging including MRI or ultrasound scans." If metal ion levels in whole blood are elevated above 7 parts per billion (ppb) for either metal ion, a second test should be performed three months after the first in order to identify patients who require closer surveillance, which may include cross sectional imaging" If MRI or ultrasound scan reveals soft tissue reactions, fluid collections or tissue masses, then revision surgery should be considered.DePuy representatives were to assist with returns of any remaining inventory.For questions regarding this recall call 574-372-7333. |  |  |  | Worldwide | Terminated | August 29, 2013 | 889 |  | Not_Computer
Z-1818-2011 | 57946 | immunochemical, thyroglobulin autoantibody | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Access Thyroglobulin Antibody II, Part Number: A32898The Access Thyroglobulin Antibody II (TgAb) assay intended use is for the quantitative determination of thyroglobulin antibody levels in human serum and plasma using the Access Immunoassay Systems. | Access Thyroglobulin Antibody II, Part Number: A32898 | 2 | March 24, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports of reagent pack to pack variability for the Access Thyroglobulin Antibody II (TgAb) reagent lot numbers identified in the recall. Erroneous results for patient and quality control samples can be generated. The ThyroglobulinAntibody II (TgAbII) reagent lots may produce:(1) Erroneously high TgAb results for both quality co |  | Beckman Coulter sent a Product Corrective Action (PCA) letter on January 26, 2011 via US Postal Service to the affected customers. The letter explained the problem identified and the action to be taken.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers identified in the tableabove;(2) If customers obtained questionable thyroglobulin antibody results, or received feedback from clinicians on reported thyroglobulin antibody results using the reagent lot numbers identified above, rerun or request a redraw of the patient sample;(3) Share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they should provide a copy of the letter to them;(4) Complete and return the enclosed response form within 10 days so the firm can be assured that you have received this important notification and are taking appropriate action.For questions regarding the recall communication, customers were told to contact Beckman Coulter Customer SupportCenter:(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter representative. | Contact the recalling firm for information | 3658 | 3658 units (1800 in US) | Worldwide Distribution - USA including AR, CA, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, MS, NC, NJ, NM, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the countries of Argentina, Armenia, Australia, Belgium, Brazil, Bulgaria, Canada, China, Czech Republic, France, French Polynesia, Georgia, Germany, Greece, Hungary, Israel, Italy, Kazakhstan, Kuwait, Lebanon, Morocco, Netherlands, New Zealand, Oman, Poland, Puerto Rico, Romania, Russian Federation, Serbia, South Africa, Spain, Switzerland, Taiwan, Tunisia, Turkey, Ukraine, United Arab Emirates, Yemen | Terminated | May 04, 2012 | 407 | 3658 | Not_Computer
Z-1819-2011 | 57962 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation (EBW) a component of the Computed Tomography X-Ray System Model Numbers EBW (CT) 728260 & EBW (Portal) 728269.The Extended Brilliance WorkplaceÂŽÂŽ (EBW) is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. The EBW is also provided as a component of PET/CT and SPECT/CT systems, performing the same function as indicated above (also referred to as the NM EBW). | Extended Brilliance Workstation (EBW) | 2 | March 24, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | During an FDA inspection on 10/1/2010 the firm was notified by FDA that their field action they initiated in Oct 2009, constituted a recall. The firm determined that the functional analysis feature in the Comprehensive Cardiac Analysis (CCA) was displaying incorrect values on the polar fields. The calculated results for the myocardial wall thickening was incorrect in that the numbers displayed f | OTHER/UNDETERMINED: Pending | Philips sent an URGENT - Field Safety Notice letter dated September 29, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware that the numbers displayed for "Wall Thickening" actually represent a reduction in wall thickness from ES to ED.Philips will be providing a sotware update to resolve the issue. For further information or support customers were instructed to contact their local Philips representative:Customer Care Service Center in the U.S. at 1-800-722-9377, option 5:Diagnostic Imaging option 1: CT or their local Philips representative. |  | 2068 | 2,068 units | Worldwide Distribution - USA including the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX , UT, VA, VT, WA, WI, WV & WY and the following countries: ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BANGLADESH, BELGIUM, BRAZIL, BRUSSEL, CANADA, CHILE, CHINA, COLUMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EDUADOR, EGYPT, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LATVIA, LEBANON, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, MIRATES, UNITED KINGDOM, VENEZUELA & YEMEN. | Terminated | September 21, 2012 | 547 | 2068 | Software | Output/Calculation | Software update | Software Update
Z-1725-2011 | 58123 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Instruments, Inc. Millar Sensors Systems Solutions Mikro-Tip Angiographic Catheters Sterile and Nonpyrogenic Model SPC-454D and SPC454F. Provide intravascular pressure readings and capability for the deliverance of high-speed injection of radiopaque contrast media used in vascular imaging to selected sites in the vascular system. | Millar disposable angiographic catheter (Left Heart) | 1 | March 24, 2011 | 2011 | Millar Instruments, Inc | Possibility of epoxy debris particles inside the catheter lumen. | PRODUCTION CONTROLS: Process Control | The firm, Millar, notified their consignees/customers by email and sent hardcopy of email on March 8, 2011. The notification described the product, problem and actions to be taken. The customers were instructed to return the catheter(s) to Millar or report back to Millar if the catheter(s) has been consumed.If you have any questions, please do not hesitate to contact Director, RA/QA at 832-667-7000. | Contact the recalling firm for information | 1080 | 1080 units | MN, NY, Washington DC, NC, MA, MN, PA, CA, MI, Australia, Germany, Canada, Japan, United Kingdom, and Norway. | Terminated | June 27, 2011 | 95 | 1080 | Not_Computer
Z-1820-2011 | 57564 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 S/N 9089, 9551, 9548.The Brilliance 64 is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Brilliance 64 Computed Tomography XRay System Model 728231, | 2 | March 24, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | During the version 2.6 software upgrade, the system was not recalibrated with the infant phantom for the infant protocol. | OTHER/UNDETERMINED: Pending | Philips sent an URGENT - Medical Device Correction letter dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers may consider reviewing any infant images that may have been affected during the Field Test to confirm reading accuracy. For further information or support concerning this issue, customers could contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 3 | 3 | Worldwide Distribution - USA including CO, IN, and TN and the countries of Norway, Israel, and Germany | Terminated | February 22, 2013 | 701 | 3 | Software | Device Operation | Instructions to review images accuracy | Safety Notice/Insructions
Z-1823-2011 | 57455 | cresolphthalein complexone, calcium | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Reagent, Calcium Test System for COBAS INTEGRA 400, 400 PLUS, 800, 800 CTS, cobas c 311, 501 and 502, Part number 20763128322, Roche Diagnostics Corporation, Indianapolis, IN.For the in vitro quantitative determination of calcium in serum, plasma and urine. | Calcium test system | 2 | March 25, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The firm has recieved complaints of imprecision and erroneous patient results for calcium on the COBAS INTEGRA and COBAS c 501 system using reagent lots 62601901 and 63020401. | OTHER/UNDETERMINED: Pending | The firm, Roche, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated October 11, 2010 to all consignees/customers via UPS(signature required) . The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the recalled product; discard from inventory according to their site's local regulations; complete and return the attached Faxback form via fax to 1-877-598-5705 even if they do not have the affected product, and file this letter for future reference. The firm added lot 63020401 to the recall in 3/2011. The firm is mailing an URGENT MEDICAL DEVICE REMOVAL notice to customers that purchased that lot of product.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have questions about the information contained in this letter. | Contact the recalling firm for information | 418 | 13, 405 units | Nationwide distribution. | Terminated | September 13, 2011 | 172 | 14946 | Not_Computer
Z-1822-2011 | 58002 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ossur Total Knee, Model Number: 2100The Total Knee 2100 is an external prosthetic knee joint. It is a polycentric knee designed for multi-speed ambulators, higher impact levels and heavier loads. | Ossur Total knee | 2 | March 25, 2011 | 2011 | Ossur Engineering, Inc | The recall was initiated based on a finding by Ossur that some units of the Total Knee 2100 produced from October 19th, 2010 until January 21st 2011, contained back links that may compromise knee function prematurely. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Ossur initiated a Safety Notice (Voluntary Product Recall) letter on February 11, 2011, via US Postal Service Registered Mail. The recall letter informed the affected customers of the reason for the recall and the product affected.Ossur requested their distributor and providers take the following action: (1) Please identify the users of each of the Total Knee 2100 model devices listed above; (2) Also, please check your inventory and isolate any recalled devices you have on hand but have not fitted on users. Customers were also informed that a representative from Ossur customer service will be contacting them to coordinate replacement of the affected devices.The letters were followed by phone calls placed by a trained Ossur Customer Service Representative to each customer after US distributors provided Ossur with customer contact information.For questions regarding this recall call 517-629-8890. | Contact the recalling firm for information | 142 | 142 units | Worldwide Distribution - USA including CA, CO, FL, LA, MD, MI, MO, NJ, NY, OH, TN, VA,, and WA and the countries of Australia, Canada, Czech Republic, Germany, France, United Kingdom, Ireland, Italy, Japan, The Netherlands, Russian Federation, Sweden and Turkey. | Terminated | February 29, 2012 | 341 | 142 | Not_Computer
Z-1821-2011 | 58017 | catheter, infusion | JCY | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ***REF 24k100***QTY 1***24k Arthroscopy Outflow/Suction Tubing Set***MADE IN USA***Rx Only***CONMED LINVATEC, CONMED LINVATEC 11311 CONCEPT BLVD, LARGO, GL 33773-4900***For use in conjunction with the ConMed Linvatec 24k Pump, for pump-induced outflow and/or suction of fluid from the joint during arthroscopic procedures. | Arthroscopy Outflow/Suction Tubing | 2 | March 25, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | On 09/13/2010 ConMed Linvatec, Largo, FL initiated a recall on 24k100 Arthroscopy Outflow/Suction Tube Set used with the ConMed Linvatec 24k Pump. The 24k100 Arthroscopy Outflow/Suction Tubing Set Lots 22747-000, 22748-000, and 2749-000 may have been incorrectly assembled and could potentially allow waste to enter the sterile field. | OTHER/UNDETERMINED: Pending | Linvatec Corp. dba ConMed Linvatec sent an Urgent Medical Device Recall Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility's inventory for the affected product. If they had any of the product they were instructed not to use and segregate the product and return for credit.Return the affected products and the completed REPLY FORM to their local ConMed Linvatec office.If the product was transferred to another facility, customers were instructed to complete the REPLY FORM documenting the transfer of the product and mail or fax the REPLY FORM to their local ConMed Linvatec office at 727-319-5701.If the customers do not have the affected product they should complete the REPLY FORM and mail in the enclosed envelope or fax to 727-319-5701. | Contact the recalling firm for information | 4598 | 4,598 | Worldwide Distribution - USA (nationwide) and internationally to Saudi Arabia, Italy, France, Canada, Portugal, South Africa, Switzerland, Denmark, Germany, New Zealand, Australia, and Spain. | Terminated | January 31, 2012 | 312 | 4598 | Not_Computer
Z-1834-2011 | 54706 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Coherence Dosimetrist, version 2.2.Intended use: Medical charged-particle radiation therapy system | Coherence Dosimetrist, v 2.2 | 2 | March 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A problem in the software has the potential to be associated with the dose; when a template is created the user has the possibility to predefine organs. For the Target the user defines the max. dose and the min dose, for the OAR (Organs at Risk) the user only defines the max. dose. The firms engineering team determined that the software defect was repeatable. | DESIGN: Software Design | Affected consignees will be notified by the Customer Service Engineer to schedule a software update beginning November 2009. |  | 60 | 60 units | Worldwide distribution: USA including: CA, KY, MI, NJ and TN; and countries of: Canada and Mexico. | Terminated | April 06, 2011 | 9 | 60 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1832-2011 | 54842 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson S2000 Ultrasound System with software and imaging.The Acuson S2000 ultrasound imaging systems are intended for these applications: fetal, abdominal, intraoperative, pediatric, small parts, transcranial, OB/GYN, cardiac, pelvic, neonatal/adult cephalic, vascular, musculoskeletal, superficial musculoskeletal, and peripheral vascular applications. Also for anatomical structures and calculation packages that provide information that provide information to the clinician...for clinical diagnosis purposes. | Acuson | 2 | March 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | "Any 2D measurement performed in Exam Review or offline, on a zoomed Retrospective clip capture may be erroneous. This applies to 2D, 2D/PW, 2D/CW and 2D/M-mode." | DESIGN: Software Design | Siemens Medical Solutions USA, Inc. sent a Consumer Safety Advisory Notification letter on February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the Customer Safety Notice to all those within their organization who need to be aware of the issue until the corrective action was completed. For questions regarding this recall call 650-694-5580. |  | 1903 | 1903 | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Australia, Belgium, Bolivia, Brazil, Bahamas, Canada, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Ecuador, Eguypt, Spain, Finald, France, Great Britian, Germany, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Ireland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Lithuania, Mexico, Malaysia, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, Slovak Republic, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, Venezeula, Yemen, and South Africa.. | Terminated | June 18, 2012 | 448 | 1903 | Software
Z-1837-2011 | 56840 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed;ACUSON S2000: 10041461;S2000 Cardio Option: 10041853;Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CAThe ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | ACUSON S2000 ultrasound systems | 2 | March 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. For questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5993. |  |  |  | Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (PerÂŽ), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) | Terminated | June 18, 2012 | 448 |  | Software
Z-1831-2011 | 57887 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The enGenTM Laboratory Automation System consists of 3 main components including Instrument Management Subsystem (Instrument Manager), Automated Track Subsystem (TCAutomation) and Configuration File developed by Ortho-Clinical Diagnostics (OCD). The Instrument Manager is a Windows-based software product that operates on a Pentium-class (or higher) PC processor. The instrument management subsystem is capable of connectivity with clinical analyzers, a Laboratory Information Systems (LIS), barcode label printers and other third party devices. The TCAutomation is a pre-analytical and post-analytical sample management system. The configuration file enables operational support of the site specific layout and instrument selection. The VITROS 5, 1 FS Chemistry System with enGenTM Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric Immunoassay methods.) | enGen TM Laboratory Automation System | 2 | March 28, 2011 | 2011 | Ortho-Clinical Diagnostics | A software anomaly that can occur with analyzers connected to the Bypass modules listed below on an enGen" Laboratory Automation System using TCAutomation" Software Version 3.1.1 and below. When this anomaly occurs, it is possible for a sample other than the intended sample to be aspirated, potentially leading to test results being assigned to the wrong patient without error notification. |  | The firm, Ortho Clinical Diagnostics (OCD), sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated January 27, 2011, by Federal Express overnight mail to all customers. Foreign affiliate consignees/customers were notified by email informing them of the issue on the same day. The letter described the problem, product and actions to be taken. The firm also included temporary procedures to help prevent the occurrence of the software anomaly. The customers were instructed to follow the enclosed Procedure to STOP the Automation System and/or Bypass Module; complete and return the attached Confirmation of Receipt - Important Response Required form by February 4, 2011, via fax to: Ortho Clinical Diagnostics, 1-888-557-3759 or 1-585-453-4110, and display this communication near the enGen Laboratory Automation System to ensure compliance by all operators until updated TCAutomation Laboratory Automation System software is installed. A Technical Bulletin will be sent to customers in the near future to be stored with user documentation. Ortho will implement a TCAutomation Laboratory Automation System software change in the near future to resolve this anomaly.If you have any questions regarding this notification, please call Customer Technical Services at 1-800-421-3311 (Options 1, 2). |  | 63 | 63 units | Worldwide distribution: USA and countries including: Canada, Chile, England, France, Germany, Italy, Mexico, Singapore and Spain. | Open |  |  | 63 | Software | Output/Calculation | Software update | Software Update
Z-1833-2011 | 58025 | guide, surgical, instrument | FZX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | REF 32-485050 VANGUARD KNEE INSTRUMENTATION MICROPLASTY SLIDEX ANTERIOR/POSTERIOR, QTY: 1The guide is used for femoral sizing - Neutral or 3 degree femoral rotation feet are attached to the slidex AP sizing guide. | VANGUARD MP SLIDEX AP SIZER | 2 | March 28, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action due to the scale on the stylus of the Vanguard MP Slldex AP Sizer being incorrect for this assembly resulting in the scale being out of position. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, Biomet, sent a "FIELD SAFETY NOTICE" dated February 11, 2011 to all customers (heads of Orthopaedic Departments/Operating Departments /Sterile Services Departments/Procurement /Supplies/ Risk Management). The notice described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the product; return to Biomet or their local Biomet distributor; ensure that the operating staff are made aware of this issue, and complete and return the FAX BACK RESPONSE FORM to Biomet UK Ltd or their local Biomet distributor via fax at +44 (0) 1656 645454. If you have any questions regarding this communication, please contact the Regulatory & Compliance Manager at +44 (0) 1793 645317 or email: edward.spearpoint@biometeurope.com. |  | 17 | 17 | Foreign only; UK, Malta, The Netherlands, Italy, Turkey, Switzerland. | Terminated | May 16, 2011 | 49 | 17 | Not_Computer
Z-1830-2011 | 57899 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Interventional Pain 6" 15 G (1.85 mm) Percutaneous Discectomy Probe, REF 407-266, Sterile, Stryker Instruments, Kalamazoo, MI.The DeKompressor Percutaneous Discectomy Probe is intended for use in the apiration of disc material during percutaneous discectomies in the lumbar, thoracic, and cervical regions of the spine | Interventional Pain 6" 15 G (1.85 mm) Percutaneous Discectomy Probe, REF 407266 | 2 | March 28, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The firm received five complaints that the stylet did not fit down the cannula. The investigation revealed that the incorrect probe assembly was packaged for two DeKompressor part numbers with the lot number of 10214012. | OTHER/UNDETERMINED: Pending | Stryker Instruments Division of Stryker Corporation sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their inventory for any recalled product and quarantine any affected product. If the product was further distributed, customers were asked to forward the letter and the completed Reply Form to all affected locations. Customers were to be contacted to coordinate return of the affected product. Upon receipt of the product, a replacement would be issued to their account. For questions regarding this recall call 1-800-253-3210. | Contact your health professional | 44 | 44 units both products | Worldwide Distribution - USA including the states of PA, NY, KS, NJ, TX, MI, OK, and FL, and the countries of Italy and Switzerland. | Terminated | June 01, 2011 | 65 | 44 | Not_Computer
Z-1839-2011 | 58104 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue, Manufacturing Part Number 01287-09, Catalog Number D128709.This cable is used to connect an Auto 10 10-pole catheter to an EP recording system. | 10 ft. Interface Cable, 10 Pin Shielded Tip Pins to 12 Pin Blue | 3 | March 28, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has confirmed that there is an inaccessibility of electronic version of the instruction for use (e-IFU) for the "10 ft. Interface Cable, 10 Pin-Shielded Tip Pins to 12 Pin-Blue", catalog number D128709, through the J&J; Gateway (ie, e-IFU website). | OTHER/UNDETERMINED: Pending | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated March 16, 2011, to all affected customers (via express mail). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the enclosed Acknowledgement Form to confirm that all customers who received the product have been notified accordingly of this concern. If customers have any questions or require additional information, they are to contact their BWI representative and return the Acknowledgement form to:Biosense Webster, Inc.l5715 Arrow HighwayIrwindale, CA 91706Attn: Recall Coordinator - Virgil EllsworthPhone: (909) 839-8500Fax Number: (909) 839-7207.For questions regarding this recall call 909-839-7207. | Contact the recalling firm for information | 606 | 606 units total | Nationwide Distribution - including GA, KY, NV, WA, PA, and UT | Terminated | May 07, 2012 | 406 | 606 | Not_Computer
Z-1835-2011 | 57857 | [6 empty] | Triathlon PKR Peg Drill, Catalog Number: 5650-4-512 (Size 1-2 Peg Drill); Catalog Number: 5650-4-536 (Size 3-6 Peg Drill);Howmedica Osteonics Corp. 325 Corporate Drive Mahwah, NJ 07430The Triathlon-PKR Knee System is indicated for primary partial knee resurfacing (PKR). This system utilizes two fixation pegs on the femoral component for fixation and alignment. To prepare for these features, a peg drill, drill guide, and drill through femoral trial are provided. The femoral Trial Drill Guide is attached to the Modular Handle and assembled into the holes on the articulation surface of the Femoral Trial. Using the appropriately sized Peg Drill (Small Drill for sizes 1-2, Large Drill for sizes 3-6), both holes are drilled. The Peg drill is advanced until the step on the drill makes contact with the front face of the Femoral Trial Drill Guide. | Triathlon PKR Peg Drill | 2 | March 28, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has identified a trend of complaints involving reports of the peg drill cold welding to the drill guide, causing the Peg Drill to cease functioning. If cold welding of these components occurs, the following potential patient and/or surgeon or surgical staff harms may results. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Stryker sent Urgent Product Recall letters on February 9, 2011, and February 10, 2011, via Fed Ex. to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069.For questions regarding this recall call 201-972-2100. |  | 13990 | 5650-4-512- 754 units; 5650-4-536- 880 units | Worldwide Distribution - USA , Sweden, Germany; France, Portugal, UK, Canada, Australia, and New Zealand, | Terminated | November 18, 2013 | 966 | 1634 | Not_Computer
Z-1846-2011 | 55905 | abnormal hemoglobin quantitation | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | The VARIANT Sickle Cell Short Program ROM Card; Distributed by and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CAThe VARIANT Sickle Cell Short Program is designed as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography. For in vitro diagnostic use. | VARIANT Sickle Cell Short Program | 2 | March 29, 2011 | 2011 | Bio-Rad Laboratories Inc | The presence of parameters for Version 40400-S on the labeled VARIANT Sickle Cell Short ROM Card instead of Version 80300-S. | OTHER/UNDETERMINED: Pending | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated April 22, 2010, to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and quarantine the product; complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE FORM via fax to: (510) 741-3954, Attn: Bio-Rad CSD Regulatory Affairs, and contact there regional Bio-Rad office for ROM Card replacement. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  | 10 | 10 units | Worldwide distribution: USA including: MD, NJ and UT; and countries of: Italy and Singapore. | Terminated | April 07, 2011 | 9 | 10 | Computer
Z-1852-2011 | 58028 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Large Port Specimen Sock, used to convert the pour/accessory port of a Cardinal Suction canister into a specimen retainer. The sock is used as an accessory in a suction collection system. Its placed in a suction canister to remove solids from the liquid aspirated during surgery. The solids may consist of patient tissue that are then sent to pathology or retained for other purposes. There is no direct contact between this device and the patient.Large Port Specimen Sock, used to convert the pour/accessory port of a Cardinal Suction canister into a specimen retainer. The sock is used as an accessory in a suction collection system. Its placed in a suction canister to remove solids from the liquid aspirated during surgery. The solids may consist of patient tissue that are then sent to pathology or retained for other purposes. There is no direct contact between this device and the patient. | MediVac Speciment Sock | 3 | March 30, 2011 | 2011 | Cardinal Health Inc | In one lot of product (lot# 1091745), the label for the small version of the sock (catalog number 65652-123) was improperly used for the larger version of the product (65652-124) | PRODUCTION CONTROLS: Labeling Mix-Ups | On February 18, 2011 Cardinal Health sent a consignee letter with the following instruction: 1) Quarantine either Catalog Number with the inticated Lot Number.2) Contact the appropiate Customer Service group for return instructions:Hospital - 800-964-5227Federal Government - 800-444-1166Distributirs - 800-635-6021All other customers - 888-444-54403) Return the enclosed acknowledge form via fax to 847-689-9101, even if you do not have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.4) Notify any customer to whom you may have distributed product affected by this recall. | Contact the recalling firm for information | 3360 | 3,360 each | Worldwide Distribution - USA including AL, AR, AZ. CA. CO,, CT, DC, DE, FL, GA, HI, IA, ID,IL,IM,KS, KY, LA, MA, MD, ME. MI MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, & Puerto Rico - and the country of Canada. | Terminated | October 04, 2011 | 188 | 3360 | Not_Computer
Z-1850-2011 | 58105 | chair, dental, without operative unit | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Pelton & Crane Spirit 3000 dental chair (Model #SP30) | Pelton & Crane Spirit 3000 | 2 | March 30, 2011 | 2011 | Pelton & Crane Company | The Link Arm Assembly that connects some rear mounted options to the Spirit 3000 or Spirit 1800 dental chairs may potentially break and fall off the dental chair. | PRODUCTION CONTROLS: Process Control | Pelton & Crane sent an "Urgent Medical Device Recall" letter dated 3/04/2011. They were instructed to: Locate the affected devices this recall is in regards to. Schedule service with the end user for Pelton & Crane representative to replace the affected components this recall is in regards to. International accounts were instructed to order link arm assemblies from Pelton & Crane and replace the affected components. All consignees were to complete the Recall Acknowledgement/Return Form and return the form to Pelton & Crane. |  | 9 | 9 units | Distribution to USA, Canada, and the Netherlands. | Terminated | August 16, 2011 | 139 | 15 | Not_Computer
Z-1845-2011 | 57613 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CellQuest Pro SW, version 4.0.2. and later;Becton Dickinson, San Jose, CA 95131CellQuest Pro software allows you to acquire and analyze data from your flow cytometer on a Macintosh computer. | CellQuest Pro SW | 2 | March 30, 2011 | 2011 | BD Biosciences, Systems & Reagents | Results and gates could be mismatched if you utilize the undo function to undo a gate that was deleted. | DESIGN: Software Design | Important Product Information letters, dated September 2007, were sent the week of 9/14/2001 via Federal Express to all affected customers. The letters identified the affected product along with the reason for recall. The letter states that "as long as customers do not use the Undo function to undo a gate deletion, the software can be used with full functionality." The letter also states that BD will change the issue in the next software version. If customers believe that they may have analyzed sample files by utilizing the Undo function, then it is recommended that the sample files be re-analyzed. If customers need more information they should contact the Customer Support Center at 877-232-8995. |  | 4552 | 4552 | Worldwide Distribution | Terminated | March 31, 2011 | 1 | 4552 | Software | Device Operation | Software update | Software Update
Z-1817-2011 | 57004 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 115V Blanketrol III Model 233 Hyper/Hypothermia System, Part number: 86107; and 115V CoolBlue Model 25-01 Hyper/Hypothermia System, Part number: 86000.Used to lower or to raise a patient's temperature and/or maintain a desired patient temperature through conductive heat transfer. | 115V Blanketroll III Model 233 and 115V CoolBlue Model 2501 | 2 | March 30, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The firm initiated this recall due to changes that have been made to the Blanketrol III Model 233 device's Operation and Operation/Technical Service Manuals, in order to stay in compliance with labeling regulations. | MISBRANDING: Labeling False and Misleading | On 10/18/2010 & 11/5/2010 the firm sent URGENT Medial Device Field Action letters, dated 10/18/2010, to their customers. The letters identified the affected product and stated the reason for recall. Customers were instructed to immediately collect and discard all affected product Operation and Operation/Technical Service manuals currently in their possession. They were to replace all the discarded manuals with the revised versions that were provided with the letter. The enclosed response form should be completed and returned by fax to 513-772-9119, scanned and e-mailed to B3manual@cszinc.com. If the product had been further distributed, then those customers should also be identified and notified of the field action via a copy of the letter. Customers should direct questions to the field action coordinator, David Niehaus at 513-773-8810 ext. 3234. |  | 1385 | 763 - 115V Blanketrol III Model 233; & 133 - 115V CoolBlue Model 25-01 | Worldwide Distribution -- US, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, & WI and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, CHILE, CROATIA, ECUADOR, EL SALVADOR, GERMANY, HONG KONG, INDONESIA, ISRAEL, ITALY, KOREA, MALAYSIA, MEXICO D.F., NEW ZEALAND, PEOPLE'S REPUBLIC OF CHINA, PERU, REPUBLIC OF SIGNAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY & VENEZUELA. | Terminated | March 08, 2013 | 709 | 896 | Not_Computer
Z-1847-2011 | 58066 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" Panel NMIC/ID-123, Catalog # 448723, packaged 25 panels/carton, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663***The BD PhoenixÂŽ Automated Microbiology System is intended for the rapid identification (ID) and antimicrobial susceptibility testing (AST) of clinically significant bacteria. The Phoenix system provides results for most aerobic and facultative anaerobic Gram negative bacteria of human origin with the use of the NMIC/ID-123 panel type. | BD Phoenix" Panel NMIC/ID123, Catalog 448723, | 2 | March 30, 2011 | 2011 | Becton Dickinson & Co. | In vitro diagnostic test kit is defective and may cause false positive or false negative results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | BD Diagnostics Systems sent an Urgent Product Recall letter dated February 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinsue use of the affected product and discard any remaining packages. BD would issue replacements for the discarded material. Customers were instructed to complete the attached form whether or not they had any inventory remaining so that the firm may acknowledge receipt of this notification. Complete and fax the form to BD Regulatory Compliance to 410-316-4258. For further assistance contact BD Customer Service at 1-800-675-0908. For other inquiries, contact BD Technical Services at 1-800-638-8663. |  | 377 | 377 cartons | Worldwide Distribution - USA including AZ, CA, FL, GA, MA, MD, NC, NY, PA, TX, WA, WI, and WY and the countries of The product was distributed to medical facilities nationwide and to foreign consignees in Argentina, Brazil, Chile, and Mexico. | Terminated | January 30, 2013 | 672 | 377 | Not_Computer
Z-1848-2011 | 57602 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HEMOCUE AB GLUCOSE 201 DM ANALYZER, in vitro diagnostic device for quantitative determination of blood glucose level. | HEMOCUE AB GLUCOSE 201 DM ANALYZER | 2 | March 30, 2011 | 2011 | Hemo Cue, Inc. | Incorrect configurations settings in software. As a result, certain error codes do not trigger appropriately. These error codes are: E01, E02 and E05 which are all related to the optronic performance of the analyzer. Possible incorrect results. | PRODUCTION CONTROLS: Process Control | Hemocue AB sent Customer Notification "Notice of Medical Device Correction on HemoCue Glucose 201 DM and HemoCue Hb 201 DM Analyzers" dated October 7, 2010, to customers notifying them of incorrect settings on error codes. The corrective action consists of an update of the configuration settings on the the affected HemoCue Glucose 201 DM and Hb 201 DM analyzers that are in the field. The affected units in the field will be corrected in accordance with a rework procedure that has been established by HemoCue AB. |  | 23 | 23 | Worldwide distribution, including USA, and Glucose 201 DM analyzers to Norway and the Netherlands; Hb 201 DM analzyers to the Netherlands, France, Ireland, Great Britain, Poland and Germany. | Terminated | June 22, 2011 | 84 | 102 | Software | Alarm/Message | Configuration update | Other
Z-1853-2011 | 54844 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound Imaging System with software versions 4.0 and 5.0. | Acuson Antares Utrasound imaging device | 2 | March 30, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Image captured from the device may be erroneous. | DESIGN: Software Design | Acuson Customer Safety Advisory notification letters were sent March 1, 2010 by certified mail. |  | 3539 | 3539 | Worldwide distribution to USA, Andorra, United Arab Emirates, Netherlands Antilles, Angola, Austria, Australia, Azerbaijan, Bosnia, Bangladesh, Belgium, Bahrain, Brunei Darussalam, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, Faroe Islands, France, Great Britain, Georgia, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Sri Lanka, Lithuania, Latvia, Morocco, Moldova, Mexico, Nicaragua, Netherlands, Norway, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Suriname, Thailand, Turkey, Trinidad and Tobago, Taiwan, Ukraine, UA, Uruguay, Venezuela, Vietnam, Yemen, and South Africa. | Terminated | June 18, 2012 | 446 | 3539 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-1854-2011 | 57773 | system, hemoglobin, automated | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HGBPROKIT which contains one meter and one container of 25 test strips. HGB TEST 100; HGB PRO TEST STRIPS -100, includes four containers of 25 test strips per container; International Technidyne CorporationThe Hgb Pro Professional Hemoglobin Testing System Consists of a portable battery operated meter and single-use test strips for the quantitative measurement of total hemoglobin from capillary and venous whole blood. The Hgb Pro Professional Hemoglobin Testing System is designed for near patient testing in a professional facility and is not intended for home use. | Hgb Pro Professional Hemoglobin Testing system | 2 | March 31, 2011 | 2011 | ITC-Nexus Dx, Inc. | ITC has determined that results obtained using certain lots of Hgb Pro Professional Hemoglobin Test Strips may differ from results obtained using a reference hemoglobin testing system. | OTHER/UNDETERMINED: Pending | ITC-Nexus Dx, Inc. sent an Urgent Medical Device Recall letter with return response forms to end users and distributors dated January 21, 2011 via UPS overnight. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to determine if they had an of the affected lots.If they have full boxes or loose containers of the affected lots customers were instructed to stop using the product and remove it from inventory.Customers were requested to fill out the return response form and send it back to ITC via fax, email or mail.For questions customers were instructed to call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (International)), Extension 4707 or e-mail techsupport@itcmed.com. |  | 2500 | 2500 containers + 50,563 addtnl containers =53063 containers | Worldwide Distribution - USA (nationwide) and the countries of India, Oman, South Africa, and Singapore | Terminated | November 02, 2012 | 582 | 2500 | Not_Computer
Z-1857-2011 | 57987 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Stanley Bed-Check 30-Day IT Chair Sensormat, With 8 Ft. Cord, Model 73010, 10 mats per inner box, 5 inner boxes per overpack box (total 50 units). The firm name on the box is Stanley, 1-800-523-7956. The mat contains only the phone number.For use in reducing patient falls. | Stanley BedCheck 30Day IT Chair Sensormat, With 8 Ft. Cord, Model 73010 | 2 | March 31, 2011 | 2011 | Stanley Security Solutions, Inc. | The cord could come loose from the mat potentially causing a short to the internal wiring. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, Stanley Healthcare Solutions, issued an "Important Technical Bulletin" dated November 25, 2009 to all customers via e-mail. In addition, the firm issued a "Compliance Notice" dated August 30, 2010 via certified mail because the firm had not collected all of the units with the defect. The bulletin described the product, problem and actions to be taken. The customers were instructed to check the cords on the mats to make sure they weren't pulled out, if the cords were pulled out, they were to return the mats for replacement; contact Customer Service at 1-800-824-2996 ext: 3440 to obtain a return authorization number and take immediate action to return the units to the firms location in Lincoln, Nebraska. The Compliance Notice informed the customers who had not sent in their units to promptly read the notice and take the necessary steps to respond appropriately for their facility. Additionally, the customers were instructed to complete and return the Bed Check Sensormat Return Form and page 2 of the "Compliance Notice" via at 402-475-4281 or mail to: Compliance Engineer, Stanley Security Solutions, Inc., 1550 N. 20th Circle, Lincoln, NE 68503. Follow-up phone calls to nonresponders were made using a phone script. A "Second Notice Compliance Notice" dated October 20, 2010 was issued via certified mail to customers asking for immediate response to the previous communication. In addition, a third notice dated October 28, 2010 which was also flagged as "Second Notice" was again used asking for a response from the customers.If you have any questions regarding the notice or any action that may need to be taken by you, please contact Quality Manager at 1-800-824-2996 extension 3419 or Compliance Officers at 402-742-9414 or 402-742-9425. |  | 19997 | 19,997 mats | Worldwide distribution: USA and the country of Canada. | Terminated | March 31, 2011 | 0 | 209183 | Not_Computer
Z-1862-2011 | 57852 | gauge, depth | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Specialty Patella Milling System Per File K2399; Catalog Number I-K2399PR00, I-K2399PR10; Stryker Howmedica Osteonics, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430The Xcelerate patella milling system (as detailed in surgical protocol LSPK35) was designed to prepare a patella to accept all Scorpio or Duracon patellar implants using either a resurfacing or recessed technique. | Xcelerate Specialty Patella Milling System | 2 | March 31, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the 10 mm on the bone remaining depth gage is incorrect thus allowing for the reamer to leave as little as 8.9mm (in a nominal condition) patella bone. | DESIGN: Device Design | Stryker sent an Urgent Product Recall letter dated January 27, 2011, via FedEX for January 28, 2011, delivery with return receipt. The letter identified the product, the problem, and the action to be taken by the customer. Customers were insructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069. Customers were also instructed to contact the hospitals in their territory that have the affected product to arrange return of the product. Product should be returned to the attention of: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey, 07430. The orange Product Remediation sticker should be attached to the return, indicating the Product Remediation # RA 2009-296. Mark the outer box with words "Product Recall."For questions regarding this recall call (201) 972-2100. |  | 52 | 52 units | Nationwide Distribution | Terminated | October 15, 2012 | 564 | 52 | Not_Computer
Z-1866-2011 | 58136 | dressing, wound and burn, occlusive | MGP |  |  | General & Plastic Surgery | 510(k) | HydroHydrofilm is for use as a post-operative and primary dressing to protect against secondary infection and to cover non-exudating wounds and for use as a secondary dressing for fixation of catheters or cannulae.film, 10cm x 12.5 cm, Paul Hartmann AG, 89522 Heidenhiem, Germany | Hydrofilm | 3 | April 01, 2011 | 2011 | Hartmann USA, Inc | Misbranded; the expiry date on the primary packing incorrectly indicates 2010-10, instead of 2015-10. Consignees were notified by letter on/about 02/11/2011. They were advised to examine their inventory and not to use any of the affected product and to return it to Paul Hartmann. A response form was included to be completed and returned to Paul Hartmann. They were further instructed to pass the no | OTHER/UNDETERMINED: Pending | Hartmann USA, Inc. sent letter to the Consignees on 2/11/11. They were advised to examine their inventory and not to use any of the affected product and to return it to Hartmann USA, Inc. A response form was included to be completed and returned to Hartmann USA, Inc. They were further instructed to pass the notice on to their consignees if distributed. For questions, call (803) 325-7600. | Contact the recalling firm for information | 100 | 100 boxes (10 x 12.5 cm per box) | Nationwide -- Including CA, GA, IL, MO, MS, TX, and VA. | Terminated | August 12, 2011 | 133 | 100 | Not_Computer
Z-1875-2011 | 58051 | tape and bandage, adhesive | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General & Plastic Surgery | 510(K) Exempt | CVS Pharmacy Plastic Sesame Street Adhesive Bandages; a sterile, latex free, assorted sizes with multiple designs adhesive bandages; 20 bandages per box, 24 boxes per case; Distributed by: CVS Pharmacy, Inc., One CVS Drive, Woonsocket, RI 02895, Made in Egypt; REF CVS 826159, UPC 0 50428 95118 7The device is used to cover and protect wounds. | CVS Pharmacy Plastic Sesame Street Adhesive Bandages | 2 | April 01, 2011 | 2011 | Pharmaplast S.A.E. | The bandages are labeled as sterile, but failed sterility testing. | PRODUCTION CONTROLS: Process Control | Pharmaplast S.A.E. called the direct account on 2/17/11 and sent a Urgent Medical device recall letter dated 2/21/11 to the direct account via overnight mail, requesting that they examine their inventory for the non-sterile adhesive bandages and quarantine them for return to Pharmaplast via sea freight. The account was instructed to contact their customers concerning the recall and return their stocks of the affected lot, which was done by telephone calls on 2/17/11.For questions regarding this Recall please call 00203-4500264. | Contact the recalling firm for information | 5336 | 5,336 cases | Nationwide Distribution - Including Illinois. | Terminated | August 25, 2011 | 146 | 5336 | Not_Computer
Z-1865-2011 | 57890 | block, beam-shaping, radiation therapy | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Axxent Flexishield MiniTo shape the beam from a low energy radiation therapy source up to 50kVp. | Axxent Flexishield Mini | 1 | April 01, 2011 | 2011 | Xoft Inc. | Xoft Axxent Flexishield Mini product may may shed particles identified as tungsten, and may look like suspicious calcifications on follow up scans if the product was used during Intraoperative radiation therapy. | DESIGN: Device Design | Xoft, Inc. sent an Urgent Field Removal Action letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and return all units and lots of Flexishield Mini Catalog Number: F5300 in their inventory. Customers were instructed to fill out the attached form and fax to 866-222-3404, then call Customers Service toll free directly at 877-963-8327 to make arrangements to return the product and receive a credit if applicable. Customers were asked to complete and return the form even if they did not have any units to return.For questions regarding this recall call 408-419-2300. |  | 165 | 165 units | Nationwide Distribution including AL, CA, CO, FL, IL, MI, MO, MS, NJ, NY, OK, PA, TN, and WI. | Terminated | June 08, 2011 | 68 | 165 | Not_Computer
Z-1864-2011 | 57742 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Cardiac Marker Control Calibration Verification Set, List number 136604/06F15-02.Used to verify the accuracy of results over the measurement range of the i-STAT cTnI test. | Cardiac Marker Control Calibration Verification Set | 2 | April 01, 2011 | 2011 | Abbott Point Of Care Inc. | Presence of CO2 in the head space of the vial can result in the generation of high results in Cardiac Marker Control Level 1, 2, and 3 which is used along with the Cardiac Marker Control Calibration Verification Set. | DESIGN: Device Design | Urgent Recall Notice letters with business reply cards were sent Fed ex on December 17, 2010 to all accounts. International notifications were sent on December 21, 2010. The letter identified the affected product and lot numbers along with a background explanation of the issue. The firm recommends that customers remove the cap from the vial; squeeze the vial 10 times, prior to thawing; replace cap on the vial; and then follow the remaining steps for use as outlines in section 14-12 of the i-STAT 1 System Manual. Customers are to complete and return the Business Reply card to acknowledge receipt of the letter. If the affected product had been further distributed then a copy of the letter should be provided to those customers as well. If additional information is needed, customers should contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or their Abbott Point of Care representative. | Contact the recalling firm for information | 557 | 557 units total | Worldwide Distribution | Terminated | October 16, 2012 | 564 | 557 | Not_Computer
Z-1877-2011 | 57445 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5, Part Number: 6607073.Allows simultaneous identification and enumeration of total CD3+, total CD4+, total CD8+, dual CD3+/CD4+, dual CD3+/CD8+ and/or total CD3+, CD19+ and CD3-/CD56+ lymphocyte percentages and absolute counts in whole blood by flow cytometry. The systems also provide the CD4/CD8 ratio when using CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, and total lymphocyte percentage when using CD45-FITC/CD56-RD1/ CD19-ECD/CD3-PC5Al. | CYTOSTAT tetraCHROME CD45FITC/CD56RD1/CD19ECD/CD3PC5 | 2 | April 04, 2011 | 2011 | Beckman Coulter Inc. | A labeling error resulted in the specimen age stability and the prepared sample stability being stated incorrectly in the instructions for use. There is no impact to the product safety or functionality. |  | An Urgent: Product Corrective Action letter dated, September 27, 2010, was sent to customers. The letter identified the affected product and stated the reasons for the recall. Customers were also asked to take the following actions: ACTION/RESOLUTIONFor issue 1: The specimen stability claims should be as follows:- 48 hours for both % positives and absolute counts for the CD45/4/8/3 tube- 24 hours for both % positives and absolute counts for the CD45/56/19/3 tubeFor Issue 2: Results produced are not impacted when samples are prepared and run within the post collection time frames provided in the product instructions for use:- Preparation within 72 hours of collection for the CD45/4/8/3 tube- Preparation within 24 hours of collection for CD45/56/19/3 tubeBeckman Coulter requested consignees complete and return the enclosed response form within 10 days so they may maintain their records. The firm also asked that they share the information with laboratory staff, and retain this notification as part of their Quality System documentation.Questions regarding the Product Corrective Action were directed to (800) 526-7694 in the United States or (800) 463-7828 in Canada. If outside the US or Canada, consignees were instructed to contact their local Beckman Coulter Representative. |  | 26822 | 26822 units | Worldwide Distribution -- United States, Albania, Australia, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Bulgaria, Burundi, Cameroon, Chile, China, Colombia, Croatia, Czech Republic, Denmark, El Salvador, Finland, France, French Guiana, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan Panama, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | June 14, 2012 | 437 | 105810 | Not_Computer
Z-1883-2011 | 55554 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit, Part Number: A78803The Access Accu Tri assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of cardiac troponin I (cTnl) levels in human serum and plasma using the Access Immunoassay Systems to aid in the diagnosis and treatment of myocardial infarction and cardiac muscle damage. Cardiac Troponin I determination aids in the risk stratification of patients with unstable angina coronary syndromes with respect to relative risk of mortality, myocardial infarction, or increased probability of ischemic events requiring urgent revascularization procedures. | Access Immunoassay Systems enhanced AccuTnl Reagent | 2 | April 04, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that under certain circumstances, the following assays may produce elevated values at the low end when run in conjunction with the enhanced Access Accu Tnl assay (P/N A 78033):(1) Access Total BhCG (P/N 33500);(2) Access HYPERsensitive hTSH/Fast hTSH (P/N 33820);(3) Access hFSH (P/N 33520).The issue can occur only when the assays are | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated December 16, 2009 to all customers. The letter described the product, problem and actions to be taken. Beckman Coulter instructed the customers to do the following mitigation to reduce the risk of encountering an erroneous result:Systems with multiple reagent pipettors (UniCel Dxl 800 and 600; UniCel DxC 880i, 860i, 680i, and 660i):1. Designate enhanced Access AccuTnl to a different reagent pipettor (s) than Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.2. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs3. Calibrate Access Total BhCG, hTSH and hFSH using fresh reagent packs4. Run QC following acceptable calibration curves. Proceed with operation following acceptable QC performanceSystems without multiple reagent pipettors (Access, Access 2; SYNCHRON LXi 725, UniCel DxC 600i):1. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs.2. Calibrate Access Total BhCG, hTSH/Fast hTSH and hFSH using fresh reagent packs3. Run QC following acceptable calibration curves.Ensure that controls and patient samples for Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH are not run immediately following enhanced Access AccuTnl controls or patient samples.For example, during routine QC runs, ensure that these assays are requested prior to enhanced Access AccuTnl.4. Proceed with operation following acceptable QC performance.NOTE: An enhanced AccuTnl sample loaded with a system priority of STAT will interrupt routine Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH sample aspiration. Avoid loading enhanced AccuTnl sample aspiration. Avoid loading enhanced AccuTnl samples with a system priority of STAT during aspiration of Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.Diagnostic test results should be interpreted in light of the clinical presentation of the patient, including: symptoms, clinical history, data from other . | Contact the recalling firm for information | 22032 | 22,032 kits | Worldwide distribution: USA and country of Canada. | Terminated | May 07, 2012 | 399 | 22032 | Not_Computer
Z-1884-2011 | 52820 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | TEE Transducer V5M, when used with the Acuson S2000 Ultrasound System, software versions: VA15, VA15A or VA15B; Model Numbers: 10041461 (Acuson S2000) and 8263703 (V5M Transducer);Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CAProduct is a medical device, a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. tis function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-mode, MMode, Pulsed(PW) Doppler mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or harmonic imaging and 3d/4d imaging on a flat panel display. The device system is intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides ability to measure anatomical structures and calculation packages that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. The device is a multi frequency, multi plane, phased sector array transducer. In the TEE (transesophageal echocardiography process where the device fails, the transducer is introduced into the esophagus of the patient to obtain images of the heart structure. | TEE Transducer V5M | 2 | April 04, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Device function failure--When performing a TEE exam with the transducer, CW (Continuous Wave) mode is not functional. Due to work flow, this may not be apparent until the transducer is inserted into the patient's esophagus. | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notification letter on April 7, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to pass the notice on to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-969-9112. |  | 61 | 61 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Great Britain, Hong Kong, Indonesia, India, Italy, South Korea, Norway, Turkmenistan, and Taiwan | Terminated | April 05, 2011 | 1 | 61 | Other |  |  | Other
Z-1879-2011 | 57810 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King System , King LTS-D Kit Size 5, Sterile, REF KLTSD405, Manufacturer: King Systems Noblesville, IN. Ventilation during anesthesia during procedures of short duration. An oropharyngeal airway is a device inserted through the mouth to provide a patent airway. | King LT(S)D Oropharyngeal Airway | 2 | April 04, 2011 | 2011 | King Systems Corp. | The firms contract manufacturer reported that the products were below the specification for hardness. This could contribute to a higher incidence of problems during the insertion and final positioning of the airway in the patient. In rare instances the softness of the tube could cause the tube to fold over. | OTHER/UNDETERMINED: Pending | King Systems Corp. sent an URGENT DEVICE RECALL NOTICE letter dated January 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The letters instructed users to cease all sales of the recalled product, to notify the consignees for any product that was further distributed, and to return all recalled product to King Systems. The notice did include a response form.For questions call 317-776-6823, ext 257. |  | 143 | 143 cases of 10 each case | Worldwide Distribution - USA and Canada | Terminated | June 29, 2012 | 452 | 281 | Not_Computer
Z-1882-2011 | 58044 | container, supplementary nitroglycerin | LLC |  |  | General Hospital | 510(k) | iSite PACS with iSite ViewForum Applications, R7.4V6 L1 L2, L3, Versions: 3.6.96, 3.6.101, 3.6.110, 4.1.96, 4.1.101, and 4.1.110; Product is manufactured and distributed by Philips Medical Systems, Foster City, CAViewForum Applications: iSite ViewForum Applications is an image processing workstation software package designed to run on standard PC hardware. The hardware required is made up of "off-the-shelf" standard computer components. iSite ViewForum Applications software receives image data from medical scanning devices, such as CT and MRI, or from image archives. iSite ViewForum Applications performs viewing, image manipulation, communication, printing, and quantification of images. | iSite PACS with iSite ViewForum Applications | 2 | April 04, 2011 | 2011 | Philips Medical Systems | A software anomaly that occurs when clinicians use Volume Vision to view images in 3D and uses the measurement tool to measure regions of interest. This could lead to an incorrect interpretation of the measurements which could cause an incorrect diagnosis and incorrect treatment plan. | DESIGN: Software Design | Philips sent an Urgent Field Safety Notice letter dated February 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to assure that all colleagues were aware of the problem.When using the Volume Vision software, check the correctness of measurements (graphics and numerical values) prior to saving a captured image containing measurements.Save captured images containing measurements immediately after creating them and before altering the measurements in the original image view.For any subsequent measurements, create a new capture and save it immediately.For further information or support concerning this issues, customers should contact their local Philips representative at +1-877-328-2808. | No consumer action necessary |  |  | Worldwide Distribution. | Terminated | August 18, 2011 | 136 |  | Software
Z-1880-2011 | 58013 | instrument, surgical, orthopedic, ac-powered motor and accessory/attachment | HWE | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | Orthopedic | 510(K) Exempt | POUCH/CARTON LABEL: ***REF 5023-138***9.5 x 25.5 x 0.4mm***QTY 1***MICRO SAGITTAL BLADE, FINE***STERILE***Rx Only***CONMED LINVATEC, CONMED LINVATEC 11211 CONCEPT BLVD LARGO, FL 33773-4908 USA***Made in USA*** To cut soft tissue or bone | Surgical Blade | 2 | April 04, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | On Nov 30, 2010 ConMed Linvatec, Largo, FL recalled their Micro Sagittal Blade, Fine Lot number 180209, Product Number 5023-138. Some units of Micro Sagittal Blade, Fine from Lot 180209 may have compromised seals in the sterile packaging. | PRODUCTION CONTROLS: Process Control | Linvatec Corp. dba ConMed sent an Urgent Medical Device Recall Notification letter dated November 30, 2010, by FedEx, email, Fax, etc) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the reply form within 30 days of receipt. Immediately check their facility's inventory for the affected product and lot number. If they had any of the product with lot number 180209. Do Not Use. Segregate and return for credit. If the product was transferred to another facility, customers were to complete the REPLY FORM documenting the transfer of the product and fax the form to 727-319-5701.For questions regarding this recall call 727-319-5701 or 800-535-8536. | Contact the recalling firm for information | 415 | 415 | Worldwide Distribution - USA including CT, FL, GA, HI, IN, MO, MS, MT, NJ, NY, OH, PA, SC, and TX and the countries of Canada, and South Africa. | Terminated | November 09, 2011 | 219 | 415 | Not_Computer
Z-1896-2011 | 58011 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Rusch Irrigation Tray w/Piston Catalogue Number 67890For irrigation according to institutional guidelines. | Rusch Irrigation Tray w/Piston | 2 | April 05, 2011 | 2011 | Teleflex Medical | The irrigation trays contain alcohol prep pads which are included in the scope of a recall initiated by Triad Group, due to a potential bacterial contamination. | OTHER/UNDETERMINED: Pending | the firm,Teleflex Medical, sent two "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letters one dated February 24, 2011 and the other dated March 30, 2011 to all consignees and customers. The letters described the product, problem and actions to be taken. The consignees were instructed to communicate the recall to any of their customers who have received product included within the scope of the field corrective action via the customer letter. The customers were informed in the letter of concerns about potential bacterial contamination of the alcohol prep pads in the affected trays. If contaminated alcohol prep pads are used, there is a potential for infection to occur. The customers were instructed "that upon opening the affected trays for use, they should discard the prep pads as medical waste". The customers were also instructed to immediately discontinue use and quarantine any products with the product numbers and lot numbers listed; return the product, and complete and return the Recall Acknowledgement Form, even if they no affected product, via fax to 1-866-804-988, Attn: Customer Service (a customer service representative contact the customer with a Return Goods Authorization number and instructions for the return of the product to Teleflex Medical). If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. | Contact the recalling firm for information | 95647 | 95,647 eaches (total for both Catalogue Numbers) | Worldwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, MS, MI, MO, MS, NC, NJ, NH, NY, OH, PA, TX, UT, VA, Puerto Rica and the country of: Canada. | Terminated | April 17, 2012 | 378 | 95648 | Not_Computer
Z-1893-2011 | 58009 | camera, surgical and accessories | KQM | Surgical camera and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ***REF VP6528***Qty 1***Articulating Arm, Secondary Flat Panel***RX Only***See Instructions for Use***CONMED LINVATEC, CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA***Catalog No ProductVPN6528 Articulating Arm, Secondary Flat Panel. 690-0024-00 Arm, Wall Mount, Flat Screen, GCXFor mounting and positioning of flat screen monitors. | Articulating Arm, Secondary Flat Panel | 2 | April 05, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | If the VHM Arm is left in the locked position, the gas spring can lose pressure and releasing the lock can result in a sudden downward movement. | DESIGN: Device Design | An Urgent Service Bulletin Notification was mailed to each ConMed Linvatec customer who received the affected product. The letter identified the affected product and stated the reason for notification. The letter asks customers to review the GCX Corporation Service Bulletins and updated Installation/User Manual. ConMed Linvatec recommends that customers periodically check the VHM Arm to see that it is functioning properly. In addition "Sudden Movement" stickers, provided by GCX, were also included to all consignees as part of the mailing. If customers in the US have questions or need technical support, they can contact ConMed Linvatec Endoscopy Customer Service at 888-292-0100 or e-mail Customerserv1@linvatec.com. |  | 373 | 373 | Worldwide Distribution -- Australia, Canada, Morocco, Saudi Arabia, Taiwan, Province of China, and throughout the U.S. | Terminated | May 17, 2011 | 42 | 373 | Not_Computer
Z-1894-2011 | 58109 | knife, ophthalmic | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Beaver XStar Safety 2.5mm Crescent Knife;REF 378234Crescent blades are used in a variety of ophthalmic surgery. Product is a cataract surgery knife designed for optimal tunnel architecture. | Knife, Opthalmic | 2 | April 05, 2011 | 2011 | Beaver-Visitec International Inc. | Loose blades in the handle of knives due to uncured epoxy. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beaver Visitec International, Inc.sent an Urgent Recall Notification letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product and to complete a Customer Response Form (to be returned via fax to 866-906-4304.For questions regarding this recall call 781-906-7917. |  | 1680 | 1680 units (1130 domestic, 550 foreign) | Worldwide Distribution - USA including AZ, MA, MI, MO, NC, NJ, NY, OH, OR, PA, RI, TN, TX, VA and the country of China | Terminated | June 14, 2012 | 436 | 1680 | Not_Computer
Z-1886-2011 | 58107 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Harmony LA/LED Monitor Arm Adapter. Model Number: LB61The Harmony LA/LED Monitor Arm Adapter is an accessory to the Harmony LA and LED Surgical Lighting System and is attached to the end of the monitor suspension arm to allow Customers to attach other medical devices, such as non-STERIS monitors, camera or other compatible items in place of the STERIS monitor. The Harmony Surgical Lighting and Media System is a family of variable pattern, variable intensity surgical lights designed to provide visible immunization of the surgical field or the patient and to provide audio-visual procedural support for the O.R. staff. The device is intended to be used by surgeons and other medical care practitioners in a surgical setting | Harmony LA/LED Arm Adapter | 2 | April 05, 2011 | 2011 | Steris Corporation | The arm adapter does not contain labelling specifying the applicable weight and moment limitations of the arm adapter | OTHER/UNDETERMINED: Pending | Steris Corporation sent an Urgent Field Correction Notice letter dated February 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Steris representative would contact customers to arrange for replacement of the product. For questions regarding the firms visit to their facility, customers were instructed to contact Steris Field Service Dispatch at 1-800-333-8828.For any questions regarding this matter call 440-392-7705. |  | 14 | 14 Units | Worldwide Distribution - USA including CO, IL, IN, MO, NC, NV, NY and OH and the countrries of Canada, China, and Australia. | Terminated | March 13, 2013 | 708 | 14 | Not_Computer
Z-1868-2011 | 57939 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Curlin 6000 IOD Ambulatory Infusion PumpProvides a means for the volumetric delivery of fluids used in Parenteral, Enteral, Epidural, Subcutaneous and Intravenous applications. | Ambulatory Infusion Pump | 1 | April 05, 2011 | 2011 | Zevex International, Inc. | Infusion pumps may exhibit false error codes which render the device inoperable until a recovery process is performed. | DESIGN: Software Design | The firm, MOOG, sent an "URGENT DEVICE RECALL NOTIFICATION" letter dated February 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following:1) Identify the affected Pumps:a) If they have an affected pump, fill out and return the enclosed Business Reply Card within five (5) working days.b) If their pump is not affected by this recall, the pump can be used for its intended purposes.2) Removal of affected Product:a) remove affected products from service and use non-affected pumpb) If you do not have an alternate pump available for use, MOOG recommended that they weigh the risks and benefits for delivering medications using the affected pump versus the risk of not using any pump for each patientWARNING: Affected pumps should not be used to deliver critical medications where a delay in therapy could result in serious adverse health effects, injury or death; and affected pumps should not be used with the neonatal population. 3)Should they Have an Affected pump which exhibits Error Code 45 or is used in therapeutic situations specified in the warning, please contact MOOG Customer Service at (800) 970-2337, Monday thru Friday, 7:00am to 5:00pm MST to arrange for a loaner pump and the return of your affected pump.If you have any additional questions regarding this recall, please contact MOOG's Customers Service Department at (800) 970-2337. | Contact the recalling firm for information | 3407 | 3,407 units | Worldwide distribution: USA and countries including: Australia, Canada, Costa Rica, Croatia, Dubai, France, Germany, Ireland, Italy, Japan, Jordan, Kuwait, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Taiwan, Turkey, UAE, United Kingdom. | Terminated | August 06, 2012 | 489 | 18306 | Software
Z-1900-2011 | 51497 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Siemens brand ACUSON S2000 Ultrasound System with software versions:VA10, VA10A, VA10B, VA10C, VA15; Model Number: 10041461;Product is manufactured and distributed by Siemens Medical Solutions USA, Inc. Ultrasound Business Unit, Mountain View, CAThe ACUSON SC2000 ultrasound imaging system is intended for the following applications: Cardiac, Neo-natal and Fetal cardiac, Pediatric, Transesophageal, Adult Cephalic, Peripheral Vessel, Intraoperative Neurological, Musculo-skeletal Conventional, and Musculo-skeletal Superficial applications. | ACUSON S2000 Ultrasound System | 2 | April 06, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. Should you have any questions, please call (650) 694-5398. |  |  |  | Worldwide distribution. | Terminated | April 06, 2011 | 0 |  | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1828-2011 | 52193 | cabinet x-ray, industrial | RCE |  |  |  |  | Phoenix x/ray nanotom series.Phoenix x¦ray nanotom series cabinet x-ray systems are used in non-destructive testing, failure analysis, process control, and metrology. | Phoenix x/ray nanotom series | 2 | April 06, 2011 | 2011 | GE Inspection Technologies, LP | Some of the systems do not have a safety interlock on the service access panel. Although the systems have two independent means for indicating when x-rays are being generated, some systems may have one indicator that is not discernible. Also, some systems lack proper placement of a caution label and others do not have this label affixed. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | The firm has completed all of the following necessary actions and no further actions have been indicated:A Field Service Memo was issued to customers. It described the affected product and described the issue. GE Sensing and Inspection provided and installed retrofit kits, and the systems in question were tested by qualified Field Service Engineers at no charge to customers. GE asked customers to post the notification letter on the affected equipment until retrofit was completed. Questions were to be directed to the GE Service Center at 800-766-6062. |  | 118 | 118 units total | Nationwide Distribution | Open |  |  | 118 | Not_Computer
Z-1898-2011 | 54684 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Device brochure, labeled in part : BIO -RAD***L70242709 VARIANT TM II TURBO***Hemoglobin A 1C***Cartridge Set, Cartridge Set, Analytical Cartridge, Guard Cartridge Insert***UNITED STATES, Bio-Rad Laboratories, Inc. Hercules, CA 94547***"Integrated method for sample preparation, separation and the percent determination of percent hemoglobin A 1 c in EDTA human whole blood. | Variant II Turbo | 2 | April 06, 2011 | 2011 | Bio-Rad Laboratories Inc | Wrong lot number was found on the Hemoglobin Cartridge Insert . The insert did not match the lot number to the device cartridge or the device cartridge label or on the Update Kit CD. | PRODUCTION CONTROLS: Process Control | The firm, Bio-Rad, sent a "MEDICAL DEVICE CORRECTION" letter dated November 29, 2009 to all customers by Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to do the following:1) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 80080, please destroy the insert and use the new insert enclosed with the letter.2) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 90170, they may continue using the insert.3) Complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to: Bio-Rad CSD Regulatory Affairs Department at (510) 741-3954.If you have any questions, please call 510-724-7000 or contact your regional Bio-Rad office. | Contact the recalling firm for information | 51 | 51 kits (Lot # 70292242; 45 kits( Lot # 70292318) | Worldwide distribution: USA including states of: CA, CT, GA, IL, IN, MA, ME, MS, NH, NV, OH, PA, TX, VA and WV; and countries including: France, Hong Kong, Italy, Korea and New Zealand. | Terminated | June 29, 2011 | 84 | 51 | Not_Computer
Z-1891-2011 | 57972 | mask, oxygen | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal F2 Anesthesia Breathing Circuit, REF 5D440-61X-17, Manufacturer: King Systems, Nobleville, IN. An oxygen mask is a device placed over a patient's nose, mouth, or tracheostomy to administer oxygen or aerosols. | Universal F2 Anesthesia Breathing Circuit | 2 | April 06, 2011 | 2011 | King Systems Corp. | The firm is conducting a subrecall of facetent masks because the masks could have a sticky substance on the inside or outside of the mask that may manifest itself 6-8 months after the manufacture date of the product. The sticky substance could result in inconvenience or in some remote cases cause patient harm. | OTHER/UNDETERMINED: Pending | The firm issued an URGENT PRODUCT RECALL letter dated 1/24/2011. The letter identified the affected product and stated the reason for recall. The letter instructed consignees to cease all sales, to immediately examine their inventory, and to quarantine the product. The letter included a response form that is to be completed and returned. Distributors were instructed to inform their customers of the recall. King Systems Customer Service Department should be contacted at 1-800-642-5464 in order to receive a Return Authorization Number. The letter stated that consignees should return the recalled products to King Systems and the firm will replace the products or provide an acceptable substitute. Customers should contact Customer Service if they have any questions regarding the recall and the letter. |  | 416 | 416 cases of 30 per case 12480 units | Nationwide Distribution -- NM, GA, TX, FL, IL, VA and LA. | Terminated | August 14, 2012 | 496 | 2042 | Not_Computer
Z-1899-2011 | 57855 | device, heat-sealing | KSD | Heat-sealing device. | Hematology | Hematology | 510(K) Exempt | 2380 MINI Hand-Held Tube Sealer Power Source(including a 23802000 Battery Pack)The Model 2380 Power Source is an instrument that generates RF power for making seals on tubing made of RF reactive thermoplastic materials typically used in blood banks, blood processing facilities and transfusion centers. | Haemonetics 2380 MINI HandHeld Tube Sealer Power Source | 2 | April 06, 2011 | 2011 | Haemonetics Corporation | Battery pack may have exposed wire becoming hot | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Haemonetics, issued a "Field Notification: 2380 Sealer battery pack" letter dated January 28, 2011 to its customers. An international English version of this letter (CL101020-IE) has been released to foreign Distribution countries. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Record the manufacturing date of battery, equipment model, serial number, and photo of the affected battery.2) Report the defect through the Haemonetics PIR system by emailing: sebra2380@haemonetics.com with all of the information listed including the appropriate contact information for their company.3) Recycle or dispose of the nickel-cadmium (NiCad) battery pack in accordance with the local, state, federal, or country specific regulations. Under no circumstances should the battery pack be incinerated as this can cause explosion and personal injury. The operation manual will be updated to provide additional information about the proper installation and removal of the battery. This updated manual will be distributed with new devices upon release.Should you have any questions or concerns regarding this matter, please contact the firm at sebra2380@haemonetics.com. |  | 50031 | 50,031 units | Worldwide distribution: USA and countries including: Australia, Austria, AEU, Belgium, Bangladesh, Canada, China, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Korea, Malaysia, Mexico, New Zealand, Pakistan, Philippines, Singapore, Sri Lanka, Taiwan and Turkey. | Terminated | March 06, 2012 | 335 | 50031 | Not_Computer
Z-1903-2011 | 58241 | needle, suturing, disposable | GAB | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | REF 906764, CURVTEK EYE NEEDLE, 12 MM LARGE, STAINLESS STEEL GAMMA STERILIZED, PKG/3, BIOMET SPORTS MEDICINE, 56 EAST BELL DRIVE, P.O. BOX 587, WARSAW, IN 46581 USA, STERILEAfter the Curvtek device is used to drill a "C" shaped tunnel through the patient's bone, the device is used to pass suture through the tunnel so that soft tissue may be tied down directly onto the bone. | CURVTEK EYE NEEDLE, 12 MM LARGE | 2 | April 06, 2011 | 2011 | Biomet, Inc. | The firm initiated a recall due to the discovery that the devices were not sterilized, as required, prior to distribution. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated March 10, 2011 to all customers. The letter described the product, problem and actions to be taken. The letter noted that -Use of a non-sterile Curvtek Eye Needle may lead to patient infection, possibly requiring a revision procedure. The customers were instructed to immediately locate, discontinue use of the product , remove the product from circulation and return the product to Biomet; carefully follow the instructions on the enclosed FAX BACK RESPONSE FORM; complete and return the response form via fax to 574-372-1683 prior to return of product; use priority carrier for shipment, and please confirm receipt of this notice by calling 800-348-9500, extension 3755 or 3756. The firm also informed physicians that had previously utilized the device to monitor patients for infection and provide treatment as appropriate. Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8am to 5pm. | No consumer action necessary | 35 | 35 | Worldwide distribution: USA including states of: IN and KY; and countries of: Canada and The Netherlands. | Terminated | July 27, 2011 | 112 | 35 | Not_Computer
Z-1901-2011 | 58208 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prismaflex System, Catalog Number 107493, Indicated for continuous solute and/or fluid removal in patients with acute renal failure or fluid overload. | Prismaflex | 2 | April 06, 2011 | 2011 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). |  | 1300 | 1300 units | Nationwide distribution. | Terminated | September 01, 2011 | 148 | 1300 | Other |  | Instructions | Safety Notice/Insructions
Z-1904-2011 | 58099 | compressor, cardiac, external | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Auto Pulse Resuscitation System; Zoll Circulation Inc., Sunnyvale, CA 94085Cardiac Compressor used as an adjunct to manual CPR. Use of AutoPulse is intended to reduce the impact of rescuer fatigue and enable rescuer to address other patient needs. | Zoll Auto Pulse Resuscitation System | 3 | April 07, 2011 | 2011 | Zoll Circulation, Inc. | Complaints of batteries discharging prematurely and that during use the Auto Pulse Units would cease operating (providing compressions). | PRODUCTION CONTROLS: Process Control | The firm, Zoll Circulation Inc., initiated a Corrective and Preventive Action Report on January 28, 2009. According to the firm, the customers were contacted by telephone calls and were provided with a Battery Maintenance program. For further questions you may call (408) 541-2140 ext.341. |  |  | not provided | Nationwide Distribution -- including, VA, NY, OH, IN, MO, UT, FL, NC, ME, WA, SD, NM, PA, MS, ID, TX, GA, CA, CO, IL, AR, WV, TN, OR, and KS. | Terminated | April 28, 2011 | 21 |  | Battery | Device Operation | Repair | Repair
Z-1905-2011 | 58112 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA versions 8.0 and 8.1Varian Medical Systems, Palo Alto, CA 94304The ARIA Radiation Oncology product is a treatment plan and image management application. It enables the authorized user to enter, access, modify, store and archive treatment plan and image data from diagnostic studies, treatment planning, simulation, plan verification and treatment. ARIA Radiation Oncology also stores the treatment histories including dose delivered to defined sites and provided tools to verify performed treatments. ARIA also includes optional tools to manage the oncology department's clinical and business information. | ARIA versions 8.0 and 8.1 | 2 | April 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified in the ARIA version 8.0 and 8.1, when an image sequence template is added to a field in RT chart and a CIAO (Completely Irradiated Aperture Opening) is calculated, the resulting CIAO display at 4DITC may be incorrect. | OTHER/UNDETERMINED: Pending | Varian Medical Systems sent an Urgent Medical Device Correction letter dated February 18, 2011, to all affected customers via Federal Express.. The letter identified the product, the problem, and the action to be taken by the customer. Varian has made design changes to address the issue in ARIA Radiation Oncology versions 8.2 and later. A corrective action will be undertaken to upgrade all affected items.For questions regarding this recall call US and Canada 1-888-VARIAN5 (888-827-4265.Europe +41 41 749 8844. |  | 645 | 645 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Azerbaijan, Bahrain, Bangladesh, Barius, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, HOng Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Lithania, Malaysia, Maurtama, Morocco, Netherlands, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | June 19, 2012 | 439 | 645 | Computer
Z-1915-2011 | 56977 | lift, patient, ac-powered | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Penner Manufacturing Transfer Electric Bariatric with Scale, Patient Transfer/Lift System, Cascade Transfer, Models 383500-1 (right entry with scale). The firm name on the label is Penner Manufacturing Inc., Aurora, NE.For use with bathing systems in nursing homes, hospitals, and assisted living facilities to transfer or lift patients under the direct supervision of trained staff. | Penner Patient Transfer/Lift System | 2 | April 08, 2011 | 2011 | Penner Mfg Inc | Incomplete weld on the pillar mounting bracket | TRAINING: Employee Error | The firm, Penner, sent an undated letter flagged "Important Safety Inspection: Action Required" via certified mail on January 19, 2009 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect the Penner Patient Transfer for any missing Welds of the Pillar Mounting Bracket; inspect the areas indicated in the enclosed photos to confirm the bracket is welded on all edges to the main frame of the base; complete the inspection, and complete and return the attached WELD INSPECTION RESULTS OF PENNER PATIENT TRANSFER form via fax to 402 694-5844 or mail to Penner Manufacturing Inc, 102 Grant Street, Aurora, NE 68818, Attn: Quality Assurance Manager. Penner stated in the letter, if welds are not complete (no weld), discontinue use and contact the Quality Assurance Manager at Penner Manufacturing immediately at 402-694-5003.If you have any questions, contact the Quality Assurance Manager at 402-694-5003. |  | 20 | 20 lifts | Worldwide distribution: USA and countries including: Canada and Mexico. | Terminated | April 08, 2011 | 0 | 2290475 | Not_Computer
Z-1929-2011 | 57288 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GEThe CARESCAPE Monitor B850 includes both new and existing subsystems interconnected to form a high acuity, multi-parameter patient monitoring system. A typical configuration would be a CARESCAPE Monitor B850 host processing unit running the CARESCAPE Monitoring platform software, a display with integrated keyboard and a frame for the insertion of parameter measurement modules. A variety of options are available to the customer including additional displays, various input devices (keyboard, mouse, bar code reader), additional modules and frames and software options. The new subsystems include a host processing unit (CPU), a 15 inch LCD display with an integrated keypad, 19 inch LCD display with an integrated keypad and touch panel interface, a five slot parameter module frame, a seven slot parameter module frame, a cabled remote control, a cabled remote keypad, and the CARESCAPE Monitoring platform software. Some of these new major subsystems include non-patient contact accessory items (e.g. cables and mounting hardware). The CARESCAPE Monitor B850 interfaces to a variety of existing physiological parameter measurement modules. In addition, the CARESCAPE Monitor B850 interfaces to a variety of existing OEM medical devices via the existing network infrastructure. The CARESCAPE Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. The CARESCAPE Monitor B850 is intended for use on adult, pediatric, and neonatal patients and on one patient at a time. The CARESCAPE Monitor B850 system is indicated for monitoring of Hemodynamic (including ECG, ST Segment, Arrhythmia Detection, ECG Diagnostic Analysis and Measurement, Invasive Pressure, Noninvasive Blood Pressure, Pulse Oximetry, Cardiac Output, Temperature, Impedance Respiration and SpO2 (Mixed Venous Oxygen Saturation)), Airway Gases (Fi/Et C02, 02, N20 and Anesthetic Agent), Spirometry, Gas Exchange (02 Consumption (VO2), C02 production (VCO2), energy expenditure (BE), and respiratory quotient (RQ)) and neurophysiological (including electroencephalography (EEG), Entropy, Bispectral Index (BIS) and Neuromuscular Transmission (NMT) Monitoring) status. The CARESCAPE Monitor B850 provides alarms, trends, snapshots and events, and calculations and can be connected to displays, printers and recording devices. The CARESCAPE Monitor B850 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for bed to bed viewing and to data management software devices via a network. The CARESCAPE Monitor B850 is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a professional healthcare facility. In addition to the healthcare practitioner, the CARESCAPE Monitor B850 is designed to provide configuration and troubleshooting capabilities to qualified service personnel. The CARESCAPE Monitor B850 is not intended for use during MRI. | GE CARESCAPE Monitor B850 | 2 | April 08, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential safety issues associated with the CARESCAPE" Monitor B850 when Patient Data Module (PDM), Tram Acquisition Module (Tram), Patient Side Module (E-PSM or E-PSMP), E- MASIMO or E-(P)RE(S)TN modules or Bed-to-bed alarming feature is used.1. When CARESCAPE Monitor B850 is used with the above listed modules, some SpO2 alarms may be delayed or missed.2. Whe | DESIGN: Software Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the recall product the problem, Safety Issues, Affected Product details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to follow the safety instructions provided in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available.If you have any questions regarding this medical device correction or the identification or affected items please contact Technical Support at 1-800-558-7044, or your local Service Representative. | Contact the recalling firm for information | 3817 | 3817 | Worldwide distribution: USA including states of:AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA,WA, WI, and WV; and countries including: JERUSALEM, UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLAND, MALTA, MALAYSIA, LEBANON, KUWAIT, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CZECH REPUBLIC, CHILE , CANADA, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | June 05, 2012 | 424 | 3817 | Software | Alarm/Message | Software update | Software Update
Z-1925-2011 | 57866 | enzyme immunoassay, inhibin-a | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA kit, Part Number: DSL-10-28100-4The Inhibin A ELISA kit is intended for the quantitative measurement of Dimeric Inhibin A in human serum or plasma. It is intended strictly for in vitro use as an aid in the diagnosis and monitoring of various hormonal reproductive disorders. | Inhibin A ELISA Kit | 3 | April 08, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the lots of Inhibin A identified above may contain microbial contamination in the conjugate diluent bottlesincluded in the kit. |  | Beckman Coulter sent an Urgent Product Corrective Action letter dated December 21, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product and discard all remaining inventory. In addition customers were requested to share the information with their laboratory staff and retain the notification as part of their laboratoryQuality System documentation. Customers were instructed to complete and return the enclosed response form within 10 days soBeckman could be assured that customers received this important notification.For product replacement:1. In the United States, please contact Customer Service at 1-800-526-3821, option 1.2. In Canada, please contact Customer Service at 1-800-463-7828.3. Outside of the United States and Canada, contact your local Beckman Coulter representative.Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. For those customers outside the United States and Canada, instructions were given to contact their local Beckman Coulter representative. |  | 740 | 740 total units for both products | Worldwide Distribution - USA including NY, IN, PA, NM, OH, TX, AZ, NC, TN, MI and CT and the countries Denmark, Germany, Spain, Turkey, United Arab Emirates, United Kingdom. | Terminated | August 31, 2012 | 511 | 740 | Not_Computer
Z-1907-2011 | 58006 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 8V3c transducer used with the ACUSON SC2000 ultrasound system. 8V3c transducer model number 10433823; ACUSON SC2000 model number 10433816.Siemens Medical Solutions USA, Inc.Business Unit Ultrasound1230 Shorebird Way,Mountain View, CA 94043High-density array transducer for pediatric and fetal echocardiography applications. | 8V3c transducer | 2 | April 08, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | When using the 8V3c transducer in color mode, the derated Ispta 3 (Spatial Peak-Temporal Average Intensity) may exceed the limit stated in the instructions for use. | DESIGN: Software Design | A Customer Safety Advisory Notification was initiated on 2/4/2011 and sent via certified mail. The letter identified the affected product and described when the problem occurs along with the associated potential risk. The firm recommends that users not use the 8V3c until the software issue is fixed. The firm is working to release a software update and customers' local Customer Service Engineer will install new software on their system when it is available. Customers should direct their questions to their local service support person. The information in the letter should be shared with all personnel within the organization who need to be aware. | Contact the recalling firm for information | 386 | 386 units | Worldwide Distribution. | Terminated | October 04, 2011 | 179 | 386 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1909-2011 | 58093 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | SPIDER Surgical Platform The SPIDER is a single port, single incision device to facilitate multi-instrument access during laparoscopic surgical procedures. The SPIDER device is provided pre-sterilized and is a single use device. | SPIDER Surgical Platform | 2 | April 08, 2011 | 2011 | Transenterix, Inc | Plastic shards may break off and fall into the abdominal space which may require retrieval during surgery. | OTHER/UNDETERMINED: Pending | TransEnterix sent an Urgent Device Recall letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the lot numbers affected, complete the attached inventory reconciliation form and return the form, along with any affected product toTransEnterix using the enclosed shipping label. Customers were asked to return the inventory form even if they had none of the lot numbers affected in their inventory.Their sales representative and customer service will work with them to replace the product in full at not cost to them.For questions regarding this recall contact TransEnterix Customer Care at 1-888-879-4111. | Contact the recalling firm for information | 305 | 305 units | Nationwide Distribution including AZ, FL, NC, NY, OH, OR, SC, TX, VA and WA | Terminated | November 07, 2011 | 213 | 397 | Not_Computer
Z-1928-2011 | 58014 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilator with software version V2.00x and Neonatal Option or V2.01x and Neonatal software. Hamilton-G5: catalog number 159001;Neonatal Software catalog number: 159700; Neonatal Option catalog number: 159187. Hamilton MedicalIntended for use in the hospital and institutional environment where healthcare professionals provide patient care. | HamiltonG5 Ventilator | 2 | April 08, 2011 | 2011 | Hamilton Medical Inc | The adaptive volume/pressure regulation can be automatically and permanently reduced to PEEP + 3 mbar when a closed suctioning maneuver is performed on the patient and may be undetected by the operator. | DESIGN: Software Design | A Medical device Safety Alert and Corrective Action along with a cover letter, was sent to all locations on February 16, 2011. The letter identified the affected product along with the reason for recall. The letter states that operators deactivate the "high tidal volume" alarm as described in appendix K.4 of the operator's manual until the ventilator is upgraded with a corrected software. Distributors will upgrade the affected devices with the corrected software as soon as possible. Questions can be directed to 800-426-6331, ext 215. |  | 97 | 97 units | Nationwide Distribution -- MA, FL, GA, NC, MI, NC, MO, TX, and WV. | Terminated | March 21, 2012 | 348 | 97 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1930-2011 | 58275 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650The CARESCAPE Monitor B650 is a multi-parameter patient monitor intended for use in multiple areas and intrahospital transport within a professional healthcare facility. The CARESCAPE Monitor B650 is intended for use on adult, pediatric, and neonatal patients and on one patient at a time. The CARESCAPE Monitor B650 is indicated for monitoring and recording of, and to generate alarms for, hemodynamic (including EGG, ST segment, arrhythmia detection, EGG diagnostic analysis and measurement, invasive pressure, non-invasive blood pressure, pulse oximetry, cardiac output, temperature and mixed venous oxygen saturation), impedance respiration, airway gases (C02, 02, N20 and anesthetic agents), spirometry, gas exchange, and neurophysiological (including electroencephalography, Entropy, Bispectral Index (BIS) and neuromuscular transmission) status. The CARESCAPE Monitor B650 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for remote viewing and to data management software devices via a network. The CARESCAPE Monitor B650 is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a professional healthcare facility. The CARESCAPE Monitor B650 is not intended for use during MRI. | CARESCAPE Monitor B650 | 2 | April 08, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue associated with the CARESCAPE Monitor B650 & B850 Cardiac Output measurement using Patient Data Module (PDM), and activation of ECG Leads Off alarm in specific cases.1. When the CARESCAPE Monitor is used to measure Cardiac Output values with the Patient Data Module (PDM) or TRAM Module, Cardiac Output values may be calculated incorrectl | DESIGN: Software Design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated March 23, 2011 to its consignees/customers. The letters were addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letters described the Safety Issues, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. The customers were instructed to follow the safety instructions documented in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available and will contact the customer to schedule the software update.If you have any questions regarding this medical device correction or the identification of affected items, please contact Technical Support at 1-800-558-7044, or your local Service Representative. | No consumer action necessary | 1443 | 1443 | AL, AZ, AR, CA, CO, CT, FL,GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA,PR, SC, TN,TX, UT, VA, WA, WV, WI.UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SWITZERLAND, SWEDEN SLOVENIA, SLOVAKIA, SPAIN, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NORWAY, NORTH KOREA, NEW ZEALAND, NETHERLANDS, MALTA, MALAYSIA, LITHUANIA, LEBANON, KUWAIT, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, DENMARK, ESTONIA, CZECH REPUBLIC, BELGIUM BAHRAIN, COLOMBIA, CHILE CANADA, BRAZIL, AUSTRIA, AUSTRALIA. | Terminated | February 04, 2012 | 302 | 8345 | Computer
Z-1927-2011 | 58143 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Inflow/Outflow Cannula;Qty: 1;distributed by Stryker Endoscopy, 5900 Optical Court, San Jose, CA 95138The inflow/outflow cannula is intended as a general purpose, concomitant suction and irrigation device for use in endoscopic surgeries. It will be used by orthopedic surgeons in arthroscopic knee and shoulder procedures as well as used in conjunction with the Stryker Flow Control Arthroscopy pump and with all 502-477-xxx scopes. The design included dual scope sheaths made of metal that offer two important features. First, it will separate the inflow irrigation to the operative site from the outflow irrigation out of the operative site. Second, the cannula will have a dual stopcock design that will separately control the inflow and outflow irrigation ports. | Inflow/Outflow Cannula | 2 | April 08, 2011 | 2011 | Stryker Endoscopy | Insufficient cleaning instruction were included for the Inflow/outflow Cannula in the user manual. | OTHER/UNDETERMINED: Pending | The firm, Stryker, issued an "URGENT: DEVICE CORRECTION" memo dated March 8, 2011 to all customers. The memo described the product, problem and actions to be taken. The customers were instructed to discard any old Instructions for Use (1000-400-816 rev. A or B) for their Inflow/Outflow Cannual; download the new IFU (1000-400-816 rev. C) from the URL link: http://stryker.com/InflowOutflowCannual; ensure they follow the new cleaning instructions provided in the IFU and complete and return the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in and email to cannula@stryker.com. Stryker stated in the memo that "It is important to send the self addressed confirmation forms back to Stryker RAQA". If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at cannula@stryker.com. | Contact the recalling firm for information | 2618 | 2618 devices | Worldwide distribution: USA including states of: AL, ALASKA, AK, AR, CA, CO, CT, DC, DE, FL, GA, ID IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: Australia, Canada, China, France, Germany, India, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Switzerland and United Kingdom. | Terminated | October 04, 2011 | 179 | 2618 | Not_Computer
Z-1924-2011 | 57141 | heat-exchanger, cardiopulmonary bypass | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Quest MPS and MPS 2 Consoles, Quest Medical, Inc., An Atrion company. Single software controlled device that incorporates a pump, a heater/cooler, temperature monitoring, pressure monitoring, a heat exchanger, an arrest agent pump and an additive pump. The MPS console is used with the MPS sterile disposables with a built in heat exchanger. The device is used with on-pump or off-pump cardiac surgery. | MPS and MPS 2 Consoles | 3 | April 08, 2011 | 2011 | Quest Medical, Inc | User update regarding installation of external heater-Cooler units with MPS and MPS2 consoles. | PRODUCTION CONTROLS: Equipment Maintenance | Quest Medical Incorporated sent a "User Update" letter in July 2010 to consignees providing information concerning use of external heater-cooler untis with their device - Myocardial Protection System ( MPS) consoles. If you are installing an external heater-water cooler unit or have questions about this user update, please contact the Field Service department at Quest Medical, Inc. The number is 800.627.0026 or 1 972.390.9800. |  | 771 | 771 consoles | Nationwide, Canada, Saudi Arabia, Kuwait, Japan, and Italy | Terminated | April 26, 2011 | 18 | 771 | Not_Computer
Z-1932-2011 | 54483 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CAThe Intended Use is as follows: The Bior-Rad II TURBO Hemoglobin A1c Program is intended for the percent determination of hemoglobin Alc in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). The Bio-Rad VARIANT II. TURBO Hemoglobin Alc Program is intended for Professional Use Only--For In Vitro Diagnostic Use. Reagent Kit Information: The VARIANT II TURBO Hemoglobin A1c program utilizes principles of ionÂŽexchange high performance liquid chromatography (HPLC). The VARIANT II Hemoglobin A1c Program is based on chromatographic separation of HbAlc on a cation exchange cartridge. System Overview: The VARIANT II TURBO Hemoglobin testing System provides an integrated method for sample preparation, separation and the percent determination of percent hemoglobin Alc in EDTA human whole blood. Technology Overview: The V ARIANT II TURBO Hemoglobin Testing is a fully automated, high-ÂŽthroughput hemoglobin analyzer. It consists of two modules- the VARIANT II Chromatographic Station (VCS) and the VARIANT II Sampling Station (VSS). In addition, a personal computer is used to control the VARIANT II TURBO System using Clinical Data Management (CDM) software. CDM is a dedicated software package of instrument control, data control, results analysis, calibration, quality control and service software. | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program | 2 | April 11, 2011 | 2011 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.Collect a Customer Medical Device Correction Response Form from every customer.Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).For questions regarding this recall call 510-724-7000. |  | 1099 | 1099 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Taiwan, Singapore, India, China, S. Korea, Hong Kong, France, Sweden, United Kingdom, Italy, Hungary, Germany, Brazil, Mexico and Colombia. | Terminated | August 30, 2011 | 141 | 1099 | Not_Computer
Z-1943-2011 | 57189 | sensor, blood-gas, in-line, cardiopulmonary bypass | DTY | Cardiopulmonary bypass in-line blood gas sensor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems Corporations, CDI 101 Hematocrit/Oxygen Saturation Monitoring System. In-line blood-gas sensor used in cardiopulmonary bypass. | CDI 101 | 2 | April 11, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Inaccurate result. Possibility for displayed values for potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other measured blood parameters to be inaccurate. | DESIGN: Device Design | On 11/1/2010, Terumo sent "URGENT MEDICAL DEVICE CORRECTION" notices to Chief Perfusionist, Department of Cardiovascular Surgery or Director of Operating Room Services. The letter informed the consignees about the incidents of the CDI 100 and 101 displaying inaccurate values. The letter states that when a displayed value is significantly different from expectations based on the clinical situation, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operators manual. | Contact the recalling firm for information | 468 | 468 monitors | US, India, Colombia, Taiwan, Venezuela, Vietnam, Puerto Rico, Malaysia, Sri Lanka, Brazil, Indonesia, Pakistan, Chile, Germany, Hong Kong, Japan, AE, Singapore, Austria, Belgium and Thailand. | Terminated | April 17, 2012 | 372 | 1500 | Not_Computer
Z-1938-2011 | 57952 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxI 800 Access ImmunoassaySystems; Part Number(s): DxI 800: 973100, DxI 800 RemanufacturedA25288, DxI 800 with spot B A71456, DxI 800Reconditioned A25285The DxI Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | UniCel DxI 800 Access Immunoassay Systems | 2 | April 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated due to Beckman Coulter receiving a report of results that were associated with incorrect patientidentifications when two sample racks advanced into the sample presentation area in one push. The firstrack of the two was not scanned. This is referred to as a "pushed pair." If a pushed pair event occurs, results may be associated with the incorrect patient identification. |  | Beckman Coulter sent an Urgent: Field Safety Notice Product Correction letter dated January 26, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. To prevent a pushed pair, be sure your sample racks are free of debris prior to use. Do not applyadditional labels to the sample racks. Extra labels can obscure the edges of a sample rack, orobstruct the gap between two racks. Be sure sample labels do not extend beyond the edges of asample rack;2. If a pushed pair occurs on the UniCel DxI system, the system generates a yellow Event Log message- Clear the presentation tray and press ROUTINE button. This Event Log message is associatedwith a variety of common Sample Presentation Unit (SPU) errors. For that reason, you must reviewthe Event Log details to confirm the pushed pair event has not occurred;3. Detach the Appendix included with this notification and store it near your UniCel DxI system. Use theAppendix as a reference if you suspect a pushed pair event has occurred on your system.Beckman Coulter also requested their customers share the recall information with their laboratory staff, and retain pages 1 and 2 of this notification as part of your laboratory Quality System documentation.For questions regarding ththe recall notification, customers were instructed to contact Beckman Coulter Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support, or call1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact theirlocal Beckman Coulter Representative. | Contact the recalling firm for information | 2639 | 2639 units (892 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | June 04, 2012 | 420 | 2639 | Not_Computer
Z-1933-2011 | 57953 | bed, hydraulic, adjustable hospital | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Adjustable (retractable) Strap and Handle, type TGN-R100GR; a component of a Lifting Pole, an accessory to long term care hospital beds such as ArjoHuntleigh's Minuet 2 Community bed and Enterprise series hospital beds; Distributed by ArjoHuntleigh, a Branch of Arjo Limited Med AB, Wednesbury, United Kingdom Lifting poles are accessories intended to assist patients in positioning themselves on a hospital bed and the lifting pole is used in conjunction with a strap and handle. The strap and handle hangs from the lifting pole and the patient grasps the handle to pull upon the strap when repositioning themselves | Rolko Adjustable Strap and Handle | 2 | April 11, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | Part of the housing of the strap retracting mechanism built into the handle may break while the patient is using the handle and strap. | OTHER/UNDETERMINED: Pending | ArjoHuntleigh sent an Urgent Field Safety Notice to all affected end users via UPS Ground on February 24, 2011. The notice informed the accounts of the potential for the strap retracting mechanism to break during use. The accounts were requested to check their inventory for the presence of adjustable straps and handles procured from ArjoHuntleigh, note if they have a red button on the retractor housing and the affected production dates as explained in the photographs in Figures 1 and 2 in the letter. Any straps and handles found as identified in Figures 1 and 2 should be removed from use and returned to ArjoHuntleigh. Any straps and handles with a yellow button design as shown in Figure 3 are not affected and can continue being used. The accounts were requested to complete and fax or mail the enclosed Customer Response Form to Arjo at 630-576-5012, indicating the number of strap and handles checked and how many were removed from use. Questions were directed to the ArjoHuntleigh Regulatory Department at 1-630-785-4885. | Contact the recalling firm for information | 90 | 90 units | Nationwide Distribution including California, Georgia, Hawaii, Michigan, South Carolina and Texas | Terminated | June 14, 2011 | 64 | 90 | Not_Computer
Z-1941-2011 | 58098 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Transmitter Sensor Cables, Part Numbers: 1002008, 1004587, 1007914-NAV, GE Healthcare Surgery, Salt Lake City, UT 84116.Cables are used with InstraTrak Navigation Systems, Models IT3000, IT2500, IT2500+, IT3500, IT3500+. InstaTrak System is intended as an aid to the surgeon for precisely locating anatomical structures anywhere on the human body during either open or percutaneous procedures. It is indicated for any medical condition that may benefit from the use of stereotactic surgery and which provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, visible on medical images such as CT, MR, or X-ray. | Sensor cables | 2 | April 11, 2011 | 2011 | GE OEC Medical Systems, Inc | Sensor cables may suffer material degradation when exposed to certain sterilization procedures. | DESIGN: Component Design/Selection | GE OEC Medical Systems, Inc sent an Urgent Recall Notice dated March 1, 2011, by certified letter on February 24, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if their Transmitter Sensor Cables or Receiver Sensor Cables have been exposed to any STERRAD sterilizer other than STERRAD 1005, discontinue use of these sterilizers for this product. The Transmitter Sensor Cables should be inspected for signs of material degradation.Further information is available at 800-874-7378. |  | 708 | 708 cables | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NY, OH, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI and WV and the countries of Austria, Canada, China, Denmark, England, France, Germany, Hong Kong, India, Ireland, Italy, Japan, San Salvador, Slovakia, Spain, and Switzerland. | Terminated | September 09, 2011 | 151 | 708 | Not_Computer
Z-1945-2011 | 58163 | radioimmunoassay, free thyroxine | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Free T4 Calibrator, Part Number: 33885, Lot Numbers: 014769The Access Free T4 Calibrators are intended to calibrate the Access Free T4 assay for the quantitative determination of free thyroxine levels in human serum and plasma (heparin) using the Access Immunoassay Systems | Access Free T4 Calibrator | 2 | April 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the stability claim for Access Free T4 calibrator lot 014769 is not met. The affected T4 calibrator lot can produce quality control and patient results that are erroneously low (by greater than 2 standard deviations). An erroneously low patient result may lead to misdiagnosis and/or delayed treatment. | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter dated November 29, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: ACTION/RESOLUTION:" Discontinue use of Access Free T4 calibrator lot 014769." Review historical quality control performance and determine whether any reported patient results were obtained while using a faulty calibration. If results were affected, you will see a shift downward in control recovery that produces out-of-range low QC results." If patient results were obtained with Access Free T4 calibrator lot number 014769, it is possible they are erroneously low. If any patient results were reported in the period between the time the erroneous calibration with the affected lot was performed and when a different calibrator lot was used, the laboratory should consider advising physicians to whom results were reported so they can evaluate whether diagnosis or treatment decisions should be revisited." No further orders of Access Free T4 Calibrator lot number 014769 will be shipped. Alternate lots will be provided for current and future replenishment orders or requests.Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed RESPONSE FORM via fax to Beckman Coulter at 786-639-4000. For product replacement:" In the United States, please contact Client Services at 1-800-526-3821, option 1." In Canada, please contact Customer Service at 1-800-463-7828." Outside of the United States and Canada, contact local Beckman Coulter representative.Questions and answers regarding this issue were enclosed for consignee/customer reference. Additional questions, were directed to the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customer. | Contact the recalling firm for information | 459 | 459 units | Worldwide distribution: USA and countries including: Argentina, Belgium, Bosnia and Herzegovina, Canada, China, Colombia, Czech Republic, French Polynesia, Hong Kong, Ireland, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Oman, Panama, Philippines, Russian Federation, South Africa, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Venezuela, and Zambia. | Terminated | July 11, 2012 | 457 | 459 | Not_Computer
Z-1961-2011 | 52928 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Auxiliary 2 MHz Continuous Wave (CW) transducer, model number 8251046, when used with ACUSON Sequoia ultrasound systems at SW revision 12.X. Manufactured by Siemens Medical Solutions USA, Inc., Ultrasound Business Unit, Mountain View, CA.The ACUSON Sequoia is a general purpose Diagnostic Ultrasound system. It is intended for the following applications: Abdominal, General Imaging and Cardiology for Fetal, Abdominal, Intraoperative (abdominal and neurological), Pediatrics, Small Organs (breast, testes, thyroid, and penis), Neonatal/Adult Transvaginal, Peripheral Vessels and Musculo-skeletal (superficial and conventional). The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. The Siemens Aux CW transducer is a non-imaging continuous wave transducer indicated for use in transthoracic/cardiac adult and pediatric applications as well as continuous wave spectral Doppler echocardiography. | ACUSON Sequoia ultra sound systems with Auxiliary 2 MHz Continuous Wve (CW) transducer | 2 | April 12, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Overheating. A software bug may cause the transducer to overheat when used with certain transducer models. The transducer can become noticeably hot within a few minutes, which may cause patient discomfort or burns. | DESIGN: Software Design | The firm, SIEMENS, sent a "Customer Safety Advisory Notification" to its customers on 3/21/2009. The notification described the product, problem and actions to be taken. The customers were instructed to disconnect any transducers that may be in an active transducer port and chose another imaging transducer prior to use of the Aux CW transducer. If they do not have another transducer to select other than those listed in the notification, they should close the study in progress and start a new study before using the Aux CW transducer-this procedure should be followed for each new patient and /or exam type. In addition, the customers were instructed to pass this notice on to all those within their organization who need to be aware of this issue until the corrective action is completed. SIEMENS will release a software update that will correct the issue. The customers local service engineer will install the updates on their system as soon as it becomes available.Should you have any questions regarding this notification, you may contact your local service support person and/or call (650) 969-9112. |  | 2591 | 2591 units | Worldwide distribution. | Terminated | March 09, 2012 | 332 | 2591 | Software | Physical Safety Hazards | Instructions to disconnect transducer/Software update | Software Update
Z-1844-2011 | 58076 | laser, ophthalmic | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Solo LT5106-S ophthalmic laser (not sold in USA) | Ellex Solo | 2 | April 12, 2011 | 2011 | Ellex Medical Pty Ltd | Unintended laser emission while the user was moving the joystick for targeting. | PRODUCTION CONTROLS: Process Control | Ellex notified all distributors that have received the affected devices per Global Customer Service Bulletin GSB-11 -01a, dated 25 February 2011. | Contact the recalling firm for information | 193 | 193 units | worldwide distribution | Terminated | October 02, 2012 | 539 | 2335 | Not_Computer
Z-1958-2011 | 58122 | chair, dental, without operative unit | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Dental chairs distributed under several brand names such as DCI Equipment; Marus MaxStar; Marus NuStar; and Pelton & Crane.Part number for the bolt: 022R065Product is labeled in part:DCI Equipment, Mfg By: Dental Equipment, LLC. DC1235, Marus Dental, Mfg By: Dental Equipment, LLC. DC1490, DC1690, and DC1700, and DC1702, Pelton & Crane, Spirit SP17. | DCI Equipment; Marus MaxStar; Marus NuStar; Pelton & Crane Spirit | 2 | April 12, 2011 | 2011 | Dental Equipment LLC | Link arm assembly that mounts the rear mounted accessories to the dental chair may potentially break and fall off the dental chair. | PRODUCTION CONTROLS: Process Control | On 03/08/11, Dental Equipment LLC. started sending out the Medical Device Recall notification letter via USPS 1st class mail to US consignees and Federal Express mail to foreign consignees.The firm also sent a recall notification letter to consignees who only purchased dental accessories such as Rear Mounted Cuspidor with Coverset, Arm Articulating Rear Mounted Cuspidor - Marus Gray, Maxstar Orbit Cuspidor utility W/ Porcelain bowl, and zinc plated bolts.The letter identified the affected dental chair as Marus brand with the following model numbers: DC1490, DC1690, DC1700 and DC1702.The firm's corrective action including the replacement of the affected link arm assembly with a new link arm assembly. The firm's field representative will contact each consignee for an appointment to replace the affected link arm assembly.Consignees can contact the Dental Equipment LLC at 503-537-3617 for any questions. |  | 66 | 55 chairs and 11 replacement bolts distributed to US and foreign consignees outside the US. | Worldwide distribution to USA, Australia, Canada, Ireland, South Africa, and Vietnam. | Terminated | September 21, 2011 | 162 | 66 | Not_Computer
Z-1959-2011 | 58110 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX and FX Plus SystemIntended use: Stationary x-ray system | Siemens Axiom Aristos FX and FX Plus System | 2 | April 12, 2011 | 2011 | Siemens Medical Solutions USA, Inc | During product monitoring, firm discovered a possible hazard to patients when using the detector alpha rotation of all Axiom Aristos FX and FX Plus systems with serial numbers up to and including 1216. When using the detector alpha rotation, the detector may descend unexpectedly during movement or in vertical or table mode. If this occurs in vertical modes, the possibility exists that the detect | OTHER/UNDETERMINED: Pending | The firm, SIEMENS, sent a "Safety Advisory Notice" dated February 15, 2011 to the affected customers via Update Instructions XP016/10/S on February 17, 2011. The notice described the product, problem and actions to be taken. The customers were instructed to immediately instruct their personnel accordingly; ensure that the safety advisory is placed in the systems's instructions for use; if they have sold this device/equipment and it is no longer in their possession, forward this safety notice to the new owner of the device/equipment; inform SIEMENS about the new owner of the device/equipment, and complete and return the Acknowledgement of receipt form. SIEMENS is preparing a modification of AXIOM Aristos FX and FX Plus that will resolve this potential malfunction. The field modification will be available from February 2011 and is expected to be carried out May 31, 2011. If you have any questions, call 09131-18-9908 or +1-888-826-9702. |  | 101 | 101 | Nationwide Distribution: USA including states of: AZ, AK, CA, CO, FL, GA. IL, IN, IA, KS, KY, ME, MI, MN, MS, MO, NE, NV, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, and WI. | Terminated | July 16, 2013 | 826 | 101 | Not_Computer
Z-1957-2011 | 58072 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Varis Vision - RT Chart version 7.3.10. For radiation therapy. | Varis Vision RT Chart | 2 | April 12, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue with the Varis Vision software where treatment can occur without the dynamic Multileaf Collimator component and may deliver a higher than expected dose of radiation. | DESIGN: Software Design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected customers on February 7, 2011. | Contact the recalling firm for information | 421 | 421 | Worldwide Distribution. | Terminated | December 23, 2011 | 255 | 421 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1952-2011 | 57830 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ArthroCare Opus SpeedLock Knotless Fixation Device; Catalog Number: OM-7500. | ArthroCare Opus SpeedLock | 2 | April 12, 2011 | 2011 | ArthroCare Corporation | During surgical procedure the product has the potential for separation of the plug guide. Separation of the plug guide may compromise the ability of the inserter to fully drive the implant completely into the bone, which may also compromise the ability to tension the suture properly. | PRODUCTION CONTROLS: Process Control | On 9/30/2010, Arthrocare notified customers via letters explaining the reason for the recall and requesting the product be returned. |  | 170 | 170 units | Product was released for distribution in US. | Terminated | August 01, 2011 | 111 | 170 | Not_Computer
Z-1953-2011 | 57990 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH750 AnalyzerThe COULTER LH 750 Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The COULTER LH 750 Hematology Analyzer provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids. | COULTER LH750 Analyzer | 2 | April 12, 2011 | 2011 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter identified a problem with some date fields that are transmitted to the Laboratory Information System (LIS). Starting Jan 1, 2011, the date fields will transmit as a blank field to the LIS if the year is greaterthan 2010. The following dates are affected: Run date, Control expiration date, Date of birth(For individuals born in 2011 or later), a | DESIGN: Software Design | The firm, Beckman Coulter, sent an "URGENT: PRODUCTION CORRECTION" letter dated January 31, 2011 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Coulter LH 750 Analyzer, Coulter LH 780 Analyzer, Coulter LH 500 Series System, and Coulter GEN S System. The letter describes the product, problem and actions to be taken. The customers were instructed to work with their US vendor to assess how this affects their laboratory. The actions they choose are dependent on the configuration of their LIS.ACTION FOR ISSUE 1:Use Sample Drawn Date and Date of Birth information available at your LIS in lieu of the analyzer transmitted fields BIRTH and UDATE.Use the current date when the results are received at the LIS in lieu of the transmitted Run Date [DATE]. Print results from the Workstation to keep a copy of the analysis date on the analyzer, if necessary.Use the control expiration date from the analyzer display or printout.ACTION FOR ISSUE 2:If you use the LIS 1G1 transmission format, use the General Information group field names in place of the Retic Parameter group field names to identify each reticulocyte run (see Table in Issue 2 section).Check with your LIS vendor to determine if it is possible to switch to the LH 750 Workstation transmission format. Format requirements are LIS driver specific.For patients with Dates of Birth in 2011 for both Issue 1 and Issue 2, results should be reviewed to determine if decision rules or age-related limits sets were triggered inappropriately.RESOLUTIONThese issues will be corrected in a future software release for the LH Series products. The GENÂŽÂŽS product has been discontinued; no software revisions will be forthcoming for the GENÂŽÂŽS product line. Customers were also instructed to complete and return the PRODUCT CORRECTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 310B06, 11. | Contact the recalling firm for information | 9419 | 9419 (total units distributed in U.S. & foreign countries) | Worldwide distribution: USA (nationwide) and countries including: Algeria, Andorra, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bermuda, Bulgaria, Burundi, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, France, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Monaco, Morrocco, Myanmar, Namibia, Netherlands, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Virgin Islands (US). | Terminated | September 25, 2012 | 532 | 9419 | Software | Output/Calculation | Software update | Software Update
Z-1934-2011 | 57450 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Waste BottleTo store waste material from completed assays. | Liquid Waste Bottle | 2 | April 12, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter confirmed Access waste bottles that ruptured and expelled liquid waste, exposing the operator to the system liquid waste. Over time, waste bottles can degrade, become brittle and potentially crack. |  | The recall communication was initiated on October 5, 2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access immunoassay Systems, UniCel DxC 600i Access, and Immunoassay Systems, Synchron LxI 725 Access Clinical Systems. The letter explains the problem identified, the impact, and actions required to be taken by customers. The letter instructs customers to inspect the waste bottle for signs of aging and deterioration. New bottles are opaque white in color.Signs of deterioration include:o Any color change to the bottle (yellow, orange or pink)o Crackingo Crazing (spider web cracking)Any defect that is not removed by routine cleaning is a sign of deterioration. It is important to routinely inspect the area nearest the handle and along all seams of the bottle. If upon inspection consignees detect any issues with their waste bottles, they were asked to contact Beckman Coulter Customer Service using the enclosed Replacement Order Form to obtain a new waste bottle, and keep this information should they require a replacement bottle in the future.Consignees were also asked to share this information with laboratory staff and retain the notification as part of laboratory Quality System documentation. They should also complete and return the enclosed Product Corrective Action Response Form within 10 days so Beckman Coulter can be certain they received this important notification.Beckman Coulter offers a service kit which may be used to replace the waste bottle and direct liquid waste to a floor drain. Beckman Coulter requires a consultation with consignees technical support representative to determine if their system can be fitted with a floor drain kit. If assistance is needed or consignees have any questions regarding this notification, they were asked to contact Technical Support at 1-800-854-3633 i. | No consumer action necessary | 7700 | 7700 units total (3178 in US) | Worldwide Distribution -- United States, Afghanistan, Algeria, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Congo, Costa Rica, Croatia, Czech Republic, Djibouti, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, Zambia, and Zimbabwe. | Terminated | July 25, 2012 | 470 | 7700 | Not_Computer
Z-1950-2011 | 57815 | electrode, electrodcardiograph, multi-function | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Leads-Out Electrode - Physio-Control, Part 16242. Indicated for use in external pacing, defibrillation and monitoring applications as a non-sterile, disposable device for single patient use only. | Medstorm Adult Radiolucent LeadsOut Electrode | 2 | April 12, 2011 | 2011 | Heart Sync, Inc | The "leads out" connector had bent pins which would not allow connection to the Medtronic Physio Control Defibrillator cable during a product demonstration. | PRODUCTION CONTROLS: Equipment Maintenance | The firm sent a recall notice to their distributors on 12/3/2008 asking them to quarantine product and return it to HeartSync. The letter stated that customers' current inventory would be exchanged for new product. Customers were to call 1-800-828-4681 to provide the necessary information. |  | 2100 | 2100 (210 cases of 10 each) | Nationwide Distribution | Terminated | April 13, 2011 | 1 | 2150 | Not_Computer
Z-1947-2011 | 57385 | instrument, cranioplasty material forming | HBX | Cranioplasty material forming instrument. | Neurology | Neurology | 510(K) Exempt | DePuy Mitek FMS Irrigation Intermediary Tube Without One-way Valve Arthroscopy tubingProduct Code: 281103Fluid Tissue Management | DePuy Mitek FMS Irrigation Intermediary Tube Without Oneway Valve Arthroscopy tubing | 2 | April 12, 2011 | 2011 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Product sterility may be compromised due to breach in sterile packaging | PRODUCTION CONTROLS: Packaging Process Control | Depuy Mitek sent an Urgent Voluntary Product Recall letter dated December 2, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to isolate all inventory of the affected product and call 1-800-382-4682, PROMPT 3, for return good authorization and replacement product. Customers were instructed to return all identified affected product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324Record their assigned return material authorization (RMA) number on the outer shipping box.Use the Prepaid Federal Express shipping label for the return of the product.Include a photo copy of the completed business reply form.For questions or concerns call 1-800-382-4682, PROMPT 3. |  | 7344 | 7344 units (372 boxes x 24 pieces) | Worldwide Distribution - USA (nationwide) and the coutnries of United Arab Emirates, Australia, Belgium, Switzerland, Czech, Denmark, France, UK, Italy, Netherlands, Poland, and Slovakia. | Terminated | April 02, 2013 | 721 | 149720 | Not_Computer
Z-1967-2011 | 58304 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix EVA Container, 3000mL, REF: 741, Sterile R, Baxa Corporation, Englewood, CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | Container, IV | 2 | April 12, 2011 | 2011 | Baxa Corporation | TPN bags may leak during filling operation, due to bag manufacturing error. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, BAXA, sent a "ExactaMix Bag Product Notice" dated March 8, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:If they do not intend to use the manual addition port, verify that the cap is firmly seated onto the port by gently tugging on the cap. If they do intend to use the manual addition port, they should snap off the additive port cap prior to compounding the bag.If they encounter a faulty additive port, contact Baxa Technical Support to receive credit and/or replacement on defective product only.To return affected bags, follow the instructions provided in the notice if the bags were purchased directly from Baxa or purchased from wholesaler, distributor or other 3rd party vendor. Additional copies of the return form can be retrieved by going to www.baxa.com/ebug OR by calling Baxa Technical Support at 800-678-2292.If you have any questions or concerns, contact the Sr. Vice President, Quality and Regulatory at (303) 400-2838 or Baxa Technical Support at 800-678-2292. | Contact the recalling firm for information | 336650 | 336,650 bags | Worldwide distribution: USA (nationwide) and countries including: Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, Hong Kong, Ireland, Japan, Malaysia, Mexico, New Zealand, Panama, Puerto Rico, Singapore, South Korea, Taiwan, UK, Uruguay, Venezuela. | Terminated | March 23, 2012 | 346 | 6561200 | Not_Computer
Z-1960-2011 | 58290 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | STERIS SYSTEM 1E Liquid Chemical Sterilant Processing SystemThe SYSTEM 1 E Liquid Chemical Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned impressible, reusable, critical and semi-critical heat sensitive medical devices, including endoscopes and their accessories. Devices processed in a SYSTEM IE have been chemically sterilized using S40 Sterilant Concentrate, a peracetic acid liquid chemical sterilant and rinsed with extensively treated, potable water. | SYSTEM 1E liquid chemical sterilant processing system | 2 | April 12, 2011 | 2011 | Steris Corporation Hopkins Facility | The printout, display and Operator Manual for the STERIS System 1E Liquid Chemical Sterilant Processing System are subject to recall/correction by the firm because in the event of an incomplete or aborted cycle, the language of the warning messages used in the printout, display, and Operator Manual omits the qualifying term: *liquid chemically* (not **sterilized), when alerting the user to inco | OTHER/UNDETERMINED: Pending | The firm, STERIS, Issued an 'Urgent Field Correction Notice' dated March 14, 2011 addressed to their Hospital Administrator, or Manager, and Risk Manager customers on March 18, 2011. The notification letter describes the product subject to field correction and the problem that exists with the language omission present in the printout display and Operator Manual messages which warn of an incomplete or aborted processing cycle. The notification goes on to state that there is no health hazard involved in this error, as the users are still made aware that the processing cycle has not been successfully completed. The notification goes on to remind the user/customers that in case of any incomplete, cancelled or aborted cycle, the users are instructed to process devices in the cycle following the directions contained in the Operator Manual. Lastly, the notification letter informs the user/customers that the company will be installing new software on the recalled systems in order to revise the language on the cycle printout and the display screen warning of a disrupted cycle. It also states that a STERIS Service Representative will be contacting each facility in order to schedule installation of the new software. The notification also provides the customers with a 1-800 telephone number for STERIS Field Service Dispatch, to be used if the customers have questions or want to request additional information. It also provides a local telephone no. 1-440-392-7455, for obtaining additional information from the firm's Director, Div. of Low Temperature Sterilization at: 1-440-392-7455. |  | 106 | 106 units | Nationwide distribution: USA including states of :CA, DE, FL, GA, IA, ID, IL, KS, MA, MO, MT, NE, NJ, NM, NY, OH, OR, SD, TX, VA, WA, WI, and WV. | Terminated | March 13, 2013 | 701 | 106 | Software | Alarm/Message | Software update | Software Update
Z-1906-2011 | 58175 | device, neurovascular embolization | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Penumbra brand Penumbra Coil System, Penumbra Coil 400, Product is manufactured and distributed by Penumbra Inc., Alameda, CA The Penumbra Coil System (Penumbra Coil 400 TM) is intended for the endovascular embolization of intracranial aneurysms and other neurovascular abnormalities such as arteriovenous fistulae. The coils are also intended for arterial and venous embolization in the peripheral vasculature. The coils function to selectively embolize aneurysms or targeted segments of the vasculature by packing a sufficient quantity of soft platinum coils to achieve occlusion. The Penumbra Coil System is compatible with the Penumbra PX 400 Microcatheter. | Penumbra brand Penumbra Coil System, Penumbra Coil 400 | 1 | April 12, 2011 | 2011 | Penumbra Inc. | The diameter of a component of the device is out of tolerance specification; the Delivery Pusher pull wire which secures the coil in place can prolapse and allow premature detachment of the coil. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Penumbra sent an Urgent Voluntary Field Removal Notice dated March 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately inspect all stock of Penumbra Coil 400 and identify product from the Lots.Remove product from these Lots from inventory and place it in quarantine for return to Penumbra Inc. (Do not use product from these Lots).Complete and return the attached product identification/ return form with in 5 business days via fax to +1 510-217-6414.Penumbra employees will contact customer directly to arrange return and replacement of product.For questions call Regulatory Affairs at 510-748-3224 or e-mail RAQA@prenumbrainc.com. |  | 2082 | 2082 | Worldwide Distribution - USA (nationwide) and the countries in the European Union (Denmark, Germany, Italy, Netherlands, Slovakia, Slovenia, Spain, United Kingdom and Switzerland) and South Africa. | Terminated | June 06, 2011 | 55 | 2082 | Not_Computer
Z-1949-2011 | 58169 | enzyme immunoassay, non-radiolabeled, total thyroxine | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Total T4 Reagent Kit, Part Number: 33800, Lot Numbers (s): All reagent pack lotsThe Access Total T4 assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of total thyroxine (T4) levels in human serum and plasma using the Access Immunoassay Systems. Total T4 assay results are used for diagnosis and confirmation of thyroid disorders such as hyper- and hypothyroidism. | Access Total T4 Reagent Kit | 2 | April 12, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the concentration of total T4 in samples from pregnant women is erroneously low (20%)when measured using the Access Total T4 assay. Erroneously low results which are actually high (outside of the normal reference range)could mask hyperthyroidism. Failure to diagnose or treat hyperthyroid disease associatedwith pregnancy could result in serious, permanent harm | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an Urgent Product Correction letter dated March 20, 2011, to all affected consignees. The letter identified the product, the problem and the action to be taken by the consignee. Consignees were issued the following instructions:(1) If you suspect pregnant patient results were evaluated at your laboratory using only Total T4 to report thyroid status, consider reviewing Total T4 results reported in the last 9 months. Consult with your Laboratory Medical Director and requesting physician toresolve any concerns you may have regarding previously reported patient results during this time period.(2) Do not use the Access Total T4 assay as the only marker for evaluating pregnant patients for thyroid disorders.(3) Please share this information with your laboratory staff and retain this notification as part ofyour laboratory Quality System documentation. If you have forwarded Access Total T4 reagent kits (Part Number 33800) to another laboratory, please provide a copy of this letter to them.(4) Please complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For questions regarding this notice, consignees were instructed to contact Beckman Coulter Customer Support Center.(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter. |  | 151740 | 151,740 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Egypt, El Salvador, Ethiopia, France, Germany, Greece, Guatemala, Hong Kong, India, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Malaysia, Mexico, Republic of Moldova, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Panama, Philippines, Poland, Puerto Rico, Reunion, Romania, Russian Federation, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, and Zambia | Terminated | August 10, 2012 | 486 | 151740 | Not_Computer
Z-1970-2011 | 53195 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LANTIS oncology Information System, an accessory to the Linear Accelerator, as part of the COHERENCE RT Therapist, but also cleared under the ONCOR Avante-Garde and Expression, as well as the syngo RT Therapist system, as part of the ARTISTE And PRIMUS system. Manufactured by IMPAC Medical Inc., Sunnyvale, CA, device distributed by Siemens Medical Solutions Oncology Care System, Concord, CA Electronic record management system as part of a linear accelerator system used to keep track of notes and database for oncology electronic medical records. | LANTIS Oncology Information System | 2 | April 13, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When editing the note type of a previously entered note, then reverting to the original type, data may be lost. | DESIGN: Software Design | Siemens sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to configure their system to lock all new notes automatically.To automatically lock new notes, open eChart Navigator, right mouse click in the notes section, select "configuration" and select "Automatically Lock New Notes."If the customer chooses not to configure their system to automatically lock new notes, follow at least one of the suggestions listed below:If the customer must edit a note type, use the append feature and document that they are making an edit to the note.Lock all new notes that are entered via eChart Notes. To lock a note manually, after creating the note, right mouse click and select "Lock,"Access and enter patient notes via eChart Notes until a LANTIS upgrade has been performed at your site.When entering patient notes via eChart Navigator, make sure that you are entering the notes in the correct section.\When entering patient notes via eChart Notes, make sure that the correct note type has been selected.Customers are to keep the Customer Safety Advisory Notice together with their LANTIS 8.30 User Manual where it should remain until the measures defined above are performed. For questions call 1-888-826-9702. |  | 30 | 30 affected units | Worldwide Distribution - USA including DE, WI, TX, TN, NJ, PA, NE, MA, and AL and the countries of South Africa, Australia, Ireland, Germany, Brazil, South Korea, India, and Japan. | Terminated | April 13, 2011 | 0 | 30 | Not_Computer
Z-1976-2011 | 57714 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic |  | SoftPerm (synergicon A) Daily Wear Contact Lenses, Rx Only, Sterile. Contents: One sterile contact lens, rigid center/soft hydrophilic skirt (25% water), in normal saline solution buffered with sodium bicarbonate. Ciba Vision Group, Duluth GA 30097 USASoftPerm contact lenses are indicted for daily wear (class II device) for correction of visual accuracy for keratoconus and in non-aphakic persons with non-diseased eyes that are myopic or hyperopic | SoftPerm (synergicon A) contact lenses | 2 | April 13, 2011 | 2011 | CIBA Vision Puerto Rico, Inc. | Out of specification values were obtained for sterility testing. The pH of the saline solution may be higher than the spec (up to pH 8.6 compared to a spec limit of pH 7.8). The osmolarity of the saline solution may be lower than specification (263 mOsm/kg compared to a spec limit of 270) | OTHER/UNDETERMINED: Pending | The firm, CIBA Vision, sent two letters of notification to its customers. First letter "IMPORTANT PRODUCT INFORMATION" dated June 2010 and the second "URGENT: Voluntary Product Recall" dated February 1, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: 1) complete the requested information on the attached SoftPerm (synergicon A) Lenses Reply Form and drop the form in the mail within 10 business days and 2) If they have any of the recalled SoftPerm (synergicon A) lenses (shipped to you after March 1, 2009) still in their possession, return as follows: Package the product and affix the prepaid UPS return label and ship within 10 business days or mail to CIBA VISION Corporation, c/o Stericycle Pharmaceutical Services, 2670 Executive Drive, Ste. A, Indianapolis, IN 46241, Atth: SoftPerm Recall. CIBA ask customers to please not include other products, lots, kits, packs, or competitor's products with their shipment. In addition, CIBA will issue a credit to the customers account upon receipt of the returned product. Please feel free to contact Stericycle at 877-542-5928 or CIBA VISION Technical Consultation Department at 1-800-241-7468 should you have any questions or concerns. | Contact the recalling firm for information | 152 | 28, 124 | Worldwide distribution: USA (nationwide) including PR; and countries including: Australia, Austria, Argentina, Bahrain, Belgium, Canada, Colombia, Czech Republic, Denmark, Dubai (United Arab Emirates), Finland, France, Great Britain, Greece, Hungary, Iceland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Netherlands, Oman, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, and Venezuela. | Terminated | November 30, 2011 | 231 | 28124 | Not_Computer
Z-1969-2011 | 56807 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand D-10 Hemoglobin A1c Program Reorder Pack, which contains Update Kit D-10 A1c Program Diskette;Model Numbers: 220-0101 (Reorder Pack); 220-0115 (Program Diskette); Distributed and manufactured by: Bio-Rad Laboratories, Inc., Hercules, CASystem Overview: The Bio-Rad D-10 Hemoglobin Testing System provides an integrated method for the separation and determination of the relative percent of specific hemoglobins (e.g., A2, F, A1 c) in whole blood. The separation is based on the principles of high performance liquid chromatography. Technology Overview: The D-10 Hemoglobin Testing System is a single module designed for hemoglobin analysis in mid-volume laboratories. The D-10 is a fully integrated system for sample dilution, processing and analysis of hemoglobin, designed for use with specific Bio-Rad reagent kits. The D-1 0 incorporates the use of a dedicated software system for instrument control, data collection and analysis. The D-10 Software performs reduction of raw data collected from each analysis. Reagent Kit Information: The D-10 Hemoglobin A1c Program utilizes principles of ion-exchange High-performance liquid chromatography (HPLC). The D-10 Hemoglobin Ale Program is based on chromatographic separation of HbA1c on a cation exchange cartridge. The Bio-Rad D-10 Hemoglobin Ale Program is intended for the percent determination of hemoglobin A1 c in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). The Bio-Rad D-10 Hemoglobin A1c Program is intended for Professional Use Only. The D-10 Hemoglobin A1c Program is for use only with the Bio-Rad D-10 Hemoglobin Testing System. | BioRad D10 Hemoglobing Testing System | 3 | April 13, 2011 | 2011 | Bio-Rad Laboratories, Inc., Hercules, CA | The firm discovered that the Update Kit Floppy Diskette included in the reorder pack was mislabeled with an incorrect Lot Number for the Calibrator/Diluent Set as: S01819; The correct Lot Number for the Calibrator/Diluent Set is: S01704. | OTHER/UNDETERMINED: Pending | The Bio-Rad Laboratories, Inc, sent a Medical Device Correction letter dated 9/8/10, to all affected customers via Federal Express addressing this mislabeling issue, a new Update Kit Floppy Diskette, Lot No. BA02210 has been produced, and Bio-Rad requests that you please destroy the Lot No. AA01815. If you have any questions please contact your regional Bio-Rad office. You may also call (510) 724-7000. |  |  |  | Nationwide Distribution including : PA, OH, MO, CA, TX, FL, MA, MD, KY, OH, MI, NJ, and WA. | Terminated | June 29, 2011 | 77 |  | Not_Computer
Z-1973-2011 | 58035 | table, radiographic, tilting | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Rapid Response H.U.T. (Head Up Tilt) Table. Model 1017. Powered electrophysiology tilt table. Table can tilt -15 degrees to 90 degrees (Trendelenburg). Medical Positioning Incorporated.800-593-3246.Used in clinical and hospital settings by trained radiologists. | Rapid Response HUT Table | 2 | April 13, 2011 | 2011 | Medical Positioning Inc | The tilt tables work intermittently and some have become stuck in an inclined or reclined position during procedures. | OTHER/UNDETERMINED: Pending | An undated Urgent - Voluntary Medical Device Product Correction letter was issued on 2/1/2010 by certified mail with return receipt requested. The letter identified the model and serial numbers of the affected product. The consignees were advised to stop using the table immediately and remove it from service until further notice. Customers were to contact Mr. Joe Hurtig at 1-800-593-3246 or jhurtig@medicalpositioning.com to confirm receipt of the notice and to provide the firm with the correct contact information. Individualized correction packages including replacement parts and instructions were sent to each consignee. The instructions consisted of (1) Rapid Response HUT Motor Replacement Instructions; (2) Control Box Replacement Instructions; and (3) ILD External Power Supply Removal Instructions. |  | 58 | 58 total | Worldwide Distribution -- USA, including Puerto Rico, and Canada. | Terminated | June 15, 2011 | 63 | 58 | Not_Computer
Z-1968-2011 | 58054 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Lube Jelly, 1 oz. water-base sterile lubricating jelly, a component of various Presource Custom Sterile Convenience Packs/Modules; Packaged by Cardinal Health Medical Products and Services, McGaw Park, IL 60085 U.S.A.; The affected lots of modules actually contain a 1 oz. tube of Non-Sterile Kendall Vaseline White Petroleum Jelly.Designed to assist hospitals in preparing for procedures by grouping all of the necessary components used for a procedure in one wrapper. | Presource Custom Sterile Convenience Packs/Modules | 2 | April 13, 2011 | 2011 | Cardinal Health | The Presource kits were assembled with non-sterile petroleum base lubricant instead of the labeled sterile water base lubricating jelly. | TRAINING: Employee Error | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated February 28, 2011 to all affected consignees via UPS next day air, informing them that their Presource kit/modules listed in the attached spreadsheet contain a non-sterile petroleum-based lubricating jelly instead of the labeled water-based sterile lubricating jelly. Cardinal instructed the accounts to quarantine the pack/modules with the indicated work order/lot number on the enclosed spreadsheet and contact the appropriate Customer Service group listed in the letter for return instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules being returned. If the product was further distributed, then those customers should also be notified. | Contact the recalling firm for information | 4236 | 4,236 modules | Nationwide Distribution | Terminated | April 19, 2012 | 372 | 4236 | Not_Computer
Z-1972-2011 | 58022 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC Clinical Systems with Serial Number less than 3412 Including: UniCel DxC 600 Part #: A10405,UniCel DxC 600 PRO Part #: A11810,UniCel DxC 600i Part #: A27318,UniCel DxC 660i Part #: A64871,UniCel DxC 680i Part #: A64903,UniCel DxC 800 Part #: A11816,UniCel DxC 800 PRO Part #: A11812,UniCel DxC 860i Part #: A64935, andUniCel DxC 880i Part #: A59102. | UNICEL DXC SYNCHRON SYSTEMS | 2 | April 13, 2011 | 2011 | Beckman Coulter Inc. | The hose clamps in the DxC (with serial numbers less than 3412, labeled behind the lower right side instrument door) hydropneumatic system are sharp and pose a possible risk of injury. An operator could be cut by the hose clamps, which could in turn result in increased exposure to any waste products (biohazards) that might be present on the assembly. |  | The recall communication was initiated with Beckman Coulter forwarding a Product Corrective Action Letter (PCA) dated November 16, 2010 to the affected customers. The recall notice explained the reason for recall and asked that customers take the following actions:(1) Customers with a service contract will have hose clamp covers installed by their Field Service Engineer during the next scheduled Preventive Maintenance (PM). A kit can be ordered through their Customer Service Representative if consignees prefer not to wait for the next Preventive Maintenance visit and(2) Customers without a service contract can obtain a kit by contacting their local Customer Service Representative.Customers were also informed that Kit (PN A68986) containing hose clamp covers and installation instructions is available at no charge from their Customer Service Representatives. One kit is required for each DxC system.The following telephone numbers were given for customers to obtain the kits: (800) 526-3821 (United States) and (800) 463-7828 (Canada).Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If any of the affected product was forwarded on to another laboratory they were asked to provide a copy of this letter to them.Questions are to be directed to Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 1535 | 1,535 units | Worldwide Distribution--United States and countries of Angola, Australia, Belgium, Bulgaria, China, Columbia, Czech Republic, Denmark, Eritrea, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Mexico, Netherlands, Reunion, Serbia, South Korea, Spain, Slovakia, Swaziland, Sweden, Switzerland, Tunisia, Turkey, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | June 28, 2012 | 442 | 1535 | Not_Computer
Z-1977-2011 | 58016 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF) | UniCel DxH 800 Coulter Cellular Analysis System | 2 | April 13, 2011 | 2011 | Beckman Coulter Inc. | There is a possibility that a specimen could be diluted if a user requests Diluent Dispense while specimens are being automatically processed in cassette presentation. The dilution may cause erroneous results.IMPACT:Erroneous results from a diluted specimen could impact clinical decision making in cases such as blood transfusion, chemotherapy or antibiotic administration- particularly for popu | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with PCA Response Form (via US Postal Service for US customers) dated October 20, 2010, to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter describes the product, problem and the following actions to be taken. The customers were instructed to remove all cassettes from the mix station prior to initiating the Diluent Dispense function to prevent the likelihood of a specimen being diluted; do not use pre-diluted samples for RETIC counts; verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and the Body Fluid results are properly validated at the LIS. This will be corrected in a future version; use the lip on the reagent drawer as shown in the picture above; avoid touching the drawer slides while replacing containers; clean all spills promptly to reduce the risk of exposure to blood borne pathogens; complete and return enclosed PRODUCT CORRECTIVE ACTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7503/7504; mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015 or email to: Regulatory.Notifications@beckmancoulter.com, and share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If you have any questions regarding this PCA, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. | Contact the recalling firm for information | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 422 | Not_Computer
Z-1982-2011 | 57776 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | SmartInfuser PainPump, Part numbers: P49220-S, P49230-S, P49230-SD, P49234-S, and P49234-SD. | SmartInfuser PainPump | 2 | April 14, 2011 | 2011 | Hsmg, Inc. D/b/a Smartinfuser Usa | Elastomeric tube in product's compression unit has the potential to burst following filling. | OTHER/UNDETERMINED: Pending | Firm notified consignees by e-mail and telephone on 3/17/2010. The correspondence identified the affected product and lot numbers; and asked that customers specify which parts, lot numbers and in what quantities they still had in their possession. |  | 1483 | 1,483 units | Nationwide Distribution | Terminated | April 25, 2011 | 11 | 1851 | Not_Computer
Z-1989-2011 | 58225 | prosthesis, knee, patello/femorotibial, semi-constrained, uncemented, porous, coated, polymer/metal/polymer | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Posterior Referencing Instrumentation for the NEXGENÂŽ COMPLETE KNEE SOLUTION QUICK-CONNECT HANDLE, REF 00-5901-034-00, NONSTERILE, Zimmer UK, Ltd.,, Zimmer Warsaw,INThis instrument is attached to the femoral implant or provisional to aid in positioning and inserting the components. | NEXGEN COMPLETE KNEE SOLUTION QUICKCONNECT HANDLE | 2 | April 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potential for breakdown of the AlTiN PVD (black) coating during surgery and the coating could fall into the surgical site. The firm has received 22 reports of the black coating coming off the instruments after cleaning and sterilization cycles and 5 resulted in a delay of surgery. | DESIGN: Device Design | The firm, Zimmer, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 14, 2011 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:Instructions for Handling of Affected Product:1) Stop using the devices and to quarantine immediately2) Carry out a physical count of all affect product in their territory and record data on Inventory Identification Form included with this letter-Note: The instruments must be returned even if the instrument is not exhibiting the black powder3)Ensure they have provided acknowledgement via "email return receipt" upon receiving this notification4) Fax copy of completed form to (574) 372-42655) Return the recalled product along with the completed Inventory Identification Form and the Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580A new or refurbished instrument will be provided as replacement for their returned product once the inventory is available.Instructions for Health Care Professional Notification1) Provide a list of all hospital that currently use or have used the affected product listed in notice by utilizing the User Facility Documentation Form and the Health Care Professional Documentation Form2) Fax a copy of the completed form(s) to: Zimmer, Inc. at (574) 372-4265Zimmer will provide a similar notification via FedEx to these hospitals making them aware of this recall and provide a "Dear Doctor" letter for their surgeons.Important: If you have distributed these affected lots further, please provide a copy of this letter to these customers when you implement these recall instructions.For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc., at 1-800-613-6131.ADDENDUM INFOThe information listed above was updated to include the additional 5 . |  | 757 | 757 | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Germany, Japan, New Delphi and Singapore. Addendum: Finland, Switzerland, UK, Denmark, Spain, Portugal, France, Netherlands, Italy, South Africa, Norway, Spain, and Sweden. | Terminated | March 19, 2013 | 705 | 9462 | Not_Computer
Z-1983-2011 | 58048 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Sky 1000/BHM Medical V10 Bariatric Ceiling Lift System; a patient transfer lift; ArjoHuntleigh, Addison, IL 60101; Model numbers LF20010, LF20011, LF20013, LF20014, LF20015, LF21015, LF21019, 9120001, 9120004, 9120005, 9120008, 9120009 and 9121009The Maxi Sky 1000 is designed for lifting patients in a homecare setting, at nursing homes and other assisted living centers, as well as in institutional or hospital settings. Patient transferring is performed under the supervision of an appropriately trained caregiver staff, in accordance with the instructions for use. All other uses must be avoided. | Maxi Sky 1000 and BHM Medical V10 Bariatric Ceiling Lift | 2 | April 14, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The battery-powered control board in the ceiling lift motor may overheat, resulting in smoke and possibly flames and the lift becoming inoperable. | DESIGN: Device Design | ArjoHuntleigh sent an Urgent Field Safety Notice dated March 4, 2011, to all affected customers via UPS Ground on March 21, 2011. The notice informed the accounts of the potential for smoke and flames coming out of the Maxi Sky 1000/V10 bariatric ceiling lifts. The accounts were requested to immediately cease use of the affected lifts, powering them down by activating the emergency stop, and tag the lifts to prevent further use. If there is a clinical need to continue use of the lift before it is retrofitted or replaced, specific instructions were given to reduce the risk of a device malfunction. The accounts were requested to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 1246 | 1,246 units | Nationwide Distribution | Terminated | February 16, 2012 | 308 | 1246 | Hardware | Device Operation | Replace device | Remove or Replace
Z-1980-2011 | 58288 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Floor Cabinet Part Number: 723335The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF) | UniCel DxH 800 Analyzer | 2 | April 14, 2011 | 2011 | Beckman Coulter Inc. | Sharp edges on drawer slides within the DxH800 Floor Stand (Cabinet) may expose operators to physical injury while replacing reagents and waste containers.IMPACT:There is potential for exposure to blood borne pathogens such as HIV, HBV, and HCV. | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated October 20, 2010, with a PCA Response Form (via US PostalService) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. This letter informed the customers about the problems identified and provided an action/resolution to this issue. Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative. | Contact the recalling firm for information | 422 | 422 units in total (175 in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | April 02, 2012 | 354 | 422 | Not_Computer
Z-1984-2011 | 55910 | abnormal hemoglobin quantitation | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020, included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories, Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. | VARIANT(tm) nbs Sickle Cell Prograin Reorder Pack. | 3 | April 14, 2011 | 2011 | Bio-Rad Laboratories Inc | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed. | DESIGN: Software Design (Manufacturing Process) | Bio-Rad Laboratories sent out the Urgent Medical Device Correction notice on 4/14/2011, by Federal Express. An Update was sent on May 13th, advising customers of a minor discrepancy in the parameters included on the affected Resin Update CD-ROMs. Customers were instructed to examine your inventory for the affected reorder pack lots. Please complete and return the Medical Device Correction Response FormQuestions regarding this recall please call (510) 724-6114. |  | 658 | 658 units | Worldwide Distribution -- USA , Brazil, Canada, Caribbean, France, Italy, and United Kingdom. | Terminated | August 08, 2011 | 116 | 658 | Other |  |  | Other
Z-1979-2011 | 58173 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER AC-T 5diff Autoloader Hematology Analyzer, Part Number: 175356, SerialNumbers(s): All Serial NumbersThe COULTER ACT-T 5diff Autoloader (AL) hematology analyzer is a 26-parameter, fully automated hematology analyzer, including a five-part leukocyte differential counter, capable of analyzing samples in a closed-vial Autoloader mode or a Manual (Stat) mode (open- or closed-vial). | COULTER ACT 5diff Autoloader Hematology Analyzer | 2 | April 14, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified a problem with occasional unexpectedly high results for RBC, HCT, PLT and MPV without instrument generated messages, for the first run of a capped patient sample. Upon repeat analysis the results are correct. Patients could be affected as described below: (1) False high RBC and HCT may result in a delay of treatment,or the delivery of |  | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated February 23, 2011 with an enclosed customer response form to all customers who have the affected instruments via US Postal Service (in US). The letter described the product, problem and actions to be taken. The customers were instructed not to analyze specimens using the Autoloader. Instead, analyze their samples uncapped (open vial) using the Manual (Stat) Mode. The end of this notice includes a Quick Reference for RunningManual Samples. For more detailed instructions, please refer to your instrument's Instructions for Use (IFU), part number 624026, Section 8.4 Running Manual (Stat) Samples. An electronic version of the IFU can be found in the instrument's On-Line Help screen as well as the Beckman Coulter web site. For inconsistent data during patient's follow-up, please consider unusual/unanticipated high or elevated platelet levels as possible falsely elevated results due to this failure. An appropriate action for the need to look-back at previous results is at the discretion of the Laboratory Director, given the fact that laboratories have a variety of measures that are employed to ensure the validity of the results. Some of these provisions include delta checks, laboratory protocols for validating results and interaction with physicians.Additionally, the customers were also requested to complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM within ten (10) days via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015, or email to: Regulatory.notifications@beckmancoulter.com.For any questions concerning this notice, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact your local Beckman Coulter representative. | Contact the recalling firm for information | 1211 | 1,211 units | Worldwide distribution: USA (nationwide) and countries of: Angola Mexico Australia, Azerbaijan, Bangladesh, Bermuda, Botswana, Bulgaria, China, Cote d'Ivoire, Finland, France, Germany, Ghana, Hong Kong, India, Italy, Mexico, Netherlands, New Zealand, Nigeria, Palestinian, Puerto Rico, Romania, Russian Federation,Saudi Arabia, Slovenia, South Africa, Spain, Switzerland,Taiwan, Tanzania, and Turkey. | Terminated | July 16, 2012 | 459 | 1211 | Not_Computer
Z-2000-2011 | 51504 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA.Medical charged-particle radiation therapy system. | LANTIS 6.1 Oncology Information System with MEDOne Module | 2 | April 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. |  | 17 | 17 | Worldwide Distribution | Terminated | April 15, 2011 | 0 | 17 | Software | Output/Calculation | Software update | Software Update
Z-1998-2011 | 56636 | cyclosporine | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Cyclosporine Sample Pretreatment Reagent, COBAS INTEGRA, IVD, Roche Diagnostics GmbH, Assembled for and Distributed by: Roche Diagnostics Corporation, Indianapolis, IN. For the quantitative determination of Cyclosporine in whole blood, using automated clinical analyzers as an aid in management of therapy in kidney, liver, and heart transplants | COBAS INTEGRA Cyclosporine Assay | 3 | April 15, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The cyclosporine sample pretreatment reagent lot was assigned an incorrect expiration date. | OTHER/UNDETERMINED: Pending | Roche Diagnostics Corporation sent an Urgent Medical Device Removal letter dated 8/13/2011, by UPS ( signature required ) to the consignees. The letter identified the product, the problem, and the action to be taken by the consignee.Consignees were instructed to discontinue use of affected product upon receipt of replacement product. Consignees were instructed to dispose of the product at their facilities. Therefore, no product returns are anticipated. Questions concerning this recall please call (317) 521-3911. | Contact the recalling firm for information | 22 | 22 kits | Nationwide Distribution -- including, IN, NJ, NY, TX, CA, PA, OH, and FL. | Terminated | May 04, 2011 | 19 | 22 | Not_Computer
Z-1999-2011 | 58202 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar SP Infusion Suite Database, Version 1.0; an Rx PC software application; Hospira, Inc., Lake Forest, IL 60045; list 13092-04-01A stand-alone PC application for administration of medications specific to the GemStar Infuser | GemStar SP Infusion Suite Database, Version 1.0 | 2 | April 15, 2011 | 2011 | Hospira Inc. | The software enables programming of the GemStar device for a Bolus Lower Lockout Limit outside of the hard limits entered into the GemStar SP Infusion Suite Database PC application. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, Hospira, sent "URGENT DEVICE FIELD CORRECTION" letters dated March 23, 2011 to their customers on the same date. The letter described the product, problem and preventative action to be taken by the customers until the software can be upgraded. The customers were requested to contact Hospira Global Product Safety and Complaints at 1-800-441-4100 to report adverse events, and Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-262-577-6921.For additional information and technical assistance, contact Hospira Advanced knowledge Center at 1-800-241-4002, option 4 (available 24 hours a day/7 days per week). | Contact the recalling firm for information | 139 | 139 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Chile, Colombia and Singapore | Open |  |  | 139 | Software | Device Operation | Software update | Software Update
Z-1873-2011 | 58134 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i, INFX-8000VThe expected usage of these devices is for fluoroscopy studies for various anatomical regions of the human body. | Toshiba Infinixi | 2 | April 15, 2011 | 2011 | Toshiba American Medical Systems Inc | Under some conditions during high doses procedures the cumulative air kerma display may reach its maximum capability and fail to display the accrued value. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | An Important: Electronic Product Radiation Warning notification was sent to all consignees. The letters described the reason for the recall and identified the affected product. The system software will be modified to allow the shift of the decimal to one position to the right allowing for a display of 99999.9 mGy. This will allow for the retention of the resolution until such time as the system exceeds the 9999.99 mGy range. A Toshiba Service person will contact customers when the software is available. Installed systems with the current version of software ( version 4.23) will begin to receive the corrected software April 2011. Previous versions of the software will begin to receive corrected software June 2011. The software will be provided free of charge. Questions related to the letter should be directed to the Regulatory Affairs Department at 800-421-1968. | Contact the recalling firm for information | 220 | 220 units Nationwide | Nationwide Distribution | Terminated | March 18, 2013 | 703 | 220 | Computer
Z-2002-2011 | 58125 | apparatus, suction, single patient use, portable, nonpowered | GCY | Nonpowered, single patient, portable suction apparatus. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Sarns Flexible Weighted Sucker in K9 Pack (custom pack), catalog number 63635-01, Sterile and Non-pyrogenic. Sterilized by ethylene oxide. Terumo Cardiovascular Systems, Ashland, MA .The Sarns weighted sucker is indicated for stationary use to remove excess fluid from the surgical field. | Sarns Flexible Weighted Sucker | 2 | April 15, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received two reports of the Terumo Weighted Flexible Sucker becoming ensnared in cardiac tissue after it was inserted into a ventricle through a valve. The firm has determined that the products's instructions for use are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However this is not stated in the IFU. The firm is conducting | DESIGN: Labeling Design | SEE UPDATE BELOW IN ALL CAPS.The firm, Terumo Cardiovascular Systems (Terumo CVS), sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were informed that the product's Indications For Use (IFU) are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However, this is not stated in the IFU. In addition, the customers were instructed to read and understand the changes and clarifications to the product's IFU; review the Medical Device Correction notice; assure that all users are aware of this notice, and confirm receipt of this communication by faxing the attached Customer Response Form via fax to: Terumo CVS Customer Service at 800-292-6551 (US only).If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. Customers Service hours: Monday-Friday, 8AM-6PM.ON 8/30/2011 THE FIRM ISSUE A LETTER ENTITLED "URGENT MEDICAL DEVICE RECALL" THAT STATED "AFTER FURTHER CONSIDERATION, TERUMO CVS DECIDED TO DISCONTINUE SALES OF THE WEIGHTED FLEXIBLE SUCKER AND IS ASKING USERS TO RETURN ALL UNUSED PRODUCT" TO THE FIRM. |  | 11690 | 11,690 | Worldwide distribution: USA (Nationwide) and countries including: BRAZIL, BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES, and MALAYSIA. | Terminated | December 10, 2012 | 605 | 11690 | Not_Computer
Z-2022-2011 | 58041 | electrode, ion specific, sodium | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CHEM 8+ Cartridge, List Number 03P91-25; 510 K053110Useful for monitoring a variety of conditions. The panel of tests is used in the hospital environment to assess kidney function, electrolyte status, acid/base balance, and blood sugar level. This test panel is also used to assess hypertension and hypokalemia. | iSTAT CHEM 8 Cartridge | 2 | April 18, 2011 | 2011 | Abbott Point Of Care Inc. | i-STAT CHEM 8+ cartridge may generate false results for sodium, chloride and hematocrit. | OTHER/UNDETERMINED: Pending | The firm, Abbott Point of Care (APOC), sent an "Urgent Recall Notice" letter dated February 2011 with business reply cards attached via Fed Ex overnight to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of cartridges from the box numbers of identified lots; discontinue use as a precaution, if their cartridges are no longer traceable to a specific box; return all unused cartridges as instructed on the enclosed Business Reply Card; if they've provided another institution with the cartridges, provide the institution with a copy of this letter, and complete and return the business reply card to APOC.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-336-8020, Option 1 or your Abbott Point of Care representative. | Contact the recalling firm for information | 6575 | 6575 cartirdges | Nationwide distribution. | Terminated | July 10, 2012 | 449 | 6575 | Not_Computer
Z-2013-2011 | 58148 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Lubricating Jelly packaged as follows: (1) DeRoyal(R) Water Soluble Lubricant, Size: 5 Gram, Rx Only, REF 31-702, STERILE EO, Distributed by: DeRoyal - 200 DeBusk Lane, Powell, TN 37849, USA; (2) DeRoyal(R) Water Soluble Lubricant, 1 / Pack, Size: 2 oz Rx Only, REF 31-704, STERILE EO, Distributor: DeRoyal, 200 DeBusk Lane, Powell, TN 37849, USA; (3) DeRoyal(R) Lubricant, Size: 5 Gram, Rx Only, REF 31-938, STERILE EO, Distributed by: DeRoyal - 200 DeBusk Lane, Powell, TN 37849, USALubricating body orifices. | Water Soluble Lubricant | 2 | April 18, 2011 | 2011 | DeRoyal Processing Center | Possible contamination due to sterilization issue. Inspection of contract manufacturer found cGMP deficiencies. | PRODUCTION CONTROLS: Process Control | The firm notified their consignees by a Re: DeRoyal Recall Of Triad Group Lubricating Jelly letter, dated Friday, March 18,2011. The letter discussed the reason for recall and the affected products. The letter asked customers to identify inventory using the attached spreadsheet and quarantine to prevent further use. Customers are to complete the spreadsheet regardless if they have effected product or not. The completed form can be faxed to 865-362-3716 or emailed to recalls@deroyal.com. All affected products should be return to DeRoyal at the address provided. Distributors are to forward a copy of the notice to their end-users. If there are questions or assistance is needed with the recall, Stephanie DeBusk should be contacted at 865-362-1037 or Sharon Cook at 865-362-6157. | Contact the recalling firm for information | 323440 | 323,440 units | Nationwide Distribution | Open |  |  | 323440 | Not_Computer
Z-2021-2011 | 58080 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Avanto, Espree, Trio A Tim System.Nuclear magnetic resonance imaging. | MAGNETOM Avanto, Espree, Trio A Tim System | 2 | April 18, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | OTHER/UNDETERMINED: Pending | Siemens Medical Solutions Inc., issued a Customer Safety Advisory Notice dated November 24, 2010, to affected customers via Update Instructions MR029/10S. Affected customers will be notified via hand delivered letter by Siemens Service Engineers at the time of inspection. A future Update instruction will be released for installation of detectors in appropriate locations on systems to switch off the power electronics in the event of a malfunction in the power supply of the gradient coils.Further questions regarding this recall please call (610) 448-3237. | Contact the recalling firm for information | 1419 | 1419 | Product was distributed domestically to hospitals and imaging centers in the District of Columbia, Puerto Rico, AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY(every State except VT). | Open |  |  | 1419 | Not_Computer
Z-2023-2011 | 55577 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Treatment Table, part number 7346534; Product is manufactured and distributed by Siemens AG, Medical Solutions, Roentgenstrasse 19-21, Kemnath, Germany 95478.The 550 TxT Treatment Table is to be used as an integral part of the Siemens linear accelerator systems to position the patient accurately according to the treatment plan for various radiation therapy treatment techniques. | Siemens brand 550 TxT Patient Treatment Table | 2 | April 19, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential for causing overheating of the power supply during treatment, and potential for unexpected table descent. | DESIGN: Component Design/Selection | Siemens issued a Customer Information letter dated March 21, 2010. There were two safety issues. There were two occurrences of smoke caused by an overheating of the power supply during patient treatment. Also, there had been a report of an unexpected lowering of the 550TxT table in patient rescue mode. With the released kit that accompanies this letter, Siemens has implemented a minor modification on the 550TxT to include an additional fuse within the power supply circuit to address the first issue described above. They also included an additional modification to address the lowering of the table by the integration of additional relays. As soon as the "move enable" button is released, the table brake will be enabled which is controlled by the implemented relays. No further action is required for the continued operations of the 550TxT table.Further questions regarding this recall please call (925) 246-8407. | Contact the recalling firm for information | 326 | 326 units | Worldwide Distribution -- Including USA, states of NJ, WI, NY, and MA. | Terminated | March 01, 2012 | 317 | 326 | Battery | Physical Safety Hazards | Repair | Repair
Z-2027-2011 | 58276 | instrument, surgical, disposable | KDC | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Hospira Primary Microbore Gemstar Split Set, Integral PAV, Yellow Striped Tubing, 110 Inch, Non-DEHP; a sterile Rx tubing set for epidural administration of fluids with the Gemstar Pump; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USA; List No. 13744-28For epidural administration of fluids with the Gemstar Pump | Hospira Primary Microbore Gemstar Split Set | 2 | April 19, 2011 | 2011 | Hospira Inc. | There is a potential for the air filter at the distal end of the set to become detached. This could result in the contamination of the sterile fluid path inside the tubing. | OTHER/UNDETERMINED: Pending | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated March 28, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return the Reply Form via fax to 1-866-324-3734, even if you do not have the affected product; inform healthcare professionals in their organization of this recall; if they have distributed the product further, notify their accounts and have them fax the reply form to 1-866-324-3724; return the affected product to Stericycle using labels provided, and call Stericycle at 1-877-877-0481, if they have not received a return label. For medical inquiries, please call Hospira Medical Communications at 1-800-615-0187. Call Hospira Global Product Safety and Complaints at 1-800-441-4100 (8am-5pm CST, M-F) or by e-mail at (ProductComplaintsPP@hospira.com). |  | 61150 | 61,150 sets | Worldwide distribution: USA (nationwide) and countries including: Canada, Korea and Singapore. | Terminated | May 09, 2012 | 386 | 147646 | Not_Computer
Z-2047-2011 | 55154 | radioimmunoassay, prolactin (lactogen) | CFT | Prolactin (lactogen) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Active Prolactin ELISA, 96 Wells, DSL-10-4500The DSL-10-4500 Active Prolactin Enzyme-Linked Immunosorbent Assay (ELISA) Kit provides materials for the quantitative measurement of prolactin in human serum. This assay is intended for in vitro diagnostic use in the diagnosis and treatment of disorders of the anterior pituitary gland or of the hypothalamus portion of the brain. | Active Prolactin ELISA | 2 | April 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Prolactin ElLISA kit, lot number 890291 over estimates the expected dose values for patient serum samples due to a shift in the assigned calibrator values.Beckman Coulter recommends that patient results generated with the affected lot be evaluated in the context of other diagnostic tests and the clinical presentatio |  | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" Letter dated August 03, 2009 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:Action/Resolution:Discontinue use and discard remaining product with lot number 890291 of Active Prolactin reagents. Contact Beckman Coulter at 1-800-231-7970 for replacement product.Newer lots (lot number 991210 or higher) of the product have been realibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, they were told to contact Technical support at the phone number listed below.Consignees were asked to share this information with their laboratory staff and retain this letter for their Quality System documentation and asked to complete and return the enclosed response form within 10 days via fax to: 786-639-4000 so that Beckman can be assured that they received the notification.If you have any other questions regarding this notification, please contact DSL Technical Support Center at 1-800-231-7970 in the United States and Canada , or contact their local Beckman Coulter Representative. |  | 30 | 30 Kits | Nationwide distribution: USA including states of: FL, GA and NC. | Terminated | April 27, 2012 | 373 | 30 | Not_Computer
Z-2045-2011 | 57750 | reagent, general purpose | LDT | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | IMMAGEÂŽ Immunochemistry Systems Wash Solution Part Number: 447060The IMMAGEÂŽ Immunochemistry Systems Wash Solution is intended for use with IMMAGEÂŽ Immunochemistry Systems and reagents for the quantitative determinations of components in biological fluids. The Beckman Coulter IMMAGEÂŽ Immunochemistry System is a fully automated, computer controlled, bench-top analyzer designed for the in vitro quantitation of biological fluid components and therapeutic drugs. The system methodologies are rate turbidimetry and rate nephelometry. | IMMAGE Immunochemistry Systems | 2 | April 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an ImmageÂż Systems Wash Solution Lot 71744 was manufactured by a Beckman European facility and not intended to be used in the U.S. However, this lot was inadvertently shipped to the U.S. and delivered to five customers. | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, contacted the customers who received the IMMAGE Immunochemistry Systems Wash Solution, lot #71744 by telephone. Beckman sent a follow-up "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 25, 2010, to all affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of ImmageÂŽ Immunochemistry Systems Wash Solution lot 71744; return the material to Beckman Coulter's Chino warehouse: Beckman Coulter Inc., 15989 Cypress Ave, Chino, CA 91710, Phone: 909-597-3967;package the material the way it arrived and use Beckmans internal Fedex account to ship the material; complete and return an enclosed FAX BACK RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulator Affairs, 250 S. Kraemer Blvd, Brea, California 92821, Mail code E2.Se.08; share this information with laboratory staff; retain the notification as part of their Quality System documentation, and provide a copy of the letter if they forwarded the affected products.If you have any questions regarding this letter, please call Beckmans Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada. |  | 9 | 9 units total in US | Nationwide distribution: USA including states of : AL, CA, MI and IN. | Terminated | June 26, 2012 | 433 | 9 | Not_Computer
Z-2030-2011 | 57340 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman Metal Disposable ICP KitProduct Code: 80-1190 This is a single use, disposable drill kit designed for gaining access to the cranium for therapeutic and/or diagnostic procedures | Codman | 2 | April 20, 2011 | 2011 | Codman & Shurtleff, Inc. | Packaging defect may compromise sterility of device | DESIGN: Packaging Design/Selection | The firm, Codman, sent an "Important Notice Urgent- Voluntary Product Recall Notification" letter dated November 30, 2010 to its customers via FedX. On January 7, 2011, Codman expanded the recall to four additional product codes.On January 31, 2011, the firm issued a supplemental recall letter and identified one additional product code: 80-11197. The letters described the product, problem and actions to be taken. The customers were instructed to identify the unused affected products in their inventory and follow the recall instructions provided: 1) If you have NO product to return, please advise Codman as soon as possible using the FACSIMILE form attached and fax to 508-828-6118 and 2) if you DO have product to return, please call 1-800-225-0460, Option 4, for instructions for return.Codman is presently working toward supplying replacement products as soon as possible. Please call Customer Service to check on the status of product availability at 1-800-225-0460.Should you have any questions regarding this recall, please call Customer Service at 1-800-225-0460 or your local Codman Representative. |  | 1671 | 1671 units | Worldwide distribution: USA including states of :AL, FL,GA, KY, MA, MI, MN, MS, OR, PA, RI, UT,WI and WV; and countries including: Argentina, Australia, Canada, Bahamas , Belgium, Dubai, Japan, Madrid, Mexico, New Zealand, Saudia Arabia, UAE, Uruguay, and United Kingdom. | Terminated | March 14, 2013 | 694 | 14792 | Not_Computer
Z-2044-2011 | 58174 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTERÂŽ HmX Hematology Analyzerwith AutoloaderPart number: 6605525, 6605526, 6605527The COULTER HmX Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential cell counter for In Vitro Diagnostic Use in clinical laboratories. | COULTER HmX Hematology Analyzer with Autoloader | 2 | April 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with Integrated Circuit Chips used in different boards throughout the Beckman Coulter LH500/HmX/HmX Autoloader Hematology analyzer. | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated March 2, 2011, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTERÂŽ LH 500 Series Analyzer, COULTERÂŽ HmX Hematology Analyzer, and COULTERÂŽ HmX Hematology Analyzer with Autoloader. The letter described the product, problem and actions to be taken. Customers were informed that the suspect Integrated Circuit Chips have been removed from all systems in inventory and manufacturing.The customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality Systemdocumentation. If customers have forwarded any of the affected product(s) listed on the PCA letter to another laboratory, please provide a copy of the letterto them.The customers were instructed to complete and return the enclosed Response Form within 10 days so Beckman are assured they have received this important communication and are taking appropriate action.The customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States and Canada or contact their local Beckman Coulter Representative. |  | 288 | 288 units total (158 in US) | Worldwide distribution: USA (nationwide) and countries of: Canada, France, Germany, Italy, Spain, United Kingdom, and Singapore. | Terminated | May 29, 2012 | 405 | 288 | Not_Computer
Z-2038-2011 | 57933 | burr, orthopedic | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ***REF MCA7-1SB*** 1 mm Fluted Ball, 7.3 cm. Use with Micro Curved Attachment (MCA).***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only***Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327 6887/ +1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126***Lot numbers: C523029652, D063031867, D113032925, D173034373 and D213035424.Intended usage: Cutting and shaping bone including bones of the spine and cranium. | Sterile Bone Cutting Bur | 2 | April 20, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Sterile Bone Cutting Burs that was manufactured with incomplete flutes which could cause degraded cutting performance and in some cases not cut at all. Anspach has not received any complaints related to this problem. | OTHER/UNDETERMINED: Pending | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated October 6, 2010, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the part numbers listed in the letter. Do not include any product received after October 7, 2010.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply Form indicating how many of each part number is being returned and confirming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. |  | 271 | 271 cutting burs affected | Worldwide distribution: USA including states of: AZ, CA, FL, GA, IL, IN, LA, MA, and MI; and countries of: Australia, Belgium, England, France, Italy, Japan, South Africa, Spain and Sweden. | Terminated | August 12, 2011 | 114 | 271 | Not_Computer
Z-2041-2011 | 57888 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 9200 DDE Panoramic and Cephalometric Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; Part numbers 110-0199G2, 110-0199G4, 110-0234G1 and 110-0234G3. The intended use of the device is to be used as an extraoral source of x-rays for imaging of the dento-maxillo-facial area. | Orthoralix 9200 DDE Panoramic & Cephalometric Dental XRay System | 2 | April 22, 2011 | 2011 | Gendex Dental Systems | There is an anomaly in the Orthoralix firmware which can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obviou | OTHER/UNDETERMINED: Pending | Gendex sent a recall letter to their accounts on February 4, 2011 via first class mail, stating that an anomaly in the Orthoralix firmware can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obvious these images cannot be used for diagnostic purposes. The letter stated that resending the image from the X-Ray system to the PC will correct the issue present in certain images sent during periods of high network traffic. Instructions were provided to download an Appendix to Orthoralix DDE User Manual that explains how to resend the images. If that does not resolve the issue, users were requested to contact Gendex Technical Support at 1-800-323-8029. Customers were asked to complete and return the enclosed Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com. For questions regarding this recall call 800-323-8029. |  | 6835 | Total 6,835 units of all products | Worldwide Distribution--United States, including Puerto Rico, Australia, Canada, Chile, Colombia, Germany, Hong Kong, India, Italy, Korea, Mexico, New Zealand, Peru, Singapore, Taiwan and Vietnam. | Terminated | February 15, 2012 | 299 | 6835 | Software
Z-2058-2011 | 58324 | prosthesis, hip, semi-constrained, metal/ceramic/ceramic/metal, cemented or uncemented | MRA |  |  | Orthopedic |  | R3 TM CERAMIC ACETABULAR LINER, BIOLOXÂŽ FORTE, 36 MM O.D. 62 MM O.D., REF 71338962, STERILE R, QTY (1), Smith & Nephew, Inc., Memphis, TN 38116, 2017/06 | R3 Forte Ceramic Liners | 2 | April 22, 2011 | 2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners, the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners. | TRAINING: Employee Error | All affected Smith & Nephew Distributors were notified via email and the hospitals were notified by letter delivered by Fed-Ex on 3/11/2011. They were instructed to immediately cease distribution, quarantine affected product for return, if product further distributed, identify customers and immediately notify them of recall.Please contact Mr. David Archer, Group Director, at 701.396.2121, ext. 5373, if you have questions. |  | 67 | 67 units | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OR, PA, SC, TN, TX, WA, and WI | Terminated | February 08, 2012 | 292 | 2271 | Not_Computer
Z-2051-2011 | 58231 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound systemManufactured by Siemens Medical Solutions USA, Inc.Business Unit Ultrasound,1230 Shorebird WayIntended for the following applications: Fetal, abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal and Peripheral Vascular Applications. | Acuson S2000 ultrasound system | 2 | April 22, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The Firm discovered a software error on the Accuson S2000 ultrasound systems affecting the measurement results in the patient report. | DESIGN: Software Design | Siemens issued a Customer Safety Advisory Notification letter to all customers on March 15, 2011, regarding a software error on the ACUSON S2000 ultrasound systems affected measurement . They are in the process of releasing a software update that will correct this issue. Siemens recommends reviewing the patient report at the end of each study to confirm that the report contains only measurements relevant to the exam just performed. No patient injury has been reported. Further questions regarding this recall please call (650) 694-5398. |  | 1533 | 1533 units | Worldwide Distribution including USA, states of AZ, VA, CA, NE, IL, MN, TN, FL, NM, and DC | Terminated | June 18, 2012 | 423 | 1533 | Software | Output/Calculation | Software update | Software Update
Z-2061-2011 | 53053 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 (model numbers 10348531, 10348532 and 10348533), when using the V5Ms transducer (model number 08263703) and the Plus option (10427795)Product is an ultrasound imaging system intended for General Radiology,Fetal, Abdominal, Intaoperatve, Pediatric, Small parts, Neunatal/Adult Cephalic, Cardiac, Transesophageal, Pelvic, Transcranial, OB/GYN, Urology, Vascular, Musculoskeletal, Superficial Musculoskeletal, and peripheral Vascular Applications. The system also provides for the measurement of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. The Arterial Health Package software provides the physician with the capability to measure Intima Media Thickness and the option to reference normative tables that have been validated and published with peer-reviewed studies. The information is intended to provide the physician with an easily understood tool for communicating with patients regarding the state of their cardiovascular system. The V5Ms is a multifrequency,multi-plane, phased sector array transducer, which can be introduced into the esophagus to obtain heart structure images. Transesophageal echocardiography is used to obtain unobstructed views of the heart surrounding tissue. | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 | 2 | April 25, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Product Failure-- When the 64-Channel System V5M is used in the 128-Channel mode, an over current fault occurs which shuts the system down. | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notification letter on July 24, 2009. The letter identified the product, the problem, and the action to be taken by the customer. A software update will be conducted in the field. Until that update occurs, customers were instructed to image without the CW doppler to avoid a system shut down.Customers were instructed to contact their service support person with any questions.For questions regarding this recall call 650-694-5993. |  | 70 | 70 units | Worldwide Distribution - USA including CA, CT, FL, MA, and TX and the countries of United Arab Emirates, Germany, Ecuador, Spain, France, Great Britian, Saudi Arabia, Sweden, Thailand, and South Africa. | Terminated | April 25, 2011 | 0 | 70 | Software | Device Operation | Software update | Software Update
Z-2065-2011 | 53209 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson X300 systems at software versions 1.00.00 to 1.0.10 and 2.0.00 to 2.0.02, model numbers 10037409, 10132987, and 10038837, manufactured by Siemens Medical Solutions ultrasound business unit, Mountain View, CAMedical device is an ultrasound imaging system for general radiology, fetal, abdominal, intraoperative, pediatric, small parts, neonatal/adult cephalic, cardiac, transesophageal, pelvic, transcranial, ob/gyn, urology, vascular, musculoskeletal, superficial musculoskeletal and peripheral vascular applications. The system also provides for measurements of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. | Acuson X300 | 2 | April 25, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Software Error with potential for misinterpretation of data-- The previous patient's last capture clip image can appear in the next patient's study. | DESIGN: Software Design | The firm, Siemens, sent a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to its customers on July 30, 2009. The letter described the product, problem and actions to be taken. The customers were instructed they need to update any of the patient registration settings to reset the cine buffer so that no data is transferred between studies, if the system software is not being updated. Siemens informed the customers that a mandatory update software is being installed on their system free of charge.If you have any questions, contact (650) 969-9112. |  | 2312 | 2312 units | Worldwide distribution: USA (nationwide) and countries including: Albania, algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile,China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Kazakhstan, Kenya, Kosovo, Lativia, Lithuania, Luxembourg, New, Zealand, Nicaragua, Norway, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam. | Terminated | April 25, 2011 | 0 | 2312 | Software | Display/Image | Software update | Software Update
Z-2060-2011 | 58261 | light source, fiberoptic, routine | FCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LuxtecÂŽ MLX 300 Watt Xenon Light Source,Model OOMLX; Surgical/Medical Light SourceThis light source is designed to supply high-intensity light to a fiber optic cable for illumination of a surgical field during a surgical and/or medical procedure. | Luxtec MLX 300 Watt Xenon Light Source | 2 | April 25, 2011 | 2011 | Integra LifeSciences Corp. | Malfunction of the MLX Control Board may cause loss of manual control and Lamp shut down. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Integra notified all customers/distributors via written communication, dated 3/08/2011. On 3/09/11, Integra Sales Representatives and Service Technicians were notified via telephone conference call and provided copies of written communications. Customers were advised to immediately discontinue use of any affected product and to place in a secure area. Note: This defect only applies to MLX units that were repaired between 12/01/2009 and 1/31/2011. This defect does not occur in manufactured units that have not been serviced, nor does it apply to MLX units that were serviced outside of this specific timeframe. Customers with devices already known to require repair and/or servicing will be contacted directly to arrange for the return of their device(s). Additionally, sales specialists were instructed to examine each customer's inventory for those device(s) that may potentially have the jumper-wire and to perform a stand-by test to determine if repair and/or servicing are required. |  | 159 | 159 units | Nationwide, Canada, Australia, Germany, Netherlands, New Zealand, Norway, Singapore and Sweden | Open |  |  | 159 | Hardware | Device Operation | Replace | Remove or Replace
Z-2062-2011 | 58320 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | S8-3t Micro TEE Transducer is an Ultrasonic Diagnostic. The Part # is 989605379471.The User information is labeled in part: "***iE33 Ultrasound System ***User Information Update***4535 614 04811 Rev A***New S8-3t TEE Transducer***This release of the iE33 Ultrasound System supports the S8-3t transesophageal echocardiographic (TEE) transducer***Manufactured by Philips Ultrasound***22100 Bothell-Everett Highway***Bothell, WA 98021-8431***USA***Telephone: +1 425-487-7000 or 800-426-2670***Fax: +1 425-485-6080***www.philips.com/ultrasound***".The S8-3t transducer is indicated for adult cardiac, pediatric cardiac and transesophageal uses, and operates in the following modes: B, M, Color M, CWD, PWD, high PRF, color Doppler, harmonic, XRES. The S8-3t Micro TEE is indicated for use on patients weighing at least 2.5 kg undergoing echocardiography exams in the ICU, echocardiography lab, interventional suite, or in the operating room during cardiac surgery. The S8-3t Micro TEE transducer can be identified by the designation "S8-3t" printed on the connector housing of the transducer. | S83t Micro TEE Transducer | 2 | April 25, 2011 | 2011 | Philips Medical Systems Nederlands | Users of the Philips model S8-3t Micro TEE Transducer have reported that significant image quality degradation may occur after a number of use/cleaning cycles. If users do not inspect the transducer for such damage prior to each use, imaging may not always be available, which could pose a risk for patients. | OTHER/UNDETERMINED: Pending | Philips Healthcare began mailing Urgent Medical Device Correction letter on March 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers are being instructed on how to inspect the distal tip for defects that cause image quality degradation, and to plan for epicardial / transthoracic ultrasound imaging as an adjunct to the S8-3t. Customers are also reminded to carefully follow cleaning described under Handling the S8-3t Transducer, of Philips' iE33 System User Information. Customers can call Philips Customer Service at 800-722-9377 and reference to FCO 79500198 for any questions about this recall. | Contact the recalling firm for information | 155 | 155 units | Worldwide Distribution - USA AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, KY, LA, MA, MD, MI, MN, MS, NC, ND, NH, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Israel, Italy, Japan, Korea Rep, Netherlands, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Ukraine, and United Kingdom. | Terminated | February 14, 2012 | 295 | 155 | Not_Computer
Z-2063-2011 | 58246 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Optical Guidance Platform | 2 | April 25, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. Varian is preparing a software patch which will be applied to affected units to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. | Contact the recalling firm for information | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 421 | 599 | Software | Device Operation | Software update | Software Update
Z-2070-2011 | 57779 | electrode, ion specific, sodium | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SynchronÂŽ LX I 725The Synchron LXi 725 System combines the Synchron LX20 PRO analyzer and the Access 2 analyzer into a single instrument presentation. Samples are loaded from a single point of entry through a Closed Tube Aliquotter (CTA) unit. The CTA functions as a sample processing manager by aliquotting and routing samples to the Access 2 and/or LX20 PRO analyzer according to programming requirements. | Unicel DXC and SYNCHRON LX Clinical Systems | 2 | April 26, 2011 | 2011 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has confirmed that for User-Defined Chemistries (UDRs), the polychrome correction is not automatically updated when wavelengths are modified. | PRODUCTION CONTROLS: Environmental Control | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased Synchron LX and UniCelÂŽÂŽ DxC Clinical SystemsOperating Software (all versions). The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:-After modifying any UDR wavelength, they MUST reboot to apply the update of the polychrome correction. -If the modified UDR requires calibration, recalibrate after rebooting. - Complete and return the enclosed RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd, PO Box 8000, Brea, California, 92282-8000; Mail Code E2.SE.08. In addition, the customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If customers forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, please call Beckman's Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada or www.beckmancoulter.com/customersupport/supportrequests. |  | 1410 | 1,410 units (503 in US) | Worldwide distribution: USA (nationwide) and countries including:Angola, Australia, Belgium, Bulgaria, China, Colombia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Eritrea, Finland, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Macao, Mexico, Namibia, Netherlands, Nez Zealand, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 1410 | Computer
Z-2068-2011 | 58042 | osmotic fragility testing | GKE | Osmotic fragility test. | Hematology | Hematology | 510(K) Exempt | Advanced Instrument Model 2020 Multi-Sample OsmometerIntended to measure the osmotic pressure of body fluids. | Advanced Instrument Model 2020 MultiSample Osmometer | 2 | April 26, 2011 | 2011 | Advanced Instruments, Inc. | A poorly soldered power cord may pose a potential electrical shock hazard. | TRAINING: Employee Error | Advanced Instruments Product Service Department contacted each customer on 1/19/11 via an Urgent Product Recall Notice. The letter identified the reason for recall along with the affected products. The letter provided instructions to safely examine the date code of the AC Power Inlet assembly on the affected product. If customers were unable to perform the inspection or had questions concerning the instructions, they were to contact Advanced Instruments Product Service Manager at 781-471-2147 or via email jorges@aicompanies.com. The attached response forms were to be completed and returned as instructed. |  | 3 | 3 units | Worldwide Distribution -- USA, including states of CA, GA IL, NJ, and MA and country of Belgium. | Terminated | December 01, 2011 | 219 | 3 | Battery | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-2052-2011 | 57744 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific's LATITUDE Patient Management System:LATITUDE 6.1 Patient Management System WebsiteServer software, Model 6488 for United StatesServer software, Model 6465 for European UnionLATITUDE 6.2 Patient Management System WebsiteServer software, Model 6441 for CanadaServer software, Model 6442 for Australia / New ZealandIntended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | LATITUDE Patient Management System | 2 | April 26, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that data within certain printed reports obtained from the LATITUDE Patient Management System may not be accurate, in that, the Most Recent Presenting Electrogram (MRP-EGM) for some of the selected patients may be either missing or incorrect. | DESIGN: Software Design | Consignees were sent a Boston Scientific "LATITUDEÂŽ Medical Records Information" letter dated January 2011. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall.The letter also included the Rate of Occurrence, Action and Further Information sections. | Contact the recalling firm for information | 25836 | 25,836 | Worldwide Distribution -- USA, including DC and PR, and countries of AUSTRALIA, CANADA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | November 15, 2012 | 569 | 25836 | Other | Output/Calculation | Notification | Safety Notice/Insructions
Z-2059-2011 | 58209 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice Computed Tomography Scanners, equipped with Software Version 1.1.1. Model #E989605657371.The MX 16-slice CT Scanner System can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. | PHILIPS MODEL MX 16slice Computed Tomography Scanners | 2 | April 26, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The Philips MX 16-slice CT scanners are subject to recall/field correction due to a problem encountered with the units when using the 'Combine Viewing' Option. The firm determined that inaccurate dimension measurements display when using the Combine Viewing Option. Moreover, if the dimension measurements from merged images created using this option are used, there is a potential for applying inac | OTHER/UNDETERMINED: Pending | Philips Healthcare issued an 'URGENT- Field Safety Notice MX 16-slice' (dated 3/10/2011) concerning a software issue to customers. The notification informed the firm's users/customers of the potential for inaccurate dimension measurements produced by the MX 16-slice CT scanners when using the Combine Function Viewing Option. Moreover, the letter notified customers of the potential for patient misdiagnosis in the event that inaccurate dimension measurements are used by a physician for reference. The Field Safety Notice described the following: 1) How to identify the affected product; 2) What actions should to be taken by the customers with regard to avoiding the issue -discontinue use of the Combine Viewing Option and instead use the MPR Batch for diagnostic reviews until a planned software upgrade is installed by the recalling firm; and 3) The firm's plans to carry out the installation of a field change order kits to correct the issue.The planned software upgrade is expected to be completed on or before 6 months after the new software is released. If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center(1-800-722-9377, option 5: Enter Site ID or follow the prompts). | Contact the recalling firm for information | 25 | 25 units were distributed domestically & 168 system units were distributed both domestically & internationally | Worldwide Distribution -- United States (AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, SC and TX) including Puerto Rico and countries of Albania, Argentina, Australia, Belgium Brazil, Chile, Columbia, Dominican Republic, Ecuador, Germany, Greece, India, Indonesia, Italy, japan, Kenya, Lebanon, Malawi, The Netherlands, Philippines, Poland, Romania, Saudi Arabia, South Africa, Spain, Switzerland, Turkey, Uganda, United Arab Emirates and Venezuela | Terminated | March 01, 2013 | 675 | 25 | Software | Display/Image | Software update | Software Update
Z-2066-2011 | 58371 | imager, ultrasonic obstetric-gynecologic | HEM | Obstetric-gynecologic ultrasonic imager. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | OB-Link release 2.0.7. distributed by Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098OBLink provides automated transmission of patient data and measurements from ultrasound systems directly to Digisonics OB-View/OB-Windows systems. | OBLink release 2.0.7 | 2 | April 26, 2011 | 2011 | Digisonics, Inc | Data/measurements from a OB report sent from ultrasound systems populated the wrong OB-View report | OTHER/UNDETERMINED: Pending | Digisonics visited consignee in April 2008 as follow-up to complaint. | Contact the recalling firm for information | 2 | 2 units | Nationwide Distribution including OK | Terminated | May 02, 2011 | 6 | 2 | Other |  |  | Other
Z-2049-2011 | 58135 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N183, REF 411 021, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, Marcy-l'Etoile France. The cards are intended for use with the Vitek 2 systems in clinical laboratories as an in vitro test to determine the susceptibility of clinically significant aerobic/anaerobic gram-negative bacilli to antimicrobial agents. | VITEK 2 Gram Negative Susceptibility Card, ASTN183 | 1 | April 27, 2011 | 2011 | Biomerieux Inc | The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BioMerieux sent an Urgent Product Correction Notice (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786," letter dated March 30, 2011, via Fed EX on April 4, 2011. to all affected customers. The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Tazobactam (TZP2). The new cards will begin shipping by July 16, 2011, but in the meantime, the cards currently in possession of the customers can be used with specific required actions that need to be followed. Until the TZP redevelopment is completed and a resolution can be implemented, BioMerieux requires their customers to immediately take the following actions: The customers will be required to suppress the TZP2 resistant and susceptible results for Escherichia coli, Morganella morganii, Proteus mirabilis, Proteus vulgaris, Providencia rettgeri, Providencia stuartii, Pseudomonas aeruginosa, and Salmonella enterica and use an alternate method of testing. In addition, the customer is to suppress the TZP2 resistant result only for Klebsiella pneumonia and use an alternate method of testing.The letter provides instructions for configuring the VITEK 2 software to suppress the results from reporting.The instructions contained in the recall notification are to be applied to all current and future product lots that contain TZP2 until further notice.The customer is to complete the enclosed acknowledgement form and return it via FAX to acknowledge receipt of the notification and is requested to file the correction notice with their VITEK 2 Systems permanent records according to their internal policy or procedure.For questions or additional information customers were instructed to call 1-800-682-2666, option 3.In addit. | No consumer action necessary | 493 | 493 boxes as of 3/31/11 | There was no U.S. distribution as the cards were for export only to Canada and Sweden. | Terminated | June 01, 2012 | 401 | 3538 | Not_Computer
Z-2101-2011 | 58289 | kit, plastic surgery and accessories | FTN | Plastic surgery kit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Argon Medical Devices, Arterial Line Kit/1, 18ga x 6", Part Number 498447, Lot number 99767634Multi-purpose Kit blister tray coupled with Tyvek lid stock heat-sealed to form a sterile barrier. | Argon Medical Devices, Arterial Line Kit/1, 18ga x 6 | 2 | April 27, 2011 | 2011 | Argon Medical Devices, Inc | Packaging for sterile tray has the potential to have a hole which could compromise sterility for the products inside. | OTHER/UNDETERMINED: Pending | Argon Medical Devices sent an Urgent Medical Device Recall letter dated March 14, 2011, by certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine and return any product remaining at their facility utilizing the instruction on the 2011 Tray Recall Response Form. Customers were instructed to complete the response form included with the letter and return by fax or email as indicated as soon as possible. Customers were instructed to pass the notice on to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. For questions regarding this recall call (903) 677-9375. |  | 10 | 10 units | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, GA, HI, IL, IN, LA, MD, MN, ND, NE, NH, NJ, NY, OK, PA, SC, TX, VA, WA, and WI and the countries of Nationwide, Japan, Australia, Canada, Singapore, and the Netherlands | Terminated | April 24, 2012 | 363 | 5780 | Not_Computer
Z-2102-2011 | 58271 | spinal channeling instrument, vertebroplasty | OCJ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConfidenceÂŽ Diamond Tip Introducer Needle 11G x 6-InchProduct Code: 2839-03-611The Confidence Needles are supplied sterile. They are used to place bone cement percutaneously at a desired location. They come in a variety of sizes (length and diameter) and configurations (Bevel tip, Diamond tip). The cement extrudes from the end of the cannulated needle with the exception of the side hole needle which extrudes cement from a portal on the side of the tube near the distal end. The needle is placed under fluoroscopy prior to cement delivery. The needle is also sometimes used in MIS cases to place the guide wire that is used to assist in the placement of the pedicle screw. | DePuy Spine | 2 | April 27, 2011 | 2011 | DePuy Spine, Inc. | Mislabeled: Package labeled (11g x 6 inch) contains a 13g x 4-inch side hole needle | PRODUCTION CONTROLS: Packaging Process Control | The firm, DePuy Spine, sent an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated March 22, 2011 to its consignees/customers ( hospital sites and domestic sales distributors). A product reconciliation form was provided to assist in the retrieval of affected product. The letter described the product, problem and actions to be taken. The customers were instructed to return any needles that they may have to their local Distributor for DePuy Spine products and complete and return the Product Reconciliation Form via fax to 508-828-3762; call Customer Service at 877-379-4871 to receive a return goods authorization number and/or mail the form with the return product to DePuy Spine Inc., 50 Scotland Boulevard, Bridgewater, MA 02324, Attn: CONFIDENCE NEEDLE RECALL COORDINATOR, RGA #.If you have any questions or concerns with regard to this notice, please contact the DePuy Spine Regulatory Compliance Department at 508-828-2790. | No consumer action necessary | 215 | 215 units | Worldwide distribution: USA including states of: AZ, IA, LA, TX, NE, NJ ,NY, PA, VA, and WA; and countries of: Canada, Australia, Belgium, France,Hungary, Italy, Malaysia, Solvenia, UK, South Africa, Switzerland, and UAE. | Terminated | February 20, 2013 | 665 | 282 | Not_Computer
Z-2082-2011 | 58291 | epstein-barr virus, other | LSE | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | BioPlex 2200 EBV IgG Calibrator Set. Catalog #663-1200.Product is labeled in part: "BIO-RAD***BioPlex 2200 EBV IgG Calibrator Set***7 x 0.5 mL***A Calibrator set for the BioPlex 2200 EBV IgG Reagent Pack***UNITED STATES, Bio-Rad Laboratories, Inc., Hercules, CA 94547".BioPlex 2200 EBV IgG Calibrator Set is intended for the calibration of the BioPlex 2200 EBV IgG Reagent Pack. | BioPlex 2200 EBV IgG Calibrator Set. | 2 | April 27, 2011 | 2011 | Bio-Rad Laboratories | Bio-Rad Laboratories confirmed a downward shift in control values for EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set Lot #46704. | OTHER/UNDETERMINED: Pending | The firm, Bio-Rad Laboratories, sent an "Voluntary Field Correction: BioPlex 2200 EBV IgG Calibrator Set, Lot# 46704 for EBV NA-1 Analyte" letter dated March 16, 2011 domestic customers via Fed-Ex and and to international subsidiaries via e-mail. The letter described the product, problem and actions to be taken. The consignees/customers were informed that Bio-Rad confirmed downward shift in control values for the EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set, Lot # 46704. If this occurs, the customers were instructed to readjust the positive control mean using the procedure presented in the customer notification "BioPlex 2200 Quality Control Values (L60013402)". If the re-adjusted mean of the EBV NA-1 controls is within the manufacturer's limit, patient results may be reported. If the new assay positive control mean still does not fall within the manufacturer's limit, the customers were instructed to contact their local Bio-Rad Technical Support and do not report the EBV NA-1 values. The customers were also instructed to complete and return the CUSTOMER RESPONSE FORM via fax to 510-741-5775.If the BioPlex 2200 EBV VCA IgG and EBV EA-D IgG positive and negative control means are within the acceptable QC range, and subsequent passing QC runs are obtained, patient results may be reported for these analytes. Bio-Rad will be manufacturing a new lot of the BioPlex 2200 EBV IgG Calibrator Set which will bring the control values within the manufacturer's limit. It may be necessary to readjust the QC ranges after the consignees receive this new calibrator set lot.Bio-Rad will send their consignees/customers the replacement calibrators and calibrator CDs when available. The consignees/customers were also instructed to discard the calibrator set lot# 46704 in accordance with national and/or local regulations.Customers can call Clinical Diagnostics Group at 510-741-4431 or email to BPD-ProductSupport@bio-rad.com for any questions . | Contact the recalling firm for information | 270 | 270 sets | Worldwide distribution: USA including the following states: AL, AZ, CA, CT, FL, GA, HI, IL, IN, LA, MA, MD, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX, VA, WA, and WI; and countries of: Australia, Czech Republic, France, Germany, Italy and Republic of Korea. | Terminated | December 22, 2011 | 239 | 270 | Not_Computer
Z-2067-2011 | 57912 | system, photopheresis, extracorporeal | LNR |  |  | Gastroenterology/Urology |  | Therakos UVAR XTS Procedural KitsIs indicated for use in the ultraviolet-A (UVA) irradiation. | Therakos UVAR XTS Procedural Kits | 2 | April 28, 2011 | 2011 | Therakos Inc | Therakos received an increased rate of complaints of centrifuge bowl leaks for certain lots of Therakos UVAR XTS Procedural Kits (Product Code XT 125). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Therakos notified all direct Therakos US and ex-US End Users and Distributor consignees via Federal Express courier with an Urgent Product Correction Notification, dated 2/1/2011. The letter identified the reason for recall along with the affected product. Customers were to immediately discontinue using and discard the identified affected lots of the product. Customers were to use alternative lots. In order to receive credit for the discarded product, customers were to complete the attached confirmation of receipt and destruction. If there were any questions, a question and answer section was provided but additional questions could be directed towards Therakos Customer Service Center at 1-800-828-6316, Option 2,1. |  | 4323 | 4,323 kits | Worldwide Distribution | Open |  |  | 4323 | Not_Computer
Z-2112-2011 | 58273 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Portable Blood Gas Analyzer System, Part Number: 393-090, 510 (k) #K092686. All units equipped with software versions: 2.4.1680.25 and all previous versions are subject to recall. The ABL90 FLEX is a portable, automated analyzer that measures pH, blood Gases, electrolytes, glucose, lactate and oximetry in whole blood. The ABL90 FLEX is intended for use by trained technologists, nurses, physicians and therapists. It is intended for use in a laboratory environment, near patient or point-of-care setting. | ABL90 FLEX | 2 | April 28, 2011 | 2011 | Radiometer America Inc | Following the receipt of two (2) reported customer complaints, the firm determined that the current operating software for their ABL 90 Analyzer does not prevent the further use of the sensor cassette when there is the possibility of iron rich plaque build-up in the reference electrode. If the device operator/user ignores the error message and continues to obtain biased measuring results using the | DESIGN: Software Design | The firm, Radiometer America, sent two Urgent "Medical Device Field Correction" letters issued on February 4, and on February 24th, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to replace the sensor cassette. In order to prevent this situation from occurring, the firm informed the customers in the recall/correction notification that they intend to upgrade all ABL 90 Flex analyzers to a new software version. The new software will perform a check during a normal system FLUSH in order to allow the analyzer to detect the bias and will notify the customer immediately of this condition and will automatically disable the cassette thus preventing further use. The letter goes on to remind the device customers that until such time as the software upgrade is installed, they should make sure to adhere to the procedures stated in the Reference Manual which call for the replacement of the sensor cassette immediately upon receiving message #1310. The letter further instructs the customer to complete and return the attached FAX RETURN FORM via fax to 440-871-0463. The letter also states that the expected release date for the new software should be sometime in the second half of February 2011. If you have any questions regarding this letter, or the upgrade, please contact Technical Support staff at 1-800-736-0600 Opt 4. |  | 32 | 32 units were distributed in the U. S. and 533 units were distributed world-wide | Worldwide distribution: USA including states of CA, DC, PA, MN, NY, NJ, TX, OH, MO, WI, UT, and MI; and country of Canada. | Terminated | May 09, 2012 | 377 | 32 | Software | Output/Calculation | Replace sensor cassette/Software update | Software Update
Z-2104-2011 | 58342 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 75 MM fixed I-Beam Tibial Plate with locking bar CO-CR-MO/TI 6AL 4V ALLOY, REF 141224, Sterile, BIOMET ORTHOPEDICS, WARSAW. IN I-Beam Tibia Plates are knee joint replacement prostheses intended for application with or without bone cement. | Biomet CC IBeam Tibial Plate 75mm | 2 | April 28, 2011 | 2011 | Biomet, Inc. | The firm is recalling the product due to the product's packaging being labeled "I-Beam 75mm. Tibial Plate", but the package actually contains the "Interlok 67mm. CR Tibial Plate". | PRODUCTION CONTROLS: Packaging Process Control | Biomet Orthopedics Inc initiated an electronic communication with Direct Accounts on March 3, 2011 and instructed them to quarantine and return all remaining/unused devices to the firm. For questions regarding this recall call 574-372-3983. | No consumer action necessary | 8 | 8 | Worldwide Distribution -USA including FL, IN, and NC and the countries of Japan and the Netherlands | Terminated | July 27, 2011 | 90 | 8 | Not_Computer
Z-2106-2011 | 58407 | reagent, occult blood | KHE | Occult blood test. | Hematology | Hematology | 510(k) | ClearviewÂŽ iFOB, ClearviewÂŽ iFOB Complete- Immunochemical Fecal Occult Blood Cassette Test Product Part Number:8111KCV (30 Test Kit ), Intended Use: Immunochemical device intended for the qualitative detection if fecal occult blood in human feces by laboratories or physician offices | Clearview iFOB Complete (30 Test) | 2 | April 28, 2011 | 2011 | Alere North Americal | Potential for false negatives test results and negative/faint positive controls | OTHER/UNDETERMINED: Pending | Alere sent a Voluntary Product Recall letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to indicate zero on the attached Recall Reply Form, sign, and fax to 1-877-877-0147, if they have no remaining product.For any remaining product customers were instructed to discontinue use and indicate the quantity of kits in the attached Recall Reply Form so that the firm can arrange for them to receive credit. Customers were instructed to destroy any affected units at their site. For any further distribution of product, customers will need to be contacted. Physicians and hospitals were advised to retest patients with negative results. Questions relating to this Voluntary Product Recall call 1-877-597-9586. |  | 6249 | 6249 kits | Nationwide Distribution | Terminated | March 07, 2013 | 679 | 15543 | Not_Computer
Z-2105-2011 | 58255 | stimulator, nerve, transcutaneous, for pain relief | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Inter x 5002 Device, Inter x Professional Sport Device, Inter x 1000 Personal Device, and Inter x Personal Sport packaged with Triad alcohol wipes. The alcohol wipes manufactured by the Triad Group were labeled under brand names Triad, Select, Select Medical Products, PSS Select or PSS World Medical Inc.Alcohol wipe is provided as a cleaning convenience for device electrodes. | InterX 5002, Sport, 1000, & Personal Sport packaged with Triad Alcohol Wipes | 2 | April 28, 2011 | 2011 | Neuro Resource Group Inc | Product package contains alcohol prep pads recalled by Triad Group because the prep pads have the potential to be contaminated with an objectionable organism. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Triad Group issued a press release on 2/7/2011, and notified consignees by e-mail 3/2/2011. Neuro Resource Group, Inc. ( NRG ) removed the recalled Triad wipes from all kits in house and obtained non recalled wipes from an alternate supplier. For replacement wipes customers may contact NRG customer service at 1-877-314-6500. | Contact the recalling firm for information | 57252 | 57,252 wipes | Worldwide Distribution | Terminated | September 23, 2011 | 148 | 57252 | Not_Computer
Z-2118-2011 | 53013 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 Base System, model number 10042832. Distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA.The ACUSON Antares is a general purpose Diagnostic Ultrasound system. Its function is to acquire primary or secondary harmonic ultrasound data and display it in: B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, Harmonic Imaging, or 3D/4D imaging on a FPD display. The Antares ultrasound imaging system is intended for the following applications: Abdominal, Intraoperative, Small Parts, Transcranial, OB/GYN, Cardiac, Transesophageal, Pelvic, Neonatal/Adult Cephalic, Vascular, Intravascular, Musculoskeletal, Superficial Musculoskeletal, Great Vessel, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. | ACUSON Antares System | 2 | April 29, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Shock Hazard. The manufacturer failed to set the system power switch to "ON" during the line level leakage and high pot current draw testing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Siemens Medical Solutions USA, Inc. sent a Customer Safety Advisory Notification letter out on May 15, 2009. Service engineers were to visit each customer to conduct necessary testing.For questions regarding this recall call 650-694-5993. |  | 1 | 1 unit | Worldwide Distribution - USA including CA, IL, MD, NC, OH, and UT and the countries of Australia, China, Denmark, Spain, South Korea, Norway, Poland, Russia and Sweden. | Terminated | April 29, 2011 | 0 | 30 | Not_Computer
Z-2017-2011 | 58188 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Shear Valve Assembly used on CELL-DYN 3200CS Analyzer and CELL-DYN 3200CS Analyzer (refurb).Abbott Diagnostics Division, Santa Clara, CA 05054Designed for in vitro diagnostic use in clinical laboratories. | Abbott CellDyn 3200 System | 2 | April 29, 2011 | 2011 | Abbott Laboratories | The Nylon Washer has an outer diameter which is undersized resulting in the inability to generate patient results. |  | Product Correction letters were sent on March 15, 2011. via Federal Express. The letter identified the affected product, explained the problem, the patient impact and the necessary action that were to be taken. Customers were to review the letter and inform their Laboratory Manager of the pending action.; complete and return the Customer Reply Form; and save a copy of the letter for their records. If there were any questions, customers were to contact Customer Support at 1-877-4ABBOTT or their local hematology customer support representative for customers outside of the US. | Contact the recalling firm for information | 174 | 174 total units, all varieties | Worldwide Distribution -- AZ, CA, FL, IA, ID, IN, MA, MI, MN, MS, ND, NY, OK, OR, SC, SD, TX, WI, WY and Germany, Singapore, Mexico, China, South Korea, Argentina, Japan, Brazil, Canada, Australia, and Chile. | Terminated | November 22, 2011 | 207 | 174 | Not_Computer
Z-2122-2011 | 58482 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6201, Single Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8063The Flo-Gard Volumetric Infusion Pumps are an electromechanical pump used for intravenous infusion of fluids at user-selected rates. | Baxter FloGard 6201 Volumetric Infusion Pumps | 2 | April 29, 2011 | 2011 | Baxter Healthcare Corp. | Pumps submitted for service may not have been properly serviced before being returned to the customer. | TRAINING: Employee Error | Baxter visited the direct accounts starting on 3/24/11 informing them that the affected Flo-Gard pumps may have been released from the preventative maintenance with incorrect accuracy testing. The affected pumps were located, taken out of service and placed in quarantine on the same date. A technician not involved with the previous testing visited both accounts on 3/25/11 and performed preventative maintenance service of the three affected devices, and were found to be within specification. | Contact the recalling firm for information | 3 | 3 units | Puerto Rico | Terminated | April 29, 2011 | 0 | 3 | Not_Computer
Z-2114-2011 | 58282 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF) | UniCel DxH 800 Coulter Cellular Analysis System | 2 | April 29, 2011 | 2011 | Beckman Coulter Inc. | For Panels containing CBC: The RBC and HCT are improperly corrected and not flagged when there is a WBC Pump Failure and the WBC ~ 140x103/uL. For Body Fluid Panels:The RBC is improperly corrected and not flagged when there is a WBC Pump Failure. IMPACT:Erroneous RBC and/or HCT could affect the interpretation and management of patients at risk with various types of anemias. | OTHER/UNDETERMINED: Pending | The recall communication was initiated on 10/20/2010, with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with PCA Response Form (via US Postal Service for US customers) to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter informs the customers about the problems identified and provides an action/resolution to this issue. When a WBC Pump Failure occurs, the WBC or TNC results will be flagged with 'R' and the WBC Pump Failure System message. The RBC and/or HCT results associated with these affected analyses also require review. This will be corrected in a future version of software.Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative. | Contact the recalling firm for information | 422 | 422 units in total (175 in US) | Nationwide in the US and worldwide to: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Open |  |  | 422 | Software | Alarm/Message | Software update | Software Update
Z-2119-2011 | 58230 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The 4D Integrated Treatment Console; Model Number: H51; Reference: 4DITC versions 8.6 and below; File Reference: CP-2007-0288; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The 4D Integrated Treatment Console function is designed to assist the operator of radiation therapy device in providing accurate treatment setups for each patient by monitoring setup parameters and preventing the radiation therapy device from commencing irradiation when any parameter is out of conformance with the treatment plan. | 4D Intergrated Treatment Console | 2 | April 29, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly, in that the Multileaf Collimator (MLC) message displayed on the 4DITC on versions 8.6 and below, does NOT interlock beam delivery. | DESIGN: Software Design | Varian Medical Systems sent an Urgent Medical Device Correction Letter on February 28, 2011, to all affected customers.Varian has created the procedure described in CTB-GE-579 to inform users on proper interfacing of the demountable BrainLAB m3 on the Varian Clinac. The procedure ( CTB-GE-579) is enclosed and posted to the customer site: http://www.MyVarian.com.This design fix in 4DITC v10.1 or later will be provided to all affected sites.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager.USA and Canada - 1-888-VARIANS (888-827-4265)Europe - + 41 41 749 8844. | Contact the recalling firm for information | 17 | 17 | Class 2 Recall - Worldwide Distribution -- USA including states of ; CA, NV, NJ, GA, IL, and TX and countries of Australasia, Belgium, Brazil, China P.R., Denmark, France, Germany, Germany ( BrachyTherapy) Hong Kong, India Pvt. Ltd, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland and United Kingdom. | Terminated | March 30, 2012 | 336 | 17 | Software | Physical Safety Hazards | Software update | Software Update
Z-2113-2011 | 58367 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Doctors Review System, OB-Windows/OB-View software release numbers 4.1 through 4.1.3, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098OB-Windows provides professional reporting and analysis of fetal growth throughout the gestational period, which can help evaluate clinical situations. OB-View provides image management for the user to capture, view, read, and edit images. | DOCTORS REVIEW SYSTEM | 2 | April 29, 2011 | 2011 | Digisonics, Inc | Study would on occasion have fetal heart rate and/or cervical length on the report that was not entered nor did it belong to that patient. Software displayed incorrect data for fetal heart rate and cervical length in OB ultrasound reports. | OTHER/UNDETERMINED: Pending | The firm, Digisonics Inc, notified the affected customers of software release via e-mail on January 24, 2007 and follow-up e-mail on February 7, 2007. The e-mails described the product, problem and the actions taken. The firm was able to trap the error and after this report the developer was able to determine the cause and a new executable was sent and installed at the customers site.If you have any question, please call 800-940-3240 or e-mail: service@digison.net. | Contact the recalling firm for information | 128 | 128 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | May 02, 2011 | 3 | 128 | Software | Output/Calculation | Software update | Software Update
Z-2121-2011 | 58286 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF) | UniCel DxH 800 Coulter Cellular Analysis System | 2 | April 29, 2011 | 2011 | Beckman Coulter Inc. | RETIC results may be invalid when the instrument user interface allows reflex processing of a pre-diluted specimen in the RETIC test panel. Beckman Coulter has validated pre-diluted specimen processing in CBC panels. RETIC test panel using pre-diluted samples has not been validated.IMPACT:The reticulocyte value is an indicator for assessing red cell production. Erroneous results could impact c | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. | No consumer action necessary | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Open |  |  | 422 | Software | Output/Calculation | Software update | Software Update
Z-2120-2011 | 58519 | pulse generator, permanent, implantable | NVZ |  |  | Cardiovascular |  | Medtronic, Kappa 600KDR600,KDR601,KDR603,KDR606,KDR651,KDR653The devices are implantable pulse generators (IPG) used to detect bradycardia events and provide pacing support. Within each product family, there are varying features, rate responsive pacing, and diagnostic data. Medtronic,Kappa 700KD700,KD701,KD703,KD706,KDR700,KDR700V,KDR701,KDR701PTO,KDR701V,KDR701VIT,KDR703,KDR703V,KDR706,KDR706V,KDR720,KDR721,KDR730,KDR731,KDR733,KVDD700,KVDD701Medtronic,Kappa 800KDR801,KDR803,KDR806Medtronic,Kappa 900KD901,KD903,KD906,KDR901,KDR903,KDR906,KDR921,KDR931,KDR933,KVDD901Medtronic, Enpulse DRE1DR01,E1DR06,E1DR21,E2D01,E2D03,E2DR01,E2DR03,E2DR06,E2DR06J,E2DR21,E2DR21PJ,E2DR31,E2DR31J,E2DR33,E2DR33J,E2VDD01Medtronic, Adapta/Versa/Sensia DRADD01,ADDR01,ADDR03,ADDR06,ADDRL1,ADDRS1,ADVDD01,SED01,SED01N,SEDR01,SEDR01N,SEDRL1,VEDR01Medtronic, ReliaRED01,REDR01,REVDD01(approved in the US but has not been sold in US)Medtronic, Vitatron Extension G and E series DRE50A1,E60A1,G70A1(not sold in the US) | KAPPA 600/700/650/800/900/RELIA/SIGMA/MEDTRONIC 350/ENPULSE/ENRHYTHM/AT500/ADAPTA/VERSA/SENIA IPGS | 2 | April 29, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | A rare measurement lock-up condition that may (at a rate of approximately 1 in 18,000 devices) inappropriately trigger ERI/RRT in the Medtronic Dual Chamber Pacemakers: Kappa 600, 700, 800, 900, EnPulse, Adapta, Versa, Sensia. This is caused by a random lock-up of the measurement system hardware that may result in an incorrect battery voltage reading of zero. This issue does NOT impact battery lon | DESIGN: Component Design/Selection | The firm, Medtronic, sent an "Important: Medical Device Correction" letter dated April 2011 to its customers. The letter described the product, problem, and actions taken. Medtronic informed the customers that this issue does NOT impact battery longevity and does NOT require device explant. Currently, the device can be reset to normal operation by a Medtronic representative. Reset devices are no more likely to experience a recurrence of this issue. Medtronic is planning to release a programmer software update that will allow the clinician to reset the device. In addition, Medtronic has consolidated their product performance data with the vision of creating a primary reference tool for performance updates.If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 1989180 | 1,989,180 | Worldwide distribution: USA (nationwide) including Guam, Puerto Rico and DC; and countries including: Algeria, Andorra, Anguilla, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei, Darussalam, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Virgin Islands (British). | Terminated | November 27, 2012 | 578 | 1989180 | Hardware | Device Operation | Programmer software update | Software Update
Z-2127-2011 | 58313 | tomography, optical coherence | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Cirrus HD-OCT Model 4000 instruments, new and refurbished, distributed with, or upgraded to, version v 1.0 - 4.5, 5.0.0 or 5.1.0The Cirrus HD-OCT is a non-contact, high resolution tomographic and biomicroscopic imaging device. | Cirrus HDOCT Model 4000 instruments | 2 | May 02, 2011 | 2011 | Carl Zeiss Meditec Inc | A software defect exists in versions 5.0.0 and 5.1.0 where when following a specific workflow, exam data may be saved under the name of the first patient | DESIGN: Software Design | Carl Zeiss Meditec, Inc. sent an Field Correction Action letter dated March 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letters requested acknowledgement of the letter and a response and the fact that a follow up packet with revised software will be sent to all customers with V 5.0.0 or 5.1.0 software.For questions, please call 1-855-853-9461. | No consumer action necessary | 6888 | 3465 US; 3423 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Boliva, Brazil, Canada, France, Chili, China, Columbia, Denmark, Ecuador,Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Japan, Lebanon, Malaysia, Mexico, New Zealand, Pakistan, Peru, Singapore,South Korea, Spain, Sri Lanka, Taiwan, Uruguay and Venezuela | Terminated | March 29, 2012 | 332 | 6888 | Software
Z-2125-2011 | 58424 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, RT Chart Application v 8.0 through 8.9Varian Medical Systems, Palo Alto, CA.Aria Radiation Oncology product is a treatment plan and image management application. It enables the authorized user to enter, access, modify, store and archive treatment plan and image data from diagnostic studies, treatment planning, simulation, plan verification and treatment. | ARIA, RT Chart Application v 8.0 through 8.9 | 2 | May 02, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue was identified with the gantry rotation direction when creating opposing field within the RT Chart function of ARIA, or possibly with third party planning systems and delivering the plan at 4DITC. | DESIGN: Software Design | Varian Medical Systems, Inc. Oncology Systems sent Urgent Medical Device Correction Field Safety Notice letters on March 9, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letters identifeid the product, the problem, and the action to be taken by the customer. Customers were instructed to:Manually enter the gantry angel 0.0 when the plan indicates angel 0.0. Manual entry is only allowed to one decimal place, so the range that triggers the issue cannot be created if the keyboard is ued to enter the gantry angel.Users with third party plan sources should verify whether gantry angles with two or more decimal places are supported. If so, customers may be affected by this issue, and should take steps suited to the device.In all cases, the display at the Treatment console is indicative of what direction the gantry will move - users are advised to be especially attentive to these values, particulary when non-zero couch rotation values are used.Users are advised to always run a simulation of the plan, including all plan motions, at the treatment unit prior to first patient treatment. For further clarification customers were insturcted to contact their local Varian Customer Support District Manager. For questions regarding this recall call 650-424-6471. |  | 612 | 612 units | Worldwide Distribution - USA including CA, DC, FL, GA, MD, MN, MS, NY, OH, and TN and the countries of Canada, Australia, Austria, Bahrain, Bagladesh, Belarus, Belgium, Brazil, Columbia, Costa Rica, Cypress, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Luxemborg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion, Romania, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tobago, Tunisia, Turkey, Ukraine, Emirates, and the United Kingdom. | Terminated | March 13, 2012 | 316 | 612 | Not_Computer
Z-2128-2011 | 58287 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF) | UniCel DxH 800 Coulter Cellular Analysis System | 2 | May 02, 2011 | 2011 | Beckman Coulter Inc. | Body Fluids results may be incorrectly interpreted by the Laboratory Information System (LIS). Body Fluid results ~ 1,000 celis/uL are reported using the US numeric separators to demark thousandths places (comma, ",") and partial units (decimal, "."). An LIS that uses non-US formats could interpret the comma as a partial units separator, which could lead to reporting erroneously low body fluid r | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. | Contact the recalling firm for information | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 422 | Software | Output/Calculation | Software update | Software Update
Z-2123-2011 | 58574 | implant, intragastric for morbid obesity | LTI |  |  | Gastroenterology/Urology |  | Lap-Band Access Port Needle 10-Pack, 20 G x 51 mm (2 in.), Product Code: 8-20302-10, Lot Number: 9L01N20 gauge, 51 mm (2 in.) length non-coring deflected tip (Huber) needle, used with the LAP-BAND System for band adjustment. | LAPBAND ADJUSTABLE GASTRIC BANDING (LAGB) SYSTEM | 2 | May 02, 2011 | 2011 | Allergan | The recall was initiated because Allergan has determined that a number of Access Port Needle pouches from one specific lot number may have small gaps in a pouch seal. Consequently the product may no longer be sterile. |  | Allegan sent an Urgent Medical Device Recall letter dated March 21, 2011, to all affected customers. Allergan requested that upon receipt of letter, that the customer check their existing inventory of product for affected lot. Customers were requested to complete Business Reply Card, stating quantity of inventory of affected lot, if any, and return to Allergan Medical. Customer with remaining inventory were instructed to contact Allergan Medical Product Support at 1-800-624-4261, x5972 or 1-805-961-5972 for return and immediate replacement. . Customers will be issued credit for returned product.For questions regarding this recall call 805-319-0589. | Contact the recalling firm for information | 1888 | 1878 x 10-packs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, MN, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Australia, Canada, Germany, Great Britian, and the Netherlands | Terminated | April 06, 2012 | 340 | 1878 | Not_Computer
Z-2124-2011 | 58470 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BardÂŽVentralex" Hernia Patch Medium Circle with Strap Diameter 6.4cm /2.5",\Catalog #0010302, lot number HUTL1283.The BardÂŽVentralex" Hernia Patch is a self-expanding, non-absorbable, sterile prosthesis. The device classification name is "Mesh, Surgical, Polymeric". It is intended for use in all forms of hernia repair requiring reinforcement with a non-absorbable support material | BardVentralex" Hernia Patch | 2 | May 02, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Medium Ventralex Hernia Patch 6.4cm /2.5" packaged in a pre-printedcarton identified for a BardÂż Ventralex Small Circle 4.3cm /1.7" | PRODUCTION CONTROLS: Labeling Mix-Ups | Davol sent an Urgent: Medical Device Recall letter on April 11, 2011, via FedeX Priority Overnight to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and remove and return the recalled lot. Distributors were requested to notify their accounts.For questions contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services & Support Department at 1-800-562-0027. | No consumer action necessary | 146 | 146 units | Nationwide Distribution - USA including AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, MT, NC, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, and WV. | Terminated | October 27, 2011 | 178 | 146 | Not_Computer
Z-2135-2011 | 58325 | starter, bone screw | HWD | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Low Profile Self-Tapping Bone Screw, 6.5 mm X 25 mm, Qty. 1, Biomet Orthopedics, 66 East Bell Drive, P.O. Box 587, Warsaw, IN 46581 USA. STERILE.To provide fixation of any acetabular cup, with appropriate sized screw holes, to the pelvic bone during total hip replacement procedures. | Low Profile SelfTapping Bone Screw, | 2 | May 03, 2011 | 2011 | Biomet, Inc. | One lot of 30 mm screws was switched with one lot of 25 mm screws. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Biomet Orthopedics sent an Urgent Medical Device Recall Notice letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove the affected lots from their inventory, and return them to Biomet. Complete and enclosed FAX Back Response Form to 574-372-1683 prior to return of product. Distributors were instructed to notify all sub-accounts of the recall.Confirm receipt of notice by calling, 800-348-9500, extension 3755 or 3983.For questions regarding this recall call 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 100 | 100 | Worldwide Distribution - USA including CA, GA, IA, IL, IN, KS, KY, MD, MI, MN, MO, NC, ND, NJ, NY, OK, PA, SD, TX, VA, WA, and WI and the country of Japan. | Terminated | February 17, 2012 | 290 | 100 | Not_Computer
Z-2134-2011 | 58402 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Patient Monitor, Model 91387. All modules may be affected except Models 90470, 90478, 90496, & 91496.Affected model number is 91387 shipped after February 10, 2010 with PCBA 670-1298-01 (serial numbers beginning with 1387-1xxxxx).The data sheet is labeled in part: "***91387-27 SL2700 Patient Monitor***91387-28 SL2800 Advanced Patient Monitor***91387-38 SL3800 Central Monitor***Spacelabs Healthcare ***Ultraview SL***The Ultraview SL 2700 and SL2800 bedside monitors include interactive networking, graphic and tabular trends, remote view and alarm watch support. They accommodate two single-high parameter modules, use external flat panel displays, and include a DC power supply that provides power for the monitor and up to three FlexoriaÂŽ system interfaces. The SL3800 central monitor includes interactive networking, graphic and tabular trends, remote view, alarm watch, and central alarm manager. The central monitor accommodates a dual-high recorder module, uses an external flat panel display, and includes a DC power supply. In addition to central monitoring, the SL3800 provides bedside monitoring operation through Biomed-level control***".Spacelabs Healthcare patient monitors, functioning as either bedside or central monitors; passively display data generated by Spacelabs Healthcare parameter modules, Flexport interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. Key monitored parameters available on the models 91367, 91369, 91370 and 91387, when employing the Spacelabs Command Module, consist of ECG, respiration, invasive and noninvasive blood pressure, Sp02, temperature and cardiac output. Additional parameters and interfaces to other systems are also available depending on the parameter modules employed. Spacelabs Healthcare patient monitors are intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other physiologic monitors via Flexport interfaces. These devices determine: a) when an alarm condition is violated; b) the alarm priority (i.e. high, medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Healthcare patient monitors may also function as a generic display or computer terminal. As a generic display or terminal, the patient monitors allow net worked based applications to open Windows and display information on other net worked monitors. Spacelabs Healthcare patient monitors are also designed to communicate with a variety of external devices such as displays, network devices, serial devices, user input devices, audio systems, and local/remote recorders. Spacelabs Healthcare patient monitors are intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. | Ultraview SL Patient Monitor, Model 91387 | 2 | May 03, 2011 | 2011 | Spacelabs Healthcare, Llc | The circuit that supplies backup power to the modules in the Medical Patient Monitor, Model 91387, is not working. If there is a power interruption, the three-minute backup may not be operational and customers may lose any changes they have made to the alarm settings. | OTHER/UNDETERMINED: Pending | SPACELABS Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use their monitors until they can be upgraded. Customers were recommended to brief their staff about the possibility of losing changes to the alarm settings and to restore them after a power interruption. Field service engineers will contact customers to schedule a mutually agreeable time to come to their hospital and upgrade their affected patient monitors at no cost to the hospital. US Consignees can call 1-800-522-7025, Select 2 for Technical Support. Consignees outside of the US can call 1-425-657-7200 x5089 for any questions about this recall. |  | 4470 | 4470 units total (1894 units in US and 2579 units internationally) | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA.AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA. | Terminated | February 15, 2012 | 288 | 4470 | Hardware | Device Operation | Notification until upgrade | Repair
Z-2133-2011 | 58278 | catheter care tray (kit) | OHR | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | Enforcement Discretion | Catheter Insertion Kit (containing, Under pad, PVP Pad, BZK Prep Pad, Sterile Lubricating Jelly, Vinyl Gloves and Poly Zipper Bag).Intended usage: Kit components to be used during Intermittent Urinary Catheterizations | Catheter Insertion Kit | 2 | May 03, 2011 | 2011 | Apogee Medical, Inc | Kits contain PVP (Providone Iodine) prep pads that are being recalled by the Triad Group due to possible micro contamination, Elizabethkingia meningoseptica. | OTHER/UNDETERMINED: Pending | The firm, Apogee Medical LLC, sent an "Urgent Medical Device Recall" letter dated March 23, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their inventory; complete and return the enclosed Product Recall Response Form as soon as possible via fax to (919) 570-9611, mail to: Apogee Medical LLC, Youngsville Commerce Center, 90 Weathers Street, Youngsville, MC 27596, or e-mail to: dpeper@apogeemed.com; cease further distribution and return any product that is the subject of this recall to Apogee Medical LLC. All returned product must be labeled with a returned goods authorization label that has been provided with the letter. Apogee will replace returned product upon receipt of the Product Recall Response Form and identified product.In addition, if the customer distributed the product, they are to identify their customer(s) and notify them at once of this product recall. The notification should include a copy of the this recall notification letter and should instruct the customer to return the recalled product so that it can be returned to Apogee.Should you have any questions regarding this notification, please call the Manager Regulatory Affairs/Quality Assurance at (919) 435-5409. |  | 5503430 | 5,503,430 units | Nationwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NJ, NY, OH, OK, TN, TX, VA and WA. | Terminated | February 28, 2012 | 301 | 5503430 | Not_Computer
Z-2129-2011 | 58383 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | APD Disposable Sets and Extension Sets used with Baxter HomeChoice and HomeChoice PRO Automated Peritoneal Dialysis Systems; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A.The HomeChoice dialysis system is a personal cycler automated peritoneal dialysis system for pediatric and adult renal patients with fill volumes ranging from 60 mL to 3000 mL. | HomeChoice and HomeChoice Pro Cyclers | 2 | May 03, 2011 | 2011 | Baxter Healthcare Renal Div | Clinicians are being reminded of the current warning in The HomeChoice APD Systems Patient At-Home Guide, 07-19-61-244 (October 2,2009), which states on page 3-13 "To reduce the risk of bums, electrocution, fire, or injury to persons: Close supervision is necessary when this product is used by, on, or near children or those unable to care for themselves." Tubing entanglement around the neck may r | OTHER/UNDETERMINED: Pending | Baxter sent Safety Alert letters dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to consider the following safeguards for at-risk patients on HomeChoice Therapy:Have a caregiver present while the patient is performing therapy.Move the HomeChoice cycler to the foot of the bed, if possible, to minimize the likelihood of entanglement that may lead to asphyxiation.Install bed railings for patients at risk for seizures or falls from bed.Complete the attached customer reply form and fax it to Baxter at t847-270-5457.For questions regarding this recall call 888-736-2543, option 2. |  |  |  | Worldwide Distribution - USA including AZ, CA, CT, DC, FL, KY, MO, NC, NM, Ny, OH, PA, TN, TX, and VA and the countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, South Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Qatar, Russia, South Africa, Saudi Arabia, Singapore, Slovania, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingom, and Venezuela. | Terminated | March 08, 2012 | 310 |  | Not_Computer
Z-2136-2011 | 58052 | cable, electrode | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | HealthPort BioAnologics Patient Electrode Sensor Test Kit, Part # SEN-0025-0, 25 kits/cartonEach kit contains 100 tab electrodes and 26 alcohol pads. The electrodes are used to measure bodyfat analysis. The alcohol pads (non-sterile) are used for cleaning the skin to help the electrodes adhere better. There are no instructions for use with the kit. | HealthPort BioAnologics Patient Electrode Sensor Test Kit, | 2 | May 04, 2011 | 2011 | Vermed, Inc. | Vermed kits contain a product component (Triad Alcohol Prep Pads) recalled due to potential Bacillus cereus contamination. | OTHER/UNDETERMINED: Pending | Vermed sent consignee a recall notification letter on February 8, 2011, followed by a phone call and email. The consignee was instructed to respond with a Recall Reply Form via fax and to notify any customers to whom the affected product was further distributed. All recalled alcohol prep pads were to be returned to Triad Group. Firm is replacing Triad alcohol pads in kits with pads from another supplier and has removed Triad from its approved vendor list.For questions regarding this recall call 802-722-9976. | Contact the recalling firm for information | 2500 | 2500 kits total; 500 from each lot except 1000 from Lot 203031 | OR | Terminated | July 25, 2011 | 82 | 2500 | Not_Computer
Z-2138-2011 | 58449 | suture, cardiovascular | DTH |  |  | General & Plastic Surgery | PMA | SharpointTM 4-0 Black Monofilament Non-Absorbable Nylon 18"/DS24mm Reverse Cutting Edged Needle.Intended for use in general soft tissue approximation and/or ligation, including use in cardiovascular, opthalmic, and neurological procedures. | Sharpoint 40 Nylon Suture | 3 | May 04, 2011 | 2011 | Surgical Specialties Corp | Mislabeled - a number of units of the product labeled as 4-0 Black Mono Nylon suture product contain a 4-0 Black Braided Silk suture. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Surgical Specialites sent an Urgent Product Recall Notification letter on March 30, 2011, along with an attached Customer Acknowledgment Form to the affected consignees. It was reported that upon opening a box labeled as 4.0 Black Mono Nylon, it was observed that the package acutally contained 4-0 Black Braided Silk suture. Customers were to review their current inventory, segregate the affected lot number and / or contact their customer accounts that have purchased this specific item / lot code requesting a return of the product to customers' facility. To ensure customers have received this notification and to arrange for credit and the return of identified product, the Customer Acknowledgment Form should be completed and returned via fax to our Quality Assurance department at (610) 404-8488. |  | 2628 | 2,628 | Class 3 Recall - Worldwide Distribution-- USA and Australia. | Terminated | August 04, 2011 | 92 | 2628 | Not_Computer
Z-2137-2011 | 58166 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | NX is the radiographer's image identification and quality control tool. It offers a broad array of features developed especially to meet the needs of the radiographer. Standard tasks are completed quickly and effortlessly with a simple touch screen. In-room integration provides a more convenient workflow for the radiographer throughout the imaging process. NX communicates seamlessly with the hospital's picture archiving and communication system (PACS), radiology information system (RIS) and hospital information system (HIS). It is DICOM compliant and meets IHE guidelines. | "NX" Modality Workstation | 2 | May 04, 2011 | 2011 | AGFA Corp. | When a radiographer is working on two open studies on an NX workstation and, in specific situations, a problem of image/annotation mix-up can occur on the NX system. | OTHER/UNDETERMINED: Pending | AGFA Corp. sent an "URGENT FIEL SAFETY NOTICE" letter via FED-EX to the customers on March 4, 2011. The letter describes the safety alert and mitigation. Customers were instructed to complete and return the Urgent Field Safety Notice Feedback Form by fax to 864-421-1664.For questions regarding this recall call 864-421-1815. | Contact the recalling firm for information | 775 | 775 units | Worldwide Distribution - USA including CA, GA, MS, ND, NM, NY, OK, PA, SC, SD, TX, and VA and the country of Canada | Terminated | March 27, 2012 | 328 | 775 | Not_Computer
Z-2139-2011 | 58326 | dental implant surgical tray (kit) | OFY | Endosseous dental implant accessories. | Dental | Dental | Enforcement Discretion | ***REF IOSS413*** OSSEIOTITE Certain Implant 4 x 13mm*** Sterile using Radiation*** Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Osseotite Certain Implant(s) & Osseotite Tapered Certain Implant(s) | 2 | May 04, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite Certain Implant Device IOSS413. The internal thread was not manufactured correctly. | OTHER/UNDETERMINED: Pending | BIOMET 3i sent an Urgent Medical Device Recall Letter on March 14, 2011, to all affected customers via emial or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told to contact Customer Service at 1-800-342-5454.Customers were instructed to respond with the attached Response Fax or email DomesticComplaints@Biomet.com.If customers still had implant(s) in their inventory, they were to return the implant(s) to BIOMET 3i using the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CD117886A replacement implant(s) will be sent to them one the product is returned.For questions regarding this recall call 561-776-6906 (EST 8:00am to 5:00pm). | Contact the recalling firm for information | 217 | 217 | Worldwide Distribution - USA including AK, FL, GA, NC, NJ, NY, SC and the countries of Canada, Europe, Japan, and South America. | Terminated | January 25, 2012 | 266 | 217 | Not_Computer
Z-2149-2011 | 53211 | [6 empty] | 4Z1c Matrix Array Transducer, model number 10033682, used with ACUSON SC2000 ultrasound systems, model number 10433816, manufactured by Siemens Medical Solutions, Ultrasound Business Unit, Mountain View, CA.The 4Z1c transducer is a special purpose ultrasound transducer designed to help clinicians perform adult echocardiography and pre-operative and intra-operative surgical assessments. Secondary applications include pediatric echocardiography, primarily benefiting older and larger pediatric populations, general abdominal and intra-operative pediatric abdominal imaging and peripheral vessel imaging. | 4Z1c Transducer | 2 | May 05, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Burn risk-- due to coolant leaks, patients may suffer burns or blisters due to increased temperature. | PRODUCTION CONTROLS: Equipment Maintenance | A customer safety advisory notification was sent out to customers on 7/10/09, and a mandatory software update will also be conducted. |  | 42 | 42 units | Nationwide. Foreign: Japan, Italy, South Korea, Singapore, Yemen, Germany, Australia, and Greece. | Terminated | May 05, 2011 | 0 | 42 | Other | Physical Safety Hazards | Software update | Software Update
Z-2146-2011 | 58181 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day Infusor 2 mL/h Portable Elastomeric Infusion System; an Rx sterile disposable drug delivery system; Baxter Healthcare Corporation, Deerfield, IL 60015; REF 2C1071KJPThe Infusor device is indicated for patients requiring slow, continuous intravenous, intra-arterial, epidural, or subcutaneous administration of medications. | Baxter Single Day Infusor 2 mL/h System | 2 | May 05, 2011 | 2011 | Baxter Healthcare Corp. | Some of the Single Day Infusor System 2 mL/h have a foil over-pouch that was incorrectly labeled as the Half Day Infusor System 5 mL/h, product code 2C1073KJP, lot 10H054. | PRODUCTION CONTROLS: Labeling Mix-Ups | Baxter Healthcare Corp. sent Urgent Product Recall letters dated March 15, 2011, via first class mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The accounts were requested to examine all of their inventory locations for both product code 2C1071KJP and 2C1073KJP, lot 10H054, and return the affected lot for replacement or credit. The accounts were also requested to complete the attached customer reply form listing the number of held for return, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Center for One Baxter at 1-800-422-9837. | Contact the recalling firm for information | 3696 | 3,696 units | Worldwide Distribution - USA including AL and the countries of Mexico and Spain | Terminated | March 12, 2012 | 312 | 3696 | Not_Computer
Z-2145-2011 | 58305 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029.The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: (1) Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood; (2) Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial). | UniCel CxH800 Coulter Cellular Analysis System | 2 | May 05, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the DxH 800 may fail to meet the carryover specification for the WBC Differential (DIFF) when the WBC is greater than 70 X 10 to the 3rd power cells/micro L (70 X 10 to the 9th power cells/L). Beckman Coulter indicated the DIFF% carryover is unlikely to have clinical impact. |  | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached Recall Response Form via US Postal Service on February 11, 2011, to all customers who have the affected instruments. Consignees were informed of the reasons for recall, the products affected and that a resolution for these issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with your laboratory staff and retain this notification as part of your Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, Consignees were told to contact Beckman Coulter, Customer Service at (800) 526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Class 2 Recall - Worldwide Distribution -- USA, and the countries of : Canada, Australia, Macao, Belgium, Malaysia, Colombia, Myanmar, Croatia, Netherlands, Czech Republic, Norway, France, Philippines, Germany, Puerto Rico, Greece, Qatar,Hong Kong, Russian Federation, Hungary, Saudi Arabia, India, Singapore, Spain, Israel, taly Sweden, Japan, Switzerland, Korea, Turkey and United Kingdom. | Open |  |  | 566 | Not_Computer
Z-2144-2011 | 58480 | susceptibility test powders, antimycobacterial | MJA | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BACTEC(tm) MGIT(tm) 2-Tube Spare AST Set Carrier, Bag of 3, Catalog 445946. ***Becton, Dickinson and Company, 7 Loveton Circle Sparks, Maryland 21152***The BACTEC MGIT AST system is a rapid qualitative procedure for susceptibility testing of Mycobacterium tuberculosis, from culture, to streptomycin, isoniazid, rifampin, ethambutol, and pyrazinamide. The routine AST set consists of a Growth Control tube and one tube for each drug, as well as a barcoded tube carrier that holds the set. The barcode provides set size and identification. A tube rack capable of holding up to 55 AST tubes in various set combinations is available for tube setup and transport. (Large or small AST carrier sets can be placed in the rack also.) The MGIT AST system is based on growth of the Mycobacterium tuberculosis strain under test as measured by the instrument. AST sets are entered into the BACTEC MGIT 960 instrument and are tested automatically. Analysis of fluorescence in a drug containing tube is compared to the fluorescence of the Growth Control tube to determine susceptibility results. | BACTEC(tm) MGIT(tm) 2Tube Spare AST Set Carrier | 2 | May 05, 2011 | 2011 | Becton Dickinson & Co. | In vitro diagnostic device components are mislabeled. | OTHER/UNDETERMINED: Pending | BD Diagnostic Systems sent an Urgent Product Recall letter on August 31, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter advised customers to discard any carriers or labels that contain the incorrect barcode for replacement. Accounts were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to 410-316-4258.For questions regarding this recall call BD Customer Service at 1-800-675-0908. | No consumer action necessary | 4123 | 4123 total labels | Worldwide Distribution - USA including CA, MA, MD, VA and the countries of Argentina, Colombia, Belgium, China, India, Hong Kong, Peru, Singapore, Thailand, and Trinidad. | Terminated | July 08, 2011 | 64 | 4123 | Not_Computer
Z-2148-2011 | 58307 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: (1) Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood; (2) Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial). | UniCel DxH Coulter Cellular Analysis System | 2 | May 05, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The DxH 800 may fail to meet the carryover specification for NRBC% when the WBC is greater than 50 X 10 to the 3rd power cells/micro L (50 X 10 to the 9th power cells/L) Beckman Coulter indicated the NRBC% carryover is unlikely to have clinical impact. |  | Beckman Coulter sent an Urgent Product Correction letter, dated February 11, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were informed the issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA including AL, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Open |  |  | 566 | Not_Computer
Z-2147-2011 | 53051 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Systems with software versions 3.0.00 and 3.0.02, model numbers 10348531, 10348532 and 10348533, manufactured by Siemens Medical Solutions, Mountain View, CA.Intended for General Radiology, Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Neonatal/Adult Cephalic, Cardiac, Transesophageal, Pelvic, Transcranial, OB/GYN, Urology, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. | ACUSON X300 | 2 | May 06, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Mistreatment risk-- when a user defined formula is used to calculate a value in a vascular exam, data is switched laterally when it is internally transferred and displayed on the summary report within the system. | DESIGN: Software Design | A field correction was initiated and a Customer Safety Advisory Notification was sent out on 7/29/09. The letter identified the affected product and stated that there was a potential for user defined calculations to be incorrect within the vascular package when using the listed software versions of the product. The letter also discussed the steps that customers could take to avoid the potential risk of the issue. A mandatory update software was developed and will be installed free of charge. Customers should have been contacted to schedule a time for the installation. Questions were to be directed to the local service support person. |  | 154 | 154 units | Worldwide Distribution -- Austria, Brazil, China, Columbia, Germany, Denmark, Egypt, Great Britain, Honduras, Ireland, India, Italy Japan, South Korea, Mexico, Malaysia, Peru, Poland, Portugal, Romania, Russia, Sweden, Singapore, Thailand, Turkey, Taiwan, United States, Uruguay, Venezuela, and South Africa. | Terminated | May 09, 2011 | 3 | 154 | Software | Treatment/Delivery/Therapy | Sofrware update | Software Update
Z-1971-2011 | 57383 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D X-ray systems with under-table collimator part number 5234960. The Precision 500D is designed to perform radiographic and fluoroscopic x-ray examinations. | Precision 500D | 2 | May 06, 2011 | 2011 | GE Healthcare, LLC | Use of high force when moving the image carriage/IDD into forward/park position can cause changes to the mechanical alignment of the collimator blades and allow the blades to open wider thanspecified, consequently allowing radiation outside the permitted field without warning to the user. Thisissue manifests with symptoms wherein the user cannot fully close down the lateral blades. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare sent an Importan Electronic Product Radiation Warning letter to customers. The letter discussed the product defect description and related hazards, affected product details, instructions, and product correction. The firm will send a GE Field Service Representative to visit each affected customer site to inspect and correct all potentially affected units in the field following the instructions in field modification instructions No: 10846.If you have any questions please call 800-437-1171. | No consumer action necessary | 440 | 440 US | Nationwide Distribution | Open |  |  | 440 | Not_Computer
Z-2152-2011 | 57044 | absorbent, carbon-dioxide | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allied Healthcare Products Inc. CarbO2Lime Carbon Dioxide Absorbent Granules for Canister Refill, Non-Hygroscopic Soda Lime USP-NF, Type: High Moisture (12% to 19%), Granule Size: 4 to 10 Mesh (4.8-2.0mm), packaged in 5-gal. pails, #55-01-0003; prepacked canisters (12 canisters/case), #55-01-0005; and replacement bags (12 bags/case), #55-01-0006, RX. The label shows the product is Manufactured by Allied Healthcare Products Inc., St. Louis, MO.The product is a non-hygroscopic soda lime compound that is specially formulated to absorb carbon dioxide from the respiratory gas stream in anesthesia applications. | Allied Healthcare Products Inc. CarbO2Lime | 2 | May 06, 2011 | 2011 | Allied Healthcare Products Inc | The granules do not absorb as long as expected in the anesthesia machines | OTHER/UNDETERMINED: Pending | Allied Healthcare Product, Inc. issued recall letters dated August 10, 2011, via certified mail beginning 8/12/10 and/or via e-mail with a read receipt. Letters were issued for both distributors and end-use customers (such as medical facilities or pharmacies) explaining the reason for recall. Distributors were requested to conduct a subrecall using the enclosed customer letter for end users. A response form was included for both distributors and end users to complete to indicate the amount of product to be returned. The customer was instructed to contact the recalling firm for return instructions.For questions regarding this recall call 314-268-1661. |  | 949 | 949 pails, 335/12-canister cases, and 706/12-refill bag cases | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MD, MS, NC, NE, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Canada and Grenada. | Terminated | March 28, 2012 | 327 | 949 | Not_Computer
Z-2131-2011 | 58567 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX20 PRO, Part Number: 476100.Designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. | Synchron LX20 PRO | 1 | May 06, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received reports of incorrect electrolyte results on Synchron LX systems due to a variety of maintenance-related hardware issues. | PRODUCTION CONTROLS: Equipment Maintenance | The recall communication was initiated on 04/05/2011 with Beckman Coulter forwarding a Product Correction letter with attached customer response form to the affected customers via US Postal Service. The letter informed customers of the products affected, the issue, the impact, and the actions to be taken.Customers were instructed to follow the steps provided in order to address the issues related to microbial contamination; and to complete and return the enclosed response form within 10 days so that receipt of this important information could be confirmed. In addition, the letter requested customers share the information with their laboratory staff and retain the notification as part of your laboratory Quality System documentation.For customers who have any questions regarding this update, the letter indicated to contact Customer Support Center Call Center via the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/. They may also contact Beckman Coulter at (800) 854-3633 in the United States and Canada or through your local Beckman Coulter representative.Beckman Coulter will issue two additional recall notification ("UPDATE TO URGENT PRODUCT CORRECTIONSynchron LX20, LX20 PRO, LXi 725 Clinical Systems") beginning on 05/20/2011, one to Laboratory Directors and one to Physicians. The notices will be mailed via US Postal Service on Friday, May 20th. Customers who have subscribed to email notifications will receive an email containing the letters on Monday, May 23th. Both letter informs the Laboratory Directors and Physicians that Beckman Coulter has received reports of erroneous/aberrant sodium, potassium, chloride and calcium results that have been released from the laboratory. Electrolyte shifts can be low or high and may occur above, below or within the normal range.Laboratory Directors were instructed to immediately disseminate the Attachment to this letter (via notification/alert procedure appropriate to your . |  | 1284 | 1284 units total for all products | Worldwide Distribution -- Australia, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Hong Kong, India, Ireland, Italy, Japan, Republic of Korea, Kuwait, Mexico, Netherlands, Oman, Panama, Philippines, Puerto Rico, Saudi Arabia, Singapore, Slovakia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | February 06, 2013 | 642 | 1284 | Not_Computer
Z-2150-2011 | 58504 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000, Model 10434730, ABVS ultrasound systems; Affected sub-components Numbers: (1) 10434968, Boom Welding (for S/Ns 250157, 250162, 250163, 250164, 250166, 250167 and 250169); (2) 10434969, Arm Welding (for S/N's 250171, 250173, 250177); Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Concord, CAThe S2000 Automated Breast Volume Scanner (ABVS) Ultrasound System for imaging the breast allows volumetric acquisition of the breast image. It provides the operator with the option of using a controlled user independent automated scan motion. The resulting 3D display information enables comprehensive image review by the medical professional. In particular, the coronal plane is scanned, which is not easily accessible with hand-held ultrasound transducers. | ACUSON S2000 ABVS Ultrasound System | 2 | May 06, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The weld connecting the ABVS arm to the system base has the potential to fail, causing the arm to fall off of the unit. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Siemens Medical Solutions USA, Inc. initiated the recall via phone calls and e-mail on March 28, 2011, to all affected customers, informing them of the affected product and providing instructions on the recall. Three units S/N 250171, 250173, and 250177 were to be removed and shipped back to Siemens. Siemens Service personnel were to be dispatched to the customer sites to perform validated stress test on units S/N 250157, 250162, 250163, 250164, 250166, 250167, and 250169 at customer sites.For questions regarding this recall call 650-694-5993. |  | 10 | 10 | Worldwide Distribution - USA, Portugal, Germany, Russia, Estonia, Brazil, France, and Saudi Arabia. | Terminated | May 06, 2011 | 0 | 10 | Not_Computer
Z-2153-2011 | 58318 | clamp, circumcision | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision Clamp, models 1.1 cm and 1.3 cm, Circumcision clamp is an instrument used to compress the foreskin of the penis during circumcision of a male infant. | Circumcision Clamp, models 1.1cm and 1.3cm | 2 | May 06, 2011 | 2011 | Surgical Design Inc | Medical device used in surgical procedure is defective and may break during use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Surgical Design, Inc. sent a Customer letter dated February 23, 2011, informing them of the affected product and providing instructions for its specific consignees. The instructions to the User is to please check your inventory to determine whether any Circumcision Clamps bought directly match those on your shelves. Please fill out the form and mail it to Surgical Design at the address on the form. If any clamps from the affected lots remain in your inventory, please ship them back to Surgical Design, noting the number being returned on the Reply form. These are the affected Lots.100210 (1.3cm)100211 (1.1cm) and100060409 (1.3cm)If you have any further questions, please call (703) 541-0196. |  | 6100 | 6100 clamps | Class 2 Recall -- Nationwide Distribution. | Terminated | June 21, 2013 | 777 | 6100 | Not_Computer
Z-2151-2011 | 58441 | device, digital image storage, radiological | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Commander, LANTIS Commander 6.1, LANTIS 8.3 Commander, LANTIS System, LANTIS 6.1 Complete, LANTIS 6.1 Essential, LANTIS 8.3 Complete and LANTIS 8.3 Essential.Distributed by Siemens, Concord, CAMfg by Impac Medical Systems, Sunnyvale, CAElectronic patient records management system | LANTIS Commander | 2 | May 06, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential safety issue in regard to data transfer between LANTIS OIS system and the hospital Information System (HIS) when the HL7 (ADT) interfact is used. | OTHER/UNDETERMINED: Pending | Siemens sent an Urgent Medical Device Correction letter on March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their LANTIS owner manual and contact there Siemens Service Engineer or Application Specialist or contact Lantis4.healthcare@siemens.com.For questions regarding this recall 925-602-8157. | Contact the recalling firm for information | 47 | 47 | Nationwide Distribution. | Terminated | April 19, 2012 | 349 | 47 | Other
Z-2154-2011 | 58379 | enzymatic method, creatinine | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Enzymatic Creatinine (ECREA) Flex(R) reagent cartridges, K1270 for the Dimension Vista(R) 500 SystemThe ECREA method is an in vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista(R) System. Creatinine measurements are used in the diagnosis and treatment of renal diseases, in monitoring renal dialysis, and as a calculation basis for other urine analytes. | Enzymatic Creatinine (ECREA) Flex(R) reagent cartridges | 2 | May 09, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Results may be falsely depressed up to 0.4 mg/dL or falsely elevated up to 0.6 mg/dL across the assay range. | DESIGN: Device Design | Siemens sent an Urgent Field Safety Notice letter dated June 2010 to all Dimension Vista(R) ECREA customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to:Discontinue use of ECREA method and transition to the Dimension(R) Jaffe creatinine (CREA) method.To notify anyone to whom they may have distributed the product.Complete and return the Effectiveness Check Questionnaire attached to the letter by fax to (302) 631-8467.For questions regarding this recall call (302) 631-6299. |  | 1399 | 1,399 | Nationwide Distribution - including AL, AZ, CA, CO, DE, FL, GA, ID, IL, IA, MD, MA, MI, MT, NH, NJ, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI. | Terminated | December 19, 2013 | 955 | 1399 | Not_Computer
Z-2157-2011 | 58308 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: (1) Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood; (2) Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial). | UniCel DxH 800 Coulter Cellular Analysis System | 2 | May 09, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The handheld barcode scanner does not read barcode labels using NW 7 symbology.Impact: Specimens with NW7 labels cannot be read by the handheld scanner. |  | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Consignees were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:If they use NW7 bar code symbiology, the labels can be read using one of the following methods:-At the single-tube presentation station (using the internal bar code scanner)-Enter the bar code information manually at the user interface of the single-tube station-Use cassette presentation.Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (nationwide) and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Open |  |  | 566 | Not_Computer
Z-2179-2011 | 55466 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc. SSH-880CV Aplio Artida; Diagnostic Ultrasound System | Toshiba America Medical Systems Inc Aplio Artida; | 2 | May 10, 2011 | 2011 | Toshiba American Medical Systems Inc | There is a possibility of the system shutting down owing to the power supply unit causing the protection function to activate thus turning off the system prematurely. | DESIGN: Device Design | The recall was initiated on 4/20/10 with Toshiba America Medical Systems (TAMS) forwarding a Medical Device Correction with a Customer Reply Form (via US Postal Service for US Customers) to all customers who purchased the SSH-880CV Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid this problem until corrective action is taken, in case this problem occurs; please contact their Toshiba Representative. If emergency use is necessary, use the following procedure which allows the system to be used temporarily: 1. Turn OFF the breaker at the back of the system and confirm that the error code near the breaker disappears. 2. Turn ON the breaker again to supply power to the system. Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email to raffairs@tams.com. If customers had any questions they can call Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 14 | 14 | USA, including Puerto Rico. | Terminated | May 25, 2011 | 15 | 14 | Other
Z-2178-2011 | 55952 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Beckman Coulter Access Free T3 Calibrators, Part Number A13430 in vitro diagnostic. | Beckman Coulter Access Free T3 Calibrator | 2 | May 10, 2011 | 2011 | Beckman Coulter Inc. | Beckman Coulter has confirmed reports that the Access Free T3 calibrator lot numbers 989902, 989903, 989904, 989905, 989906 and 989907 will show a decline in performance before reaching their stated expiration date.The affected calibrator lots will cause one or more quality control samples to produce out-of-range high result. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent a letter dated December 29, 2009, providing the customers with an explanation of the reason for recall and asked them to take the following actions:- Discontinue use of the lots of reagent listed above.- Review your historical quality control to ensure that the assay performance of these lots has been within your laboratory specifications.NOTE: Out-of-range high quality control results may indicate invalid test results.Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed response form within 10 days so that Beckman can be assured that they received this important notification.Assistance or questions regarding this notification, were directed to TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States andCanada, contact consignees were told to contact their local Beckman Coulter Representative. |  | 2320 | 2,320 kits | Distribution Nationwide and Canada. | Terminated | March 29, 2012 | 324 | 2320 | Not_Computer
Z-2169-2011 | 57446 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Statspin Express 4 Centrifuge Model Number: SSH4Intended Use: Rapid Separation of plasma and serum from primary gel collection tubes | Statspin Express 4 Centrifuge | 2 | May 10, 2011 | 2011 | StatSpin, Inc., dba Iris Sample Processing | Rotors may fail due to wire spokes breaking | DESIGN: Device Design | Iris notified customers by letter in August 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to install the new rotors and discard the older rotors with the wires. There was no need for customers to return the rotors, but the firm requested customers fax or e-mail the enclosed form to verify that they have installed the new rotors. For any questions customers should call 1-800-782-8774. |  | 2090 | 2090 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia Belgium, Czech Republic, Finland, France, Greece, Kuwait, Norway, Turkey,UK, and Norway. | Open |  |  | 2090 | Not_Computer
Z-2162-2011 | 57472 | cannula, ophthalmic | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD VisitecHigh Viscosity Injector 6mm .90mm (20G) [Hammer] Catalog Number: 585228High Viscosity Injectors are used during retinal surgery for fluid infusion. A High Viscosity Injector is placed into the sclera and temporarily sutured into place. | BD Visitec High Viscosity Injector 6mm .90mm (20G) Hammer | 2 | May 10, 2011 | 2011 | Becton Dickinson and Company | Restricted or no fluid flow for Visitec High Viscosity Injector. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Beaver Visitec (BVI) notified customers via telephone and by sending an Urgent Recall Customer Notification letter dated December 7, 2010. The letter identified the product, the problem, and the action to be taken by the customer. ,Customers were requested to return all unused product in their inventory for credit or refund and complete the attached Customer Response Form and return it by fax to 781-893-7957. Questions can be directed to Inside Sales Manager at 781-906-7952. |  | 3990 | 3990 units | Worldwide Distribution - USA (nationwide) and the countries of Israel, Jordan, Russia, Saudi Arabia, South Africa, Turkey, Iraq, Lebanon, Morocco, Tunisia, Syria, India, Argentina, Australia, Austria, New Zealand, China, UK, Germany, Russia, Belgium, Italy, France, Sweden, and Ireland. | Terminated | July 03, 2013 | 785 | 19129 | Not_Computer
Z-2167-2011 | 57594 | [6 empty] | Philips NM3 Monitors, Model 7900PHILIPS NM3 Monitor, English - P/N 1051674PHILIPS NM3 Monitor, English Refurbished - P/N U1051674PHILIPS NM3 Monitor, English, Demo Unit - P/N DU1051674PHILIPS NM3 Monitor, Spanish - P/N 1060462NM3 is a multi-parameter patient monitor that utilizes the FloTrak Elite module to perform some of the calculations. | Philips | 2 | May 10, 2011 | 2011 | Respironics Novametrix, LLC. | Software: two displayed respiratory parameters are incorrect due to errors in the system software: MValv (alveolar minute ventilation) and Vt/kg (tidal volume/patient weight). In both instances, the displayed parameters are higher than actual. | DESIGN: Software Design | Respironics issued Field Action Notification Letter via Fed'x on 1/4/11, notifying customers of the issue and to instruct them on conditions for continued use and return of the device. The device will be updated with revised software to correct the calculation errors. Any questions or need additional information about this recall action, Contact Respironics US Customer Service at 1-800-345-6443 for customers residing in the USA or +011 (203)697-6488 for International customers. |  | 182 | 182 units | NationwideForeign: Belgium, China, Columbia, Egypt, England, India, Italy, Saudi Arabia, and Spain | Terminated | March 07, 2013 | 667 | 220 | Software | Output/Calculation | Software update | Software Update
Z-1678-2011 | 58047 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | 2.0 mL Cartridge/Syringe Piston component of Insulin PumpInsulin infusion | 2.0 mL Cartridge/Syringe Piston component of Insulin Pump | 2 | May 10, 2011 | 2011 | Animas Corporation | Action is being undertaken to address the occurrence of adverse events associated with leaking cartridges. Firm became aware of this issue through customer complaints from pump users reporting the cartridges were leaking. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Animas Corporation sent an Urgent Recall letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Change their cartridge immediately using the instructions provided. Return all of the unused affected product using the enclosed box. Shipments must be returned using the United States Postal Service. For technical questions call Animas Customer Support at 1-855-254-5668 or call the toll free number at 1-8777-937-7867.For any questions about returning the product call 1-877-280-2339.Customers were instructed to notify the firm by logging on to the following website: www.animasnotification.com or call the Call Center at 1-877-280-2339 if they have no product to return. |  | 384180 | 384,180 | Worldwide Distribution - USA (nationwide) and France. | Open |  |  | 384180 | Not_Computer
Z-2165-2011 | 58060 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 4000 PN 00025000000The GEM Premier 4000 is a portable critical care system for use by health care professionals to rapidly analyze whole blood samples at the point of health care delivery in a clinical setting and in a central laboratory. The instrument provides quantitative measurements of pH, pC02, p02, sodium, potassium, chloride, ionized calcium, glucose, lactate, hematocrit, total bilirubin and CO-OXimetry (tHb, 02Hb, COHb, MetHb, HHb) parameters. Total bilirubin can also be quantitated from heparinized plasma samples when analyzed in the tBili/CO-Ox mode. These parameters, along with derived parameters, aid in the diagnosis of a patient's acidlbase status, electrolyte and metabolite balance and oxygen delivery capacity. Total bilirubin measurements are used in the diagnosis and management of biliary tract obstructions, liver disease and various hemolytic diseases and disorders involving the metabolism of bilirubin. In neonates, the level of total bilirubin is used to aid in assessing the risk of kernicterus | Instrumentation Laboratory GEM Premier 4000 | 2 | May 10, 2011 | 2011 | Instrumentation Laboratory Co. | Measurement (amperometric) spike can occur early in cartridge life on the glucose and lactate sensors during patient blood analysis, leading to erroneously highresults | DESIGN: Device Design | Instrumentation Laboratory notified accounts by a Field Safety NotifIcation Letter dated Feb 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use whichever of the following mitigation options works best for their institution. Both options will have the same effect to eliminate the potential impact of a rare spike on the glucose and lactate sensors. Customers were instructed to complete the enclosed Response Tracking Form and return immediately fax to 781-861-4207 or e-mail to ra-usa@ilww.com.Customers were advised to contact their local representative with any questions. |  | 4379 | 4379 units | Worldwide Distribution - USA (nationwide) and the countries of CANADA ISRAEL ITALY APAN KUWAIT MALAYSIA MEXICO NETHERLANDS NORWAY PANAMA POLAND PORTUGAL PUERTO RICO SAUDI ARABIA SLOVENIA SPAIN SWEDEN SWITZERLAND TAIWAN TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and URUGUAY. | Terminated | May 23, 2012 | 379 | 4379 | Other |  |  | Other
Z-2172-2011 | 58234 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System,Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: (1) Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood; (2) Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial). | UniCel DxH 800 Coulter Cellular Analysis System | 2 | May 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: Specimen tubes may fall out of a Type A cassette (PN A53551, kit of 20) when the analyzer inverts the cassette in the mixing station, or when an operator manually inverts a cassetteThe impact:- Specimen tubes may fall out of cassettes and be misplaced inside the instrument- Specimen tubes may break upon impact, causing possible biohazard exposure- Patien |  | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form being sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Customers were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:-Periodically inspect the mix station drip tray for misplaced tubes that may have fallen from the cassette.- Periodically examine the Type A cassettes to determine if the expandable grippers (clips) within the lavender-colored insert are stuck in the open position (tubes will be loose). Consignees were told to contact their Beckman Coulter representative for assistance with questionable grippers.- Clean any spills according to their laboratory standard operating procedures. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain the notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (Nationwide)and to the countries of Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, & United Kingdom | Terminated | November 21, 2012 | 561 | 566 | Not_Computer
Z-2176-2011 | 58252 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03976; Model Number: HZ1; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Optical Guidance Platform, version 2.6 and 2.6.1 | 2 | May 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The anomaly that has been identified with the Optical Guidance Platform (OGP) software may not be always displaying the correct transfer date on the patient file. | DESIGN: Software Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the Optical Guidance Platform (OGP) Instructions For Use on Loading a treatment plan; before positioning the patient, compare the plan details loaded into the OGP with the plan details printed from the treatment planning systems; verify that the patient information, isocenter, gantry and couch angles are identical as requested in the Instructions For Use, and advise the appropriate personnel working in their radiotherapy department of the content of this recall letter.Varian has made changes to the OGP Instructions For Use -Varian strongly recommends you download the most recent version at www.MyVarian.com or request a copy from your local Varian Customer Support District or Regional Manager. Varian also is preparing a software correction which will be applied to affected unites to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager at the follow: USA and Canada - 1-888-827-4265; Europe +41 41 749 8844; E-mail-North America: support-americas@varian.com; Australia/New Zealand: support.anz@varian.com; Europe: support-emca@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com, and Latin America: soporte.al@varian.com; Internet: Oncology Systems customer site - www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 406 | 599 | Software | Output/Calculation | Instructions for verification/Software update | Software Update
Z-2182-2011 | 58466 | bed, patient rotation, powered | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | KCI RotoProne Critical Care Therapy System buckles for securing patient to the surface during proning therapy, Model 209500.Indicated for treatment and prevention of pulmonary complications. | KCI RotoProne Critical Care Therapy | 2 | May 10, 2011 | 2011 | KCI USA, Inc. | Buckles used to secure patient to surface during proning therapy can come open if not properly engaged by user. | DESIGN: Device Design | Firm sent a letter with the "Warning: Improper Insertion of the RotoProne Therapy System Buckle Can Lead To Unintentional Bed Exit" to all consignees on 2/23/2011. The letter identified the affected product and the reason for the recall. Customers were instructed to read the attached product warning regarding the proper method to attach the RotoProne buckles and to post it with the device to ensure continued proper use of the buckles. If customers had questions regarding the proper use of the buckles, they were to contact KCI Customer Technical Service at 1-800-275-4524. |  | 229 | 229 units | Worldwide Distribution - USA and Kuwait. | Terminated | October 24, 2011 | 167 | 229 | Not_Computer
Z-2175-2011 | 58182 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Philips IntelliVue Clinical Information Portfolio software as follows:865047 IntelliVue Clinical Information Portfolio Release D.03; 865209 IntelliVue Clinical Information Portfolio Release E.0 and E.01; 865300 IntelliVue Clinical Information Portfolio NICU Link | Philips IntelliVue Clinical Information Portfolio | 2 | May 10, 2011 | 2011 | Philips Healthcare Inc. | Edits to an order can result in unintended scheduled interventions and a patient may receive additional interventions or medication doses that are not intended. | DESIGN: Software Design | Philips Healthcare notified accounts on 3/16/11 via the Urgent Medical Device Correction Notification. Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel.All customers potentially affected by this issue will receive the software correction as part of the field action. |  | 387 | 387 units | Worldwide distribution: USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Kenya, Latvia, Malaysia, Netherlands, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Open |  |  | 387 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2180-2011 | 58395 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, DigiView and ERS software release 3.6.4.4 through 3.6.5.2+, 2008 Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098, 800-940-3240.DigiView is a digital image management and reporting system with an associated clinical database. It provides capture, review, and storage of images together with report creation and data basing. | DigiView software | 2 | May 10, 2011 | 2011 | Digisonics, Inc | Data from one patient's study had the potential for being put into another patient's report. | DESIGN: Device Design | Digisonics notified end users via e-mail January 20, 2009 and January 21, 2009, that Data from one patient's study was being put into another patient's report. The problem occurred when a System bug was found in DigiView Release 3.6.4.4, and higher versions in which under certain circumstances (specific workflow) a failure to switch to a patient's images occurs. Two software patch versions will be made available to sites on or after February 5, 2009, to correct the reported problem.If you have any further questions, please contact Digisonics Customer Support at (800 ) 940-3240. |  | 70 | 70 units | Worldwide Distribution -- USA including States of : OH, NC, LA, TX, CA, MO, NY, MD, MA, DC, IA, AZ, PA, MI, NC, FL, IN, SC, ID, MN, AR, OK, IL, CT, and countries of Canada, Lebanon, and New Zealand. | Terminated | May 31, 2011 | 21 | 70 | Software | Output/Calculation | Software update | Software Update
Z-2170-2011 | 58168 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Beckman Coulter HbA1c APT (Hemoglobin A1c, Whole Blood Application), Part Number: OSR61177, Lot Numbers: 9390 & 9472Immuno-inhibition test for the quantitative determination of HbA1c (Hemoglobin A1c), in human whole blood, on the Beckman Coulter AU680 with whole blood automated pretreatment (APT) capability only. For in vitro diagnostic use only. The absolute HbA1c and Total Hemoglobin (THb) values generated as part of the HbA1c assay are intended for use in the calculation of the HbA1c/Total Hemoglobin ratio, and must not be used individually for diagnostic purposes. | Beckman Coulter HbA1c APT | 2 | May 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with two lots of HbA1c APT reagent OSR61177.When HbA1c APT reagent lots 9390 and 9472 are calibrated with HbA1c calibrator ODR3032,elevated recovery may be observed for patient samples. Falsely elevated HbA1c values may beinterpreted as increased risk for diabetes or poorly controlled diabetes. Consequent medicalacti |  | Beckman Coulter initiated the recall communication with a Product Corrective Action (PCA) letter with attached recall response form on March 1, 2011 to the affected customers. The letter explained the problem and identified the action to be taken. The PCA also instructed the consignees to:(1) Please discontinue the use of lots 9390 and 9472 and discard according to your local procedures.(2) Please contact your local Beckman Coulter support Organization for replacement reagent.(3) You may need to re-evaluate samples for HbA1c which were tested with either of the listed lots of reagent.In addition, customers were instructed to:(1) share the information with their laboratory staff and retain this notification as part of their Quality System documentation. If any of the affected product was forwarded to another laboratory, a copy of the letter should be provided to them.(2) complete and return the enclosed response form within 10 days so that the firm was assured the customer received the important communication. For any questions regarding this notice, customers were instructed to contact Customer Technical Support at 1-800-854-3633 in the U.S.A and at 00353656831111 in Europe. |  | 373 | 373 kits | Worldwide Distribution - USA (Nationwide) Austria, Croatia, France, Germany, Korea, Philippines, Russian Federation, Thailand, Venezuela, and Vietnam. | Terminated | June 28, 2012 | 415 | 373 | Not_Computer
Z-2181-2011 | 58590 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical System TrueBeam system for stereotactic radiosurgery and radiotherapy. Model number H19; | Varian Medical System TrueBeam | 2 | May 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Position sensor failure mode may result in an inaccurate position calculation. The error manifests itself in one of two possible ways.1. The video returned by one of the two cameras inside the Spectra is all white or all black. The consequence of this failure symptom is that the Spectra stops tracking immediately.2. The video returned by one of the two cameras inside the Spectra is returned b | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian Urgent Medical Device Correction, Urgent Field Safety Notice letters were sent to all affected facilities by certified mail on April 5, 2011. |  | 20 | 20 | Nationwide distribution. | Terminated | June 19, 2012 | 406 | 20 | I/O | Output/Calculation | Notification | Safety Notice/Insructions
Z-2177-2011 | 58380 | enzymatic method, creatinine | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension(R) EZCR Flex(R) Reagent Cartridge. Catalog number DF270B. In vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista(R) System. | Dimension(R) EZCR Flex(R) Reagent Cartridge | 2 | May 10, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Patient results may be falsely depressed or falsely elevated across the assay range. | DESIGN: Device Design | Siemens issued an Urgent Field Safety Notice dated June 2010 to all Dimension(R) EZCR customers; to notify them that Siemens confirmed an issue which many cause EZCR results to be falsely depressed or falsely elevated across the assay range. The letter asks users to discontinue use of this lab test kit. |  | 860 | 860 | Nationwide distribution. | Terminated | March 19, 2013 | 679 | 860 | Not_Computer
Z-2166-2011 | 58400 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DigiView software release 3.4.4.4 through 3.5.5.2. , 2008 Digisonics, Inc, 800-940-3240.Digiview is a digital image management and reporting system with an associated clinical database. It provides capture, review and storage of images together with report creation and data basing. | DigiView software | 2 | May 10, 2011 | 2011 | Digisonics, Inc | Failure of software to update some measurements when tracings or points are amended after measurement calculations have been performed. | DESIGN: Device Design | Digisonics, Inc., notified consignees by telephone and e-mail on 10/17/2008, informing them that a problem had been found in the Save Linked Images functionality of DigiView Release 3.6 (and all releases which start with 3.6 such as 3.6.1.). The problem is the failure of the software to update some measurements when tracings or points are amended after measurement calculations have been performed. A utility program, called Linked Image Measurement Verify was being supplied to all sites as a patch to the DigiView program. Release 3.6.6 was to be provided to all customers having previous 3.6 version. Problem was fixed in version 3.6.6 and 3.7.1.If you have any questions, please contact Digisonics Customer Support at (800) 940-3240. |  | 70 | 70 units total | Worldwide Distribution -- USA including States of : TX, OH, NC, MA, NY, IL, FL, OK, MO, ID, AZ, AR, MI CA, MD, and country of Canada.*** Due to the age of this recall, district is recommending no audit checks*** | Terminated | May 31, 2011 | 21 | 70 | Software | Output/Calculation | Software update | Software Update
Z-2171-2011 | 58216 | drills, burrs, trephines & accessories (manual) | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Ventricular Bolt Monitoring Kit W/Cranial Access, Drain, 1104HMC, Integra LifeSciences, 3498 West 2400 South, West Valley City, Utah 84119.Kit contains components which are used to monitor Intra Cranial Pressure and provide access to the cerebral ventricles for Cerebro-spinal fluid sampling, drainage and fluid injection. | Micro Ventricular Bolt Pressuring Monitoring, Cranial Access, and Drainage Kit | 2 | May 10, 2011 | 2011 | Integra LifeSciences Corp. | Outer box of Monitoring Kit was mislabeled with the incorrect part number. | PRODUCTION CONTROLS: Error in Labeling | Integra LifeSciences Corp. sent an Urgent Product Recall Notification letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine any affected product and arrange for return to Integra by contacting Integra Customer Service Returns at 1-800-654-2873 (choose option 2) to obtain a Return Merchandise Authorization number and place an order for replacement product. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified. For any questions regarding this recall call 609-936-2485. |  | 45 | 45 kits | Nationwide Distribution including CA, CT, GA, HI, IL, KY, MA, NC, NJ, NM. NV, OH, TX, WA, and WI. | Terminated | May 10, 2011 | 0 | 45 | Not_Computer
Z-2174-2011 | 58180 | prosthesis, knee, patello/femorotibial, semi-constrained, uncemented, porous, coated, polymer/metal/polymer | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NEXGEN COMPLETE KNEE SOLUTION FEMORAL REPLACEMENT JAW FOR USE WITH FEMORAL AND PROVISIONAL IMPACTOR/EXTRACTOR , NONSTERILE, REF 00-5901-026-40, Zimmer UK, Ltd., Zimmer Warsaw, IN.Intended use: The instrument is attached to the femoral implant or provisional to aid in positioning and inserting the components. | NEXGEN COMPLETE KNEE SOLUTION REPLACEMENT JAW FOR USE WITH FEMORAL AND PROVISIONAL IMPACTOR/EXTRACTO | 2 | May 10, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potiential for the spring clip to break off during use and fall into the surgical site. The firm has received 24 complaints of the spring breaking, four of which were reported that the patient had to undergo additional x-rays to ensure the fractured clip had not been left in-vivo. | DESIGN: Device Design | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" notice dated March 14, 2011 to its Consignees/customers. The letter described the product, problem, and the actions to be taken. The customers were instructed to stop using the device and quarantine immediately; carry out a physical count of all affected product in their territory and record this data on the Inventory Return Certification Form and fax to (574) 372-4265; ensure they have provided acknowledgement via "email return receipt" upon receiving notification; return the recalled product along with the completed Inventory Return Certification Form and Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580-Note: The instrument must be returned even if the spring clip is not missing from the instrument; and if they distributed these affected lots further, provide a copy of this letter to the those customers when they implement these recall instructions. The letter stated that a new or refurbished instrument with a redesigned spring clip will be provided as replacement for the recalled device. For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 27 | 27 devices | Worldwide distribution: USA (nationwide)including states of: HI, CA, NY, MI, IN, GA, FL, WI, MN, NJ, IL, VA, TX, NC, PA, MS, WA, OR, OH, KS, TN, NH, AZ, NV, UT and AK and countries of : Australia, Canada, China, Germany, India, Japan, Singapore and Sweden. | Terminated | August 28, 2013 | 841 | 829 | Not_Computer
Z-2184-2011 | 58583 | radioimmunoassay, human growth hormone | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 ELISA Reagent Kit; Part Number" DSL-10-6600; Lot Numbers: 992743, 992287, 991916, 991530, 991432, 990220 & 890830This assay is intended for in vitro diagnostic use. | IGFBP3 ELISA Reagent Kit | 2 | May 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was was initiated because patient results with the IGFBP-3 ELISA kit may be elevated by 70% compared to stated expected ranges in the Instructions for Use (IFU). | PRODUCTION CONTROLS: Process Control | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached customer response form on March 30, 2011 to all customers who have the affected product, IGFBP- 3 Enzyme-Linked immunosorbent Assay ( ELISA) kit .This kit provides materials for the quantitative measurement of IGFBP-3 in serum. The letter informed the customers of the reason for recall, the products affected and actions to be taken.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they were asked to provide a copy of the letter to them.Beckman further asked consignees to complete and return the enclosed Response Form within 10 days so they are assured consignees have received this important communication.Questions were directed to the Beckman Coulter Customer Support Center.- Website: http://www.beckmancoulter.com/customersupport/support- Phone, call 1-800-854-3633 in the United States and Canada- Outside the United States and Canada contact local Beckman Coulter representative. |  | 1212 | 1,212 units | Worldwide -- USA, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Italy, Korea, Republic of, Lebanon, Luxembourg, Morocco, Netherlands, New Zealand, Puerto Rico, Syrian Arab Republic, Taiwan, Thailand, Turkey & United Kingdom | Terminated | December 02, 2011 | 205 | 1212 | Not_Computer
Z-2183-2011 | 58479 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | TEMPUS IC PROFESSIONAL Patient Monitor, Part No. 00-1002, Manufactured by: Remote Diagnostics Technologies Ltd., The Old Coach House, Farleigh Wallop, Hants, UK RG25 2HT.Indicated for 3 Lead ECG monitoring and 12 Lead ECG recording, non-invasive blood pressure (NIBP), respiration, end-tidal C02 (ETCO 2), pulse oximetry (SP02) and tympanic temperature. | TEMPUS IC PROFESSIONAL | 2 | May 11, 2011 | 2011 | Remote Diagnostic Technologies Ltd. | It was discovered that when a 12 Lead ECG is recorded and monitoring is quickly restarted, the ECG can cease functioning. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Remote Diagnostic Technologies, Ltd notified its consignees of the problem and the work-around via letter with the subject line "ECG Issue - Field Action Required" on 03/25/2011. The letter stated that customers can avoid the issues by viewing the recorded ECG for at least 13 seconds. Two diagrams were attached with the letter depicting the conditions under which the issue can occur. Customers can continue to use the device normally in all other respects. The letter states that the firm is working on a software modification which will permanently fix the issue and prevent it from occurring. Customers are to contact the Regulatory Affairs and Operations Manager if they have any questions. | Contact the recalling firm for information | 19 | 19 units | Nationwide Distribution -- KY, MD, VA, and FL. | Terminated | September 16, 2011 | 128 | 19 | Software | Device Operation | Software update | Software Update
Z-2185-2011 | 57194 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe, Catalog 500AVHCT, Terumo Cardiovascular Systems Corporations, Ann Arbor, MI. The CDI 500 is indicated for use to display patient blood gas values during extracorporeal procedures. | CDI 500 Blood Parameter Monitoring System | 2 | May 12, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received 234 complaints of inaccuracy of displayed values on the CDI 101 and 500 monitors. In some cases the customer reported inaccurate potassium values. Other inaccuracies include CO2, hematocrit, O2 saturation, pH, hemoglobin and other miscellaneous inaccuracies. The CDI 101 and CDI 500 system are stand alone systems that are used during extracorporeal procedures to continuou |  | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual.The letter also states "Please note that we are not recommending that you stop using your CDI monitor system."If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST. | Contact the recalling firm for information | 3260 | 3260 units all models. | Worldwide distribution: USA (nationwide) and countries of : Australia, Barbados, Belgium, Brazil, Columbia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Guatamala, Hong Kong, Honduras, Jordan, Japan, Libya, Mexico, Malaysia, Panama, PH, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad, Taiwan and Uruguay, Germany, Saudia Arabia, Paraguay, South Korea, Dominican Republic, Phillipines, South Africa and UAE. | Open |  |  | 3260 | Other | Display/Image | Instructions | Safety Notice/Insructions
Z-2190-2011 | 58517 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE / RT Therapist v4.1 system.The ONCOR and PRIMUS system may be updated to utilize the RT Therapist v4.1 software.Siemens Healthcare, Concord, CA.The intended use is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ARTISTE / RT Therapist v4.1 system | 2 | May 12, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When using the ARTISTE Linac and RT Therapist v4.1 along with an imported cropped reference image, the image maybe processed incorrectly, thus resulting in a mistreatment. |  | Distribution of the Urgent: Medical Device Correction Customer Advisory Notice began on March 24, 2011 via certified mail or by the Siemens Customer Service Group. The letter identified the issue and when it occurs. Users were asked to verify the isocenter position of the reference image in the Portal Review application before using a reference image imported from the treatment planning system. Siemens is currently working on a solution and will notify customers as soon as the update is available. The letter should be included with the Owner Manual of the system in the chapter for "Safety Advisory Letters" until an update is provided. | Contact the recalling firm for information | 228 | 228 active devices | Worldwide Distribution | Terminated | February 13, 2012 | 277 | 228 | Software | Display/Image | Software update | Software Update
Z-2191-2011 | 58219 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Forteo [teriparatide (rDNA origin) injection] black starter kits containing Triad alcohol pads. The starter kits look like insulated lunch boxes and contain alcohols pads, needles, instruction guides, and freezer gel packs. Note: the kits do not contain the drug, Forteo. Intended use: Travel starter kits to hold the drug Forteo and supplies needed to deliver the subcutaneous injection. | Forteo teriparatide (rDNA origin) injection starter kits | 2 | May 12, 2011 | 2011 | Eli Lilly and Company | Eli Lilly and Company announced 3/17/2011 that patients should not use the alcohol prep pads made by the Triad Group that are contained in the black starter kits for Forteo [teriparatide (rDNA origin) injection] in the United States.The Tri-ad Group is recalling the alcohol prep pads due to potential contamination with the bacteria, Ba-cillus cereus, which could result in life threatening infecti | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Eli Lilly and Company, sent two letters entitled "URGENT MEDICAL DEVICE RECALL TRIAD GROUP ALCOHOL PREP PADS INCLUDED IN FORTEO STARTER KITS ALL ALCOHOL PREP PAD LOTS" to its customers. The first letter was dated March 18, 2011 and second follow-up letter was dated March 28, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to examine their inventory of Forteo starter kits; if they have any Forteo starter kits with a black bag, remove and appropriately discard the Triad Group alcohol prep pads; discontinue providing the product with the starter kits; inform any patients that may have received Forteo starter kits containing the Triad Group alcohol prep pads to discontinue use; note: after removing the Triad Group alcohol prep pads, all other components of the Forteo starter kits are acceptable for use: travel bag, pen case, gel packs, needles and instructional materials, and complete and return the enclosed business reply card. "It is important to note that this recall does not affect or involve the Forteo delivery device."For questions pertaining to the recall, contact Triad Group at 262-538-2900. (Please refer to the Triad Group press release issued on January 5, 2011 on the FDA website.) If you have product specific medical questions related to Forteo please call 1-800-LILLYRX. |  | 47000 | approximately 47,000 kits | Worldwide distribution: USA (nationwide) and countries of: Czech Republic, Germany, Slovak Republic, and Switzerland. | Terminated | August 30, 2011 | 110 | 47000 | Not_Computer
Z-2192-2011 | 58451 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Nokor Admix Needle 16 G x 1 TW (1.6mm x 25mm) REF 305216 SterileBD, Franklin Lakes, NJ MADE IN USAFor use in aspiration of medications, typically in the pharmacy setting. | BD 16 G x 1 in. Nokor" noncoring vented needle | 2 | May 13, 2011 | 2011 | Becton Dickinson & Company | Incorrect unit package labels which reflect Reorder #305216 instead of the correct item, Reorder #305213, were used. | PRODUCTION CONTROLS: Packaging | BD sent an Urgent: Product Recall letter with return response forms to direct consignees on 03/01/11 and 03/02/11 via UPS. The letter identified the affected product and lot along with the reason for the recall. The letter asked customers to complete the enclosed Distributor Packing List and to follow the instructions provided on the Packing List for returning the affected product. The firm will issue a credit upon receipt of the affected product. Customers can contact BD Customer Service if they need assistance or have questions regarding the recall. |  | 161000 | 161,000 units | Worldwide Distribution -- USA, including the states of NJ, IL, LA, NM, OH, NV, and VA, and country of Canada. | Terminated | November 30, 2012 | 567 | 161000 | Not_Computer
Z-2193-2011 | 58421 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, Digiview and ERS software releases 3.6.44, 3.6.4.6, and 3.6.5, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098, Tel: (800) 940-3240, Fax: (713) 529-7999Software used to make on-screen (2-D, M-Mode and Doppler) measurements of left ventricle, right ventricle, etc. | Digisonics, Digiview and ERS software releases 3.6.44, 3.6.4.6, and 3.6.5 | 2 | May 13, 2011 | 2011 | Digisonics, Inc | Left ventricle volume was inaccurate when the volumes were measured from images having different scales. | OTHER/UNDETERMINED: Pending | The firm, Digisonics, notified consignees/customers via e-mail entitled "Distribution of Notification of Biplane Error" dated February 6, 2009. The email described the product, problem and actions to be taken. The customers were instructed to immediately contact Digisonics if their site calculates biplane volumes-a utility will be provided which will run on their database and identify all studies in which biplane volumes are calculated. The customers were also instructed to share the following important information with all users of DigiView at their site. The email stated -until the software patch has been installed, biplane volumes should be calculated only from pairs of images having the same depth scale. Replacement DigiView/ERS software that prevents further occurrence of this problem will be provided if it appears their site could be impacted.The Distribution of New Software Release was sent on March 6, 2009.If you have any questions, please call Digisonics Customer Support at 800-344-4266. |  | 74 | 74 units | Worldwide distribution: USA (Nationwide) and countries of: Canada, Lebanon and New Zealand. | Terminated | May 31, 2011 | 18 | 74 | Software | Output/Calculation | Software update | Software Update
Z-2198-2011 | 58198 | tape and bandage, adhesive | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General & Plastic Surgery | 510(K) Exempt | Curad Flex-Fabric/Tissu Souple, 3/4 x 3 in., 100 sterile adhesive bandages, Made in China, Manufactured for Medline Industries, Mundelein IL 60060 | Curad Flexible Fabric Bandage | 2 | May 13, 2011 | 2011 | Jiangsu Wujin Hualian Health Dressing Co., Ltd. | The sterility of the product has the potential to be compromised due to weak sealing of the bandages. | PRODUCTION CONTROLS: Process Control | Changzhou Hualian Health Dressing Co., Ltd notified sole consignee by e-mail and a Medical Device Recall letter, dated 3/18/2011. The letter identifies the affected product and the reason for the recall. The customer was to examine their inventory and quarantine the affected product. If the product had been further distributed, those customers were to be identified and notified of the recall. The enclosed response form was to be completed and returned to Medline Industries, Inc. Questions could be directed to Customer Service, Monday through Friday between 8:30 A.M. and 4:00 P.M. | Contact the recalling firm for information | 1092 | 1,092 cases | Nationwide Distribution -- IL. | Terminated | April 23, 2012 | 346 | 1092 | Not_Computer
Z-2197-2011 | 58167 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray" Anterior Cervical Plate (ACP) System, Intended use: The Manta Ray ACP System is an anterior cervical plate that is indicated for temporary stabilization of the cervical spine from C2-C7. | Manta Ray" Anterior Cervical Plate (ACP) System | 2 | May 13, 2011 | 2011 | Integra LifeSciences Corp. | Based upon the results of the firm's literature review, they determined that the Package Insert (Information for Use) and information in the Surgical Technique should be revised to include additional warnings, precautions and possible adverse events. |  | The firm, Integra Spine, sent an "URGENT: FIELD CORRECTIVE ACTION" letter dated January 10, 2011 with a Field Corrective Action Acknowledgement and Return Form sent via Fed Ex/ priority mail to all customers including physicians and distributors. The letter describes the product, problem and actions to be taken. The customers were instructed to review the current inventory to identify product from the affected part numbers; replace the existing IFU (22-PI-01 rev C) with the revised IFU (22-PI-01 rev D) and the existing Surgical Technique (ST-22-001 rev A) with the revised Surgical Technique (ST-22-001 rev B), and complete and return the Field Corrective Action Acknowledgement and Return Form via fax to: 877-558-6227 or email.If you have any additional questions, contact the Manager Quality Systems Integra Spine at 330-475-8646. | No consumer action necessary | 244 | 244 IFU | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, GA, ID. IL, KS, LA, MD, MO, NH, NM, NV, NY, OH, OK, OR, PA, PR, SD, TN, TX, UT, VA, WI, and WV. | Terminated | October 27, 2011 | 167 | 244 | Not_Computer
Z-2195-2011 | 58438 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Medical Linear Accelerator, All Varian Clinac with Model Numbers: H14, H18, H26, H27, H28, H29, Hcx; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Trilogy Radiotherapy Delivery System is a radiation therapy accelerator intended to deliver megavoltage x-ray treatments for conventional radiotherapy (three dimensional conformal radiotherapy and intensity modulated radiotherapy) and stereotactic radiosurgery and radiotherapy. Stereotactic treatments are intended for therapy of lesions, e.g., arteriovenous malformations, primary tumors and metastases. | VARIAN HIGH ENERGY LINEAR ACCELERATOR | 2 | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly whereby, following prolonged use, the screw fastener holding the wedge body to the tray may fail. | PRODUCTION CONTROLS: Process Control | The firm, Varian Medical, sent an "Urgent Medical Device Correction" Letter dated March 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect all hard wedges for missing or sheared screw fasteners, if any wedge is found to have one or more missing or sheared screw fasteners, remove the wedge from use and notify Varian immediately; inspect all hard wedges for loose screw fasteners, and retighten as necessary, and implement a program of monthly inspection of wedges for loose or missing hardware. The customers were also instructed to advise the appropriate personnel working in their Radiotherapy Department of the content of this letter.Varian Service shall immediately begin an interim action to replace screw fasteners on all hard wedges for Clinacs installed prior to January 2004. Varian Service will contact the customers to arrange for the inspection visit.Varian also stated that they shall continue their Investigation of this issue to further clarify the exact conditions under which the wedge screw fasteners may fail, and to analyze possible final solutions. Varian will notify the customers when final determinations are made.If you require further clarification, please feel free to contact your local Varian Customer Support Services Manager and/or USA and Canada at 1-888-4265; Europe: +41 41 749 8844; or email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.Helpdesk@varian.com; Japan: Japan.apps helpdesk@varian.com; Latin America: soporte.al@varian.com, a d Internet: Oncology Systems-www.myvarian.com or Varian Medical Systems - www.varian.com. |  | 5616 | 5616 units | Worldwide distribution: | Terminated | June 07, 2013 | 756 | 5616 | Not_Computer
Z-2194-2011 | 58490 | drills, burrs, trephines & accessories (manual) | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra NeuroSciences, Combo Kit, Lot 171237, Integra LifeSciences 3498 West 2400 South, West Valley City, UT 84119.A kit for cranial access containing multiple necessary components to create the burr hole which is used to gain access to the cranial vault for the purpose of providing a portal for catheter placement. | Integra NeuroSciences Combo Kit | 3 | May 13, 2011 | 2011 | Integra LifeSciences Corp. | Outer packaging of Combo Kit was mislabeled with an incorrect lot number. | PRODUCTION CONTROLS: Labeling Mix-Ups | Integra LifeSciences Corp. sent a " Urgent Product Recall Notification" letter dated February 22, 2011, to the only affected consignee. It was reported the Combo Kits of SPOO98 / INS- 4500 ( Hermetic Large Style Ventricular Catheter set and drill set for NY) - Catalog # 31046 Lot Number 171237, Outer packages were labeled incorrectly. The customer was advised to return this, and they would be replaced with correctly labeled product. The issue is easily recognizable, and does not pose a safety risk to patients.An acknowledgement form was included to be returned to Integra. Further information is available at 609-936-2485. |  | 7 | 7 Kits | Nationwide Distribution -- including the state of New York. | Terminated | May 13, 2011 | 0 | 7 | Not_Computer
Z-2196-2011 | 57951 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Aortic Perfusion Cannula with Duraflow Coating, REF: DARH201190TA, 20 Fr. x 28 cm, Sterile EO, Edwards Lifesciences LLC, Irvine, CA 92614.The cannula perfuses blood to the ascending aorta during short-term ( | Aortic Cannula | 2 | May 13, 2011 | 2011 | Edwards Lifesciences Llc | Bonding of the tip to the aortic cannula tubing may decline over time and may contribute to a drop in bond strength. Possible tip separation. |  | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated January 21, 2011 it customers via Fed-Ex. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory for any of the lot numbers; return all unused product from their stock to Edwards; contact their Customer Service Organization at (800) 424-3278 to obtain a Returned Goods Authorization number and replacement product; complete and return the Aortic Perfusion Cannula Recall Response form via fax to: Edwards at (949) 250-3489, and return unused product to Edwards Lifesciences, 12050 Lone Peak Drive, Draper, UT 84020. If you have questions that have not been answered by this letter, call Edwards Customer Service at (800) 424-3278 from 8:00AM - 4:30PM Pacific Time. |  | 20 | 20 | Nationwide distribution: CA only. | Terminated | May 13, 2011 | 0 | 20 | Not_Computer
Z-2064-2011 | 58160 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Defibtech Lifeline AED and ReviveR AED, Model: DDU-100 Series AEDsSoftware version 2.004 or earlierThe product under correction is the DDU-100 AED. The AED is indicated for use on victims of sudden cardiac arrest (SCA) when the patient is unconscious, unresponsive and not breathing. The indications for use also state that for patients under 8 years old, child/ infant electrode pads are to be used and that therapy should not be delayed to determine the patients exact age or weight. The AED must also be used by or on the order of a physician. These AEDs are sold under the brand names Lifeline AED and ReviveR AED. | Defibtech Lifeline AED and ReviveR AED | 1 | May 16, 2011 | 2011 | Defibtech, LLC | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process | DESIGN: Software Design (Manufacturing Process) | The firm, Defibtech, sent an "URGENT MEDICAL DEVICE SAFETY INFORMATION AND CORRECTION" letter dated March 11, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove the pads from the storage pouch on the back of their DDU-100 series AED and inspect the 9-digit serial number label.2)Carefully note the 9-digit serial number on the back of the AED, then visit the Defibtech website at www.defibtech.com/fa11 and enter their serial number to determine if their AED is affected; email their serial number and contact information to fa@defibtech.com or call Defibtech directly at 1-877-453-4507 or 1-203-453-4501 and Defibtech will inform them if their unit is affected.3) If their AED is affected, confirm receipt of this notification with Defibtech as soon as possible by website and/or confirmation card. Complete and return the confirmation form via fax to 1-203-738-1072. Note: "it is very important for you to contact us either through the Defibtech website or return this card to confirm that you have received this notice. This information will expedite the process of upgrading your AEDs."If your DDU-100 series AED serial number DOES NOT fall within the listed range, your AED is not affected by this recall.A letter was also sent to the distributors and direct accounts that have received the affected product by first class mail with a prepaid, return written acknowledgement card. The field action consists of a software change to the AED. All affected AEDs operating with software version 2.004 or earlier are to remain in the field and in service throughout the field action process.Upon receipt of contact information from affected customers, Defibtech will send an upgrade kit to customers consisting of a software data card, instructions on how to perform the upgrade, and a pre-paid upgrade confirmation postcard that must be returned to Defibtech upon completion. |  | 563 | 91, 472 units | Worldwide distribution: USA (nationwide) and countries including: AUSTRALIA, AUSTRIA , BAHRAIN, BRAZIL, CANADA, CHILE, CHINA , CYPRUS, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA, REPUBLIC OF LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND URUGUAY. | Open |  |  | 91472 | Software | Device Operation | Software update | Software Update
Z-2199-2011 | 58596 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | The Operating Manual contains the Operating Instructions for AS 104 Blood Cell Separator and has a Software Version 4.71. The AS 104 Blood Cell Separator Device has a Catalog Number 9007031.Used for depletion or exchange of red blood cells during therapeutic apheresis procedures on the Fresenius AS 104 Blood Cell Separator Device. This set is used with a single stage separation chamber. | Fresenius AS 104 Blood Cell Separator 2/03.11 (OPUSA) | 2 | May 16, 2011 | 2011 | Fresenius Kabi, LLC | Operating Manual of the AS 104 Blood Cell Separator is being updated to eliminate any reference to Red Blood Cell (RBC), C4Y, P1Y and BMSC sets. This is a corrective action and recall expansion of Z-0616-2011. | MISBRANDING: Labeling False and Misleading | Fresenius Kabi sent a "Urgent Notification" letter dated May 6, 2011, to the affected consignees they are recalling the Operating Manual for the Fresinius AS104-Blood Cell Separator.The Red Blood Cell (RBC) sets have been recalled according to recall number Z-0616-2011, and will permanently not be used with the AS 104 Blood Cell Separator Device. The RBC set was never approved by the FDA according to Warning Letter (Reference # GEN1001508) dated 31 January 2011. The customers are being instructed to complete the self addressed postage pre-paid " Product Response Card." Place a check mark next to the disposition choice used.Fresenius plans to send the revised Operating Manual and a notification letters to consignees at the end of April 2011.Consignees who have questions about this recall can call the firm at 1-425-242-2074. | Contact the recalling firm for information | 158 | 158 units | Nationwided Distribution -- AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MI, MN, NC, NH, NJ, NM, NY, OH, PA, SD, TN, TX, VT, and WI. | Terminated | December 06, 2011 | 204 | 158 | Not_Computer
Z-2200-2011 | 58505 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, Version 2.6 and 2.6.1, Model Number: HZl, Manufactured and Distributed by: Varian Medical Systems Inc., Palo Alto, CA.For use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Optical Guidance Platform, Version 2.6 and 2.6.1 | 2 | May 16, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Optical Guidance Platform (OGP) Software v2.6 and v2.6.1 where the transfer of datasets from treatment planning systems other than the FastPlan system result in a lateral offset error. | DESIGN: Software Design | On 3/30/2011, Varian distributed an Urgent Medical Device Correction Urgent Field Safety Notice to all affected users, with a description of the problem and identification of the affected product. The letter also informed customers of how they can determine if the problem affects them. The letter states that there is no work around for the issue unless they possess a FastPlan system. Varian is preparing a software patch to correct the issue and it will be applied to all OGP v2.6 and v2.6.1 systems. Customers are to advise the appropriate personnel of the content of the letter. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 598 | 598 units | Worldwide Distribution | Terminated | June 19, 2012 | 400 | 598 | Software | Output/Calculation | Software update | Software Update
Z-2212-2011 | 55207 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate TM Sample Processing System, P/N A20745 Basic Assembly - AutomateThe AutoMate system prepares sample tubes for diagnostic testing and for storage. The AutoMate system contains an automated track that connects to Input/Output, Centrifuge, Serum Level Detection, Decapper, and Output modules. The AutoMate TM 800 model also contains and Aliquot module. | AutoMate TM Sample Processing System | 2 | May 19, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the AutoMate Sample Processing Systems. If during the recovery from Error 81201001, a stuck sample carrier with a tube is moved backwards to the loading position, sample tube/input tray information will become mismatched for samples being loaded onto the AutoMate. | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated May 7, 2009, with a Fax Back Response Form (sent by US mail) to all customers who received the AutoMate Sample Processing System. The letter identified the product, the problem and a workaround to correct the problem. Customers were instructed to take the following steps:If the stuck sample carrier is moved backwards during recovery from this error, the sample tube/input tray information will become mismatched.Figure 2 in the PCA letter shows the AutoMate viewed from the rear. Sequence 1 to 3 depicts the error recovery issue. The sample is loaded, but cannot move along the track. As the sample attempts to move along the track, it is captured by the track stopper (not seen) below the track. Error 81201001 will be displayed at the console with the following description: The stopper (SL01) did not stop the sample tube carrier. If the stopper or an operator forces the sample backward along the line, it will cause the information mismatch.If you encounter this scenario, do not move the carrier backward to the loading position. See Figure 3 in the PCA letter.Recovery from Error 81201001:1. The normal recovery for error 81201001 is to move the sample carrier forward. If there is no room to do so in the Input module, extra steps must be taken for the carrier to move forward.a. Resume the module by performing error recovery on the Hardware Error Recovery screen at the console. b. Continue the console error recovery process until a carrier is released from BR01, thereby creating space for the stuck carrier to be pushed forward.c. Move the stuck sample carrier forwardd. Once again, resume the module by performing error recovery on the Hardware Error Recovery screen. (The error should not return again)Recovery if Symptoms are Observed: 1. If the incorrect processing described above is observed, the system must be stopped immediately and shut down. All samples m. |  | 5 | 5 units total | Nationwide Distribution including VA KY and CA | Terminated | June 28, 2012 | 406 | 5 | Not_Computer
Z-2204-2011 | 55542 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64; System, X-Ray, Tomography Computed | CHVH001A/1B Variable Helical Pitch (vHP) Software for TSX101A: Aquilion 64 | 2 | May 19, 2011 | 2011 | Toshiba American Medical Systems Inc | The recall was initiated by Toshiba America Medical Systems (TAMS) because the software associated with TSX-101A: Aquilion 64 can prevent scanning from being performed at the specified helical pitch (couch movement speed). | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) forwarded an Urgent: Medical Device Correction letter, dated 4/6/10, with attached Customer Reply Form, via US Postal Service to all customers who purchased the CHVH-001A/1B variable helical pitch (vHP) software for TSX-I0IA: Aquilion 64. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that a Toshiba service representative would contact them to schedule an appointment in order to install a corrective measure that will prevent the occurrence of the problem. Toshiba would like customers to set a scan range different from that for the previous scan for the second and subsequent vHP scans, if multiple vHP scans with the same scan conditions are performed in Auto Mode.Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. Customers with any questions can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Representative at (800) 521-1968. |  | 35 | 35 | Nationwide Distribution -- MN, OH, MO, MI, NY, KS, MD, TX, NJ, IL, ID, TN, AZ, PA, HI, GA, MT, WA, FL, NC, WI, ME, & CA. | Terminated | August 17, 2012 | 456 | 35 | Software | Device Operation | Software update | Software Update
Z-2203-2011 | 55543 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Software for SSH-880CV Aplio Artida; System Diagnostic Ultrasound | SSH880CV Aplio Artida | 2 | May 19, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has become aware that the FLEX M software has reversed left and right annotation and post FLEX M measurement error, which may result in incorrect values. | DESIGN: Software Design | The recall communication was initiated on 1/6/10 with Toshiba forwarding an Urgent: Medical Device Correction letter with attached Customer Notification Reply Form (via US Postal Service for US Customers) to all customers who purchased the FLEX-M mode on Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid performing M image measurement using images saved in FLEX-M mode until the corrective action is completed. If customers need to perform measurement for images saved in FLEX-M mode, they were instructed to use manual calibration. Customers were instructed to note that even though the corrective action is taken, the existing saved images (saved images before the corrective action is taken) are not modified. Therefore, when referring to the measured values on the images which are saved in FLEX-M mode before the corrective action is taken, customers were instructed to carefully handle the measurements results because they may be incorrect.Customers were instructed to complete the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com.If customers have any questions regarding the medical device correction letter, they can contact the Director of Regulatory Affairs at (800) 421-1968 or their local Toshiba representative. | Contact the recalling firm for information | 28 | 28 | Nationwide Distribution -- Including states of IA, MT, PA, CA, OR, GA, LA, OH, TX, MA, OH, WI, and Puerto Rico. | Terminated | April 20, 2012 | 337 | 28 | Software | Output/Calculation | Software update | Software Update
Z-2201-2011 | 56389 | enzyme linked immunoabsorbent assay, rubella | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Software CD is labeled in part:"BLOOD VIRUS DIVISION***BIO-RAD EVOLIS System APF Version 4.4 Software Version 2.00 2009/05***"The EVOLIS Assay Protocol Files (APF) Software contains assay-specific instructions necessary for the EVOLIS Microplate processing system to process enzyme immunoassays on the EVOLIS instrument, perform data reduction and generate reports. | EVOLIS Assay Protocol Files ( APF) Software | 2 | May 19, 2011 | 2011 | Bio-Rad Laboratories | The EVOLIS Assay Protocol Files Software is configured in such a way that may result in a miscalculation of quantitative test results and/or an error in reporting positive/negative sample values. | DESIGN: Software Design | Bio-Rad Laboratories sent an Urgent Product Correction letter dated July 19, 2010, to all customer sites via FedEx. There are two issues related to The EVOLIS Assay Protocol Files (APF) Software that could potentially affect patient results. Issue 1 involves quantitative sample test/calculations/results interpretation for assays with linear regression, point to point or 4PL data reduction. Issue 2 involves sample test results with " Result Report Flags''. A software patch to correct these two issues will be sent to each customer site. Bio-Rad Laboratories will work with each customer to complete documentation confirming that the software patch has been successfully installed on each instrument. If you have any questions or concerns regarding these two issues, please contact Technical Support at 1-800-224-6723, option 2 and then option 3. | Contact the recalling firm for information | 272 | 272 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 25, 2011 | 98 | 272 | Software | Output/Calculation | Software update | Software Update
Z-2215-2011 | 57469 | system,test,thyroglobulin | MSW | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access Immunoassay Systems Thyroglobulin,Part Number: 33860This device is intended to aid in monitoring for the presence of local and metastatic thyroid tissue in patients who have had thyroid gland ablation (using thyroid surgery with or without radioactivity) and who lack serum thyroglobulin antibodies. | Access Immunoassay Systems Thyroglobulin | 2 | May 19, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because In addition to Access Tg Sample Diluent, an additional diluent, Access Wash Buffer II, was added to Thyroglobulin instructions for use (IFU) for evaluation of samples having Tg levels above the reportable range of the assay. When customers interchange between diluents there is an unacceptable shift in results. A mediandifference of 16% (ranging from 14% to 23%) b | OTHER/UNDETERMINED: Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 18, 2010, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems Thyroglobulin. The letter describes the product, problem and actions to be taken. The customers were instructed to:1. Do not dilute elevated thyroglobulin samples with Wash Buffer II.2. Do not run thyroglobulin patient samples using the onboard dilution feature on the UniCel DxI instrument.3. Dilute elevated thyroglobulin samples only with Access Thyroglobulin Sample Diluent (REF 33866).4. Review your patient result archives. Manually dilute and repeat the analysis of all elevated thyroglobulin samples that had previously been diluted with Wash Buffer II, using Access Thyroglobulin Sample Diluent instead.5. A forthcoming version of the UniCel DxI Assay Protocol File (APF) will remove the Access Thyroglobulin onboard dilution assay (dTg) from the list of available onboard dilution assays.6. The Access Tg product Instructions for Use will be updated to eliminate Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples. Customers should update procedure manuals and other documents that your laboratory may have created which list Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples.In addition, customers were instructed to complete and return the enclosed PCA Response Form with in 10 days via fax to: 786-639-4000 or mail to: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs. Customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman's Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact their local Be. |  | 6414 | 6,414 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | August 10, 2012 | 449 | 6414 | Not_Computer
Z-2214-2011 | 57591 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel Select, BrightSpeed Edge Select ,BrightSpeed Elite Select and BrightSpeed Elite Computed Tomography X-ray systems with software version 09BW35.11 or 09HW30.4.The GE BrightSpeed Delight Series Computed Tomography X-ray system is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes, for patients of all ages, including Axial, Cine, Helical, Cardiac, and Gated acquisitions. These images may be obtained either with or without contrast. | GE BRIGHTSPEED DELIGHT CT SCANNER SYSTEM | 2 | May 19, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential set of circumstances that could cause X-ray continuation during an unexpected table stop on certain BrightSpeed systems (BrightSpeed Excel/Edge/Elite Select and BrightSpeed Elite Computed Tomography X-ray Systems ) if this event were to occur on your system, patient safety may be impacted. | DESIGN: Software Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION GE Ref: 22932" letter dated January 19, 2011 to its consignees/customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Contact Information, and Additional Dosimetric Information.A second letter dated March 11, 2011 was sent to the consignee/customer letter consisting of a cover letter and an attachment. "This letter is to notify you of an error in the January 19, 2011, Urgent Medical Device Correction Letter, that you have already received (GE Ref: 22932). In the Safety Issue section of that letter it is written: "If this issue were to occur, it would result in the remainder of the exposure being delivered at the patient position where the table stopped up to the prescribed scan time or 30 seconds, whichever is shorter." The stated 30 seconds limitation is incorrect for your particular model. This information has been corrected on the attached Urgent Medical Device Correction letter, so please discard the original letter if it is still in your possession. "GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction.The customers were also instructed as follows: prior to GE's installation of a software update, continue to monitor patient scanning closely. Should the table stop unexpectedly and X-ray remains on, manually stop the scan, and ensure that all potential users in your facility are made aware of this safety notification and the recommended actions.If you have any questions, please contact GE immediately at the following: United States: 800-437-1171; Japan: 0120-055-919; Korea: 1544-6119; Australia/New Zealand: 800-659-465 and China: 800-810-8188. For other countries please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notifi. | Contact the recalling firm for information | 479 | 479 | Worldwide distribution: USA (nationwide) including states: AL, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PR, RI, SC, SD, TN, TX, and WV; and countries including: URUGUAY, TURKEY, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, PUERTO RICO, PORTUGAL, POLAND, PARAGUAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LEBANON, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, EGYPT, COLOMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BELARUS, BAHRAIN, AUSTRALIA, ANGOLA, ALGERIA. TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, PHILIPPINES, PARAGUAY, OMAN, NEW ZEALAND, MOROCCO, KUWAIT, KAZAKHSTAN, ISLAMIC REPUBLIC OF IRAN, ARGENTINA. | Terminated | July 10, 2012 | 418 | 479 | Software | Device Operation | Software update | Software Update
Z-2208-2011 | 58037 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek GRYPHON T BR DS Anchor w/ORTHOCORDProduct Code: 210811Gryphon BR Suture Anchors are intended for: Shoulder: Rotor Cuff Repair, Bankart Repair, SLAP Lesion Repair, Biceps Tenodesis, Acromio-Clavicular Separation Repair, Deltoid Repair, Capsular Shift or Capsulolabral Reconstruction; Foot/Ankle: Lateral Stabilization, Medial Stabilization, Achilles Tendon Repair; Knee: Medial Collateral Ligament Repair, Lateral Collateral Ligament Repair, Posterior Oblique Ligament Repair, lliotibial Band Tenodesis; Elbow: Biceps Tendon Reattachment, UInar or Radial Collateral Ligament Reconstruction. | Depuy Mitek GRYPHON T BR DS Anchor w/ORTHOCORD | 2 | May 19, 2011 | 2011 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Metal debris in the shaft component of the device | PRODUCTION CONTROLS: Manufacturing Material Removal | DePuy Mitek sent an Urgent Voluntary Product Recall Notification letter dated February 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check their inventory to determine if they have any affected product. Customers were asked to complete the enclosed business reply form and fax it to 508-828-3750 or email to mitekcomplaints@its.jnj.com whether or not they have any of the affected product remaining in their inventory. Customers were asked to utilize the return good authorization number provided by customer service and return all identified product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324For any questions call 1-877-379-4871 or contact DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. | Contact the recalling firm for information | 1126 | 1126 units | Worldwide Distribution - USA (nationwide) and the countries of Austrailia Austria Belgium Canada China Estonia Finland France Germany Great Britain Indonesia Ireland Italy Japan Lebanon Lithuania Netherlands Poland Singapore Slovenia South Africa Spain Sweden and Switzerland | Terminated | February 20, 2013 | 643 | 10139 | Not_Computer
Z-2219-2011 | 58429 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | PercuNav Field Generator Stand is a Computed tomography, X-Ray System. Part Number 453561431911.Designed to be used only with the PercuNav" image guided intervention system. The Field Generator Stand (with articulating Mounting Arm) is used to position the Field Generator near the area of interest. | PercuNav Field Generator Stand | 2 | May 19, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | The PercuNav Field Generator Stand could tip over during set up, during use, or during transport. | DESIGN: Device Design | On 4/4/11, Philips Healthcare began sending out the URGENT-Medical Device Correction notification letter to their consignees.Philips informed their customers of a potential tipping hazard of the PercuNav Field Generator Stand. The notification letter also provides instructions to customers to ensure the articulating Mounting Arm is positioned so the Field Generator is not further than the specified distance from the back of the clamp on the stand's central pole during set up or during use. Customers are instructed to apply a provided label on stand to warn of a potential for tipping due to excessive force. For transport, customers are instructed to remove the Field Generator from the Mounting Arm and place the Field Generator in the storage compartment on the system cart. The Mounting Arm should then be collapsed and folded into a 90 degree angle.Customers are informed that a Philips field service representative will contact them to schedule installation of a service kit on the stand, free of charge, once service kits are available.Customers can call Philips Customer Service at 1-866-767-7822 and reference to FCO 79500197 for any questions about this recall. | Contact the recalling firm for information | 65 | 65 units total (33 units distributed to US and 32 units to foreign consignees) | Worldwide Distribution -- USA, Australia, Austria, Belgium, China, France, Germany, Hong Kong, India, Korea, Netherlands, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | March 15, 2012 | 301 | 65 | Not_Computer
Z-2213-2011 | 58235 | products, red-cell lysing products | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER LYSE S III diff Lytic Part Number: 8546796.COULTER LYSE S III diff Lytic Reagent For use as an erythrocyte lytic agent for the quantitative determination of hemoglobin and for leukocyte counting and sizing on COULTER GEN S, MAXM, HmX, LH500, LH 750, and LH 780 hematology analyzers. Use with only ISOTON III or LH Series diluent. Refer to your instrument product manuals and/or online help, as applicable. | Coulter Lyse S III diff Lytic | 2 | May 19, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter because the COULTER LYSE III diff Lytic Reagent Lot number 101554F was released for distribution to customers between November 17 and December 1, 2010 prior to completion of QC testing. |  | Beckman Coulter sent a Product Corrective Action ( PCA) letter dated December 28, 2010, to all affected consignees with an attached PCA Response form via US Postal Service for US Customers) who purchased the Coulter Lyse S III diff Lytic Lot # 101554F. The letter provides customers with an explanation of the problem identified and an action to be taken. If customers currently had any of the lot mentioned above in their laboratory, they were asked to discontinue use and discard remaining product according to their laboratory chemical safety procedures. And to complete and return the enclosed response form within 10 days to acknowledge that they received this notice and are taking appropriate action.For the population at most risk, there is no serious consequences and only remote probability of injury and medically reversible consequences due by the small change for Hgb and WBC results.Consignees were also asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation.Technical questions regarding this Product Corrective Action were directed to Beckman Coulter Customer Service at( 800) -526-7694 in the United States and Canada. | Contact the recalling firm for information | 1296 | 1,296 units | Worldwide Distrubution-- USA ( nationwide) including states of : AL, AR, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, RI, SC, TN, TX, VA, WI, & WV and country of Canada. | Terminated | June 04, 2012 | 382 | 1296 | Not_Computer
Z-2217-2011 | 58667 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Audio-Video Engineering PAS-210 Remote Active Speakers Kit used with the IntelliVue Information CenterThe intended use of the Information Center Software is to display physiologic waves, parameters and trends, formal data for strip recordings and printed reports and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. The Information Center Software provides for the retrospective review of alarms, physiologic waves and parameters from its database. An additional intended use is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips AudioVideo Engineering PAS210 Remote Active Speakers Kit | 2 | May 19, 2011 | 2011 | Philips Healthcare Inc. | PAS-210 Kit Remote Active Speaker Assembly used with the IntelliVue Information CenterSpeaker cable connection may lead to intermittent or loss of audio alarm annunciation | DESIGN: Component Design/Selection | Philips Healthcare issued an Urgent - Medical Device Correction letter dated April 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the "Action to be taken by Customer/User" section of the instructions while they await the correction, which will likely come in the form of a replacement audio cable for all affected speaker assemblies.For further information or support concerning this issue, please contact Philips Healthcare Customer Care Service Center at 800-722-9377 # 3, # 1 and reference the letter or your local Philips representative in your area. |  | 15062 | 15062 units | Worldwide Distribution - USA (nationwide) and the country of Canada | Open |  |  | 15062 | I/O | Alarm/Message | Replace cable | Remove or Replace
Z-2069-2011 | 57473 | veterinary laser | RGB |  |  |  |  | Lockheed Martin Aculight Capella-r 1850 Laser. This device has the Accession number 0612339.The Capella R-1850 is a Class 4 solid state diode laser product designed to emit 1850mn laser radiation at maximum pulse energy of 5 mJ (@1ms). The output port connects to a user-selected 200-600 micron diameter optical fiber. The Capella R-1850 is labeled in part: "***LASER APERTURE***Aculight Corporation 22121 20th ave SE Bothell, WA 98021***Feb 2007***Model R-1850, SN: 007***". | Lockheed Martin Aculight Capella R1850 Laser | 2 | May 19, 2011 | 2011 | Lockheed-Martin Aculight | It has been determined that there are certain software driven operating sequences involving off-on cycling of the key-switch that do not result in a reset of the Stim Enable button, and emissions can resume after a standard 5 second countdown without pressing this button a second time. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | On 11/22/10, Lockheed Martin Aculight sent a Memo to their Customers and Users informing them of the "Stim Enable" button reset error. The defect involves the function of the "Stim Enable" button on the front panel of the Capella R-1850.In December 2010, the firm sent a second lletter with instructions on how to return the Capella R-1850 Laser back for repair.The firm's plan to correct the defect involved removal of existing software, installation of revised software, and documented verification testing, per written instructions, to confirm resolution of the Stim-Emit button failure to reset. Corrections were to be made at no cost to the customer. A Lockheed Martin Aculight customer service representative will contact you shortly and provide detailed information on how to return the Capella-R 1850 laser product in your possession for testing and repair on this defect.Should you have any questions, please call (425) 482-1100. |  | 17 | 17 units | Devices were distributed Universities, Institute, and government consignees in CT, IL, MA, MD, OH, TN, TX, and UT.One unit was distributed to Germany. | Terminated | August 23, 2011 | 96 | 17 | Software | Device Operation | Software update | Software Update
Z-2221-2011 | 58614 | implant, dermal, for aesthetic use | LMH |  |  | General & Plastic Surgery |  | Radiesse, Injectable Implant - 1.3cc Syringe, Sterile, Sterilized Using Steam, Use Before 2013-01, Ref 8046M7, Lot 1024417, Manufactured by: BioForm Medical Inc, 4133 Courtney Rd #10 Franksville, WI 53126.Subdermal implantation for the correction of moderate to severe facial wrinkles and folds, such as masolabial folds and it is also intended for restoration and/or correction of the signs of facial fat loss (lipoatrophy) in people with human immunodeficiency virus. | Radiesse Injectable Implant | 3 | May 19, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics, Inc. is recalling 1 lot (1034417) of Radiesse 1.3 cc Implant. The lot was supposed to be scrapped because it had failed to meet its Loss On Drying (LOD) specification. Unfortunately it had been marked as approved, released and shipped. | TRAINING: Employee Error | The firm, Merz Aesthetics, Inc., contacted its consignee/customer via e-mail on April 6, 2011. The e-mail states the product, problem and action to be taken. The customer was instructed to follow up with the concerned and inform the status time to time to the Recall Coordinator. The product was confirmed to be quarantined and has been requested to be sent back.If you need any additional information regarding this issue or if you have any questions, contact the International Regulatory Affairs Specialist at 262-835-3300 ext. 3081. | Contact the recalling firm for information | 695 | 695 units | International only: Taiwan. | Terminated | November 17, 2011 | 182 | 695 | Not_Computer
Z-2216-2011 | 58534 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS accelerators that utilize the RT Therapist v4.2 system in combination with syngo RT Oncologist 4.2.The ONCOR and PRIMUS may be updated to utilize the RT Therapist v4.2 software. The RT Oncologist v4.2 is currently only on Non-US systems. Distributed by Siemens Healthcare, Concord, CA 94520Intended use: Radiation therapy | ARTISTE , ONCOR and PRIMUS accelerators | 2 | May 19, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential safety issue that may result in mistreatment when customers are configured with the ARTISTE Linac RT Therapist v4.2 in combination with RT Oncologist v.4.2 and use the remote "Online 2D Image Review. | DESIGN: Software Design | The firm, Siemens, sent an "Urgent: Medical Device Correction Customer Safety Advisory Notice" on March 31, 2011 via Siemens Customer Service Group or certified mail to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to not use the "Online 2D Image Review" on the Syngo RT Oncologist 4.2; to always calculate the offset directly on the RTT 4.2 before it is applied for the treatment with the patient on the treatment table; include this notice in their syngo RT Therapist 4.2 System Owner Manual chapter "Safety Advisory Letter" where it should remain until the measures defined in the notice are performed, and in the interest of safety, perform the preventative measures noted in this notice and inform all affected personnel immediately.Siemens is currently working on a solution for the incremental loading workflow in syngo RT Therapist 4.2 and will provide a software update as soon as possible. Siemens will notify the customers as soon as an update for their system will be available. UPDATE: The release of Safety Update UI TH010/11/S to distribute the software update for RTTv4.2.94O04 began on April 15, 2011.If you have any questions, contact Regulatory Affairs at 925-602-8157. | Contact the recalling firm for information | 45 | 45 active devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, OH, PA and UT; and countries including: Australia, Belgium, Croatia, Germany, Japan, South Africa, Sweden and United Kingdom. | Terminated | March 30, 2012 | 316 | 45 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2220-2011 | 58581 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, Client software versions 8.0 through 8.6; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | May 19, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly in Eclipse Treatment Planning System, versions 8.0 to 8.6, whereby Actual SSD (source surface distance) value may not correspond to the Planned SSD value after Planning Approval. | DESIGN: Software Design | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 4, 2011, to its customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Set the Planning Approved status in Eclipse (Eclipse displays both the Actual and Planned SSD values during plan approval) and verify both SSD values prior to approving the plan.2)Verify the SSD values in the Eclipse Treatment Plan Report prior to approving the plan for treatment or proceeding with treatment delivery (this report lists both the Actual and Planned SSD values).3) Prior to approving the plan in RT Chart, verify both the Actual and Planned SSD values in the Parameters Workspace by selecting Field Parameters |  | 5288 | 5288 units | Worldwide distribution. | Terminated | June 27, 2012 | 405 | 5288 | Software | Output/Calculation | Software update | Software Update
Z-2206-2011 | 58510 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64, Computed Tomography X-Ray System, Model Numbers: 728231 and 728326, including the Ingenuity CT. The recalled units are equipped with software versions: 3.5.21028 and 3.5.25028. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Brilliance 64 CT XRay System | 2 | May 19, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | A program software error in the algorithm produces image artifacts that could possibly lead to a misdiagnosis of patients who have small metal implants. | OTHER/UNDETERMINED: Pending | A 'Field Safety Notice - Urgent - Medical Device Correction' notification was sent to the firm's customers via Certified mail on 3/8/2011. The notification letter explains the nature of the Metal Artifact Reduction/Metal Artifact Correction software error; the Brilliance CT X-Ray systems subject to correction and how to identify them; and the circumstances under which this error can occur. It also asks that the user stop running the MAR application on any of the affected Brilliance systems identified with the recalled serial numbers and software versions alluded to in the medical device correction notification. Lastly, the notification states that Philips Medical will provide a solution that will fix the issue and the notification lists contact names; telephone numbers; and E-mail addresses which the customer can contact for further information concerning this device issue, or customer support assistance. |  | 10 | 10 units | Nationwide Distribution -- DC, IL, IN, MI, MO, NY, OH, OR, PA, TX, and VT. | Terminated | February 22, 2013 | 645 | 23 | Software | Output/Calculation | Software update | Software Update
Z-2223-2011 | 58514 | kit, i.v. start | LRS |  |  | General & Plastic Surgery | Enforcement Discretion | V-CATH Insertion Kits, Sterile; Product Code Number: 301-00;Manufactured by Centurion Medical Products for Neo Medical;42514 Albrae Street, Fremont, CA 94538 | VCATH Insertion Kits | 2 | May 19, 2011 | 2011 | Neo Medical, Inc. | Firm received recall notice from their contract manufacturer due to visible particulates in Bacteriostatic Sodium Chloride Injection, USP 0.9%, 30 mL vials that are included in insertion kits. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Neo Medical forwarded Urgent Recall Notices to Dealers and customers on April 5, 2011 via fax. Neo Medical Customer Care also contacted customers by phone on April 7, 2011, followed by Urgent Product Recall Notification letters were sent on April 11, 2011 along with a copy of the Recall Notification from Centurion. The letter requested that customers with affected lots immediately discontinue distribution and contact the end users instructing them to stop dispensing the product. Customers were to complete the accountability information even if inventory had been depleted and fax it to 510-353-9025. In order to return the affected product, customers should contact the Complaint Coordinator at 925-719-2197 or pjduvall@neomedicalinc.com. Replacement or credit would be issued for all product returned. | Contact the recalling firm for information | 1990 | 1990 | Worldwide Distribution -- USA, including states of MO, WA, PA, CA, OH, UT, TX and Canada. | Terminated | November 22, 2011 | 187 | 1990 | Not_Computer
Z-2222-2011 | 58586 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague, Colleague CX and Colleague CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8151, 2M8161 and 2M9161. Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. Electronic infusion pumps indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration, such as intravenous (IV), intra-arterial (IA), subcutaneous, epidural or irrigation of fluid spaces applications. | Colleague, Colleague CX and Colleague CXE Infusion Pumps | 2 | May 19, 2011 | 2011 | Baxter Healthcare Corp. | The Hi-Pot safety testing may not have been conducted on the affected pumps during service, prior to being returned to the end user. | TRAINING: Employee Error | The firm, Baxter Healthcare Corporation, telephoned the customers on April 14, 2011, informing them that that the Hi-Pot test,and sent a follow up "URGENT DEVICE CORRECTION" letter dated April 27, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the pumps and take them out of service as soon as possible (Baxter will then send a shipping container for retrieval, re-inspection and return of the device back to the customer at no cost.); and return their colleague pumps for full functional testing including the HiPot test to the Baxter Buffalo Grove Global Technical Services (GTS) Center at Baxter GTS, 900 Corporate Grove Drive, Buffalo Grove, IL 60089.Baxter will perform full functional testing, including the HiPot test, on affected Colleague pumps for which the customer has continuing medical necessity. Should Baxter identify any issues through repair or service prior to returning the Colleague pumps to the customer.If you have questions regarding this communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CT, Monday through Friday. |  | 117 | 117 pumps | Worldwide distribution: USA (nationwide) including states of :Alabama, Arizona, California, Connecticut, Florida, Georgia, Illinois, Kentucky, Louisiana, Massachusetts, Maryland, Mississippi, North Carolina, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Texas, Virginia, Washington and West Virginia, and country of: Canada. | Terminated | May 09, 2012 | 356 | 117 | Not_Computer
Z-2218-2011 | 58549 | oximeter | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | AVOXimeter 1000, AVOXimeter 40001000- portable bedside whole blood oximeter that directly measures oxyhemoglobin fraction, the total hemoglobin concentration, and oxygen content. It has enhanced software that optimizes its use int he cardiac catheterization laboratory through calculations of physiologic equations. 4000-a portable bedside whole blood CO-oximeter that directly measures total hemoglobin concentration and the relative concentrations of oxyhemoglobin carboxyhemoglobin concentration methemoglobin. In addition, oxygen content oxygen capacity and oxygen saturation indices are calculated. | AVOXimeter 1000, AVOXimeter 4000 | 2 | May 19, 2011 | 2011 | International Technidyne Corp. | An incorrect calibration code that affects the total hemoglobin measurement was applied to affected Avoximeter instruments. | DESIGN: Device Design | International Technidyne Corp. sent a letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to check their inventory to determine if they have any of the affected product. If they have any of the affected product, customers were instructed to stop using it and remove it from their inventory. Customers were also instructed to complete the attached Customer Account Tracking Form and return it by fax to (732-635-0144) or e-mail (techsupport@itcmed.com) or mail to:ITC Technical Support20 Corporate Place SouthPiscataway, New Jersey 08854For any questions call the Distributor or ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International), ext 4707. |  | 17 | 17 | Worldwide Distribution - USA including TX, KS, OH, OR, FL, GA, NY, MA, and DC and the countries of Taiwan and Thailand | Terminated | October 25, 2012 | 525 | 17 | Other | Output/Calculation | Remove | Remove or Replace
Z-2202-2011 | 58179 | spinal vertebral body replacement device | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Theken L"POD" Titanium Vertebral Body Replacement System, Vu c"POD" PEEK-OptimaÂŽ Vertebral Body Replacement System, c"POD" Titanium Vertebral Body Replacement SystemThe Theken CPOD / LPOD VBR is indicated for use in the thoracolumbar spine (T1 -L5) to replace a collapsed, damaged, or unstable vertebral body due to tumor or trauma (e.g., fracture). The Theken CPOD / LPOD VBR is designed to restore the biomechanical integrity of the anterior, middle, and posterior spinal column even in the absences of fusion for a prolonged period. Bone graft material is recommended to be packed into the interior openings of the device prior to implantation.The Theken Vu C POD VBR System is indicated for use in the thoracolumbar spine (T1 -L5) to replace a collapsed, damaged, or unstable vertebral body due to tumor or trauma (e.g., fracture). | THEKEN CPOD/LPOD VERTEBRAL BODY REPLACEMENT SYSTEM | 2 | May 19, 2011 | 2011 | Integra LifeSciences Corp. | Certain Theken Spine Vertebral Body Replacement Systems were distributed (i) without certain components (spacers and set screws) that are necessary for their use as vertebral body replacements, and (ii) with a component (insertion tool) that is not relevant to the devices' intended use. | OTHER/UNDETERMINED: Pending | The firm, Integra Spine, sent an "URGENT: RECALL NOTIFICATION" letter dated January 18, 2011, to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their current inventory and return all remaining kits; complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via fax to: 877-558-6227 or email: melissa.niosi@integralife.com, and contact Integra Spine Customer Service at 866-942-8698 regarding availability of suitable replacement products.Integra Spine will be taking immediate action to restore your inventory and ensure minimal impact on availability for surgical procedures.Should you have any additional questions, please feel free to contact me at 609-936-2495 or email at jon.caparotta@integralife.com. | No consumer action necessary | 93 | 93 Kits | Nationwide distribution: USA including states of: AL, AZ, CA, FL, ID, IL, MD, MI, NH, NM, NV, OH, OK, OR, PA, TX, UT and WI. | Terminated | October 27, 2011 | 161 | 93 | Not_Computer
Z-2228-2011 | 58675 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Kits and Trays containing both alcohol swabstick/prep pad and lubricant: DeRoyal (R) AMNIOCENTESIS TRAY, LINK DREW LACO HOSP, REF 50-11656.02, Rx Only, STERILE EO, Manufactured by: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA. | DeRoyal (R) AMNIOCENTESIS TRAY | 2 | May 20, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contained either one or more of the following possible contaminated recalled products: Triad lubricating jelly, alcohol prep pads, and alcohol swabs. |  | DeRoyal notified their consignees by letter with the subject line "Re: DeRoyal Recall of Triad Medical Lubricating Jelly, Alcohol Prep Pads and Alcohol Swabs" on 02/07/2011 that kits in their possession contained recalled Triad products that were possibly contaminated. Customers were to use the attached spreadsheet to identify and quarantine the affected product. The spreadsheet should have been completed even if there was no affected product. The form should have been faxed back to 865-362-3716 or emailed to recalls@deroyal.com. Customers' DeRoyal Sales Representative would re-label the affected product in order to prevent further use. The DeRoyal Sales Representatives are also providing in-service training on effectuating this recall, including the importance of identifying, retrieving and discarding the recalled products as well as the documentation necessary to account for these products. Distributors were directed to recall down to the retail level. If there are questions or if further assistance is needed, customers can contact their DeRoyal Sales Representative or customer support department at 1-800-251-9864. |  | 340976 | 340,976 units total | Nationwide Distribution | Open |  |  | 340976 | Not_Computer
Z-2227-2011 | 58584 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UnicelÂŽ DxH 800 CoulterÂŽ Cellular Analysis System.A quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. | UniCel DxH 800 Coulter Cellular Analysis System | 2 | May 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the auto prune will not remove patient results if the patient record includes demographics where the results are less than 30 days old. | DESIGN: Software Design | The recall communication was initiated on 4/12/11 with Beckman Coulter forwarding an Urgent Product Correction (PC) letter with attached PC Response Form (via US Postal Service for US customers) to all customers who purchased the UniCelÂŽ DxH" 800 CoulterÂŽ Cellular Analysis System PN 629029. The letter provides customers with an explanation of the problem identified and an action to be taken. For Issue 1 & 2, customers were instructed to perform manual database maintenance procedures to locate and delete unnecessary patient result and demographic information that extends beyond the life (in days) of the control lots. Ensure an alternative method of analysis is available during database maintenance.For Issue 3, customers were instructed to not attempt to transmit the files again after receiving an incomplete report. Instead, submit the missing file(s) by one of the following methods:- Submit your data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)- Fax printouts of the missing files to 786-639-3999- Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.- Include the IQAP number on any communication. A corrected report will be returned to the customers' laboratories.For issue 4, customers were instructed to submit their control data to IQAP by one of the following methods:-Submit their data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)-Fax printouts of the missing files to 786-639-3999-Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.-Include your IQAP number on any communication.Customers were instructed to complete and return the enclosed Response Form within ten days. Customers were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. | Contact the recalling firm for information | 673 | 673 units | Worldwide Distribution -- US, Canada, Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Republic of Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 369 | 673 | Software
Z-2231-2011 | 58569 | products, red-cell lysing products | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTERÂŽ Ac"T" 5diff Fix Reagent. Part Number: 8547171COULTERÂŽ AC"T" 5diff Fix, Instructions for Use, PN 772264. Intended for use as a leukocyte reagent to differentiate subpopulations of cells on COULTER Ac"T 5diff hematology analyzers. Intended for use only with specific Ac"T 5diff reagents. Refer to the instrument product manuals and/or online help, as applicable. | COULTER Ac"T" 5diff Fix Reagent. | 2 | May 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that an issue has been identified with one of the raw materials used to manufacture the Coulter A T 5 diff PN 8547171. | OTHER/UNDETERMINED: Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter on April 8, 2011, with attached PCA Response Form to all customers who purchased the COULTERÂŽ Ac"T" 5diff Fix Reagent. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discard the product according to their laboratory safety procedures and use the attached response form to request replacement product. All flagged results should be reviewed as outlined in the Instructions for Use manual and your laboratory protocol. A retrospective review of results is not required if all recommendations for review have been followed as shown in the Instructions for Use.Customers were instructed to share the information with their laboratory staff and to retain the notification as part of their Quality System documentation. If the product was further distirbuted to another laboratory, customers should provide them a copy of the notification letter.The raw material issue has been corrected beginning with COULTERÂŽ A T 5diff Fix lot 16902E.Customers were instructed to complete and return the enclosed Response Form within ten (10) days.Customers with any questions concerning this notice, were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. | Contact the recalling firm for information | 4662 | 4662 units total (2278 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Botswana, Bulgaria, Czech Republic, Estonia, Finland, France, Germany, Greece, Guyana, Italy, Japan, Kenya, Mexico, Morocco, Namibia, Netherlands, Nigeria, Philippines, Puerto Rico, Russian Federation, Slovakia, South Africa, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Zambia | Terminated | June 04, 2012 | 381 | 4662 | Not_Computer
Z-2232-2011 | 58464 | camera, surgical and accessories | KQM | Surgical camera and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | MAKO PN 204588 Camera204588 Label With Camera Stand Number ***PN 200294 SN CAM XXX 2011-01.*** Weight 86.2kg.***Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995, UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.*** Manufactured in USA*** MAKO SURGICAL CORP. 2555 Davie Road, Ft. Lauderdale, FL 33317.***204589 Robot Arm Label***PN 203999 SN ROB XXX 2011-01.***V 100/120/230 A 9.6/8.0/4.2 Hz 50/60***Weight 39.4kg***Class I Equipment Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995, UL 60601-1:2003 CAN/CSA-C22.2 No. 601-1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 Davie Road, Ft. Lauderdale, FL 33317.***204588 Label With Guidance Module Number***PN 201251 SN GUD XXX 2011-01***Weight 51.7kg***Class I Equipment. Conforms to IEC 60601-1/A2: 1995, En 60601-1/A2: 1995, UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90***Manufactured in USA***MAKO SURGICAL CORP. 2555 Davie Road, Ft. Lauderdale, FL 33317.***THE RIO uses an infra-red camera for the purpose of spatial navigation. This camera is present on the Camera Stand and communicates with the Guidance Module to provide the locations of the tracked instruments (i.e. robotic-arm, patient trackers, probes, etc). The information is then used to provide the virtual representation of the patient anatomy and the tools on the monitor, and provides necessary information to the robotic-arm to ensure proper bone preparation. | MAKO PN 204588 Camera | 2 | May 20, 2011 | 2011 | Mako Surgical Corporation | MAKO Surgical Corporation recalled ONE infra-red camera MAKO PN 204854 which was a component of their RIO Robotic Arm Interactive Orthopedic System with SSN P7-03127, on November 22, 2010. This component can function improperly and cause the accuracy of the system to be outside of the specification listed by the manufacturer. The recall was initiated after the camera manufacture, NDI, Waterloo, On | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, MAKO, sent a "RECALL NOTICE" letter dated November 22, 2010, to its customer. The letter described the product, problem and action to be taken. The customer was instructed to complete and return the attached Acknowledgement Form to their MAKOplasty Specialist (MPS) via mail to: MAKO Surgical Corp., Attn: Jim Pomeroy, 2555 Davie Road, Fort Lauderdale, FL 33317. The letter stated that "It is important that the camera be replaced immediately, and MAKO is currently addressing the situation with the camera manufacturer." In addition, MAKO will replace the camera with a camera that is not affected by the recall and verify the customers system to meet the functional specifications.If you have any questions or concerns, please contact your MPS or Sr. Director of Quality Assurance at 954-927-2044 ext 604, and /or Sr. Director, Customer Support at 954-927-2044 ext. 442. | Contact the recalling firm for information | 1 | 1 | Nationwide distribution: NY only. | Terminated | November 09, 2011 | 173 | 1 | Hardware | Output/Calculation | Replace component | Remove or Replace
Z-2225-2011 | 58333 | dermatome | GFD | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Universal Power Supply Units, Part Number DP0007 and DP0800 , for use with the Integra Padgett Dermatomes - Model S, SB, and S6, Device Listing #E208841, 510 (k) exempt.The Integra Padgett Dermatomes are intended for use in the cutting of skin grafts for use in plastic surgery and/or burn treatment. The Model S and S6 dermatomes are intended for professional use in hospitals, and similar medical facilities, where the patient will be under the supervision of trained personnel. The Model SB dermatomes are used in skin banks for the harvesting of cadaver tissues. | Padgett Dermatome power supply units | 2 | May 20, 2011 | 2011 | Integra LifeSciences Corp. | Dermatome power supply units distributed by the recalling firm may not meet the required medical device electrical safety regulations which can result in an increased potential for electric shock to the device user and/or patient. | OTHER/UNDETERMINED: Pending | Integra LifeSciences Corp. initiated their recall by contacting their Sales Representatives via telephone conference calls on 3/10/2011. Separate written recall communications, in the form of an 'URGENT PRODUCT RECALL NOTIFICATION' were provided via E-mail and FedEx to the firm's Sales Representatives on 3/11/2011; and to their hospital and distributor customers on 3/11/2011; 3/14/2011; and on 4/12/2011. The recall notification addressed to the non-Integra Sales Rep customers informs the customers of the problem with the power supply units (that they may not meet the required medical device electrical safety standards which may result in an increased potential for electrical shock). The notification goes on to ask that the customers examine their inventories and determine if they have any affected product on-hand; if so, the customers are instructed to discontinue use of the power supply unit and return the unit(s) to Integra LifeSciences Corporation immediately for reworking. The letter also provides photographs of the affected units which the customer can use to locate and verify the serial number(s) of the device (if applicable). In addition, the notification is accompanied by a 'RECALL ACKNOWLEDGMENT AND RETURN FORM' which the customer is asked to sign and complete in order to document receipt of the recall notification and whether or not the customer has any of the recalled power supply units available for return. The notification letter provided to the Integra Sales Representatives describes the device models affected; the reason for the recall; and asks that the Reps. examine their inventories for the presence of any of the affected power supply units. If found, the Reps are instructed not to use, or to further distribute any of the affected units in their possession, but to immediately make arrangements to return the unit(s) to Integra for rework. The notification supplied to the Sales Reps provides them with a synopsis of the firm's field action. | Contact the recalling firm for information | 892 | 892 units were distributed in the U.S & 155 units were distributed in foreign countries | Worldwide Distribution - USA including FL, CO, IL, AR, MD, IN, CA, MT, AZ, NY, LA, TX, MA, MO, CT, MI, NE, OH, NC, WV, VA, PA, NJ, TN, UT, MN, OK, WI, GA, AL, WA, KS, OR, AK, DC, SC, NE, MS, RI, NV, IA, and ND and the country of Canada | Terminated | April 17, 2012 | 333 | 892 | Not_Computer
Z-2237-2011 | 58563 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Nasal Cannula (Adult) Salter Style with modified nasal prongs, 7" (2.1M) Q supply tube and CO2 sample tube with male luer connector 22mm (I.D) x 6mm (O.D) Adaptor. Adult Divided Cannula.Single Patient Use. Do not SterilizeSalter Labs, 100 W. Sycamore Road, Arvin, CA 93293Used to deliver supplemental oxygen or airflow to a patient or person in need of respiratory help. | Nasal Cannula (Adult) Salter Style | 2 | May 20, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | Complaint received that the product in the box was different from labeled product. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Salter Labs, sent an "URGENT PRODUCT RECALL" subtitled Adult Divided Cannula Product Code 4707 Lot 012411, letter dated April 19, 2011 via certified mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory; return any of the lot in their possession to Salter; contact Salter Customer Service at 1-800-235-4203 to arrange for the return of the product, and complete and return the Customer Reply Form via fax to 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 22150 | 22,150 pieces | Nationwide distribution. | Terminated | August 23, 2011 | 95 | 22150 | Not_Computer
Z-2226-2011 | 58555 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Systems Calibrator 1, Part Numbers: 468405, Lots prior to M005558The Beckman Coulter SYNCHRON Calibrator 1 (CAL 1), used in conjunction with SYNCHRON reagents, is intended for the calibration of immunoprotein tests on SYNCHRON LX and UniCel DxC Systems. | SYNCHRON Systems Calibrator 1 | 2 | May 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because standardization of the Synchron Systems Calibrator 1, used for both the Synchron LX and UniCel DxC systems, has been established to the new IFCC Reference Preparation for Plasma Proteins, ERM-DA470k (formerly CR470) beginning with lot number M005558.During the standardization to the new IFCC standard, ERM-DA470k, we confirmed that Complement C4 (C4) had been re |  | Beckman Coulter initiated a Product Corrective Action (PCA) letter on the week of March 14, 2011, to the affected customers. Though Beckman is recalling the LX Cal 1, the recall letter was sent to the C4 reagent customers, as these are the only customers who will be impacted by the problem noted in the PCA letter. The letter provided the customers with an explanation of the problem identified and an action to be taken.The letter recommended consignees take the following actions:(1)Notify their clients to review and evaluate the impact on patients whose Complement C4 recovered near the clinical decision points.(2) Re-evaluate their laboratory's reference ranges for C4.The letter further requested that consignees share this information with laboratory staff and retain this notification as part of your laboratory Quality System documentation.Questions regarding this Product Announcement letter, or requests for the correlation graphs should be directed to Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Request or call (800) 854-3633 in the United States and Canada. |  | 96173 | 96,173 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Czech Republic, Denmark, Eritrea, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Namibia, Netherlands, New Caledonia, New Zealand, Oman, Panama, Poland, Puerto Rico, Qatar Reunion, Russian Federation, Rwanda, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Viet Nam | Terminated | May 04, 2012 | 350 | 96173 | Not_Computer
Z-2224-2011 | 58558 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac and TrueBeam, High Energy Linear Accelerator, Model Numbers: H14, H19, H29; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CAIndicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Clinac | 2 | May 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The coolant may leak. The resulting coolant leak presents a risk of electrical shock to any person working within the protective housing. | DESIGN: Process Design | On 3/31/2011, all consignees were notified via an Urgent Medical Device Correction letter and on-site visit by the firm's sales representative, informing them of the affected products, the issue, and recommended actions. If non-Varian personnel service the device, then they should have been immediately notified of the content of the letter. Anyone who performs work on the device should exercise extreme care to identify any possible coolant leak and immediately locate and correct the source of the leak. Customers were to immediately notify Varian Service or their internal service personnel of any coolant leakage. Varian was to contact customers to arrange for inspection if one had not been performed. The appropriate personnel should have been advised of the content of the letter. If further clarification was needed, customers were to contact their local Varian Customer Support. | Contact the recalling firm for information | 718 | 718 units | Worldwide Distribution | Terminated | March 13, 2012 | 298 | 718 | Not_Computer
Z-2242-2011 | 58564 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs Nasal Cannula (Adult) Salter Style with 14' (4.3) supply tube.Single Patient use. Do not sterilize.Used to deliver supplemental oxygen or airflow to a patient or person in need of respiratory help. | Salter Laboratories Adult Nasal Cannula | 2 | May 23, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | One complaint received of product in box different from what was on label. | OTHER/UNDETERMINED: Pending | Salter Laboratories sent a URGENT PRODUCT RECALL NOTIFICATION letter dated April 19, 2011, certified mail to all of the affected consignees, along with a RECALL EFFECTIVENESS CHECK FORM to acknowledge receipt of this communication.It was reported that a package labeled 1600-14, Adult Nasal Cannula actually contained a 25-foot length of tubing and has determined that additional product may have been incorrectly labeled in this manner.No illness or injury would result from this product mix-ups.Should you have any questions, please contact Salter Laboratories at 1-800-235-4203. |  | 3500 | 3500 pieces | Worldwide Distribution - USA (nationwide) and the country of South America. | Terminated | August 23, 2011 | 92 | 3500 | Not_Computer
Z-2240-2011 | 58487 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Stryker, Impression, Non-Powered Mattress, Model 2980, Stryker Medical, Portage, MI.The Impression non-powered mattress is recommended for use in the prevention of pressure ulcers. It is intended for use on Stryker frames and on other frames supporting a 84" x 35" x 6" mattress. | Styker Impression Mattress | 2 | May 23, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Affected mattresses may not conform to federal flammability standards. | OTHER/UNDETERMINED: Pending | Stryker Medical sent an URGENT MEDICAL DEVICE RECALL letter dated April 29, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate all affected mattresses , remove the existing mattress cover, and replace it with an enclosed mattress cover. Customers were then to notify Stryker to confirm repairs. If the mattresses were loaned or sold, Stryker requested customers to notify the new user of the recall. Customers with questions or concerns, were asked to call 269-389-6689. |  | 213 | 213 | Worldwide Distribution - USA including California and the country of Canada. | Terminated | October 07, 2011 | 137 | 213 | Not_Computer
Z-2243-2011 | 58533 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system;Therapist Express Basic, syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS.The ONCOR and PRIMUS may be updated to utilize the RTT v4.x software.Distributed by Siemens, Healthcare, Concord, CA 94520Radiation therapy | ARTISTE , ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system; | 2 | May 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens found a potential safety issue related to patient treatment with table offsets on syngo RT Therapist (RTT) 4.1 and 4.2 which could result in dose being delivered to wrong location. | DESIGN: Software Design | Siemens initiated an Urgent Medical Device Correction letter on March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:use only relataive table setup with all planned table values '0' if applicable.In cases where the treatment plan contains at least one table parameter other than '0' to consider the following:Use only one PVG plus resulting offset per fraction.Once an offset has been applied do not use 'apply offset' at RTT for any further offset.Select the PVG and treatment beams in the same auto-sequence. In casethis is not possible, the table position for each treatment beam, in the fraction, needs to be verified separately (especially for treatments with beams having different table posiitons. If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Scenario 2 - Manual change of table posiiton followed by a PVG calculated offset.Always deliver the PVG's at the planned table position. Avoid manual change of table posiiton prior to applicaiton of PVG.Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, need to be verified separately (especially for treatments with beams having different table positions).If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Customers were informed that a software update would be provided as soon as possible.Customers were asked to include the Customer Sa. | Contact the recalling firm for information | 164 | 164 active devices | Worldwide Distribution - USA including CA, DE, FL, KY, LA, MA, MI, MO, NC, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV and the countries of Australia, Belgium, Canada, Croatia, Egypt, France, Germany, Hungary, India, Ireland, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Turkey, and the United Kingdom. | Terminated | March 30, 2012 | 312 | 164 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2238-2011 | 58789 | table, radiologic | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Maquet Magnus Carbon-fibre table top; Device Part Number 1180.16xx (x represents a digit from 0-5 and F-for the US market)The carbon-fibre MAGNUS table top serves to support and position patients for surgical procedures and interventional examination immediately before, during and after the surgical phase. The table top is radio translucent and enables intra-operative use of x-ray equipment. | Maquet Magnus Carbonfibre table top | 2 | May 23, 2011 | 2011 | Maquet Inc. | Maquet GmbH's quality monitoring detected that incorrect cover sheets were installed with some Magnus carbon fibre table tops. | DESIGN: Device Design | Maquet sent an Urgent Medical Device Corrective Action letter dated 4/15/2011, along with a Action Response Form, via Federal Express to their three customers informing them the incorrect cover sheets were installed with some of the affected "MAGNUS 1180.16F0 carbon-fibre table tops." The firm will have a Service Technician contact you to service the OR table, and will be exchanging the cover sheet free of charge. To the customers, please follow the instructions on the response form. At this time there have been no reports that injury has occurred.If you have additional questions, please contact yur local MAQUET Customer Service at 1-888-627-8383. |  | 3 | 3 | LA, MO, VA | Terminated | August 21, 2013 | 821 | 3 | Not_Computer
Z-2241-2011 | 58582 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501.The device is labeled in part: "***Neusoft***NeuViz16 MULTI-SLICE CT SCANNER SYSTEM***MANUFACTURED FOR NEUSOFT MEDICAL SYSTEMS CO., LTD***MANUFACTURED BY PHILIPS AND NEUSOFT MEDICAL SYSTEMS CO., LTD***ADDRESS: NEUSOFT PARK, HUNNAN INDUSTRIAL AREA, SHENYANG***CHINA***". | NeuViz 16 MultiSlice CT Scanner System | 2 | May 23, 2011 | 2011 | Philips And Neusoft Medical Systems Co., Ltd. | Dimension measurement displays on a combined image were found inaccurate after the clinician merged multiple image series into one image. | OTHER/UNDETERMINED: Pending | Neusoft Medical Systems sent an Urgent Device Correction Notification letter dated February 9, 2011, via certified mail to their distributor NEUISYS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of combined images. Instead, use MPR Batch for diagnostic reviews until the planned upgrade is provided to the product.Neusoft Medical System and Philips Medical System released the mandatory Field Change Order (FCO) by March 31, 2011. All nine affected systems will be updated with aforementioned kits in six months after the release.Consignees were instructed to contact Service Support Department by e-mail to nms-service@neusoft.com. | Contact the recalling firm for information | 9 | 9 units distributed in US | Nationwide Distribution including CA, GA, LA, NC, OH, TN and TX. | Terminated | August 24, 2011 | 93 | 9 | Other | Output/Calculation | Upgrade | Repair
Z-2239-2011 | 58685 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BardÂŽ Soft Mesh 4"x6" (10cm x 15cm) Flat, Product Code: 0117010. The BardÂŽ Soft Mesh patch is a non-absorbable, sterile prosthesis. It has a large pore design and is constructed of knitted polypropylene monofilaments. The device classification name is "Mesh, Surgical, Polymeric". | Bard Soft Mesh patch | 2 | May 23, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeling: complaints of product labeled as BardÂż Soft Mesh 4" x 6" containing BardÂż Soft Mesh 6" x 6" | PRODUCTION CONTROLS: Labeling Mix-Ups | Davol issued an URGENT: MEDICAL DEVICE RECALL notification letter to customers dated April 25, 2011, informing them the recall involving a single lot of the BARD Soft Mesh - Lot# HUUF0530. This recall has been initiated because this lot number may contain a different size patch than is indicated. Customers were instructed to immediately examine inventory and quarantine product subject to recall. Customers are asked to identify and notify their customers to whom recalled product may have been further distributed. A Recall Effectiveness Check reply form should be completed and returned to Davol via fax (401-825-8753). For directions on returning product or for questions regarding the recall, contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services & Support Department at 1-800-562-0027. |  | 198 | 198 units | Nationwide Distribution | Terminated | October 28, 2011 | 158 | 198 | Not_Computer
Z-2246-2011 | 58565 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 6.5 Client Software.IMPAX 6.5 combines the traditional activities of RIS informatics management with PACS image management to provide a powerful platform for imaging-based planning, interpretation, and results distribution. The IMPAX 6.5 Client focuses on the integration of PACS, RIS and Speech applications into a single delivery of information. A single IMPAX 6.5 Client application can be used by a range of users on any appropriate, networked workstation they have access to. | IMPAX 6.5 | 2 | May 24, 2011 | 2011 | AGFA Corp. | Software issue that occurs after dictation of a study. | OTHER/UNDETERMINED: Pending | Agfa Corp. sent an "URGENT SAFETY NOTICE" customer letter via Agfa email to customers on 4/14/2011, along with a "URGENT SAFETY NOTICE FEEDBACK FORM," The letter identified the product , the problem, and the action to be taken by the customer.It was discovered a software issue in the IMPAX 6.5. The upgrade has already been released to address the software problem.Customers were instructed to fax back the feedback form as soon as possible to (864) 421-1664.If you have any questions, please contact your local Agfa Healthcare organization. | Contact the recalling firm for information | 95 | 95 units | Nationwide Distribution -- including states of WI, and SC. | Terminated | September 15, 2011 | 114 | 95 | Software |  | Software update | Software Update
Z-2245-2011 | 58328 | dental implant surgical tray (kit) | OFY | Endosseous dental implant accessories. | Dental | Dental | Enforcement Discretion | ***REF INT411***OSSEOTITE Tapered Certain Implant 4 x 11.5mm***Sterile Using Radiation***Rx Only***BIOMET 3i Dental lberica S.L.WTC Almeda Park, Ed, 1 Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain***Phone +34 934-705-500; Fax +34 933-717-849.Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Osseotite NT Certain | 2 | May 24, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite NT Certain Dental Implant INT411 . The internal thread was not manufactured correctly. | OTHER/UNDETERMINED: Pending | Biomet 3i, LLC sent an Urgent Medical Device Recall letter dated March 14, 2011, via email or fax to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A small percentage of the implant lots may have a condition affecting the internal thread. The condition, if present, will prevent the cover screw, temporary healing abutment, final restorative abutment, etc. from engaging the implant. Customers were instructed to contact Customer Service at 1-800-342-5454 if this condition is recognized upon attempting to assemble any of the screws with the implant.Customers were instructed to return the implant to the firm for replacement to the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CC117886Customers were also instructed to respond with the attached Response Fax or email to DomesticComplaints@Biomet.com.For any questions regarding this recall call 561-776-6906. | Contact the recalling firm for information | 236 | 236 | Worldwide Distribution - USA (nationwide) and the countries of Europe, Japan, Asia, and South America. | Terminated | January 25, 2012 | 246 | 236 | Not_Computer
Z-2244-2011 | 58550 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT - FA Reagent BottleThe BacT/ALERT Microbial Detection System is used to determine if microorganisms are present in blood or other normally sterile body fluid samples taken from a patient suspected of having bacteremial/fungemia. The BacT/ALERT System and culture bottles provide both a microbial detection system and a culture media with suitable nutritional and environmental conditions for organisms commonly encountered in blood infections and other normally sterile body fluid infections. An inoculated bottle is placed into the instrument where it is incubated and continuously monitored for the presence of microorganisms that will grow in the BacT/ALERT FA bottle. BacT/ALERT FA Culture Bottles are used with the BacT/ALERT Microbial Detection System in qualitative procedures for enhanced recovery and detection of aerobic and facultative anaerobic microorganisms (bacterial and fungi) from blood and other normally sterile body fluids. | BacT/ALERT FA Reagent Bottle | 2 | May 24, 2011 | 2011 | bioMerieux, Inc. | The charcoal suspension in a portion of the lot does not meet the intended density. | OTHER/UNDETERMINED: Pending | BioMerieux sent an Urgent Product Removal Notice dated April 13, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Evaluate positive results adhering to their institution's protocols for determining false versus true positives. Negative bottles should be checked by smear and/or subculture at some point prior to discarding as negative. Any remaining inventory of BacT/ALERT FA culture bottle lot 1027481 should be destroyed in accordance with their facility's policies and procedures. An Acknowledgement and Product Replacement Form was included to be completed and returned to bioMerieux or credit or replacement of unused bottles. |  | 327700 | 327,700 bottles | Worldwide Distribution - USA (nationwide) and the countries of China, Tokyo, South Korea, Canada, Sweden, Austria, Poland, and Belgium | Terminated | April 03, 2012 | 315 | 327700 | Not_Computer
Z-2293-2011 | 58545 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics version 9.0, Siemens Medical Solutions, USA, Inc., Ann Arbor, MI.The syngo Dynamics is a picture archiving and communications system intended for acceptance, transfer, display, storage, archiving and manipulation of digital medical images, including quantification and report generation. It is not intended to be used for the reading of mammography images. | syngo Dynamics version 9.0 | 2 | May 25, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | If one or more graphs are in a table on the report, and if any of those graphs are deleted, it may lead to reports failing to save and/or marking the study as read might leave the study in an unread state or might mark the study as read but leave the report with a "Preliminary" watermark displayed. | DESIGN: Software Design | The firm, SIEMENS, sent a "Customer Safety Advisory Notice" letter dated March 24, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to place graphs in sections and not tables in the report; avoid deleting graphs from the reports in run-time and include this safety advisory notice in their operator's manual where it should remain until the service patch is applied. SIEMENS will be issuing a service patch to fix the defect. The letter states the patch will be available in the next few weeks.If you have any questions, call +1 (734) 205-2400. | Contact the recalling firm for information | 29 | 29 devices | Nationwide distribution: USA including states of: CA, CT, FL, ID, MD, MI, MN, NJ, NY and PA. | Terminated | July 10, 2012 | 412 | 29 | Software | Device Operation | Software update | Software Update
Z-2233-2011 | 58373 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | MicroStream Filterline Infant/Neonatal products: Part # 006285 - FILTERLINE H SET INF/NEO 25 UN PHYSIO (Inf/Neo Airway Adapter). The adult/pediatric airway adapter and the adult/pediatric FilterLine and VitaLine sets are not affected by this recall. The intended use of Microstream Filter Lines is to convey a sample of a patient's expired breath to a capnograph in order to measure the CO2 partial pressure in intubated patients. It is a component of the FilterLine H Set Inf/Neo, Part # 006324 and the VitaLine H Set Inf/Neo, Part # 010807. | MicroStream Filterline ICU | 1 | May 25, 2011 | 2011 | Oridion Medical 1987 Ltd. | fine plastic strands on inner surface of the infant neonatal airway adapter may become dislodged and inhaled by the patient | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Oridion, sent an "Microstream Filterline Field Correction Notice" letter dated March 25, 2011 to its US and foreign customers via telephone, email with the documentation discussed in the call, and letter. The letter described the product, problem and actions to be taken. The customers were instructed to place all affected product held in their inventory on ship hold; gather these units in their inventory and at the customer sites and exchange with the new product; inspect filter line starter (sample) kits shipped with Capnostream 20 and Microcap patient monitors for affected product, if affected Infant/Neonatal filter lines are found, remove them from the starter (sample) kit; send a purchase order in order to obtain replacement product for units in their stock and at customers sites, and complete and return the Inventory Reporting Form Excel Spread Sheet within two weeks of receiving the email even if they or their customers do not have the affected inventory. Oridion shall conduct a removal of the products from the related production lots from the field and an immediate replacement with product from subsequent production lots that meet all requirements. This action will be conducted by Oridion's distributors or by Oridion directly in some cases. Oridion anticipates that new product will be available for shipment during the week of April 4th.Questions regarding the recall can be directed towards the Customer Service Coordinator at: 1-888-674-3466 x1234 or NA-FCA@oridion.com. | Contact the recalling firm for information | 350 | 350 units in US (14 boxes with 25 units ea) | Worldwide distribution. | Terminated | March 19, 2012 | 299 | 700 | Not_Computer
Z-2291-2011 | 58420 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. | Logix Order Entry (OE) Software | 2 | May 25, 2011 | 2011 | Baxter Healthcare Corp. | The LOGIX OE automated backup does not retain the three most recent backup files and instead will only retain the last three backups of the previous year. The software stores backup files in numeric order rather than chronological date order. | DESIGN: Software Design | Baxter Healthcare Corp. sent an URGENT DEVICE CORRECTION letter dated March 29, 2011, via first class mail to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. It was discovered there is an issue with the LOGIX OE automated backup not retaining the three most recent backup files and instead will only retain the last three backups of the previous year. The accounts were requested to turn off the automated backup function on the LOGIX OE Software, and perform manual backups following the instructions included in the Urgent Device Correction letter, and retain a copy of the letter at each computer on which LOGIX OE Server is installed. The customer were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. | Contact the recalling firm for information | 84 | 84 units | Worldwide Distribution -- USA including Puerto Rico, and the country Canada. | Open |  |  | 84 | Software | Device Operation | Instructions on turning off auto backup | Safety Notice/Insructions
Z-2249-2011 | 58021 | tray, catheterization, sterile urethral, with or without catheter (kit) | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Case Label: McKESSON Medi-Pak Irrigation Tray With Bulb Syringe Sterile Single Use Disposable Latex-Free 20 Trays Per Case Reorder No. 37-103 Marketed By McKesson Corporation McKesson Medical-Surgical, Richmond, VA 23228 Made in the U.S.A.Tray Label: McKESSON Medi-Pak Irrigation Tray With Bulb Syringe Sterile Single Use Disposable Latex-Free Contents: 1 - Graduated Tray, 1200 cc 1 - Green Bulb Syringe (Catheter Tip), 60 cc 1 - Graduated Solution Container, 500 cc 1- Underpad, 1 - Sterile Alcohol Prep Pad Reorder No. 37-103 Contents STERILE in unopened, undamaged package. Marketed by McKesson Corporation McKesson Medical-Surgical, Richmond, VA 23228 Made in the U.S.A.Usage: The trays are irrigation trays for catheters | McKesson | 2 | May 25, 2011 | 2011 | Nurse Assist, Inc | The irrigation procedure trays contain Triad Group Alcohol Prep Pads recalled by Triad Group because they have the potential to be contaminated with an objectionable organism. | OTHER/UNDETERMINED: Pending | The firm, Nurse Assist Incorporated, sent a letter dated January 20, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory immediately; quarantine the product; identify their customers and notify them at once about the product recall (Nurse Assist will provide bright labels to affix to your affected inventory); provide their customers with labels as well, and complete and return the attached RECALL ACKNOWLEGEMENT form via fax to: 817-231-1500 or mail to: Nurse Assist, Inc., 3400 Northern Cross Blvd., Forth Worth, TX 76137.Only the product and lots identified on the attached list are affected.If you have any questions or to request labels, please call Nurse Assist Customer Service Monday through Friday, from 8:00 am to 5:00 pm Central Time at 1-800-649-6800. |  | 2840 | 2840 cases | Worldwide distribution: USA (nationwide) and countries including: Canada, Dubai and Saudi Arabia. | Terminated | October 25, 2012 | 519 | 124789 | Not_Computer
Z-2294-2011 | 58600 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Zoll E Series BLS Defibrillator/Pacemaker/MonitorSoftware versions 7.10 and 7.11The product is used by qualified medical personnel for the purpose of converting ventricular fibrillation to sinus rhythm or other cardiac rhythms capable of producing hemodynamically significant heart beats. In addition, the product may be used in synchronized mode to terminate certain atrial and ventricular tachycardias and other arrhythmias resistant to drug therapy. The E Series may be used for cardiac pacing in conscious or unconscious patients for up to a few hours duration as an alternative to endocardial stimulation. The E Series may be used for the monitoring a patients electrocardiogram (ECG). | Zoll | 2 | May 25, 2011 | 2011 | ZOLL Medical Corporation | Zoll E Series BLS Defibrillator may fail to auto charge and deliver energy | DESIGN: Software Design | The firm, Zoll, sent an "Urgent Medical Device Correction" letter dated April 22, 2011 via UPS to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately alert all E Series BLS users to this problem; instruct users to not touch patient and keep the patient still during ECG analysis, Instruct users that if they experience this problem, to press the Analyze button again, and contact ZOLL Technical Service for the software upgrade kit. The letter also states that affected units can remain in service until the software has been updated. If you have any questions, please contact 24/7 technical support at 1 (800) 348-9011 or +1 (978) 421-9460. |  | 874 | 874 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Finland, Germany, Malaysia and United Kingdom. | Terminated | January 23, 2014 | 974 | 874 | Software |  | Software update | Software Update
Z-2292-2011 | 58601 | keratome, ac-powered | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | WaveLight FS200 Femtosecond Laser System Alcon Product Code 8065990714 WaveLight Model 1025 Software Version: Version Blue, Service Pack 2.The WaveLight FS 200 Laser System is an ophthalmic surgical laser indicated for use in refractive surgery for LASIK flap cutting, lamellar and perforating keratoplasty, and intrastromal surgery for implants. | WaveLight FS 200 Femtosecond Laser System | 2 | May 25, 2011 | 2011 | Wavelight AG | The software calculating the patterns were found to has a rounding error for certain parameter ranges. These rounding errors could result in an incomplete cut in the donor or patient tissue. | OTHER/UNDETERMINED: Pending | WaveLight AG sent an URGENT MEDICAL DEVICE CORRECTION letter April 18, 2011 to the affected customers. The letter identified the product, problem and the action to be taken by the customer. The letter advised the customers todiscontinue keratoplasty cuts with the WaveLight FS200 until the new version can be installed. Alcon will contact you to schedule a visit of a service engineer at your site, to disable the keratoplasty function in your device, until a new software version can be installed to fix this problem. If you have experienced a product problem or adverse event related to this medical device correction, please contact Alcon Medical Safety at 1-800-561-6466, option 3.Customers were requested to complete the attached reply form to acknowledge receipt of this notification and fax it to ( 817) 916-9087. | Contact the recalling firm for information | 55 | 55 units | Worldwide Distribution - USA including states of TX, KS, NC, CA, FL, OH, SC, MD, VA and NY and countries of Switzerland, Germany, Greece, United Kingdom, Ireland, Spain, Italy, Australia, India, Malaysia, France, Finland, Saudi Arabia, Czech Republic, Poland, China and South Korea and Iran. | Terminated | September 23, 2011 | 121 | 55 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2295-2011 | 58612 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50ÂŽ Ventilator. Model #HT50 H1-B The device is intended to provide continuous or intermittent mechanical ventilator support for the care of individuals who require mechanical ventilation. Specifically, the NEWPORT HT50 is applicable for adult and pediatric (i.e. infant, child and adolescent) patients, greater than or equal to 10 kg (22 Ibs), who require the following general types of ventilatory support. | Newport HT50 Ventilator | 2 | May 26, 2011 | 2011 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical Instruments confirmed that the Newport HT50 ventilators that were shipped in December 2008 were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. | OTHER/UNDETERMINED: Pending | Newport Medical Instruments, Inc. sent an FIELD CORRECTION letter dated January 13, 2009, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. The Newport HT50 ventilators that were shipped in December 2008, were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. The customer were also requested to complete the attached customer REWORK INSTRUCTIONS form and fax it to (714) 427-0489. If you have any further questions, please call (714) 427-5811 x 344. | Contact your health professional |  |  | Worldwide Distribution-- USA including states of OH, VA, WA, WI and country of Turkey. | Terminated | June 01, 2011 | 6 |  | Not_Computer
Z-2383-2011 | 58650 | accessories, catheter, g-u | KNY | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus ACMI, REF 25 BX, PTFE Coated 3 cm Flexible Tip Guidewire, .035" x 150 cm, Sterile EO, Rx Only, Gyrus ACMI, Inc. 136 Turnpike Road,Southborough, MA 01772-2104 USAUreteral guidewires are used to establish access in the urinary system for the passage of urological devices. An introducer is first used to establish entry to various portions of the urinary tract (urethra, bladder, ureter), then a guidewire is threaded through the introducer. Following placement of the guidewire, various urological devices can be introduced into the targeted portion of the urinary tract for various therapeutic and diagnostic procedures. | PTFE Coated 3 cm Flexible Tip Guidewire | 2 | May 26, 2011 | 2011 | Gyrus Acmi, Incorporated | Gyrus ACMI, Inc. wants to bring a recall issue to their customers attention concerning one lot of their 25BX-PTFE Coated 3 cm Flexible Tip Guidewire. As a result of a manufacturing error in the production this lot of guidewires, some units of product were loaded in the dispenser backwards. In such cases, the stiff end of the guidewire would be dispensed from the product first and, if undetected, | PRODUCTION CONTROLS: Process Control | The firm, Olympus ACMI, sent an "Urgent: Medical Device Recall" letter dated April 15, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Inspect their stock to confirm if any of the affected lot of products remains in their possession2)Immediately cease any further use of the affected product, remove, quarantine and return it. 3) Call their Gyrus ACMI customer service representative at 1-888-542-7266 to obtain a Returned Goods Authorization 4) Return the product to: Attn: PTFE Guidewire Recall RGA#_________Gyrus ACMI LP, 6845 Wedgwood Court, Maple Grove, MN 553115) Complete and return the RECALL REPLY FORM via fax, even if they are returning no product, to: Regulator Affairs (508) 804-2624. The letter states that the customer will receive repeat mailings until they have faxed in the form. Additionally, If they have further distributed the product they are requested to identify and notify their customers at once.If you have any questions, please do not hesitate to contact customer service at 1-888-524-7266. | Contact the recalling firm for information | 145 | 145 | Nationwide distribution: USA including states of: CA, GA, HI, IN, KY, MT, NY, NC, PA, TX, UT, and WA. | Terminated | August 23, 2012 | 455 | 145 | Not_Computer
Z-2248-2011 | 58471 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Philips MicrostreamÂŽ Model 989803159581 VITALINE H SET INFANT/NEONATAL 25 UN PHILIPSThese products are used with Philips Monitors for the measurement of endtidal CO2 in infants and neonates. They provide a means to transport samples of patient respiratory gases from the patient's airway to the Microstream capnograph. This by performed by placing the airway adapter portion of the FilterLine H Set Infant/Neonate an VitaLine H Set Infant/Neonatal in the patient airway system of a respirator or an anesthesia machine. | MICROSTREAM FILTERLINE ICU | 1 | May 26, 2011 | 2011 | Philips Healthcare Inc. | Fine plastic strands on inner surface of the infant/ neonatal airway adapter may become dislodged and inhaled by the patient. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Philips, sent a Field Safety Notice titled "URGENT MEDICAL DEVICE RECALL" letter dated April 2011 to its customers. Customer notifications outside the US will be managed by Philips representatives in each affected geography. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all products from affected lots; remove from inventory and dispose in accordance with local regulations (new products will be sent proactively to each affected customer). Direct accounts were instructed to notify additional customers to whom the product has been distributed. These customers are instructed to follow the instructions in the "Action to be taken by Customer/User" section of the Notice. Questions regarding the recall can be directed towards the customer's local Philips representative at 1-800-722-9377. |  | 75 | 75 units (3 boxes) | Worldwide distribution. | Terminated | March 14, 2013 | 658 | 25475 | Not_Computer
Z-2382-2011 | 58629 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator. Model No. HT50-H, HT50-HB, HT50-H1, and HT50-H1B.The device is intended to provide continuous or intermittent mechanical ventilator support for the care of individuals who require mechanical ventilation. Specifically, the NEWPORT HT50 is applicable for adult and pediatric (i.e. infant, child and adolescent) patients, greater than or equal to 10 kg (22 Ibs) | Newport HT50 Ventilator | 2 | May 26, 2011 | 2011 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the Newport HT50 hard shell Emergency Preparedness (EP) Storage Cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. | OTHER/UNDETERMINED: Pending | Newport Medical sent an FIELD CORRECTION NOTICE to all affected customers dated May 14, 2010. The letter identified the product, the problem and the action needed to be taken by the customer. Newport Medical also provided a revised labeling for the Emergency Preparedness (EP) Storage Case and an instruction sheet describing where to apply the new labels. Customers were instructed with the information that all the EP storage cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. Customers were instructed to complete the attached acknowledgement form and fax it to " Regulatory@NewportMedical.com " Customers with any questions should call Technical Support Department at (800) 451-3111 for US customers, and (714)427-5811 x500 for international customers. | Contact your health professional | 1113 | 1,113 EP cases | Worldwide Distribution -- USA including states of PA, TX, AK, GA, CA, MD, KY, FL, VA, LA, NC, OK and countries of Canada, Colombia, Armenia and Japan. | Terminated | June 01, 2011 | 6 | 1113 | Not_Computer
Z-2386-2011 | 58431 | breast pump kit | OHH | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | AC Power Adaptor (transformer), TENPAO Model U090100Q31; used with the Pump In Style Advanced family of breast pumps; Medela Inc., McHenry, IL 60051-0650; Made in China; Article #9207010.Converting 120 Volt AC power to 9 Volt DC power used for providing power to the Pump In Style Advanced Breast pumps. | AC Adaptor Article 9207010 | 2 | May 27, 2011 | 2011 | Medela Inc | The power adaptor (transformer) used with the Pump In Style Advanced family of breast pumps have power cords that may easily become damaged if not properly handled. | DESIGN: Device Design | Medela Inc., sent an " IMPORTANT FIELD ACTION " letter dated April 4, 2011, to all affected customers via The United States Postal Service ( USPS) United Parcel Service ( UPS) and E-mail .The letter identified the product, the problem, and the action to be taken by the customer. It was discovered the power adaptor/transformer used with the Pump in Style Advanced family of breast pumps can develop damage in the form of twists, kinks, bends, bulges or exposed wires in the power cord. Use of the power adaptor with a damaged cord could result in a short in the cord resulting in sparks, smoke, burning and melting of the transformers. If the users have a damaged power adaptor or power adaptor cord, they were requested to discontinue use of the adaptor and visit Medela's website at www.poweradaptorhelp.com to request a replacement adaptor. Questions about the recall process were directed to Medela Customer Service at 1-855-367-0104 or e-mail at poweradaptor@medela.com. Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelapoweradaptorhelp.com on April 4, 2011. | Contact the recalling firm for information | 478823 | 17,681 units alone, 461,142 units with products | Worldwide Distribution - Nationwide Distribution and Puerto Rico and countries of Canada, Bermuda, Cayman Islands, Hong Kong and Haiti.CHI-DO proposes phone call audit checks of 2 consumers, 2 retailers and 1 hospital account by each district. | Terminated | October 16, 2012 | 508 | 478823 | Battery
Z-2387-2011 | 58652 | enzyme immunoassay, opiates | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | DrugCheck, Manufactured by: Express Diagnostics Int;l Inc. 1550 Industrial Drive, Blue Earth MN 56013 USA, Model numbers :70801-4(NxScan OnSite, AMP500-COC-MDMA-MET500-MTD-OP1300-TCH-BUP-pH,SG-NI,CR, Lot 4-801 102910, Exp APR 2012, & Lot 4-801 112210, Exp MAY 2012, Export Only); 65500-4A (AMP-COC--OP1-TCH-MET-ALC-pH,SG-NI,CR, Lot 4A-500 110310, Exp MAY 2012, For forensic use only);30705 ( DIP DRUG TEST, AMP500/COC/MDMA/MET500/MTD/OP1300/TCH, Lot 705 112210, Exp MAY 2012);60702-6A AUS (AMP300-COC300-OP1300-OP12000-MET300-TCH50-BZ0200-ALC-pH,SG,OX-NI,CR-GL, Lot 6A-702 113010 & 6A-702 112910, Exp MAY 2012, Export Only); 60600-A (AMP-COC--OP1-TCH-MET-BZO-ALC, Lot A600 110310, Exp MAY 2012, For forensic use only);61205 (AMP-BAR-BZO-COC-MDMA-MET-MTD-OPI-OXY-PCP-TCA-THC, Lot 205 110410, Exp MAY 2012); 70701-4 (NxScan OnSite, AMP500-COC-MDMA-MET500-MTD-OP1300-TCH-pH,SG-NI,CR, Lot 4-701 102910, Exp APR 2012, Export Only); 31012-6 (DIP DRUG TEST, AMP/BAR/BZO/COC/MDMA/MTD/OPI/OXY/ PCP/THC/pH SG, OX/NI, CR, GL, Lot 6-012 110310, Exp MAY 2012);31003 (AMP/COC/OP10300/THC/MET/BZO/BAR/PCP/MTD/TCA/MET, Lot 003 102510, Exp APR 2012);Intended use: The DrugCheck Dip Drug Test is a one-step immunoassay for the qualitative detection of multiple drugs and drug metabolites in human urine. | DrugCheck | 2 | May 27, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc. has initiated a medical device recall on various model numbers of DrugCheck devices,This recall was due to incorrect product expiration dating after findings from an FDA inspection. Use of expired product may not work properly, and may give incorrect screening results. | EXPIRATION DATING: Incorrect or No Expiration Date | The firm, Express Diagnostic Int'l, Inc., sent an "Urgent Recall of DrugCheck Product" letter dated April 19, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine the product; notify Express Diagnostic of the quantity even if no product is remaining; complete and return the DISTRIBUTOR/CUSTOMER RECALL CONFIRMATION FROM via fax to 1-507-526-2252 as soon as possible. In addition, if they have further distributed the product to identify their customers and notify them of the recall. Provided procedures to be followed for them to destroy the quarantined product. Express Diagnostics will replace the DrugCheck Drug Screen devices.If you have any questions, please contact the Quality Assurance Manager at 507-526-3951, ext. 117. | Contact the recalling firm for information | 6925 | 6925 devices | Worldwide distribution: USA (nationwide) including states of: FL, IA, MD, OK, and WA; and countries including: AUSTRALIA and ITALY. | Terminated | July 06, 2012 | 406 | 6925 | Not_Computer
Z-2385-2011 | 58199 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | ***BIOMET MICROFIXATION BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***6 HOLE REGULAR DOUBLE Y PLATE***"1.5MM SYSTEM PLATE DOUBLE Y-SHAPE, REGULAR 0.6 MM TITANIUM***NON-STERILE PRODUCT***This titanium plate is used in the 1.5mm Lorenz Midface System which is indicated for use in trauma, neurosurgery, craniomaxillary reconstruction and orthognathic procedures in the midface. | 6 HOLE REGULAR DOUBLE Y PLATE | 2 | May 27, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling their 1.5 6 Hole Reg Double Y, Lots 785710, 753120, 693500 and 733600. This product is labeled as 01-7110-K (1.5 6 Hole Reg Double Y), however the package contains a 01-7112-K (1.5.6 Hold Long Double Y). | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Biomet Microfixation, sent an "URGENT MEDICAL SAFETY ALERT NOTICE" letter dated July 28, 2010 to its customers/distributors. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their stock and ensure that the plates were properly distributed into their trays and that they have the sufficient quantity of plates that they need. In addition, they were instructed to complete and return the attached Inventory Reconciliation form via fax to: Attn: Customer Relations Specialist at 904-741-9425 and notify the recipient of this notice, if they have further distributed the product.If you have any questions, please contact the Customer Relations Specialist at 800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. | Contact the recalling firm for information | 180 | 180 | Worldwide distribution: USA (nationwide) including states of: FL, IN, KY, MT, NC, NE, NY, OH, TX, VA, and WI; and countries including: Argentina, Canada, Colombia, Czech Republic, Japan, Korea, Mexico, Norway, South Africa, Spain and Turkey. | Terminated | June 24, 2011 | 28 | 180 | Not_Computer
Z-2384-2011 | 58632 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator, Model # HT-50-H, HT50-HB, HT50-H1, and HT50-H1BProduct Code: NOUThe device is intended to provide continuous or intermittent mechanical ventilator support for the care of individuals who require mechanical ventilation. Specifically, the NEWPORT HT50 is applicable for adult and pediatric (i.e. infant, child and adolescent) patients, greater than or equal to 10 kg (22Ibs), who require the following general types of ventilatory support: positive pressure ventilation with Assist/Control, SIMV and SPONT/CPAP modes of ventilation. The NEWPORT HT50 is a restricted medical device intended for use by qualified, trained personnel under the direction of a physician; it is suitable for use in hospital, sub-acute, emergency room, and home care environments, as well as for transport and emergency response applications. | Newport HT50 Ventilator | 2 | May 27, 2011 | 2011 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the HT50 Dual Pac Ventilator is not compatiable with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. The Dual Pac system has a slightly higher amperage draw than the single battery and as a result can cause the fuse in the external battery pack to blow making the battery inoperable. | OTHER/UNDETERMINED: Pending | Newport Medical sent an IMPORTANT PRODUCT NOTICE on July 22, 2009, to all affected customers who purchased the Newport HT50 Dual Pac Ventilator with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. Newport Medical also sent an upgrade kit (FSK3301) which includes instructions for changing the fuse on the external battery. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to not use an external battery with the new Dual Pac Ventilator until they have completed the fuse replacement. Customers with any questions about this fuse replacement or require additional upgrade kits, were instructed to contact Newport Medical's Technical Support Department at (800) 451-3111 or (714) 427-5811 x500. |  | 333 | 11, 333 units total | Nationwide Distribution | Terminated | June 01, 2011 | 5 | 11333 | Battery
Z-2359-2011 | 58346 | kit, i.v. start | LRS |  |  | General & Plastic Surgery | Enforcement Discretion | Material ECVC1140, TRIPLE LUMEN INSERTION KIT, CENTURION TRIPLE LUMEN INSERTION KIT, REORDER ECVC1140, STERILE, LATEX FREE, CENTURION MEDICAL PRODUCTS CORP. WILLIAMSTON MI 48895, USA 800.248.4058 www.centurionmp.comFACILITATES LINE INSERTION | CENTURION TRIPLE LUMEN INSERTION KIT | 2 | May 27, 2011 | 2011 | Centurion Medical Products Corporation | This voluntary sub-recall was initiated because the vials of Bacteriostatic Sodium Chloride Injection, USP, 0.9% utilized to make the products may contain visible particulates. |  | Centurion Medical Products Corporation issued press on April 1, 2011, and sent an Urgent Voluntary Recall letter dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the kits immediately and instructions were provided for the return of the kits.Customers were also asked to complete a form to fax or email back to the firm regarding the number of kits the firm has in their possession.Customers were instructed to provide recall notification to other firm's who received the recalled kits from their direct account.For any questions regarding this recall call 631-924-4000, ext 142. | Contact the recalling firm for information | 10 | 10 | Nationwide Distribution | Terminated | September 13, 2011 | 109 | 71456 | Not_Computer
Z-2392-2011 | 58665 | collector, urine, (and accessories) for indwelling catheter | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | URI-DRAIN Medium Male External Catheter With Two-Sided Adhesive Strap and Skin Protective WipeProduct ID: 8884732000Intended use: Collector urine from male patients (and accessories) for indwelling catheter | URIDRAIN Standard Male External Catheter | 2 | May 27, 2011 | 2011 | Covidien LP | A component of the Uri-Drain device recalled the Skin-Prep (manufactured by H&P; Industries-Triad) a protective wipe due to bacterial contamination | OTHER/UNDETERMINED: Pending | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated April 29, 2011 to its customers via US Postal Service. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the product, place in quarantine, and return any inventory. Distributors are requested to notify their customers. All customers are required to respond to the notice regardless of remaining inventory. Responses will be tracked for follow-up and effectiveness checks.For Customer Service questions regarding the recall please telephone: 1-800-962-9888 option 1 (weekdays, 8am to 6:30pm Eastern) and email to: sdfeedback@covidien.com. Should you have any medical questions concerning this recall, please call 508-261-6393. | No consumer action necessary | 5760 | 5760 units | Worldwide distribution: USA (nationwide) and countries including: Canada and Bermuda. | Open |  |  | 62352 | Not_Computer
Z-2395-2011 | 58518 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ES0101, AEM Disposable Curved Scissors, 1/2", 35cm, Sterile R, Encision Inc., Boulder, CO 80301 Cutting tissue during electrosurgical procedures | AEM Disposable Scissors | 2 | May 31, 2011 | 2011 | Encision, Inc. | Sterility may be compromised due to holes in packaging. | DESIGN: Packaging Design/Selection | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection.Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Contact the recalling firm for information | 598 | 574 Ten-Count Boxes, 24 Kits | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand | Terminated | August 09, 2011 | 70 | 2052 | Not_Computer
Z-2401-2011 | 58568 | nucleic acid amplification assay system, group b streptococcus, direct specimen test | NJR | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Xpert Infinity Software, V 4.1a.Cepheid, Sunnyvale, CA 94089system software for GeneXpert Infinity 48 | Xpert Infinity Software | 2 | May 31, 2011 | 2011 | Cepheid | A mismatch of Patient ID or Sample ID and the Host Order from the hosted LIS and its requested test results could occur when running in batch mode and cartridges are introduced in rapid succession. | DESIGN: Software Design | Cepheid sent an Corrective Action Notification letter dated April 13, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cepheid recommended customers using 4.1a software on the INFINITY 48 System that is interfaced to a LIS with Host Ordering to make these alterations in operationof the system.After scanning the Patient ID (optional). Sample ID and Cartridge barcode, the host test order information should apear highlighted in the Order Test workspace.Once the highlighted host order information appears select SUBMIT.Warning. If the host test order informaiton is not highlighted, cancel the order by selecting CLOSE. Re scan Patient ID (optional). Sample ID and cartridge barcode and verify highlighted host test order before SUBMIT. Contact Cepheid Technical Support at 888-838-3222 in the US or 33 5 63 82 53 19 outside of the US for any questions regarding this Corrective Action Notice. | No consumer action necessary | 19 | 16 US, 3 ROW | Worldwide Distribution - USA including CA, UT, TX, NJ, FL, NC, CO, IN, VT, MN and the countries of the UK, Denmark and Saudi Arabia. | Terminated | June 20, 2011 | 20 | 19 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-2403-2011 | 58636 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron UMPC is a component in the DataCaptor" Connectivity System. Catalog Number: DC-NU-MPCCapsule Neuron" UMPC is a bedside device that enables automatic collection of vital signs data that can local cache data and connect to the hospital network through a wiredEthernet connection or via a standard 802.11 b/g/n wireless network, available in a fixed or mobile mounting option, running either DataCaptor continuous or periodic connectivity softwareThe DataCaptor" System is indicated for use in data collection and clinical information management either directly or through networks with independent bedside devices. DataCaptor" is not intended for monitoring purposes, nor is the software intended to control any of the clinical devices (independent bedside devices / information systems) it is connected to | Capsule Neuron UMPC | 2 | June 01, 2011 | 2011 | Capsule Tech Inc. | Software: Data cached on the Neuron due to server unavailability is not being delivered to the server after server availability is restored.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes.Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power source | DESIGN: Software Design | Capsule Tech initiated three corrective action notifications for the Neuron" UMPC units three failure modes as follows:a) Data cached on the Capsule Neuron" UMPC due to server unavailability is not being delivered to the server after server availability is restored:i) A customer advisory was provided to a subset of consignees, along with DataCaptor software version 6.6.5 on December 1, 2010.ii) A customer advisory was provided to the remaining consignees, along with DataCaptor software version 6.6.7 on January 27, 2011.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes:i) A customer advisory was provided to all consignees on March 23, 2011, via phone call, email and FedEx;.The correction will be provided as a software update. A formal recall notification letter will be provided to all consignees alerting them to the availability of the correction once it is released, via phone call, email and FedEx. c) Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power. Note: the Capsule Neuron" UMPC operates normally when connected to power.i) A customer advisory was provided to all consignees on April 23, 2010, via FedEx.For questions regarding this recall contact Capsule technical support team at (800) 260-9537 or e-mail support@capsuletech.com.. | Discard product or return to place of purchase for a full refund or exchange | 3144 | 3,144 Capsule Neuron UMPC units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | March 20, 2012 | 293 | 3144 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2402-2011 | 58557 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cholestech LDX Lipid Profile"GLU test cassettes, Catalog Number 10-991, Lot # 245104Cholestech LDX Lipid-Glucose tests are intended for the quantitative determination of total cholesterol, HDL (high-density lipoprotein) cholesterol, triglycerides, and/or glucose in whole blood, serum, and plasma. A TC/HDL (total cholesterol/HDL cholesterol) ratio and estimated values for LDL (lowdensity lipoprotein) and non-HDL cholesterol are calculated by the Cholestech LDX. | Cholestech LDX Lipid Profile"GLU test cassettes | 2 | June 01, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated due to an issue related to an error in the encoding of calibration coefficients in the sample cassette that cause the Cholestech LDX analyzer to report glucose values as |  | Alere San Diego, Inc ( formerly known as Biosite Incorporated sent an URGENT MEDICAL DEVICE RECALL letter on February 18, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were requested to:(1) Discontinue use and/or sale of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 245104.(2) Discard in accordance with your local regulations, all units of the above listed product.(3) Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation.(4) If you have forwarded the product listed above to another laboratory, please provide a copy of this letter to them.(5) Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice, to indicate the number of boxes of test cassettes discarded from your inventory and replacement product will be sent to you.For questions about the information contained in this notification, customers were instructed to contact Alere Technical Services at 1-877-308-8289 (24 hours a day, 7 days a week) or via e-mail at service.cholestech@alere.com. |  | 5305 | 5,305 Boxes | Worldwide Distribution - USA ( nationwide ) including the countries of South America ( EL Salvador) , Europe ( Denmark, France, Norway and Switzerland), Bahrain, Kuwait, South Africa, and United Arab Emirates. | Terminated | April 02, 2012 | 306 | 5305 | Not_Computer
Z-2407-2011 | 58267 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 02, 2011 | 2011 | Computerized Medical Systems Inc | XiO Software: For a specific clinical setup, the MLC leaf positions on the exported DRR images might not match the MLC leaf positions displayed in XiO. | DESIGN: Software Design | The firm, Elekta, issued an "important Safety Notice" undated, to its customers. The notice indicates the problem and the clinical impact. There is no reasonable workaround for this issue. The notice states a future XiO Release will resolve the issue and the customers will be notified when this solution is available. A return postcard was included with the mailing to verify consignee received the notice. For sites not returning the postcard within one month, the recalling firm will contact them via phone or email to confirm receipt of the notice. The customers were instructed to distribute this notice to any and all users of the CMS software at their organization who are potentially affected by this issue.If you have any questions, please call 314-993-0003 or 800-878-4267. |  | 415 | 415 | Nationwide distribution, including Puerto Rico, to medical facilities. Distribution was also made to Military and Government consignees. Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, and United Kingdom. | Open |  |  | 415 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-2405-2011 | 58664 | bag, bile collecting | EXF | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Drainage Bag, Model 600D, 600 mL, Single use, Ethylene Oxide sterilized, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005.The Biliary Drainage Bag is used to aid with fluid drainage involving abscess, biliary and nephrostomy. Mechanical failure of this device would result in patient discomfort. This device is not life sustaining no life supporting. The medical device is not intended to be implanted. This medical device has surface contact with the patient. Improper use of this device could cause fluid back-up and result in invasive contact through the catheter. The drainage bag will be used in conjunction with a nephrostomy/biliary drainage catheter. | Drainage Bag | 3 | June 02, 2011 | 2011 | Remington Medical Inc. | There have been complaints of the Drainage Bags leaking from the outlet and the inlet port. | OTHER/UNDETERMINED: Pending | Remington Medical Inc.sent an Urgent Medical Device Recall letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine product subjec to recall. If they further distributed this product, they were to identify their customers and notify them at once of the product recall. Notification to their customers should include a copy of the recall notification letter. Customer were to receive a Return Authorization Form, reference RGA #600-D.For questions regarding this recall call 800-989-0057, ext. 213. |  | 89294 | 89294 units | Nationwide Distribution | Terminated | September 13, 2011 | 103 | 89294 | Not_Computer
Z-2406-2011 | 58619 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 Ventilator. Model #HT70, The Newport HT70 ventilator is intended to provide continuous or intermittent positive pressure mechanical ventilatory support for the care of individuals who require mechanical ventilation through invasive or noninvasive interfaces. Specifically, the Newport HT70 ventilator is applicable for infant, pediatric and adult patients greater than or equal to 5 kg (11 Ibs). The Newport HT70 ventilator is suitable for use in hospital, subacute, emergency room, and home care environments as well as for transport and emergency response applications. | Newport HT70 Ventilator | 2 | June 02, 2011 | 2011 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed a limited number of power cords recently received from their supplier, Glob-Tek, have a slightly oversized plug connector which prevents thecord from locking into place. These cords are fully functional as supplied. However, the cord can be easily disconnected. If the cord is disconnected, the HT70 will alert the user with a power di | OTHER/UNDETERMINED: Pending | Newport Medical sent an PRODUCT NOTICE LETTER, dated September 20, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to discard the original cord and only use the replacement cord to avoid unintentional disconnection of the power cord. Each affected Customer was sent a replacement cord.Customers with any questions regarding the Product Notice were instructed to contact the Technical Support Department at (800) 451-3111. |  | 22 | 22 units total | Worldwide Distribution--USA (nationwide) including the states of MA, PA, OH, IA and Puerto Rico and the countries of Colombia, UAE, Hong Kong, Hungary, India, Japan, Nepal, China, Philippines, Singapore and South Africa. | Terminated | June 03, 2011 | 1 | 22 | Battery
Z-2404-2011 | 58637 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are:ARTISTE MV System;ONCOR Avant GardeONCOR ImpressionONCOR Impression PlusONCOR Expression.Manufactured by Siemens AG, Medical Solutions, Kemnath, Germany;Registered by Siemens Medical Solutions, Concord, CA.ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC | 2 | June 02, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment. | OTHER/UNDETERMINED: Pending | Siemens sent a Consumer Safety Advisory and Urgent Medical Device Correction letter to all affected customers beginning on April 15, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Updated instructions for using the Motion Stop button with the 160-MLC were provided with the letter. Customers were requested to use the Motion Stop button and the corresponding reset procedure only when required. In the unlikely event that the Motion Stop button needs to be used repeatedly during treatment of the same beam or segment, customers were advised to perform an "in session resumption of the beam by selecting 'abort' at the syngo RT Therapist and then re-selecting the beam again for the remaining delivery. In this case, the beam will be re-loaded with the originally prescribed leaf positions. Please consider the already delivered dose for resumption.Customers were asked to include the Customer Advisory Notice in their LINAC System Owner Manual in chapter 'Safety Advisory Letters' where it should remain until the measures defined above were performed. Customers were asked to inform all affected personnel immediately. | Contact the recalling firm for information | 221 | 221 active devices | Worldwide Distribution - USA including CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV. | Terminated | March 16, 2012 | 288 | 221 | Not_Computer
Z-2409-2011 | 58681 | catheter, straight | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SureCath Set with Accessories, Coloplast Corp.-MN 06093, Sterile EO., Manufacturer Coloplast -A/S 3050 Denmark.28022S 1400SureCath Set,acc CH8 Male,bulk28023S 1400Surecath Set,acc CH10 Male,bul28024S 1400Surecath Set,acc CH12 Male,bul280320 1400SureCath Set,acc CH14 Male 28032S 1400Surecath Set,acc CH14 Male,bul28036S 1400Surecath Set,acc CH14/1,2 Male28037S 1400Surecath Set,acc CH16 Male 280370 1400Surecath Set,acc CH16 Male 28033S 1400Surecath Set,acc CH16 Male280360 1400Surecath Set,acc CH14/1,2 Male 280340 1400SURECATH SET W.ACC FEM14`The SureCath Set is indicated for use by patients for intermittent catherization for the purpose of bladder drainage. The insertion supplies provided are to be used to insert the catheter using sterile technique. | SureCath Set with Accessories | 2 | June 03, 2011 | 2011 | Coloplast Manufacturing US, LLC | Coloplast has initiated a recall on SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies because they contain Iodine prep pads which may expire before the overall product expiration date. SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies are both packaged with a set of accessories which includes Iodine prep pads. The Iodine prep pads included i | EXPIRATION DATING: Incorrect or No Expiration Date | Coloplast Manufacturing US, LLC sent a Coloplast "Product Field Correction Notice" dated April 21, 2011. The letter was addressed to "Dear Valued Customer". The letter described the problem and the product involved in the recall. Customers were advised to provide the current inventory levels of the lots affected and the firm will be providing them with new iodine prep pads with sufficient shelf life. This product will include customer instructions, informing the customer to discard the original prep pads and use the new prep pads instead. If they prefer, they could themselves send the affected product back to Coloplast and they in return will attach new prep pads to the product for them.If they have further distributed the product, they advised consignees to contact their customers and to include the letter provided to them to inform them of the recall. | Contact the recalling firm for information | 438412 | 438,412 | Nationwide Distribution including AL, AZ, CA, CO, CT, DE, FL,GA, ID, IL, IA, MD, MA, MI, MN, MS, NJ, NY, NC, OH, OK, OR, PA,PR, SC, TN, TX VA, WA, WI. | Terminated | April 24, 2012 | 326 | 564759 | Not_Computer
Z-2411-2011 | 58785 | bed, hydraulic, adjustable hospital | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eleganza 3 Hospital BedIntended use: AC Powered adjustable Hospital Bed | Eleganza 3 Hospital Bed | 2 | June 03, 2011 | 2011 | Linet Americas Inc | Linet has been made aware that there has been breakage of the Mobi-lift handle on a few beds which may cause failure during patient assist use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Linet America, notified US customers only via phone on May 5, 2011. A follow up "service advisory information" letter dated May 11, 2011 was sent to all customers. The letter described the product, problem and action to be taken. The customers were instructed to remind/retrain staff to ensure that the Mobi-lift handles are always in the stowed position during patient transport and after use, and complete and return the attached Inspection procedure for each affected Eleganza 3 bed via mail to: Linet Americas, 10420-R Harris Oaks Blvd., Charlotte, NC 28269 or fax to: 704-248-5655. Linet Americas will be receiving a supply of new Mobi-lift design and will begin the process of executing service calls to replace the Morbi-lift handles with the new design.If you have any questions, contact the Director of Quality and Service Operations at 704-248-5650 or 980-406-6446. | Contact the recalling firm for information | 1601 | 1601 units | Worldwide distribution: USA (nationwide) including states of: KY, MA, NC, and NY; and country of: Canada. | Terminated | July 19, 2012 | 412 | 1601 | Not_Computer
Z-2408-2011 | 58205 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.0.2 and above.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 03, 2011 | 2011 | Computerized Medical Systems Inc | When RT Plans are exported from XiO and RT Structure Sets and Images are exported from Monaco, the DICOM iso-center coordinates may not be accurate. | DESIGN: Software Design | Elekta / Computerized Medical Systems Inc began mailing an Important Safety Notice to all affected customers in September 2010. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised not to mix treatment plan information from multiple planning systems. Customers were asked to return the postcard as proof of receipt and their understanding of the Safety Notice. For questions regarding this recall call 408-830-8000. |  | 1377 | 1377 | Worldwide Distribution - USA (nationwide and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1377 | Other |  |  | Other
Z-2412-2011 | 58735 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DuraspanTM Dialysis Catheters, Ultra and Long Term, 15.5 Fr, Sterile EO, Manufactured for r4 Vascular, Inc . 7550 Meridan Circle North, Suite 150, Maple Grove, MN 55369. Description Model #Duraspan Kit - 24cm 62152420Duraspan Kit - 28cm 62152820Duraspan Kit - 32cm 62153220Duraspan Exchange Kit - 32cm62153228Duraspan Kit - 36cm 62153620Duraspan Exchange Kit - 36cm62153628Duraspan Kit - 40cm 62154020Duraspan Exchange Kit - 40cm62154028Duraspan Kit - 55cm 62155520Duraspan Ultra Kit - 24cm621524E0Duraspan Ultra Kit - 28cm 621528E0Duraspan Ultra Kit - 36cm 621536E0Duraspan Ultra Kit - 40cm621540E0Duraspan Ultra Kit - 55cm621555E0The r4 DuraspanTM and the DuraspanTM Ultra long-term dialysis catheter are indicated for attaining short and long term ( | DuraspanTM Dialysis Catheters, Ultra and Long Term | 2 | June 06, 2011 | 2011 | R4 Vascular Inc | r4 Vascular, Inc. is conducting a voluntary recall of specific lots of Duraspan Dialysis Catheters. Duraspan catheters are 15.5 Fr dual lumen chronic hemodialysis catheters. Over a six month period, seven patients at five hospitals experienced separation of the luer connector from the extension tube on the Duraspan catheter. Testing during failure analysis disclosed that these lots and other lot | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | R4 Vascular Inc sent an Urgent Medical Device Recall letter dated May 3, 2011, to all affected customers. The letter described the problem and the product being recalled. Consignees were instructed to stop distributing the product and to quarantine the recalled lots and if they had already distributed the affected product to seek its return from their customers. A template letter to physicians and healthcare facilities was enclosed for Distributors/Retailers to send or hand deliver to their customers.For questions regarding this recall call 1-704-905-2105. | No consumer action necessary | 2489 | 2489 kits | Nationwide Distribution including CA, CO, CT, FL, GA, IL, LA, MD, MA, MN, MS, MO, NJ, NC, OH, PA, SC, TX, and WA. | Terminated | July 27, 2012 | 417 | 2489 | Not_Computer
Z-2415-2011 | 58677 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Pinnacle Destination Guiding Sheath, 6 French, 90 cm, Straight Tip, Tuohy-Borst Valve, Lot MD24, MD31, ME07, & ME14. 510(k) K012812Guiding Sheath is used to facilitate the introduction of interventional and diagnostic devices into the human vasculature, including but not limited to the renal arteries for the Renal Guiding Sheath and the carotid arteries for the Carotid Guiding Sheath. | Pinnacle Destination Guiding Sheath, 6 French, 90 cm | 2 | June 06, 2011 | 2011 | Terumo Medical Corporation | Molding defect in destination product that render the device unusable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo Medical Corporation sent an Urgent Product Bulletin Recall letter dated June 24, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo requested customers to immediately discontinue use, and quarantine any available stock of the listed product codes. In order to facilitate the exchange of these devices, a Terumo Territory Manager contacted all customers by July 9th to confirm receipt of the Terumo Destination Bulletin and scheduled an appointment to visit each customer in person by July 31st to personally manage the exchange. For questions or for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 214 | 214 units | Worldwide Distribution - USA including AL, AK, AZ, TX, LA, MS, MN, MO, CA, GA, FL, HI, KS, IA, ID, IL, TN, NM, OR, OK, WA, and WI and the countries of Belgium, Brazil, and Panama | Terminated | July 07, 2011 | 31 | 258 | Not_Computer
Z-2413-2011 | 58591 | hepatitis a test (antibody and igm antibody) | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ARCHITECT HAVAB-M Assay Reagent; a chemiluminescent microparticle immunoassay (CMIA) for the qualitative detection of IgM antibody to hepatitis A virus (IgM anti-HAV) in human adult and pediatric serum and plasma (dipotassium EDTA, lithium heparin, and sodium heparin) and neonatal serum; Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA, List 6L21-25 | ARCHITECT HAVABM Assay Reagent | 2 | June 06, 2011 | 2011 | Abbott Laboratories | The HAVAB-M reagent lot has reduced specificity and may cause an increase in false grayzone and/or false reactive results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Laboratories issued a Product Recall letters dated April 20, 2011 to all affected customers that received ARCHITECT HAVAB-M reagent, lot 93794HN00. The letters informed the accounts of the increased grayzone and/or reactive results and were instructed to discontinue use of the lot and destroy any remaining inventory of the lot according to laboratory procedures. If the account did not have an alternate lot of reagent, they were told to continue using the recalled lot until their lab receives the replacement lot, and to perform supplemental testing with another HAVAB-M method to confirm the results for all specimens that are grayzone or reactive. The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. For further questions regarding this please contact Abbott at 1-877-422-688. | Contact the recalling firm for information | 788 | 788 kits | Nationwide Distribution including Puerto Rico | Terminated | February 10, 2012 | 249 | 788 | Not_Computer
Z-2420-2011 | 58658 | apparatus, autotransfusion | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics OrthoPAT -Orthopedic Perioperative Autotransfusion SystemUS Only: 1050-110-ENGOutside US:1050-110-ENG, I050-240-XXX (generic), 1050-240-ENG, 1050-240-SPN, 1050-240-ITL, 1050-240GER,1050-240-DUT, 1050-240-FRN, 1050-240-CZE, 1050-IIO-JPN.Provides perioperative autotransfusion ( salvaging and reinfusing red blood cells (RBC's) from blood shed from patients undergoing orthopedic surgery. The portable system is used in operating rooms to recycle blood lost during surgery and in the recovery room to recycle blood lost after surgery. It can be transported to the patient floor for postoperative salvage. | ORTHOPEDIC PERIOPERATIVE AUTOTRANSFUSION | 2 | June 07, 2011 | 2011 | Haemonetics Corporation | 2 Models Model 1:OrthoPAT devices with no centrifuge drain spill containment system may have potential for fluid contamination leading to potential spark or smoke may occur are recalledModel 2:OrthoPat equipped with a spill collection drainage system updated with revised Instructions for Use and Quick reference Guide Addendum | DESIGN: Device Design | The firm, Haemonetics, sent an "Urgent Field Recall" letter dated April 29, 2011 to all customers. The letter describes the product, problem and actions to be taken. This recall covers the 2 models. Model 1: OrthoPAT devices without drain ports (Serial numbers 00001 through 02279-T). All OrthoPAT devices without drain ports worldwide, will be removed from distribution and scrapped. Model 2: OrthoPAT devices with drain ports (serial numbers beginning with 02280-T and higher). All OrthoPAT devices with drain ports will remain in clinical use and will receive updated labeling.The customers were instructed to follow the enclosed directions to confirm their possession of any units of this model and arrange for their return to Stericycle; immediately cease use of devices not equipped with a spill collection drainage system ( to determine whether their OrthoPAT device is equipped with the spill collection drainage system, please refer to the attached illustrations/instructions); complete and return the IMPORTANT-CUSTOMER CONFIRMATION sheet via fax to: Stericycle at 1-866-792-5450, and RETURN the OrthoPAT devices NOT equipped with spill collection drainage system (for instructions on how to arrange return of OrthoPAT Devices, call a Stericycle Representative at 1-866-918-8736). A Haemonetics representative will be in contact with the customers facility within 14 working days to assist with the recall actions.The firm, Haemonetics, will update the OrthoPAT labeling. The labeling will have specific instructions to permanently deploy the centrifuge spill collection bag at all times. Device cleaning instructions in the event of a blood spill are also clarified and include specific instructions to unplug the device from the AC outlet in the event of a fluid spill on the device. If a device does experience a fluid spill, new instructions direct the user to return the device to Haemonetics for evaluation and repair. Operator Manuals will be replaced with a new revisi. | Contact the recalling firm for information | 2976 | 2976 units | Worldwide distribution: USA (nationwide) and countries including: Austria, Australia, Belgium, Switzerland, Canada, China, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Guatemala, Ireland, Israel, India, Italy, Japan, Lithuania, Luxembourg, Latvia, Monaco, Netherlands, Norway, Panama, Poland, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, Slovak Republic, EI Salvador, Turkey and Taiwan., | Open |  |  | 2976 | Not_Computer
Z-2467-2011 | 58703 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Hitachi HIVISION Preirus/Avius Ultrasound equipped with V03.02 Software - Doppler Trace Measurement program. 510 (k) Numbers: Preirus - K093466 (6/1/10), Avius - K102901 (10/27/10).The Preirus and Avius are intended for use by trained personnel (doctors, sonographers, etc) for the diagnostic ultrasound evaluation of Abdominal, Cardiac, Intra-operative, Feta, Pediatric, Small Organ, Peripheral vessel, Biopsy, Trans-rectal, Trans-vaginal, Musculoskeletal, Neonatal Cephalic, Adult Cephalic, Endoscopy, Intra-luminal, Genecology, Urology and Laparoscopic clinic applications. The Modes of Operation of the HIVISION Preirus are B mode, M mode PW mode (Pulsed Wave Doppler), CW mode (Continuous Wave Doppler) Color Doppler, Amplitude Doppler (Color Flow Angiography), TDI (Tissue Doppler Imaging), 3D Imaging, 4D Imaging, RealTime Tissue Elastography, and Real Time Virtual Sonography. | HIVISION Preirus, Avius ultrasound systems | 2 | June 08, 2011 | 2011 | Hitachi Medical Systems America Inc | Software anomaly. The Doppler Trace Measurement feature on certain models of the firm's diagnostic ultrasound systems, may, under certain conditions whereby multiple measurements are done sequentially, record incorrect values. The recorded value is displayed to the user on a report; however depending upon the measurement area, displayed velocity values could be lower than at the true velocity lead | DESIGN: Software Design | The firm, Hitachi, sent a "DEVICE CORRECTION NOTICE" letter dated April 15, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were inform of the nature of the software error in the Doppler Trace Measurement feature; the circumstances under which this error can occur; and the safety issue involved should incorrect values generated lead to a false diagnosis of stenosis in a patient. The letter provides the customer with a workaround solution in order to avoid this issue. The customer were also instructed to call the HMSA Applications Helpline at 877-871-2777 to reprogram their system to temporarily eliminate the noted safted issue and to immediately have all selections unchecked on the Doppler trace measurement result store setting set up screen. The notification letter also states that Hitachi Medical Systems America (HMSA) will be having their service personnel schedule a visit at each customer location and install a software upgrade on each affected system unit in order to directly address, and permanently resolve this issue. If you need any further information regarding this notice, please contact the Manager of Regulatory Affairs at 800-800-3106 x3720 or email to: thistlewaited@hitachimed.com. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 24 | 24 system units | Nationwide distribution: USA including states of: CA, MD, KY, FL, MN, NJ, IA, OH, TX, IL, and IA. | Terminated | February 19, 2013 | 622 | 24 | Software | Output/Calculation | Software update | Software Update
Z-2468-2011 | 58620 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Model Number: e360The e360 Ventilator System is intended to provide continuous (endotracheal or tracheostomy [ET] tube) or non-continuous (mask) ventilatory support and monitoring for infant, pediatric, and adult patients requiring tidal volumes equal to or greater than 20 milliliters (mL). The device is for use by prescription only. The intended environments include hospital, hospital-type, and intra-hospital transport environments. Hospital use typically includes general care floors, operating rooms, special procedure areas, emergency rooms, and intensive and critical care areas within the hospital. Hospital-type use includes facilities such as or similar to surgicenters, sub-acute centers, and special nursing facilities outside of the hospital. Intra-hospital transport includes patient transport within the hospital or hospital-type facility. | Newport e360 Ventilator | 2 | June 08, 2011 | 2011 | Newport Medical Instruments Inc | Cregarding e360 ventilators that were shipped with a different software version from what the customers ordered. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Newport Medical Instruments Inc. initiated a field correction on October 27, 2009, to address complaints received regarding e360 ventilators that were shipped with a different software version from what the customers ordered. Newport Medical Instruments provided the correct software to the Newport sales or service providers who were responsible for the units affected. They installed the correct software into the ventilators at the customer sites. No written Correction Notice was provided since the corrections were done at the customers' request.Should you have any questions or require additional information, please call (714) 427-5811 extension 344. | No consumer action necessary | 18 | 18 | Worldwide Distribution--USA (nationwide) including states of CA, KY, and Silver City, New Mexico and country of Quito, Ecuador. | Terminated | June 08, 2011 | 0 | 18 | Software |  | Software update | Software Update
Z-2423-2011 | 58663 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | Ringloc + Acetabular Shell limited hole finned 50 mm Size 23 porous coat, for uncemented use, sterile, REF 16-104150, Biomet Orthopedics, Warsaw, IN.Intended Use: The Porous Plasma Spray (PPS) RinglocÂŽ Acetabular System is intended to replace the natural acetabulurn damaged by disease, trauma or revision of previous arthroplasty. Indications For Use: 1 . Non-inflamnmatory degenerative joint disease including osteoarthritis and avascular necrosis. 2. Rheumatoid arthritis 3. Correction of functional deformity 4. Treatment of non-union, femoral neck fracture, and trochanteric fractures of the proximal femur with head involvement, unmanageable by other techniques. 5. Revision procedures where other treatment or devices have failed. | Porous Plasma Spray (PPS) Ringloc Acetabular System | 2 | June 08, 2011 | 2011 | Biomet, Inc. | The firm has initiated this action following an investigation which found that the locking ring may be assembled incorrectly in the acetabular shell. If the locking ring is incorrectly assembled in the acetabular shell, the liner cannot be completely seated in the shell. | TRAINING: Employee Error | Biomet Orthopedics sent an URGENT MEDICAL DEVICE RECALL NOTICE DATED April 14, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Immediately locate and remove the affected product.Follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. Customers were to confirm receipt of the notice by calling 800-348-9500, ext 3755 or 3756. For any questions regarding this recall call 574-371-3755 or 574-371-3756. | Contact the recalling firm for information | 762 | 762 all products | Worldwide Distribution - USA (nationwide) and the countries of Australia, Netherlands and Japan. | Terminated | September 16, 2011 | 100 | 2069 | Not_Computer
Z-2541-2011 | 58571 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-modular reverse shoulder humeral tray with lock ring, 44 mm 10 degree angle, sterile, single use, REF 11-113685, Biomet Orthopedics, Warsaw, IN.Biomet Comprehensive Reverse Shoulder products are indicated for use in patients whose shoulder joint has a grossly deficient rotator cuff with severe arthropathy and/or previously failed shoulder joint replacement with a grossly deficient rotator cuff. The patient must be anatomically and structurally suited to receive the implants and a functional deltoid muscle is necessary. The Comprehensive Reverse Shoulder is indicated for primary, fracture, or revision total shoulder replacement for the relief of pain and significant disability due to gross rotator cuff deficiency. | Shoulder humeral tray | 2 | June 09, 2011 | 2011 | Biomet, Inc. | The firm is initiating this recall following an investigation which identified the rare possibility that certain units of Comprehensive Reverse Shoulder Humeral Trays may contain a locking ring that is incorrectly assembled. | TRAINING: Employee Error | The firm, Biomet Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 22, 2011 to their customers. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately locate the recalled product; discontinue use; return the product back to Biomet; confirm receipt of the notice by calling 800-348-9500 extension 1368 or 1325, and complete and return the attached form via fax to: 574-372-1683 or mail to: Biomet, 56 East Bell Drive, Warsaw, IN 46580.Customers with questions regarding this recall should call 800-348-9500 extension 1368 or 1325. |  | 1575 | 1,575 units all products | Worldwide distribution: USA (nationwide) including: Puerto Rico and countries including: Australia, Canada, Chile, Columbia, Costa Rica, Korea and Netherlands. | Terminated | February 21, 2013 | 623 | 1575 | Not_Computer
Z-2552-2011 | 58263 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV (Cardio Vascular)The IMPAX Cardiovascular suite is a cardiovascular information system, providing archiving, image display and modality/study specific structured reporting. | Cardio Review Station (CRS) | 2 | June 10, 2011 | 2011 | AGFA Corp. | Customer error with the angiography (XA: X-Ray Angiography) measurement calibration functionality within the Cardiovascular Review Station resulted in improper calibration. | OTHER/UNDETERMINED: Pending | The firm, AGFA Healthcare, sent an "URGENT SAFETY NOTICE" letter dated March 21, 2011 via FED-EX to its customers.. The letter describes product, problem, and actions to be taken. The customers were instructed to distribute this information within their facility to all individuals who need to be aware and to complete and return the URGENT SAFETY NOTICE FEEDBACK FORM via fax to: 864-421-1664 or email. AGFA is supplying the customers with a CRS User Manual Addendum to emphasize the proper steps required to achieve an accurate measurement when using the calibration tool within CRS. Mitigation will provide the customer/end-user with awareness of the correct XA measurement calibration technique and the potential consequences if proper technique is not followed. Additionally, AGFA will offer customer training on the use of the calibration tool. This training will be provided remotely via the web at no cost. Customers were ask to register before May 1, 2011 by contacting AGFA Customer Support Center toll-free at 877-777-2432. Training sessions are from May-June 2011.If you have any questions about this matter, please contact your local AGFA HealthCare organization or the AGFA Customer Support Center at 877-777-2432. |  | 536 | 536 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | November 29, 2011 | 172 | 536 | Not_Computer
Z-2419-2011 | 58405 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Atlantis ÂŽ SR Pro 2, 40 MHz Coronary Imaging Catheter, 3.6F (1.18mm) x 135 cm, USE BY 2011-05, REF 39014, UPN H749390140, Sterile R Sterilized using irradiation, Made in USA 47900 Bayside Parkway, Fremont, CA 94538The intended use as stated in Japan Shonin 21200BZY00484000: This product is an imaging catheter with built-in ultrasonic transducer that sends ultrasound waves inside the blood vessel to perform B-mode ultrasonagraphy. The indications for use as stated in the 510(k) and DFU 90606867: The Atlantis coronary catheters are intended for ultrasound examination of coronary intravascular pathology only. Intravascular Ultrasound imaging is indicated in patients who are candidates for transluminal coronary interventional procedures. | iCross Coronary Imaging Catheter | 1 | June 10, 2011 | 2011 | Boston Scientific Corporation | Catheter tip detachment that is not associated with elongation of the catheter typically caused by excessive pulling forces. EXPANDED: Boston Scientific is initiating a recall of all iCross Coronary Imaging Catheters because they have received and confirmed 8 occurrences of catheter tip detachments due to embrittlement. | DESIGN: Process Design | Boston Scientific " Important Customer Notification" letters dated March 14, 2011 were hand delivered to customers by Boston Scientific sales representatives starting on March 28th, 2011. The sales representatives verbally reviewed information. A tracking log will be maintained document completion of the customer notifications.The letter described the product, problem and possible root cause as well as one corrective action, a change in sterilization method from Gamma to E-beam implementation.Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter dated May 26, 2011, to all affected customers. Customers were instructed to Identify and discontinue use of any affected product within their inventory and to segregate it immediately and return it to Boston Scientific. | Contact the recalling firm for information | 30669 | 30,669 | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Japan, Trinidad, Tobago, and American Virgin Islands. | Terminated | July 30, 2012 | 416 | 96702 | Not_Computer
Z-2551-2011 | 58736 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte@ Sodium Bicarbonate Liquid Concentrate, Dialysate Concentrate for Hemodialysis LiquidPart number: 08-4000-LBFormulated to be used with a three stream hemodialysis machine which is calibrated for acid and bicarbonate concentrates | NaturaLyte Sodium Bicarbonate Liquid Concentrate | 3 | June 10, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Mislabeled: incorrect lot number printed on case labels | TRAINING: Employee Error | Fresenius sent an CUSTOMER NOTIFICATION LETTER letter dated April 18, 2011, to all affected customers via Certified Mail, Return Receipt Requested. Fresenius also contacted affected customers by telephone April 18, 2011, The letter identified the product, the problem and the action needed to be taken by customer. Customers were instructed to examine their stock to determine whether they have the incorrect case label and are asked to remove thecorrectly labeled product from the incorrectly labeled case and discard the case.Acknowledgement forms were Fax-back forms for clinics and mail-back forms (including self-addressed, stamped envelopes), for home patients.Question or concerns should be directed to Fresenius Customer Service Team at 1- 800.323.5188. |  | 3190 | 3,190 cases | Nationwide Distribution--AL, CA, CT, FL, GA, NJ, NV, NY, OR, PA, and WA | Terminated | September 12, 2012 | 460 | 3190 | Not_Computer
Z-2553-2011 | 58730 | shield, nipple | HFS | Nipple shield. | General Hospital | General Hospital | 510(K) Exempt | Contact Nipple Shield (Sterile) - 24mm (Standard); Individually wrapped silicone contact nipple shields for immediate use for latch-on difficulties while breastfeeding the baby under professional care only; Manufactured for Medela, Inc., P.O. Box 660, McHenry, IL 60051-0660, Made in Switzerland; article #67203S Individually wrapped for immediate use. For latch-on difficulties. Made of silicone. For use under professional care only while breastfeeding the baby. | Contact Nipple Shield (Sterile) 24mm (Standard) | 2 | June 10, 2011 | 2011 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | PRODUCTION CONTROLS: Packaging | The firm, Medela, sent an "Important Device Recall" letter dated May 9, 2011 to all of their customers/direct accounts. Medala also began telephoning or visiting the direct accounts/customers on May 18, 2011. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) not to use the affected lots and place them in quarantine2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form inlcude with this letter4) complete and return the Product Inventory Form via fax to 800-995-7867 or email jim.hughes@medela.com, even if they have no affected product on hand5)If you have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. Each affected product will be replaced with (2) non-sterile nipple shields.5) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. Questions about the recall process were directed to Medela Customer Service at 1-888-835-5968 or e-mail at medelareturns@medela.com. Distributor/retailers were requested to notify their customers of the recall, using the enclosed customer letter. |  | 4320 | 4,320 units | Nationwide distribution: USA including states of: AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, MA, MD, ME,MI, MN, MS, MT, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TX, UT, VA, VT, WA, WI, and WV. | Terminated | February 16, 2012 | 251 | 4320 | Not_Computer
Z-2388-2011 | 58794 | pad, alcohol, device disinfectant | LKB |  |  | General Hospital | 510(k) | Churchill Medical SystemsAMS-7080CP (Dressing Change Kit).The device intended use is a dressing change kit. | Churchill Medical Systems | 1 | June 11, 2011 | 2011 | Vygon Corporation | Skin-Prep wipes purchased by Churchill (a Vygon company) for inclusion into Dressing Kits and PICC Insertion Kits were recalled by the manufacturer due to potential bacterial contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A recall letter printed on Vygon letterhead dated April 20, 2011 and specific to each customer was sent via overnight courier mail. The letter described the affected product, issue and actions to take. Customers were instructed to check their inventory for the recalled products, cease use and distribution, and quarantine all affected product. Each customer received a custom copy of the letter with the specific product codes and lots that were shipped to them described in the Recall Acknowledgement and Inventory Return Form. Distributors were provided with instructions on how to handle product that was further distributed by them. The firm issued an updated Recall Notice dated May 31, 2011 to two consignees effected by the expanded recall.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@vygonus.com. |  | 7900 | 7,900 units | United States (FL, MA, MD, and VA). | Terminated | June 18, 2012 | 373 | 38579 | Not_Computer
Z-2414-2011 | 58661 | incubator, neonatal | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Datascope/ Maquet IABPs. System 98/98XT Intra-Aortic Balloon Pump; CS100/CS100i Intra-Aortic Balloon Pump; CS300 Intra-Aortic Balloon Pump; Manufactured by Datascope Corp., 1300 MacArthur Blvd., Mahwah, NJ 07430The Intra-Aortic Balloon Pump (IABP) is an electromechanical system used to inflate and deflate intra-aortic balloon. It provides temporary support to the left ventricle via the principle of counterpulsion. The IABP is placed in the descending aorta, via the principe of counterpulsation. The IABP is placed in the descending aorta, just distal to the left subclavian artery. Once the balloon is positioned, the IABP is adjusted to trigger in synchrony with the ECG or arterial pressure waveform to ensure that inflation and deflation occur at the appropriate points during the cardiac cycle. The CS300 has the additional capability to automatically calibrate a fiber-optic pressure sensor placed in the tip of the intra-aortic balloon. The IABP will recalibrate every 2 hours or sooner should the patient or environmental conditions change. The target populations are adult and pediatric. The IABP is intended for use in the health care facility setting. | System 98/98XT, CS100/CS100i, CS300 IntraAortic Balloon Pump | 1 | June 13, 2011 | 2011 | Maquet Datascope Corp - Cardiac Assist Division | Several models of the Intra-Aortic Balloon Pumps are being recalled due to a potential problem with the fan's assembly which may cause the power supply to detect an overheating event and shut down. | DESIGN: Device Design | Maquet sent an Urgent Product Field Correction letter dated March 16, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to note the WARNING in the IABP Operating Instruction. Customers were advised to refer to the IAB Instructions for Use, Section D (Manually inflating and Deflating a Catheter). Section D of the IAB instruction for Use reiterates that an IAB should not remain inactive for more than 30 minutes and also provides detailed instruction on how to keep an IAB active during IABP failure. Customers were advised that their Service Representatiave would arrange to replace their supply which would contain a new fan assembly. Upon completion of the power supply replacement, customers were asked to sign a service repair order to verify satisfactory completion of the work. For technical questions contact Technical Support Department at 1-800-777-4222 and press 4.For scheduling a service visit by a Field Service Representatiave, contact the firm's US Call Center at 1-800-777-4222 and press 3.Maquet issued a press release on June 6, 2011 advising consumers of the issue and that their recall was considered a class one. |  | 885 | 885 units | Worldwide Distribution | Terminated | October 25, 2012 | 500 | 885 | Not_Computer
Z-2557-2011 | 58455 | joint, hip, external limb component | ISL | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Accolade TMZF Plus Hip Stem:Catalog # 3020-0537 Accolate d TMZF Hip Stem # 5Catalog #6020-3535 Accolade 132 Size 3.5 Catalog # 6021-0230 Accolade Plus TMZF Hip Stem #2 Catalog # 6021-0335 Accolade Plus TMZF Hip Stem #3 Catalog #3021-0435 Accolade TMZF Hip Stem #4 Catalog # 6021-0435 Accolade Plus TMZF Hip Stem #4 Catalog #6021-2530 Accolade (127 deg) Size 2.5 Accolade (127 deg) Size 2.5 Catalog # 6021-4535 Accolade (127 DEG) Size 4.5 K994366; k023105; K032300Manufactured by " Stryker Ireland, Carrigtwohill Industrial Estate Carrington County Cork, IrelandThe subject hip stem is a single-use device intended for the reconstruction of the head and neck of the femoral joint. This hip stem is intended for primary reconstruction of the proximal femur or revision total hip arthroplasty. | Accolade TMZF Plus Hip Stem | 2 | June 13, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Accolade TMZF Plus Hip Stems are being recalled due to grit blast media which was observed in the drive hole. | TRAINING: Employee Error | Stryker Orthopaedics sent an Urgent Product Recall letter dated March 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069.Customers were also instructed to:Contact the hospitals in their territory that have the affected product to arrange return of the product.Return the affected product via and Inter-Org Transfer to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey 07430. Attach the fluorescent orange PRODUCT REMEDIATION sticker to their return, indicating the Product Remediation # RA 2009-459. Mark the outer box with the words "Product Recall."For any questions call (201) 972-2100. |  | 44 | total of all units is 44 | Nationwide Distribution | Terminated | October 10, 2012 | 485 | 44 | Not_Computer
Z-2561-2011 | 57760 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. The Monaco RTP system is used to create treatment plans for any cancer patient for whom external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. | 2 | June 14, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: For VMAT plans, the measured dose of radiation delivered is lower than the software calculated dose by about 5% in Monaco Release 2.03.00 and 2.04.00. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent a "USER NOTICE" on January 13, 2011, to all affected users. This notice identified the product, the problem, and the action needed to be taken by the customer. A return postcard is included for the customer to confirm receipt of the user notification. The firm intends for correction in the dose calulation engine to appear in the Monaco Release 3.00 in March 2011. Any further questions please call (403) 830-8023. |  | 94 | 94 | Worldwide Distribution-- USA (nationwide) and countries of Austria, Australia, Belarus, Canada, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, Poland, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand and UK. | Open |  |  | 94 | Software | Output/Calculation | Software update | Software Update
Z-2562-2011 | 57845 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning Workstation. Focal Release 4.51.00 and above.Allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data form a Windows-based PC and transfer the information back to the RTP system. It is intended for use in radiation treatment planning using generally accepted contouring methods. | Focal Radiation Treatment Planning Workstation | 2 | June 14, 2011 | 2011 | Computerized Medical Systems Inc | Focal: If a plan involving Dynamic MLC Delivery is sent from XiO to Focal and called up there in Plan Review, the angle of the gantry is set to zero for all beams. If this plan is then sent from Focal to the R&V; System, the gantry angles all remain at zero which could lead to patient mistreatment. | DESIGN: Software Design | Elekta CMS Software sent a USER NOTICE dated November 5, 2010, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends to correct the software so it does not modify anything about the plan sent from XiO.Customers were instructed to contact their local customer support office with any questions. |  | 155 | 155 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Belarus, Belgium, Brazil, Canada, Chile, China, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Latvia, Libya, Lithuania, Mexico, Morocco, Netherlands, New Zealand, Philippines, Poland, Romania, Singapore, Slovenia, South Africa, Spain, Suriname, Switzerland, Taiwan, Turkey, Turkmenistan, and the United Kingdom. | Open |  |  | 155 | Software | Output/Calculation | Software update | Software Update
Z-2559-2011 | 58226 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.34.00 through 4.50.00.For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Focal Release 4.34.00 through 4.50.00 software | 2 | June 14, 2011 | 2011 | Computerized Medical Systems Inc | Focal Software: When a patient is defined as prone and the shifts are manually entered and the LOCK SHIFT box is checked the X coordinate (left/right) and the Z coordinate (anterior/posterior) are reversed (left to right, anterior to posterior). This could potentially lead to incorrect isocenter definitions from prone patients and thus a mislocation of dose. | DESIGN: Software Design | Computerized Medical System sent a CUSTOMER ADVISORY dated June 12, 2009, to all affected users. This advisory identified the product, the problem, and the action needed to be taken by the customer. There were no patients mistreated as a result of this issueThe firm will send out a User Notice to all sites still runningaffected software to remind them that the issue resolution is available.For further questions, please call (403 ) 830-8023. |  | 462 | 462 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and the countries of Algeria, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Egypt, Estonia, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, Nicaragua, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, and Venezuela. | Open |  |  | 462 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-2560-2011 | 58229 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 1.0.0 through 1.0.2. Used to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Monaco RTP System | 2 | June 14, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: If the user selects Monte Carlo Photon as the Secondary Algorithm (to be used for final dose calculation through segments), and selects SKIP Forward button at any time during stage 2 optimization, the system may only use the pencil beam algorithm for final computation of dose even though Monte Carlo was selected. Other than specific output to the console window, there will be no other in | DESIGN: Software Design | Computerized Medical Systems, Inc. sent a "MONACO CUSTOMER ADVISORY" letter dated June 2008 to all affected customers. The letter describes the product, problem, and the action to be taken. The customer was instructed to use caution when using the Skip Forward button in stage 2 according to written instructions provided in the letter.This defect exists in Monaco Release 1.0.0 through 1.0.2 and will be resolved as soon as possible. Note: The issue was resolved in Monaco Release 2.0.1 and above.Questions should be directed to the manufacturer's representative telephone number at 408-830-8023. | No consumer action necessary | 47 | 47 | Worldwide Distribution-- USA (nationwide) including the states of IL, MI, MO, OR, SC, TX, VA, and Puerto Rico and the countries of Canada, China, Cyprus, Estonia, France, Germany, Greece, Italy, Japan, Netherlands, Philippines, Poland, Portugal, Slovenia, South Korea, Spain, and United Kingdom. | Terminated | April 04, 2013 | 660 | 47 | Software | Output/Calculation | Instructions/Update | Safety Notice/Insructions
Z-2558-2011 | 58272 | electrode, ion-specific, chloride | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTI LION E-Plus Cassette, Model BP7507, Opti Medical (25 Cassettes with Samples Probes per box). Cassettes are consumables used in the OPTI LION Electrolyte Analyzer intended to be used for the measurement of sodium potassium, chloride, ionized calcium and pH in samples of whole blood, serum and plasma. . Cassettes contain one-time use sensors that are used with the OPTI-LION analyzer to perform in-vitro measurements of electrolytes. OPTI LION E-Plus Cassettes are supplied in boxes of 25 and may only be used with the OPTI LION Electrolyte Analyzer. | OPTI LION EPlus | 2 | June 14, 2011 | 2011 | OPTI Medical Systems, Inc | The values reported by the chloride sensors may report erroneous values below specification on aqueous controls and potentially with patient samples of whole blood and serum/plasma. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | OPTIMedical Systems sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 15, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were notified by letter via e-mail and fax. Customers were instructed to discontinue use of the product and to return any unused product to OPTI Medical Systems, Inc. Customers outside the U.S. were instructed to destroy unused product and certify the destruction on the enclosed Recall Response Card. The Response Card was to be completed and returned to OPTI Medical Customer Service via fax. OPTI Medical Systems will replace the cassettes they have with product from another lot.Customers were instructed to contact the Technical Support Department at 1-800-490-6784 for questions relating to this notice. |  | 86 | 86 boxes (25 units (per box) | Worldwide Distribution-- USA (nationwide) including states of CA and VA, and countries of India, Lebanon, Philippines, Sweden, Venezuela and Vietnam. | Terminated | August 08, 2011 | 55 | 86 | Not_Computer
Z-2556-2011 | 58762 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 804159, Terumo Coronary Ostia cannula, 14 Fr, 10 (25 cm) long A cardiopulmonary bypass vascular catheter, cannula, or tubing is a device used in cardiopulmonary surgery to cannulate the vessels, perfuse the coronary arteries, and to interconnect the catheters and cannulas with an oxygenator. The device includes accessory bypass equipment. | Terumo Coronary Ostia Cannulae | 1 | June 14, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The presence of adhesive and/or skiving of plastic material observed inside the distal tip of the cannula, which could lead to particulate and/or decrease in the internal diameter. The plastic skiving may detach and enter the patient, embolizing the coronary vasculature. The potential risk due to decrease in the internal tip diameter is inadequate delivery of cardioplegia and/or prolonged procedu | PRODUCTION CONTROLS: Process Control | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 16, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken. The letter described the problem, the potiential hazard, and the action to be taken by the consignee. Consignees were instructed to discontinue use of the product and return all product to Terumo. The letter provided a list of alternate products to be used instead of the recalled products. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM. | Contact the recalling firm for information | 960 | 960 | Worldwide Distribution - USA (nationwide) and the countries of BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES (UAE), SINGAPORE, MALAYSIA, AND SOUTH AFRICA. | Terminated | March 21, 2012 | 281 | 6400 | Not_Computer
Z-2563-2011 | 58676 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface to Data Innovations Instrument Manager software; GE Healthcare IT Solutions, Barrington, IL 60010. The Centricity Laboratory is designed to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. It integrates with other Hospital Information Systems (HIS) through HL7 interface by order entering and report processing. The centricity is intended to interface with various lab devices through standard HL7 interfaces to record, report and distribute lab results. | GE Centricity Laboratory Instrument Interface Data Innovations Instrument Manager Software | 2 | June 14, 2011 | 2011 | Ge Healthcare It | The GE Centricity Laboratory Instrument Interface to Data Innovations is processing a preliminary result instead of the final result from the Process System Manager when there is a test rerun conducted. | DESIGN: Software Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 3, 2011 to all affected customers. The letter describes the product, problem, and action to be taken by the customer. The letter provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. | Contact the recalling firm for information | 14 | 14 units | Worldwide Distribution-- USA (nationwide) including states of Georgia, New York, Ohio, Oregon, and Tennessee and countries of Canada, England, and Qatar. | Terminated | October 04, 2011 | 112 | 14 | Software | Device Operation | Software update | Software Update
Z-2565-2011 | 58747 | prothrombin-proconvertin and thrombotest | JPF | Prothrombin-proconvertin test and thrombotest. | Hematology | Hematology | Enforcement Discretion | Factor II (Prothrombin) G20210A Kit Light Cycler Instrument; Manufactured in Germany for Roche Molecular Systems, Inc. Branchburg, NJ 08876 USA DisFor use with Light Cycler 2.0 instrument; The Factor II (Prothrombin) G20210A Kit allows the detection and genotyping of a single point mutation (G to A at position 20210) of the human Factor II gene from DNA isolated from human whole peripheral blood. The test is performed on the LightCyclerÂŽ 2.0 Instrument utilizing polymerase chain reaction (PCR) for the amplification of Factor II DNA recovered from clinical specimens and fluorogenic target-specific hybridization for the detection and genotyping of the amplified Factor II DNA. The Factor II (Prothrombin) G20210A test is an in vitro diagnostic test for the detection and genotyping of the Factor II (Prothrombin) G20210A mutation as an aid to diagnosis in the evaluation of patients with suspected thrombophilia. The test is intended to be used on the LightCyclerÂŽ 2.0 Instrument using the Roche Confidential Page 3 of 9 LightCyclerÂŽ Software 4.05 or 4.1. The specimen preparation must be performed according to the workflow procedures described in the package insert.tributed by Roche Diagnostics Indianapolis, IN 46256 USA; 03610195001 | Factor II (Prothrombin) G20210A Kit | 2 | June 14, 2011 | 2011 | Roche Molecular Systems, Inc. | Roche Molecular Systems has decided to discontinue marketing the Factor II (Prothrombin) G20210A Kit and the Factor V Leiden Kit for use with the Light Cycler 2.0 Instrument in the US Market | OTHER/UNDETERMINED: Pending | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to discontinue use of the affected product by August 31, 2011.Complete the attached faxback form ( 4293-00-0311 ) and fax it to 1-888-345-5359.For questions or technical support, please contact Roche Molecular at 1-800-526-1247,. | Contact the recalling firm for information | 13864 | 13,864 units | Nationwide Distribution | Terminated | August 13, 2012 | 426 | 30890 | Not_Computer
Z-2564-2011 | 58258 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 2.0.1 and aboveUsed to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Monaco Radiation Treatment Planning System | 2 | June 14, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: due to a software defect, the Radiological Depth (cm) and SSD to DRP (cm) parameters are reported incorrectly on the Beam Summary Report. This occurs for all Monaco IMRT plans. | DESIGN: Software Design | Computerized Medical System sent a "USER NOTICE" letter dated June 24, 2010, to all affected users. This notice identified the product, the problem, and the actions to be taken. There is no current workaround. All customers will be mailed a postcard to be returned as proof of receipt and understanding of the notification. Consignees will be contact via telephone or e-mail after June 30, 2010 if the postcard is not returned This notice is to be distributed to any and all users of CMS software at your organization who are potentially affected. The notice states the issue will be resolved in Monaco Release 2.04, but no date for this release was given.For further questions, please call (403) 830-8023. |  | 62 | 62 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Austria, Belgium, China, Finland, France, Germany, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, South Korea, Spain, Switzerland, Taiwan, Thailand and United Kingdom. | Open |  |  | 62 | Software | Output/Calculation | Software update | Software Update
Z-2569-2011 | 57843 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.04.00 and above.Used to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Monaco Radiation Treatment Planning Workstation | 2 | June 15, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: In XiO IMRT (Dynamic MLC Delivery) software, the radiation beam angle is set to zero. If different beam angles are selected via Monaco software, the beam angles will remain at zero. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent a User Notice to the affected consignees on October 27, 2010, explaining the reason for correction, the clinical impact, and the workaround method. (VERIFY USE NOTICE DATE) A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 11 | 11 | Worldwide Distribution - USA including CA, IL, MN, NM, OH, NY, TX, and WA and the Netherlands. | Open |  |  | 11 | Software | Output/Calculation | Software update | Software Update
Z-2567-2011 | 58227 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.1 through 4.40.00.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO RTP System | 2 | June 15, 2011 | 2011 | Computerized Medical Systems Inc | XiO Software: Under certain conditions, when the port shape is altered, dose is not recalculated. The port-change flag is not being set correctly when the port shape is regenerated. The dose distribution will not match the blocking displayed and the mis-match might not be obvious. Monitor Unit calculations will also be based on an incorrect dose distribution. This problem has existed since Re | DESIGN: Software Design | Computerized Medical Systems, Inc. sent a "XiO CUSTOMER ADVISORY" letter dated December 8, 2008 to all affected customers. The advisory describes the product, problem, and the actions to be taken. The customers were instructed to set the prompt to "Yes" to avoid the problem from occurring according to written instructions provided in the advisory. This problem has existed since Release 4.3.1 and is expected to be resolved with XiO Release 4.50.00. Questions should be directed to the manufacturer's representative at 408-830-8023. |  | 784 | 784 | Worldwide Distribution-- USA (nationwide) including Puerto Rico and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Tajikistan, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 784 | Software | Output/Calculation | Software update | Software Update
Z-2572-2011 | 58233 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.1 through 4.34.02.For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Focal Radiation Treatment Planning System, | 2 | June 15, 2011 | 2011 | Computerized Medical Systems Inc | Focal: In absolute dose mode a beam with a port may be calculated with open normalization, with the result that the displayed dose reflects open normalization while the monitor units and QA plan reflect blocked normalization. This is particularly important when these beams also include compensators. The problem is detectable by looking at the delivered dose to the weight point in the patient as | DESIGN: Software Design | Computerized Medical Systems Inc sent a "CUSTOMER ADVISORY" letter dated March 2008 to all affected customers. The letter describes the product, the problem and the actions to be taken. The firm provided a workaround for customers that states "Alternatively the plan can be imported directly into XIO". A copy of the Advisory was shipped with software within the affected range to those requesting a software upgrade. The issue was resolved in Focal Release 4.40.00 introduced in June 2008.The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available.For further questions, please call (403) 830-8023. |  | 307 | 307 | Worldwide Distribution -- USA (nationwide) and the countries of Albania, Argentina, Australia, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Thailand, Turkey, United Kingdom, Uruguay, and Venezuela. | Open |  |  | 307 | Software | Display/Image | Software update | Software Update
Z-2574-2011 | 58214 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.5.0 and above.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 15, 2011 | 2011 | Computerized Medical Systems Inc | In XiO, when Telepathy is exited, re-entered, and the temporary plan re-called, the OLD beam weight will be displayed but the dose distribution will reflect the modified beam weight. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter to all affected customers in May 2008. The letter identified the product, the problem and the action needed to be taken by the customer.The issue was resolved in XiO Release 4.50.00 which has been available since June 2009.A User Notice was sent in March 2011, to all the customers reminding them of the availability of the correction. This notice was sent hardcopy via US Mail but there was no return receipt.For further questions, please call ( 408) 380-8023. |  | 936 | 936 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 936 | Software | Display/Image | Software update | Software Update
Z-2568-2011 | 58215 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 through 4.50.00.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 15, 2011 | 2011 | Computerized Medical Systems Inc | In XiO, when the beams are renumbered, it is possible to get into a state where the dose displayed for a beam does not match the beam placement on the patient. This can occur when dose is calculated, beams are renumbered, and the dose is re-summed only. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter dated May 2009, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.Computerized Medical Systems, Inc have created "patches" to resolve the problem. A User Notice was sent in March 2011, to all users still using the affected software notifying them of the updated software available to correct the problem. This letter was sent by US Mail but no return receipt was included.For further questions, please call (408) 380-8023. |  | 1116 | 1116 | Worldwide Distribution -- (USA) nationwide Distribution, including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia, Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1116 | Software | Display/Image | Software update | Software Update
Z-2571-2011 | 58657 | electrode, ion specific, calcium | JFP | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ca++ Membrane Units, Model #942-046, and Model #942-060 used on ABL505/555, EML105, ABL600, ABL700 Series, and ABL800 Analyzers. The membranes are packaged in a plastic box which contains four (4) membranes per box. Calcium membrane jackets that are used in the ABL505/555, EML105, ABL600, ABL700 Series and ABL800 Analyzers for InVitro testing of samples of whole blood for the parameters pH, PO2, cK+, cNa+, cCa2+, cCl-, cGlu, cLac, CCrea, ctBil, and co-oximetry parameters (ctHb, sO2, and the hemoglobin fractions: FO2Hb, FCOHb, FMetHb, FHHb andFHbF). | calcium membrane jackets | 2 | June 16, 2011 | 2011 | Radiometer America Inc | Based on internal testing and user/customer complaint reports received, the firm determined that wrinkles may arise on the outer membrane of the calcium membrane units upon installation. Retention of fluids between the wrinkles can cause a positive bias resulting in calcium measuring errors. This situation poses a risk to the safety of patients, particularly when the measurement results are used |  | The firm, Radiometer America, sent a "MEDICAL DEVICE FIELD CORRECTION" letter dated April 14, 2011, to all affected customers. The letter describes the product, problem and the actions to be taken. The letter provides the customers with a series of steps to be taken when replacing the Calcium Membrane (Until further notice the above procedure must be carried out when replacing the Calcium membrane). The letter also provides quality control instructions for verifying calcium performance. Additionally, the customer is required to manually adjust the upper and lower QC limits on the analyzer in relation to the limits printed on the insert sheet in order to detect the bias during quality control measurements (This must be done for every lot of QC ampoules used, and for any new lot of ampoules installed).The letter asks that each customer complete and return the attached form via fax to 440-871-0463, in order to verify that the notification letter was received. Questions should be directed to the firm's Technical Support staff at 800-736-0600, Opt. 4. |  | 30323 | 30323 boxes | Worldwide Distribution--USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV, and the countries of: Algir, argentina, Austria, Austrailia, Bahrain, Belarus, Belgium, Brazil, Brunei Darussalam, China, Columbia, Cuba, Greece, Denmark, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greenland, Guatemala, Hungary, India, Iraq, Ireland, Isreal, Italy, Jamaica, Japan, Kasakhstan, Kenya, Lebanon, Lithuania Korea, Malaysia, Malta, mexico, Moldavia, Nepal, New Zealand, Norway, Netherlands, Peru, Philipines, Poland, Portugal, Qatar, Oman, Rumania, Russia, Saudi Arabia, Singapore, Serbia and Kosvo, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Syria, Thailand, The Netherlands, Tunesia, Turkey, UK, Ukraine, United Arabia Emirates, Urguary, Usbekistan, and Venezuela. | Terminated | February 19, 2012 | 248 | 30323 | Not_Computer
Z-2579-2011 | 58197 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 16, 2011 | 2011 | Computerized Medical Systems Inc | XiO: When the beam template is recalled, XiO prematurely initiates a dose calculation, re-snaps the MLC or Port to the structure of the current patient , but then fails to re-calculate dose for the new MLC or Port position. If no subsequent change is made to force a dose recalculation, XiO will display the dose from the original MLC/Port shape dose instead of the dose from the re-snapped MLC/Port | DESIGN: Software Design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 14, 2010 to its affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to distribute this notice to any all users of CMS software at their organization who are potentially affected by this issue, and complete and return the attached postcard to confirm receipt of the Safety Notice. The notice also states that this problem has existed since the XiO Release 4.3.0 and is resolved in the XiO release 4.26.00. Elekta will notify the customers when this solution is available.If you have any questions, please call 314-993-0003 or toll free 800-878-4267. | Contact the recalling firm for information | 1208 | 1208 | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1208 | Software | Device Operation | Software update | Software Update
Z-2578-2011 | 58200 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation treatment Planning System, XiO Release 4.2.0 and aboveUsage: Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation treatment Planning System | 2 | June 16, 2011 | 2011 | Computerized Medical Systems Inc | When a dialog box is displayed and the dose status is changed to "Dose not current", beam edits can be made in the graphics area using the mouse (rotate collimator, reposition collimator, move isocenter, move weight point, etc.). Then if CANCEL is selected on the dialog box, any edits made on the dialog box are undone, the dose is left unchanged, but the beam edits made in the graphics area with t | DESIGN: Software Design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 6, 2010 to its affected customers. The notice describes the product, problem and actions to be taken (workaround). The customers were instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue. The notice also states that this problem has existed since XiO Release 4.2.0 and will be resolved in XiO Release 4.70.00. Elekta will notify the sites when this solution is available. If you have any questions, please call 314-993-0003 or Toll Free: 800-878-4267. | Contact the recalling firm for information | 1293 | 1293 | Worldwide distribution: USA (nationwide) including Puerto Rico, and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1293 | Software | Device Operation | Software update | Software Update
Z-2577-2011 | 58210 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.0 through 4.40.00Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 16, 2011 | 2011 | Computerized Medical Systems Inc | XiO: When the patient is positioned with their feet toward the CT gantry, the coordinates exported to the LAP IsoMark laser positioning system are not correct. | DESIGN: Software Design | The firm, CMS, Inc., sent a "XiO Customer Advisory" letter dated April 8, 2008 to all affected customers. The letter describes the product, the problem, and the actions to be taken ( workaround). The letter also states the problem will be resolved in a future release. No release number or date was given.If you have any questions, please call 408-380-8023. |  | 75 | 75 | Worldwide distribution: USA (nationwide) and countries of: China, France, Germany, Japan, Mongolia, Netherlands, Nicaragua, Poland, Russia, South Korea, Switzerland, and Turkey. | Open |  |  | 75 | Software | Output/Calculation | Software update | Software Update
Z-2529-2011 | 58910 | stent, superficial femoral artery | NIP |  |  | Cardiovascular | PMA | Boston Scientific Innova Over-the-Wire Self-Expanding Stent System, Catalog No. 39181-07403. 7 mm x 40 mm x 130 cm (stent diameter x stent length x delivery system length). Made in USA, Two Scimed Place, Maple Grove, MN 55311 USA. The Innova Self-Expanding Stent System is comprised of two components: the implantable endoprosthesis and the stent delivery system. The stent is a laser cut self-expanding stent composed of a nickel titanium alloy (Nitinol). On both the proximal and distal ends of the stent, radiopaque markers made of tantalum increase visibility of the stent to aid in placement. The stent is constrained within a 6F (2.1 mm maximum OD) delivery system. The delivery system is a triaxial design with an outer shaft to stabilize the stent delivery system, a middle shaft to protect and constrain the stent, and an inner shaft to provide a guidewire lumen. The delivery system is compatible with 0.035 in (0.89 mm) guidewires. When ready to be implanted, the stent is deployed by retracting the middle shaft of the delivery system. A radiopaque marker at the distal end of the delivery system aids in visibility during deployment. As the stent is exposed to body temperature, it expands to appose the vessel wall.The Innova Self-Expanding Stent is available in a variety of diameters and lengths.The delivery system is also offered in two working lengths (75 cm and 130 cm).The Innova Self-Expanding Stent System is indicated for the treatment of peripheral vascular lesions | Innova SelfExpanding Stent System | 1 | June 16, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a recall of the Innova Self-Expanding Stent System because they have received 6 complaints involving no deployment/partial deployment of the Innova Stent. Potential health hazard events resulting from this type of failure include increased procedure time, vessel wall injury, and/or emergency surgery to remove the partially deployed stent. This recall d | DESIGN: Process Design | Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter, dated May 13, 2011, to all affected customers via registered mail. The letter explained the issue, identified affected product, stated that further distribution and or use of the devices should cease immediately, and asked that devices in inventory to be segregated and immediately returned to Boston Scientific. An Account Reply Verification Tracking Form was also asked to be returned. Boston Scientific Sales Representative were assist in the return process by providing a Returned Goods Authorization (RGA) number.Customers were instructed to contact their local Sales Representative for any questions. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Product removal in US is being managed through SuperNOVA Clinical Trial suspension in accordance with IDE #G100291. OUS distribution to include: Austria, Finland, France, Germany, Italy, Netherlands, Spain, and Switzerland. | Terminated | February 04, 2014 | 964 |  | Not_Computer
Z-2582-2011 | 57807 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. XiO Release 4.50.00 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy of brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 17, 2011 | 2011 | Computerized Medical Systems Inc | XiO: When a new electron beam is added to a plan, the calculation Algorithm is set to Pencil Beam, and an Aperture is added to the beam, the user can select Port-Port Properties and edit the Material Thickness values. If the algorithm is then changed to Monte Carlo, the calculation will proceed and the user is not warned the Material/Thickness values just entered are not used in the dose calculati | OTHER/UNDETERMINED: Pending | Elekta CMS Software drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was distributed to all affected sites in November 2010. The issue will be resolved in the XiO Release 4.70.00, expected to be available in July 2011. |  | 556 | 556 | Worldwide Distribution - USA (nationwide, Washington, DC and Puerto Rico and the countries of Albania, Algeria, Australia, Austria, Bahamas, Belarus, Belgium, Brazil, Canada, Chile, China,Croatia, Cyrus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Phillipines,Poland, Portugal, Romoania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, Turmenistan, Ukraine, United Kingdom, and Vietnam | Open |  |  | 556 | Software | Treatment/Delivery/Therapy | Software Patch | Software Update
Z-2580-2011 | 58787 | equipment, traction, powered | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | MTD 4000 - Mettler Traction Decompression System, Model: ME 4000Product Usage: Provide traction and mobilization of skeletal structures and skeletal muscles | Mettler Traction Decompression System, Model: ME 4000 | 2 | June 17, 2011 | 2011 | Mettler Electronics Corp | The recall was initiated because Mettler has confirmed the possible failure of internal mechanical component and software detection for potential failure regarding Electronics Traction Device - MTD4000. The firm are initiating the recall because component failure might result in patient injury. Use of the device should cease immediately. |  | Mettler Electronics sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 14, 2011 to all affected consignees. The letter identified the product, the problem and the actions to be taken. Customers were instructed to immediately cease use of the recalled device, examine their inventory and quarantine product subject to recall. In addition, consignees were told if they have further distributed the affected product to please identify their customers and notify them at once of this product recall. The letter asks that that each includes a "Recall Return Response Form" that consignees are required to complete and return. If you have any questions please contact Mettler Electronics Corp 714-533-2221 x331, Monday through Friday, between the hours of 8:30 A.M. and 4:00 P.M. Pacific Time. | Contact the recalling firm for information | 221 | 221 | Worldwide Distribution --USA (nationwide) and countries of Taiwan, Ireland and Haiti. | Terminated | July 10, 2012 | 389 | 221 | Hardware |  | Remove | Remove or Replace
Z-2590-2011 | 58717 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ENDO-FUSE(TM) FUSION ROD, 7 mm X 70 mm IMPLANT MATERIAL: STERILE R, Ti6A14V, CPTi, REF: 500005970, LOT 088605422, Wright Medical Technology, Inc , 5877 Airline Road, Arlington, TN 38002The ENDO-FUSE(R) Intra-Osseous Fusion Rods are indicated for the reduction and fixation of fractures. They are indicated for the integral fixation of fractures, bony fusions, and non-unions. They are also indicated for reconstructive procedures where reduction and fixation of bone fragments are required (e.g. osteotomies). | ENDOFUSE(R) Fusion Rods | 2 | June 17, 2011 | 2011 | Wright Medical Technology Inc | Two lots- one lot of Endo-Fuse Fusion Rod, 7x50mm was commingled one lot of Endo-Fuse(R) Fusion Rod, 7x70mm. Not all parts in each lot are affected. The commingled rod sizes might not be detected prior to surgery. | PRODUCTION CONTROLS: Process Control | The firm, Wright Medical, sent an "URGENT: Medical Device Voluntary Recall" letter dated May 6, 2011 to their customers via FedEx . The letter describes the product, problem, and actions to be taken. The customers were instructed to examine their inventory for the affected products; immediately return any affected product to Wright Technology as soon as possible (contact Customer Service at 800-238-7117 for return instructions and replacement inventory), and immediately complete and return the attached response form via fax to: 901-867-7401.If you have any questions regarding this matter, please contact 800-874-5630. |  | 121 | 121 | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, ID, IN, IL, LA, MI, MN, MO, NC, NE, NH, NY, OH, PA, SC, TN, TX, VA, VT, WA, and WI. | Terminated | December 27, 2012 | 559 | 186 | Not_Computer
Z-2586-2011 | 58784 | assay,hybridization and/or nucleic acid amplification for detection of hepatitis c rna,hepatitis c virus | MZP |  |  | Microbiology |  | COBASÂŽ AmpliPrep / COBASÂŽ TaqManÂŽ HCV Test, US-IVDThe COBASÂŽ AmpliPrep/COBASÂŽ TaqManÂŽ HCV Test is an in vitro nucleic acid amplification test for the quantitation of Hepatitis C viral (HCV) RNA in human plasma or serum of HCV-infected individuals using the COBASÂŽ AmpliPrep Instrument for automated specimen processing and the COBASÂŽ TaqManÂŽ Analyzer or the COBASÂŽ TaqManÂŽ 48 Analyzer for automated amplification and detection. Specimens containing HCV genotypes 1 6 have been validated for quantitation in the assay. The COBAS AmpliPrep/COBAS TaqMan HCV Test is intended for use as an aid in the management of HCV-infected individuals undergoing anti-viral therapy. The assay measures HCV RNA levels at baseline and during treatment and can be utilized to predict sustained and non-sustained virological response to HCV therapy. | COBAS AmpliPrep / COBAS TaqMan HCV Test, USIVD | 2 | June 17, 2011 | 2011 | Roche Molecular Systems, Inc. | The COBAS Ampliprep/COBAS TaqMan HCV Test has been shown to under-quantitate a subset of genotype 4 patient specimens by approximately 1.0-1.5 log 10 in the absence of any sequence mismatches. | OTHER/UNDETERMINED: Pending | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2011 and faxback form to their customers UPS ground. The letter described the product, problem and actions to be taken. The customers were instructed to review prior test results if there have been indications of discrepancies between test results generated with COBAS AmpliPrep/COBAS TaqMan HCV Test as compared to other relevant clinical and laboratory findings; complete and return the attached Product Advisory Notice: Affiliate Feedback Fax-Form via fax to: 1-866-792-5445, and file this letter for future reference, in particular for use with their 2010 CAP Survey Results to explain any under-quantitation of sample HVL 2-01.Please contact Roche Molecular Diagnostics Technical Support at 1-800-526-1247 if you have any questions about the information contained in this letter. |  | 50993 | 50,993 units | USA (nationwide) and countries of: Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudia Arabia, Sinagpore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | August 27, 2012 | 437 | 339392 | Not_Computer
Z-2584-2011 | 59037 | electrocardiograph | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS3 Vital Signs Monitor Models:VS3863069, 863070, 453564041251, 453564041261, VS3863071, 863072, 863073, 863074, 453564041271, 453564041281, 453564041291, 453564041301The SureSigns VS3 vital signs monitor and VM4, VM6 and VM8 patient monitors are for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. The SureSigns VS3 vital signs monitor and VM4, VM6 and VM8 patient monitors are for monitoring, recording, and alarming of multiple physiological parameters of adults, pediatrics, and neonates in healthcare environments. Additionally, the monitor is intended for use in transport situations within a healthcare facility. | Philips Healthcare | 2 | June 17, 2011 | 2011 | Philips Healthcare Inc. | Speakers on the SureSigns Patient Monitors and Vital Signs, failure may be identified by the display of a Speaker Malfunc message in the alarm message area or an AUDIO FAILED message in the main screen of the device or by the absence of audible sound | DESIGN: Software Design (Manufacturing Process) | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 2011 to all affected customers. The letter identifies the products, problem, and actions to be taken.The letter states that customers may continue to use the affected products while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP. Additionally, the letter states that if the customer experience this INOP or there is no sound from the SureSigns Monitor, the customer is to remove the device from use and contact their local Philips service representative.The letter explains that Philips is completing the design for a replacement speaker assembly, which is expected to be available within 8 weeks. When available, Philips will replace the affected speaker assemblies free of charge.Customers are instructed to contact their local Philips Customer Care Center at 1-800-722-9377 if further information or support is required. | Contact the recalling firm for information | 5225 | 5225 units | Worldwide Distribution--USA (nationwide) and countries of ARAB JAMAHIRI, AUSTRIA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, CANADA ,CHILE, COLOMBIA, CROATIA, CYPRUS, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, KUWAIT, LEBANON, LIBYAN, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND, OMAN, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA ,SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | August 15, 2013 | 790 | 5740 | I/O |  | Replace | Remove or Replace
Z-2591-2011 | 58853 | instrument for treatment of hyperhidrosis | OUB | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MiraDry System consisting of the DTS3000 Console, MiraDry Handpiece(MD3001-HP), the miraDry biotip and the miraDry Treatment Template system.Miramar Labs, Inc.445 Indio WaySunnyvale, CA 94085Intended use: Treatment of primary axillary hyperhidrosis. | MiraDry System | 2 | June 17, 2011 | 2011 | Miramar Labs, Inc. | Reports of skin burns on axilla after treatment with the miraDry System. | DESIGN: Device Design | The firm,Miramar, has a service provider in Japan that assists in dealing with customers. The service providers were provided a copy of the investigation report and communicated directly with the customers to set up appointments for the field correction. The customers were personally visited by a representative from Miramars service provider in Japan. All devices had field correction performed on April 26-28, 2011. |  | 8 | 8 units | International only: Japan | Terminated | June 17, 2011 | 0 | 8 | Not_Computer
Z-2588-2011 | 58808 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Kinamed SuperCable Grip and Plate System Driver, 3.5mm Hex, Catalog Number: 35-860-2060The driver is intended to be used in conjunction with compression and locking bone screws for fixation of long bone fractures when Kinamed Trochanteric Grips & Cable-Plates are used.SubsequentProduct Codes: HWC, JDQ. | SuperCable ScrewDriver, 3.5mm Hex | 2 | June 17, 2011 | 2011 | Kinamed Inc | The recall was initiated due to inadequate mechanical strength of these screwdrivers' hexagonal tipfeature, causing it to strip during final tightening. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Kinamed Inc., sent a recall letter dated August 18, 2010 to consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The customers were instructed to Immediately examine their stock and return any Kinamed SuperCable Screwdriver, 3.5mm Hex bearing these lot numbers, along with the completed attached Effectiveness Check Form and shipment under RMA #74567 to the attention of: "Director of Quality Assurance"Kinamed, Inc.820 Flynn RoadCarnarillo, CA 93012-8701Note: The Kinamed sales agent shipped and had the product replaced for the customers.If you have any questions, please contact (303)-681-2208, (805) 384-2748 or(800) 827-5775. |  | 8 | 8 units | Worldwide distribution: USA (nationwide) state of: CO and country of: Columbia. | Terminated | October 25, 2011 | 130 | 8 | Not_Computer
Z-2585-2011 | 58837 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Integra CUSA NXT Ultrasonic Tissue Ablation System Service ModuleThe CUSA NXT Ultrasonic Surgical Aspirator System is indicated for use in surgical procedures where fragmentation. emulsification and aspiration of soft and hard (e.g. bone) tissue is desirable. including Neurosurgery. Gastrointestinal and affiliated organ surgery. Urological surgery. Plastic and Reconstructive surgery. General surgery, Orthopedic surgery. Gynecological surgery, Thoracic surgery. laparoscopic surgery and Thoracoscopic surgery. | Integra CUSA Selector NXT Ultrasonic Surgical Aspirator | 2 | June 17, 2011 | 2011 | Integra Limited | The insulated covering of a wire cable that powers the vacuum pump in the CUSA NXT Service Module can become abraded over time resulting in the exposure of bare wire. If this bare wire were to come in contact with the metal surface of the pump, it could result in a short circuit and a loss of aspiration. | DESIGN: Device Design | Integra sent a "FIELD ACTION NOTIFICATION" dated May 2, 2011 to all affected customers. The letter identified the product, the problem and the actions to be taken. Customers were informed of a Field Corrective Action that is being initiated. All devices (customer owned, Integra Loaner Pool, and Sales demos) will undergo rework of the vacuum pump assembly. The rework will be scheduled, at customer sites by Integra Field Service Technicians or returned to a designated Integra Service Center for rework. A loaner will be presented to customer accounts upon request. The letter asks that each customer complete and return the enclosed Recall Acknowledgement and Return Form to Integra, maintaining a copy of the letter and form for your records. For any additional information please call 609-936-2485. |  | 105 | 105 units | Worldwide -- USA (nationwide) and countries of France, Germany, Hungary, Ireland, Italy, Netherlands, Spain, Sweden and Taiwan | Open |  |  | 105 | Hardware | Treatment/Delivery/Therapy | Rework | Repair
Z-2603-2011 | 58763 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, porous | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encompass Hip Stem, Plasma Spray, 12/14 Taper, Extended, REF 109-XXXX, Sterile R, Ortho Development, Draper, Utah 84020. Intended for use in total hip arthroscopy. Device fixation is achieved by an optimal, press-fit in the medullary canal, which maximizes contact between the stem and the bone. | Encompass Hip Stem, Plasma Spray | 2 | June 20, 2011 | 2011 | Ortho Development Corporation | Tests on PVC packaging material resulted in alert levels for cytotoxicity. | DESIGN: Packaging Design/Selection | Ortho Development Corporation notified consignees by letter on May 11, 2011. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to return all remaining affected units to Ortho. Replacement product was to be shipped by May 31. |  | 113 | 113 units | Nationwide Distribution - USA including CA, FL, and TX. | Terminated | November 09, 2011 | 142 | 113 | Not_Computer
Z-2599-2011 | 58250 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.1.1 through 4.40.00.Product Usage: for use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | FOCAL Workstation | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | Focal Software: if individual MLC leaf position edits are made, the original MLC aperture is retained. A subsequent aperture edit should not affect the MLC leaf positions however it was found that instead the leaves will then snap (unintentionally) back to the original aperture. Also, if the patient is sent to XiO after the MLC is edited, an isocenter shift made in XiO will also cause the MLCs to | DESIGN: Software Design | Elekta CMS Software sent an "ADVISORY NOTICE" dated February 6, 2009 to its affected customers. The notice identified the product, the problem, and the actions to be taken. Customers were informed that the workaround method available is to (edit the MLC pattern using the aperture first and the individual MLC leaf edits second). The firm states that this issue will be resolved in Focal version 4.50.00 and above. A "User Notice" will be sent to all users still using the affected software notifying them of the updated software available to correct the problem. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. | No consumer action necessary | 760 | 760 | Worldwide Distribution -- USA (nationwide) including the state of: Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, Uruguay and Venezuela. | Open |  |  | 760 | Software | Output/Calculation | Software update | Software Update
Z-2598-2011 | 57801 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation, Monaco Release 2.03.00 and aboveTo create treatment plans for any cancer patient for whom external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Monaco Radiation Treatment Planning Workstation | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: When the isocenter location is edited in the IMRT Activity screen, the shift coordinates are not updated in the IMRT Plan report. | DESIGN: Software Design | Computerized Medical Systems Inc drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was sent January 13, 2011. A return postcard is included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. | No consumer action necessary | 137 | 137 | Worldwide Distribution - USA (Nationwide distribution to AK, AZ, CA, CO, FL, IL, IN, MA, MI, MN, MO, NC, NM, NY, OH, OR, PA, SC, TX, VA, WA, and WI and the countries of Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Phillipines, Poland, Portugal, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and the United Kindom. | Open |  |  | 137 | Software | Display/Image | Software update | Software Update
Z-2597-2011 | 58502 | processor, tissue, automated | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Peloris II Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd, Mount Waverly, 3149 VIC, Australia; catalog number 26.0005.This tissue processor prepares tissue samples for sectioning by transforming fixed samples into wax embedded samples. This is achieved by exposing the tissue samples to a sequence of reagents in the processing retorts. | Peloris II Rapid Tissue Processor | 2 | June 20, 2011 | 2011 | Leica Microsystems, Inc. | The processor potentially has a faulty impellor which could lead to reagent carry-over from one processing step to another, | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Leica Microsystems sent a MEDICAL DEVICE CORRECTION letter dated April 11, 2011, to the affected customers via first class mail. The letter identified the product, the problem, and the action needed to be taken by the customer.The Customers were instructed to discontinue use of the device until their local Leica Microsystems representative contact them to arrange the replacement of the faulty impellor. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. | No consumer action necessary | 22 | 22 units | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Canada, Germany Japan, Portugal and United Kingdom. | Terminated | June 20, 2011 | 0 | 22 | Not_Computer
Z-2595-2011 | 58260 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.2 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | XiO Software: when Digitally Reconstructed Radiographs (DRRs) are generated in XiO, and DICOM is used to export the DRRs (File-Print-DRR), the exported images are associated with the Beams, but NOT the Plan. The Plan association had always been done manually by the user. This is functionality that has never been provided on XiO. The problem occurs because XiO is not storing a persistent DICOM | DESIGN: Software Design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available.No date was given for the mailing or the solution. | No consumer action necessary | 1292 | 1292 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1292 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-2592-2011 | 58269 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.50.00 and abovefor use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Focal Radiation Treatment Planning System | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | Focal: for XiO plans calculated in relative dose mode, the Focal Plan Review application will display the DVH and maximum dose correctly. However, the isodose display is not scaled correctly. If the user evaluated the plan based only on the isodose display, they might misinterpret the actual dose delivered to the patient. | DESIGN: Software Design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to consignees beginning on May 21, 2010. The notice identified the product, the problem, and a workaround. |  | 308 | 308 | Worldwide Distribution - USA (Nationwide including Puerto Rico) Foreign distribution was made to Albania, Australia, Austria, Belgium, Canada, Chile, China, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, and Turkmenistan. | Open |  |  | 308 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-2604-2011 | 58270 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above | XiO Radiation Treatment Planning System | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | XiO Software: when a treatment machine is defined in Source File Maintenance (SFM), the MLC Model is set to Elekta, and beams using this treatment machine are added to a new plan in Teletherapy, XiO automatically adds an MLC to the beam. This is because the Elekta linac will always have an MLC in place when the plan is actually delivered. | DESIGN: Software Design | A Safety Notice, dated 5/20/10, was sent to Elekta / Computerized Medical Systems Inc clients in May 2010. The notice indicates the problem, the clinical impact. The notice states a future XiO Release will resolve the issue. A return postcard is included with the mailing to verify consignee received the notice. For sites not returning the postcard by 5/28/10, the recalling firm will contact them via phone or email to confirm receipt of the notice. The notice also informed the consignees they would be contacted when a solution was available. | No consumer action necessary | 57 | 57 | Nationwide distribution to medical facilities including Puerto Rico. Distribution was also made to two Military and Government consignees. Foreign distribution was made to Australia, Belgium, Brazil, Canada, Colombia, France, India, Ireland, Israel, Italy, Malaysia, Netherlands, Philippines, Poland, Spain, and United Kingdom. | Open |  |  | 57 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2573-2011 | 58811 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Ultraview SL 91517 Capnography module is a mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture, to aid in determining the patient's ventilatory, circulatory, and metabolic status.Mainstream monitoring uses a self-calibrating CO2 sensor to ensure continuous monitoring without interruption of connection. Reusable, lightweight ( | Ultraview SL Capnography Module | 2 | June 20, 2011 | 2011 | Spacelabs Healthcare, Llc | Capnography Module, Model 91517, resets 20 seconds after power failure, instead of after three minutes as specified. This reset clears the RAM and the module returns to default settings. | DESIGN: Software Design | Spacelabs Medical sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 12, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. Customers were recommended to set all their Capnography Module Configuration Manager alarm settings to "ON". The letter also recommended that customers brief their staff about the possibility of losing changes to the default settings and to restore them after a power interruption.A field service engineer will be contacting customers to set up a date and time when affected modules can be upgraded at no cost to the customer. For questions about this corrective action program, call Spacelabs Medical at 1-800-522-7025, Select 2 for Technical Support. For outside the US, customers can call 1-425-657-7200, x5089. |  | 5883 | 5883 units | Worldwide Distribution--USA (nationwide) and countries of ALGERIA, ARGENTINA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA,COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FRANCE,GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, ISRAEL, ITALY,JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEPAL,NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PALESTINE, PANAMA,PERU, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND,SYRIAN ARAB REPUBLIC, TAIWAN, REPUBLIC OF CHINA, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | December 02, 2011 | 165 | 5883 | Software | Device Operation | Upgrade | Repair
Z-2596-2011 | 58256 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.0 through 4.34.01For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Focal Radiation Treatment Planning System | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | Focal: when a plan is sent to Focal from XiO and the structure properties are edited in Focal CT Sim activity, these changes result in changes to the MLC or field size for the plan when they should not. XiO will recalculate dose when modified plan is returned to XiO however the user may not notice that it has changed form the plant they had approved earlier. Editing the structure properties (col | DESIGN: Software Design | Computerized Medical Systems Inc sent a Customer Advisory, dated 11/30/07, was sent to consignees December 2007 to notify them of the issue and how to avoid the workaround. The advisory listed the problem, clinical impact, and a workaround. The firm will be sending each site a reminder that the issues is resolved. | No consumer action necessary | 233 | 233 | Worldwide Distribution - USA (Nationwide AL, AR, CA, DE, FL, GA, IL, IN, MA, MD, MI, MOP, NE, NY, OH, PA, RI, TX, VA, WA, WI, and WV and the countries of Albania, Argentina, Belgium, Bosnia and Herzegovina, Brazil, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Turkey, United Kingdom, Uruguay, and Venezuela. | Open |  |  | 233 | Software | Output/Calculation | Software update | Software Update
Z-2594-2011 | 58253 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.40.00 and above.Product Usage: for use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | FOCAL Workstation | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | Focal: the couch position can be moved between scans of the same CT study. When the images are imported into Focal, the software is not checking the DICOM Image Position (Patient) tag and therefore is not aligning the images sets if there are different couch positions indicated. A shift between image sets will exist of the same amount that the CT couch was moved. | DESIGN: Software Design | Elekta CMS Software sent an "ADVISORY NOTICE" dated January 19, 2009 to its affected customers. The notice identified the product, the problem, and the acton to be taken. Customers were informed that the workaround method available is to (use the image fusion feature already in Focal to relign the scans). The firm states that this issue is resolved in Focal version 4.62.00, which was released in March 2011. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. | No consumer action necessary | 31 | 31 | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, FL, IL, IN, KS, MO, NC, OH, OK, PA, TN, and WA. and countries of Australia, Austria, Canada, India, Netherlands, Portugal, Singapore, Spain, and United Kingdom. | Open |  |  | 31 | Software | Display/Image | Software update | Software Update
Z-2593-2011 | 58262 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 2.1.0 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | June 20, 2011 | 2011 | Computerized Medical Systems Inc | XiO Software: The Patient Scan position is being confused with the treatment orientation of the patient. The Patient Scan position is determined from the data entered in PFM. Subsequently, when a New Teletherapy Plan is created for a patient in Teletherapy, the user is prompted with Head/Foot Gantry on the text page. This prompt is asking for the orientation of the patient when they are treated. | DESIGN: Software Design | Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available. | No consumer action necessary | 1400 | 1400 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Open |  |  | 1400 | Software |  | Instructions | Safety Notice/Insructions
Z-2608-2011 | 58906 | culture media, antimicrobial susceptibility test, mueller hinton agar/broth | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL" OXACILLIN SCREEN AGAR, catalog #221952, box of 10 plates.Oxacillin Screen Agar (originally named MRSA Screen Agar) was developed for the detection of methicillin-resistant Staphylococcus aureus (MRSA). These strains are resistant to penicillinase-resistant penicillins (PRPs), such as methicillin, oxacillin and nafcillin. Since the method to detect MRSA uses the same inoculum as the Bauer-Kirby antimicrobial disc susceptibility test procedure, the oxacillin screen test may be conveniently performed on isolates at the same time as routine susceptibility testing. | BBL" OXACILLIN SCREEN AGAR | 2 | June 21, 2011 | 2011 | Becton Dickinson & Co. | Oxacillin Screen Agar may not have been produced according to specification and could result in false antibiotic susceptibility test results in patient specimens. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2011 to all affected customers. The letter identifies the product, problem and actions to be taken by the customers.The letter instructed customers to discontinue use of the affected lot number and discard any remaining packages. Customers were requested to complete and return an attached response form via fax at 410-316-4258. The letter states that all discarded product will be replaced.Questions concerning the recall process were directed to BD Customer Service Department at 1-800-675-0908. |  | 5440 | 5440 plates | Worldwide Distribution--USA (nationwide) and the countries of Canada, Columbia, Jamaica, South Korea, and Taiwan. | Terminated | August 15, 2011 | 55 | 5440 | Not_Computer
Z-2601-2011 | 59000 | controller, infusion, intravascular, electronic | LDR |  |  | General Hospital | 510(k) | "***DOSI-FLOW 1 I.V. EXTENSION SET WITH I.V. FLOW REGULATOR REGULADOR DE CAUDAL PARA SETS DE INFUSION I.V.***REF 259-0250***50 UNITS***STERILE OE***LATEX-FREE***LEVENTON LEVENTON, S.A.U. C/Newton, 18-24 08635 Sant Esteve Sesrovires BARCELONA-SPAIN***Distributed by: Victus www.victus.com 4918 S.W. 74th Court Miami, FL 33155***" To control intravascular infusion of medication. | DOSIFLOW 1 I.V. EXTENSION SET | 2 | June 21, 2011 | 2011 | Leventon S. A. U. | Leventon, S.A.U. Barcelona, Spain is recalling DOSI-FLOW REGULATOR, Part Number 259-0250, Lot Numbers 101444L and 101494L. Product has potential to leak. | DESIGN: Device Design | The firm, Leventon, S.A.U. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to the affected customer. The letter identifies the product, problem, and actions to be taken by the customer.The letter instructs the customer to examine their inventory and quarantine any product subject to the recall. They were also instructed to identify their customers who have received the product, notify them of the recall, send a copy of the recall notification letter and request that the customers return any unused product. The customers were also instructed to complete and return the attached RECALL RESPONSE FORM within 5 working days via fax to 00 34 93 6532556 or e-mail to dsalvatierra@leventon.es.Questions concerning this recall should be directed to the QA Manager at 00 34 93 8176316 or e- mail to dsalvatierra@leventon.es. | Contact the recalling firm for information | 1000 | 1000 cases (50,000 units) | Nationwide Distribution: FL only. | Terminated | June 28, 2011 | 7 | 1000 | Not_Computer
Z-2606-2011 | 58833 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay SystemsQuality Control Kits, Part Number: 37309The Access Ostase QC is used for monitoring system performance of the Access Ostase assay. | Access Ostase QC | 2 | June 21, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that Access Ostase Calibrator and QC vials from the affected lots have shown an increased rate of fracturing and may leak the contents in the box during shipment to customers.Operator injury may occur due to contact with broken glass, exposure to infectious agents, and/or exposure by skin contact to agents that may cause sensitizatio | DESIGN: Packaging Design/Selection | Beckman Coulter sent an PRODUCT CORRECTIVE ACTION (PCA) letter on May 10, 2011 to all affected customers who purchased Access Ostase Calibrators and Access Ostase QC. The letter identified the customers with an explanation of the problem and the potential impact should the problem occur and the action needed to be taken. Customer were instructed to:(1) Inspect all Access Ostase Calibrator and QC vials and packaging for signs of failure prior to use, including:-Cracking, particularly along the bottom of the vial,-Absence of reagent and/or leaking of contents into the box(2) Human source material used in the preparation of Access Ostase Calibrator and QC material has been tested and found non-reactive for Hepatitis B (HBV), Hepatitis C (HCV), and Human Immunodeficiency Virus (HIV-1 and HIV-2). No known test method can offer complete assurance that infectious agents are absent. Access Ostase Calibrator and QC vials should be handled in accordance with Good Laboratory Practices, and should be handled as if capable of transmitting infectious disease.(3) Skin contact with Access Ostase Calibrator and QC material may cause sensitization. Wear suitable gloves when handling vials. If contact with skin occurs, wash immediately with plenty of soap and water.Customers were told to share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected products listed above to another laboratory, please provide a copy of the letter to them.In addition, customers were requested to complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For replacement of defective vials:(1) In the United States, please contact Client Services at 1-800-526-3821, option 1.(2) In Canada, please contact Customer Service at 1-800-463-7828.(3) Outside of the United States and Canada, contac. | Contact the recalling firm for information | 702 | 702 kits | Nationwide in US and worldwide to: Algeria,Argentina,Australia,Austria,Belgium,Canada,China,Colombia,Czech Republic,France,Georgia,Germany,Greece,Hong Kong,Hungary,India,Italy,Japan,Jordan,Lebanon,Liechtenstein,Mexico,Morocco,New Zealand,Panama,Reunion,Romania,Russian Federation,Saudi Arabia,Serbia,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | July 11, 2012 | 386 | 891 | Not_Computer
Z-2570-2011 | 58896 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHECK PERFORMA STRIP NO. 05987288031, 100 BANDELETTES REACTIVES, ROCHE DIAGNOSTICS GMBH, WWW.ACCUCHEK.COM glucose monitoring test strips. | ACCUCHECK PERFORMA STRIP | 1 | June 21, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Erroneous results: Internal investigations revealed that the test strips exhibited a performance issue that can cause test strip errors or a bias in test result that exceed the tolerance range according to EN ISO 15197. | DESIGN: Component Design/Selection | Direct accounts were notified by a Roche Diagnostics press release on 5/18/2011 that included the product's name, lot number, and information to withdraw the remaining inventory of lot 470049 from storage. |  | 27019 | 27,019 units | No US distribution, in France Only. | Terminated | October 25, 2011 | 126 | 27019 | Not_Computer
Z-2607-2011 | 56787 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Eon Mini Neurostimulation (IPG) System Model 3788System is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. | Eon Mini Neurostimulation Sysem | 2 | June 22, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | Defective battery was identified as cause for reports of inability to communicate or recharge the product. | PRODUCTION CONTROLS: Process Control | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers.St. Jude Medical recommends no action be taken if the product serial numbers do not match the numbers listed in the letter. For product remaining in unimplanted inventory, customers are instructed to contact their SJM representative so that product may be returned and replacements be provided. For implanted product that matches the serial numbers, the Medical Advisory Board recommend not to unnecessarily explant the devices if they are functioning as intended. Replacement devices will be provided at no additional cost to the customer. A Product Reconciliation Form was attached to the letter for customers to complete and return any unused product to the St. Jude Medical Neuromodulation Division at 6901 Preston Road., Plano, TX 75024. Questions regarding this recall are directed to the customer's St. Jude Medical Neuromodulation Division Representative. |  | 17141 | 17,141 | Worldwide Distribution-- USA (nationwide) including PUERTO RICO and countries of AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, FINLAND,FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA,SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Open |  |  | 17141 | Battery | Device Operation | Replace | Remove or Replace
Z-2575-2011 | 58995 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Integra Pain Management Disposable, Convenience Kit PICC INSERTION TRAY, Reorder Number 3403203, Sterile/EO, Integra LifeSciences Corporation, West Valley City, Utah 84119. | Integra PICC INSERTION TRAY | 1 | June 22, 2011 | 2011 | Integra LifeSciences Corp. | Kits contain protective wipes that may be contaminated with Bacillus cereus. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra LifeSciences consignees were notified by Urgent Medical Device Recall letter on May 20, 2011 and told to quarantine any remaining product and to notify customers to return any remaining product to them. Further information is available at (800) 241-2210. | Contact the recalling firm for information | 520 | 520 kits | UT | Terminated | August 01, 2011 | 40 | 560 | Not_Computer
Z-2610-2011 | 58779 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI TM 101 Monitor, CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System, The CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where you need continuous monitoring of the blood's oxygen saturation, hematocrit and hemoglobin levels. | CDITM 101 Monitor | 2 | June 22, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that the presence of electromagnetic interference (EMI) in the operating room may affect the accuracy of the valuesdisplayed by the CDI"101 Hematocrit/Oxygen Saturation Monitoring System: hematocrit (HCT), hemoglobin (Hgb) and Oxygen Saturation. | DESIGN: Device Design | Terumo Cardiovascular Systems, Corp issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 28, 2011, to all affected customers. The letter identifed the product, the problem and the action needed to be taken by the customer. After a solution has been verified, Terumo CVS will implement a correction for all affected units. Customers were instructed to review this Medical Device Correction notice, Assure that all users are aware of this notice. Confirm receipt of this and to fax back the Customer Response form.For further question please contact Terumo CVS Customer Service at 1-800-521-2818, Monday - Friday, 8 AM - 6 PM EST. | Contact the recalling firm for information | 622 | 622 | Worldwide Distribution -- USA ( nationwide) and the countries of Colombia, Vietnam, India, Indonesia, Pakistan, Canada, Chile, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Singapore, Taiwan, Belgium, Thailand, Vietnam and India | Terminated | November 06, 2012 | 503 | 622 | Other | Output/Calculation |  | Other
Z-2625-2011 | 58822 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | 12-lead ECG Cable Assembly contains the following cable and attachments:Main Cable with Limb Leads: The main cable includes the proximal connector which attaches to the monitor and four incorporated limb lead wires that terminate with ECG electrode snaps. All signals travel through this cable for cardiac monitoring and for 12-lead ECGs.Precordial Lead Attachment: The precordial lead attachment includes the precordial connector which attaches to the main cable with six precordial leads (also terminated with ECG electrode snaps). This attachment is used when a 12-lead ECG is needed. The main cable includes a cover over the connector for the precordial lead attachment to protect it from contaminants if the precordial lead attachment is notconnected.This 12-lead ECG Cable Assembly is designed for use in the Pre-hospital and hospital environment. The 12-lead ECG cables are currently available for the LIFEPAK 12 Defibrillator/Monitor and LIFEPAK 15 Monitor/Defibrillator. The 12-lead electrocardiogram is used to identify, diagnose and treat patients with cardiac disorders and is useful in the early detection and prompt treatment of patients with acute ST elevation myocardial infarction (STEMI). Documentation of transient or intermittent arrhythmias and other electrophysiologic events that occur in the prehospital setting can assist in diagnosis and treatment decisions in the Emergency Department. | 12Lead ECG Cable | 3 | June 22, 2011 | 2011 | Physio Control, Inc. | The 12-lead ECG cables have experienced premature failures which cause the shielding and conductors to break internally creating ECG noise or "lead off" indications. | PRODUCTION CONTROLS: Storage | Physio-Control, Inc. sent an "URGENT-MEDICAL DEVICE ACCESSORY INFORMATION" letter dated May 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The customers were advised to follow the instructions:-Use standard hard paddles or therapy electrodes for ECG monitoring if monitoring cannot be performed using the ECG cable. Refer to Troubleshooting Tips for ECG Monitoring in the LIFEPAK 12or LIFEPAK 15 Operating Instructions-Monitoring chapter.-Disconnect the ECG cable from the monitor after use and coil it loosely for storage. Please refer to the Storage Instructions for 12-lead ECG Cable.-Inspect cables daily for bulges which may indicate damaged internal wires. If bulges are visible, replace the ECG cable.-Perform periodic 6 month function testing: Refer to Maintenance and Testing Schedule in the LIFEPAK 12 or LIFEPAK 15 Operating Instructions- Maintaining the Equipment chapters for instructions on how to perform a patient ECG cable function check to confirm proper operation ofthe cable.Customers can visit Physio-Control's website at www.physio-control-notices.com/12LeadECG for additional information or todownload copies of current LIFEPAK 12 defibrillator/monitor or LIFEPAK 15 monitor/defibrillator Operating Instructions and Operator's checklists.Customers can also call Physio-Control's Technical Support Team at 1-800-442-1142, option 5, from 6:00 A.M. to 4:00 P.M. (Pacific), Monday-Friday. | Contact the recalling firm for information | 21825 | 21,825 units distributed worldwide | Worldwide Distribution--USA (nationwide) including Puerto Rico and countries of Australia, Brazil, Chile, Germany, Hong Kong, Mexico, The Netherlands, and New Caledonia. | Terminated | August 12, 2011 | 51 | 21825 | I/O
Z-2612-2011 | 59072 | sterilization wrap containers, trays, cassettes & other accessories | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | 45cm Laparoscopic Instruments Sterilization tray.Manufactured by: Stryker Endoscopy, 5900 Optical Court,San Jose, CA 95138Designed to hold laparoscopic surgical manual instruments which are intended for cutting, holding, grasping and manipulating tissue, organs, organ area, surgical auxiliaries such as suturing material. The 45cm Lap tray can hold up to 11 instruments in the top tier and a combination of instruments in the bottom tier provided that the total weight of the loaded tray does not exceed 6.53 kg. | 45cm Laparoscopic Instruments Sterilization tray | 2 | June 22, 2011 | 2011 | Stryker Endoscopy | During testing and review, Stryker Endoscopy discovered that the sterilization parameters for the 132 degree C gravity steam sterilization cycle in the IFU may not consistently sterilize all instruments in the 45cm Laparoscopic Instruments Tray. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker Endoscopy sent a "URGENT: DEVICE CORRECTION" letter dated June 8, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to discard all old Instructions for Use sheets for their 45cm Laparoscopic Tray and download the new IFU from the URL link:http://stryker.com/45cm tray. Additionally, customers were instructed to complete and return an attached "acknowledgement of receipt" form and fax to 408-754-8378 or scan and e-mail to trays@stryker.com.Questions can be answered by calling 1-800-624-4422 or by e-mail at trays@stryker.com. |  | 603 | 603 | Worldwide Distribution--USA (nationwide) and countries of Brazil, Canada, Chili, Germany , Korea, Mexico, Philippines, Poland, Portugal, Singapore, Sweden, Switzerland, and Taiwan. | Terminated | October 25, 2011 | 125 | 603 | Not_Computer
Z-2609-2011 | 58767 | defibrillator, implantable, dual-chamber | MRM |  |  | Cardiovascular |  | Paradym DR Model 8550 and Paradym VR Model 8250, Sterile EO, Manufactured in Italy by Sorin CRM S. r.l., Via Crescentino s.n., 13040 Saluggia(VC) ItalyPARADYM DR 8550 and PARADYM VR 8250 are indicated for use in patients who are at high risk of sudden cardiac death due to ventricular tachyarrhythmias and who have experienced one of the following situations: Survival of at least one episode of cardiac arrest (manifested by the loss of consciousness) due to ventricular tachyarrhythmia, recurrent, poorly tolerated sustained ventricular tachycardia (VT). | Paradym DR Model 8550 and Paradym VR Model 8250 | 2 | June 22, 2011 | 2011 | ela Medical Llc | On February 28, 2011, the company discovered that since the beginning of the US commercial distribution of Paradym DR 8550 and Paradym VR 8250 ICDs in October 2010, the devices included an incorrect version of the Patient Booklet.The company determined that there is no risk of adverse event associated with this labeling error. | PRODUCTION CONTROLS: Error in Labeling | Sorin CRM USA instructed its field representatives on April 27, 2011 via a conference call to attach the correct Patient Booklet (ref. N582,) as an addendum to the device's external packaging, along with a cover letter "Quality Note" dated May 2, 2011 explaining the discrepancy. The field representatives were fedexed the information on May 2, 2011.The letter instructed the Field Representative to attach the correct patient booklet and an addendum, to the outside of the device box.Field Representatives were asked to complete the acknowledgement form by identifying S/N's in which they attached the pouch and their signature of completion and fax the form to 866-500-6096 or e-mail a scanned copy to sorin.ptusa.sorin.com.For questions regarding this recall call 877-663-7674." For affected units that were already implanted as of April 27, 2011, the patient booklet given to the patient may include information written by the physician or hospital staff that is related to the implanted system (ICD and leads), physician references and patient medications. For this reason and given the fact that the company determined that there is no risk for the patient associated with this labeling error, the company decided to not replace the patient booklet for this population. | Contact the recalling firm for information | 275 | 275 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, FL, GA, IA, MA, MI, MO, MS, NJ, NY, NC, OH, OK, PA, SC,A TN,T X, and WA. | Terminated | April 26, 2012 | 309 | 275 | Not_Computer
Z-2624-2011 | 58877 | implant, eye sphere | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Product Brand Name: Coated Bio-Eye Orbital Implant w/Conformer.Product Generic Name: Implant, Eye Sphere Wrap, Implant, Orbital.Model Number: I0016C, I0018C, I0020C, I0022C, I0024C.Description of the product: The Coated Bio-Eye consists of a coralline hydroxyapatite sphere encased in a resorbable shell that aids inimplant insertion and provides a means of securing muscles to the implant.The Coated Bio-Eye Orbital Implant is indicated in orbital implantation following enucleation, or as a secondary orbital implant following extrusion, migration or rotation of primary orbital implants. It is indicated in any situation where materials such as silicone, acrylic, polyethylene or glass orbital implants would be used. | Coated BioEye Orbital Implant w/Conformer. | 2 | June 22, 2011 | 2011 | Integrated Orbital Implants Inc | The recall was initiated because Integrated Orbital Implants has confirmed certain lots of the Coated Bio-Eye were manufactured without expiration dating visible on the packaging. All unlabeled product is currently beyond its expiration date. | OTHER/UNDETERMINED: Pending | Integrated Orbital Implants Inc initiated a recall communication on 5/20/2011 with Integrated Orbital Implants (IOI) forwarding a Urgent Safety Notification letter to all their customers who had purchased the Coated Bio-Eye implants. The letter asked the customers to look for evidence of a printed expiration date on Coated Bio-eye package. If any Coated BioEye packages were not labeled with an expiration date the letter instructed customers to immediately sequester the products and not to use the products in surgery and contact IOI by telephone at (858) 6779990 or 1 (800) 4246537 to request an RMA number and shipping instructions.Also stated in the letter, IOI will replace Coated BioEye implants without expiration dates free of charge upon receipt of the implants labeled withoutexpiration dates. | Contact the recalling firm for information | 4237 | 4237 | Worldwide Distribution - USA (nationwide) and the countries of nationwide and worldwide to: Japan, Greece, Germany, Romania, UAE, Australia, Canada, Philippines, Korea, Taiwan, Chile, Denmark, UK, Italy, Pakistan, Turkey, Philippines, India, South Africa, and Singapore. | Terminated | January 11, 2012 | 203 | 4237 | Not_Computer
Z-2623-2011 | 58834 | stopcock, i.v. set | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Large Bore 4-Way Stopcock, Volume 0.65 mL, Injection Site, Rotating Male Luer Lock; a sterile, nonpyrogenic fluid pathway; product code 2N5601. Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. | Baxter Interlink System Large Bore 4Way Stopcock | 2 | June 22, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for pinholes and tears in the individual packaging of the stopcocks, compromising the sterility of the product. | DESIGN: Packaging Design/Selection | Baxter Urgent Product Recall letters dated May 17, 2011 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that potential pinholes and tears in the individual packaging could compromise the sterility of the stopcock. The accounts were requested to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/17/11, and instructed to notify their customers and ministries of health of the recall. |  | 4800 | 4,800 units | California and internationally to Japan and Australia | Terminated | June 18, 2012 | 362 | 4800 | Not_Computer
Z-2611-2011 | 59004 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Pathfinder Holter Analyzer is a Programmable Diagnostic Computer. The brochure is labeled in part: "***Pathfinder Holter Analyzer ***Spacelabs Healthcare Company Headquarters***5150 220th Ave. SE, Issaquah, WA 98029***PO Box 7018, Issaquah, WA 98027-7018***Telephone: 425-657-7200***In North America: 800-522-7025***www.spacelabshealthcare.com***".The Reynolds Pathfinder 700 Holter Analyzer is intended to be used to analyze magnetic tapes of ambulatory electrocardiograms made on compatible Holter Recorders. The system will detect various arrhythmias and will measure ST elevation or depression. The system acts in both a automatic mode (tape analyzed without operator intervention) and in an interactive mode; (operator can intervene and affect analysis). | Pathfinder Holter Analyzer | 2 | June 22, 2011 | 2011 | Spacelabs Healthcare, Llc | A potential hazard associated with the use of Spacelabs Healthcare Pathfinder Holter Analyzer software. The second 24 hours of a 48 hours EVO recording will not be analyzed. | DESIGN: Software Design | Spacelabs Healthcare sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 3, 2011 to all customers. The letter describes the product, problem, and the actions to be taken by the customers. Spacelabs Healthcare recommended that until software upgrades can be done, the customers should weigh the benefits versus the risks when deciding whether or not to continue using the product. Spacelabs Healthcare also recommended that customers advise their staff that the second 24 hours of a 48 hour EVO recording would not be analyzed.Field service engineers will be contacting customers to set up a time and date to install the revised version of the Pathfinder software. Spacelabs will upgrade Pathfinder software regardless of whether the consignees are using EVO recorders to ensure this hazard does not happen in the future.Consignees can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the US, 44 1992 507740 for any questions about this recall. |  | 58 | 58 units total (5 units in US and 53 units internationally) | Worldwide Distribution--USA (nationwide) including states of CA and NC and countries of Czech Republic, Denmark, Hong Kong, Israel, The Netherlands, and United Kingdom. | Terminated | October 26, 2011 | 126 | 58 | Software | Output/Calculation | Software update | Software Update
Z-2616-2011 | 58838 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit Part number 70367-07 Mission St. Joe's - Asheville X-Coated Low Prime Pack, Sterile, Terumo Cardiovascular Systems, Ashland, MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | Pressure Monitoring Kit | 2 | June 22, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process. | DESIGN: Component Design/Selection | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 3, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were sent 10 replacement pressure monitoring kits with the removal notice as immediate replacements. The firm will send additional replacements to customers that have more than 10 units of affected product as noted on the Customers' returned reponse form. Customers that received the custom packs were provided stickers to identify the recalled packs. The customers were instructed not to open the packs to remove the affected pressure monitoring kit prior to clinical use, but to replace the pressure domes at the time of use.For more information please contact Terumo Cardiovascular Customer Service at 800- 521-2818. |  | 1515 | 1515 all custom packs | Worldwide Distribution -- USA (nationwide )including the states AR, CA, FL, GA, MO, MT, NC, ND, OK, OR, PA, TN, TX, UT, VA, WA and WY, and the countries of Argentina, Belgium, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Guatamala, Honduras, Hong Kong, Indonesia, Kuwait, Malaysia, Mexico, Pakistan, Panama, Philippines, Singapore, South Korea, Thailand, United Arab Emirates, Uruguay and Vietnam | Terminated | July 18, 2012 | 392 | 21010 | Not_Computer
Z-2581-2011 | 58854 | blood pressure cuff | DXQ | Blood pressure cuff. | Cardiovascular | Cardiovascular | 510(k) | Cuffable Blood Pressure Cuffs; Manufactured for Vital Signs, Inc., a GE Healthcare Company 20 Campus Road, Totowa, NJ 07512 Made in Mexico.The cuff is used to determine a subject's blood pressure. | Cuffable Blood Pressure Cuffs | 2 | June 22, 2011 | 2011 | GE Healthcare, LLC | BP cuffs may not properly inflate due to a leak | DESIGN: Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 26, 2011 to all affected customers. The letter identified the product, the problem and the action to be taken. Customers were instructed to discontinue use of the product, isolate all affected products, return only unused products and distribute this letter to any other healthcare institutions that are potentially affected by this issue. The letter asks that each customer complete and return the enclosed confirmation form via fax to 800-535-7923, in order to confirm receipt of the notification letter. For any questions, contact your local Vital Signs Account Manager or the International Customer Service at 1-800-932-0760. |  | 2684996 | 2,684,996 units | Worldwide Distribution -- USA (nationwide) including Puerto Rico and countries of: Argentina, Australia, Belgium, Canada, Chile, Ecuador, Finland, Germany, Hong Kong, Israel, Italy, Latvia, Lebanon, Nicaragua, Panama, Peru, Saudi Arabia, Slovenia, UK, Uruguay and Venezuela | Terminated | May 28, 2013 | 706 | 2684996 | Not_Computer
Z-2633-2011 | 55276 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tubing Assembly #142, V2, V12 Mixer/Wash a component ofthe DxC instruments.Tube, Assembly, #142 V2, V12 Mixer/Wash delivers appropriate amount of wash concentrate on the UniCel DxC SYNCHRON Clinical Systems. The UniCel DxC Synchron Clinical Systems are fully automated, computer-controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent) | Beckman Unicel DxC 600 and DxC 800 Clinical Systems | 2 | June 23, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the Wash Concentrate Tube Assembly 142 may fail to deliver the appropriate amount of wash concentrate on some SYNCHRON DxC Systems due to an accumulation of crystals.The impact of this issue would include:(1) Quality Control and patient results can potentially be affected.(2) If system has been calibrating properly and Quality Control is recovering within es |  | Beckman Coulter sent out an "Urgent Product Corrective Action" letter dated July 23, 2009 and mailed the week of July 27, 2009 to all affected customers. The letter provides the cusotmers with explanation of the problem identifed and asks them to take the following actions:" If consignees are not able to resolve these issues through routine troubleshooting they were told to contact BCI Customer Technical Support at (800) 854-3633 in the United States and Canada." This issue will be corrected with an upcoming software version and tubing modification." Complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain their records.Consignees were also asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation.Questions regarding this Product Corrective Action letter were directed to their Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada. | No consumer action necessary | 2392 | 2392 units | Worldwide Distribution - USA (Nationwide) and Canada | Terminated | July 02, 2012 | 375 | 2392 | Computer
Z-2622-2011 | 58228 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Interplant Radiation Treatment Planning System, Interplant Release 3.4.0Used when planning and performing prostate brachytherapy implant procedure. It provides ultrasound probe alignment and radioactive seed implantation. | Interplant | 2 | June 23, 2011 | 2011 | Computerized Medical Systems Inc | Interplant Software: If the trajectory needle with a non-zero retraction is adjusted in any of the planning views (Transverse, 3D) and the program is closed and reopened, the Needle Definitions page in the Dose Plan Summary (Plan View) displays the needle with a 0.00 retraction even when the needle is displayed correctly in the planning view. If the Eliminate Leading Spacers function is applied t | DESIGN: Software Design | Computerized Medical Systems, Inc. sent a "CUSTOMER ADVISORY" letter dated December 1, 2008, to all affected customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to discontinue use of the product. There is no workaround at this time. This issue was resolved in Interplant Release 3.4.1 software. All possibly affected users have been shipped the Interplant Release 3.4.1 software in May 2010. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 41 | 41 | Worldwide Distribution -- USA (nationwide) including states of: CA, KS, KY, MA, MD, ME, MO, NC, NE, NY, OH, PA, TX and VA and countries of: Australia, Brazil, Chile, China, Germany and Japan. | Terminated | April 04, 2013 | 651 | 41 | Software | Display/Image | Software update | Software Update
Z-2627-2011 | 58873 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BACTEC" FX Instrument-Top, catalog #441385. The BD BACTEC" FX instrument is designed for the rapid detection of bacteria and fungi in clinical specimens. Samples are drawn from patients and injected directly into BACTEC culture vials, which are placed into the instrument for incubation and testing. | BACTEC" FX InstrumentTop | 2 | June 23, 2011 | 2011 | Becton Dickinson & Co. | In vitro diagnostic instrumentation may exhibit random operational errors that could delay effective patient treatment. | DESIGN: Software Design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated August 2010 to all affected customers. The letter describes the product, problem and action to be taken.The letter described the type of error messages the customers may experience. All affected units will be scheduled for on site software update as well as a disk drive replacement. A BD field service engineer will contact the customers to schedule the on site visit.Customers experiencing issues prior to the scheduled visit or have questions should contact the BD Diagnostics Service Department at 800-638-8663. |  | 907 | 907units | Worldwide Distribution--USA (nationwide) including countries of Australia, Belgium, Canada, China, Colombia, New Zealand, Hong Kong, Japan, Korea, Philippines, Singapore, Taiwan, and Thailand. | Terminated | July 05, 2011 | 12 | 907 | Software
Z-2629-2011 | 58793 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-s Ventilator Systems; Article number 66 40 440.Intended for treatment and monitoring of patients in the range of neonates, infants and adults with respiratory insufficiency; these are only to be used by healthcare providers in hospitals or healthcare facilities and for in-hospital transport. These are not to be used with any anesthetic agents. | Servos Ventilator Systems | 2 | June 23, 2011 | 2011 | Maquet Inc. | Through customer complaints, Maquet has been informed that SERVO-i ventilators have stopped to ventilate and have generated three different technical error codes. | OTHER/UNDETERMINED: Pending | Maquet Inc. sent an URGENT DEVICE SAFETY ALERT letter dated April 20, 2011, to all affected customers via Federal Express. The letter identified the product, the problem and the action needed to be taken by the customer.The letter informs the user of the issue and provides additional instruction and how the existing manual should be updated. Customers are asked to fill out the attached response form and fax it to 1-973-807-9210. The following warning will be included in all user's manuals delivered with new ventilators:'Warning ! The Servo-i Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.Warning! The Servo-s Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.'Additionally the warning stated above will be included in the format of a sticker with the yearly preventive maintenance kit for Servo-i and SErvo-s to update exiting User's manuals for the installed base of Servo-i and SErvo-s ventilators. A maintenance report form will also be included in the kit which serves as an acknowledgement from the user that the sticker has been added to an affected ventilator unit's operating manual.For further questions, please call (973) 709-7660. | Contact the recalling firm for information | 17881 | 17039 units worldwide; 842 units US | Worldwide Distribution - USA (nationwide) | Terminated | September 17, 2013 | 817 | 61091 | Not_Computer
Z-2636-2011 | 58813 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Mevatron K2, camera imaging system, Part Number: 1940753;Product is manufactured and distributed by SIEMENS MEDICAL SOLUTIONS USA, INC. Concord, CA; and SIEMENS AG, MEDICAL SOLUTIONS Roentgenstrasse, Kemnath, GERMANY, and SIEMENS AG, MEDICAL SOLUTIONS Henkestrasse, Erlangen, GERMANY The intended use of the SIEMENS branded linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | camera imaging system | 2 | June 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc - Radiation Oncology | The firm has become aware of a potential issue that may result in a service engineer receiving an electrical shock should the connector cable housing become unfastened from the chassis. | OTHER/UNDETERMINED: Pending | The firm, SIEMENS, sent a "Customer Information" letter to its customers on November 4, 2010. The letter describes the product, problem and actions to be taken. The customers were instructed to include this letter in their BEAMVIEW TI System Owner Manual chapter "Safety Advisory Letters" where it should remain. Siemens has developed an updated kit to install an additional grounding cable on the camera's central control unit to eliminate the problem. The updated kit was delivered with the Customer Information letter.If you have any questions, contact the Service Manager- Regulatory Affairs at 925-246-8407. |  |  |  | Worldwide distribution: | Terminated | November 14, 2011 | 144 |  | Not_Computer
Z-2645-2011 | 58206 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ERGO++ Stereotactic Radiation Treatment Planning System, ERGO++ Release 1.6.3 and 1.6.3.1Product Usage: Used to create treatment plans for any cancer patient for who external beam radiation therapy has been prescribed. It is an accessory to linear accelerators used for radiation therapy. It is indicated for use in the planning of 3-dimensional radiation therapy. | ERGO | 2 | June 23, 2011 | 2011 | Computerized Medical Systems Inc | ERGO Release 1.6.3 is overestimating the MU values. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent an "USER NOTIFICATION" dated December 11, 2009, to all affected customers. The notice identified the product, problem and suggested a workaround to minimize the risk. The notice stated that this workaround does NOT substitute the need for regular Quality Assurance. Customers are instructed to review page 39 of user manual for further explanation of warnings. Customers were asked to complete the Notice for Customer section and return to 3D Line Medical System. This issue has been resolved in Release 1.7.0 and above. No follow-up notice will be distributed to customers. For further information please contact: CMS Software (314) 812-4460 (St Louis Office). | Contact the recalling firm for information | 45 | 45 | Nationwide distribution to medical facilities in CA, MO, NC, ND, NV, NY, OH, OK, TX, and WA. Foreign distribution was made to Brazil, Cuba, Czech Republic, France, Germany, India, Italy, Japan, Malaysia, Poland, Portugal, Serbia, and Spain. | Open |  |  | 45 | Software | Output/Calculation | Software update | Software Update
Z-2630-2011 | 58859 | [6 empty] | 1500T12 Cool Path Ablation System - GeneratorThe generator is a microprocessor-controlled device that produces a continuous unmodulated radiofrequency (RF) output of 485 kHz. The generator has a maximum power output of 100 Watts. The catheter delivers the RF power from the generator in a monopolar mode between its distal electrode (tip electrode) and a large indifferent electrode (dispersive pad or "DIP"), which is in compliance with ANSUAAMI Standard HF-18-60601-2-2. The Generator operates in Temperature Control mode only. The generator is a temperature controlled system, where temperature measured by the temperature sensor in the Therapy 1300 Series Steerable Ablation Catheter or other compatible catheters is monitored and the power delivered by the generator adjusts within the selected limits until the desired temperature is achieved. The generator's thermocouple temperature control circuitry is based on a Type T thermocouple. The front panel displays the actual real-time power output, impedance, duration, and measured tip electrode temperature. The amount and duration of RF power delivery is user selectable, The desired catheter tip electrode temperature is also user selectable. A low pass filter is enabled during ablation to permit recording of intracardiac electrograms and alternately disables the filter during stimulation (pacing). The output power of the Generator will shut off if the measured temperature exceeds 80ÂŽC or if the measured temperature exceeds the user selected temperature set point by more than 5ÂŽC for more than 3 seconds. The Generator has built-in safety features, which include a self test at power up and automatic RF power shut off if the measured tissue impedance falls below 50 Ohms or exceeds 300 Ohms or the preset impedance value for more than 2 seconds. The 1500Tl2 Cardiac Ablation Generator is only compatible with the Cool Point Irrigation Pump. | 1500T12 Cool Path Ablation System Generator | 2 | June 23, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed its T14 and T12 RF Generatormay malfunction upon encountering an unexpected electric current (e.g. static electricity).The resulting malfunction results in a failure of the generator to respond to key pressesproperly, and may cause unexpected energy delivery or inability to discontinue energydelivery with the start / stop butt | OTHER/UNDETERMINED: Pending | The firm, St. Jude Medical (SJM), sent a "Product Correction" letter dated May 9, 2011 to it's affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the T14/T12 generators and remote controls until the units have been properly serviced; complete and return the attached field action form, and utilize the units in accordance with the Operators Manual. In case of system malfunction, the customers were instructed to shut off the RF power delivery by turning off the generator via the rocker switch, dislodging the power cord, or disconnecting of the patient cable or indifferent electrode connection. An SJM sales representative will contact the customers in the near future to facilitate repair or exchange of the impacted unit.If you have questions regarding this notice, contact the Marketing Manager, AF at (905-286-4015), Marketing Director, AF/CRM, at (905-286-4107); Product Manager (Tokyo) 03-6255-5740, and/or your SJM sales representative. |  | 154 | 154 | International distribution only: Canada and Japan. | Terminated | April 03, 2012 | 285 | 168 | Not_Computer
Z-2634-2011 | 58887 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***Processed at: Regeneration TECHNOLOGIES, INC. Alachua, FL, USA 32615***Distributed by: SpinalGraft Technologies, LLC 4340 Swinnea Road STE 39 Memphis, TN 38118***OSTEOFIL DBM Paste, 5cc Tissue ID: 321933 Reorder: 002405 Lot #(Donor ID): 101050353 Expiration: 02/21/2013***STERILE***"Bone Void Filler used in spinal surgery. | OSTEOFIL DBM Paste, 5cc | 2 | June 23, 2011 | 2011 | RTI Biologics, Inc. | RTI Biologics issued a recall of Osteofil DBM Paste, 5cc that was distributed on 03/04/2008 to one customer. The product was labeled as 5.0cc but should have been labeled as 0.5cc. All customers were notified of this recall on 04/09/2011. | PRODUCTION CONTROLS: Labeling Mix-Ups | RTI Biologics sent a "VOLUNTARY RECALL" letter dated April 9, 2008 to the one distributor, Spinal Graft Technologies. LLC. The letter describes the product, problem, and action to be taken by the distributor.The letter states that some of the product was issued a return authorization and requests that the distributor recover the remaining product from distribution. The recovered product will be issued a return authorization. RTI states the grafts can be returned in a non frozen state.Recall effectiveness will be determined by return of the product. |  | 26 | 26 units | Domestic Distribution--TN only. | Terminated | July 24, 2012 | 397 | 26 | Not_Computer
Z-2647-2011 | 57906 | mask, oxygen | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | First Breath ÂŽ Adult Oxygen Mask with Tubing, 213cm (7ft)Reorder Number: 001421Oxygen/Aerosol Mask For acute use, e.g. Post-op recovery, minimal trauma, patient transfer; used for patients requiring more oxygen than is delivered via a standard nasal cannula ( | Portex First Breath Adult Oxygen Mask with Tubing | 2 | June 24, 2011 | 2011 | Smiths Medical | A sticky substance on the inside and/ or outside of the Mask | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smiths Medical notified consignees by letter on 2/14/11 titled Urgent Medical Device Correction Notice via U.S. Mail (USA only) This Notice includes an Urgent Medical Device Correction Notice Confirmation Form that provides instructions to consignees to return unused affected product to Smiths Medical and to complete the effectivity tracking form. On 3/17/11 Smiths Medical issued an additional recall letter to extend the recall to 3 additional lots not included in the original recall letter dated 2/14/11. Also, the firm issued a recall letter on 3/17/11 to include the Value Pack (VP1302) General Anesthesia Kit to include 7 kit lots. Contact Recall Coordinator 1-800-348-6064 prompt 4, ext 2273. |  | 8652 | 8,652 units | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Armenia, Brazil, Canada, France, Hungary, Italy, Jordan, Japan, Libya, Poland, Russian Federation, and Saudi Arabia. | Open |  |  | 88653 | Not_Computer
Z-2657-2011 | 58828 | tube, tracheostomy (w/wo connector) | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone CTS (Close to Shaft) Cuff Adjustable Neck Flange Tracheostomy Tube (Wire Reinforced), Arcadia Medical Silicone Air Cuff Adjustable Tracheostomy Tube, Sterile, Latex-Free, Provides direct tracheal access for airway managementArcadia Medical Corporation1450 East American Lane, Schaumburg, IL 60173 USAArcadia Medical Corporation Tortola, British Virgin Islandswww.arcadiamedical.com, Made in USA | Silicone CTS Tracheostomy Tube | 2 | June 24, 2011 | 2011 | Arcadia Medical Corporation | The product does not contain a punched hole above the cuff access port (a lumen that enables suctioning of secretions proximal to the inflated cuff) as intended by the devices design. This eliminates the user's ability to suction secretions through this suction port, necessitating that the user revert to standard of care methods for suctioning secretions from around the cuffed trach tube shaft. | PRODUCTION CONTROLS: Process Control | Arcadia Medical Corporation sent a "MEDICAL DEVICE RECALL" letter on May 9, 2011 and email on May 8, 2011 to all affected costumers. The letter included; description of the product name, afffected lot numbers, description of the problem, instructions to discontinue use of the recalled devices and to return them to Arcadia Medical for replacement . The letter instructed the customers to email Arcadia Medical with the number of devices in their inventory. For additional information contact Arcadia Medical Corporation at 219-779-9431. |  | 118 | 118 | Worldwide Distribution --USA including the state of Ohio, and the countries of Denmark, Italy, Germany, Turkey, Slovakia, Isreal, France, Sweden, and Romania | Terminated | August 30, 2011 | 67 | 118 | Not_Computer
Z-2658-2011 | 58875 | dna probe, nucleic acid amplification, chlamydia | MKZ | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | BD ProbeTec" CT/GC/AC Reagent Pack, catalog #440450, labeled in part ***Becton Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152***The BD ProbeTec" ET Chlamydia trachomatis and Neisseria gonorrhoeae Amplified DNA Assays, when tested with the BD ProbeTec ET System, use Strand Displacement Amplification (SDA) technology for the direct, qualitative detection of Chlamydia trachomatis and Neisseria gonorrhoeae DNA in endocervical swabs, male urethral swabs, and in female and male urine specimens as evidence of infection with C. trachomatis, N. gonorrhoeae, or of co-infection with both C. trachomatis and N. gonorrhoeae. Specimens may be from symptomatic or asymptomatic females and males. A separate Amplification Control is an option for inhibition testing (BD ProbeTec ET CT/GC/AC Reagent Pack). The BD ProbeTec ET CT/GC assays may be performed using either the BD ProbeTec ET System or a combination of the BD ProbeTec ET System and BD Viper" instrument. | BD ProbeTec" CT/GC/AC Reagent Pack | 2 | June 24, 2011 | 2011 | Becton Dickinson & Co. | In vitro diagnostic reagent packs may be underfilled and contribute to invalid quality control testing or false positive patient results. | PRODUCTION CONTROLS: Process Control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated September 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to discontinue use of the product and discard any remaining product. BD will replace any discarded products. Additionally, customers were requested to complete and return an attached form whether or not they have any inventory remaining, to confirm receipt of the notification. The letter instructed the customers to fax the completed form to: Attention: Regulatory Compliance at 410-316-4258.For assistance regarding replacements, contact BD Customer Service Department at 1-800-675-0908. All other inquiries should contact BD Technical Services Department at 1-800-638-8663. |  | 130 | 130 kits | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, KY, LA, ME, MN, MO, MS, MT, NE, OH, PA, RI, TX, and WA and the countries of Australia, Canada, Europe, and Japan. | Terminated | June 24, 2011 | 0 | 130 | Not_Computer
Z-2663-2011 | 58442 | infusor, pressure, for i.v. bags | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | Infusable and Infusa Scan Single Patient Use Pressure Infusors; Vital Signs Inc.A pressure infusor for an IV bag is a device consisting of an inflatable cuff which is placed around and IV bag. When the device is inflated, it increases the pressure on the IV bag to assist the infusion of the fluid. | Infusable and Infusa Scan Single Patient Use Pressure Infusors | 2 | June 24, 2011 | 2011 | GE Healthcare, LLC | Infusable and InfusaScan Pressure Infusors may leak and not apply desired pressure to their IV solution bag and subsequently the rate of infusion may be inadequate. | DESIGN: Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 with an attached return response form to its customers via Fed Ex. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) DO NOT use infusable or InfusaScan pressure infusors (listed lot numbers).2) Isolate all affected product and contact Vital Signs Customer Service at +1-973-956-5300 to arrange for the return and replacement of the infusor.3) If they have forwarded any of the affected product to any other healthcare institution, forward a copy of this letter to those institutions.4)Return only unused product.In addition, the customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions, and to complete and return the attached confirmation form via fax to +1-973-956-5440.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 954888 | 954,888 units | Worldwide distribution: USA (nationwide) states including: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY; and countries of: Argentina, Australia, Chile, Colombia, Canada, Denmark, Dominican Republic, Finland, Germany, Greece, Hong Kong/Macao, India, Israel, Italy, Korea, Lebanon, Lativia, Mexico, Malaysia, Nicaragua, Netherlands, New Zealand, Paraguay, Philippines, Portugal Panama, Puerto Rico, Peru, Russia, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | October 31, 2012 | 495 | 954888 | Not_Computer
Z-2661-2011 | 59064 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | 0200086, Triad Alcohol Prep Pads, 100 box0200432, Alcohol Prep Pads, packaged in the Alere INRatio 2 PT/INR Home Monitoring Kit; orHemosense INRatio 2 Prothrombin Time/lNR Testing Kit, Self-Test System;0100007, Alcohol Prep Pads, packaged in the Hemosense INRatio Prothrombinme/lNR Testing Kit, Self-Test System;0100072, Alcohol Prep Pads, packaged in the INRatio Starter Kit | Triad Alcohol Prep Pads | 2 | June 24, 2011 | 2011 | Alere San Diego | The recall of the Triad Group alcohol prep products is due to potential contamination of these products with the bacteria Bacillus cereus, which could lead to life-threatening infections. | OTHER/UNDETERMINED: Pending | Alere sent a CUSTOMER NOTIFICATION letter on February 28, 2011 to all affected custumers.The letter included; description of product, affected lots, statement of potential for contamination which could lead to life-threatening infections. Alere asked its customers to discard the Triad Alcohol Prep Pads provided and complete and FAX the enclosed Verification Form within 10 days to confirm receipt of notice and to indicate the number of discarded/required replacement alcohol prep pads.The letter instructed customers to call Alere Technical Services for additional information at 888-246-7483. | Contact the recalling firm for information | 4829 | Estimated 4829 total units | Worldwide Distribution--(USA) Nationwide and countries of Canada, Australia, Japan and Puerto Rico | Terminated | February 29, 2012 | 250 | 4829 | Not_Computer
Z-2660-2011 | 58864 | prosthesis, elbow, constrained, cemented | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Condyle Kit with Hexalobular, REF 114700 DISCOVERY ELBOW SYSTEM, HUMERAL CONDYLE SET - HEXALOBULAR, CO-CR-MO / YI-GAL-4V ALLOY, IF USED FOR REVISION SURGERY, INSTRUMENTS 414926 & 414923 MAY BE REQUIRED LOT 183300, 1 SET, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE, P.O. BOX 587 WARSAW, IN 46581 USA EXPIRY DATE: 2021-03Elbow joint replacement prosthesis intended for primary and revision joint arthroplasty for use in cemented applications. | Discovery Elbow System with Humeral Condyle | 2 | June 24, 2011 | 2011 | Biomet, Inc. | The firm initiated this recall after becoming aware that a Discovery Condyle Kit with Hexalobula contained two male condyles instead of a male and female condyle. | PRODUCTION CONTROLS: Packaging Process Control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated May 25, 2011 to all affected customers. The notice describes the product, problem, and actions to be taken by the customers.The notice instructed customers to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Additionally, customers were instructed to ship the package priority mail with a copy of the response form.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday 8 AM to 5 PM. |  | 10 | 10 KITS | Worldwide Distribution--USA (nationwide) including the states of AR and SD and the country of Japan. | Terminated | August 29, 2011 | 66 | 10 | Not_Computer
Z-2659-2011 | 58770 | pack, hot or cold, water circulating | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precision Ice One-Patient-Use Control Unit and Wrap System (Model # 4BB02BK).The device is intended to reduce edema and promote healing of soft-tissue injuries and lymphedema through stimulation of the lymphatic system. The device is a therapeutic system which consists of a control unit and fluid circulating pads. The pads are mounted in wraps designed for specific parts of the body such as ankle, knee, elbow, and shoulder. The pads provide cold treatment and pressure which help promote healing.The owner's manual is labeled in part: "PRO TRAINERS' CHOICE COMPANY, INC.***www.precisionice.com***info@precisionice.com***Mailing Address PO Box 27***Kingston, WA 98346 USA***Phone: 360-297-3902***Shipping Address***26129 Calvary Lane NE, Suite 200***Kingston, WA 98346 USA***Fax: 360-297-3903".The sticker which is placed on the back of the control board is labeled in part: "MOD#4BB02BK***SER#100***".The shipping box is labeled in part: "Pro Trainers' Choice' LLC***Model #4BB02BK***Serial #100____***".The knee wrap is labeled in part: "***Pro Trainers' Choice Co. Inc./Precision Ice One Patient Use Knee Wrap***#1BK01051KN***This device is to be used for one patient only; do not re-use!***". | Precision Ice | 2 | June 24, 2011 | 2011 | Pro Trainers' Choice Company | The control panel of the Precision Ice One Patient Use Control Unit and Wrap System has a overheating issue. The control panel printed circuit boards were missing a jumper connection across resistors R28 A, B, C, and D. | OTHER/UNDETERMINED: Pending | Pro Trainers' Choice sent a Mandate-Medical Device Recall on April 18, 2011 and April 20, 2011 to all affected costumers.The letter included a description of the product, affected model and lots. It instructed the costumers to discontinue using the Precision Ice, Model 4BBK02BK until the device is returned and corrected by Pro Trainers' Choice. The device should be returned to PO Box 27, Kingston, WA 98346.The customers were instructed to notify their sub-account customers of this recall if the devices were further distributed.Customers can call Pro Trainers' Choice Company at 360-297-3902 or visit the website www.precisionice.com for any questions about this recall. | Contact the recalling firm for information | 336 | 336 units | Nationwide distribution including CA, CO, FL, IN, KS, NY, OR, TX, and WI. | Terminated | March 01, 2012 | 251 | 336 | Hardware | Physical Safety Hazards | Return | Remove or Replace
Z-2662-2011 | 58186 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify & Record System, for MOSAIQThe intended use of Sequencer (generically referred to as a verification system) is to assist in the process of process of patient treatment on a radiotherapy treatment machine. It gives the capability to notify users of actions that need to take place prior to treatment, to display reference images for set up purposes and to auto setup the machine to predefined settings. | Elekta Impac Software Sequencer | 2 | June 24, 2011 | 2011 | Impac Medical Systems Inc | There's a problem in the database conversion program that is used when upgrading from the Multi-ACCÂżS to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on, the patient will be mistreated. | DESIGN: Software Design (Manufacturing Process) | The firm, Elekta, sent a "USER NOTICE" to its customers with an attached confirmation of receipt form. The notice describes the product, problem and actions to be taken. The customers were instructed to follow the Interim mitigation instructions noted in the notice -verify the Arc start angle of the Simple ARC with MLC treatment field definition prior to delivery and if the customer identifes a field with an errant gantry start angle, to cease using that plan immediately and consult with the original plan for that field and make the correction to the gantry start angle. In addition, the customers were instructed to complete and return the Confirmation of Receipt via fax to: North America and Asia at 702-992-5002 or Europe and Africa +44 (0)1293 654401, Attn: Elekta OIS Su[pport and/or scan and email to: support@impact.com, europe.support@impact.com or by mail; and distribute this notice to any all users of IMPAC software at their organization who are potentially affected by this issue.For questions or requests for upgrade, email support@impac.com (North America/Asia) or europe.support@impac.com (Europe) or call USA 1-800-488-4672. | Contact the recalling firm for information | 840 | 840 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | September 16, 2011 | 84 | 840 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-2646-2011 | 58894 | sterilization wrap containers, trays, cassettes & other accessories | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Smith & Nephew Hip Arthroscopy Repair Instrument TrayPart Number: 72202732Device Description-The Smith & Nephew Hip Arthroscopy Repair Instrument Tray is designed to contain and protect reusable surgical instruments during transport, sterilization, and storage and to allow optimal exposure of the tray's contents to sterilant during the sterilization process. The instrument tray is a perforated stainless steel case with a latchable lid and handle. The tray is fitted with nylon coated aluminum holders, silicone instrument holders, and a protective mat. The instrument tray is marked to facilitate organized instrument placement. | Smith & Nephew Hip Arthroscopy Repair Instrument Tray | 2 | June 24, 2011 | 2011 | Smith & Nephew, Inc. Endoscopy Division | Hip Arthroscopy Repair Instrument Tray shipped with an incorrect Instructions For Use (IFU) | PRODUCTION CONTROLS: Labeling Mix-Ups | Smith & Nephew sent an "URGENT FIELD CORRECTION NOTIFICATION" letter dated May 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to confirm receipt of the corrected IFU document (10600729 rev B.) and communicate to appropriate facility individuals. Once appropriate individuals are informed, customers were instructed to complete and return information requested at the bottom of the letter. The completed information should be returned to Smith& Nephew via fax at 978-749-1654 or by mail at Smith & Nephew Endoscopy, 150 Minuteman Road, Andover, MA 01810.Questions may be answered by contacting the Manager, QA and Compliance at 978-749-1397. |  | 11 | 11 units | Worldwide Distribution--No USA Distribution; countries of Italy, Portugal, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 320 | 11 | Not_Computer
Z-2670-2011 | 58716 | enzyme immunoassay, amphetamine | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | SCREENERS Dip Drug Test, Catalog Number: SCD-7MB300The SCREENERS Dip Drug Test is a one-step immunoassay for the qualitative detection of multiple drugs and drug metabolites in human urine at a given cutoff concentrations (refer to PN: 42131-SC). The SCREENERS Dip Drug Test is used to obtain a visual, qualitative result and is intended for professional use only. This assay provides only a preliminary result. Clinical consideration and professional judgment must be applied to any drug of abuse test result, particularly in evaluating a preliminary positive result. In order to obtain a confirmed analytical result, a more specific alternate chemical method is needed. Gas Chromatography/Mass Spectroscopy (GC/MS) is the preferred confirmation meth Drugs covered by the Screeners Dip Drug Test (SCD-7MB300): BAR, BZO, COC, MET, OPI300, PCP, THC. | SCREENERS Dip Drug Test | 2 | June 27, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech's internal review of regulatory files found that the two affected lots were manufactured without proper regulatory clearance. | PRODUCTION CONTROLS: Process Control | Ameditech, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 18, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer.The letter requested the customer to:(1) Cease distribution of the affected lot and provide remaining inventory count to Ameditech.(2) Destroy any remaining inventory of these lots and document destruction per your site requirements.(3) Ensure that all users of the device receive a copy of the Urgent Medical Device Recall letter.A recall verification form was attached to the letter to be completed and returned by the customer via fax at 1-858-677-0243.Any questions about the information contained in the recall notice should be directed to Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 6000 | 6000 test cards (120 kits) | Nationwide Distribution | Terminated | August 11, 2011 | 45 | 6000 | Not_Computer
Z-2671-2011 | 58718 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | B. Braun CelsiteÂŽ Venous Access Ports , Distributed by Aesculpa Inc.Model ST301, ST305V, ST301G, ST501G and ST315Reference # Model # 04430425 ST301 04430095 ST305V 04433823 ST301G 04437025 ST501G 04436725 ST315 | B.Braun Celsite | 2 | June 27, 2011 | 2011 | B. Braun Interventional Systems | The affected products contain the incorrect winged SurecanÂż needle accessory component in certain CelsiteÂż Access Port package configurations. The winged SurecanÂż needle is used to deliver medications and fluids into implanted CelsiteÂż Access Ports.Two winged SurecanÂż needle sizes are provided in CelsiteÂż U.S. package configurations: 20G (PN 4058631) or 22G (PN 16006593A). A 20G or 22G needle | DESIGN: Process Design | B.Braun Interventional Systems Inc, (BIS) sent a "Urgent Medical Device Recall" letter on April 16, 2011 to all affected customers.The letter described the problem and the product involved in the recall. Informed consignees that their local B.Braun Interventional System Inc Product Specialists will contact them with the specifics for return/replacement of the product upon return of qualified removed product. They also requested that a CelsiteÂŽ Access Port Patient Implant Form be filled out for each patient implanted with one of the CelsiteÂŽ Access Ports from the affected lots. This Form to be provided to them by their local BIS Product Specialist.For additional information contact 800 523-9676 x4459. | Contact the recalling firm for information | 83 | 83 | NY, ND, OK ,PA | Terminated | July 06, 2012 | 375 | 83 | Not_Computer
Z-2664-2011 | 58884 | catheter, electrode recording, or probe, electrode recording | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | ISMUS CATH Mapping Catheter with Auto ID Technology, Part Number D-1171-34-S, Catalog #D7R20P14CTThe affected catheters are indicated for multiple electrode electrophysiological mapping of cardiac structures of the heart, i.e. recording or stimulation only. | ISMUS CATH Mapping Catheter with Auto ID Technology, | 2 | June 27, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster, Inc. has recently become aware that 20 electrodes catheters with Auto ID Technology are not being recognized when connected to the CARTO 3 System. | DESIGN: Software Design | Biosense Webster, Inc. sent a Customer Notification Letter, dated May 24, 2011, with an Acknowledgement Form and Instructions for Manually Entering the Catheter Identification in the CARTO 3 System to all US and Outside US (OUS) customers who purchased these products. For customers in the US, two letters and forms will be sent via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively. Instructions for manual entry of the catheter identification in the CARTO 3 System are being included in the communication to customers. These instructions will allow customers to input the identity of the catheter into the CARTO 3 System and be able to utilize the catheter.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. Customers were instructed to contact their BWI representative for any questions. |  | 429 | 429 units (333 in US) | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Japan, and Singapore | Terminated | October 12, 2012 | 473 | 1531 | Not_Computer
Z-2672-2011 | 58966 | bag, urine collection, leg, for external use, sterile | FAQ | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RUSCH Urinary Drainage Bag, 2000 mlProduct Usage: For one time use per the Directions for Use with urinary catheter for drainage of urine. | RUSCH | 2 | June 27, 2011 | 2011 | Teleflex Medical | Sterile integrity cannot be guaranteed. | OTHER/UNDETERMINED: Pending | Teleflex Medical sent an " URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 2, 2011. The letter identified the product, the problem and the actions needed to be taken. The customers were instructed to immediately discontinue distribution, quarantine any affected products, communicate this recall to any of their customers using the provided customer letter, and complete and fax the enclosed Recall Acknowledgement Form to Customer Service at 1-866-804-9881. Once the fax is received instructions for product return will be provided to customers. In the event that an alternative approach is needed, customers are instructed to contact Teleflex Medical Customer Service at 1- 866-246-6990. For further questions contact your sales representative or Customer Service at 1-866-246-6990. The recall was expanded and an additional letter was sent out on July 20, 2011 to the affected consignees with the above instructions.*****RECALL EXPANDED***** Product 2. Teleflex sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated July 20, 2011, to all affected customers.The letter identified the product, the problem, and the action needed to be taken by the customer. On July 19, 2011 this recall was expanded to include additional product. Please see above strategy information for details. |  | 14488 | 14,488 eaches | Nationwide Distribution: AZ, CA, FL, GA, IL, KS, KY, LA, MA, MS, NC, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WI including Puerto Rico. | Terminated | April 05, 2013 | 648 | 14489 | Not_Computer
Z-2673-2011 | 58878 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 3500 Blood Gas AnalyzerPN 00026000000The GEM Premier 3500 blood gas analyzer is a portable system for use by health care professionals to rapidly analyze whole blood samples at the point of health care delivery in a clinical setting. The instrument provides quantitative measurements of whole blood for pH, pC02, p02, Na+, K+, Ca++, glucose, lactate and hematocrit. These parameters, along with derived parameters, aid in the diagnosis of a patient's acid/base status, oxygen delivery capacity, and electrolyte and metabolite balance. | Instrumentation Laboratory GEM Premier 3500 analyzer | 2 | June 27, 2011 | 2011 | Instrumentation Laboratory Co. | Software upgrade: misuse of the barcode gun by repeatedly clicking can cause the instrument to attach test results to an incorrect patient ID. | DESIGN: Software Design | Instrumentation Laboratories sent a "FIELD SAFETY NOTIFICATION" letter dated April 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.Included with the letter was an upgrade package containing Software V7.2.2 disk along with installation instructions. Also included with the upgrade kit was a Upgrade Tracking Form to be completed and returned via fax at 781-861-4207 or via e-mail at ra-usa@ilww.com.Customers with questions were directed to contact their local representative. |  | 840 | 840 units | Worldwide Distribution-- USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE , INDONESIA, ISRAEL, ITALY, JAPAN , KOREA, REPUBLIC OF CHINA, MALAYSIA, NETHERLANDS , PHILIPPINES, POLAND, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, THAILAND ,TURKEY, and UNITED KINGDOM. | Terminated | March 15, 2013 | 627 | 840 | Software | Output/Calculation | Software update | Software Update
Z-2669-2011 | 58832 | device, incontinence, mechanical/hydraulic | EZY | Implanted mechanical/hydraulic urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology |  | AMS 800 Urinary Control System Control Pump, REF 72400098, (720172-01 Japan only ), Sterilization Method Steam, American Medical systems Inc., 10700 Bren Road West, Minnetonka, MN 55343InhibiZone Antibiotic Surface Treatment, AMS 800 Urinary Control Pump with InhibiZone, REF 72404127, Sterilization Method ETO, American Medical Systems Inc., 10700 Bren West, Minnetonka, MN 55343The AMS Urinary Control System (of which the pump is a component) is used to treat urinary incontinence due to reduced outlet resistance (intrinsic sphincter deficiency) following prostate surgery. | AMS 800 Urinary Control System Control Pump | 2 | June 27, 2011 | 2011 | American Medical Systems, Inc. | American Medical Systems is recalling all unexpired AMS 800 and AMS Acticon Neosphincter Control Pumps.Based upon a review of our product test procedures, we are unable to confirm that all Control Pumps have met our requirements. Test errors could potentially result in a failure to identify a malfunction in the Control Pump activation mechanism. This malfunction may lead to inadvertent activa | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | American Medical System sent an "URGENT MEDICAL DEVICE RECALL letter dated May 9, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customerCustomers were instructed to return any existing stock . Please contact American Medical System at 1 (800) 328-3881 to obtain a Return Authorization number. Customer Service will also provide a FedEx shipping number for your use to return affected product.For further questions please contact American Medical Systems at 1- (800)328-3881. | Contact the recalling firm for information | 15466 | 15,466 | Worldwide Distribution--USA (nationwide) including Washington D.C. and Puerto Rico and the countries FRANCE, GERMANY, CANADA, ISRAEL, SOUTH AFRICA, LEBANON, AUSTRIA, ARGENTINA, SAUDI ARABIA, SPAIN, BELGIUM NETHERLANDS, PERU, VENEZUELA, SRI LANKA, PORTUGAL, TUKEY, BRAZIL AUSTRALIA, CZECH REPUBLIC, TUNISIA. LUXWMBOURG, QATAR, SOLOMON ISLANDS, BOSNIA & HERZEGOVIA, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, ECUADOR, ESTONIA, FINLAND, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PANAMA, POLND, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVAKIA, SLOVENIA, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM and URUGUAY. | Terminated | July 06, 2012 | 375 | 15691 | Not_Computer
Z-2685-2011 | 56408 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxI 800 Access Immunoassay System with Unicel DXI system software version 4.3 | Beckman Coulter UniCel DxI 800 | 2 | June 28, 2011 | 2011 | Beckman Coulter Inc. | Test results may have incorrect ranges. If a default sample type setting on the tests screen is changed while running Unicel DXI system software version 4.3, the new setting is not saved to the system database. The changed sample type setting will reset to the previously saved setting when power is interrupted to the system, or if an operator reboots the system. If the default sample type i | DESIGN: Software Design | Beckman Coulter forwarded on 7/6/10 a Product Corrective Action letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the UniCel Dxl 600 and Dxl 800 Access Immunoassay systems. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to:A revised version of UniCel DxI System Software that corrects this issue is currently in development. Until the new software version is available, Beckman Coulter has provided an alternate method for changing the default sample type setting. See the appendix included with this letter for additional information.Refer to the attached Questions and Answers document for additional details and instructions, and to determine if their laboratory is running UniCel DxI Sytem Software version 4.3. Customers may wish to confirm that proper ranges and units of measure are being reported for any Access assays with sample type settings that were changed while running software version 4.3.Customers were asked to share this information with their laboratory staff, and retain the notification as part of their laboratory Quality Ststem documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that consignees received the notification.If they need assistance or have any questions regarding this notification, customers were told to contact Technical Support at 1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. | Contact the recalling firm for information | 626 | 626 units | Nationwide in the US and to Canada | Terminated | May 07, 2012 | 314 | 794 | Software | Output/Calculation | Software update | Software Update
Z-2676-2011 | 58807 | tube, tracheostomy (w/wo connector) | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Jet-Ventilator Catheter for Adults, 14G, STERILE, NO LATEX, REF 30-02-918-1, VBM Medizintechnik GmbH, Einsteinstrasse 1 7212 Sul z a.N. GermanyAllows a safe and atraumatic transtracheal access for oxygenation/ventilation with a manual jet ventilator or an automatic high frequency jet ventilator. | Jet Ventilation Catheter (transtracheal) | 2 | June 28, 2011 | 2011 | VBM Medical Inc. | Product labeled with incorrectly as being appropriate for use in adults when the intended population for use of the product is children and infants. | PRODUCTION CONTROLS: Error in Labeling | VBM Medical Inc sent an "URGENT FIELD SAFETY NOTICE" dated May 9, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The reference and lot numbers, details of the error, along with the incorrect and correct box label, and instructions to the customers to discard the box with the incorrect label or send the complete product back to VBM Medical Inc. for re-packaging. For any further or detailed questions, please contact VBM Medical Inc. at (317) 776-1800.On 7/12/11 upon receiving the recall classification letter, the firm sent out an additional letter to its direct accounts informing them of the recall's classification and reiterating the instructions listed above. | Contact the recalling firm for information | 29 | 29 | Worldwide Distribution -- USA ( nationwide ) including the country Montreal QC | Terminated | March 14, 2012 | 260 | 417 | Not_Computer
Z-2692-2011 | 58560 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software versionsEos software versions: 1.2.4, 1.2.5, 1.2.6,1.2.7,1.2.8, 1.2.9,1.2.10, .2.11,1.2.12, 1.2.13, 1.3.2,1.3.3, 1.3.4, 1.4.0, 1.4.1, 1.4.2.sterEOS software versions: 1.2.3.3788, 1.2.4.4286, 1.3.2.3373, 1.3.3.3671, 1.3.4.3740, 1.3.5.3915, 1.3.6.4272,1.4.1.4813,1.4.2.4850.EOS system is a digital radiography system used in general radiographic examinations. sterEOS workstation is a medical image review station for acceptance, transfer, display, storage, and digital processing of 2D X-ray images. | EOS System | 2 | June 28, 2011 | 2011 | Eos Imaging | Print true size, print requests from the EOS acquisition station or the sterEOS image review station may, in certain cases, be affected by a different enlargement factorduring printing. T | DESIGN: Software Design | EOS imaging sent a "FIELD SAFETY NOTICE" letter dated April 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to wait for the printing queue to empty before sending a second "print true queue" request with different print settings. Customers should then click the button shown on the instruction sheet to view the print queue.EOS imaging will provide a software update to correct all affected systems at no cost to the customers. An EOS representative will contact customers to schedule a time to upgrade the software.Questions can be answered by calling 33 1 55 25 63 90. |  | 15 | 5 EOS Systems, 10 sterEOS workstations | Worldwide Distribution--USA (nationwide) including the states of CA, IA, MO, VA, and WI, and countries of Canada and Europe. | Terminated | January 18, 2012 | 204 | 15 | Software | Output/Calculation | Software update | Software Update
Z-2677-2011 | 59005 | culture media, antimicrobial susceptibility test, excluding mueller hinton agar | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Remel Spectra MRSA, Ref R01821, 10 plates per package, and Ref R01822, 100 plates per package. The firm name on the label is Remel, Lenexa, KS.The product is a selective and differential chromogenic medium recommended for use in the qualitative detection of nasal colonization of methicillin-resistant Staphylococcus aureus (MRSA) to aid in the prevention and control of MRSA in healthcare settings. It is also intended for use in the qualitative detection of MRSA from positive blood cultures demonstrating gram-positive cocci on Gram stain. | Remel Spectra MRSA | 2 | June 28, 2011 | 2011 | Remel Inc | The product failed to prevent growth of methicillin-sensitive Staphylococcus aureus (MSSA) | PRODUCTION CONTROLS: Process Control | Remel, Inc. sent an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letter, dated May 10, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to review their inventory and discard all remaining product lots. Additionally, customers are to complete the Product Inventory Checklist and indicate whether or not they desire replacement of the product or credit. A self-addressed , stamped envelope was enclosed in the letter to use to return the checklist. Questions can be answered by contacting the Technical Services Department at 800-255-6730 or 913-888-0939. |  | 558 | 187/10-packs and 261/100-packs | Nationwide Distribution--including Puerto Rico. | Terminated | June 18, 2013 | 721 | 448 | Not_Computer
Z-2693-2011 | 58975 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | THD155428- 15.5F X 28cm Titan HD(TM) Catheter KitProduct Usage: Hemodialysis catheter insertion kit | Titan Hemodialysis Catheter | 2 | June 28, 2011 | 2011 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect valved peelable introducer. The catheter sized is 15.5F and the introducer packaged in the kit is a 15F. The kit should contain a 16F introducer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Medcomp sent a "PRODUCT ALERT" letter dated May 19, 2011 to all affected customers. The letter described the product, problem, and actions to be taken. The letter is requesting the return of all affected product. Customers are instructed to contact their customer service representative for an RGA (returned good authorization) number. If you have any questions or concerns contact Medcomp at 215-256-4201 ext. 2225. | Contact the recalling firm for information | 90 | 90 | Worldwide Distribution -- USA (nationwide) including states of AZ, CA, CO and RI and countries of Australia, Brazil, and Netherlands | Open |  |  | 90 | Not_Computer
Z-2687-2011 | 59019 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | The product is the "FASTx Sternal Intraosseous Device".The expiration date for the recalled products is end of June, 2012.The device is used to insert an infusion tube into the manubrium (upper portion of the sternum) of a patient. It is intended for intraosseous infusion as an alternative to intravenous access, to facilitate emergency resuscitation through the use of drugs and fluids. It is indicated for use in establishing a sternal intraosseous access route in adult and adolescent patients (12 years of age and older) requiring vascular administration of drugs or fluids to facilitate emergency resuscitation. | FASTx Sternal Intraosseous Device | 2 | June 28, 2011 | 2011 | Pyng Medical Corporation | Device failed to operate; the insertion of the infusion tube of the FASTx Sternal Intraosseous Device was not achieved in live subjects. | OTHER/UNDETERMINED: Pending | Pyng Medical Corp. sent a "FSN/RECALL FOR FASTx STERNAL INTRAOSSEOUS DEVICE, P/N FASTx -H" letter dated November 5, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are to cease distribution and use of all devices. All devices remaining in inventory should be quarantined. Customers are asked to complete the information at the bottom of the letter and return the product by FEDEx. Questions regarding this notice should contact the Director of Operations at 1-604-303-7964 ext. 211. |  | 402 | 402 devices | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, England, Italy, Mexico, Poland, and Sweden. | Terminated | August 10, 2011 | 43 | 402 | Not_Computer
Z-2680-2011 | 58532 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems, CDI blood parameter monitoring system 500 with arterial and venous blood parameter modules, catalog 500AV. | CDI 500 | 2 | June 28, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Inaccurate readings may result after Methylene Blue, an intravascular dye, is administered to the patient. The presence of Methylene Blue may interfere with the readings from the shunt sensor for the CDI 500 system in a way that would cause inaccurate pH values. Because the pH value is used to calculate other values, these parameters could also be affected by the presence of Methylene Blue: K | DESIGN: Device Design | Terumo sent URGENT MEDICAL DEVICE CORRECTION letters to Chief of Perfusion, Department of Cardiothoracic Surgery or Director of Operating Room Services, dated 4/21/2011. The letter explained the reason for the correction and the potential hazard. The letter stated that Terumo was updating the Operators Manual for all CDI systems and would provide the manuals to all users when available. Consignees with questions should call 1-800-521-2818. Monday- Friday 8 am- 6 pm. |  | 4267 | 3767 all CDI 500 models | Worldwide distribution, including USA, Australia, Barbados, BELGIUM, BRAZIL, CANADA, Chile, Colombia, Costa Rica, Dominican Republic, ECUADOR, El Salvador, France, Germany, Guatemala, Holland, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Korea, Libya, MALAYSIA, MEXICO, PAKISTAN, PANAMA, Paraguay, Saudi Arabia, SINGAPORE, South Africa, Spain, TAIWAN, THAILAND, The Phillipines, TRINIDAD, UNITED ARAB EMIRATES, Uruguay, Venezuela and Vietnam. | Terminated | August 06, 2012 | 405 | 6082 | Not_Computer
Z-2690-2011 | 58444 | coagulator-cutter, endoscopic, unipolar (and accessories) | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LiNA Loop Endoscopic Monopolar Loop for LSH (Laparoscopic Supracervical Hysterectomy). Large loop, 8mm connector, 1 X REF: EL-200-8 and 10 X REF: EL-200-8, Manufactured by LiNA Medical ApS, Surgery Division. | LiNA Loop | 2 | June 28, 2011 | 2011 | LINA Medical ApS | Potential for the LiNA Loop to break during operating procedure. | DESIGN: Labeling Design | Lina Medical Aps notified customers by newsletter in December 2009 followed by a letter to distributors including warnings added to instructions for use in January 2010. |  | 1950 | 1,950 pieces | GA, MN, PA, WA, TN, MA, IN, Ontario Canada | Terminated | February 23, 2012 | 240 | 2720 | Not_Computer
Z-2674-2011 | 58909 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, CaridianBCT, Lakewood, CO 80215.The Spectra Optia Apheresis System, a blood component separator, is intended for use in therapeutic apheresis applications. The system may be used to perform Therapeutic Plasma Exchange, Mononuclear Cell Collection, and Red Blood Cell Exchange. | Spectra Optia Apheresis System | 2 | June 28, 2011 | 2011 | Caridian BCT, Incorporated | An additional air detection system is being added to all existing systems to detect air in the system's return line. | DESIGN: Software Design | CaridianBCT sent a "MEDICAL DEVICE RECALL" letter dated December 2010 to all customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to continue using their systems as long as they followed the instructions in the Operator's Manual. An Acknowledgement of Receipt form was included with the letter to be completed and faxed to 1-303-876-9277 or e-mailed to Regulatory.Affairs@caridianbct.com. A CaridianBCT representative will schedule a visit to facilities to upgrade their systems within the next 6 months.Further information is available by contacting the CaridianBCT Support Center at 877-339-4228 or 303-231-4357. |  | 594 | 594 units | Worldwide Distribution--USA (nationwide) and countries of Australia, Belgium, Canada, China, Hong Kong, Japan, Malaysia, Mexico, Pakistan, Saudi Arabia, Singapore, South Korea, Taiwan, Thailand, and Vietnam. | Terminated | September 12, 2011 | 76 | 594 | Software
Z-2691-2011 | 59039 | ligator, esophageal | MND | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Disposable Hemostasis Clip, DCH-7-230, Rx only, Single Use only, Cook MedicalProduct Usage: This device is used for endoscopic clip placement within the gastrointestinal tract for the purpose of endoscopic marking, hemostasis for mucosal sub-mucosal defects less thank 3cm in the upper GI tract, bleeding ulcers, arteries les than 2 mm, and polyps less than 1.5cm in diameter in the GI tract,. This device is not intended for the repair of GI tract luminal perforations. | WilsonCook Endoscopic Clipping Device | 2 | June 28, 2011 | 2011 | Cook Endoscopy | The disposable clip may not deploy after being positioned inside the patient. | DESIGN: Device Design | Wilson-Cook Medical, Inc. (Cook Endoscopy) sent an "URGENT: PRODUCT RECALL" notice dated June 7, 2011 and a "PRODUCT RECALL EFFECTIVENESS SURVEY' dated June 21, 2011 to all affected consignees. The notice described the product, problem, and actions to be taken. Medical users were instructed to quarantine any affected products and return to Cook Medical. The distributors were instructed to contact their customers and coordinate the return of the affected products to Cook Medical. The recalled notice and survey should be completed and returned to Cook Medical via fax to 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. If you have any questions regarding this matter, please contact the Customer Relations Department at toll free-1-800-457-4500, press 4 then enter extension 2146 or by email at CustomerRelationsNA@cookmedical.com . |  | 307 | 307 units | (USA) Nationwide Distribution including the states of: CA, HI, MA, ME, NC and NY. | Terminated | May 15, 2012 | 322 | 307 | Not_Computer
Z-2703-2011 | 58936 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***LINVATEC SRS SHOULDER RESTORATION SYSTEM CrossFT 5.5mm CrossFT Suture Anchor with Two#2 (5 metric) Hi-Fi Sutures***SRS.Linvatec.com***ConMed Linvatec Biomaterials, Ltd. Hemiankatu 6-8L 33720 Tampere, FINLAND***STERILE EO***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA Made in USA***REF CFP-5502***"The non absorbable suture anchor is intended to reattach soft tissue to bone in orthopedic surgical procedures. | CrossFT Suture Anchor with Two 2 HiFi Sutures | 2 | June 29, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling CrossFT Suture Anchor with Two #2 HI-Fi sutures, Lot 211581, Product Number CFP-5502. There is a possibility that the product may have been incorrectly assembled. The disposable driver may have the incorrect color handle (red versus black). | DESIGN: Device Design | ConMed Linvatec sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated January 27, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to check their inventory for affected lots and return for credit. A Reply Form and FedEx Air bill was enclosed with the letter for customers to complete and return with the product.Questions regarding the recall can be answered by contacting the Customer Service Department via e-mail at Custserv1@linvatec.com. | Contact the recalling firm for information | 45 | 45 each | Worldwide Distribution--USA (nationwide) and countries of Canada and South Africa. | Terminated | July 18, 2011 | 19 | 45 | Not_Computer
Z-2708-2011 | 58913 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set for Epidural Use, 110" (2.8 m), Male Luer Lock Adapter; a sterile, single use, nonpyrogenic fluid pathway for use with 6200 and 6300 series Flo-Gard Volumetric Infusion Pumps; Baxter Healthcare corporation, Deerfield, IL; product code 2C7554sThe tubing set is intended for use with Flo-Gard Infusion Pumps to provide epidural delivery of anesthetic or analgesic drugs for periods up to 96 hours. | Baxter Solution Set for Epidural Use | 2 | June 29, 2011 | 2011 | Baxter Healthcare Corp. | The lot of Baxter Solution Set for Epidural Use is being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | TRAINING: Employee Error | Baxter Healthcare sent Urgent Product Recall letters dated May 31, 2011, via UPS overnight delivery to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/31/11, and instructed to notify their customers and ministries of health of the recall. *** The recall was expanded to include two additional lots of product code 2C7554s on July 11, 2011 via letter of the same date. The recall strategy remains the same as that of May 31, 2011. ***. |  | 13920 | 13,920 units | Worldwide Distribution - USA including Alabama, Arizona, California, Illinois, Kentucky, Louisiana, New Mexico, North Carolina, Ohio, Oregon, Tennessee and Washington, and internationally to Canada and South Africa. | Terminated | June 18, 2012 | 355 | 13920 | Not_Computer
Z-2698-2011 | 58702 | prosthesis, knee, femorotibial, semi-constrained, cemented, metal/polymer | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic |  | Oxford Femoral Knee System REF 154600, LOT 2272736, WARNING DO NOT RESTERILIZE, STERILE UNLESS PACK DAMAGED, BIOMET UK Ltd, (510K# K011138)This device is used in partial knee replacement surgery as a result of patients who are experiencing degenerative arthritis. | Oxford Femoral Knee System | 2 | June 29, 2011 | 2011 | Biomet U.K., Ltd. | This recall was initiated due to a customer complaint which reported that while peeling off the Tyvek lid on the blister pack, very little effort was required to remove it and this could lead to contamination of the device. | PRODUCTION CONTROLS: Packaging | Biomet sent a "FIELD SAFETY NOTICE" dated April 21, 2011, to all affected accounts. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, locate and remove from inventory and quarantine pending return to their Bioment distributor. Customers were also instructed to complete and return the attached "Fax Back Response Form" to their local Biomet distributor.For questions customers should contact their local Biomet Contact. | Contact the recalling firm for information | 12 | 12 devices | Outside the US to Austria, Belgium, and Germany | Terminated | August 03, 2011 | 35 | 12 | Not_Computer
Z-2701-2011 | 58145 | intracranial neurovascular stent | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | CODMAN ENTERPRISEÂŽ Vascular Reconstruction Device and Delivery System 4.5mm x 28mmProduct code: ENF452812 | Codman | 2 | June 29, 2011 | 2011 | Codman & Shurtleff, Inc. | Mislabeled: One of the stick on stickers on the sterile pouch inside of the product box indicates the stent size to be 22mm in length, when in fact the product is 28mm in length. | DESIGN: Labeling Design | Codman & Shurtleff, Inc.sent an "Urgent Medical Device Correction Notice" on March 21, 2011.The letter recommend the incorrect label be removed and discarded. There are 3 additional peel-off labels, 2 on the outer box and one on the pouch. All 3 of these contain the correct printed size information (28 mm) that can be used for applying to the inventory sheet, patient record or patient card, as needed. It is NOT required to return any of this product to Codman.Customer Service at 1-800-225-0460 or contact your local Codman Neurovascular Representative. | No consumer action necessary | 380 | 380 units | Worldwide - Nationwide (USA), Canada, Korea. | Open |  |  | 380 | Not_Computer
Z-2705-2011 | 58376 | instrument, surgical, non-powered | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Sequoia final driver closure top retention, non-sterile, REF 3384-2, Zimmer Spine, Austin, Texas, The device is used to apply final torque to closure tops of the implant.The Sequoia Final Driver is used in conjunction with a counter torque tube to apply final torque to closure tops that secure the rod component to the screw head component until the torque limiting handle pops once, indicating the implant has been locked. This process is repeated with all closure tops in a construct (assembly). The Sequoia Final Driver is also used to tighten the set screw feature and lateral locking cams on the SpeedLinkII transverse connectors when used and is used to remove previously inserted closure tops during construct removal. | Sequoia final driver closure top retention | 2 | June 29, 2011 | 2011 | Zimmer Inc. | A rounded or partially rounded tip on the screwdriver may not provide the required locking torque to lock the closure top to the polyaxial screw. When the hex tip of the screwdriver is not fully engaged and/or the screwdriver is used off axis, the hex tip of the screwdriver can become rounded or partially rounded. | DESIGN: Device Design | The firm, ZIMMER, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 4, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that each instrument tray had both final drivers (with and without retention feature); contact their customer service representative if they do not have both versions; ensure proper seating of the drive by moving the driver side to side, and complete and return the Distributor Certification Form via fax to: Zimmer Spine at 512-258-0995. The letter states the firm will be distributing the version of the screwdriver with no retention feature beginning in March 2011 and distributing the new surgical technique in April 2011. The firm is including an additional final driver in the instrument tray for final tightening of the closure tops. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance & Regulatory Affairs at 800-655-2614. | Contact the recalling firm for information | 505 | 505 devices | Worldwide distribution: USA (nationwide) including the states of: AL, AZ, CA, CT, CO, FL, GA, LA, MI, MN, MO, NC, NY, OH, OK, PA, TN, TX, WA and WI; and countries of: Czech Republic, Denmark, France, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | August 28, 2013 | 791 | 505 | Not_Computer
Z-2707-2011 | 58576 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxPlus bifuse extension set (minibore) needleless connectors, Model Number: MP9232-CProduct Usage: The MaxPlus bifuse extension sets are used administer fluids from a container to a patient's vascular system through a needle or catheter inserted into the patient's artery or vein. | Intravascular Administration Set and Extension Set | 2 | June 29, 2011 | 2011 | Carefusion 303, Inc. | Carefusion has initiated a product recall for the MaxPlus bifuse extension set (minibore) needleless connectors Model MP9232-C as a result of a potential for occlusion. The product may be occluded at the bifused component. | DESIGN: Device Design | CareFusion initiated on May 3, 2011, two distinctive "URGENT MEDICAL DEVICE RECALL NOTIFICATION" via mail delivery, to all affected customers and distributors. The notices identified the product, the problem, and the actions to be taken. Customers were instructed to 1.) Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card to CareFusion. Note on the card your distributor, the MaxPlus bifuse extension set lot numbers and quantities that will be returned.2.) All recalled product should be returned directly to the distributor from whom it was purchased. Customers will receive their return goods number through their distributor or CareFusion. 3.) Also, enclosed is a Frequently Asked Question (FAQ) Letter to address any questions or concerns related to the use of the MaxPlus" bifuse extension sets.The second notification was specifically addressed to distributors. The distributors were provided specific instructions for product returned. For customers who may have questions, please contact your distributor directly, or call the CareFusion Support Center at 1-800- 562-6018. The Support Center will be in operation from 7 a.m. - 4 p.m. PST. For distributors who may have any questions or require additional materials, please contact the CareFusion Support Center at 1800- 562-6018. Additionally, instructions were given to report adverse reactions experienced via the FDA's Med Watch Program. |  | 26150 | 26,150 units | Nationwide Distribution -- USA including states of: CA, IA, IN, and TX | Terminated | November 28, 2011 | 152 | 26150 | Not_Computer
Z-2706-2011 | 59028 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic |  | Brand NameSTAAR Surgical CollamerÂŽ Ultraviolet-Absorbing Posterior Chamber SinglePiece Foldable Intraocular LensCommon NamesCollamer single piece IOLCollamer plate haptic IOLThe Device is intended to correct aphakia in persons 60 years of age or older in whom a cataractous lens has been removed by cataract extraction. The Device is to be implanted in the posterior chamber and in the capsular bag through a tear-free capsulorhexis (circular tear anterior capsulotomy). | STAAR Surgical Collamer UltravioletAbsorbing Posterior Chamber Single | 2 | June 29, 2011 | 2011 | Staar Surgical Co. | The recall was initiated by Staar Surgical due to an error in the Directions for Use (DFU) for Staar Surgical's Collamer Ultraviolet Absorbing Posterior Chamber Single Piece Foldable Intraocular Lens (Collamer IP IOL), alsoknown as the nanoFLEX IOL. The Device Description erroneously indicates that the lens could beimplanted in the ciliary sulcus. | OTHER/UNDETERMINED: Pending | Staar Surgical send out an URGENT - Medical Device Correction Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to acknowledge receipt of the letter by signing, dating, and faxing the form to Staar Surgical. For any questions call (800) 292-7902, ext 2345. |  | 16750 | 16,750 devices distributed to the field | Worldwide Distribution - USA (nationwide) and the countries of South Africa, Slovakia, Mexico, Hong Kong, and Sri Lanka | Open |  |  | 16750 | Not_Computer
Z-2697-2011 | 58592 | tube tracheostomy and tube cuff | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | RUSCH Percutwist Set with Crystal Clear SterileCrystal Clear Tracheostomy Tube is a device inserted into a patient's trachea surgically via the neck to maintain an open airway. | RUSCH Percutwist Set with Crystal Clear Sterile | 2 | June 29, 2011 | 2011 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention. | OTHER/UNDETERMINED: Pending | Teleflex Medical sent an URGENT FIELD SAFETY NOTICE dated January 4, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock for product, cease use and distribution of stock from the affected lot number and quarantine all products from the affected lot number immediately. Customers were instructed to complete the field safety corrective action acknowledgement and stock status form and fax to their local Teleflex Medical Customer Service Sales Rep.If the customers did have stock from the affected lot number, they were to contact their local Teleflex Medical Customer Service/Sales Rep or distributor as indicated on the attached field safety corrective action acknowledgement and stock status form for a Return Autorization Number.Distributors were to communicate the field safety corrective action notice to their customers and forward the completed Acknoweldgement Form to orders.intl@teleflexmedical.com or fax to +353 1 437 0773.For questions regarding this recall call +353 906 460 838. |  | 1 | see line 1. | Worldwide Distribution - Arabia, Asia, Austria, Belgium, Czech Republic, Denmark, Estonia, Germany, Greece, Ireland, Italy, Spain, Saudi Kuwait, Portugal, Russia Federation, Sweden, Slovenia and South America. | Terminated | March 02, 2012 | 247 | 1 | Not_Computer
Z-2704-2011 | 58857 | mask, oxygen | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System, Vital Signs Inc., 20 Campus Road, Totowa, NJ 07512 USAThe kit contains oxygen masks and blood pressure cuffs use on pediatric patients | Broselow/Hinkle Pediatric Emergency System | 2 | June 29, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | Broselow Pediatric Emergency System Kit may contain oxygen masks that may contain a potentially sticky substance on the inside and/or outside of the mask. | OTHER/UNDETERMINED: Pending | GE issued an URGENT MEDICAL DEVICE CORRECTION letter dated April 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the affected product, isolate all affected product, complete the attached confirmation form and fax back per the instructions and return any unused product.For any product forwarded to any other healthcare institution, customers were instructed to forward a copy of the URGENT MEDICAL DEVICE CORRECTION letter to those institutions. For any questions contact Vital Signs Customer Service at 1-800-932-0760. |  | 33002 | 33,002 units | IL | Terminated | January 14, 2013 | 565 | 33002 | Not_Computer
Z-2700-2011 | 58810 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | LIFECATH S PICC with Microflash IntroducerCatheter, Percutaneous, IntravascularProduct Code 2296.214Product Usage: Intravascular catheter | VYGON LIFECATH S PICC | 2 | June 29, 2011 | 2011 | Vygon Corporation | Peripherally inserted central catheter has the potential to leak or break at the hub while in use, which could lead to a loss of blood or catheter migration. | OTHER/UNDETERMINED: Pending | Vygon Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 26, 2011. The letter identified the affected products, problem, and actions to be taken. The customers were instructed to check inventory and quarantine all affected products; immediately cease use and distribution of all affected products; return any affected products using the instruction from Customer Service involving product pickup (contact Customer Service at 1-800-473-5414 for a Return Merchandise Authorization number); and immediately complete and return the attached form via fax to 215-390-2019. Distributors were provided with instructions on how to handle products that were further distributed.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-,ail at customerservice@yygonus.com. |  | 600 | 600 individual units | Nationwide Distribution--USA including states of: AZ, CA, FL, MO, NE, NJ, NY, PA, TX, UT, VA, and WA. | Terminated | June 01, 2012 | 338 | 1170 | Not_Computer
Z-2702-2011 | 58836 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Radiation Treatment Planning System, model H48;Varian Medical Systems Inc.. | Eclipse Radiation Treatment Planning System | 2 | June 29, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Analytical Anisotropy Algorithm versions when modeling the second source, the collimator jaw positions are read from the first control point of the field. This can cause the delivered dose to be higher than planned locally by up to 10%. | DESIGN: Software Design | Varian Medical Systems, Inc. Oncology Systems sent a "Urgent Medical Device Correction" letters on May 3, 2011***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc. at 888- 827-4265. |  | 5201 | 5201 units ***12-12-12 AMENDED TO: 5448 units*** | Worldwide distribution; Nationwide (USA) and the countries of Vietnam, Venezuela, United Kingdom, United Arab Emirates, Ukraine, Turkey, Tunisia, Trinidad and Tobago, Thailand, Taiwan, Syria, Switzerland, Sweden, Sri Lanka, Spain, South Korea, South Africa, Slovenia, Singapore, Saudi Arabia, Russia, Romania, Qatar, Portugal. Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Malaysia, Macau, Lithuania, Libya, Latvia, Kuwait, Kazakjstan, Japan, Italy, Israel, Ireland, Iraq, Indonesia, India, Hungary, Hong Kong,, Guam, Greece, Germany, France, Finland, Egypt, Ecuador, Denmark, Czech Republic, Colombia, China, Canada, Brazil, Belgium, Belarus, Bangladesh, Bahrain, Austria, Australia, Argentenia, Algeria and Canada. | Open |  |  | 5201 | Software | Treatment/Delivery/Therapy | Instructions/Software Update | Software Update
Z-2694-2011 | 58912 | system, communication, powered | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Tobii C-Series Lithium batteries Type TB8A, TB8B, TB12A used with the Tobii C8 and Tobii C12 Speech Generating Communication Devices.Communication device for the impaired and functions as a speech generating device; text or symbols can be used to generate synthesized or digitized speech. Perfect for asking or answering questions, having a conversation. | Tobii CSeries Lithium Batteries | 2 | June 29, 2011 | 2011 | Tobii Assistive Technology, Inc. | Lithium Batteries may overheat and become combustible | PRODUCTION CONTROLS: Environmental Control | Tobii Technology AB sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to stop using the batteries and to remove them from their devices. They were instructed to dispose of the batteries in an environmentally responsible manner as hazardous materials. Tobii will replace the affected batteries. Customers were asked to complete and return an acknowledgement form via fax at 781-461-8213 or e-mail at recall@tobiiati.com.Contact the firm by E-mail: recall@tobiiati.com or telephone: 1-800-793-9227. | Contact the recalling firm for information | 2430 | 2430 units | USA (nationwide) | Terminated | October 01, 2012 | 460 | 2430 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-2711-2011 | 58809 | replacement, ossicular prosthesis, total | ETA | Total ossicular replacement prosthesis. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Kurz, Heinz Kurz Gmb, Medizintecnik, Tuebinger Strasse 3, 7214 Dusslingen, Germany, Surgical Implant, Clip - Piston MVP Titanium, REF: 1006 711, Lot 2 09 13 06.Product is designed for ossicular replacement to functionally restore the middle ear in cases of pathologic changes of the conductive hearing mechanism. | Total Ossicular Replacement Prosthesis, CliP Piston MVP | 2 | June 30, 2011 | 2011 | Heinz Kurz GmbH Medizintechnik | Outer package labels product .75 mm longer than actual size. | MISBRANDING: Labeling False and Misleading | Kurz Medical, Inc. sent an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" letter dated May 9, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to return unused product for replacement by using an enclosed FedEx label. If the product has been implanted , the letter requests the customers send the date of surgery and size of the implant. | Contact the recalling firm for information | 40 | 32, only 8 consigned with U. S. hospitals | Nationwide Distribution--USA (nationwide) including the states of MN, NC, and WA. | Terminated | November 23, 2011 | 146 | 40 | Not_Computer
Z-2709-2011 | 59043 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product Brand Name: Versafitcup Double Mobility Liner Inserter Product Generic Name: Versafitcup Double Mobility Liner Inserter Model Number: 01.26.10.0018The hip prosthesis is designed for cementless use in total hip arthroplasty in primary or revision surgery. The patient should be skeletally mature. The patients condition should be due to one or more of: " Severely painful and/or disabled joint: as a result of osteoarthritis, post-traumatic arthritis, rheumatoid arthritis or psoriactic arthritis, Congenital hip dysplasia, Ankylosing spondylitis. " Avascular necrosis of the femoral head. " Acute traumatic fracture of the femoral head or neck. " Failure of previous hip surgery: joint reconstruction, internal fixation, arthrodesis, hemiarthroplasty, surface replacement arthroplasty, or total hip replacement where sufficient bone stock is present. " Dislocation risks. | Versafitcup Double Mobility Liner Inserter | 2 | June 30, 2011 | 2011 | Medacta Usa Inc | Medacta International has recently realized that the base of the Liner Inserter Ref 01.26.10.0018Lot 105549 of the Versafitcup Double Mobility system could become rusty after the washingprocedures. This is due to an incorrect raw material used by one of our suppliers during theproduction phase of this lot. For this reason, Medacta International has made the decision toremove all the pieces | MISBRANDING: Labeling False and Misleading | Medacta USA sent an "URGENT SAFETY NOTIFICATION-VERSAFITCUP DM LINER INSERTER" letter dated May 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected device, immediately sequester it and not use the device in surgery. The letter instructs customers to pack the affected device for shipment to Medacta USA under Return Material Authorization 1590. Customer will receive a replacement within a few days.Questions regarding this recall are directed to the Director of Regulatory and Quality at 805-437-7085 ext. 26. | Contact the recalling firm for information | 18 | Worldwide: 18 pieces. In the USA: 9 pieces | Worldwide Distribution--USA (nationwide) including the states of CA, ID, IL, and OH, and 9 products to foreign countries. | Terminated | July 12, 2011 | 12 | 18 | Not_Computer
Z-2713-2011 | 58544 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Shoulder Instruments Fracture Stem Inserter w/Version Control, REF 31-406901, Biomet Orthopedics Warsaw, IN, Instrument used to introduce the stem into the canal.The instrument is used to insert the shoulder implant into the shoulder. | Comprehensive Shoulder Instruments Fracture Stem Inserter w/Version Control | 2 | June 30, 2011 | 2011 | Biomet, Inc. | When the surgeon attempts to remove the inserter from the stem, the inserter may become stuck not allowing it to disengage from the stem. The stem may have to be removed from the patient in order to disengage the inserter from the stem. If the stem is removed from the patient, it could potentially cause a significant delay in the procedure or damage/fracture of the humerus which may require inter | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Biomet Orthopedics, sent two "URGENT MEDICAL DEVICE RECALL NOTICE" letters one dated April 7, 2011 and the second letter dated April 12, 2011 which included an additional lot number, to its customers. The letters described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the noted product; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the Fax Response Form within three (3) business days via fax to 574-372-1683; use priority carrier for their shipment; If the product was further distributed, they MUST notify and give letter to Hospital personnel, and confirm receipt of this notice by calling, 800-348-9500, extension 3755 or 3756.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 469 | 469 both products | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Australia, Canada, Chile, Columbia, Costa Rica, Italy, Japan, Korea, Netherlands, New Zealand, Spain and Sweden. | Terminated | April 08, 2013 | 648 | 469 | Not_Computer
Z-2715-2011 | 59174 | filter, infusion line | FPB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pall Total Nutrient Admixture Filter, 1.2 micron filterThis filter is a 1.2 micron low protein binding membrane indicated for the removal of inadvertent particulate debris and microorganisms (Candida Albicans) from total nutrient admixtures and undiluted intravenous fat emulsions. Product is designed for use with intravenous infusion pumps only. | Pall "TNA1" Total Nutrient Admixture Filter | 2 | June 30, 2011 | 2011 | Pall Life Sciences Puerto Rico Llc | Incorrect filter media used (0.2 um filter used instead of the required 1.2 um) within the TNA1 filter used with total parenteral nutrition mixtures. | PRODUCTION CONTROLS: Labeling Mix-Ups | Pall Life Sciences LLC, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 15, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructed customers to maintain the product in quarantine. If distributed, the letter request that customers notify the user to discontinue use. Additionally, customers were asked to complete and return an attached Recall Response Form. The letter also included instructions on how to return the product.Please contact the Customer Service Department at 516-801-9923, if you have questions. |  | 15575 | 15,575 units | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, IL, KS, LA, MA, MO, MS, NC, NY, OK, SD, TX, WA, and WY, and Puerto Rico and the countries of Colombia and Costa Rica. | Terminated | January 10, 2013 | 560 | 15575 | Not_Computer
Z-2714-2011 | 56848 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftReports versions 1.1.6.x and 1.1.7.x SoftReports version 1.1.6 released 2/2008; Version 1.1.7 released 6/2008.SCC Soft Computer 5400 Tech Data Drive Clearwater, FL 33760.SoftReports is a report designer and web-based report launching tool to be used by knowledgeable, trained, and experienced personnel. It provides the ability to create ad hoc, user defined queries, and reports using data from other SCC products. SoftReports enables users to customize standard reports generated by SCC systems to meet their specific needs. This system performs the following functions: 1. The ability to customize report layouts using beginner, intermediate, or advanced options. 2. The ability to import and export reports. 3. The ability to provide custom desktops to launch SQL's within other SCC applications. 4. The ability to preview and modify layouts specific to users' needs. 5. The ability to modify elements of reports using the comprehensive control formatting toolbar. 6. The ability to incorporate charts and graphics into reports. 7. The ability to modify existing query templates. 8. The ability to create independent and embedded queries. 9. The ability to pars SQL statements and notify users of errors within the syntax. 10. The ability to easily update criteria parameters to indicate input needed for executing queries. 11. The ability to view and print any report that is created using the Designer. 12. The ability to schedule, execute, and track ad-hoc reports. This system does not perform the following functions: 1. SoftReports is not configured to write back to the database. 2. SoftReports does not support user-created Ad Hoc reports being used for diagnostic purposes. 3. SoftReports is not intended to alter the meaning of patient result reports (test reports). | SoftReports versions 1.1.6.x and 1.1.7.x | 2 | June 30, 2011 | 2011 | SCC Soft Computer | SCC Soft Computer is issuing a correction for SoftReports Versions 1.1.6.x released 02/2008 and 1.1.7.x released 06/2008. Potential hazard - when using SoftReports Layout Designer, results displayed in the report may not match the patient information in the report header. | DESIGN: Software Design | SCC Soft Computers sent a "RISK TO HEALTH NOTIFICATION TASK" letter dated June 9, 2010 to affected customers. The letter identified the product, problem, and actions to be taken by the customers.Customers were offered an upgrade or hot fix. The letter requested acknowledgement of receipt of the task notification and permission to install the hot fix. |  | 49 | 49 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 26, 2012 | 392 | 49 | Software | Display/Image | Software update | Software Update
Z-2710-2011 | 58695 | system, hyperthermia, rf/microwave (benign prostatic hyperplasia),thermotherapy | MEQ |  |  | Gastroenterology/Urology |  | The following kits contain the UROLOGIX, Coolant Bag, CA1111B, Sterile EO, Manufactured for: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447.UROLOGIX, CTC Advance Short, (2.5-3.5 cm), PT1034, Sterile EO,Manufactured by: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447.CTC Advance Standard , (3.0-5.0 cm), PT1036, Sterile EO,, Manufactured by: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447.CTC Advance Long, (4.5+ cm), PT1038, Sterile EO,Manufactured by: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447.Targis Standard, (3.0-5.0 cm), PT1013, Sterile EO,Manufactured by: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447.The Urologix Targis system is a non-surgical device intended to relieve symptoms and obstruction associated with Benign Prostatic Hyperplasia (BPH) and is indicated for men with prostatic urethral lengths greater than 2.5 cm. | Urologix Targis System, Coolant Bag | 2 | June 30, 2011 | 2011 | Urologix, Inc. | Urologix is recalling certain lots of TargisÂż System Coolant Bag devices distributed over a two week period. The devices are packaged into disposable Procedure Kits containing a microwave catheter, a Coolant Bag and a RTU PlusÂż balloon. The reason for this Recall involves three lots of Coolant Bags that contain an increased risk of leakage of water at the Sensor Module. This leak is detectabl | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Urologix, Inc. sent an "IMPORTANT RECALL INFORMATION" letter dated April 28, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs the customers to identify the affected kit lots, remove the Coolant Bag pouch(s) and to complete and return a Return Material Authorization form with the product(s). Additionally, an enclosed Customer Acknowledgement Form should be completed and faxed to 763-475-1443.Customers with questions should contact the Director of Regulatory and Quality at 763-475-1400. | Contact the recalling firm for information | 119 | 119 | Nationwide Distribution--including the states of AL, CA, CO, FL, GA, MA, MI, MN, NJ, NY, OR, PA, TN, WV, and including Puerto Rico. | Terminated | April 26, 2012 | 301 | 119 | Not_Computer
Z-2721-2011 | 58971 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heater Model 333W, Cardiopulmonary bypass temperature controllers. This device is intended to provide temperature-controlled warm water flow to a blood heat exchanger in order to maintain a patient's blood temperature. It can be used during extracorporeal circulation procedures. It is intended to use with commercially available heat exchangers. | Cardiopulmonary bypass temperature controllers | 2 | July 01, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The probe jack contacts (terminals) on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with each other. If this situation occurs, the service technician, user, and/or patient could be exposed to electric shock during servicing and/or use. | DESIGN: Component Design/Selection | Cincinnati Sub-Zero Products Inc issued an "Urgeant Medical Device Field Action" letter on May 2, 2011 to all affected customers.The letter includes; description of product and problem, affected lots and instructs customers to discontinue its use. The product is to be returned to the Service Department for an update and the customer is instructed to complete and mail Field Action Form.For additional contact the firm at 1-800-989-7373 or (513) 772-8810. |  | 256 | 212 units - U. S. distribution & 44 units foreign distribution | Worldwide Distribution and countries of The Netherlands, Kuwait, India, Saudi Arabia, Taiwan and Australia. (USA) Nationwide including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, LA, MA, MD, MI, MIN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, and WI. | Terminated | March 07, 2013 | 615 | 372 | Not_Computer
Z-2719-2011 | 58465 | passer, wire, orthopedic | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS SYSTEM, CATALOG NO. 9045-02 A sterile, medical device kit consisting of the components listed in Table 1-1 below. The kit is used by medical professionals to stabilize the spine as an aid to fusion through immobilization of the facet joints. A Surgical Technique brochure is provided in Appendix A and other labeling, including instructions for use, is provided in Appendix B. | PERPOS PLS SYSTEM | 2 | July 01, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to K-wires with a distal tip diameter that may be greater than specified causing an interference fit with the canulated screw. Removing the K-wire during a procedure may dislocate the implant resulting in the implant loosening or being pulled out or position.Use of this product may result in a delay of the surgical procedure and/or pull-out of the facet screw | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Interventional Spine sent an RECALL NOTIFICATION letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to take the following actions: Immediately examine inventory and quarantine product subject to recall. In addition, if this product was further distributed, please notify the Vice President Quality Assurance at once. Consignee were asked to also complete and return the enclosed response form as soon as possible. Consignee with questions were directed to the Consignees' Regional Director or the Vice President Quality Assurance.For further questions, please call (949) 472-0006. |  | 864 | 864 units total for all products | Nationwide Distribution including the states of to NC, NY, TN & VA | Terminated | February 14, 2013 | 594 | 864 | Not_Computer
Z-2717-2011 | 58855 | immunohistochemistry antibody assay, estrogen receptor | MYA | Immunohistochemistry reagents and kits. | Hematology | Pathology |  | iView DAB Detection Kit, Part number: 760-091(05266157001)This detection kit is intended for in vitro diagnostic use. Ventana Medical Systems' (Ventana) iVIEW DAB Detection Kit is an indirect biotin streptavidin system for detecting mouse IgG, mouse IgM and rabbit primary antibodies. The kit is intended for laboratory use to identify targets by immunohistochemistry (IHC) in sections of formalin fixed, paraffin embedded and frozen tissue on Ventana Medical Systems (Ventana) automated slide stainers by light microscopy. The clinical interpretation of any staining, or the absence of staining, must be complemented by morphological studies and evaluation of proper controls. Evaluation must be made by a qualified pathologist within the context of the patient's clinical history and other diagnostic tests. | iView DAB Detection Kit | 2 | July 01, 2011 | 2011 | Ventana Medical Systems Inc | The recall was initiated because Ventana Medical Systems, Inc. (Ventana), hasidentified the presence of precipitate in some DAB chromogen dispensers in two lots of ultraViewUniversal DAB Detection Kit (part number 760-500, lot A05885 and lot B00988) and one lot of iViewDAB Detection Kit (part number 760-091, lot B01057A.) |  | The firm, Ventana, sent two letters dated May 16, 2011, an "URGENT Medical Device Recall" notice to customers who received the affected lots and an "URGENT Medical Device Correction" to customers of unaffected lots. The letters described the product, problem and actions to be taken. The customers were instructed to discontinue use of any kits from the lots identified above; follow their internal standard operating procedures for assessing potential impact to reported patient results, If their laboratory had already used kits from these lots; conduct visual inspections of any DAB chromogen dispensers in use as part of their daily maintenance routine; inspect their current inventory and contact their local customer support team (800-227-2155) to replace any unused or partially used kits from the affected lots.Ventana highly recommends use of same-slide controls, especially when testing for targets of therapy, to mitigate any impact on patient results.If you find a dispenser containing precipitate in lots not identified in this letter or have any questions and concerns, please notify/contact your local customer support team (800-227-2155) . | Contact the recalling firm for information | 25 | 25 units | Worldwide distribution: USA (Nationwide) and countries of Canada, Germany and Israel. | Terminated | February 10, 2012 | 224 | 4015 | Not_Computer
Z-2732-2011 | 58733 | saw, powered, and accessories | HAB | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | DYONICS POWER Shaver Blades,4.5mm Long Full Radius, package of 3Catalog Number: 7205341Product Usage: Smith &#38; Nephew DYONICS Disposable Endoscopic Blades are indicated for resection of soft and osseous tissues in both large and small articular cavities. | Smith & Nephew | 2 | July 01, 2011 | 2011 | Smith & Nephew, Inc. Endoscopy Division | Product has insufficient weld between the shaft tube and tip, could result in the tip separating from the shaft tube during use. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Smith & Nephew notified customers by an "URGENT - PRODUCT RECALL 1st NOTIFICATION" letter dated May 6, 2011. The letter describes the product, problem, and actions to be taken. The letter asked customers to call 800-343-5717 (option 3) or send an email to endo.andreturn@smith-nephew.com for a Return Authorization (RA) Number to return product, obtain credit or order replacement blades. Customers are to complete the form and return to the address provided being sure to reference the RA number. For any questions regarding this recall customers were instructed to call 508-337-3731. |  | 555 | 555 units | Worldwide Distribution -- USA (nationwide) and countries of: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong Kong, Italy, Japan, Mexico, Netherlands, Turkey, and United Kingdom. | Terminated | May 09, 2012 | 313 | 1299 | Not_Computer
Z-2734-2011 | 58678 | broach | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Foundation Knee System Instrumentation, Tibial Broach Assembly.Used in conjunction with the Tibial Sizing Template and Tibial Broach Guide to prepare the resected tibial plateau for acceptance of a keeled tibial baseplate stem. | Foundation Knee System Instrumentation,Tibial Broach | 2 | July 05, 2011 | 2011 | Encore Medical, Lp | Strike plate may dislodge from the broach stem handle during impaction of the device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Encore Medical, Lp (d.b.a. DJO Global) sent out a "Urgent Field Safety Notice" dated April 20, 2011 to all affected customers. The firm also e-mailed the consignees as well. The letter described the product, affected lots, reason for recall and provided recommendations including; asking customers contact firm for a replacement, provided a "Field Safety Response Form" to fill out and return. The firm also e-mailed the consignees as well. For additional information please contact the Director of Commercial Logistics at (512) 834-6330 or the International Customer Service at 512 834-6275. | Contact the recalling firm for information | 15 | 15 units | Louisiana and Arkansas and the countries of Germany, Italy, United Kingdom and Japan. | Terminated | April 06, 2012 | 276 | 15 | Not_Computer
Z-2733-2011 | 58686 | recorder,event,implantable cardiac,(without arrhythmia detection) | MXC | Medical magnetic tape recorder. | Cardiovascular | Cardiovascular | 510(k) | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE).This product is a lead less non-therapeutic implantable cardiac monitor and is implanted to aid clinicians in diagnosing intermittent symptoms which may be due to arrhythmia | SJM Confirm, Implantable Cardiac Monitor (ICM) | 2 | July 05, 2011 | 2011 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774. |  | 83 | 83 units | Worldwide distribution: (USA) Nationwide including the States of CA, WA, FL, WV, NY, MD and LA, and the country of Germany | Terminated | July 19, 2011 | 14 | 83 | Software | Output/Calculation | Software update | Software Update
Z-2744-2011 | 58919 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Cambridge Endoscopic Autonomy Laparo-Angle 5mm Metzenbaum Scissors with Monopolar Cautery, 45 cmCatalog Number:LA1101-12Intended for use in a variety of minimally invasive procedures to facilitate grasping, mobilization, dissection, and transection of tissue. | Cambridge Endoscopic Autonomy LaparoAngle 5mm Metzenbaum Scissors with Monopolar Cautery | 2 | July 06, 2011 | 2011 | Cambridge Endoscopic Devices, Incorporated | PET shrink sleeve insulation material dislodgingfrom the distal tips of Autonomy Laparo-Angle (ALA) Metzenbaum Scissors | PRODUCTION CONTROLS: Process Control | Cambridge Endoscopic Devices, Inc. sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 20, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are instructed to return unused inventory using RGA Number 301. Returned devices will be replaced.Questions regarding this recall, call 508-405-0790 for assistance. |  | 255 | 255 units | Worldwide Distribution--USA (nationwide) and the countries of Armenia , Australia, Austria, Chile, China, Denmark, Hong Kong, India, Ireland , Italy, Japan, Saudi Arabia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 24, 2012 | 384 | 902 | Not_Computer
Z-2751-2011 | 58525 | solvent, adhesive tape | KOX | Surgical skin degreaser or adhesive tape solvent. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Brand REMOVE Universal Adhesive Remover Wipe - Pouch Label and 50 count box Label: ***smith&nephew;***#403100*** REMOVE Universal Adhesive Remover Wipes Gently removes tapes and adhesives Reduces risk of irritation and skin stripping Contains Aloe***1 wipe and 50 wipes*** For external use only. Avoid eyes - Use with adequate ventilation. Keep out of reach of children. INGREDIENTS: Dipropylene Glycol Methyl Ether, C10-11 Isoparaffin, Aloe Extract, Benzyl Alcohol, Fragrance. ------------------- Brand: Universal Adhesive Remover Wipes - Carton Label: ***smith& nephew Universal Adhesive Remover Wipes ***Quantity 20 Cartons of 50 Code #403100***. Product Usage: Adhesive removers gently clean all types of adhesive residue from the skin, including hydrocolloid-based, acrylic-based and ruber-based adhesives. | Adhesive Remover Wipe | 2 | July 06, 2011 | 2011 | Smith & Nephew Inc. | Smith & Nephew, Inc., St. Petersburg, FL is recalling multiple lot numbers of the following products: REMOVE Universal Adhesive Remover Wipes, UNI-SOLVE Adhesive Remover Wipes, SKIN-PREP Protective Wipes, PERI-PREP Protective Wipes, and NO-STING SKIN-PREP Protective Wipes. These products were manufactured by H&P; Industries dba The Triad Group who has initiated a recall of products manufactured i | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Smith & Nephew Inc. sent an "URGENT DEVICE CORRECTION" letter dated April 8, 2011 to all customers who received the recalled products. The letter described the product, problem and the appropriate actions to be taken. Customers are instructed to immediately discontinue using the recalled products, examine their inventory, and follow the return or disposition instructions provided. The letter instructed customers to seek an alternative product for use. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. | Contact the recalling firm for information | 50 | 118,896 boxes of 50 count | Worldwide Distribution -- USA (nationwide) including the states of: NC, PA, NM, OR, VA, TN, NY, OR, AL, IL, MN, TX, and Puerto Rico and the countries of: Australia, Canada, England, France, Germany, Hong Kong, Japan, Mexico, New Zealand, Singapore, South Africa, Taiwan, and United Kingdom | Terminated | July 18, 2011 | 12 | 408316 | Not_Computer
Z-2747-2011 | 58845 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Quinton BETA-CAP Adapter; a Peritoneal Dialysis accessory to the PD catheter used to connect the catheter to the PD transfer set; Covidien product code 8814-661001, Baxter product code 5K4560 | Quinton BETACAP Adapter | 2 | July 06, 2011 | 2011 | Baxter Healthcare Renal Div | A change made in the Quinton BETA-CAP adapter used to connect the patient PD catheter to the Baxter transfer set has resulted in a reduction in thread engagement which may lead to an increased risk of separation and the potential to develop peritonitis. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Baxter sent an URGENT PRODUCT RECALL letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter described the problem of loose connections when using the Quinton BETA-CAP with Baxter transfer sets. Baxter requested the dialysis centers to contact their patients currently using a Quinton BETA-CAP adapter to schedule an office visit with their attending physician as soon as practical tohave the connection inspected or replaced. If they choose to continue use of the Quinton BETA-CAP adapter, Baxter recommends the clinician exercise extra caution and vigilance to ensure the connection between the Quinton BETA-CAP adapter and the Baxter Transfer Set is secured and tight. The clinicians were instructed to also check their inventory for the Quinton Beta Cap adapter 8814-661001 purchased directly from Baxter, and contact Baxter at 888-229-0001 to arrange for the return and credit of the unused product. Baxter recommends that the Covidien BETA CAP adapter no longer be used and to instead exclusively use the Baxter Titanium Adapter (5C4129) which has sufficient thread geometry for maintaining a secure connection with the Baxter Transfer Set. The customers were also instructed to complete and fax the enclosed customer reply form to (847) 270-5457, indicating the quantity of Quinton Beta Cap adapters being returned and acknowledging the understanding and dissemination of the information in the letter. | Contact the recalling firm for information | 13942 | 13,942 units | Class II Recall -- Nationwide Distribution | Terminated | May 21, 2012 | 320 | 13942 | Not_Computer
Z-2741-2011 | 58655 | lift, patient, ac-powered | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | JASMINE Full Body (Patient) Lift, electrically-powered device with a 500 lb. weight capacity.The recalled product is an electrically-powered device either fixed or mobile, used to lift and transport patients in the horizontal or other required position from one place to another. | JASMINE Full Body (Patient) Lift | 2 | July 06, 2011 | 2011 | Invacare Corporation | Component Defect. The rear mast connector on certain patient lifts distributed by the firm, may break. Failure of the rear mast connector, while in use would result in a patient being dropped, possibly resulting in serious injury or death to the patient or care provider. | OTHER/UNDETERMINED: Pending | INVACARE contacted their customers via telephone on 4/8/2011 and 4/11/2011, and informed them of this recall. During these telephone conversations, the customers were asked to immediately cease use of the JASMINE patient lifts and to quarantine all product on-site until further notice from Invacare. The firm followed up on the telephone recall notification with a recall letter dated: April 18, 2011. The recall letter was sent via registered mail or Federal Express so as to verify customer's receipt of the recall notification. The recall letter refers the customer/care provider back to the telephone conversations. and notifies the customer of the possibility of failure of the rear mast connector which could result in a patient's possibly being dropped from the patient lift. The letter asks that the customer remove the suspect patient lift(s) from service and quarantine the device in order to prevent further use until such time as an Invacare Rep. contacts the customer by telephone and makes arrangements to pick-up the affect product. The letter is accompanied by a Response Form and a Product Tracking Sheet which lists all of the JASMINE devices shipped to the customer's facility. The customer is asked to complete the attached Response Form and to fax the form back to Invacare at the telephone number provided on the bottom of the form. The signed Response Form provides information to Invacare documenting the customers receipt of the form and understanding of the instructions contained therein, and also; the number of affected devices remaining on-hand at the customer's location, and the number of devices further distributed by the customer. In those instances in which the products have been further distributed, the customer is asked to 'carry out the instructions of the recalling firm and extend the recall to their sub-account customers. |  | 274 | 274 units | Worldwide Distribution - USA including DC, FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NJ, NV, NY, OH, OK, PA, RI, SC, TN, VA, WA, and WI and the countries of Canada, Ireland, United Kingdom, Norway, Austria, Germany, and Australia. | Terminated | February 28, 2013 | 603 | 274 | Not_Computer
Z-2723-2011 | 58536 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pathfinder NXT Closure top driver - short, REF 3566-2, Zimmer Inc., Warsaw, INProduct Usage: The Closure Top Driver is used in conjunction with compressors, distractors, and reducers, as needed, to manipulate and hold the spine into the desired position via loosening/tightening maneuvers on the provisionally inserted closure tops of the construct assembly. After confirmation of correct screw and rod placement, the Closure Top Driver is used to apply final torque to the provisionally tightened closure tops to secure the construct assembly. The Closure Top Driver is also used to remove previously inserted closure tops during construct removal. | Pathfinder NXT Closure Top Driver Short | 2 | July 06, 2011 | 2011 | Zimmer Inc. | Zimmer Spine, Inc. is initiating a voluntary recall regarding an increased risk of rounding or partial rounding of the hex tip on the 3566-1 (long) and 3566-2 (short) Closure Top Driver used for insertion, provisional tightening, and final locking of Closure Tops during the screw and rod construct assembly of the Pathfinder NXT Pedicle Screw System.When the hex tip of the Closure Top Driver is | DESIGN: Device Design | Zimmer, Inc. issued an "URGENT MEDICAL DEVICE RECALL" letters dated April 2011, to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed that (THE GOLD CLOSURE TOP DRIVERS 3566-1 and 3566-2 ARE NOT TO BE USED FOR FINAL TIGHTENING); an alternate instrument for final construct tightening will be included in the instrument set 3570-1 short and 3570-2 long (These instruments are silver instead of gold for easy identification and do not include a retention feature for holding the Closure Top); to ensure final drivers are added to each PathFinder NXT set and inspect Closure Top Drivers for rounded hex wear; contact their customer service representative if they do not have both versions; and complete and return all attached forms via fax to: Zimmer Spine at 512-258-0995. The letter states that the firm will revise the surgical technique to show the updated uses of the screwdrivers. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance & Regulatory Affairs at 800-655-2614. |  | 135 | 135 devices | Worldwide Distribution -- USA (nationwide) including states of: AZ, CA, WI, MI, KS, OK, TX, GA, TN, FL, NJ, CT, PA and the countries of: Australia and Italy. | Terminated | October 07, 2011 | 93 | 269 | Not_Computer
Z-2739-2011 | 58562 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reverse Shoulder Prosthesis System RSP Humeral Socket Shell DJO 9800 Metric Blvd. Austin, TX 78758 STERILE R Model numbers: 508-00-000 and 508-00-008.The Reverse Shoulder Prosthesis (RSP) is indicated for use in patients with a grossly rotator cuff deficient shoulder joint with severe arthropathy or a previously failed joint replacement with a gross rotator cuff deficient shoulder joint. The patient's joint must be anatomically and structurally suited to receive the selected implant (s), and a functional deltoid muscles in necessary to use the device. | Reverse Shoulder Prosthesis System | 2 | July 06, 2011 | 2011 | Encore Medical, Lp | The device has the potential to not appropriately mate with the humeral stem. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DJO Surgical sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructed customers to quarantine affected products and to contact Customer Service to obtain replacement products and a RMA.Questions are directed to the Director of Commercial Logistics at 512-834-6330. |  | 81 | 81 units | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, LA, MA, MD, ME, NM, NY, OK, PA, SC, TN, TX, UT, and WA and the countries of Germany, Italy, and United Kingdom. | Terminated | January 25, 2012 | 203 | 81 | Not_Computer
Z-2738-2011 | 58547 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Large Roller Pump 6" Diameter, Catalog 801041, Terumo Cardiovascular Systems, Corp., Ann Arbor, MI The product is indicated for up to 6 hours of extracorporeal circulation of blood for arterial perfusion, regional perfusion and cardiopulmonary bypass procedures when used by a qualified medical professional. | Terumo Advanced Perfusion System 1 | 2 | July 06, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received multiple reports of the lower roller pump tube clamp mechanism failing to function in a small population of devices, making it difficult to remove or insert the tubing. If the failure occurs, the clamp dial will no longer be able to control opening/closing of the tube clamp mechanism. |  | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated March 21, 2011. The letter identified the product, the problem and the action needed to be taken by the customer. An additional URGENT MEDICAL DEVICE RECALL correction notification and cover letter was sent to all consignees on XX/XX/XX to explain the reason for the expansion of the recall and will provide an update to the timing of implementation of the field correction.Expansion now includes devices manufactured between 2006 and March 2012.The correction strategy is to issue the customers a letter and addendum to the Operator's Manual that specifies what the user should do in the event of a tube clamp failure for any reason. The letter states that the firm is still investigating the issue and the correction. The letter states that in the coming months when a resolution is identified, consignees may be contacted to make arrangements for a upgrade to the system. There is no specific timeline for the correction to be made.If you have any further questions, please call a Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours Monday - Friday, 8 Am - 6 Pm EST. |  | 111 | 111 devices. | Worldwide Distribution -- USA (nationwide) and the countries of Argentina, Australia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, Indonesia, Japan, Jordan, Korea, Kuwait, Malaysia, Mexico, Pakistan, Republic of Georgia, Russia, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, The Phillipines, Turkey, United Arab Emirates and Vietnam. | Terminated | August 28, 2013 | 784 | 3679 | Not_Computer
Z-2745-2011 | 58908 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | HALO Ultra Ablation Catheter:Rx only; for use only with the Halo Flex Energy generator and accessories;Manufactured in the USA at BÂŽRRX Medical, 540 Oakmead Parkway, Sunnyvale, CA 94085.Product Usage: Indicated for use in the coagulation of bleeding and nod-bleeding sites in the gastrointestinal tract including but not limited to the esophagus. | HALO90 ULTRA Ablation Catheter Model 909200 | 2 | July 06, 2011 | 2011 | Barrx Medical, Inc. | Complaints of malfunction of the HALO ULTRA Ablation Catheter related to the device pivot mechanism's inability to return the electrode cap to a flat neutral position. | DESIGN: Component Design/Selection | BARRX Medical, Inc. sent a "PRODUCT RECALL NOTIFICATION" letter dated May 25, 2011 via Federal Express to all of their customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the recalled product. BÂŽRRX Medical Territory Managers will be in contact with each facility and arrange for pick up of the recalled product. The "Steps for Voluntary Recall" form that is enclosed with each letter provides customers with complete instructions on segregating the product, completing the Tracking/Verification Form, obtaining Returned Good Authorization (RGA) number and preparing product for pick up. Customers are to complete and return the Tracking/Verification Form by fax to 408-738-1741 which should include amount of product on hand, contact name and hospital/distributor information. Questions should be directed to BÂŽRRX Medical Customer Service at 888-662-2779. |  | 936 | 936 devices | Worldwide Distribution -- USA (nationwide) and the countries of the United Kingdom and China | Terminated | September 07, 2011 | 63 | 936 | Not_Computer
Z-2746-2011 | 58907 | instrument, manual, surgical, general use | MDM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker Wire Caddy, Stryker Instruments Wire Caddy, REF 505-505, STERILE, QTY 1, MANUFACTURED FOR STRYKER INSTRUMENTS, KALAMAZOO, MI USA 49001 (269) 323-7700 (800) 253-3210.This device is intended for use in the storage and retrieval of guide wires up to 9.8 ft. in length and recommended for guide wires of .018 to .038 inches in diameter. | Stryker Wire Caddy | 2 | July 06, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The towel clamps packaged with the wire caddy have been found to have blood stains on them which could result in infection or blood borne pathogen transmission. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Stryker Instruments sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructs customers to quarantine affected product and return the recalled product using the shipper supplied to them by Stryker. Replacements will be issued to customers upon receipt of the product.For questions regarding this recall please call 1-800-800-4236 ext. 3808. | No consumer action necessary | 522 | 87 boxes, 435 untis | Nationwide Distribution-- including the states of Alabama, Arizona, California, Florida, Illinois, Kansas, Maryland, Michigan, Missouri, Pennsylvania, Texas and Virginia. | Terminated | August 03, 2011 | 28 | 435 | Not_Computer
Z-2740-2011 | 58792 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | NucliSens EasyQ Enterovirus serotype 68 v1.1: Catalog number: 200300In Vitro Diagnostic Test | NucliSens EasyQ Enterovirus serotype 68 ref.200300 v1.1 | 2 | July 06, 2011 | 2011 | bioMerieux, Inc. | The NucliSENS EasyQ Enterovirus Reagents are unable to detect Enterovirus serotype 68, in respiratory samples. | DESIGN: Device Design | BioMerieux Vitek, Inc. sent an "Urgent Product Correction Notice" letter on/about May 10, 2011.The letter included; a description of the product, problem, possible impact to the patient and recommendations. Recommendations included instructing customers to discontinue the use of the kit for the detection of Enterovirus 68 and Rhinovirus 87 serotypes and noted the corrected IFU will be available the first week of April 2011. An Acknowledgement form was attached to be returned via fax.For additional questions please contact your local BioMerieux Clinical Customer representative at (800) 682-2666, option 3, option 2. | Contact the recalling firm for information | 42 | 42 kits | (USA) Nationwide distribution including the states of MD, MI, MN, MT, NY and PA | Terminated | October 06, 2011 | 92 | 42 | Not_Computer
Z-2742-2011 | 58847 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG Hemostasis System Level II Control:Biological Quality Control Level II Kit for Thrombelastograph Coagulation Analyzer TEG 5000 SeriesUsed as the operational check and calibration verification for the TEG 5000 Thrombelastograph Hemostasis Analyzer System | TEG Hemostasis System Level II Control: Biological Quality Control Level II Kit | 2 | July 06, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemoscope received a report that during the process of Validation, Level II QC was running with an out-of-range low MA (Maximum Amplitude). | OTHER/UNDETERMINED: Pending | Haemoscope, a Division of Haemonetics Corp, sent a letter dated March 2, 2011 to all of the affected customers. The letters informed the customers that Haemoscope had received a higher than normal inquiry rate for the Level II Quality Control Kits, part 07-008, lots HMO196 and HMO 197. The accounts were requested to return their stocks of those two lots for further testing and evaluation and to complete the enclosed acknowledgement form, confirming receipt of the letter and the number of kit boxes being returned, and fax it back to Haemoscope at 847-588-0455. The returned kits will be replaced at no charge.For additional information please contact Haemoscope at 847-588-0453, ext-217 . | Contact the recalling firm for information | 568 | 568 kits | Worldwide Distribution: (USA) Nationwide and the countries of Canada, Czech Republic and the Philippines | Terminated | September 25, 2012 | 447 | 568 | Not_Computer
Z-2753-2011 | 59024 | marker, radiographic, implantable | NEU | Implantable clip. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Bariatric Target Ring (Nitinol) (a thin, circular nitinol ring) Model # K-9536-N, Lot #08-001, Lot #08-002 & Lot #08-003. One ring per plastic pouch, three rings are then placed into a plastic box with lid. This box is then placed into a solid white box. Item is shipped and sold non-sterile.It is intended to be used to radiographically (with X-Ray) mark the site of the "gastropexy" performed during gastric bypass surgery, thereby making it accessible for future diagnostic, therapeutic and feeding purposes. | Bariatric Target Ring (Nitinol) | 2 | July 07, 2011 | 2011 | Kapp Surgical Instrument Inc | During an FDA Inspection on 3/21/11 - 3/22/11 at Kapp Surgical an uncontrolled instruction for use was discovered which contained sterilization parameters which differed from the instructions contained in the device master record. Kapp Surgical decided to conduct the recall because we cannot be sure that incorrect instructions were not shipped. | PRODUCTION CONTROLS: Labeling Mix-Ups | Kapp Surgical Instrument, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to check their inventory and quarantine any product affected by this recall. The firm will make arrangements for returning the products. Replacements will be made by Kapp Surgical.Questions about this recall are directed to the QA Manager at 800-282-5277. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 222 | 222 units (rings) | Product was distributed to the following states: OH, PA & SC. Product was also shipped to Saudia Arabia. | Terminated | February 19, 2013 | 593 | 222 | Not_Computer
Z-2758-2011 | 58335 | endoscope and/or accessories | KOG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | J-Vac Bulb Suction Reservoir, Manufactured for: Ethicon, Inc. Somerville, New Jersey, 08876-0151.Closed Wound Drainage Systems have been used as an adjunct in the following specialties to evacuate potentially detrimental collection of certain fluids (e.g. pus, extravascular blood, bile) from wounds in body cavities and to reduce the risk of infection. | JVAC Bulb Suction Reservoir | 2 | July 07, 2011 | 2011 | Ethicon, Inc. | There is a potential for the sterile barrier to be compromised in multiple lots of Blake silicone drains, Blake silicone drain kits, Blake cardio connectors, J-VAC reservoirs, J-VAC drain adaptors. | PRODUCTION CONTROLS: Packaging Process Control | Ethicon, Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated March 24& 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the products and to return the products to Stericycle. The letter includes shipping labels for product return. Ethicon will send replacement product. Customers are asked to return a completed Business Reply Card to Stericycle regardless of whether or not they have the products subject to this recall.Customers with questions regarding this recall are directed to call 1-866-664-1401. |  | 846990 | 846,990 | Worldwide Distribution-- USA (nationwide) and Puerto Rico and the countries of Armenia, Australia, Austria, Azerbaijan, Columbia, Cyprus, Czech Republic, Denmark, French Guinea, French Polynesian, Germany,Israel, Italy, Japan, Kuwait , Netherlands, New Caledonia, New Zealand, Norway, Singapore, Slovenia, South Africa, Spain, Belarus, Belgium, Brazil, Canada, Chile, Dominica Republic, Egypt, Estonia, Finland, France, Greece, Guadeloupe, Hong Kong, Hungary, Ireland, Latvia, Lithuania, Malaysia, Martinique, Mozambique, Portugal, Reunion, Russia, Saudia Arabia, Sweden, Switzerland, Ukraine, and United Kingdom. | Terminated | January 10, 2014 | 918 | 3389796 | Not_Computer
Z-2730-2011 | 59110 | wire, orthodontic | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Stainless Steel, Natural, Lower, .018": Arch Wires.Highland Metals, San Jose, CA 95125 | Stainless Steel, Natural, Lower, .018": Arch Wires | 3 | July 07, 2011 | 2011 | Highland Metals | Boxes of stainless steel orthodontic wires were mislabeled with the incorrect size. | TRAINING: Employee Error | Highland Metals sent out "Highland Metals Advisory Notice Mislabeld Products Alert" and "Advisory Notice" letters from January 11, 2008 to April 18, 2008 to all affected costumers.The letter included: product name and affected lots. It advised them of the mislabeling and instructed them to return the product for replacement.For additional information contact Highland Metals at (408) 271-2955. |  | 40 | 40 packs of 100 | Worldwide distribution including the countries of Colombia, France, Scotland, Germany, Australia, Mexico, England, United Kingdom, Thailand; and (USA) Nationwide including the states of OK, CA, NJ, AZ, MD, LA, IL, PR, FL, NE, IN, TX, NH, NY, CO and PA. | Terminated | July 07, 2011 | 0 | 146 | Not_Computer
Z-2755-2011 | 58850 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ElevateÂŽ System with InteProÂŽ Lite", Anterior and Apical Prolapse Repair System, REF720093-01, Sterile EO, Manufactured by:American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USAProduct Usage: The AMS Elevate System is intended for use where the connective tissue has ruptured or for implantation to reinforce soft tissues where weakness exists in the urological, gynecological, and gastroenteriological anatomy. This includes but is not limited to the following procedures: pubourethral support, including urethral slings, vaginal wall prolapse repairs including anterior and posterior wall repairs, vaginal suspension, reconstruction of the pelvic floor and tissue repair. As configured, the Elevate Anterior device is used solely for prolapse repair | Elevate Anterior and Apical Prolapse Repair Systems | 2 | July 07, 2011 | 2011 | American Medical Systems, Inc. | This is an important notification regarding the ElevateÂż Anterior and Apical Prolapse Repair Systems, Part Number 720093-01, manufactured between August 3, 2010 and January 17, 2011. We have recently received complaints stating that the eyelets were missing from the center graft of some units within the scope of this issue.The complaint rate to-date for the affected product is 0.045%, however th |  | American Medical Systems (AMS) sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 9, 2011 to all customers who received the recalled product. The letter identified the product, problem, and emphasize that the recalled product should not be implanted. The customers were instructed to examine their inventory for the recalled product and call AMS Customer Service at 1-800-328-3881 (select option 1) for an exchange of the product. For assistance in evaluating inventory the customers are instructed to contact their local AMS sales representative. | Contact the recalling firm for information | 8563 | 8563 | Worldwide Distribution -- USA (nationwide) including DC and Puerto Rico and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, FRANCE, GERMANY, PORTUGAL SPAIN, UNITED KINGDOM, ARGENTINA, COLOMBIA, ESTONIA, FINLAND, GREECE, IRELAND, ISRAEL, ITALY, LEBANON, NETHERLANDS, NEW ZEALAND, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SWEDEN, SWITZERLAND, TAIWAN REPUBLIC OF CHINA, VENEZUELA | Terminated | June 28, 2012 | 357 | 8563 | Not_Computer
Z-2761-2011 | 55227 | radioimmunoassay, cortisol | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Cortisol EIA, 96 Wells, DSL-10-2000The DSL-10-2000 Active Cortisol Enzyme Immunoassay (EIA) kit provides materials for the quantitative measurement of cortisol in serum, plasma or urine. This assay is intended for in vitro diagnostic use. | Active Cortisol EIA | 2 | July 08, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the listed lot numbers of Active Control ELISA reagent will result in falsely high values for patient samples due to a shift in the calibration of the assigned standard values.Beckman Coulter recommends that patient results, generated with the affected lots, be evaluated in the context of other diagnostic tests and the clinica | OTHER/UNDETERMINED: Pending | Beckman Coulter sent a Product Corrective Action (PCA) letter dated July 13, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use until further notice and discard remaining affected product. Customers were instructed to contact Beckman Coulter at 1-800-231-7970.Newer lots (lot number 991069 or higher) of the product have been recalibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, please contact Technical Support at the phone number listed below.Consignees were asked to share this information with laboratory staff and retain the PCA letter for their Quality System documentation, and complete and return the enclosed Response Form within 10 days so that the firm can be assured that they received the notification.Questions regarding this notification were directed to the DSL Technical Support Center at 1-800-231-7970 in the United State and Canada, or their local Beckman Coulter Representative. |  | 1369 | 1369 Kits | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | April 27, 2012 | 294 | 1369 | Not_Computer
Z-2764-2011 | 58976 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerator;Model numbers H14, H27 and H29.Product Usage: Varian High Energy Linear Accelerator is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | Varian High Energy Linear Accelerator | 2 | July 08, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Coolant System input water supply manifold may leak in some Clinac Linear Accelerators, posing a risk of electrical shock to any person working within the protective housing. | OTHER/UNDETERMINED: Pending | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 18, 2011, via certified mail, return receipt requested. The letter identified the product, problem and recommended actions to be taken. The letter states that customers will be contacted by Varian Service to arrange for replacement of the recalled product. For additional information contact Varian Medical Systems, Inc at 888- 827-4265. | Contact the recalling firm for information | 93 | 93 | Worldwide Distribution -- USA (nationwide) and countries of: Australia, Columbia, Denmark, France, Ireland, Italy, Japan, Morocco, Poland, Portugal, Saudi Arabia, South Korea, Spain, and United Kingdom. | Terminated | March 13, 2012 | 249 | 93 | Not_Computer
Z-2763-2011 | 58974 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, versions 8.0 and 8.1;Varian Medical Systems, Palo Alto, CA.Product Usage: Aria Radiation Oncology product is a treatment and image management application. It enables the authorized user to enter, access, modify, store and archive treatment plan and image data from diagnostic studies. | Image Database | 2 | July 08, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue has been detected when using the RT Chart function in ARIA versions 8.0 and 8.1. When users select the "Convert to Setup Field" option, a Set-up field is created, that removes the DMLC plan components in order to create an open collimator field for setup purposes. | DESIGN: Software Design | Varian Medical Systems, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 17, 2011, via Certified Mail to US; Federal Express to Canada. Foreign distribution will be handled by Varian Managers. The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc at 888- 827-4265.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***. |  | 357 | 357 | Worldwide Distribution -- USA (nationwide) and countries of Algeria, Australia, Austria, Belgium, Brazil, Chile, China, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Iraq, Ireland, Italy, Japan Jordan, Kuwait, Malaysia, Morocco, Netherlands, Oman, Pakistan, Panama, Portugal, Romania, Russia, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunsia, Turkey, Ukraine, United Arab Emirates, and United Kingdom, | Terminated | March 15, 2012 | 251 | 357 | Software
Z-2762-2011 | 58830 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving Communication SystemThe software is a picture archiving and communication system intended for acceptance, transfer, display, storage, archive, and manipulation of digital medical images including quantification and report generation. | syngo Dynamics 9.0 | 2 | July 08, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | For systems upgraded to syngo Dynamics version 9.0, if a legacy report generated by version 3.x or earlier is opened in the version 9.0 portal, the legacy report may be deleted and replaced with a new report from the version 9.0 reporting system or a blank report. | DESIGN: Software Design | Siemens Medical Solutions, USA, Inc sent a letter entitled "Customer Safety Advisory Notice" to all direct accounts on May 10, 2011. The letter stated the name and version of the software, problem noted, a work-around to be implemented until a service patch is created to fix the problem, and that the letter should be placed in the software's operator's manual to be seen by all who utilize the software. |  | 38 | 38 | Nationwide Distribution | Terminated | July 09, 2012 | 367 | 38 | Software | Output/Calculation | Software update | Software Update
Z-2765-2011 | 58566 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost 94, Diagnost 96, Diagnost 97 Universal X-Ray System The Diagnost is a multifunctional, universal X~ray system with remote controlled facilities. Routine examinations that can be performed with the Diagnost include: Gastro-Intestinal examination, Urograpy, Bronchography, Arthrography, Hysterography, Myelography, Lymphography, Skeletal radiography, Tomography and Venography/phlebography. Special procedures that require subtracted and non subtracted digital imaging include: Angiography, Peripheral angiography (bolus chasing) and Vascular and non~vascular interventional procedures | Philips Diagnost 94, Diagnost 96, Diagnost 97 Universal XRay System | 2 | July 08, 2011 | 2011 | Philips Healthcare Inc. | Updated Instructions For Use: During an examination where the patient is positioned on a -rotatable- chair with knees pointed to the rear of the system, the patients lower legs or ankles may be injured/broken | DESIGN: Device Design | Philips Healthcare notified accounts by letter on April 18, 2011. The letter advised users of the problem and provided updated instructions for proper placement of the patient during the examination. An Updated Instruction For Use was provided.For further information and support customers were instructed to contact the Philips representative. |  | 52 | 52 units | Nationwide Distribution | Terminated | February 15, 2013 | 588 | 52 | Not_Computer
Z-2766-2011 | 58994 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Linear Accelerators (aka Trilogy Mx)Varian Medical Systems, Palo Alto, CAProduct Usage: Intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors , and conditions anywhere in the body where radiation treatment is indicated. | Trilogy Mx Radiotherapy Delivery System | 2 | July 08, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Imaging arms of the TrueBeam Accelerator may have loose encoder pulleys that could lead to inaccurate readout of arm geometry. | DESIGN: Device Design | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 19, 2011, via certified mail, return receipt requested in the US. The letter identified the product, problem and recommended actions to be taken. For additional information contact Varian Customer Support at 1-888- 827-4265. | Contact the recalling firm for information | 77 | 77 | Worldwide Distribution -- USA (nationwide) and countries of: Canada, Denmark, France, Germany, Isreal, Italy, Netherlands, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 13, 2012 | 249 | 77 | Not_Computer
Z-2768-2011 | 58981 | susceptibility test cards, antimicrobial | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 System, VITEK 2 XL 110V (Catalog #27227) and VITEK 2 XL 220V (Catalog #27228).VITEK 2 is an automated system consisting of instruments, software, and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. The system is intended for use to perform in vitro diagnostic organism identification and antimicrobial susceptibility testing in human applications. | VITEK 2 XL | 2 | July 09, 2011 | 2011 | Biomerieux Inc | Cards may not load into the incubator correctly causing them to be loaded in the wrong order which could result in the results being linked to the wrong accession number and reporting results erroneously. | DESIGN: Device Design | Biomerieux, Inc. sent an "URGENT PRODUCT CORRECTION NOTICE" dated June 2, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers.The letter explains how to recognize the issue and instructs customers to take certain steps to resolve the problem. Additionally, an Acknowledgement Form is attached to the letter for customers to complete and return via fax to 919-314-4539. A firmware update will be applied by Biomerieux, Inc. technical support representatives.Questions regarding this recall are directed to the Customer Service Department at 1-800-682-2666, option 3. | Contact the recalling firm for information | 1107 | 1,107 instruments | Worldwide Distribution-- USA (nationwide) and the countries of Austria, Australia, Belgium, Brazil, Canada, China, Colombia, Germany, France, India, Israel, Italy, Japan, Korea, Mexico, The Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirate, and United Kingdom. | Terminated | June 06, 2013 | 698 | 1734 | Not_Computer
Z-2769-2011 | 58993 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The SIU (System Interface Unit) is a component in the Tool Connection Unit of the PercuNav system. It is part of the position sensor system used to calculate the position of the instrument.The Tool Connection Unit is a component of the PercuNav Image Guided Intervention System. It is used to connect tracked instruments to the PercuNav system unit.It has a System Part Number: 989605392501, 989605396032, 989605395201.Tool Connection Unit Part Number: 453561426002, 453561443201The model number: Stand-alone PercuNav: TX2-120, PercuNav integrated on iV22: TXV2.PercuNav is a stereotaxic accessory for Computed Tomography (CT), Magnetic Resonance, (MR), Ultrasound (US), Positron Emission Tomography (PET), Single Photon Emission Computed Tomography (SPECT), Fluoroscopy, Endoscopy and other imaging systems. It displays the simulated image of a tracked insertion tool such as a biopsy needle, guidewire or probe on a computer monitor screen that shows images of the target organs and the current and the projected future path of the interventional instrument taking into account movements of the patient. This is intended for treatment planning and intra-operative guidance for surgical procedures. The device also supports an image-free mode in which the proximity of the interventional device is displayed relative to another device. The system can support a diagnostic mode in which fused images are used to help locate and evaluate targets visible on one modality on a second modality. The device is intended to be used in clinical interventions and for anatomical structures where imaging is currently used for visualizing such procedures. The device is also intended for use in clinical interventions to determine the proximity of one device relative to another. | PercuNav, Tool Connection Unit | 2 | July 09, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | PercuNav Tool Connection Unit may contain an electronic component that does not meet specification, which could cause registration difficulties and/or inaccurate instrument tracking. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare sent an "MEDICAL DEVICE CORRECTION" letter dated May 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The notification letter instructs customers to follow all directions in the Instructions for Use. A Philips Healthcare representative will schedule an appointment to replace the Tool Connection Unit at no charge to the customers.If further information is needed concerning this recall please call Philips Customer Service at 866-767-7822. |  | 10 | 10 units (6 in US and 4 internationally). This was later changed to 7 US and 3 International) all serial codes remain the same | Worldwide Distribution--USA (nationwide) including the states of CA, GA, OH, WA, and Washington DC, and the countries of Australia, India, Netherlands, and South Korea. | Terminated | February 15, 2012 | 221 | 10 | Not_Computer
Z-2781-2011 | 54621 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DICOM Query/Retrieve, version 1.1.4. Individual part # 9733808, or as part of a system order, DICOM Q/R Software Kit, part # 9733933. DICOM Query/Retrieve software is installed on Medtronic StealthStation systems.Product Usage: DICOM Query/Retrieve software enables the user to query a hospital Picture Archiving and Communication System (PACS) for patient DICOM scans. The user can then transfer the identified scan to a Medtronic StealthStation system. | DICOM Query/Retrieve, version 1.1.4 | 2 | July 11, 2011 | 2011 | Medtronic Navigation, Inc. | Software allows incorrect patient or patient history to be displayed. | DESIGN: Software Design | Medtronic Navigation sent a "FIELD SAFETY NOTICE" letter dated June 2, 2011 to all affected customers. The letter identified the product, problem and description of the corrective action. The letter states that Medtronic field representatives will visit each consumer/user and upgrade the version of DICOM Q/R to 1.1.5. Further information is available by contacting Medtronic Navigation Technical Services at 1-800-595-9709. |  | 90 | 90 | Worldwide distribution - USA (nationwide) -- including states of: TX, NC, ID, FL, OH, UT, CA, TN, IL, AZ, NE, LA, PA, OR, AR, MI, NJ, MD and countries of: Canada, Denmark, Finland, Great Britain, Hungary, Ireland, Japan, Korea, Netherlands, Poland, South Africa, Spain, and Taiwan. | Terminated | July 12, 2011 | 1 | 90 | Software | Display/Image | Software update | Software Update
Z-2782-2011 | 57012 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Site Rite Vision Ultrasound System, REF 9770032, Bard Access Systems, Inc.Intended for diagnostic ultrasound imaging or fluid flow analysis of the human body. | Site Rite Vision Ultrasound System, | 2 | July 11, 2011 | 2011 | Bard Access Systems | Purchasing erroneously ordered the incorrect fan using the distributor's part number instead of using the specified manufacturer part number. | PRODUCTION CONTROLS: Process Control | Bard Access Systems, Inc sent out an "Urgent" Medical Device Recall (User Level)" dated June 2011 to all affected customersThe letter included: product name, description of problem and requested the return of units bearing "B50" on the fan. Bard Access Systems, Inc. will replace the incorrect fan with the correct fan. For additional information call 1-724-779-3000. | Contact the recalling firm for information | 86 | 86 units | Nationwide (USA) distribution including the states of AL, AR, AZ, CA, CO, FL, IA, IN, MD, MI, MO, MS, NC, NJ, NM, NY, OH, PA, RI, SC, TX, UT, VA, VT, and WA. | Terminated | July 11, 2012 | 366 | 86 | Not_Computer
Z-2788-2011 | 59006 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System;Varian Medical Systems, Palo Alto, CARadiotherapy treatment for patients with malignant and benign diseases. | Eclipse Treatment Planning System | 2 | July 11, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Eclipse Monte Carlo (eMC) algorithm where not-equidistant data points in the open beam profile used during configuration may lead to inaccurate results. | DESIGN: Software Design | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail dated May 20, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product from use and advise users to manually initiate the configuration process for resampling to be applied when a previous calculation model is copied. Customers were asked to sign and return the notice to confirm receipt. For questions call USA and Canada 1-888-827-4265For Europe call +41 41 749 8844. |  | 10356 | 10,356 | Worldwide Distribution - USA (nationwide) | Terminated | July 13, 2012 | 368 | 10356 | Software | Output/Calculation | Remove | Remove or Replace
Z-2771-2011 | 55164 | radioassay, intrinsic factor blocking antibody | LIG | Vitamin B12test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxC 880i, 860i,680i and 660i Synchron Access Clinical SystemsThe Unicel DxC Integrated Workstation takes samples loaded from a single point of entry through a UCTA (UniCel Closed Tube Aliquotter). The UCTA functions as the sample processing manager by aliquotting and routing samples to the DxC and the Dxl analyzers according to the programming requirements | UniCel DxC 880i, 860i,680i and 660i Synchron Access Clinical Systems | 2 | July 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of inaccurate Access Intrinsic Factor Ab test results displayed on, or printed from, the UniCel Dxl portion of an integrated workstation listed above. The problem is limited to numerical Intrinsic Factor Ab results produced on integrated systems running UniCel Dxl system software version 4.0, 4.1, or 4.2. The DxC portion of your inte | DESIGN: Software Design | Beckman Coulter, Inc. sent out an "Urgent: Product Corrective Action" letter dated May 19, 2009 to all affected customers.The letter explained the reason for recall and provided recommendations including: -System software currently under development will correct this problem, and will be made available soon. Until then, please refer to UniCel DxC reports, user interface information, and LIS transmissions, which contain accurate Access Intrinsic Factor Ab test results. -Refer to the UniCel DxC system documentation for information on displaying and printing results from the DxC portion of an integrated workstation, or from your laboratory information system (LIS) -Past results for these assays only need to be reviewed if they were reported from the Dxl portion of an integrated workstation -Complete and return the enclosed response formFor additional information please contact the Customer Technical Support at 1-800-854-3633, or your local Beckman Coulter representative. |  | 5 | 5 units | Worldwide Distribution: (USA) Nationwide including KY; and the country of Canada. | Terminated | May 07, 2012 | 301 | 5 | Software | Output/Calculation | Software update | Software Update
Z-2783-2011 | 59115 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Bard PowerLine 6 Fr. Dual-Lumen Polyurethane Catheter with SureCuff Ingrowth Cuff, REF 0700610, Sterile, Contents...1 Each - PTFE Introducer and Dilator, 6.5 FR. x 10 cm..., Assembled in Mexico. Product Usage: Designed for short or long term access to the central venous system. They are designed for the administration of I.V. fluids, blood products, drugs, parenteral nutrition solutions, as well as blood withdrawal and power injection of contrast media. | Cuffed central venous catheter | 2 | July 11, 2011 | 2011 | Bard Access Systems | Kit label indicates the wrong dimension for an enclosed introducer. | DESIGN: Labeling Design | Bard Access Systems, Inc. notified customers by letter dated June 16, 2011. The letter notified customers of the mislabeling issue and offered a replacement product if desired. For additional information contact Bard Access Systems at 1-800-443-3393. | Contact the recalling firm for information | 6485 | 6,485 units | Worldwide distribution - USA (nationwide) including the states of: HI, KY, TX, VT and the country of Canada. | Terminated | March 01, 2012 | 234 | 6485 | Not_Computer
Z-2780-2011 | 58969 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; Release 6.5 through 10 (build 7.3.10 through 10.0.34)Varian Medical Systems, Inc.Used to plan radiotherapy treatments for patients with Malignant or benign diseases. Plans beam irradiation with photon, electron and proton beams, as well as for internal radiation treatments. | Eclipse Release 6.5 through 10 (build 7.3.10 through 10.0.34) | 2 | July 11, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, user editing of the MLC Leaf positions may not automatically invalidate calculated dose distributions. | DESIGN: Software Design | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter included the description of the product, problem and recommendations and includes the release of a service pack in the future. For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265). |  | 10664 | 10,097***12-12-12 AMENDED TO: 10664 units*** | Worldwide distribution: Nationwide (USA) and the country of Canada. | Open |  |  | 10664 | Software
Z-2787-2011 | 58897 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens YSIO - solid state x-ray imagerSolid state x-ray imager (flat panel/digital imager), image-intensified fluoroscopic x-ray system | Siemens YSIO solid state xray imager | 2 | July 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential hazard to staff or patients when lowering the table height on the Ysio system. If staff or patients lower extremities are placed underneath the table and the user lowers the table top, it is possible that parts of the table may collide with extremities and cause injuries. | OTHER/UNDETERMINED: Pending | Siemens sent a Safety Advisory Notice to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to place the Safety Notice in the System's instruction for use. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. Customers were instructed to complete the Acknowledgement of Receipt Form and return to the firm. |  | 141 | 141 | Nationwide Distribution including Washington, DC and the following States: AR, CA, CO, FL, ID, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Open |  |  | 141 | Not_Computer
Z-2778-2011 | 58970 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac iX, 23EX, CX, DHXC, Novalis Tx, Trilogy and TrueBeam Linear Accelerators Drive Stands Indicated for sterostatic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Clinac iX, 23EX, CX, DHXC, Novalis Tx, Trilogy and TrueBeam Linear | 2 | July 11, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Some drive stand bolts may fail to meet mechanical specifications. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers.The letter included a description of product, problem and recommendations. This includes performing daily QA to detect isocenter shift and the use of independent patient position system. Customers are advised to discontinue its use if a shift is detected and to contact Varian Medical Systems for a replacement stand bolts and adjustment if necessary.For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265). | Contact the recalling firm for information | 36 | 36 | Worldwide distribution: (USA) Nationwide including the states of PA. NJ, NY, MA, VA, TX, KS, FL, OH, SC, MN, CA, TN and NV; and the countries of Canada, Colombia, France, Italy, Japan, Netherlands and United Kingdom. | Terminated | July 13, 2012 | 368 | 36 | Not_Computer
Z-2789-2011 | 59055 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist v4.1 system with the LANTIS OIS.The Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist | 2 | July 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential safety issue of possible data loss for the transfer of the same patients treatment data between the LANTIS OIS and the RTT v2.x software and the RT Therapist version 4.1 system | DESIGN: Software Design | Siemens Customers Service group began providing Update Instructions and installing the software update on affected machines beginning May 25, 2011. A customer information letter was also provided. | Contact the recalling firm for information | 45 | 45 customers with affected configuration in US, 37 customers ROW. | Worldwide Distribution - USA including distribution to LA, CA, PA, MA, NE, NJ, DE, WI, NC, NY, WI, MD, FL, KY and to Poland, Netherlands, Ireland, G=Belgium, Germany, New Zealand, P.R. China, Saudi Arabia, Australia, Malaysia, Egypt, Canada, and Spain. | Terminated | March 01, 2012 | 234 | 45 | Software | Device Operation | Software update | Software Update
Z-2779-2011 | 59100 | screw, oral | KBW | Ear, nose, and throat manual surgical instrument. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Labeled in part:"***BIOMET MICROFIXATION***BIOMET MICROFIXATION***1520 TRADEPORT DRIVE***JACKSONVILLE, FL 33218 USA***www.biometfixation.com***2.0 MM X 7.0 MM TITANIUM***#01-6577***LOT XXXXXX***QTY 00001***X-DRIVE FOSSA SCREW 5/PKG***2.0 MM X 7.0 MM***".Bone screws intended to aid in the alignment and stabilization of bone in the oral cranio-maxillofacial skeletal system. These screws are to be used with our TMJ (Temporal Mandibular Joint) implants. | 2.0x7mm Fossa Xdrive screws | 3 | July 11, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling part number 01-6577, 2.0x7 mm Fossa X-drive screws, 5 pack, lot # 233270. This product is being recalled due to the possibility that the pack may have contained another part, 01-6581 2.0x11 mm Fossa x-drive screws, 5 pack. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Biomet Microfixation, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to complete and return an Inventory Reconciliation form via fax to 904-741-9425. The form requests that customers indicate whether they want to keep the product or return for credit.Customers can contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400 ext. 468 for questions regarding this recall. | Contact the recalling firm for information | 25 | 25 units | Product was distributed to one distributor in Argentina. | Terminated | November 09, 2011 | 121 | 25 | Not_Computer
Z-2776-2011 | 59135 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Endopledge Sinus Catheter Kit, REF: EP, 9 Fr x 18.9", Sterile EO, Sterile R, Edwards Lifesciences LLC, Irvine, CA. Device is a three-lumen, 9 Fr catheter with an elastomeric balloon near its tip for occluding the coronary sinus in order to retrograde perfuse the coronary circulation. An 11 Fr catheter introducer sheath, which is compatible with the catheter contamination guard, is packaged with these catheters. | Endopledge Sinus Catheter Kit | 2 | July 11, 2011 | 2011 | Edwards Lifesciences, LLC | Possibility of separation of the introducer sheath from the hub on Coronary Sinus Catheters. | PRODUCTION CONTROLS: Process Control | Edwards Lifesciences, LLC sent an Urgent Product Recall - Action Required letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusotmers were instructed to check their inventory for any affected product and return all unused product from their stock to Edwards at:Edwards Lifesciences12060 Lone Peak PkwyDraper, UT 84020Attention: Suzanne Carpenter, RGA #Customers were instructed to contact Customer Service to obtain a Returned Goods Authorization number and replacement product . Once customers verified their inventory, they were to complete the attached form and fax it to Edwards at (949) 250-3489. For any questions customers were instructed to call Edwards Customer Service at (800) 424-3278. | Contact the recalling firm for information | 2798 | 2798 units | Nationwide Distribution. | Terminated | June 01, 2012 | 326 | 2895 | Not_Computer
Z-2786-2011 | 58885 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Systems Symphony, Sonata and Symphony a TIM SystemNuclear magnetic resonance imaging system | Siemens MAGNETOM Systems Symphony, Sonata and Symphony a TIM System | 2 | July 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm has discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens initated their recall of the device on March 31, 2001, by sending a Customer Safety Advisory Notice/Letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer.Customers were notified that detectors will be installed in appropriate locations on the system, which will switch off the system's power electronics in the event of a malfunction in the power supply of the gradient coils, preventing damage to the system and its environment. |  | 440 | 440 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and Puerto Rico. | Terminated | January 08, 2014 | 912 | 440 | Not_Computer
Z-2775-2011 | 58720 | enzyme immunoassay, inhibin-a | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter - Manual Inhibin A ELISA Reagent Kit; Part Number DSL-10-28100t-4The Inhibin A enzyme-linked immunosorbent assay (ELISA) kit provides materials for the quantitative measurement of dimeric inhibin A in human serum and plasma. This assay is intended for in vitro diagnostic use as an aid in the diagnosis and monitoring of various hormonal reproductive disorders. | Manual Inhibin A ELISA Reagent Kit; Part Number DSL10281001 | 2 | July 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because patient results are falsely elevated by approximately 30% because the reference standard of the Manual Inhibin A ELISA assay has shifted from the assays original calibration.Erroneously elevated results may lead to unnecessary ultrasound testing to confirmgestational age, patient psychological stress until diagnostic testing is completed, and/ormedical decis | DESIGN: Process Design | Beckman Coulter sent an "Urgent Product Correction" letter dated March 31, 2011 to all affected customers.The letter included an explanation of the problem and recommendations including; discontinue use of the kit and a Response Form to fill out and return.For additional information contact the Beckman Coulter Customer Support Center at 1-800-854-3633 or your local Beckman Coulterrepresentative. | Contact the recalling firm for information | 5521 | 5521 units total for all products | Worldwide distribution: (USA) Nationwide and the countries of Armenia, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Jordan, Republic of Korea, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Peru, Portugal, Romania, Russian Federation, Serbia, Slovenia, Spain, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | July 11, 2012 | 366 | 5521 | Not_Computer
Z-2770-2011 | 59003 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, REF 5515, 7F, Sterile EO, Rx ONLY, 1 x Catheter and 5 x Catheter, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369.Langston dual lumen catheters are indicated for delivery of contrast medium in angiographic studies and for simultaneous pressure measurement from two sites. This type of pressure measurement is useful in determining transvalvular, intravascular and intraventricular pressure gradients. | Langston Dual Lumen Catheters | 2 | July 11, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions, Inc investigation of a recent Device Experience Report has made them aware of a potential problem with their 7F Langston dual lumen pigtail catheter (Model 5515) with lot number 548264.Recently it was reported that pouches containing Langston catheters (Model 5515) were missing a seal,posing a risk of product contamination. It is possible that a compromised sterile barrier | PRODUCTION CONTROLS: Packaging Process Control | Vascular Solutions, Inc. sent an "URGENT MEDICAL DEVICE FIELD ACTION- LOT SPECIFIC" letter dated June 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructed customers to check their inventory for the affected product and provided information on how to examine the product. Customers are to complete and return a Field Action Customer Inventory Form via fax at 763-656-4250 or e-mail at www.vascularsolutions.com. VSI Customer Service will contact customers to set up return shipment of affected product. All affected product will be replaced upon receipt of the shipped product.Customers should contact their local Sales Representative for questions concerning this recall. | Contact the recalling firm for information | 23 | 23 (5 pack) boxes | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, MA, MI, MO, NM, NC, OK, and WA and the countries of Austria and Canada. | Terminated | April 05, 2012 | 269 | 23 | Not_Computer
Z-2794-2011 | 59057 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker 4.0 mm. Angled Aggressive 12 Degree Blade, Precision Cutters, Stryker Instruments, Kalamazoo, MI 49001 USA, (269) 323-7700 (800) 253-3210, Sterile, Single Use, Rx ONLYFor Functional Endoscopic Sinus Surgery for the incision of soft and osseous tissue in the sinus cavities, open plastic, and reconstructive surgery of the head and neck. | STRYKER 4.0 mm ANGLED AGGRESSIVE 12 DEGREE BLADE | 2 | July 12, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | Recalled device was packaged with the incorrect angled blade. The labeling indicates a 12 degree angle but the product has a 40 degree angle. | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker Instruments Division of Stryker CorporationOn 6/13/11 sent an 'URGENT MEDICAL DEVICE RECALL NOTIIFICATION" letter to all affected customers that included the device name, lot number, dates of distribution, reason for the recall, risks associated with use of the device, a photo of the device, instructions to quarantine and return the devices, and a form to fax back to the firm to confirm receipt and adherence to the recall letter. | Contact the recalling firm for information | 95 | 95 Units | Worldwide Distribution - USA including CA, LA,, MI, NJ, OH, OR, AR, IN, and Internationallly to Australia, Duisburg, Newbury and Hong Kong. | Terminated | August 03, 2011 | 22 | 95 | Not_Computer
Z-2790-2011 | 59035 | humidifier, respiratory gas, (direct patient interface) | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Humidifier Adaptor, Rx only, Teleflex Medical, Research Triangle Park, NC 27709.Use to add water vapor to a dry medical gas for administration to a patient. | HUDSON RCI, Humidifier Adaptor | 2 | July 12, 2011 | 2011 | Teleflex Medical | The packaging may not have been properly sealed, and therefore the sterility of the product cannot be guaranteed. | PRODUCTION CONTROLS: Packaging | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 6, 2011 to all affected customers. The letter described the product, problem, and the action to be taken.The letter instructs customers to discontinue use and quarantine the affected product. Customers must complete and return a Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service. A representative will contact customers with instructions on how to return the affected product. Customers with no affected product are also instructed to complete and return the form to verify receipt of the notification.Contact Customer Service at 1-866-246-6990 if you have questions regarding this recall. and a record of returns. A third letter was addressed to the Risk Manager with the above instructions. | Contact the recalling firm for information | 960 | 228, 960 eaches | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, MI, MN, MO, MS, NM, NE, NY, OH, OK, PA, TX, VA, WI and WM and the countries of Japan, Philippines, and Thailand. | Open |  |  | 228960 | Not_Computer
Z-2793-2011 | 58979 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Guidant RAPIDO Cut-Away Guiding Catheter, 6F. Part Numbers: 66676-201 and 667552-100. For use in cardiac diagnostic and therapeutic procedures. | Guidant RAPIDO CutAway Guiding Catheter | 2 | July 12, 2011 | 2011 | Abbott Vascular | Five lots of Guide Catheters and Accessory products were processed through Ethylene Oxide (ETO) sterilization with additional plastic material covering the pallet, which is not consistent with validated packaging configuration for ETO sterilization. | PRODUCTION CONTROLS: Process Control | Abbott Vascular sent an Urgent Device Recall letter, dated January 7, 2011, noting the products being recalled and the reasons for recall. The firm also provided the following reconciliation instructions to its territory manager:-Help customer identify aflected product.- CaIl Abbott Vascular Customer Service (800) 221-9902 to receive RGA number. Record RGA number below.- Fax this completed form to (951) 914-3826.- Return a copy of this completed form with the returned product. |  | 389 | 389 units total | Nationwide distribution, and Ireland. | Terminated | June 26, 2012 | 350 | 389 | Not_Computer
Z-2795-2011 | 58824 | system and accessories, isolated heart, transport and preservation | MSB | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | XVIVO Perfusion AB (Part of Vitrolife), THAM Solution, 3.3 mmol/ml, 50 mL, REF 99019, Sterile, Used to increase the pH of the Perfadex used for perfusion of donor lungs. Perfadex with THAM kits, 8 x 1000mL REF 19017, and 2 x 2800mL REF 19018. | Perfadex with THAM | 2 | July 12, 2011 | 2011 | Vitrolife, Inc. | Precipitation observed in some bottles of Tham buffer used for adjusting pH of Perfadex prior to perfusion of donor lungs. | PRODUCTION CONTROLS: Process Control | XVIVO Perfusion, part of the Vitrolife Group, sent out a "Partial recall of Perfadex with THAM" letter on May 13, 2011 to all affected customers.Customers were instructed to to quarantine all Perfadex with THAM. Replacement THAM was to be shipped by Vitrilife within a few days and the exchanged THAM was to immediately sent to Vitrolife, Englewood, CO.For additional Information please call (303) 762-1933. |  | 1457 | 1457 bottles | Worldwide distribution including USA and Europe. | Terminated | July 15, 2011 | 3 | 1457 | Not_Computer
Z-2799-2011 | 58768 | restoration, base metal | NSQ |  |  | Dental | Enforcement Discretion | "***REF IWTH56***Certain EP HEALING ABUTMENT 5mm(D) X 5mm(P) X 6mm(H).*** Healing abutment*** STERILE.***Rx Only.*** Made in USA*** Implant Innovations, Inc. Global Headquarters 4555 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 342-5454 www.3i.online.com. Outside US 561-776-6700***Implant Innovations Siberica S.L., WTC Almeda Park, Ed. 1 Planta 1a Pl. De la Pau s/n 08940-Cornella de Llobregat (Barcelona) Spain. Phone: +34 934 705500. Fax +34 933 717 849. Lot Number: 602691.***"Biomet 3i healing abutments are temporary healing abutments for use in maintaining the soft tissue opening throughout the healing process of dental restoration. | EP HEALING ABUTMENT | 3 | July 13, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i of Palm Beach Gardens, FL is recalling their Certain EP Healing Abutment (IWTH56). The Healing Abutment dimensions did not match outside label and laser etching on part. The product was actually 8mm but the package indicated 6mm. | PRODUCTION CONTROLS: Error in Labeling | The firm, BIOMET 3i, initially contacted each US customer by telephone using a prepared phone script. Following telephone contact, the firm followed up each call with a faxed "Urgent: Medical Device Recall" letter dated March 1, 2010. International customers were notified first by email then by fax. The letter and email described the product, problem and actions to be taken. The customers were instructed to return the healing abutment(s) for replacement if they have not already opened and used the product; respond with the attached response fax and return the product(s) to the: Regulatory Affairs Department, BIOMET 3i, 4555 Riverside Drive, Palm Beach Gardens, FL 33410, Incident No. IC95863. BIOMET will send a replacement healing abutment(s) once the returned product is received.If you have any questions or concerns, please contact the Recall Coordinator at 561-776-6906. | Contact the recalling firm for information | 49 | 49 | Worldwide distribution: USA (nationwide) and countries of: Australia, Europe and Japan. | Terminated | November 09, 2011 | 119 | 49 | Not_Computer
Z-2798-2011 | 58863 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Terumo(tm) Hypodermic Needle, sterile, individually packaged into a "blister-pack" flexible plastic bottom-web thermo-formed to create a pocket, heat-sealed with paper top-web,1,000 needles per case.Product Usage: Aspiration and injection of fluids for medical purposes. | DISP. HYPODERMIC NEEDLE | 2 | July 13, 2011 | 2011 | Terumo Medical Corporation | Hypodermic needle packaging defect could compromise sterility. Affected lot numbers are: MC1725, MC2025, MC2025 | PRODUCTION CONTROLS: Packaging | The firm, Terumo Medical Corporation sent an "URGENT PRODUCT BULLETIN: RECALL OF TERUMO PRODUCT" letter dated June 8, 2010. Consignees were contacted via telephone on June 7 thru 9, 2010 to confirm receipt of the Product Recall Bulletin. The letter describes the product, problem and the actions to be taken. The letter instructed customers to immediately discontinue use and quarantine any available stock of the listed product codes. The letter stated that distributors must not distribute any of the listed product codes, quarantine any available stock and send the recall bulletin immediately to each customer that received the recalled products to inform them of this issue. Additionally, Terumo Medical Corporation sent a "Terumo Customer Instructions for Returning Recalled Product and Customer Confirmation of Product Recall Notification" letters dated June 8, 2010 to all affected facilities. Customers were instructed to fill-out the requested information and return to Terumo Medical Corporation (TMC) Product Recall Coordinator. Consignees were contacted via telephone on June 8 and 9, 2010 to confirm receipt of the recall information and compliance with the requested action. For questions or need for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. | Contact the recalling firm for information | 453000 | 453,000 needles | Nationwide Distribution (USA) -- AZ, CA, CO, FL, GA, IL, KS, KY, MA, MO, NC, NJ, NY, OH, PA, TN, TX, WA, UT, and VA | Terminated | July 14, 2011 | 1 | 453000 | Not_Computer
Z-2796-2011 | 58806 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague 3 and 3 CX Infusion Pumps, Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2 M8153 and 2M8163. Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. | Colleague 3 and 3 CX Infusion Pumps | 2 | July 13, 2011 | 2011 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | TRAINING: Employee Error | Baxter telephoned all seven affected customers beginning May 4, 2011, informing them that that a post repair test was not performed correctly on their Colleague pumps during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will coordinate with the affected customers, the return of the device for the necessary testing to be completed at their facility and return of the device back to the customer at no cost.Please call ( 800 ) 437-5176 for a complaint that includes any indication of the failure of the product to meet a customer's expectations for quality or to meet performance specifications. | Contact the recalling firm for information | 12 | 12 pumps | Nationwide Distribution including the states of NY, Oh, KY,SC, PA, FL and Puerto Rico. | Open |  |  | 12 | Not_Computer
Z-2797-2011 | 58916 | products, red-cell lysing products | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A.The product is a lytic reagent used to lyse (destroy) red blood cells for hemoglobin determinations or aid in the counting of white blood cells. | Sysmex StromatolyserIM (SIM220A) | 2 | July 13, 2011 | 2011 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent was found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | PRODUCTION CONTROLS: Environmental Control | Sysmex America, Inc. sent a "MEDICAL DEVICE RECALL" letter dated June 2009 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax at 800-218-2711. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the recall letter are directed to the Sysmex Technical Assistance Center at 1-888-879-7639.The firm sent a second notice to non-responding accounts on September 14, 2009. A copy of the June 26, 2009 recall letter and the fax back form were included with the letter. The customers were informed that Sysmex had not received their completed fax back form and were requested to respond to confirm the notification was received. | Contact the recalling firm for information | 4005 | 4,005 cubetainers | Worldwide Distribution--USA (nationwide) and the countries of Brazil, Canada and Honduras. | Terminated | July 13, 2011 | 0 | 4005 | Not_Computer
Z-2724-2011 | 56016 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IQecg IQiC Interface, Part Numbers: 4-100-1407, Software Version 6.0IQecg is to be used, under the supervision of a Physician, to obtain Electrocardiograms from the human body surface. | IQecg | 2 | July 14, 2011 | 2011 | Brentwood Medical Technology Corp | The recall was initiated because when a user of the IQmanager, IQdevices SDK, and IQiC (Centricity) software versions listed above accesses the QT dialogue box to either accept or edit the QT value, the QTc value is recalculated using a Hodges QTc equation, regardless of patient's age. The software does this without notifying the user of the change from use of the Bazett equation, which is utiliz | DESIGN: Software Design | The firm, Midmark, sent an "URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter on May 19, 2010 with an attached Corrective Return Material Authorization (RMA) form via US Postal Service Certified Mail to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their IQecg, IQmanager, or lQiC Centricity software to determine the software version currently in use. If it is a listed product, it is subject to this field corrective action. A sample software label was attached to aid the consignees/customers in locating and identifying the software CD. For customers who have already installed the software, the version number can be identified in the "About" page.Consignees/customers were instructed to complete the enclosed Field RMA form and return it with the affected software to:Quality Assurance DepartmentMidmark Diagnostics Group1125 West 190 S t.Gardena, CA 90248In addition, if they further distributed this product, identify customers and notify them at once of this product field corrective action.Replacement software was to be shipped to the consignees/customers upon receipt of the returned software CD. If they require a UPS return shipping label or an on-line download of replacement software, contact Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or 5108. If you have any questions, call Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or QA/RA Manager, Midmark Diagnostics Group 310-516-5116. |  | 90 | 90 units | Worldwide distribution: USA (Nationwide) and countries of: Algeria, Australia, Canada, France and Saudi Arabia. | Terminated | August 22, 2011 | 39 | 1621 | Computer
Z-2802-2011 | 59207 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | iSert Intraocular Lens (Model PC-60AD) | iSert Intraocular Lens (Model PC60AD) | 2 | July 14, 2011 | 2011 | Hoya Surgical Optics, Inc. | The recall was initiated because Hoya Surgical Optics has confirmed that the iSymm Intraocular Lens (Model FC-60AD) and iSert Intraocular Lens (Model PC-60AD) have been sold and marketed with out an FDA approval. | PREMARKET APPROVAL: No Marketing Application | Hoya Surgical Optics Inc. sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated June 22, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken.The letter instructs customers to quarantine affected product and return to HSO at their earliest convenience. Shipping instructions are included in the letter. Customers are to complete an Effectiveness Check Questionnaire and fax it to 1-909-680-3986.Questions regarding the recall process are directed to Hoya Customer Service at 1-866-750-5870 or e-mail at customercare@hoyasurgopt.com. |  | 33743 | 33,743 units total | Nationwide Distribution | Terminated | November 29, 2012 | 504 | 33743 | Not_Computer
Z-2800-2011 | 59018 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Model 134, WarmAir Convective Warming Device, 115V, Part #86180, Device Listing #D111865, The WarmAir 134 is used in conjunction with a single-patient use blanket to warm a patient's body temperature. The device draws in room air through a filter by means of an electric blower. It heats the air to the desired temperature by controlling the air temperature entering a flexible connecting hose, and conveys it to the disposable blanket (i.e. FilteredFlow Blankets, Warming Tube) in contact with the patient, in order to warm the patient. This device provides a continuous flow of heated air as a means of warming patients suffering the common, but significant problems of hypothermia. The system can be used in Post Anesthesia Care Units (PACUs), Intensive Care Units (ICUs), Surgical Intensive Care Units (SICUs), Emergency Rooms (ERs), Operating Rooms (ORs), or any other departments or hospital facilities requiring patient temperature management. | Model 134, WarmAir Convective Warming Device | 2 | July 14, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | A component defect. The recalling firm determined that one of the connectors attached to the firm's WarmAir 134 device's Printed Circuit Board (PCB) can overheat. Should this event occur, the excess heat build up within the device may result in damage to the device's external case and/or melting of some of the internal components of the unit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CSZ Cincinnati Sub-Zero sent an "URGENT MEDICAL DEVICE RECALL FIELD ACTION" notice to all affected customers. The letter identifies the product, problem, and actions to be taken.CSZ instructs customers to complete and return an enclosed recall response form and to state whether or not their facility is in possession of the device. The completed form should be returned via fax at 513-772-9119 or via e-mail to WA134Recall2011@cszinc.com. Additionally, customers are instructed to discontinue use of the device and return to their Biomedical Engineering and/or Service Department for device exchange and return.Questions regarding this notice are directed to the WarmAir 134 field action (recall) administrator at 1-800-989-7373. | No consumer action necessary | 463 | 463 units were distributed | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the country of Canada. | Terminated | February 26, 2013 | 593 | 463 | Hardware | Physical Safety Hazards | Return | Remove or Replace
Z-2835-2011 | 59078 | lancet, blood | FMK | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | OneTouch Delica Lancing Device;LifeScan Inc., Milpitas, CATesting of Blood Glucose Levels | OneTouch Delica Lancing Device | 2 | July 15, 2011 | 2011 | Lifescan Inc | Individuals attempting to remove a lancet from the OneTouch Delica lancing device without following instructions for use may experience an accidental needle stick. | OTHER/UNDETERMINED: Pending | Lifescan sent an Urgent Medical Device Correction letter dated May 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told that no action on their part is needed and were asked not to return any of the affected product. Customers were instructed to continue to distribute the product. Healthcare professionals and pharmacies were asked to share the information with their patients who use the OneTouch Delica Lancing Device. Patients with questions should contact Customer Care Line at 1-800-574-6138. For questions regarding this recall call 1-800-600-7226. | Contact the recalling firm for information | 1032515 | 1,032,515 in US | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 05, 2012 | 174 | 1032515 | Not_Computer
Z-2833-2011 | 59096 | prosthesis, tracheal, expandable | JCT | Tracheal prosthesis. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Boston Scientific, WALLGRAFTÂŽÂŽ OVER-THE-WIRE, Tracheobronchial Endoprosthesis, Made in IRELAND, Ballybrit Business Park, Galway, IRELAND, STERILE, EO.The WALLGRAFT Tracheobronchial Endoprosthesis is indicated for use in the treatment of tracheobronchial strictures produced by malignant neoplasms. Material Catalog #DescriptionM00170502070-502WALLGRAFT PKG T/B/10X20 10FR COMMLM00170503070-503WALLGRAFT PKG T/B/10X30 10FR COMMLM00170505070-505WALLGRAFT PKG T/B/10X50 10FR COMMLM00170507070-507WALLGRAFT PKG T/B/10X70 10FR COMMLM00170523070-523WALLGRAFT PKG T/B/12X30 11FR COMMLM00170525070-525WALLGRAFT PKG T/B/12X50 11FR COMMLM00170527070-527WALLGRAFT PKG T/B/12X70 11FR COMMLM00170545070-545WALLGRAFT PKG T/B 14X50 12 FRM00170547070-547WALLGRAFT PKG T/B 14X70 12FRM00170562070-562WALLGRAFT PKG T/B/6X20 9FR COMMLM00170563070-563WALLGRAFT PKG T/B/6X30 9FR COMMLM00170565070-565WALLGRAFT PKG T/B/6X50 9FR COMMLM00170567070-567WALLGRAFT PKG T/B/6X70 9FR COMMLM00170572070-572WALLGRAFT PKG T/B/7X20 9FR COMMLM00170573070-573WALLGRAFT PKG T/B/7X30 9FR COMMLM00170575070-575WALLGRAFT PKG T/B/7X50 9FR COMMLM00170577070-577WALLGRAFT PKG T/B/7X70 9FR COMMLM00170582070-582WALLGRAFT PKG T/B/8X20 9FR COMMLM00170583070-583WALLGRAFT PKG T/B/8X30 9FR COMMLM00170585070-585WALLGRAFT PKG T/B/8X50 9FR COMMLM00170587070-587WALLGRAFT PKG T/B/8X70 9FR COMMLM00170592070-592WALLGRAFT PKG T/B/9X20 10FR COMMLM00170593070-593WALLGRAFT PKG T/B/9X30 10FR COMMLM00170595070-595WALLGRAFT PKG T/B/9X50 10FR COMMLM00170597070-597WALLGRAFT PKG T/B/9X70 10FR COMML | WALLGRAFT TRACHEOBRONCHIAL ENDOPROSTHESIS AND UNISTEP DELIVERY SYSTEM | 2 | July 15, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating an important field correction related to WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis products. This letter is intended to clarify and re-emphasize that U.S. WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis stents are NOT currently cleared by the FDA for use in benign tracheobronchial strictures. |  | The firm, Boston Scientific, sent an "Urgent Medical Device Field Correction" letter dated May 25, 2011 to its customers. The letter described the product, the problem and actions to be taken. The letter states that no product is being recalled and asks consignees not to return the product. The customers were instructed to read the enclosed information and post it in a way to ensure it is easily accessible to all users of this device; complete and return the Reply Verification Tracking Form via fax to: CV Field Action Center, 1-866-213-1806, and if they are a distributors to forward this notification to their customers. If you have any questions or would like assistance with this field correction, please contact your local Sales Representative and/or call 763-494-1133. | Contact the recalling firm for information | 23218 | 23,218 FOR ALL PRODUCTS | Worldwide distribution. | Terminated | August 10, 2012 | 392 | 23218 | Not_Computer
Z-2812-2011 | 59104 | suture, nonabsorbable, silk | GAP | Natural nonabsorbable silk surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Syneture Sofsilk 5/0 BLACK 5 X 18" CV-22 D*TACHnon-absorbable silk sutures GS43MSofsilk is intended for use in general soft tissue approximation and/or ligation, including use in the cardiovascular, ophthalmic, microsurgery and neurological surgery | Syneture Sofsilk | 2 | July 15, 2011 | 2011 | Covidien LP | Potential for the sterile barrier to be breached compromising the sterility of the product | PRODUCTION CONTROLS: Equipment Maintenance | Covidien sent an URGENT MEDICAL DEVICE RECALL letter dated June 3, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify and quarantine all affected inventory. Distributors who further distributed the affected product were instructed to forward the letter to their customers who received the affected product. Customers were instructed to contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722- 8772, option 1, to obtain a Return Material Authorization (RMA) prior to returning the affected product. Customers were also asked to complete the attached product recall form in its entirety and fax the form to 203-822-6009 and include a copy with products being returned. If customers did not have any of the affected product they were to simply return the recall form via fax or email indicating they had zero units. For any questions regarding this recall customers were to contact their Covidien Representative or Covidien Quality Assurance at 203-492-5232. |  | 528 | 528 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Denmark, France, Germany, Italy, Japan, Korea, and Singapore | Terminated | February 10, 2014 | 941 | 93600 | Not_Computer
Z-2834-2011 | 59099 | light, operating, dental | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | LuxStar Dental Light have the following Agency model or Catalog Model numbers as ML1000, OL1000, CL1000, WL1000, CB1000, TL1000, and UL1000.The product is labeled in part: " Marus Dental***Mfg By: Dental Equipment LLC.***705 S. Springbrook Rd Building C Newberg, Oregon, 97132 USA***Dental Light***MN ML1000***".Product Usage: The intended use of the dental light is used for illuminating the oral cavity during the performance of dental procedures. | LuxStar Dental Light | 2 | July 15, 2011 | 2011 | Dental Equipment LLC | Lens heat shield (birdcage) of the Luxstar Dental Light could fall and injure a patient. | DESIGN: Device Design | Dental Equipment, LLC sent an "URGENT MEDICAL DEVICE RECALL" notification letter via USPS 1st class to US consignees and via Federal Express or Fax to foreign consignees on June 24, 2011. The letter informed customers about the product, problem and tether upgrade corrective action. A Field Correction Acknowledgement Form and the Field Correction Completion Form was attached for customers to complete and return via the fax number 503-537-3617. The letter states that customers are to review their records in order to identify where the affected devices are located and which are still in service. Customers can call the Customer Service at 800-304-5332 and order the self-install Upgrade Kits with the Part Number 30792 at no charge. Customers are instructed to complete the Rework Correction Form once all of the lights have been updated and fax or mail to Marcus, attention Bonnie Dundas, at (503) 537-3617. |  | 13226 | 13,226 units distributed worldwide | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico and the countries of: Australia, Belgium, Canada, France, Germany, Ireland, Italy, Mexico, Norway, Poland, Russia, United Kingdom, Vietnam, and Trinidad. | Terminated | October 24, 2012 | 467 | 13226 | Not_Computer
Z-2831-2011 | 59184 | sheath, corrugated rubber, for non-indwelling catheter, non-sterile | NNY | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Rusch, Sure Seal Golden Drain, One Piece Urinary Incontinence Device, Teleflex Medical.Used for urine collection. | Rusch, Sure Seal Golden Drain, One Piece Urinary Incontinence Device | 2 | July 15, 2011 | 2011 | Teleflex Medical | The Kit contains Skin-Prep protective wipes that were manufactured by the Triad Group and are being recalled by Smith & Nephew, due to possible bacterial contamination. | OTHER/UNDETERMINED: Pending | Teleflex Medical sent an "Urgent Medical Device Recall Notification" dated June 27, 2011 to all affected customers.The firm provided recommendations including all affected product be discontinued and quarantined. A Recall Acknowledgement Form was provided to be completed and returned. If the product was further distributed, they were urged to contact their customers and communicate the recall.For additional information please contact Customer Service at 866-246-6990 or your local sales representative. | No consumer action necessary | 6880 | 6880 units (total for both products) | Nationwide Distribution (USA) including the states of FL, GA, OH and SC. | Terminated | March 20, 2012 | 249 | 6880 | Not_Computer
Z-2837-2011 | 59263 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Non-DEHP I.V. Catheter Extension Set, 6.6 (17 cm), Volume 0.8 mL, Injection Site, Male Luer Lock Adapter; an Rx sterile, single use nonpyrogenic fluid pathway; 200 sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 2N3375Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. | Baxter Interlink System NonDEHP I.V. Catheter Extension Set | 2 | July 15, 2011 | 2011 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | PRODUCTION CONTROLS: Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated July 6, 2011 via first class mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from their facility; call Baxter Healthcare's Center for Service at 1-888-229-0001 between 7:00 am and 6:00 pm Central Time, Monday through Friday to return affected product and receive credit; complete and return the attached CUSTOMER REPLY FORM to Baxter via fax at 1-847-270-5457 or scan and email to fca@baxter.com and if you are a dealer, wholesaler or distributor/reseller, notify your customers of this recall. If you have questions regarding the communication, call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. | Contact the recalling firm for information | 9800 | 9,800 sets | Worldwide distribution: USA (nationwide) including: AL, LA, MO, SD, TX and WV; and country of: Canada. | Terminated | April 18, 2012 | 278 | 9800 | Not_Computer
Z-2838-2011 | 58769 | restoration, porcelain | NSP |  |  | Dental | Enforcement Discretion | "***3i Incise Custom-Made Label. QP157 Approved 07/09/10*** 3i inc se Custom-made Dental Framework***Biomet 3i 4555 Riverside Drive Palm Beach Gardens, FL 33410 800-342-5454 www.biomet3i.com Please consult instructions for use and recommendation***For Dental Lab use ONLY Non-sterile/For single Use***"Indicated for use in restorative procedure to affix the final prosthesis to the abutment. Coping and frameworks provide the base structure for final porcelain application as the final step of creating a restoration. The devices are intended to be used to create final restoration for dental reconstruction with either a prepared tooth or dental implant abutment. | Incise Zirconia Coping (CBZR0102) | 3 | July 16, 2011 | 2011 | Biomet 3i, LLC | On December 1, 2010 Biomet 3i recalled their product Incise Zirconia Coping (CBZR0102). The Coping contains a manufacturing defect. Coping is oversized. All affected direct accounts were notified by letter on December 14, 2010. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomet 3i sent an "MEDICAL DEVICE RECALL" letter dated December 14, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken.The letter instructs customers to complete and fax the Product Recall Response form to 561-514-6316. In addition, a shipping label is enclosed with the letter for returning the affected product.Customers with questions and concerns should contact Biomet 3i recall coordinator at 561-776-6906. | No consumer action necessary | 16 | 16 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 18, 2011 | 2 | 16 | Not_Computer
Z-2840-2011 | 59155 | [6 empty] | ENCORE(R) UNDERGLOVE, Latex, Powder-Free, Smooth Surgical Gloves, REF 2018455, Size 5 1/2, Sterile, Manufactured for Ansell Healthcare, 1635 Industrial Road, Dothan, AL 36303, Made in Sri Lanka Surgical gloves | ENCORE(R) UNDERGLOVE | 2 | July 18, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves do not meet the tensile requirements of ASTM D 3577 reducing the shelf life of the product from 5 years to 2 years. |  | Ansell Healthcare Products LLC sent an Urgent Voluntary Recall letter dated June 9, 2011, to all affected customers on June 10 and 13, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue use and distribution of any of the affected product and return them to the firm. If the customer distributed gloves further, customers were advised to forward a copy of the notification letter along with their tracking number to their customers. Customers were instructed to complete and return the attached Response Form to Ansell by fax at 1-800-722-8155. Returned product would be replaced or credited to the customer. |  | 9580 | 9580 pairs | Nationwide Distribution | Terminated | January 10, 2013 | 542 | 1176176 | Not_Computer
Z-2839-2011 | 57518 | system, photopheresis, extracorporeal | LNR |  |  | Gastroenterology/Urology |  | Therakos Cellex Procedural Kits.Designed to interface with the Cellex Photopheresis System to perform cell separation and photoactivation in a single, closed and sterile circuit. | Therakos Cellex Procedural Kits | 2 | July 19, 2011 | 2011 | Therakos Inc | Cellex Procedural kits may have minor tears in their Tyvek covers. | PRODUCTION CONTROLS: Process Control | Therakos, Inc. sent an "Important Product Notification" dated December 2, 2010 to all affected customers.The letter included description of problem, affected lot and recommendations. Customers were asked to dispose any product with affected lot and to contact them for for replacement, and to complete and return the attached Fax reply form.For additional information contact Therakos Customer Technical Services at 877-865-6850. |  | 110 | 110 units | Nationwide (USA) distribution including the state of MI, MO, WI, CA, OH, NH, MA, and OR. | Open |  |  | 110 | Not_Computer
Z-2852-2011 | 58980 | kit, surgical instrument, disposable | KDD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | "***REF: 24-4050 TMJ SYSTEM ArthroSimplicity KIT***STERILE PRODUCT***STERILE EO***KIT CONTAINS BOTH STERILE AND NON-STERILE COMPONENTS DISTRIBUTED BY: BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" Kit Contains: Drape with pouch; 18Gx1.5in needle; OR towel; Towel clips; Skin marker with ruler; Band-aids; Cotton ball; 5ml syringe; Gauze sponges; Extension tubing; Cholarprep swabsticks; Utility Bowl. All-inclusive kit that alleviates the burden of piecing together each of the items necessary for a Arthroscopy using the OnPoint Scope. | Arthrosimplicity Kit | 2 | July 19, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling TMJ System OnPoint Scope Procedure Kit REF 24-3050; TMJ System ArthroSimplicity Kit REF 922070 & 24-4050; TMJ System InnerVue Scope Procedure Kit REF 922090. The OnPoint Scope Kits and ArthroSimplicity Kits contain Chloraprep Swabsticks that do not contain any warnings regarding the use of the product on eyes, ears or mouth. |  | Biomet Microfixation (formerly W. Lorenz) sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter to the consignees/customers on October 7, 2010. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to fill out the attached Inventory Reconciliation sheet and fax to 904-741-9425 and if the consignees/customers further distributed this product, they should notify their customers and provide them with a copy of this safety alert (the letter must be given to the personnel responsible for receiving field action notices). If you have any questions, please contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. | Contact the recalling firm for information | 103 | 103 | Worldwide distribution: USA (nationwide) including states of: AL, AR, CA, CO, CT, DC, CR FL,GA, HI, IA, IL, IN, KS, KY, LA, MA, MD,ME, MI MN, MO, MS, MT, NC, ND, NJ, NY, OH, O K, OR, PA, RI, SC, TN, TX, UT, VA and WA; and countries of: Argentina, Bahrain, Canada, Colombia, Costa Rica, and Finland. | Terminated | April 19, 2012 | 275 | 828 | Not_Computer
Z-2853-2011 | 59270 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161. Product Usage: Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. Electronic infusion pumps indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration, such as intravenous (IV), intra-arterial (IA), subcutaneous, epidural or irrigation of fluid spaces applications. | Colleague CX Infusion Pump | 2 | July 19, 2011 | 2011 | Baxter Healthcare Corp. | While replacing a defective pump head mechanism and performing calibration, Baxter's Service center may have failed to apply GLYPTOL (a red insulating varnish) on a potentiometer that helps maintain pump accuracy for detecting an air bolus. | PRODUCTION CONTROLS: Process Control | Baxter Healthcare Corporation telephoned its customers on July 1, 2011 and sent a follow up "URGENT PRODUCT RECALL" letter dated July 6, 2011, informing them that the Glyptol (a red insulating varnish) may not have been applied on the Air-In-Line (AIL) circuit board. The letter described the product, problem and actions to be taken. The customers were instructed to locate the affected pumps, remove from service and return to Baxter for inspection as soon as possible. If GLYPTOL is found to be missing, recalibration will be performed and GLYPTOL will be applied per requirement before returning the pumps back to the customer at no cost. The customer was instructed to complete and sign the attached customer reply form. The letter states that if the COLLEAGUE Pumps were distributed to other facilities, please forward communication to the pump owner as appropriate. For any questions contact The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time, Monday through Friday. |  | 2 | 2 units | Nationwide (USA) Distribution -- including the state of Pennsylvania. | Terminated | January 20, 2012 | 185 | 2 | Not_Computer
Z-2858-2011 | 59051 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Runthrough NS Tapered PTCA Guidewire, item code number TW*DP418FHProduct Usage The Runthrough NS guidewire is used to facilitate placement of balloon dilation catheters for Percutaneous Transluminal Coronary Angioplasty (PTCA) and/or Percutaneous Transluminal Angioplasty (PTA) | Runthrough NS Guidewire | 2 | July 20, 2011 | 2011 | Terumo Medical Corporation | Baloon catheter guidewires may become uncoiled during angioplasty procedures. | DESIGN: Component Design/Selection | Terumo Medical Corporation initiated contacting consignees on June 3, 2011 via an in-person visit to hand deliver an "URGENT PRODUCT BULLETIN: TERUMO PRODUCT SAFETY ALERT" letter dated June 3, 2011. The letter emphasized to consignees the importance of following the current instructions and recommendations provided to minimize the potential for damage during use and removal of the recalled product. For further questions, please contact your Terumo Territory Manager or call (800) 283-7866 and ask to speak with the Runthrough NS Guidewire Product Manager, at extension 4129. |  | 70 | 70 units | Nationwide (USA) Distribution -- including the states of: CT, GA, and SC. | Terminated | July 27, 2011 | 7 | 70 | Not_Computer
Z-2804-2011 | 58996 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GEM PREMIER 4000 CARTRIDGE K+ REPORTING:GEM 4000 BG/Hct/Lytes/G/L/CO-Ox 75 test iQM Part Number : 00027407510150 test iQM Part Number:00027415010300 test iQM Part Number: 00027430010450 test iQM Part Number: 00027445010600 test iQM Part number: 00027360010The GEM Premier 4000 is a portable critical care system for use by health care professionals to rapidly analyze whole blood samples at the point of health care delivery in a clinical setting and in a central laboratory. The instrument provides quantitative measurements of pH, pC02, p02, sodium, potassium, chloride, ionized calcium, glucose, lactate, hematocrit, total bilirubin and CO-Oximetry (tHb, 02Hb, COHb, MetHb, HHb) parameters | Instrumentation Laboratory | 1 | July 21, 2011 | 2011 | Instrumentation Laboratory Co. | K+ results on the GEM Premier 4000 are too low when compared to a reference analyzer, withbiases exceeding allowable error claim of aÂż 0.5 moUL | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Instrumentation Labs, issued a Revised "Urgent Field Safety Notification" letter dated July 14, 2011, to GEM Premier 4000 customers. The letter describes the product, problem, updated issue description and actions to be taken. The customers were instructed to follow the instructions to Disable K+ Parameter ,provided in the letter, to eliminate the potential of erroneous K+ result reporting and to complete and return the MANDATORY RESPONSE TRACKING FORM immediately via fax at 781-861-4207 or email to ra-usa@ilwww.com. The letter also noted an Important Change to Customer Actions from June Urgent Field Safety Notification.Prior strategy :Instrumentation Laboratories (IL) notified U.S and Canadian customers via the Urgent Field Safety Notification mailed USPS between June 2 - 3,2011. International Affiliates / Dealers notified through e-mail on June 1, 2010. A response (tracking) form was included. A Worldwide Regulatory Notice was e-mailed on May 31.2011.to IL country affiliates and distributors, notifying them of the issue, providing a risk summary and the support documentation for the notification. The customer notification will be handled per their local regulatory requirements.If you have questions regarding this notification, please contact IL at 800-678-0710, press 2 at prompt. |  | 24787 | 24787 | Worldwide distribution: USA (nationwide) and countries of: CANADA, ARGENTINA, AUSTRALIA ,AUSTRIA, BELGIUM, BRAZIL, CHILE, CHINA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PANAMA, POLAND, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SLOVENIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VIET NAM. | Terminated | November 27, 2012 | 495 | 28616 | Not_Computer
Z-2859-2011 | 59103 | injector, contrast medium, automatic | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiomat Illumena Syringe, 150mL Linden Luer Syringe with Handi-Fil Straw (p/n 900103), in a sealed molded plastic tray.The Liebel Flarsheim Illumena 150 mL Linden Luer Syringe with Handi-Fil Straw is designed for use with the Illumena Injector System. The Angiomat Illumena is designed to inject a radiopaque contrast medium into the vascular system for Angiographic or CT procedures as prescribed by qualified health care professionals. | Angiomat Illumena Syringe, 150mL Linden Luer Syringe with HandiFil Straw (p/n 900103) | 2 | July 21, 2011 | 2011 | Mallinckrodt Inc | Syringes packaged and labeled as the Illumena 150mL Linden Luer Syringe with Handi-Fil Straw, product 900103, contained the Illumena150mL Syringe with Handi-Fil Straw, product 900101. | PRODUCTION CONTROLS: Labeling Mix-Ups | COVIDIEN sent an Urgent Device Recall letter dated May 20, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use the attached business response form and report the quantity of lot number 0253304 currently in their stock. If the customer did not have any of the affected lot they were to indicate by checking the appropriate box on the form. When completed, customers were asked to fax to (314) 654-8206 or email to productmonitoringhazelwood@covidien.com. All product from lot number 0253304 was to be returned to Covidien. Once the business notification form is received a Return Goods Authorization for the quantity of product reported will be sent to the customer. Customers would receive credit for all return product from affected lot.For any questions regarding this recall customers were asked to contact the Product Monitoring department at 800-778-7898. For questions about credit for the product, call Customer Service at 888-744-1414, menu option 1, and then menu option 2. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 11000 | 11,000 units (200 cases/50 units per case) | Worldwide Distribution - USA including CA, FL, IL, IN, MI, MO, NY, OH, TN & TX. and the country of Australia. | Terminated | February 19, 2013 | 579 | 11000 | Not_Computer
Z-2836-2011 | 59161 | drape, surgical | KKX | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Kimberly-Clark, Three Quarter Drape, Non sterile/Disposable, Surgical drape. | KimberlyClark, Three Quarter Drape | 2 | July 21, 2011 | 2011 | Kimberly-Clark Corporation | Metal shavings were found within the folded portion of the surgical drape. |  | Kimberly-Clark Corporation sent an "Urgent: Voluntary Product Recall" letter dated June 6, 2011 to all affected customers and a subsequent letter dated June 22, 2011.The letter notified them of the recall and provided recommendations. They were instructed to evaluate their facilities unused inventory of the affected drapes to determine if any of the impacted product lots remain with their facility. If present, they were to quarantine and discontinue use. A Customer Product Recall Response Form was included to be completed and faxed to KC. They will be contacted by a representative regarding the return and replenishment of the product. If the product was further distributed, the consignee was advised to contact the end user, make them aware of the recall and to retrieve any of the affected product in inventory. For additional information please contact Kimberly-Clark Regulatory Affairs at 770-587-7400. |  | 223 | 223 cases (300 units per case) | Nationwide (USA) distribution, including the states of CA, FL, GA, IA, IL, MO, NC, NY, OH, TN, TX, VA and WI. | Terminated | October 14, 2011 | 85 | 223 | Not_Computer
Z-2855-2011 | 59291 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MatriStemÂŽ Plastic Surgery Matrix XS Intended for implantation to reinforce tissue where weakness exists in urological, gynecological, and gastrointestinal anatomy, including, but not limited to the following procedures: pubourethral support, tissue repair, body wall repair, and esophageal repair | MatriStem Plastic Surgery Matrix XS, | 2 | July 21, 2011 | 2011 | ACell, Inc | Surgical mesh for implantation may have elevated endotoxin levels. |  | ACELL, Inc. sent an "Urgent: Medical Device Safety Alert" letter dated July 07, 2011 to all affected customers.The letter notified customers (physicians) of the possibility of elevated endotoxin levels in the product and affected lots. They were instructed to monitor their patients receiving the affected implants.For additional information please contact the Director of Regulatory Affairs & Quality Assurance at (410) 953-8508. |  | 30 | 30 units | Nationwide (USA) Distribution including the states of CA, CT, FL, ID, IL, MD, MI, MS, NC, NJ, NV, NY, OH, OR, PA, TX and VA | Terminated | October 21, 2011 | 92 | 30 | Not_Computer
Z-2856-2011 | 59259 | gonococcal antibody tests | LGB | Gonococcal antibody test (GAT). | Microbiology | Microbiology | PMA | PALL Medical LeukoGuardÂŽ LGB Arterial Line Leukocyte Reduction Filter with Bypass Loop, STERILE. CATALOG/MODEL/Reorder Number: LGB. CLASSIFICATION NAME: Filter, Blood, Cardiopulmonary Bypass Arterial Line. COMMON/USUAL NAME: Arterial Line Filter with Bypass Loop. -- Manufactured by: PALL NEWQUAY, CORNWALL, TR9 6TT, UK -- Device Listing # D000745. --- Designed to reduce the levels of circulating leukocytes and remove microemboli greater than 40 microns in size from the perfusate during extracorporeal circulation including: gas emboli, fat emboli and aggregates composed of platelets, red blood cells and other debris. | Pall LeukoGuard LGB Arterial Line Leukocyte Reduction Filter with Bypass Loop | 2 | July 21, 2011 | 2011 | Pall Corporation | Potential for the housing of the arterial line Leukocyte Reduction Filter to separate before or during the priming procedure. | TRAINING: Employee Error | Pall Medical sent an "Important Customer Notification" letter dated May 26, 2011 to all affected customers.The letter included a description of the problem, affected lots and provided recommendations including the return of the product.For additional information, please contact the firm at (516) 484-0263. | Contact the recalling firm for information | 232 | 232 units | Nationwide (USA) Distribution including the states of CA, LA, MI, TX and WI. | Terminated | October 03, 2012 | 440 | 232 | Not_Computer
Z-2854-2011 | 59189 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | VirtuoSaph Plus Endoscopic Vessel Harvesting System , REF VSP550, Sterile, Terumo Cardiovascular SystemIndicated for use in minimally invasive surgery allowing access for vessel harvesting, and is primarily indicated for patients undergoing endoscopic surgery for arterial bypass. It is indicated for cutting tissue and controlling bleeding through coagulation, and for patients requiring blunt dissection of tissue including dissection of blood vessels and dissection of blood vessels of the extremities. Extremity procedures include tissue dissection and/or vessel harvesting along the saphenous vein for coronary artery bypass grafting and peripheral artery bypass grafting. | VirtoSaph Plus Endoscopic Vessel Harvesting System | 2 | July 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems received four (4) complaints from a single customer about tunnel breathing/inadequate insufflation with the VirtuoSaph Plus Endoscopic Vessel Harvesting System during cases in a limited production launch situation. Teumo's investigation identified issues with the devices's capability to maintain pressure within the tunnel. | DESIGN: Device Design | Terumo Cardiovascular Systems contacted the one customer to this device on June 10, 2010 by telephone.There was one consignee. The firm contacted this one consignee by phone on 6/10/2010. There were 2 contacts at this one consignee. The recalling firm contacted both of the contacts on 6/10/2010. The consignee was asked to dispose of remaking inventory of devices. The consignee said had disposed of the product.For information on this recall call (734) 741-6173. | No consumer action necessary | 35 | 35 units | Nationwide (USA) Distribution including the state of New York. | Terminated | July 27, 2011 | 6 | 35 | Not_Computer
Z-2158-2011 | 58419 | monitor, patient position, light-beam | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare DISCOVERY MODEL NM/CT 670NM/CT system | GE Healthcare DISCOVERY MODEL NM/CT 670 | 2 | July 21, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a missing laser aperture warning label (i.e. Laser Aperture - DoNot Stare Into Beam) affecting certain Discovery Systems. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will bring affected systems into compliance. A GE Healthcare Service Representative will provide a product label correction to address this issue. This will be carried out via a Field Modification Instruction (FMVI 408311 planned to be released in June 2011 with an estimated completion date of August 2011.For questions call 262-513-4122. | Contact the recalling firm for information | 4 | 4 | Nationwide Distribution | Terminated | August 31, 2012 | 407 | 4 | Not_Computer
Z-2864-2011 | 59201 | stretcher, wheeled | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Eye Surgery Stryker Medical Stretchers Model 1089Intended to transport patients in a horizontal position. | Eye Surgery Stryker Medical Stretchers Model 1089 | 2 | July 22, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | The firm has determined that the brake rod assemblies of various stretchers that were built with a 5th wheel were manufactured with roll pins that did not meet the materials specifications. Some roll pins may have fractured during the assembly process and could fall out of the brake rod assemblies, rendering the brake or steer systems inoperable from certain pedals or causing a reduction in overa | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker Medical sent an "Urgent Medical Device Correction" dated July 19 ,2011 to all affected customers.The letter included; the name and model numbers of the stretchers, the reason for the correction and dates of manufacture. It contained the instructions to check the braking system and if found to be inoperable, to remove it from service, and if operable to check prior to each use. Also, information regarding contacting the recipients of any loaned or sold stretchers to relay the information in the letter and Medwatch reporting information. A self addressed postcard to confirm receipt of the letter was also included.For additional information call (269) 389-6604. | Contact the recalling firm for information | 5 | 5 | Worldwide Distribution: Nationwide (USA) including the states of NY, MI, LA and CA ,and the countries of Canada, China, Spain, Japan, Korea, United Kingdom, South Africa. | Terminated | August 29, 2013 | 769 | 112 | Not_Computer
Z-2873-2011 | 59002 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***REF: 91-6704***THINFLAP SYSTEM CROSS-DRIVE, SELF-DRILLING SCREW 1.5X4MM (.059X.159") (WXL)***TITANIUM***NON-STERILE PRODUCT***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA*** www.biometmicrofixation.com***QTY 00001***".Internal fixation screws intended to aid in the alignment and stabilization of bone in the oral cranio-maxillofacial skeletal system. | THINFLAP SYSTEM CROSSDRIVE, SELFDRILLING SCREW 1.5X4MM (.059X.159") (WXL)TITANIUM | 2 | July 22, 2011 | 2011 | Biomet Microfixation, Inc. | On May 17, 2010, Biomet Microfixation issued a Safety Alert on their Thinflap System Cross Drive, Self Drilling Screw, 1.5x4mm, Model #: 91-6704, Lot#: 173770. The product is labeled incorrectly. Some packages of 91-6704 lot 173770 are labeled as cross drive screws, but contain center drive screws. All affected customers were notified of the recall on May 17, 2010. | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet Microfixation notified distributors and customers by letter, email and telephone calls. Distributors and customers were instructed to fill out and return the Inventory Reconciliation sheet, via Fax to 904-741-9425. Distributors were also instructed to notify their customers if the product was distributed further.For questions regarding this recall customers were instructed to call 1-800-874-7711 or 904-741-4400, ext. 468. | Contact the recalling firm for information | 927 | 927 | Worldwide Distribution - USA including AL, CA, GA, HI, IL, MO, NC, NY, TX and the country of Japan. | Terminated | October 12, 2011 | 82 | 927 | Not_Computer
Z-2866-2011 | 59167 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Avanta Fluid Management Injection System, Catalog numbers: AVA 500 TABL, AVA 500 PEDLProduct Usage: Administration of intravascular radio-opaque contrast media and common flushing solutions during angiographic procedures. | MEDRAD Avanta Fluid Management Injection System | 2 | July 22, 2011 | 2011 | Medrad Inc | Event was initially intended to be a product enhancement of the MEDRAD Avanta Fluid Management Injection System. Upgrades were initiated in October, 2007. Observance of a possible increase in complaints and adverse event reporting resulted in the firm issuing "Important Product Use Reminder" to remind users of proper procedures for using the system and the need for diligence in reducing air embo | DESIGN: Software Design | Medrad Inc., sent two types of "MEDICAL DEVICE FIELD CORRECTION" notifications on June 2, 2011 to all affected customers. 1. To those customers who Avanta Systems have already undergone market withdrawal and thus have been upgraded, the notification provided customer awareness that prior action should be considered a field correction. The letter stated that NO action is required since their system(s) has already been upgraded. If you have any questions, concerning these activities contact the Director of Quality Systems at (724) 940-6905. 2. To those customers whose Avanta Systems have not yet been upgraded through market withdrawal, the notification informed customers that the recalling firm is undertaking a Field Correction to upgrade their system(s). The letter included a step by step guide that will allow customers to verify that their system(s) has NOT been upgraded. Customers are instructed to complete and fax this form to MEDRAD at (412) 406-0941. If you have any questions, concerning these activities contact the Product Manager at (412) 767-2400, extension 8401. | No consumer action necessary | 842 | 842 | Worldwide Distribution -- (USA) (nationwide) including the state of Puerto Ricoand the countries of: AE, AL, AT, BE, BG, BR, CA, CH, CO, CY, DE, DK, DO, ES, FI, FR, GB, HK, HR, HU, IE, IL, IS, IT, JO, JP, KR, KW, LY, MA, MT, MX, MY, NL, PA, PH, PK, PT, QA, RS, RU, SA, SE, SG, SI, TR, VI, and ZA. | Terminated | June 05, 2012 | 319 | 842 | Software
Z-2871-2011 | 59077 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips brand BrightView Gamma Camera System,SYST, BRIGHTVIEW-TILT, 3/8", Model Number: 2170-3002A;Product is manufactured and distributed by Philips Medical Systems (Cleveland), Inc.,Cleveland, OHThe BrightView Gamma Camera System is intended to produce images depicting the anatomical distributions of single photon emitting radioisotopes within the human body for interpretation by medical personnel. | Philips brand BrightView Gamma Camera System, SYST, BRIGHTVIEWTILT, 3/8" | 2 | July 22, 2011 | 2011 | Philips Medical Systems | The gap between the two parts of the table may permit pinches to occur on either patient's skin or other body parts; it has the potential for serious injury if a patient's finger is caught between the pallet and patient table resulting in broken finger bone. | DESIGN: Device Design | Philips Healthcare issued a Customer Information letter (Field Change Order (#88200402)) dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instruction for use for their system, which states "Monitor the patient and the system vigilantly during any system motions to make sure that ...accessories, equipment, and all ojbects such as hair, jewelry, or neckties-remain clear of any moving camera and imaging table parts. For additional information, customers were istructed to see the BrightView SPECT instructions for Use, section 2.2 - Safety Warnings and Precautions and Safety and Information Labels. Philips Healthcare will notify all affected customers and will install a sticker through an Action for Performance - Proactive Field Change Order (FCO) to resolve the issue. For customers in North America who require further information or support contact Customer Care Solutions Center at 1-800-722-9377, Select Option 5. In all other countries the local Philips Healthcare office should be contacted. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, China, Chile, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Greece, India, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, Peru, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Switxerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, and the United Kingdom. | Terminated | November 06, 2012 | 473 |  | Not_Computer
Z-2879-2011 | 59146 | instrument, surgical, orthopedic, ac-powered motor and accessory/attachment | HWE | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | Orthopedic | 510(K) Exempt | Arthrex Reciprocating Saw AttachmentType: V-600SRDistributed by: Arthrex Inc., Naples, Florida United States The system Is a modular electric system, consisting of a handpiece drive, various accessories, and a charging station for machining bones. It can be used in the following areas of application: orthopaedic and A&E; Interventions, such as osteotomies, large and small bone operations, and joint replacement operations. The charging station is intended exclusively for charging V 600 Li-Ion and V 300 LiIon rechargeable batteries | Arthrex Reciprocating Saw Attachment | 2 | July 25, 2011 | 2011 | W & H DentalWerk Buermoos GMBH | Reciprocating saw attachment V 600 Sr (REF: 17306210), screws can become loose or come off (the attachment) and fall into the surgical wound and may not be noticed due to its small size. | PRODUCTION CONTROLS: Equipment Maintenance | Madrimed GmbH (Distributor partner) of W&H; Dentalmechanik GmbH sent a letter dated May 24, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were ask to advise their users to check the adapters thoroughly before and after each application, until the conversion of all attachments have been completed and to send all products that need to be converted to the nearest service centers.W&H; decided to convert all products that might be affected in their service centres and equip them with an additional security device.If you have any questions, call +43627462360 or email: regulatory@wh.com. | Contact the recalling firm for information | 194 | 194 units | Worldwide distribution: USA (nationwide) and countries of: Austia, Australia, Canada, Czech Republic, Denmark, France, Germany, Italy, Kuwait, Malaysia, Netherlands, Russia, Slovakia, South Africa, Spain, Taiwan, Turkey and United Kingdom. | Open |  |  | 194 | Not_Computer
Z-2875-2011 | 59021 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | CD HORIZON (R) LEGACY (TM) MAS, REF 75445540, pedicle screw, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132Intended to help provide immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic, lumbar, and/or sacral spine. | CD HORIZON (R) LEGACY (TM) MAS | 2 | July 25, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are longer then size etched on the screw. | PRODUCTION CONTROLS: Error in Labeling | Medtronic sent a Voluntary Recall letter dated June 28, 2011, directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. The letter identified the product, the problem, and the action to be taken by the affected customers. If appropriate, personal follow up with Sales Representatives will be made to assure that the recall communication is understood by Risk Managers and to facilitate return of product and the associated Recall response documentation. Foreign consignees will be notified by email on June 29, 2011.For any questions or concerns regarding this recall customers were to contact their Medtronic Sales Representative or the Global Quality Department at 800-876-3133, ext 3173. |  | 199 | 199 units | Worldwide Distribution - USA (nationwide) and the countries of MN, AK, FL, PA, ME, GA HI, CA, TX, OH, IN, MO, WI, MI, NY, IN, LA, MD, MA, NC, OR, NE WA, ID, CO, OR, UT, IL, WV, Canada, Singapore, Netherlands, Brazil, Columbia, Hong Kong, and Mexico. | Terminated | December 24, 2012 | 518 | 460 | Not_Computer
Z-2880-2011 | 59178 | stimulator, muscle, powered | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | eSwallow Multi-Program Dysphagia Therapy Test unitIntended Use: Powered Muscle Stimulator for treatment of dysphagia. | eSwallow MultiProgram Dysphagia Therapy Test unit | 2 | July 25, 2011 | 2011 | eSwallow USA LLC | It was discovered that the firm was distributing medical devices without FDA approval. | OTHER/UNDETERMINED: Pending | eSwallow USA LLC sent an URGENT MEDICAL DEVICE RECALL letter dated June 22 2011, via US mail and email notification to each distributors. The letter identified the product, the problem, and the action to be taken. Distributors were instructed to immediately examine inventory, quarantine all product subject to recall, and identify customers that received these products and notify them at once of the recall. The letter advised distributors to forward a copy of the recall notification letter to their customers. The letter states that the distributors are to return all inventories to the eSwallow Corporate Office eSWALLOWusa, 106 Hidden Drive, Scottsboro, Al 35769. The recall should be carried out to the customer/user level. Customers are instructed to complete and return the attached Response Form as soon as possible. If you have any questions, call (256) 571-0443. | Contact the recalling firm for information | 123 | 123 units | Nationwide (USA) Distribution - including states of AL, AZ, CA, GA, FL, LA, MS, and VA. | Terminated | July 03, 2012 | 344 | 123 | Not_Computer
Z-2882-2011 | 57968 | radioimmunoassay, dehydroepiandrosterone (free and sulfate) | JKC | Dehydroepiandrosterone (free and sulfate) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DHEA RIA ReagentThe DSL8900 DHEA radioimmunoassay is intended for quantitative determination of dehydroepiandrosterone in serum or plasma. This assay is intended for in vitro diagnostic use. | DHEA RIA Reagent | 2 | July 26, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that DHEA measurements using DHEA RIA (REF DSL-8900) kit lots identified may overstate actual DHEA levels in patient samples. | PRODUCTION CONTROLS: Process Control | Beckman Coulter sent an "URGENT: PRODUCT CORRECTION ACTION" letter dated November 15, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue using the affected products and discard them according to local regulations. Customers were asked to review patient DHEA results ran on the affected lots and interpret them in light of additional tests (such as LH, FSH, testosterone, DHEA-S, cortisol, 17-hydroxyprogesterone, and estradiol) and the patient's clinical status. A response form was enclosed for customers to complete and return.Contact the Beckman Coulter Technical Support Center at 1-800-854-3633 in the United States, or their local Beckman Coulter Representative for questions regarding this notification. |  | 3542 | 3,542 units (1181 in US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Hungary, India, Ireland, Italy, Japan, Mexico, Netherlands, Poland, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 04, 2012 | 283 | 3542 | Not_Computer
Z-2885-2011 | 59109 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. This software algorithm analyzes the patient's electrocardiographic (ECG) rhythm and indicates whether or not it detects a shockable rhythm. The LIFEPAK 500 AED requires operator interaction to defibrillate the patient.The LIFEPAK 500 AED is intended for use in the pre-hospital and workplace and community environments. It has been tested to RTCA/DO-160D; "Environmental Conditions and Test Procedures for Airborne Equipment" (refer to Specifications on page 5-15 of the Operating Instructions 3005338-001). The LIFEPAK 500 AED is intended for use by personnel who are authorized by a physician/medical director and have, at a minimum, the following skills and training: -CPR training -AED training equivalent to that recommended by the American Heart Association -Training in the use of the LIFEPAK 500 AED. | LIFEPAK 500 AED | 2 | July 26, 2011 | 2011 | Physio Control, Inc. | The potential reliability issue is the result of a design characteristic of a SRAM (Static Random AccessMemory) electronic component on the main printed circuit board assembly. | DESIGN: Component Design/Selection | Physio-Control, Inc. sent an "URGENT- ACTION REQUIRED" letter dated June 2011 to all affected customers. The letter describes the product, problem, and the action to be taken by the customer.The letter recommends that users who store their device(s) in hard-shelled carrying cases remove them from the case(s)immediately and discontinue the use of the case(s). Additionally, the letter recommends that customers perform regular maintenance in accordance with the Operating Instructions. A Confirmation Sheet was enclosed with the letter for customers to complete and return via fax at 1-866-448-9567. Customers can call the Technical Support Team at 1-800-442-1142, Option 5. |  | 12197 | 12197 units distributed worldwide (5169 units outside US; 7028 units in the US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Bahamas, Brazil, Canada, Chile, Columbia, Germany, Hong Kong, Iran, Japan, Mexico, Netherlands, Peru, and Uruguay. | Terminated | February 23, 2012 | 212 | 12197 | Hardware |  | Instructions | Safety Notice/Insructions
Z-2884-2011 | 58921 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | LPCCC2: Low Profile Abutment Non-Hexed Castable Cylinder by Biomet 3iCastable cylinders are intended to be used to produce a definitive casting for dental restoration onto which porcelain can then be cast | LPCCC2: Low Profile Abutment NonHexed Castable Cylinder | 3 | July 26, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling the Model LPCCC2 Low Profile Abutment, Non-Hexed Castable Cylinder, Lots 2010101036 and 2010101037  Exp. 10/2015. Package labeling may not correctly describe the package contents. Product was distributed in Europe only. | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet 3i sent a "Medical Device Advisory Notice" letter dated February 17, 2011 to all affected customers.The letter described the problem, instructed customers to returned any unused product and to confirm receipt of the notification.For additional information, please contact Customer Service at 800-342-5454. | No consumer action necessary | 232 | 232 unit | Worldwide distribution: http://webapps.cdrh.fda.gov/division-tracking/findTrackable.do?docNum=RCL110656&reviewType;=RCL | Terminated | July 27, 2011 | 1 | 232 | Not_Computer
Z-2881-2011 | 59101 | enzyme immunoassay, amphetamine | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Calloway LabsMulti-Drug 9 Panel Assay A rapid, one step screening test for the simultaneous for the qualitative detection of multiple drugs and drug metabolites in human urine. | Calloway LabsMultiDrug 9 Panel Assay | 2 | July 27, 2011 | 2011 | Alere San Diego | Alere is initiating a voluntary recall on the listed lots of MULTI-DRUG 9 PANEL INTEGRATED CUP manufactured by Alere for sole use by Calloway Labs due to a mistake that occurred while processing the order, these products were incorrectlylabeled as CLIA Waived, when in fact the devices are classified as Moderately Complex. Therefore these devices are misbranded. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Alere San Diego sent an "Urgent Medical Device Recall" letter dated June 9, 2011 to all affected customers.The letter included: products affected, reason for recall and instructed them to return all affected product. An acknowledgement form was also included for the customer to return to Alere.For questions contact Technical Services at 888.246.7483, option 2 or by e-mail at InnovaconComplaints@alere.com. |  | 2400 | 2400 kits (60,000 devices) | (USA) Nationwide Distribution including the states of MO and MA. | Terminated | April 04, 2012 | 252 | 2400 | Not_Computer
Z-2887-2011 | 59016 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, Part Number: 447070, Lot Number: M908398 & M911529.Rheumatoid Factor (RF) Reagent, when used in conjunction with IMMAGE Immunochemistry Systems and Calibrator 5 Plus, is intended for the quantitative determination of rheumatoid factor in human serum or plasma by rate nephelometry. | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent | 2 | July 27, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the listed lots of Rheumatoid Factor (RF) Reagent might produce false positive results. A false positive result might lead to additional unnecessary testing. |  | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE/PRODUCT CORRECTION" letter dated May 31, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products. Customers were asked to complete and return an attached FAX BACK Response Form to 714-961-4234. Replacement products may be obtained by calling Order Assistance at 800-526-3821 in the United States, (800) 463-7828 in Canada or contacting your local Beckman Coulter representative.For any questions regarding this notice, contact the Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the United States and Canada; or contact your local sales representative. |  | 6302 | 6302 units (850 in the US) | Worldwide Distribution--USA (nationwide) and the countries of Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Ireland, Italy, Japan, Kuwait, Macao, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, and United Kingdom. | Terminated | June 03, 2013 | 677 | 6302 | Not_Computer
Z-2857-2011 | 59288 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Global Focus Marketing and Distribution, Silencer S2200 Centrifuge | Silencer S2200 Centrifuge | 1 | July 27, 2011 | 2011 | Global Focus Marketing & Develop | Potential for bucket and insert not seated on the rotor to cause the lid to break and buckets to be ejected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Global Focus Marketing and Distribution, Ltd issued a press release on 07/19/11 and a URGENT: DEVICE RECALL letter was mailed to the firm's customers on 07/15/11 and 07/21/11. Discontinue use. For questions, contact: (800) 323-4306 or silencer@gfmd.com. |  | 859 | 859 units | Worldwide to USA, United Kingdom, Ireland, Canada, and China. | Terminated | November 06, 2013 | 833 | 859 | Not_Computer
Z-2889-2011 | 59346 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program Reorder Pack;contains 400 HbA1c or 200 HbA2/F/A1c tests. Kits include Elution Buffer 1 and 2; Calibrator 1 & 2.Bio-Rad Laboratories Inc. Hercules, CA.Product Usage: Short Proghram: D-10 Dual Program is intended for the percent determination of hemoglobin A1c in human whole blood using ion-exchange high-performance liquid chromatography (HPLC).Extended Program: The Bio-Rad D-10 Dual Program is intended for the percent determination of hemoglobins A2, F, and A1c, and for the detection of abnormal hemoglobins in human whole blood using ion-exchange high-performance liquid chromatography (HPLC).The Bio-Rad D-10 Dual Program is intended for Professional Use Only: For In-Vitro Diagnostic Use. | D10 Dual Program Reorder Pack | 3 | July 27, 2011 | 2011 | Bio-Rad Laboratories Inc | Reports of calibration failure and late retention time using Elution Buffer AA10434 and AA10435. | OTHER/UNDETERMINED: Pending | Bio-Rad Laboratories Inc sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected D-10 Dual customers and Bio-Rad subsidiaries on July 11, 2011 via FedEx or fax. The letters identified the product, the problem, and the action needed to be taken. The letters ask that inventory be examined and affected reorder pack lots be quarantined and destroyed according to local, state and federal regulations. All D-10 Dual customers were instructed to review their chromatograms generated from these lots. If the calibration has passed with good chromatography, the controls are within specification, and the F, A1c, and A2 peaks elute in their specified retention time windows for the patient samples, the customer may continue to use and report results with the affected buffer lots. The letters include a chart showing Good Patient Chromatogram with Correct Retention times and Failed Calibrator Chromatogram with late retention times. If calibration fails and/or late retention times cause one or more peaks to be misidentified, affected reorder packs should be destroyed. The correction letters and response form should be translated, if necessary, and distributed to all affected customers. Tracking documentation should be compiled for all customers who receive the Medical Device Correction letter. The Medical Device Correction Response Forms should be completed and returned by fax to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or a PDF copy to Jackie_Buckley@bio-rad.com. |  | 418 | 418 | Worldwide -- USA (nationwide) including the states of: AR, CA, CT, FL, HI, IN, MD, MN, MO, MS, NJ, NY, OH, PA, RI, TX. and the countries of: Brazil, Canada, China, England, France, India, Italy, Mexico, New Zealand, and S. Korea. | Terminated | March 12, 2012 | 229 | 418 | Not_Computer
Z-2888-2011 | 59023 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir Model 135 Convective Warming Device, Cincinnati Sub-Zero Products, Inc.Used to prevent hypothermia and/or reduce cold discomfort before, during, and after surgical procedures through convective heat transfer from the controller to a warm-air heated blanket | WarmAir Model 135 Convective Warming Device | 2 | July 27, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Due to its design the circuit breaker on the device could trip causing it to shut down and not provide heated air to the blanket. | DESIGN: Component Design/Selection | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 12, 2011 to all affected customers.The letter includes: description of product, problem, and information to request a repair tubing kit free of charge. It also asks customers to complete and return a "Recall Response Form".For additional information e-mail WA135-2011@cszinc.com or call (800) 989-7373 or (513) 772-8810. | Contact the recalling firm for information | 5474 | 5474 units | Worldwide Distribution: (USA) Nationwide, and the countries of Australia, Brazil, Bulgaria, Canada, Chile, Ecuador, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sri Lanka, Taiwan, Thailand, The Netherlands, Turkey, U.A.E., UK, Uruguay, Venezuela, and Vietnam. | Terminated | August 02, 2013 | 737 | 5474 | Hardware |  | Return | Remove or Replace
Z-2893-2011 | 59088 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerators;Varian Medical Systems, Palo Alto, CA.Radiation Therapy intended to deliver megavoltage x-ray treatments for conventional radiotherapy and stereotactic radiosurgery and radiotherapy. | Clinac Linear Accelerators | 2 | July 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The throat cover on the High Energy Clinac may detach if not properly installed and possibly strike a patient. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter provides directions on how to re-install the gantry center throat cover.Customers can contact their Varian Customer Support Manager at 1-888-827-4265 for further clarification. | Contact the recalling firm for information | 278 | 278 | Worldwide Distribution-including USA (nationwide) | Terminated | March 14, 2012 | 230 | 278 | Not_Computer
Z-2890-2011 | 59036 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(tm) MGI (tm) System Software, Catalog 445922, Lot V5.02A.Product Usage: The BACTEC(tm) MGIT (tm) System Software is used in the MGIT 960 and MGIT 320 instruments for the rapid detection of mycobacteria from clinical specimens. In addition, an Antimicrobial/Drug Susceptibility Testing (AST or DST) system is used as a rapid qualitative procedure for susceptibility testing of Mycobacterium tuberculosis, from culture, to streptomycin, isoniazid, rifampin, ethambutol, and pyrazinamide. The routine AST set consists of a Growth Control tube and a tube for each drug, as well as a barcoded tube carrier that holds the set. The carrier barcode provides set size and identification. A tube rack capable of holding up to 55 AST tubes in various set combinations is available for tube setup and transport. Large or small AST carrier sets can be placed in the rack. The MGIT AST system is based on growth of the Mycobacterium tuberculosis strain under test as measured by the instrument. AST sets are entered into the BACTEC MGIT instrument and are tested automatically. Analysis of fluorescence in a drug containing tube is compared to the fluorescence of the Growth Control tube to determine susceptibility results. | BD BACTEC(tm) MGIT (tm) 960 System | 2 | July 28, 2011 | 2011 | Becton Dickinson & Co. | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults. | DESIGN: Software Design (Manufacturing Process) | BD Diagnostic Systems verbally notified all affect customers via the telephone on March 28, 2011. The notification identified the product, the problem, and the action to be taken by the customer. An "URGENT PRODUCT RECALL" follow-up letter was mailed to all customers on April 8, 2011 via UPS. Customers were instructed to complete and fax the attached reply form to acknowledge their receipt of the notification to Regulatory Compliance, 410-316-4256. The letter instructs customers to refer to the modified release notes and IMPORTANT NOTICES outlined in the attached notification letter. If further assistance is needed regarding replacements and all other inquires, please contact BD Technical Services and Support at 1-800-638-8663. | Contact the recalling firm for information | 421 | 421 disks | Nationwide Distribution (USA) including states of: AK, AL, AR, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | July 29, 2011 | 1 | 421 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-2891-2011 | 59272 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | DeRoyal (R) Incision and Drainage Tray, RES 47-763.02, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 20641391, 21092240, 21950988, 22594153, 22739623, 22773346, 23068497, 23098901, 23435082, 24203902, 24725655, 24781729, 24901874, 25403527DeRoyal (R) Incision and Drainage Tray, RES 47-773, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 20432290, 20605752, 20922248, 21092178, 22615161, 23231338, 23567595, 24526171, 24781745DeRoyal (R) Incision and Drainage Tray, OCEAN SPRINGS HOSP, RES 50-12437.08, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 22871229, 23187891, 23282362, 23335049, 23693022, 24004508, 24289321, 24599046, 24853850, 25159510, 25609040, 20675048, 20886013, 21316397, 21529982, 21861715, 22203984, 22412903, 22729708, 22799765DeRoyal (R) Incision and Drainage Tray, ST JOSEPHS HOSPITAL, RES 89-6818.01, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 23606275, 23941604, 24336115, 24867670, 25220026, 25324301, 25814617, 21364022, 21372541, 21420961, 21469650, 21527602, 21641296DeRoyal (R) Universal Insertion Tray, RES 89-5649.01, ST FRANCIS MED CAPE GIRARDEAU, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA Lot Numbers: 22981727, 23234934, 21320071, 21374634, 21651531, 21817255, 22019861, 22187911, 22588503, 21651531 DeRoyal (R) St. Suture Removal Kit, RES 32-1695, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA, made in Dominican RepublicLot Numbers: 20170218, 21096486, 21114411, 21242931, 21331394, 21344427, 21381711, 23683529, 23767212, 23883758DeRoyal (R) Stainless Suture Removal Set, RES 32-351, Rx Only, STERILE EO, Manufactured by: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA, made in Dominican RepublicLot Numbers: 20188629, 20738627, 20838062, 21114437, 21410754, 21667171, 21790558, 22052557, 22389467, 22572739, 22865427, 23169950, 23452384, 23627906, 23883820, 24255913, 24408502, 24957430DeRoyal (R) CRICOTHYREOTOMY AMBULANCE KIT, WISHARD MEMORIAL HOSPITAL, RES 89-6752.01, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 20835339, 21511803, 23485821DeRoyal (R) DR MATA'S PLASTIC PACK, OUR LADY OF LOURDES, RES 89-7212.01, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Numbers: 22408453, 22653627, 22718814, 22777961, 23218782, 23942181, 24285119, 24571254, 24597462, 25414921, 25527677, 25646027, 25683611, 25744313, 25744321Custom Surgical kits. | DeRoyal Incision and Drainage Tray | 2 | July 28, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contain recalled Triad PVP prep pads | PRODUCTION CONTROLS: Process Control | DeRoyal sent a Customer letter dated July 6, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The customer was instructed to reference the attached spreadsheet to identify your accounts kits/trays and the corresponding lot numbers affected by this recall.Please complete and return the Acknowledgement Form by fax ( 865-362-3716) confirming you have been notified of this action.For further questions, please call ( 865) 362-2334. |  | 11856 | 11,856 kits/trays | Nationwide Distribution--including the states of CA, FL,GA, HI, IL, IN, LA, MA, MO, MS, NY, PA, and RI. | Open |  |  | 11856 | Not_Computer
Z-2892-2011 | 59377 | [6 empty] | Aquamantys 2.3 Bipolar SealerModel #23-113-1.Product Usage: The Aquamantys Bipolar Sealer is a sterile, bipolar electrosurgical device intended to be used in conjunction with the Aquamantys Pump Generator for delivery of RF energy and saline for hemostatic sealing and coagulation of soft tissue and bone at the operative site. It is intended for, but not limited to orthopaedic, spine, thoracic, and open abdominal surgery. The device is not intended for contraceptive tubal coagulation (permanent female sterilization | Salient Surgical | 2 | July 28, 2011 | 2011 | Salient Surgical Technologies, Inc | Insulating material on the shaft of the device may separate and expose the electrical conductor and result in burning the patient skin | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Salient Surgical Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" notification letter dated July 18, 2011 to all affected customers and distributors. The letter describes the product, problem and the actions to be taken. The letter requested that customers and distributors immediately identify and discontinue use of all Aquamantys 2.3 Bipolar Sealer devices in inventory and quarantine them until they can be returned to Salient. The letter states that a Salient sales representative will arrange return of the recalled product to Salient. The letter provides information for customers about the replacement of the recalled product. If you have any questions, please contact Customer Service, at 866-777-9400, or your sales representative. |  | 1011 | 48, 963 units | Worldwide Distribution - USA (nationwide) and countries ofChina, Denmark, France, Germany, Greece, Italy, Japan, , South Africa, Spain, and Turkey. | Terminated | March 09, 2012 | 225 | 48963 | Not_Computer
Z-2860-2011 | 59148 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow ON-Q with Select-A-Flow Pumps, Intended for continuous and/or intermittent delivery of medication to and around surgical wound sites and/or close proximity to nerves for preoperative, perioperative and postoperative regional anesthesia and pain management. It is a controller that allows the user to regulate the amount of medication that is dispensed. | IFlow ONQ with SelectAFlow Pumps | 2 | July 28, 2011 | 2011 | I-Flow Corporation | The recall was initiated because I Flow became aware that a small percentage of ON-Q Pump with Select A Flow may develop a minor leak with a flow rate controller. There may be consequences for the patient, such as under delivery of pain medication or the need to prematurely remove the pump, which would result in loss of the pain treatment provided by the pump. | PRODUCTION CONTROLS: Process Control | I-Flow Corporation sent an "Important Voluntary Market Withdrawal Notice" dated May 12, 2011 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to quarantine and return all impacted products within their facilities. Customers were instructed to respond with in 5 days of receiving the Market Withdrawal Notice and fax it to (949) 206-2663. Customers with additional questions, were instructed to call I-Flow Customer service at (800) 448-3569 (8-5pm, Pacific Standard time) or I-Flow 24 hour technical support line (800) 444-2728. |  | 96892 | 96,892 units total (95,637 units in US) | Nationwide (USA) Distribution and worldwide to: Australia, Canada, Cyprus, Italy, Germany, Finland, Greece, Israel, South Africa, Sweden, UAE, and UK. | Terminated | August 09, 2012 | 378 | 96892 | Not_Computer
Z-2897-2011 | 57856 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio Artida Ultrasound Diagnostic System SSH-880CV.Toshiba American Medical Systems, Inc.A diagnostic ultrasound system. | Aplio Artida Ultrasound Diagnostic System SSH880CV | 2 | July 29, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is initiating a field correction on the Aplio Artida Ultrasound Diagnostic System because the panel operation for the Aplio Artida may become disabled (i.e., system may hang up) when processing to display patient information (heart rate, time, VCP counter, ect.). | DESIGN: Software Design | Toshiba America Medical Systems, Inc. sent out an "Urgent Medical Device Correction" dated December 21, 2010 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Toshiba will issue a field modification to correct the software, and if the problem occurs, they were instructed to shutdown the system holding down the power supply button for 5 seconds for 5 seconds and then restart the system. A form to send back to Toshiba is included. For questions on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. | Contact the recalling firm for information | 46 | 46 units | Worldwide Distribution: (USA) Nationwide distribution including the states of MD, OH, PA, NY, MA, GA, AL, TX, IL, WI, IA, MN, KY, MT, OR, LA, and CA; and the country of Puerto Rico. | Terminated | April 20, 2012 | 266 | 46 | Software | Device Operation | Software update | Software Update
Z-2894-2011 | 59113 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Arrow ASK-20608-CMC StimuCath Continuous Nerve Block Kit. Arrow International, Inc.Permits placement of catheters next to nerves and nerve plexuses for continuous nerve block anesthesia or analgesia techniques for periods not exceeding 72 hours. | Arrow ASK20608CMC StimuCath Continuous Nerve Block Kit | 2 | July 29, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging were found to have an insufficient seal. Therefore, packaging may have been compromised, and sterility of the product cannot be guaranteed. | PRODUCTION CONTROLS: Packaging | Arrow International, Inc. sent an "Urgent Medical Device Recall" letter dated June 17, 2011 to all affected customers.The letter describes the problem and provides recommendations including; quarantine affected product, return of product, and send an acknowledgement form back to manufacturer.For additional information call 800 233-3187. | Contact the recalling firm for information | 95 | 95 kits | Nationwide Distribution including the state of North Carolina. | Terminated | October 01, 2012 | 430 | 95 | Not_Computer
Z-2896-2011 | 58031 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One: Aquilion TSX-301B, 64 System, X-Ray, Tomography.Toshiba American Medical Systems, Inc.A multislice CT system that supports whole-body scanning. | Aquilion One: Aquilion TSX301B | 2 | July 29, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction for the TSX-301A and TSX-301B: Aquilion 64 System, X-Ray, Tomography because communication between reconstruction units in the system console may be broken due to failure of the optical conversion connector, resulting in a break of communication between units. | DESIGN: Software Design | Toshiba America Medical Systems, Inc. sent an "Urgent Medical Device Correction" letter dated January 4, 2011 to all affected customers.The letter provides the customers with an explanation of the problem and the firms plan to install preventive measures and provided them with a Customer Reply form to send Toshiba. A representative will contact the customers for an appointment to install the measures to their CT system. For question on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 5 | 5 | (USA) Nationwide Distribution including the states of WI, MO, TN, IA, AL, AZ, TX, and GA. | Terminated | August 17, 2012 | 385 | 11 | Software
Z-2883-2011 | 59212 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Arrow NextStep Antegrade Chronic Hemodialysis Catheter.Product is used for attaining long-term vascular access for hemodialysis and apheresis. | Arrow NextStep Antegrade Chronic Hemodialysis Catheter | 1 | July 30, 2011 | 2011 | Arrow International Inc | Firm has received complaints indicating breakage and/or separation of the stylet within the Arrow(R) NextStep(R) Antegrade Catheters. | DESIGN: Device Design | An "Urgent Medical Device Recall" letter dated June 28, 2011 was issued to consignees. The letter requested that customers check their stock, cease use and distribution, and quarantine all affected product. Consignees were further directed to complete the "Recall Acknowledgement & Stock Status Form", return the form to their sales representative, and return affected product as instructed on the letter. For questions or concerns, call 1-800-233-3187. | No consumer action necessary | 41 | 41 | United States (states of CA, DE, FL, MI, NC, and TN). | Terminated | April 16, 2012 | 261 | 41 | Not_Computer
Z-2848-2011 | 59047 | condenser, heat and moisture (artificial nose) | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Vital Signs- A GE Healthcare Company - Hygroscopic Condenser Humidifier (HCH)/Anesthesia Breathing Circuit; GE Medical Systems SCS 283 rue de la Miniere 78530 BUC, France p +33.1.30.70.40.40; Vital Signs Colorado Inc 11039 E. Lansing Circle Englewood, CO 80112 P +1.973.956.5300.The intended use of the device is to provide proper flow of the medical gases or oxygen to patients under anesthesia or ventilation. | Vital Signs A GE Healthcare Company | 1 | July 30, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | The Hygroscopic Condenser Humidifier (HCH) device may have an occlusion that can prevent proper flow of the medical gases or oxygen, possibly resulting in insufficient oxygen delivered to the patient. | DESIGN: Device Design | An "Urgent Medical Device Correction" letter with the return confirmation form was sent to affected customers. Phone calls were also placed to customers. Consignees were informed of the affected device, issue, and actions to take. Customers were instructed to isolate all affected product, arrange return by contacting Customer Service at 1-800-932-0760, and return only unused product. Customers were asked to complete the attached confirmation form and fax back as per the instructions. For questions, contact your local Vital Signs Account Manager or call Customer service at 1-800-932-0760. |  | 222009 | 222,009 | Worldwide Distribution including the United States. | Open |  |  | 222009 | Not_Computer
Z-2913-2011 | 58920 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***BIOMET MICROFIXATION***CENTER DRIVE STAINLESS STEEL***2.0MM SYSTEM SCREWDRIVER BLADE***Non-Sterile Product***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***"Blade used to implant screws. | Center Drive Stainless Steel 2.0MM System Screwdriver Blade | 2 | August 01, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling Center Drive Blade, 2.0, Model Number 01-9047, Lot # 064980. Product is non-conforming (blades do not retain screws). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Biomet Microfixation, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" undated letter to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the affected products and immediately return them to Biomet Microfixation, Attn: Global Complaint Coordinator, 1520 Tradeport Drive, Jacksonville, FL 32218-2480; complete and return the Inventory Reconciliation sheet via Fax to the attention of: Global Complaint Coordinator at 904-741-9425, and notify their customers if product was further distributed. If you have any questions, contact the Global Complaint Coordinator at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. | Contact the recalling firm for information | 35 | 35 | Worldwide distribution: USA (nationwide) and countries of: Canada, Israel and South Africa. | Terminated | November 09, 2011 | 100 | 35 | Not_Computer
Z-2915-2011 | 59210 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476, Gantry S/N 9001-9023.The GEMINI TF Big Bore is a diagnostic imaging systems for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other organ imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode, dynamic, and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI TF Big Bore | 2 | August 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur, it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects. | DESIGN: Software Design (Manufacturing Process) | Philips Medical Systems (Cleveland) Inc. sent an "GEMINI TF BIG BORE CPE CUSTOMER LETTER" dated October 20, 2011 to all affected customers. The letter describes the product, problem, and updates that are available to the customers. The letter provides additional information about acquisition workflows for SUV Calibrations and Validations.For support concerning this upgrade, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 23 | 23 units | Worldwide Distribution-USA (nationwide) including the states of CO, LA, MD, MI, MO, MN, PA, TX, VA, VT and WI and the countries of Belgium, Germany, Italy, Saudi Arabia and Switzerland. | Terminated | September 21, 2012 | 417 | 23 | Software | Output/Calculation | Software update | Software Update
Z-2902-2011 | 59136 | analyzer, gas, nitrous-oxide, gaseous phase (anesthetic conc.) | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Ultraview Spacelabs Medical Multigas Analyzer, Model 91518, V1.00.14.Spacelabs Healthcare Llc.Monitors gas concentrations and alerts clinical personnel when the concentration of anesthetic agents, oxygen, carbon dioxide, or nitrous oxide falls outside of defined limits. The anesthetic agent being administered is automatically identified. | Ultraview Spacelabs Medical Multigas Analyzer, Model 91518 | 3 | August 01, 2011 | 2011 | Spacelabs Healthcare, Llc | Spacelabs received a report of the Model 91518 Multigas Analyzer trending low levels of an anesthetic agent that is not present during a particular case. The error is limited to only the trend table data. | DESIGN: Software Design | Spacelabs Healthcare, Llc. sent a "Customer Bulletin" dated June 20, 2011 to all affected customers. The customers were notified of the trend problem and are advised to follow the instructions in the Perioperative section and the ADMIT/DISCHARGE section of the users manual.For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support or 1-425-657-7200 x5089. | Contact the recalling firm for information | 481 | 481 units worldwide (333 units internationally and 148 units in the US) | Worldwide Distribution: (USA) Nationwide distribution, and the countries of ARGENTINA, BAHRAIN, BOLIVIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, ECUADOR, EGYPT, FRANCE, INDIA, MEXICO, NICARAGUA, PARAGUAY, PERU, POLAND, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, CHINA, THAILAND, UNITED ARAB EMIRATES, U.K., URUGUAY, VENEZUELA, and VIETNAM. | Terminated | September 21, 2011 | 51 | 481 | Software
Z-2901-2011 | 59046 | electrode, cutaneous | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | AxioBionics Wearable Therapy BioVest/Unitard that includes 2-16 BioGel Velcro Electrodes in sizes: 2"x2", 2"x2.5", 2"x3", 2"x3.5", 2"x4", 2"x5", 2"x6", 2"x8", 1.5"x3", 1.5"x5", 1.5"x6", and 1"x1.5".Cutaneous electrode garments for neuromuscular electrical stimulation. | AxioBionics Wearable Therapy BioVest/Unitard | 2 | August 01, 2011 | 2011 | Axiobionics, LLC | Premature Device Failure : Failure of the electrode sites to conduct and provide muscle stimulation due to silver fabric component failure. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Axiobionics, LLC is notifying all affected customers by phone. Three attempts will be made to contact the customer by phone. Contact by telephone will be made to prescribing physician to determine if Axiobionics has latest patient contact information. If information is different, new contact information will be recorded and 3 additional attempts will be made with the new information.If the patient cannot be reached by telephone, a certified letter will be sent to the patient's address. If contacting by mail, the letter template will be provided to the FDA recall office . All patients should have been contacted by telephone, certified mail, or in person by June 10, 2011.All patients affected by the field correction will be asked the following questions to determine if their electrodes are functional and resistances are not exceeding the output capabilities of the stimulator they are using. 1. Are you using your garment?2. Is your garment delivering the stimulation required for your therapy?3. If the patient is using an NP4 stimulator: What are the resistance readings on each channel?4. If the patient answers No to question 2 or states that the resistance is reading greater than 200 ohms on at least one channel in question 3, the garments will be returned to Axiobionics for repair.Garments that are not functioning correctly will be repaired first. Functioning garments will be tested by immersion into a salt bath under the given protocol to determine if the silver fabric is properly coated.Questions regarding this correction should be directed to the Axiobionics representative at 800-552-3539. | Contact the recalling firm for information | 7 | 7 | Nationwide Distribution--USA (nationwide) including the states of AZ, GA, IL, IN, MI, MO, NY, TX, and WY. | Terminated | August 29, 2013 | 759 | 53 | Not_Computer
Z-2910-2011 | 58698 | drills, burrs, trephines & accessories (simple, powered) | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ***REF MCA5-06SD, QTY: 1***0.6 mm Diamond Ball, 5.6 cm. Use with Micro Curved Attachment (MCA).***Sterilized using Irradiation***RX Only***ANSPACH Manufacturer***Made in USA***The Anspach Effort, Inc. 4500 Riverside Drive Palm Beach Gardens, FL 33410, USA Tel: (800) 327 6887 / + 1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP 11 1LA United Kingdom Tel: +441494 616126******REF MCA7-06SD, QTY: 1***0.6 mm Diamond Ball, 7.3 cm. Use with Micro Curved Attachment (MCA).***Sterilized using Irradiation***RX Only***ANSPACH Manufacturer***Made in USA***The Anspach Effort, Inc. 4500 Riverside Drive Palm Beach Gardens, FL 33410, USA Tel: (800) 327 6887 / + 1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP 11 1LA United Kingdom Tel: +441494 616126***Cutting and Shaping bone including bones of the spine and cranium | Anspach Single Use, Sterile Bone Cutting Burrs | 2 | August 01, 2011 | 2011 | The Anspach Effort, Inc. | On July 30th, 2010 the Anspach Effort, Inc. initiated a voluntary recall on their Anspach Single Use, Sterile Bone Cutting Burrs. A complaint was recieved stating that the product had failed during a procedure. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE PRODUCT REMOVAL" letter dated October 20, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customer to screen their inventory for the affected product and not to include any product received after October 18, 2010. Customers are instructed to contact Anspach Customer Support at 800-327-6887 to arrange for return and replacement of product. Customers are to complete and return a Customer Reply Form via fax at 800-327-6661.Questions regarding this notice should be directed to Anspach Customer Support at 800-327-6887.Customer notification began on 10/20/2010 in writing (Fed Ex) with instructions to return the required devices. Anspach utilized their sales representatives to facilitate the product replacement and return activities. | Contact the recalling firm for information | 68 | 68 | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, France, Italy, Japan, South Africa, Spain, Sweden, and United Kingdom. | Terminated | November 07, 2011 | 98 | 908 | Not_Computer
Z-2914-2011 | 59166 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs component of the Philips MX8000 Series Dual v. Exp. Computed Tomography System, Serial Numbers: 8008-8413, 80001-800022, Model #728130, 510 (k) #K010817.Product Usag: The MX 8000 series Dual v. Exp are Computed Tomography X-Ray system units intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | MX8000 Dual v. Exp CT XRay System | 2 | August 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems decided to conduct a recall/field correction of the patient support vertical brake hub component in all of their CT MX8000 Dual v. Exp Computed Tomography X-Ray System scanners due to a component defect which could cause an unexpected downward motion of the table (couch) while the patient is on the table. This event could occur because the short key connecting the vertical | DESIGN: Device Design | Philips Medical Systems (Cleveland) Inc. issued an "URGENT - Medical Device Correction - Field Safety Notice" letter dated June 1, 2011 to all affected customers, via Certified Mail for tracking purposes. The notification identified the product, the problem involved, and the action(s) to be taken by the customers with regard to this device issue. The notification instructs the customers to monitor their units and if the recalled product malfunctions, to immediately notify their Phillips service representative. The letter also states that a Philips Field Service Engineer will visit each customer location and perform a brake repair or replacement on their system to resolve the issue. For further information or support, the customers are instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada the customers are instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompt(s). |  | 412 | 412 system units were distributed. | Worldwide Distribution -- USA (nationwide) including the states of: AR, AZ, CA, CO, FL, FA, HI, IL, IN, KY, MA, MI, MO, MS, NC, NV, NY, OH, OK, SC, SD, TN, TX, VA, VT, WA and Puerto Rico and the country of Canada. | Terminated | August 07, 2013 | 737 | 412 | Not_Computer
Z-2916-2011 | 59441 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sign elance Monitor, Model 93300Spacelabs Healthcare, Llc.A vital signs patient monitor with arrhythmia detection. | Vital Sign elance Monitor, Model 93300 | 2 | August 02, 2011 | 2011 | Spacelabs Healthcare, Llc | Model 93300 Vital Signs Monitor may stop operating while in use, and when this occurs, the monitor may also fail to power ON. The resulting situation may cause a loss of monitoring for the patient. | DESIGN: Process Design | Spacelabs Healthcare, Llc sent an "Urgent-Medical Device Correction" letter dated 18 July 2011 to all affected customers. The letter notified the customers of the problem and recommended customers make their staff aware. It also informed them a Field Service Engineer will contact them to check, and if necessary make adjustments.Customers can contact Spacelabs Healthcare at 800-522-7025 and select 2 for Technical Support or outside the U.S. at +86512-8717-1888, extension 304. |  | 1970 | 1970 units | Worldwide Distribution: (USA) Nationwide distribution including the state of WI, and the countries of Puerto Rico, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Chile, China, Colombia, Cyprus, Dominican Republic, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Kenya, Korea, Kuwait, Lebanon, Malaysia, Napel, Netherlands, Nicaragua, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Republic of Macedonia, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Tunisia, Turkey, UAE, United Kingdom, Venezuela, and Vietnam. | Terminated | January 25, 2012 | 176 | 1970 | Other | Device Operation | Notification | Safety Notice/Insructions
Z-2919-2011 | 59102 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO versions equal to or greater than 2.4.0 and less than 4.63.0The XiO RTP System is used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. | XiO Radiation Treatment Planning System | 2 | August 03, 2011 | 2011 | Computerized Medical Systems Inc | The XiO Software currently prompts for entry of CT to ED data for a studyset. No indication is given that the Electron Density data should only be used for Photon and Electron calculations and either Relative Stopping Power or Mass Density should be used for Proton calculations (depending on the algorithm). | DESIGN: Software Design | The firm, Eletka, Inc., issued an "Important Safety Notice" dated May 18, 2011 via mail to all affected sites. The notice explains the issue and identifies a work around to the problem until the new release is available. The customers were instructed to complete and return the enclosed "Confirmation of Receipt" form via FAX to 314-993-1175, attn: QA; e-mail or mail: CMS Software, Attn: QA, 13723 Riverport Drive, Suite 100, Maryland Heights, MO 63043.The notice states that "this problem has existed since XiO Release 2.4.0 and has been resolved in XiO Release 4.63.0 by adding warning messages and explanations to the XiO On-Line Help. Your site will be notified when this release is available."If you have any question, call 314-993-0003, 800-878-4267 or 408-830-8023. | No consumer action necessary | 7 | 7 | Worldwide distribution: USA (nationwide) including: IL, IN, MA, NJ and OK; and country of Germany. | Terminated | April 04, 2013 | 610 | 7 | Other | Display/Image | Software update | Software Update
Z-2929-2011 | 58684 | culture media, antimicrobial susceptibility test, excluding mueller hinton agar | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel Haemophilus Test Medium (Agar), 150mm plates, 10/pk, Ref. #R04033. The firm name on the label is Remel, Lenexa, KS.For use in qualitative procedures by the Clinical and Laboratory Standards Institute (CLSI) as the growth medium for antimicrobial disk diffusion tests with Haemophilus species. | remel Haemophilus Test Medium (Agar) | 2 | August 03, 2011 | 2011 | Remel Inc | Organism strains may fail to consistently produce confluent growth | OTHER/UNDETERMINED: Pending | The firm, Remel, Inc., sent two "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letters one dated April 20, 2011 to: Microbiology Lab Supervisors and the second dated April 25, 2011 to:Distributors via regular mail. The letters described the product, problem and actions to be taken. The customers were instructed to to review their inventory for the listed product lot(s); discard all remaining units of the product lots; and complete and return the Product Inventory Checklist by May 9, 2011 regardless if they have affected product in the enclosed self-addressed, stamped envelope. The distributors were further instructed to notify their customers if they further distributed the product. If you have inquiries concerning this advisory, please contact our Technical Services Department at 800-255-6730 or 913-888-0939. |  | 548 | 538/10-pack boxes | Nationwide distribution. | Terminated | June 21, 2013 | 688 | 781 | Not_Computer
Z-2918-2011 | 58978 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Aruba Wireless Network Access Points, Model 65 utilized by the Acuity Central Monitoring System.Welch Allyn Protocol, Inc.Intended for use by clinicians for the central monitoring of neonatal, pediatric and adult patients in health care facilities. The system connects to a network of patient monitors to record and analyze physiological data being acquired by the monitors. | Aruba Wireless Access Points, Model 65 | 2 | August 03, 2011 | 2011 | Welch Allyn Protocol, Inc | Aruba Wireless Network Access Points, Model Number 65, experienced a higher than expected incidence of dropouts due to external RF interference. | DESIGN: Software Design | Welch Allyn Protocol, Inc. sent a letter dated May 16, 2011 to affected customer.The letter stated there were concerns over drop-outs in the system recommended the system at this hospital to revert back tothe Symbol wireless infrastructure. For questions call Welch Allyn at 315-685-4100. | Contact the recalling firm for information | 12 | 12 units were affected | Worldwide distribution: (USA) Nationwide distribution including the state of California; and the countries of Singapore; Philippines, France, Italy, Malaysia, Poland, UK, England, UAE, Germany, Spain, Ireland, Portugal, New Zealand, Hong Kong, Thailand, Australia, Canada, and the Netherlands. | Terminated | August 12, 2011 | 9 | 12 | Other
Z-2931-2011 | 58852 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Pinnacle Pelvic Floor Repair Kit--Anterior/Apical STERILE, and Pinnacle Pelvic Floor Repair Kit--Posterior STERILE.Boston Scientific, Corp.Indicated for tissue reinforcement and stabilization of fascial structures of the pelvic floor for vaginal wall prolapse, where surgical treatment is intended, either as mechanical support or bridging material for the fascial defect. | Pinnacle Pelvic Floor Repair KitAnterior/Apical, and Pinnacle Pelvic Floor Repair KitPosterior | 2 | August 03, 2011 | 2011 | Boston Scientific Corporation | The device may exhibit low tensile strength between the needle and suture and lead to needle detachment during mesh leg placement. | DESIGN: Device Design | Boston Scientific Corporation sent on May 10, 2011 an "URGENT MEDICAL DEVICE RECALL-IMMEDIATE ACTION REQUIRED" letter to all affected customers.The letter included; reason for recall, instructions for the quarantine, discontinued use, and return of the devices, and information for Medwatch reporting. The letter also includes a form that is to be faxed to the firm.For information call (508) 683-4678. | Contact the recalling firm for information | 540 | 540 | Worldwide Distribution; (USA) Nationwide distribution, and the countries of Puerto Rico, France, Great Britain, Ireland, Italy, Spain, Sweden, Canada, and Australia. | Terminated | November 15, 2011 | 104 | 540 | Not_Computer
Z-2932-2011 | 59356 | catheter, urological | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, GB: Female Catheter, All Purpose PVC, 08 Fr. 18 cm, Distributed by Teleflex Medical, Research Triangle Park, NCProduct Usage: Catheter that is inserted through the urethra and used to pass fluids to or from the urinary tract. | RUSCH | 2 | August 03, 2011 | 2011 | Teleflex Medical | Pin holes were detected in the packaging therefore sterility of the device can not be guaranteed. | OTHER/UNDETERMINED: Pending | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 30, 2011 to all affected consignees. The letter described the product, problem and actions to be taken. The letter instructs the customers to discontinue use and quarantine any of the affected products. Customers were asked to complete the attached Recall Acknowledgement Form and return as per the instructions. A separate letter was sent to distributors instructing them to communicate the recall to any of their customers who may have receive the affected products, using the recall letter and form as a template. For questions, contact your local sales representative or call Customer service at 1-866-246-6990. | Contact the recalling firm for information | 14378890 | 14,378,890 units | Worldwide Distribution and (USA) Nationwide Distribution. | Open |  |  | 14378902 | Not_Computer
Z-2927-2011 | 59214 | stretcher, wheeled, powered | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Stryker Medical Stretcher Model 1105 Prime 5th WheelA wheeled stretcher consists of a platform mounted on a wheeled frame that is designed to transport patients in a horizontal position. The device has siderails and supports for fluid infusion equipment. | Stryker Medical 1105 Prime 5th Wheel | 2 | August 03, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that some Prime Brake Pad Assembly components were manufactured with brake shoes that did not meet specifications. Accordingly some brake shoes used within the Prime Brake Pad Assembly may wear prematurely, over the life of the product. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical intervention. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker Medical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 18, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to locate the affected stretchers and to test the brakes. If the brakes are found to be inoperable, customers are instructed to remove the stretchers from service and contact the firm at 800-STRYKER, option 3. Additionally, customers are instructed not to put the stretchers back in service until they are repaired by the firm's field service representative. For units with properly functioning brakes, the letter also instructs the user to test the brakes before each use until the unit has been serviced. A post card was enclosed to confirm receipt of the letter. Questions or concerns may be addressed by calling Stryker at 269-389-6604. | Contact the recalling firm for information | 23 | 23 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, FL, GA, HI, IL, MA, MN, MO, NH, NJ, NV, NY, PA, RI, TX, UT, and VA and the country of THE NETHERLANDS. | Terminated | August 28, 2013 | 756 | 164 | Not_Computer
Z-2920-2011 | 59419 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set with Duo-Vent Spike; 109" (2.8 m), Male Luer Lock Adapter 10 drops/mL.Baxter Healthcare Corporation.A sterile, single use, nonpyrogenic fluid pathway for use with rigid nonvented and flexible containers. | Baxter Solution Set with DuoVent Spike | 2 | August 03, 2011 | 2011 | Baxter Healthcare Corp. | The Solution Sets are being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | TRAINING: Employee Error | Baxter Healthcare Corporation sent a "Urgent Product Recall" letter dated July 21, 2011 to all affected customers.The letter included a description of the problem and requested they contact Baxter at 1-888-229-0001 to return all affected product. They were asked to complete and return the attached customer reply form, and to contact and notify their customers.For questions contact Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 4752 | 4,752 units | (USA) Nationwide Distribution including the states of Alabama, California, Iowa, Illinois, Indiana, Massachusetts, Maryland, Michigan, North Carolina, New Hampshire, Ohio, Pennsylvania, Tennessee and Virginia. | Terminated | June 18, 2012 | 320 | 4752 | Not_Computer
Z-2930-2011 | 59300 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Intuitive Surgical da Vinci Si Surgical System (IS3000);Mfg by: Intuitive Surgical, Sunnyvale, CA 94086Product Usage: Intended to assist in the accurate control of Intuitive Surgical Endoscopic Endo Wrist Instruments and Accessories. The system is indicated for adult and pediatric use (except for transoral otolaryngology surgical procedures). It is intended for use for trained physicians in an operating room environment in accordance with the representative, specific procedures set forth in the Professional Instructions for Use. | Intuitive Surgical da Vinci Si Surgical System (IS3000) | 2 | August 03, 2011 | 2011 | Intuitive Surgical, Inc. | Potential failure of the retention component of the Master Tool Manipulator (MTM) of the Intuitive Surgical da Vinci Si Surgical System which could cause uncontrolled movement. | DESIGN: Device Design | Intuitive Surgical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 1, 2011 to all affected customers. Customers were asked to complete the attached response form and return as per the instructions. The letter identifies the product, problem, and actions to be taken by the customers. The letter reinforce to customers to follow the product labeling WARNING. The letter states that all affected customers sites will be receiving a call from a Customer Support staff to schedule a retrofit of the component. A letter was also sent to non affected customers to inform them of the urgent medical device correction and actions to be taken. If you have any questions, please contact Customer Service at 1-800-876-1310 or for Outside United States call +41 21 821 2020. |  | 183 | 183 | Worldwide Distribution -- USA (nationwide) and the countries of: Belgium, Czech Republic, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Netherlands, Norway, Philippines, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2011 | 30 | 183 | Not_Computer
Z-2917-2011 | 59301 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Automated External Defibrillator (non-wearable). Model number 9200, 9200N, 9200DN, 9200RD, 9210N, 9210RD, and CardioVive 92530.The battery component is warranted for 5 years, and electrodes have a shelf life of 2 years.The Powerheart AED G2 and CardioVive 92530 are intended to acquire the ECG rhythm for the detection of, and to provide treatment for, ventricular tachyarrhythmias of patients who are at risk of sudden cardiac arrest. The device is intended to be used in medically supervised environments by trained personnel, in which patients are under the direct care of physicians and/or medical persons authorized by the state, province, or country regulations in which they practice. | Powerheart AED, G2 | 2 | August 03, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) intent to end service support and discontinue accessory sales of Powerheart G2 Automated External Defibrillators. Customers are recommended to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. | PRODUCTION CONTROLS: Manufacturing Material Removal | Cardiac Science Corporation (CSC) sent two "9200 SERIES CARDIAC SCIENCE AED NOTICE" letters dated August 2010 and March 2011 to all affected customers. The letters identifies the product, problem, and actions to be taken by the customers.Customers are encouraged to upgrade their devices through special offers by calling CSC at 800-426-0337, Option 2.For questions about the end of factory support for the affected devices, call the Technical Support team at 800-426-0337, Option 1.*****In August 2011, Cardiac Science sent out another updated recall notification letter to their consignees. This letter stated that once a device has entered the "Service Required" state, please contact the firm at 800.426.0337 Option 2. Per prior notifications, the 9200 series AEDs have reached the end of service life and can no longer be repaired. Information about these notifications may be found as an announcement in the Services and Support section of Cardiac Science website.Cardiac Science advised the customers to upgrade to a Powerheart AED G3 (this device supports the 2010 AHA/ERC Guidelines). It is stronglyrecommended that customer to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. For information about special offers to help you upgrade please call 800.426.0337. Option 2. | Contact the recalling firm for information | 7304 | 48,700 units (41396 units in US; 7304 units Internationally) | Worldwide Distribution-USA (nationwide) and the countries of AE, AT, AU, AW, BE, BR, CA, CH, CL, CN, CO, CY, CZ, DE, OK, EG, ES, FI, FR, GB, HK, HR, HU, 10, IE, IL, IN, IR, IS, IT, JP, KR, KW, KY, LB, LT, LV, MT, MX, NG, NL, NO, NZ, PE, PF, PL, PR, PT, QA, RO, RU, SA, SG, SK, SY, TH, TR, TT, TW, VG, ZA, and ZW. | Terminated | December 20, 2011 | 139 | 48700 | Not_Computer
Z-2945-2011 | 59058 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard PS Open Box Femoral 70 MM Left Interlock, " FOR CEMENTED USE ONLY" CO-CR-MO ALLOY.Biomet, Inc.A knee joint replacement prosthesis non-coated femoral indicated for cemented application only. | Vanguard PS Open Box Femoral 70MM Left Interlok | 2 | August 04, 2011 | 2011 | Biomet, Inc. | The recalled device was manufactured with two distal pegs but package labeling indicates the device is pegless. | TRAINING: Employee Error | Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to quarantine and return the devices, and a return form.For information or questions on this recall call Biomet at (574) 371-3756. |  | 8 | 8 | Worldwide Distribution: (USA) Nationwide including the states of NJ, PA, LA, SD, TX, and KS; and the country of The Netherlands. | Terminated | September 26, 2012 | 419 | 8 | Not_Computer
Z-2948-2011 | 59252 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Gamma Camera, 3/8".Philips Medical Systems, Inc.Intended to produce images depicting the anatomical distributions of single photon emitting radioisotopes within the human body for interpretation by medical personnel. | BrightView XCT Gamma Camera, 3/8". | 2 | August 04, 2011 | 2011 | Philips Medical Systems | Non-conformances with BrightView X/XCT systems running version 2.0.1 including data transfer, XCT acquisition, Pixel Mask for image reconstruction and table moving incorrectly. | DESIGN: Software Design | Philips Medical Systems, Inc. set a "Customer Information" letter dated July 8, 2011 to all affected customers.The letter informed them of an upcoming software updated (version 2.5).For questions please contact 800-722-9377 (option 5). |  | 213 | 213 units, all systems | Worldwide Distribution: (USA) Nationwide distribution and the countries of Greece, Australia, Belgium, China, Czech Republic, France, Germany, Indonesia, Israel, Japan, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Puerto Rico, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan U.A.E., and the UK. | Terminated | February 15, 2013 | 561 | 213 | Software
Z-2950-2011 | 59454 | reservoir, blood, cardiopulmonary bypass | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | CAPIOXÂŽ Flexible Venous Reservoirs, 1500 mL, product codes: CXFVR1500X and CXFVRI500XC, used singly or packaged as components in cardiovascular procedure trays labeled in part ***Terumo Caerdiovascular Systems Corporation, Elkton, MD 21921***Product Usage: The reservoir bag is intended for use as a blood collection and storage device during cardiopulmonary bypass procedures. The device is intended for use in conjunction with blood-gas oxygenators in the extra-corporeal circuit. The device may be used in procedures lasting up to 6 hours. | CAPIOX Flexible Venous Reservoir | 2 | August 04, 2011 | 2011 | Terumo Cardiovascular Corporation | Reservoir blood bag used in cardiovascular procedures may leak. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated July 20, 2011, to all affected consignee. The letter described the problem, the potential hazard, and the actions to be taken. The letter instructed customers to identify affected inventory and segregate from non-affected product. The recalling firm advises users to not use affected product and return for replacement. The customers are instructed to call the recalling firm Customer Service for a Returned Goods Authorization (RGA) number. Customers were asked to complete the attached Response Form and return as per the instructions. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM EST. |  | 253 | 253 reservoirs | (USA) Nationwide Distribution including states of: FL, MI, MT, NY, TN and VA. | Terminated | February 10, 2012 | 190 | 253 | Not_Computer
Z-2946-2011 | 59368 | catheter, embolectomy | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, ProntoÂŽ V4, Extraction Catheter, Sterilized using ethylene oxide, Rx Only. The following sizes/model numbers are affected: 5.5F( 4005), 6F(4006, 7F(4007, 8F(4008) are involved in the recall.The Pronto catheter is indicated for the removal of fresh, soft emboli and thrombi from vessels in the coronary and peripheral vasculature. | Vascular Solutions, Pronto V4, Extraction Catheter | 2 | August 04, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions , Inc recently discovered that there was a manufacturing defect with Pronto V4 extraction catheters (Models 4005, 4006, 4007, 4008). Use of this defective product could lead to a tip separation. This maycause injury and/or procedural delay, resulting in a deteriorating hemodynamic condition of the patientwhich may require medical intervention. | PRODUCTION CONTROLS: Process Control | Vascular Solutions, Inc. sent an "URGENT: MEDICAL DEVICE FIELD ACTION" letter dated July 8, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify the affected product and remove from current inventory. Included with the letter was a Field Action Customer Inventory Form for customers to complete and return. VSI will contact customers for replacement product upon receipt of the form and returned product.Contact your local representative for questions concerning this notice. | Contact the recalling firm for information | 1086 | 1086 | Nationwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, FL, ID, IL, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK ,OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | April 03, 2012 | 243 | 1086 | Not_Computer
Z-2952-2011 | 59382 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Booster Laser Guide Catheter, 8 Fr, REF 518-063, Sterile EO, Spectranetics Corporation, 96 Talamine Court, Colorado Springs, CO 80907.Designed for directing and supporting Spectranetics laser catheters for use in the treatment of infrainguinal stenoses and occlusions. | Laser Guide Catheter | 2 | August 04, 2011 | 2011 | Spectranetics Corporation | Packaging integrity may be compromised, thus sterility can not be assured. | DESIGN: Packaging Design/Selection | Spectranetics sent an "PRODUCT RECALL NOTICE: EU SHIPMENTS OF TURBO-BOOSTER LASER GUIDE CATHETERS" letter dated July 15, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and set aside affected product. Customers are to contact their sales representative for information on how to return the affected product and receive replacement product.Further information is available at +011.49.719.447.2539. | Contact the recalling firm for information | 123 | 123 units | Foreign Distribution-including the countries of Belgium, Cyprus, Germany, Greece, Italy, Spain, Switzerland, and United Kingdom. | Terminated | August 12, 2011 | 8 | 180 | Not_Computer
Z-2953-2011 | 58043 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron JR. Microcoagulation Systems Prothrombin Time (PT). International Ethidine Corp (ITC)-Nexus Dx.A unitized mlcrocoagulation test intended for use in performing one-stage prothrombin time. | Hemochron JR. Microcoagulation Systems Prothrombin Time (PT). | 2 | August 05, 2011 | 2011 | International Technidyne Corp. | Certain lots of Hemochron Jr. PT Cuvettes (J201) demonstrate a higher bias than historically observed when compared to laboratory reference instruments. | DESIGN: Device Design | International Technidyne Corp. (ITC) sent "Urgent Medical Device Recall" letters dated May 10, May 25 and June 09, 2011 to all affected customers.The letter included a description of the problem, asked customers fill out the attached form, and return affected product for replacement.For questions, please contact ITC Technical Support at (800) 631-5945. |  | 4713 | 6,218 boxes in US ; 4713 boxes outside US | Worldwide Distribution: (USA) Nationwide Distribution including the countries of; Austria, Australia, Canada, Switzerland, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Korea, Liechtenstien, Libya, Netherlands, Norway, Oman,Portugal, Qatar, Saudi Arabia, Sweden, Turkey, Taiwan, and South Africa. | Terminated | November 02, 2012 | 455 | 10931 | Not_Computer
Z-2956-2011 | 59379 | apparatus, autotransfusion | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT -Cardiovascular Perioperative Autotransfusion SystemModel 2050 USModels Out-Side the US:02050-UK, 02050-ES, 02050-IT, 02050-NL, 02050-GE, 02050-FR, 02050-JAThe HaemoneticsÂŽ Cardiovascular Perioperative Autotransfusion Systemfor Use (cardioPAT) is indicated for use to salvage red blood cells from blood lost intraoperatively and postoperatively during cardiovascular surgicalprocedures, where the expected rate of processing of salvaged blood and fluidaspirated from the surgical site is less than or equal to two liters per hour. | Haemonetics cardioPAT Cardiovascular Perioperative Autotransfusion System | 2 | August 05, 2011 | 2011 | Haemonetics Corporation | Updated Operations Manual for the deployment of the spill collection system and cleaning | DESIGN: Component Design/Selection | Haemonetics sent a "FIELD NOTIFICATION: UPDATE OF CardioPAT OPERATIONS MANUAL FOR DEPLOYMENT OF SPILL COLLECTION BAG AND CLEANING" letter dated July 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter requests that customers complete an enclosed Customer Confirmation sheet indicating that the affected devices are in their possession and to return the sheet via fax to Haemonetics at 781-987-9252. Additionally, instructions for deployment of the spill collection drainage system, revised Operations Manual and Quick Reference Guide Addendum are enclosed with the letter.Contact Haemonetics at 1-800-537-2802 for questions regarding this notice.. | Contact the recalling firm for information | 316 | 316 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, France, Germany, Great Britain, Italy, Japan, Spain, and The Netherlands. | Open |  |  | 316 | Not_Computer
Z-2955-2011 | 59392 | cable, electrode | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Supreme Electrophysiology Extension Cable, REF 401984, For use with Hexapolar Catheters, Sterile EO, Rx only, St Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55345.Supreme Electrophysiology Extension Cable used to connect Electrophysiology catheters to external diagnostic or therapeutic equipment. The section indexing connector located a t the distal end of the cable will only connect to SMJ Electrophysiology catheters. | Supreme Electrophysiology Extension Cable | 2 | August 05, 2011 | 2011 | St. Jude Medical | St. Jude Medical is recalling a single lot (3334664) of Supreme Electrophysiology Extension Cable which were packaged with a Quadripolar cable instead of a hexapolar Cable. St. Jude Medical has determined that the incorrectly labeled product poses no safety risk since the extension cable can be safely used with a corresponding Quadripolar catheter but cannot be incorrectly connected to a Hexap | PRODUCTION CONTROLS: Labeling Mix-Ups | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE FIELD ADVISORY NOTICE" letter dated June 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate the affected product and to complete and return an attached SJM AFD Reconciliation Form. The form will be used to replace product and as documentation that the customers received the notice. Fax the completed form to 877-497-031.Contact the Director of Product Marketing at 512-329-1054 for questions regarding this notice. | Contact the recalling firm for information | 32 | 32 | Nationwide Distribution-USA (nationwide) including the states of FL, NC, NY, and OR. | Terminated | August 16, 2012 | 377 | 32 | Not_Computer
Z-2959-2011 | 59119 | catheter, peritoneal, long-term indwelling | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Aspira Drainage Kit, Product Code 4991502. The Aspira Peritoneal Drainage system is indicated for intermittent drainage of recurrent and symptomatic malignant ascites. The catheter is intended for long-term access of the peritoneal cavity in order to relieve symptoms such as dyspnea or other symptoms associated with malignant ascites. | Aspira Peritoneal Drainage system | 3 | August 08, 2011 | 2011 | Bard Access Systems | Labeling for drainage kit contained incomplete information. | PRODUCTION CONTROLS: Error in Labeling | The firm, BARD, sent a letter entitled "Labeling Discrepancy on Aspira Drainage Accessory Kits" dated June 12, 2011 to its customers. The letter provides the customers with the information missing from the labeling and/or IFU stating that the product can be safely used for peritoneal and pleural applications.If you have any questions please do not hesitate to contact BARD at 1-800-290-1689. If there are any clinical questions please contact Medical Services and Support at 1-866-893-2691. |  | 39290 | 39,290 units | Nationwide distribution. | Terminated | November 02, 2011 | 86 | 68530 | Not_Computer
Z-2957-2011 | 59352 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | HOLDEX Single-Use Holder PP closed pack. Greiner Bio-One Austria Gmbh. A sterile single-use disposable holder with an off-center female luer at the top of the holder and epoxy bonded stainless steel needle on the underside of the holder, recommended for use with commercially available butterfly and luer needles in routine venipuncture procedures. | HOLDEX SingleUse Holder PP | 2 | August 08, 2011 | 2011 | Greiner Bio-One North America, Inc. | At removal of the tube from the holder, the needle may dislodge and blood spill can occur. | OTHER/UNDETERMINED: Pending | Greiner Bio-One North America, Inc. sent a "Recall HOLDEX Single Use Holder PP" letter dated June 29, 2011 to all affected customers.The letter instructed customers stop distributing the product immediately. A Product Disposition form was provided to be completed and returned to Greiner Bio-One at which time upon receipt, arrangements for product return to Greiner Bio-One will be made for replacement or credit. They were requested to provide a list of their customers to Greiner Bio-One that received the affected item.For additional information call Greiner Bio-One at (888) 266-3883. | Contact the recalling firm for information | 142400 | 142,400 pieces | Nationwide Distribution including the states of CA, GA, IL, KS, MA, IL, PA and VA. | Terminated | November 03, 2011 | 87 | 142400 | Not_Computer
Z-2958-2011 | 59449 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015. Product Code: 2M8161. Product Usage: Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. Electronic infusion pumps indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration, such as intravenous (IV), intra-arterial (IA), subcutaneous, epidural or irrigation of fluid spaces applications. | Colleague CX Infusion Pump | 2 | August 08, 2011 | 2011 | Baxter Healthcare Corp. | Baxter's Service Center tested and returned to the user the Colleague infusion pump that did not meet PSIG (Pounds-force per Square Inch) specifications | OTHER/UNDETERMINED: Pending | Baxter Healthcare Corp. notified the consignee by telephoned on July 22, 2011, with the call documented per the approved call script. A follow-up "URGENT PRODUCT RECALL" notification was sent dated July 25, 2011, reiterating the telephone communication informing them that the COLLEAGUE pump may not have met certain PSIG (Pounds-force per Square Inch) specifications after testing by the Baxter Global Technical Service Center. Both forms of notification identified the affected product, problem description, hazard involved, and actions needed to be taken. The consignee was instructed to locate the device, remove it from service and return it to Baxter by contacting Baxter Global Technical Services at 1-800-843-7867 option 1, to obtain a prepaid UPS label and shipping box. Upon receipt of the returned unit, Baxter will perform the required testing and address any issues through repair or service. The notication instructs the consignee to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter as per the instructions provided. If the recalled product have been distributed to another facility, or if you are a dealers, wholesalers and distributors, the letter request that this communication be forward to the pump owner as appropriate. For any questions regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central time. |  | 1 | 1 unit | (USA) Nationwide Distribution including the state of Minnesota. | Terminated | February 14, 2012 | 190 | 1 | Not_Computer
Z-2965-2011 | 59330 | excimer laser system | LZS |  |  | Ophthalmic |  | STAR S4 IR Excimer Laser System; Software Version 5.00Mfg by : AMO Manufacturing USA, LLC510 Cottonwood DriveMilpitas, CA 95035Iris registration is designed to compensate for pupil centroid shift and cycloration, both of Which may occur because of patient characteristics and the different environments at the time of measurement and at the time of treatment. | STAR S4 IR Excimer Laser System; Software Version 5.00 | 2 | August 09, 2011 | 2011 | AMO Manufacturing USA, LLC | Reports of lower success rates with Iris Registration (IR) image capture with the STAR S4 IR Excimer Laser System v 5.3 that with previous version. | DESIGN: Software Design | Abbott Medical Optics (AMO) sent an "ADVISORY NOTICE" dated July 18, 2011 to all customers using the AMO WAVESCAN Wavefront System. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to create a separate patient file for each eye on the WAVESCAN System (STAR S4 IR Excimer Laser Systems,software versions 5.00-5.30) according to the directions in attachment A. Additionally, customers are to complete and return an attached form to AMO via fax at the regional fax number.Contact an AMO representative for questions regarding this notice. |  | 673 | 673 total v 5.00 - 5.22 | Worldwide Distribution-USA (nationwide) including Puerto Rico and countries of Africa, Asian Pacific, Canada, Europe, Japan, Latin America, and the Middle East. | Terminated | September 10, 2013 | 763 | 673 | Software |  | Instructions | Safety Notice/Insructions
Z-2964-2011 | 59426 | instrument, measuring, corneal radius | HJB | Corneal radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System.Haag-Streit USA, Inc.A non-invasive, non-contact OLCR (Optical Low Coherence Reflectometer) Biometer used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of IOL (Intraocular Lens) for implantation after removal of the natural crystalline lens following cataract removal. | Lenstar LS 900 EyeSuite Biometry version i2.000, | 2 | August 09, 2011 | 2011 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. |  | Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.For additional information contact the Haag-Streit at (513) 398-3937. |  | 216 | 216 | Nationwide (USA) Distribution. | Terminated | March 08, 2013 | 577 | 216 | Software | Output/Calculation | Software update | Software Update
Z-2963-2011 | 59349 | wire, orthodontic | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Heat Activated Niti Arches, Natural, Upper, .019x 25", Arch Wires.Packaged in 19 packages of 100 and 6 packages of 10.Highland Metals, Inc.Wire used for orthodontic treatment to place in the brackets to shape teeth. | Heat Activated Niti Arches, Natural, Upper, .019x 25", Arch Wires. | 3 | August 09, 2011 | 2011 | Highland Metals | Heat Activated Niti, Natural, Upper, .019 x 25" Arch wires mixed with size 021x0.25" Arch Wires. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Highland Metals, Inc. notified their affected customers by e-mail and followed with an "Important Alert" letter dated July 12, 2011.The letters informed customers of the mix up of wires. Customers were asked to contact Highland Metals at 800-368-6484 to schedule a call tag to return any remaining product. Upon receipt credit will be issued.For additional information contact Highland Metals at (800) 448-7971 or (408) 271-2955. |  | 10 | 10 packages of 100 | Worldwide Distribution; Nationwide (USA) distribution including the states of NV, CT, and FL; and the countries of Australia and Turkey. | Terminated | September 14, 2011 | 36 | 10 | Not_Computer
Z-2981-2011 | 55941 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Aplio XG model SSA-790A Aplio XG (SSA-790A) is a premium-class diagnostic ultrasound system incorporating state-of-the-art technologies. It is provided with the highest level of image quality and operability (reflecting the XG (next Generation) in its name), and can be used with ease for research as well as for routine clinical studies. | Aplio XG model SSA790A | 2 | August 10, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) is initiating a field correction on the the Aplio XG model SSA-790A because it did not meet the requirement of clause 15b, IEC60601-1 which requires that the residual voltage shall not exceed 60V after 1 second from taking off the AC plug. | DESIGN: Device Design | Toshiba America Medical Systems (TAMS) sent an "URGENT:MEDICAL DEVICE CORRECTION' letter dated June 2, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to be careful not to touch the contacts or prongs on the AC plug within two seconds after unplugging the system from a wall socket. This precaution is to be adhered to until a Toshiba Service Representative replaces the power supply. Customers are to complete and return a Customer Reply Form and fax to 877-349-3054.Customers with questions regarding the notification should contact the Director of Regulatory Affairs at (800) 421-1968 or call their local Toshiba Representative at (800) 521-1968. | Contact the recalling firm for information | 78 | 78 | Nationwide Distribution | Terminated | April 20, 2012 | 254 | 78 | Not_Computer
Z-2980-2011 | 58918 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia Radiation Therapy Pallet. Siemens Medical Solutions USA, Inc.An accessory to the Symbia T Series Emission Computed Tomography System/X-Ray Computed Tomography System.This specialty pallet is used for performing SPECT and/or SPECT/CT Scans on oncology patients with the goal of fusing the SPECT and/or SPECT/CT data with CT simulation data for planning radiation therapy. The pallet attaches to the patient beds carbon fiber pallet. | Symbia Radiation Therapy Pallet | 2 | August 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | There a potential patient safety issue when using the Symbia Radiation Therapy Pallet. When a patient is not iso-centered before starting a scan with the Symbia T Series Systems and a non-circular orbit is selected, it is possible that one of the detectors may come into contact with the patient. | DESIGN: Device Design | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated June 06, 2011 to all affected customers.The letter informs customers of the potential safety issue if the patient is not iso-centered before starting a scan and a non-circular orbit is selected and provided recommendations including ensure the patient is iso-centered before starting a scan. If the patient cannot be positioned at the iso-center of the RTP, instructions were provided. The customers are asked to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter.For questions or assistance contact Siemens Medical Solutions USA, Inc. at (800) 888-7436 (USA). |  | 18 | 18 pallets | Worldwide Distribution: Nationwide (USA) including the states of Florida and Texas; and the countries of Canada, China, Italy, Mexico, Poland, Portugal and Russia. | Terminated | November 27, 2012 | 475 | 18 | Not_Computer
Z-2977-2011 | 59107 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems.Manufactured by Siemens AG, Medical Solutions, Kemnath, Germany.Delivery of X-ray photon and electron radiation for the therapeutic treatment of cancer. | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems | 2 | August 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety rick related to transferring reference images for offset calculation between a syngo RT Therapist v 4.2 or a syngo RT Oncologist 4.2 and other versions of the COHERENCE/syn go RT Therapist or Oncologist or vice versa. An incorrect offset may be applied for patient positioning and result in dose being delivered to the wrong location. | DESIGN: Software Design | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION/CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter provides preventive measure instructions for customers to use until an update for the system is available. | Contact the recalling firm for information | 121 | 121 devices | Worldwide Distribution-USA (nationwide) including the states of CA, FL, NC, NY, OH, PA, SD, UT, and WI and countries of Australia, Belgium, Croatia, Czech Republic, Germany, India, Ireland, Japan, Poland, Saudi Arabia, South Africa, Spain, Sweden, and United Kingdom. | Terminated | March 30, 2012 | 233 | 121 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2994-2011 | 58744 | system, photopheresis, extracorporeal | LNR |  |  | Gastroenterology/Urology |  | XT125 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT; XT001 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENTSingle-use, disposable, closed system with sterile fluid pathways. The device collects the buffy coat (leukocyte-enriched blood) from the patient in a discontinuous flow process and intermittenly returns the remaining plasma and erythocytes to the patient. | THERAKOS UVAR XTS Procedural Kits | 2 | August 10, 2011 | 2011 | Therakos, Inc. | Increased complaints were received of tubing leaks at the junction between the tubing and Fluid Logic Module plastic ports of the kits. | DESIGN: Device Design | The firm, Therakos Photopheresis, sent a "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated March 31, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately discontinue using the affected products; deface and discard their remaining inventory of affected procedural kits, following their internal procedures for destroying and discarding of medical waste, and complete and return the attached Confirmation of Receipt Form no later than 08-April-2011 via fax to: Safety, Risk Management and Surveillance at 1-888-557-3759 or 1-585-453-4110. If you have any additional technical questions, please contact Therakos Customer Service Center at 1-800-828-6316. If you have technical questions regarding this communication, please call Therakos Technical Support at 1-877-865-6850. |  | 13239 | 13239 kits | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD,TN, TX, UT, VA, VT, WA, and WI; and countries of: Argentina, Austria, Australia, Belgium, Canada, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Korea, Kingdom of Saudi Arabia, Jordan, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 13239 | Not_Computer
Z-2990-2011 | 59042 | ventilator, non-continuous (respirator) | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP with Heated HumidifierDevice is intended to treat Obstructive Sleep Apnea (OSA) in patients weighing 30 Kg and above. The device creates a positive pressure in the patient's airway to prevent collapse of the airway while the patient is sleeping. The heated humidifier provides additional humidity to the air supplied to the patient. The heated humidifier has a removable water chamber for filling with water and ease of cleaning. | DeVilbiss Model DV5x Series CPAP with Heated Humidifier | 2 | August 10, 2011 | 2011 | Sunrise Medical HHG Inc. dba DeVilbiss Healthcare | Some units were manufactured with a defective humidifier chamber o-ring. Unit could leak, allowing water to escape during use, causing damage to furniture. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, DeVilbiss Healthcare, issued a "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 11, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to verify their inventory of CPAPs with Heated Humidifiers and isolate any units that are included in the list; replace the humidifier chamber o-rings with the new ones provided; discard the original o-ring removed from the humidifier chamber; complete and return the URGENT: MEDICAL DEVICE CORRECTION Confirmation letter via mail to Quality Manager, DeVilbiss Healthcare LLC, 100 DeVilbiss Drive, Somerset, PA, 15501; affirm that all units have been corrected and that all end users have been sent replacement parts, when all of the CPAPs on the serial number list have been corrected, and return a copy of the sign serial number sheet as confirmation.If you have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800-388-1988 from 8:00AM-5:00PM EST. |  | 210 | 210 | Worldwide distribution; USA (nationwide) including states of AL, AR, CA, FL, IL, IN, IA, KS, KY, MD, MA, MI, MS, NE, NJ, NY, OH, PA, WA, WV, and WI; and countries of Europe. | Open |  |  | 210 | Not_Computer
Z-2978-2011 | 59087 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac, Trilogy. Trilogy Tx and Novalis Tx Linear Accelerators.Varian Medical Systems, Inc. Oncology SystemsIntended for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions in the body where radiation treatment is indicated. | Clinac, Trilogy. Trilogy Tx and Novalis Tx Linear Accelerators | 2 | August 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | On Varian Linear Accelerators where the customer may calibrate the collimator angle position readout in reverse, switching the 90 degree and 270 degree positions. | DESIGN: Device Design | Varian Medical Systems sent an "Urgent Medical Device Correction-Urgent Field Safety Notice" letter dated May 26, 2011 to all affected customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters identifies the product, describes the problem and provides recommendations. Varian recommends maintenance only be performed by qualified personnel with appropriate training, always verify collimator calibration and anytime a collimator posthion readout calibration is performed, and always verify that the collimator angle is not calibrated in reverse by performing the appropriate checks. For questions contact Varian at (888) 827-4265. |  | 5685 | 5685 | Worldwide Distribution: Nationwide (USA) Distribution and the countries of; Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Belorussia, Bolivia, Bosnia and Herzegoniva, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Rep., Denmark, D.R. Ecuador, Egypt, Estonia, Finland, France, French Poly., Georgia, Germany, Greece, Guam, Guatemala, H.K., Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Martinique, Mexico, Moldova, Monte Negro, Morocco, Nepal, Netherlands, N.Z., Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, P.R., Qatar, Reunion, Romania, Russia, S.A.,Serbia, Singapore, Slovakia, South Africa, S.K., Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad/Tobago, Tunisia, Turkey, Ukrain, U.A.E., U.K., Uruguay, Venezuela, Vietnam, V.I., and Zimbabwe. | Terminated | July 13, 2012 | 338 | 5685 | Not_Computer
Z-2979-2011 | 59170 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera/Skyra System - Nuclear magnetic resonance imaging | MAGNETOM Aera/Skyra | 2 | August 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary.The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 20, 2011 to all affected customers. The letter describes the product, problem, and actions that will be taken.The letter informs customers that an inspection of the gradient cable connections and an emergency shutdown mechanism will be installed by a field service representative during a visit to all affected locations. |  | 26 | 26 | Nationwide Distribution-USA (nationwide) including the states of AK, FL, GA, IL, MA, MI, MN, MT, NC, NJ, NY, OH, TX, UT, VA, WA, and WV. | Terminated | January 14, 2014 | 888 | 26 | Not_Computer
Z-2992-2011 | 59401 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0mm diameter, length 200 mm.Remote controlled radionuclide applicator system.Varian Medical Systems, Inc.Mfg by Varian Medical Systems Haan GmbH, Haan, GermanyProduct Usage: The Interstitial Plastic Needles with 2 mm diameter are designed for interstitial treatments in areas such as head and neck, gynecology, breast or prostate. In order to avoid bending and to facilitate the introduction of the needles each needle is supplied with a Mandrin | GammaMed Plastic Needle with Mandrin | 2 | August 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Needle tip may separate from the needle shaft such that the tip may detach in the patient upon withdrawal. | DESIGN: Device Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated July 13, 2011 to all possibly affected customers. This notification supplements and modifies the previous "Customer Technical Bulletins" originally sent in September 2008. The initial Customer Technical Bulletins informed customers of the possibility that needle tips cans crack and separate. It also reminded users that the product has a limited life and to check needles prior to use as per instructions. Current correction notice letter instruct customers to stop using any plastic needles in inventory with lot numbers C01 to D01 and those with lot numbers starting with H or I. The letter also request that the recalled products be returned to Varian BrachyTherapy for replacement. The letters include information as to where the lot number is located on the device. All appropriate personnel should be notified of the correction notice. Contact information is provided for all geographical areas. USA contact for questions or concerns is 1-800-360-7909. |  | 601 | 601 - all sizes. | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Austria, Bangladesh, Belgium, Brazil, Canada, Columbia, Cyprus, Czech Republic Denmark, Dominican Republic, Estonia, France, Germany, Greece, Hungary, India, Japan, Kazahastan, Kuwait, Lebanon, Macedonia, Mexico, Moldova, Nepal New Zealand, Norway, Pakistan, Philippines, Poland, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain Sweden, Switzerland, Taiwan, Thailand, Trinidad, Tobago, Turkey, Ukraine, United Arab Emirates Uruguay, United Kingdom, and Venezuela. | Terminated | June 20, 2012 | 315 | 601 | Not_Computer
Z-2989-2011 | 59362 | system, therapeutic, x-ray | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Shielding Material Flat 20 x 20 cm, Sterile/FO0.05 mm Pb Equivalent (50keV)Carl Zeiss Meditec AGIndicated for radiation therapy treatments. | INTRABEAM Shielding Material | 2 | August 10, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec received a report that the INTRABEAM 20 x 20 cm sterile shield's packaging looked wet inside. Investigation found the substance to be Triethylene Glycol, | PRODUCTION CONTROLS: Process Control | Carl Zeiss Meditec, Inc. sent an "Important Field Correction Action" letter dated July 13, 2011 to all affected customers.The letter included description of affected product and problem. Customers are asked to notify the firm if they have experienced any problems with the product and to return any remaining product for replacement upon completion of required testing. Customers are asked to fill out and return an Acknowledgement Form. For additional information contact Carl Zeiss Meditec at (925) 570-4844. |  | 8 | 8 boxes of 10 sterile packages. | Nationwide (USA) Distribution including the states of FL, CA, NY, VA and AL. | Terminated | October 13, 2011 | 64 | 8 | Not_Computer
Z-2983-2011 | 59347 | device, digital image storage, radiological | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Oncology Information System (OIS) Manufactured by IMPAC/Elekta medical as Multi -Access OIS for Siemens.An oncology digital imaging information management system. Lantis OIS in combination with Sequencer and a non-Siemens linac interface.System components include: LANTIS Commander, LANTIS system, LANTIS 6.1 Commander, LANTIS 8.3 Commander.Subcomponents include: Linac Interface, other vendor, LANTIS Varian Clinac, LANTIS GE Saturn, LANTIS Philips SLS,LANTIS Philips 75 ASU, LANTIS Elekta (Philips) and LANTIS Varian EX I/F. | IMPAC / Elekta Medicals Sequencer | 2 | August 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue exists for data transfer between LANTIS OIS system (containing PRIMEVIEW or PRIMEVIEW Data Gateway) and the Sequencer V&R; system via a third party interface is used. | DESIGN: Software Design | Siemens sent an "Update Instructions/Customer Safety Advisory Notice" to all affected users, beginning June 30, 2011.The letter includes description of the product and problem and provides updated instructions to ensure data is stored in the system.If you have questions contact Siemens at (925) 602-8157. |  | 59 | 59 devices | Worldwide Distribution: Nationwide (USA) including the states of MI, DE, WI, FL, CA, TN, NJ, PA, NE, IL, WV , MA and VA; and the countries of Poland, France, Germany, Canada, Netherlands, India, Ireland, Egypt, P.R. china, Australia, Saudi Arabia, Turkey, Republic Korea, Belgium, Hungary, New Zealand, Italy, Malaysia, and Spain. | Terminated | March 09, 2012 | 212 | 59 | Software | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-2986-2011 | 59367 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Direct Drive Grasper, Model numbers C4130, 5MM X 35CM DIRECT DRIVE GRASPER 10/BXIntended Usage: The grasper is indicated for grasping and manipulating tissue during general or laparoscopic surgery. The instrument is not intended for use when minimally invasive techniques are contraindicated. Complications which have been reported secondary to surgical grasping and manipulation include, but are not limited to: local trauma, injury, and tissue necrosis. | Direct Drive Grasper | 2 | August 10, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed specific lot numbers of the DirectDriveÂż Grasper (model numbers C4130 and C4140) due to a potential weakness in the locking trigger, which could result in the trigger fracturing and the jaw locking in the closed position. | DESIGN: Device Design | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 22, to all affected consignees. The letter described the product, problem and actions to be taken. Consignees were instructed to check their inventory for the recalled products and returned immediately as per the instructions. The letter instructs consignees to complete the attached Voluntary Recall Notification Confirmation Form and fax it to Regulatory Affairs at 949-713-8832 or email it to recall60342148@appliedmedical.com. For questions pertaining to product return contact Chris Jacobson, Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. For regulatory questions contact Susan Fehrenbacher at (949) 713-8041 or by email at sfehrenbacker@appliedmedical.com. | Contact the recalling firm for information | 38864 | 38,864 units total (24,868 units in US) | Worldwide Distribution - USA (nationwide) including the state of Puerto Rico and the countries of: Australia, Bahrain, Canada, Chile, Colombia, Ecuador, Europe, India, Israel, Jordon, Kuwait, Lebanon, Libya, Malaysia, Mexico, Peru, Qatar, Singapore, , South Africa, South Korea, Saudi Arabia. Taiwan, Thailand, Turkey, and Venezuela. | Terminated | March 26, 2012 | 229 | 38864 | Not_Computer
Z-2982-2011 | 59386 | [6 empty] | Biopsy Trays (CT, Puncture & Prep), Packaged individually in header bags, 10 per case.Sterile, Disposable Convenience Kits used for diagnostic radiology procedures. A diagnostic radiology kit used in CT and general biopsies | North Coast MediTek Inc Biopsy Trays | 2 | August 10, 2011 | 2011 | North Coast Medi-Tek Inc | During routine bioburden testing, unusually high results were found for a Bioburden Test done on 4/19/11. Bioburden Level reported as | OTHER/UNDETERMINED: Pending | North Coast Medi - Tek Inc. sent a URGENT PRODUCT RECALL letter dated May 13, 2011 to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The customers were instructed to stop shipment of all remaining product and return to recalling firm. They were also requested to contact their customers and notify them of the recall and retieve all remaining product.For further questions please call (440) 974-0750. |  | 2190 | 2,190 trays | Nationwide Distribution CA, CO, FL, GA, ID, IL, MA, NC, NE, NY, OH, OK, PA, TN, NC and TX | Terminated | February 16, 2012 | 190 | 2190 | Not_Computer
Z-2985-2011 | 59221 | [6 empty] | ATLANTIS(R) Anterior Cervical Plate System Screw Caddy; Self Drilling Fixed Screw Caddy (976-516).Medtronic Sofamor Danek USA, Inc.This instrument is a precision device which incorporates a measuring function. | ATLANTIS(R) Anterior Cervical Plate System Screw Caddy; Self Drilling Fixed Screw Caddy | 2 | August 10, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. | PRODUCTION CONTROLS: Process Control | Medtronic Sofamor Danek USA, Inc. sent "Voluntary Recall - Medtronic" letters on July, 2011 to all affected customers.The letter included; description of product and problem, and asked customers to stop using the product. Customers are asked to contact their Sales Representatives for replacement.For question on this recall call Medtronics Global Quality Department at (800) 876-3133. | Contact the recalling firm for information | 191 | 191 units | Worldwide Distribution: Nationwide distribution including the states of MN, OK, IL, AK, CA, PA, UT, LA, SC, NY, MI, GA, WI, KY, AL, MO, AZ, TN, TX, MA, NE, OH, FL, CO, MT, CT, VA, NC, NH, WA, WV, WY, IN, IA, KS, MT, ME, DE, DC, NJ, MD, MS, NV, HI, ID, ND, and AR; and the country of Uruguay. | Terminated | April 26, 2013 | 625 | 1064 | Not_Computer
Z-2976-2011 | 59296 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper Systems Models:Allura Xper FD10 (F/C)-Biplane Fixed Allura Xper FD10/10 (incl. OR table)Allura Xper FD20- High resolution fixed-x rayAllura Xper FD20/10Allura Xper FD20/20Allura CV20The Allura Xper F010, FD10/10, FD20, FD20/10 and FD20/20 Use From Labeling): is intended for use in cardiovascular and vascular X-ray imaging applications. including diagnostic,interventional procedures (such as PTCA, stent placement and atherectomies), pacemaker implantations and ElectroPhysiology (EP). The Allura Xper is intended for: Dedicated vascular and neurovascular imaging applications, including diagnostic and interventional procedures. This includes, e.g. peripheral, cerebral. thoracic and abdominal angiography, as well as PTCAs. stent placements, embolisations and thrombolysis. and Cardiac imaging applications including diagnostics, interventional procedures (such as PTCA, stent placing, atherectomies), pacemaker implantations and Electrophysiology (EP). Non-vascular interventions such as drainages, biopsies and vertebroplasties procedures. Use of the Allura Xper FD Series in the OR environment is restricted; the system should only be used for minimally invasive procedures. (Endovascular) | Philips Medical Allura Xper Systems | 2 | August 10, 2011 | 2011 | Philips Healthcare Inc. | System freezes from 20 seconds up to one minute. No fluoroscopy exposure and no geometry movements during this time period. | DESIGN: Software Design | Philips Health Care North America sent a FIELD SAFETY CORRECTION ACTION letter dated July 4, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.This FCD will be identified as FCD72200200 and will be free of charge. This FCD consists a CD with the correct firmware and a installation procedure to implement the FCD correctly.Please contact your local Philips representative. For further follow up please call (978 ) 687-1501. | Contact the recalling firm for information | 1423 | 1423 units | Nationwide Distribution | Open |  |  | 1423 | Software | Device Operation | Firmware update | Software Update
Z-3001-2011 | 58322 | [6 empty] | Micro-Fuel Cell Oxygen Sensor, Class RI5Sensor Part Number: C41131-R15Teledyne Instruments - Analytical InstrumentsDisposable oxygen sensors with limited useful life. These sensors are accessories to other oxygen diagnostic devices, and are not in themselves suitable for therapeutic use. | MicroFuel Cell Oxygen Sensor, Class RI5 | 2 | August 11, 2011 | 2011 | Teledyne Instruments - Analytical Instruments | The recall was initiated because during a test on a batch of sensors on the company shipping dock, Teledyne Instruments discovered electrolyte loss from the sealing cylinder on a few sensors. This test was done pursuant to a complaint from one customer of label discoloration caused because of the electrolyte loss. The riskof a leak form the oxygen sensor while in use is the disruption of therape | DESIGN: Process Design | Teledyne Instruments sent emails and "Urgent--Device Field Corrections/Removal" letters in December 2010 to all affected customers.Th letter includes description of device and problem. Customers are advised to check them before use for any signs of electrolyte loss and if discoloration is noted to discontinue its use. and provides recommendation. Teledyne-desplaines to generate a product field corrections/removal log that lists names/addresses of the consignees, quantity of sensors shipped, sensor serial numbers and the date of shipment.For additional information or questions contact Teledyne Instruments at (626) 934-1504. |  | 2813 | 2813 | Worldwide Distribution: Nationwide (US) distribution including the states of CA, TN and WI; and the countries of , Canada, Brazil, Chile, Saudi Arabia, and the U.K. | Terminated | August 22, 2011 | 11 | 2813 | Not_Computer
Z-3010-2011 | 59112 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Evacuation Chair Model 6254It Is intended to aid in the transportation of passengers up and/or down a flight of stairs in an evacuation situation. | Stryker Evacuation Chair | 2 | August 11, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Some of the red release bars used during the time frame of 22 December 2010 through 31 January 2011 were reworked. The rework subjected the parts to excessive heat and as a result, the bars can bend during the action of squeezing the release bar against the black cross tube to release the track. The red release bars may bend to the degree that they permanently engage the track locking mechanism | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Stryker Medical, sent an "Urgent-Medical Device Recall" letter dated June 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate the listed product; inspect each unit prior to use; return the enclosed post card to confirm receipt of this notification; if they have loaned or sold any of the chairs listed, forward a copy this letter to the new users and advise Stryker of new users names and locations, and if they have disposed of any of the affected chairs and they are no longer in use, advise Stryker of their obsolescence by providing them with the serial numbers. Stryker stated in the letter that, their goal is the timely upgrade of all affected units within the next 90 days.If you have any questions or concerns, please contact us at 800-STRYKER, option 2. Please reference RA-2011-027. Normal business hours are Monday-Friday 8am-5pm EST. |  | 220 | 220 | Worldwide distribution: USA (nationwide) and countries of: Canada, China, France, Germany, India, Netherlands, and Sweden. | Terminated | August 29, 2013 | 749 | 434 | Not_Computer
Z-3003-2011 | 59122 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE Linear Accelerator with RT Therapist v4.1 system.Siemens Medical Solutions, Kemnath, Germany.Distributed by Siemens Medical Solutions, Concord, CA 94520.Product Usage: Linear Accelerators used to deliver X-ray photon and electron radiation for therapeutic treatment of cancer. | Linear Accelerator | 2 | August 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue was discovered when a pause for an imaging segment is inserted (where an imaging segment does not exist) with the ARTISTE and RT Therapist 4.1 system will continue to the next image. | DESIGN: Software Design | Siemens sent an "Urgent: Medical Device Correction/ Customer Safety Advisory Notice" letter to all affected customers starting June 3, 2011. The letter describes the product, problem, and instructions to prevent mistreatment. The letter states that Siemens is working on an updated version of RTT 4.1 that will fix the issues. Users will be notified as soon as the update for the system is available. The customers were instructed to include Customer Safety Advisory notice in their syngo RT Therapist 4.1 System Owner Manual chapter "Safety Advisory Letters" where it should remain. |  | 223 | 223 active devices | Worldwide Distribution - USA (nationwide) including the states of: DE, FL, KY, LA, MA, MI, MO, NE, NJ, OH, PA, TN, WI, UT, WV and Puerto Rico and the countries of: Australia, Belgium, Brazil, Canada, China, Egypt, France, Germany, Hungary, India, Ireland, Italy, Lebanon, Malaysia, Netherlands, New Zealand, Philippines, Poland, Republic of Korea, Russian Fed, Saudi Arabia, South Africa, Norway, Spain, Turkey and United Kingdom. | Terminated | February 17, 2012 | 190 | 223 | Software | Device Operation | Software update | Software Update
Z-2997-2011 | 59404 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Titanium Fletcher-style Applicator Set - Defined Geometry - CT Compatible with Titanium intrauterine probe with stopper length 80 mm, 30 degree angle.Varian Medical Systems, Palo Alto, CA;Manufactured by Varian Medical Systems Haan GmbH, Haan, Germany.To treat cancer of the uterus, cervix, endometrium and vagina. | Titanium Fletcherstyle Applicator Set Defined Geometry CT Compatible | 2 | August 11, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential of weld failure of the Titanium intrauterine probe at the cervical stopper. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to return all affected product to Varian Brachy Therapy by using the contact information provided in the letter. Replacement product will be provided free of charge to the customers.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 437 | 437 - all sizes | Worldwide Distribution. | Terminated | June 20, 2012 | 314 | 437 | Not_Computer
Z-3008-2011 | 59445 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Deerfield, IL; product code 2M8161 | Colleague CX Infusion Pump | 2 | August 11, 2011 | 2011 | Baxter Healthcare Corp. | A review of Baxter service records indicates that the Baxter Colleague CX Infusion Pump has not met certain voltage specifications. This issue could potentially cause a delay in initiation of therapy. | OTHER/UNDETERMINED: Pending | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump, and promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com.If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. | Contact the recalling firm for information | 1 | 1 pump | Nationwide distribution: Missouri only. | Terminated | October 19, 2011 | 69 | 1 | Hardware | Treatment/Delivery/Therapy | Return | Remove or Replace
Z-3002-2011 | 59378 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System;Varian Medical Systems, Palo Alto, CA Product Usage: Eclipse Treatment Planning System is used to plan radiotherapy treatments for patients with malignant or benign diseases. It is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (Brachytherapy) treatments. | ECLIPSE TREATMENT PLANNING SYSTEM | 2 | August 11, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, the treatment field energy may change from the current selection to a previous selection without use intervention and without invalidating the MU or calculated dose distribution. | DESIGN: Software Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 30, 2011 to all affected consignees. The letter identifies the affected product, problem, and actions to be taken. The letter states that Varian will contact all affected customers to apply a database configuration script to the database and is preparing a correction for Eclipse Version 10 to address the issue. A Varian representative will contact each site to schedule correction when it is available. The letter also provides RECOMMENDED USER ACTIONS to follow. The Letters go on to reiterate caution and warnings from Eclipse instructions for use. All appropriate personnel should be notified of the correction notice. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. For further clarification, contact your local Varian Customer Support District or Regional Manager. |  | 2436 | 2436 | Worldwide Distribution. | Open |  |  | 2436 | Software | Treatment/Delivery/Therapy | Upgrade | Repair
Z-3007-2011 | 59448 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Deerfield, IL. Product Code: 2M8161. Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. Electronic infusion pumps indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration, such as intravenous (IV), intra-arterial (IA), subcutaneous, epidural or irrigation of fluid spaces applications. | Colleague CX Infusion Pump | 2 | August 11, 2011 | 2011 | Baxter Healthcare Corp. | The Colleague infusion pump may not have met certain Air in Line (AIL) specifications after testing by the Baxter Global Technical Service Center. This issue could potentially cause a delay in initiation of therapy. | DESIGN: Process Design | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump; promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com, and forward this communication to the pump owner as appropriate, if they distributed the pump to another facility or if they are dealers, wholesalers, or distributors/resellers.The letter states "Upon receipt of the returned device, Baxter will perform required testing on the affected COLLEAGUE pump for which you have indicated there is continuing medical necessity. Should Baxter find any issues during testing, Baxter will address those issues through repair or service prior to returning the pump to you at no charge".If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. | Contact the recalling firm for information | 1 | 1 Pump | Nationwide distribution: MN only. | Terminated | November 09, 2011 | 90 | 1 | Not_Computer
Z-3005-2011 | 59416 | clamp, vascular | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Surgical Clamp Inserts, Model#G-8650, 86MM Insert Soft/Fibra 15/BxSurgical Clamps and inserts are indicated for surgical clamping during cardiovascular, peripheral vascular, and general surgery. The clamps are available in 33, 45, 61 and 86 mm jaw lengths and the inserts are designed to snap easily into the clamp jaws. | Surgical Clamp Inserts | 2 | August 11, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed that Surgical Clamp Inserts, Model Numbers: G-8600, G-8650, G-8655, has the potential packaging nonconformance which could result in a compromise to the sterile barrier. | PRODUCTION CONTROLS: Packaging Process Control | The firm, Applied Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to its customers. The letter provides customers with an explanation of the problem identified and actions to be taken. The customers were instructed to check their inventory; complete and return the attached Voluntary Recall Notification Confirmation Form via fax to Regulatory Affairs at 949-713-8832, and return the recalled Standard 86mm Inserts immediately to Applied Medical, Attn: RGA# 60308783, 22872 Avenida Empresa, Rancho Santa Margarita, CA 92688.Customers with product return questions, were instructed to contact the Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. Customers with regulatory questions, were instructed to contact the Director, Regulatory Affairs, at (949) 713-8041 or by e-mail at sfehrenbacher@appliedmedical.com. | Contact the recalling firm for information | 20247 | 20,247 units total | Worldwide distribution: USA (nationwide) including states of: Arizona, California, Colorado, Florida, Georgia, Hawaii, Iowa, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, and Michigan; and countries of: Japan, Israel, and Singapore. | Terminated | March 26, 2012 | 228 | 20247 | Not_Computer
Z-3009-2011 | 59275 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis U mobile x-ray systemProduct Usage: Mobile x-ray system | Axiom Artis U | 2 | August 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When the heat unit of the angiography system reaches or exceeds 65%, the cool down time display will appear with an acoustic signal. If users ignore the warning and continue the radiation, the tube may be autoprotected when it reaches its temperature limitation causing the X-Ray to lock until the tube cools down. Following a tube auto-protection, an error message "Door is Open" or "Err 421/240: | MISBRANDING: Labeling False and Misleading | Siemens sent a "Customer Safety Advisory Notice" letters to all affected users dated June 13, 2011. The letter identified the product, the problem, and the actions to be taken. Siemens included in the notice an addendum to the instruction for use to each system and request customers to take special attention. Customers were instructed to place the Safety Notice in the System's instruction for use. All appropriate personnel should be notified of the correction notice. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. |  | 18 | 18 | Nationwide Distribution (USA) including the states of: CA, CT, FL, KS, MN, MO, NJ, NY, NC, PA, and UT. | Terminated | January 24, 2014 | 897 | 18 | Not_Computer
Z-3014-2011 | 59463 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Metha(R) Hip System Cobalt Chrome (CoCr) Metha Modular Neck Adapters used with the Metha Modular Short Hip Stems.Aesculap AG.Intended to replace a hip joint | Metha(R) Hip System Cobalt Chrome (CoCr) Metha Modular Neck Adapters used with the Metha Modular Sho | 2 | August 12, 2011 | 2011 | Aesculap Inc | Aesculap Implant Systems received a notification from the manufacturer, Aesculap, AG of a fractured CoCr neck adapter. |  | AESCULAP Implant Systems sent an "Important Recall Notification" letter dated to all affected customers.The informed customers of a complaint received where fractured CoCR adapter was found and to return any affected product to the firm. The reason for the fracture is currently under investigation. While the firm is trying to determine the root cause of the fracture, they will be offering the Metha non-modular short hip implants to their customers. Customers are asked to return an Inventory Sheet to AESCULAP.For additional information on this recall call (610) 984-9239 or (610) 984-9300. |  | 149 | 149 | Nationwide (USA) Distribution including the states of AL, FL, GA, MD, NM, NY, OK, and TN. | Terminated | March 19, 2013 | 585 | 149 | Not_Computer
Z-3018-2011 | 57583 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-CAT Classic (X-Ray Tube head and Power Supply)X-Ray imaging device that constructs a three dimensional model from images taken during a rotational X-Ray sequence. Intended to be used whenever a dentist, oral surgeon or other physician needs 3D information of high contrast objects. | iCAT Classic (XRay Tube head and Power Supply) | 2 | August 12, 2011 | 2011 | Imaging Sciences International, LLC | Recalling firm received a customer complaint regarding a defective X-ray tube head and power supply | OTHER/UNDETERMINED: Pending | The firm, Imaging Sciences, sent an "Urgent Medical Device Field Correction" letter dated December 15, 2011 via first class USPS to its customer. The letter noted the product and described the problem and action to be taken. The customer was instructed to complete and return the ACKNOWLEDGEMENT FORM via fax to 215-997-5665 or email: darin.oppenheimer@imagingsciences.com. Imaging Sciences stated in the letter that an X-ray tube head and power supply will be provided and installed by one of their technicians. Should you have any questions regarding this information, call 215-622-1545. | Contact the recalling firm for information | 1 | 1 | X-ray tube and power supply were distributed to one consignee: Dr. Alan D'Allessandro, 2500 W. Higgins Road, Suite 665, Hoffman Estates, IL 60169 | Terminated | March 01, 2013 | 567 | 1 | Not_Computer
Z-3012-2011 | 59059 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration, GE Healthcare 3000 North Grandview, Waukesha, WI 53188Seno Advantage is a medical image review station that allows easy selection, processing, filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. When interpreted by a trained physician, mammographic images displayed on the high-resolution monitors may be used as an element for diagnosis. Furthermore filmed images from all modalities may also be used as an element for diagnosis. | GE Healthcare, Seno Advantage 2.1 and 2.2 Workstations | 2 | August 12, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that patient mismatch could occur on some Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration. This issue may impact patient safety.When Seno Advantage workstation is synchronized with RIS using CCOW (an HL7 standard protocol designed to enable disparate applications to synchronize in real-time), the following issues could occur: 1.Two different | DESIGN: Software Design | GE Healthcare sent a URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated May 31, 2011. The letter identified the product , the problem and the action needed to be taken by the customer.GE Healthcare will upgrade all affected workstations at no cost. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. | Contact the recalling firm for information | 391 | 391 | Worldwide Distribution--USA (nationwide) including the states of: AZ, CA, CO, CT, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OK, OR, PA, SC, TN,TX, and WI, and the countries of UNITED KINGDOM, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, PORTUGAL NORWAY, NEW ZEALAND, LUXEMBOURG, LEBANON, KOREA, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, DENMARK, CANADA, BELGIUM AUSTRIA, AUSTRALIA FINLAND. | Terminated | June 29, 2012 | 322 | 391 | Software |  |  | Other
Z-3017-2011 | 59517 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion Pump.Animas Corp.An insulin infusion. | Animas 2020 Insulin Infusion Pump | 2 | August 12, 2011 | 2011 | Animas Corporation | Incorrect component (1nF capacitor) was used to produce 191 printed circuit boards used in six pump lots of which 380 pumps were distributed. This issue may result in the pump displaying a "Replace Battery" alarm not related to the condition of the battery. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Animas Corporation sent an "Urgent: Insulin Pump Recall" letter dated July, 29 20011 to all affected customers.The letter included affected product and problem. They requested anyone who received the pump to send it back to Animas and a replacement pump will be sent at no charge.For additional information contact the Customer Support Department at (877) 937-7867. | Contact the recalling firm for information | 380 | 380 | Worldwide Distribution: Nationwide (USA) distribution including the states of CA, CO, MA, and MD; and the countries of Australia, Finland, France, Israel, Italy, Sweden and the U.K. | Terminated | December 07, 2011 | 117 | 380 | Hardware | Alarm/Message | Replace device | Remove or Replace
Z-3015-2011 | 59398 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm, Models: FSIII, FSIIIC and OfficemateThe Officemate, FSIII and FSIIIc are compact mobile fluoroscopic C-arm systems specifically designed for x-ray imaging of extremities. The systems are used by trained medical personnel, but at times are used for non medical applications. The functionality of the units is identical for the OfficeMate, FSIII and FSIIIc. Also the construction of the different models is identical. That includes the mounting of the flex arm to the cabinet and the flex arm to the C-arm. The difference between the models resides in the monitors: " Officemate has one small monitor " FSIII has two larger monitors " FSIIIc is an international version of the FSIII (power kit) | Fluoroscan Mini Carm | 2 | August 12, 2011 | 2011 | Hologic, Inc. | The Flex arm connector to C-arm can break and the C-arm that is supported by the flex arm may fall down | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Hologic, Inc., sent an "Urgent: Medical Device Recall" letter dated July 25, 2011 via FedEx to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to visually check the the Flex arm/C-arm on their unit for any sign of damage or cracking, if any damage or cracks are observed, remove the unit from service until they receive the retrofit kit; and to complete and return the Confirmation of Notification form via fax to Hologic at 781-280-0662 or mail using the enclosed pre-addressed, pre-paid return envelope. After the customer confirms that they have the unit in their possession, Hologic will expedite the user installable repair kit. The kits will be available starting the first week of August. The kits will also be dispatched via FedEx. Confirmation of installation will be requested for all dispatched kits.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 540 | 540 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Spain, and Viet Nam. | Terminated | June 19, 2012 | 312 | 540 | Not_Computer
Z-3016-2011 | 59376 | nebulizer (direct patient interface) | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLife Nebulizer Heater, Catalog Code 2M8021The AirLife Nebulizer Heater is part of the heated nebulizer system intended to provide warm bland aerosol therapy to respiratory patients. The system is used in conjunction with a standard oxygen flowmeter, which by providing pressurized oxygen powers the nebulizer creating ultra-fine particles in the therapeutic aerosol. The heater warms the solution before it reaches the aerosol jet in the nebulizer and heats the dilution air before it is entrained into the aerosol stream. An aluminum heat exchanger on the nebulizer cap transfers the thermal energy to a nebulizer solution bottle warming the solution prior to aerosolizing. | AirLife Nebulizer Heater | 2 | August 12, 2011 | 2011 | Carefusion 211 Inc | The recall was initiated because CareFusion's Investigation has confirmed the Nebulizer Heater can, over time, develop a failure mode that results in an internal electrical short. The short can result in smoke and sparks being emitted from the device. In the presence of oxygen enriched environments and a fuel source, this failure represents a potentialrisk of flame. CareFusion is voluntarily rec | OTHER/UNDETERMINED: Pending | Carefusion sent a URGENT PRODUCT RECALL letter dated July 15, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to examine their inventory for the affected product as indicated on the sample shown in the enclosed Attachment A-1 and remove any affected heaters present in their facility. Confirm the total number of affected units in their possession and enter the quantity and respective serial number on the enclosed Customer Response Card (Attachment A-2). Customers were instructed to contact CareFusion's Technical Support Department at 800-554-8933, (6:30 am to 5:00 pm PDST), to obtain an RMA (returned materials authorization) number for the return of the affected product. Upon receipt of their response, CareFusion will ship to their customers new replacement Nebulizer Heaters in exchange for the affected units.Customers were instructed to acknowledge receipt of this communication by completing the enclosed Customer Response Card and faxing or sending a copy by return mail to the address indicated on the Customer Response Card, (Attachment A-2). | Contact the recalling firm for information | 22670 | 22,670 units total | Nationwide Distribution | Terminated | April 30, 2012 | 262 | 22670 | Hardware | Physical Safety Hazards | Remove | Remove or Replace
Z-3013-2011 | 59388 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator probe; 3.2mm diameter, braced PVDF.Varian Medical Systems, Palo Alto, CA.Mfg by Varian Medical Systems, Haan GmbH.Flexible Applicator Robes are designed for brachytherapy intracavitary treatments in areas such as vagina, vagina stump or rectum. | GammaMed Flexible Applicator | 2 | August 12, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Certain GammaMed Flexible applicator probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | PRODUCTION CONTROLS: Process Control | Varian Medical Systems, Inc. sent out Customer Technical bulletins in August 2008 and Urgent Medical Device Correction/Urgent Field Safety Notices on July 13, 2011 to all affected customers.The letter included description of device and problem. The current notice supplements and modifies previous instructions and ask that they stop using any flexible applicator probes in inventory. They are to be returned to Varian Brachy Therapy for replacement. Letter go on to list possible scenarios if the length deviation is not detected and advised all personnel working in the radiotherapy department should be notified of the recall. For additional information contact (800) 360-7909. |  | 1343 | 1343 | Worldwide Distribution: Nationwide (USA) and including the countries of France, Germany, India, Trinidad and Tobago, United Kingdom, Canada, Nepal, Ukraine, Czech Republic, Bangladesh, Thailand, Turkey, Brazil, Estonia, Switzerland, Russia, Greece, Philippines, Spain, Belgium, Kuwait, Romania, Morocco, Mexico, Columbia, Austria, Poland, Kazahstan, Vietnam, Cyprus, New Zealand, Dominican Republic, Norway, Uzebestan, Hungary, Lebanon, South Africa, South Korea, Slovenia, Venezuela, Serbia, Tunisia, Pakistan, Macedonia, Ireland, Azerbajian, Ukraine, Morroco, Italy, Singapore, Nambia, Turkey, Sudan, and Taiwan. | Terminated | June 20, 2012 | 313 | 1343 | Not_Computer
Z-3024-2011 | 56839 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical elance Vital Signs Monitor, Model 93300.The Spacelabs elance Vital Signs Monitor and elance Central Station is a prescription device intended for use by healthcare professionals in all areas of a healthcare facility. It is indicated for use in adult and pediatric populations. The device facilitates the monitoring of: ECG with arrhythmia detection; respiration; non-invasive blood pressures; invasive blood pressures; body temperature; functional arterial oxygen saturation; and entidal CO2. | Vital Signs Monitor, Model 93300 | 2 | August 15, 2011 | 2011 | Spacelabs Healthcare, Llc | The monitor has a 3 second delay in the waveforms and respiration alarm; the delay is not disclosed in the IFU. | PRODUCTION CONTROLS: Error in Labeling | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated September 24, 2010 to all affected customers. The letter describes the product, problem, and the action to be taken.The letter includes the revised Operator's Manual as a replacement for customers current manual.For Technical Support call Spacelabs Medical at 1-800-522-7025 and select 2. | Contact the recalling firm for information | 2517 | 2517 | Worldwide Distribution-USA (Puerto Rico) and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KENYA, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MACEDONIA, MALAYSIA, NEPAL, NETHERLANDS, NICARAGUA, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU, PUERTO RICO, PHILIPPINES, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | August 17, 2011 | 2 | 2517 | Not_Computer
Z-3019-2011 | 57543 | table, radiographic, stationary top | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Pro-Lok Platform, REF MTSBRT002. Firm name on label is CIVCO, Orange City, Iowa Product Usage: The recalled product is intended to support and aid in positioning a patient during radiologic and other medical procedures. | CIVCO ProLok Platform, REF MTSBRT002 | 2 | August 15, 2011 | 2011 | Civco Medical Instruments Inc | Systems sold before December 2009 displayed a scale of 0-40cm going from left to right, while systems after December 2009 reversed this and were displayed as 40-0cm from left to right. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CIVCO issued a "FIELD SAFETY NOTIFICATION" letter dated 12/20/2010 beginning 12/30/2010, via FedEx or UPS to distributors and customers. The notification identifies the affected device, problem, and field safety corrective action. The letter states that if customers have more than one system purchased from different time periods that they should examine their systems and assess if there is any mismatch. The recalling firm will rework the system to correct the difference and send it back to the customer. To take advantage of this correction customers are advised to contact the recalling firm within thirty days of receiving the safety notification. The letter includes a response form to be completed and returned as per instructions, to indicate that the notice has been read and understood by all operators of the equipment. The letter to distributors requested that a copy of the notification letter be forward to all customers who have received this component. For assistance contact your CIVCO insides sales representative by calling CIVCO Medical Solutions, Radiation Oncology, at 800-842-8688. |  | 20 | 20 platforms | Worldwide Distribution - USA ( nationwide) including the state of PA. and countries of: Canada, China, England, Israel, Italy, Japan, Korea, The Netherlands, Russia, Spain, Sweden, Switzerland and Turkey. | Terminated | June 28, 2013 | 683 | 134 | Not_Computer
Z-3023-2011 | 59235 | radioimmunoassay, estriol | CGI | Estriol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AccessÂŽ Immunoassay Systems UNCONJUGATED ESTRIOL Part # 33570The Access Unconjugated Estriol assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of unconjugated estriol levels in human serum using the Access Immunoassay Systems. | Access Immunoassay Systems Unconjugated Estriol | 2 | August 15, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that a single lot of Access Unconjugated Estriol reagent (PIN 33570) may produce suppressed patient results and an increased incidence of IND (Indeterminate) flags. These issues have been associated with reagent lot 913348. This reagent met all testing specifications prior to release. | OTHER/UNDETERMINED: Pending | Beckman Coulter sent a PRODUCT CORRECTIVE ACTION (PCA) letter dated December 8, 2009, to all affected customers. The letter was sent via US Postal Service for all US customers with attached PCA response form. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of Access Unconjugated Estriol reagent lot number 913348. " No further orders of Access Unconjugated Estriol reagent lot number 913348 will be shipped." Abnormally low Access Unconjugated Estriol results from samples tested previously with the affected lot should be reviewed in the context of other diagnostic tests." Access Unconjugated Estriol results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Rerun any tests that produce a questionable result.In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at (800) 854-3633 in the US and Canada. Outside of Canada and the US, customers were instructed to contact their local Beckman Coulter Representative. | Contact the recalling firm for information | 2786 | 2786 units total (2776 in the US) | Worldwide Distribution -- USA (nationwide) including the states of CA, CT, MA, NC, NM and TX and the country of Canada. | Terminated | April 16, 2012 | 245 | 2786 | Not_Computer
Z-3022-2011 | 59394 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, 4D Structure Merge Tool Option in Eclipse Release 8.5 -10 (Client Build 8.2.2x through 10.0.34) with 4D Planning; Model Number: H48; Reference/FSCA Identifier: CP-04931; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. | Eclipse Treatment Planning System | 2 | August 15, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified in the 4D Structure Merge option of certain versions of Varian's Eclipse Treatment Planning System. The treatment field energy may change from the current selection to a previous selection without user intervention and without invalidating the MU or calculated dose distribution. | DESIGN: Software Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/FIELD SAFETY NOTICE" letter dated July 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to cease using the 4D Structure Merge tool option in the Eclipse until further notice. Additionally, it is recommended that customers ensure that a qualified physician reviews the accuracy and placement of all patient structures used for plan development and evaluation and to conduct standard quality assurance procedures prior to patient treatment. Customers will be contacted by a Customer Service representative when the technical correction for this issue is available.Contact your local Varian Customer Support District or Regional Manager if you have questions regarding this notice. |  |  |  | Worldwide Distribution | Open |  |  |  | Computer
Z-3025-2011 | 59661 | tenaculum, uterine | HDC | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GYNDisposables Inc, IUD INSERTION KIT 935K, STERILE, SINGLE USE, Tenaculum 356T, Uterine Sound 906S, Scissors, Povidone-Iodine Solution, (3) Swabsticks, Paper Towel, STERILE EO, REF 935K, QTY 1, GYN Disposables,Inc. 2671 Appling Road, Memphis, TN 38133, ph 901.377.2200, fx 901.377.2207Intended for intrauterine device (IUD) insertion. | IUD INSERTION KIT 935K | 2 | August 15, 2011 | 2011 | GYN Disposables Inc | Lack of sterility assurance. | DESIGN: Process Design | GYN Disposables Inc. sent out an "Urgent: Medical Device Recall" letter dated July 29, 2011 to all affected customersThe letter includes: product affected, description of problem and asks customers quarantine all affected product. They can contact GYN Disposables to have the product picked up and to receive credit. A Response Form is also included.For questions on this recall call (888) 372-2203, extension 101. |  | 3987 | 3987 units | Nationwide Distribution. | Terminated | October 29, 2012 | 441 | 3987 | Not_Computer
Z-3028-2011 | 59218 | system, blood collection, vacuum-assisted, manual | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer Push Button Blood Collection Set; REF 367324 Rx only, 23 G x 3/4" x 12 " (0.6 x 19 mm x 305 mm);BD Vacutainer Push Button Blood Collection Set without Multiple Sample Luer Adaptor, Sterile, Rx only, 21 G x 3/4 x 12" 0.8 x 19 mm x 305 mmBlood collection system. | BD Vacutainer Push Button Blood Collection Set | 2 | August 16, 2011 | 2011 | Becton Dickinson & Company | BD Vacutainer Push Button Blood Collection Set can potentially undergo cannula separation from the hub. | DESIGN: Device Design | BD Diagnostics sent a "Lot Specific Product Recall" letter dated April 22, 2011 to all affected customers.The letter asked customers return any affected product and provides instructions for its return. The letter ask customers return the attached Packaging List with the shipment.For question on this recall contact BD Diagnostics at (201) 847-5117. | Contact the recalling firm for information | 1744000 | 1,744,000 | Nationwide (USA) Distribution. | Terminated | June 06, 2012 | 295 | 1778960 | Not_Computer
Z-3026-2011 | 59408 | exerciser, powered | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AlterG Anti-Gravity Treadmill;Alter G, 48349 Milmont Drive, Fremont, Ca 94538Product Usage: Rehabilitation of lower body injury &#38; Neurological conditions. | GTrainer | 2 | August 16, 2011 | 2011 | Alter-G, Incorporated | Potential installation error on AlterG Anti-Gravity Treadmill that could cause greater than expected unweighting. | PRODUCTION CONTROLS: Equipment Maintenance | Alter-G, Inc. notified all affected customers by telephone, e-mail and followed by a "Medical Device Customer Advisory Notification" letter dated July 15, 2011. Customers were provided information on the installation error, the potential harm to patients, and the actions to be taken. The letter instructs customers to discontinue use of the Alter G until the recalling firm can evaluate the machine and correct any installation error. The letter states that the recalling firm is planning to visit affected customer's facility as quickly as possible and will contact customers to schedule this visit. Questions or concerns should be directed to David Tjon at 510-270-5369. | Contact the recalling firm for information | 16 | 16 potentially affected units | Nationwide Distribution (USA ) - including the states of: OK and TX. | Terminated | August 16, 2011 | 0 | 16 | Not_Computer
Z-3027-2011 | 59239 | enzyme immunoassay, non-radiolabeled, total thyroxine | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AccessÂŽ Immunoassay Systems Total T4 Calibrators, Part Number: 33805,Subsequent Product Code: JISThe Access Total T4 Calibrators are intended to calibrate the Access Total T4 assay for the quantitative determination of total thyroxine (T4) levels in human serum and plasma using the Access Immunoassay Systems. | Access Immunoassay Systems Total T4 Calibrators | 2 | August 16, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stability claims for the Access Total T4 calibrators may not be met. This issue is associated with calibrator lots 917074 and 008534. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 13, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected lot numbers and review historical quality control performance. Additionally, a response form was attached to the letter for customers to complete and return. Alternate lots will be provided for current and future replenishment orders or requests.Customers with questions regarding this notification can call Technical Support at (800) 854-3633 or contact their local Beckman Coulter representative. | Contact the recalling firm for information | 2495 | 2495 units total (2492 in US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | July 11, 2012 | 330 | 2495 | Not_Computer
Z-3032-2011 | 59456 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, & 91393 are sold with with Perioperative Mode (option D), which includes Start case/End case functionality.Spacelabs Healthcare patient monitors, functioning as either bedside or central monitors; passively display data generated by Spacelabs Healthcare parameter modules, Flexports interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. Key monitored parameters available on the models 91367, 91369, 91370 and 91387, when employing the Spacelabs Command Module, consist of ECG, respiration, invasive and noninvasive blood pressure, Sp02, temperature and cardiac output. Additional parameters and interfaces to other systems are also available depending on the parameter modules employed.Product Usage: When a clinician enters END CASE mode, there is the expectation the operation is over and the patient is being disconnected from the monitor. Removing the patient cables will trigger multiple alarms so clinicians will enter this END CASE mode to prevent having to listen to false alarms. Spacelabs Healthcare patient monitors are intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other physiologic monitors via Flexport interfaces. These devices determine: a) when an alarm condition is violated; b) the alarm priority (ie. high, medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Healthcare patient monitors may also function as a generic display or computer terminal. As a generic display or terminal, the patient monitors allow network-based applications to open Windows and display information on other net worked monitors. Spacelabs Healthcare patient monitors are also designed to communicate with a variety of external devices such as displays, network devices, serial devices, user input devices, audio systems, and local/remote recorders. Spacelabs Healthcare patient monitors are intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. | Spacelabs Medical Patient Monitor | 2 | August 17, 2011 | 2011 | Spacelabs Healthcare, Llc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally. | DESIGN: Device Design | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated July 25, 2011 along with the Perioperative section of the Operations Manual to all consignees in the US. This letter includes recalled model 91369, 91370, 91387, & 91393 (option D). On July 27, 2011, Spacelabs sent out an updated recall notification letter to include one additional model 91367 to all consignees in the US. The letter identifies the products, problem, and actions to be taken. On August 2, 2011, a recall notification letter (translated as necessary) and the Perioperative section of the Operations Manual was emailed to all international subsidiaries and distributors of record. The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular, they should know how to confirm the state of the alarm tones prior to starting a case. The letter states that the firm is in the process of upgrading the software and will be contacting customers to update monitors at no cost. If customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. | Contact the recalling firm for information | 2812 | total 2812 units distributed worldwide | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, Bahrain, Bolivia, Brazil, Canada, Caymen Island, Chile, China, Colombia, Cyprus, Ecuador, Finland, France, Germany, Guam, India, Indonesia, Israel, Italy, Jordan, Mexico, Morocco, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, Uruguay, and Venezuela. | Terminated | May 23, 2012 | 280 | 2812 | Software | Alarm/Message | Software update | Software Update
Z-3030-2011 | 59488 | hemoglobin a2 quantitation | JPD | Hemoglobin A2assay. | Hematology | Hematology | 510(k) | Hemoglobin F & A2 Calibrator, 0.2 mL, 4 brown glass vials/box. Mfr. By: Tosoh Bioscience, Minato-ku, Tokyo, Japan.Hemoglobin F & A2 Calibrator is used as in-vitro calibration material for Beta-thalassemia testing on the Tosoh G7. | Hemoglobin F & A2 Calibrator. | 3 | August 17, 2011 | 2011 | Tosoh Bioscience Inc | The value after reconstitution appeared to be lower than the assigned value on the label. The measured value for Lot BCAL05R was 5.3%, whereas the value on the label is 5.7%. The criteria of the value is +/- 0.3%. | PRODUCTION CONTROLS: Process Control | TOSOH Bioscience sent an "Urgent Recall Notification" letter dated July, 25, 2011 to all affected customers.The letter described the problem and asks customers to immediately discontinue use of the product and return all remaining product back to the firm. Customers can contact TOSOH for replacement.For additional information or question on this recall please contact TOSOH Technical Support at (800) 248-6769. | Contact the recalling firm for information | 6 | 6 boxes/4/2 mL btls | Worldwide Distribution: Nationwide (USA) distribution including the states of FL and IL; and the country of Canada. | Terminated | February 24, 2012 | 191 | 6 | Not_Computer
Z-3036-2011 | 56806 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Ultraview Spacelabs Compact monitor, models 91367, 91369, and 91370, with touch screen display and wireless option.Product Usage: Spacelabs Medical PCIS patient monitors passively display data generated from Spacelabs Medical patient parameter modules, Flexport interfaces, gas analyzers, and other SDLC based products. Spacelabs Medical PCIS patient monitors may also function as a generic display or computer terminal. As a generic display or terminal, the PCIS patient monitors allow network based applications to open windows and display information on other networked monitors. PCIS patient monitors is intended to sound various alarm tones that are generated by the specific PCIS patient parameter modules. Each patient parameter module determines (a) when an alarm condition is violated; (b) the alarm priority (i.e. high, medium or low); (c) alarm limits; and (d) when to generate and terminate alarms. PCIS patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Medical PCIS patient monitors are designed to communicate with a variety of external devices such as display devices, network devices, serial devices, user input devices, audio system, local/remote recorders and Spacelabs Medical patient parameter modules. In all applications, Spacelabs Medical PCIS patient monitors provide access to that data on the Spacelabs Medical network either via hardwired or wireless transmission.. Through this network, a patients data may be displayed at the bedside monitor, central station or a clinical information system. | Portable patient monitor with arrhythmia detection or alarms. | 2 | August 18, 2011 | 2011 | Spacelabs Healthcare, Llc | Monitor may reset repeatedly which can result in a loss of patient monitoring for several minutes. Cause is a software problem. | DESIGN: Software Design | Spacelabs Medical sent an "URGENT - MEDICAL DEVICE CORRECTION" notification letter to U.S. consignees by mail on September 17, 2010 and to international consignees by e-mail on September 24, 2010. The letter identified the recalled product, warns customers of the potential hazard and identifies some ways to mitigate the risk if they decide to continue use. The letter states to correct this situation, the recalling firm Field Service Engineers will be contact customers to schedule a time to install a free software update. If customers have any questions about this corrective action program, please contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 579 | 579 total (517 to US acc'ts; 62 to international acct's. | Worldwide Distribution - USA (nationwide) and the countries of: International distribution to: Canada, China, France, Germany, Italy, and the United Kingdom. | Terminated | August 22, 2011 | 4 | 579 | Software | Device Operation | Software update | Software Update
Z-3033-2011 | 58381 | system, facet screw spinal device | MRW |  |  | Orthopedic | 510(k) | PERPOS PLS, CATALOG #'s 9045-01 & 9045-02A sterile, medical device kit consisting of the components listed in Table 1-1 below. The kit is used by medical professionals to stabilize the spine as an aid to fusion through immobilization of the facet joints. | PERPOS PLS, CATALOG 's 904501 & 904502 | 2 | August 18, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to Pull Pins that may disengage from the Facet Screw during the compression step, requiring that compression be completed with a device driver rather than the compression tool.Use of this product may result in less than optimal compression leading to pseudoarthrosis, non-union or facet pain. | OTHER/UNDETERMINED: Pending | Interventional Spine, Inc. sent a recall letter dated March 15, 2011, informing consignees of the product, the problem, and the action to be taken. Consignees were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if they distributed the product further they were asked to notify Jane Metcalf, the Vice President of Quality Assurance at once.Interventional Spine further informed consignees that the recall was being carried out to the individual consignee level. And their assistance was appreciated and necessary to prevent use of this product.They were then asked to complete and return the enclosed response form as soon as possible. Questions were directed to call their regional directors.For questions regarding this recall call 949-472-0006. |  | 553 | 553 units | Nationwide Distribution including CO, CT, FL, KY, LA, MA, MO, NC, NY, OK, OR, TN, TX, VA & WI | Terminated | February 11, 2013 | 543 | 960 | Not_Computer
Z-3035-2011 | 59365 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging versions 11.5, 11.5.1, 11.5.2, 11.6, 11.7 and 11.8.Horizon Medical Imaging is a medical image and information management application that is intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from various medical imaging systems. | Horizon Medical | 2 | August 18, 2011 | 2011 | Mckesson Information Solutions LLC | Images for one patient may appear in the study of another patient. | OTHER/UNDETERMINED: Pending | Mckesson Medical Imaging Group sent a URGENT FIELD SAFETY NOTICE letter dated June 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On February 2, 2011, the McKesson Medical Imaging company initiated a process for applying an image display software update to the HMI systems at the hospital site that initiated the complaint. They were advised to contact McKesson Customer Support at 1-800-663-2533, to indicate whether they would like to receive the software update. They were advised to distribute the notification to all those within their organization who need to be aware and to alert other parties affiliated with their organization that are affected by this advisory notice. |  | 438 | 438 devices | Worldwide Distribution--USA (nationwide) Australia, Canada, Guam and the United Kingdom. | Terminated | October 04, 2012 | 413 | 438 | Software | Output/Calculation | Software update | Software Update
Z-3037-2011 | 59465 | catheter, electrode recording, or probe, electrode recording | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | ISMUS Catheter Mapping Catheter with Auto ID Technology, Manufacturing Part Number D-1171-34-S, Catalog Number D7R20P14CTThe ISMUS Catheter with Auto ID Technology has been designed for electrophysiological mapping of cardiac structures. | ISMUS Catheter Mapping Catheter with Auto ID Technology | 3 | August 18, 2011 | 2011 | Biosense Webster, Inc. | The recall was expanded because between February and April 2011, Biosense Webster, Inc. notified customers that a voluntary Customer Notification had been Initiated due to the inaccessibility of the electronic version of the Instructions for Use (e-IFU) for the "10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue", Catalog Number D128709,through the J&J; Gateway (i.e., e-IFU websi | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Biosense Webster sent an "URGENT FIELD SAFETY NOTICE" letter dated April 29, 2011 to all affected customers. The letter describes the product, problem, and the action taken by the firm.The letter notifies and advises customers that the updated IFUs for ISMUS Catheter with Auto ID Technology and the Duo-Decapolar catheters have been released and posted on the e-IFU website as of April 22, 2011 and April 21, 2011. Customers were instructed to complete the Acknowledgement Form and return the form to Biosense. Customers with questions or require additionial information were instructed to contact their Biosense Webster representative. |  | 325 | 325 units | Nationwide Distribution | Terminated | October 12, 2012 | 421 | 1855 | Not_Computer
Z-3044-2011 | 59375 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TSRH(R) Spinal System SCREW, REF 837-645, Size: 6.5mm x 45mm, Mat'l: TITANIUM ALLOY, Rx only, QTY: 1 EA, Medtronic Sofamor Danek USA. Product Usage: TSRH(R) Spinal System is intended to help provide immobilization of spinal segments as an adjunct to fusion of the thoracic, lumber, and/or sacral spine. | TSRH(R) Spinal System SCREW | 2 | August 19, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are 5 mm longer than the size etched on the screw. | PRODUCTION CONTROLS: Error in Labeling | On 07/12/2011, the recalling firm, Medtronic notified the Sales Managers and Sales Force by conference call and email regarding the imminent notification and follow-up activities. A "VOLUNTARY RECALL" notification letter was sent to Risk Managers and Sales Representatives on 07/13/2011 with a delivery date of 07/14/2011. The letter states to immediately remove the recalled product from use and return to Medtronic. The letter includes a questionnaires form to be completed and returned as per the instructions. If appropriate, personal follow up with Sales Representatives will be made to assure that Risk Managers understand the recall communication and to facilitate return of product and the associated questionnaire documentation. For questions, contact your Medtronic Sales Representative or the Global Quality Department at 800-876-3133, extension 3173. |  | 17 | 17 units | Nationwide Distribution (USA) including the states of: OH, NY, and WI | Terminated | July 20, 2012 | 336 | 17 | Not_Computer
Z-3039-2011 | 59391 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Calcium Flex Reagent is used on the Dimension clinical chemistry system as an in vitro diagnostic test intended for the quantitive determination of calcium activity in human serum, plasma and urine.Dimension(R) CA Flex(R) reagent is used on the Dimension(R) clinical chemistry system as an in vitro diagnostic test intended for the quantitative determination of calcium activity in human serum, plasma and urine. | Dimension Calcium Flex Reagent | 2 | August 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostic has confirmed a high frequency of "Abnormal assay" flags and imprecision. Test results across the assay range may be falsely depressed or elevated, as much as, 1.93 mg/dL at a concentration of 9.26 mg/dL. | OTHER/UNDETERMINED: Pending | An Urgent Field Safety Notice datey July 2011 was issued to all affected customers. Customers are being advised to immediately discontinue use and discard any remaining inventory of Dimension(R) CA Flex(R0 reagent cartridge - lot EA2153. Replacement order for the affected lot is in process, and will be sent to impacted customers free of charge.Customers were instructed to forward the Field Safety Notice to anyone whom they may have distributed the affected product.For questions regarding this recall call 302-631-6299. |  | 2500 | 2500 cartons | Worldwide Distribution to Guam, Puerto Rico, Chile, Costa Rica, Venezuela, South Korea, Vietnam and domestically to AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MN, MI, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WV, WI, and WY. | Terminated | December 27, 2013 | 861 | 2500 | Not_Computer
Z-3041-2011 | 59520 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | AmscoÂŽ V-PRO 1 and AmscoÂŽ V-PRO 1 Plus Low Temperature Sterilization Systems Model #MLR, Serial # 033250701 - 031581115. The STERIS AmscoÂŽ V-PRO and V-PRO 1 Plus Low Temperature Sterilizer are self-contained stand alone devices, using vaporized hydrogen peroxide as the sterilant. The sterilizer is intended for use in the terminal sterilization of cleaned, rinsed, and dried, reusable metal and nonmetal medical devices used in healthcare facilities. The sterilization cycle operates at low pressure and temperatures, and is therefore suitable for processing medical devices sensitive to heat and moisture. | Amsco VPRO 1 and Amsco VPRO 1 Plus Low Temperature Sterilization Systems | 2 | August 19, 2011 | 2011 | Steris Corporation | STERIS learned that AmscoÂż V-PRO 1 and V-PRO 1 Plus Low Temperature Sterilization Systems users may experience premature wear of the vacuum pump and filtering assembly components. Users may also experience incomplete aeration of the device's chamber following a failed leak test. | OTHER/UNDETERMINED: Pending | Steris Corporation sent an URGENT VOLUNTARY FIELD CORRECTION NOTICE dated August 5, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.A STERIS Service Representative will contact your facility to arrange for STERIS to upgrade your V-PRO System(s). You can continue to use your system prior to the installation of the system upgrades. In the interim period should you experience any alarms or aborted cycles, the load mustbe handled in accordance with the AmscoÂŽ V-PROTM 1 and V-PROTM 1Plus Low Temperature Sterilization Systems Operators Manual and your internal hospital procedures.For further information or if you have questions regarding this matter, please contact Field Service Dispatch at 1-800-333-8828. |  | 663 | 663 units | Worldwide Distribution -- USA (nationwide) including the states of: AK, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, WI, WV & WY, and countries of: Australia, Bangladesh, Belgium, Belarus, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, India, Iraq, Ireland, Italy, Japan, Kuwait, Lebanon, Mexico, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, South Korea, Spain, Switzerland, Turkey, Vietnam, United Arab Emirates, and United Kingdom. | Terminated | March 27, 2013 | 586 | 663 | Not_Computer
Z-3042-2011 | 59525 | accessories, solution, ultrasonic cleaners for lenses | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Muti-Purpose Solution, For Soft Contact Lenses, Sterile, 12 Fl oz. (355 mL), Distributed by Wal-Mart Stores, Inc., Bentonville, AR 72716Daily cleaning, removing protein deposits, rinsing, chemical (not heat) disinfection, and storage of soft (hydrophilic) contact lenses. | Equate MutiPurpose Solution, For Soft Contact Lenses, Sterile | 3 | August 19, 2011 | 2011 | Bausch & Lomb Inc | Muti-purpose solution bottles, lot GC1107 were packaged in a carton that was marked with an incorrect lot code and expiration date. | PRODUCTION CONTROLS: Packaging | Bausch & Lomb Inc. notified the affected customer via teleconference and e-mail on June 30, 2011. The e-mail describes the product, problem, and actions to be taken by the customer.The customer was instructed to hold and stop shipments of any inventory of the affected product. An official recall was initiated with the customer on July 8, 2011 via e-mail and teleconference.Customer was asked to review their inventory and store returns, and send any of the affected product to Bausch & Lomb Returns Dept., 130 Commerce Drive, Greenville, SC 29615. | Contact the recalling firm for information | 1200 | 1200 bottles | Distribution-in the state of AK only | Terminated | November 16, 2011 | 89 | 1200 | Not_Computer
Z-3040-2011 | 59688 | warmer, thermal, infusion fluid | LGZ |  |  | General Hospital | 510(k) | enFlow IV Fluid Warmer Part number 980105VS;(510K name: enFlow IV Fluid Warmer Model 100, EFlow Disposable Cartridge, Model 200 and EFlow AC Power Supply Model 120)Intended for warming blood, blood products, and intravenous solutions, prior to administration. Intended to be used by healthcare professionals in hospital, clinical and field environments to help prevent hypothermia. | enFlow IV Fluid Warmer | 2 | August 19, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | enFlow IV Fluid Warmer may have an issue which results in excessive external surface temperature on the underside of the warmer. | DESIGN: Device Design | GE Healthcare sent UMDC Letters/UMDC Confirmation Letters to customers on June 30, 2011. GE Healthcare sent UMDC Letters/UMDC Confirmation Letters on October 24, 2011 reflecting a new recall strategy. | Contact the recalling firm for information | 1307 | 1307 units | Worldwide Distribution - USA (nationwide) and the country of Vietnam | Terminated | June 14, 2013 | 665 | 1307 | Not_Computer
Z-3051-2011 | 56251 | antinuclear antibody, indirect immunofluorescent, antigen, control | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | "BIO-RAD Kallestad Mouse Stomach/Kidney 8 Well Slides, Bio-Rad Laboratories, STORE FROZEN"The test kit contains slides and substrate.For use as an indirect fluorescent antibody procedure for the detection and semi-quantitation of human autoantibodies. | Kallestad Mouse Stomach/Kidney (MSK) Test Kit | 3 | August 22, 2011 | 2011 | Bio-Rad Laboratories | KALLESTAD Mouse Stomach/Kidney (MSK) test kits contained Crithidia slides instead of the MSK slides. | TRAINING: Employee Error | Bio-Rad Laboratories notified affected customers by phone on 6/30/2010 and followed up with a 2nd call on 7/13/2010.The call informed customers of the problem with the kit and requested they destroy the Crithidia slide packs. Bio-Rad send out replacement slide packs. For information or questions on this recall call Bio-Rad at (425) 498-1709. |  | 224 | 224 kits were distributed; 35 of those kits were determined to contain the wrong slides. | Worldwide Distribution: (USA) Nationwide distribution including the states of CT, OK; and the country of Puerto Rico. | Terminated | August 22, 2011 | 0 | 224 | Not_Computer
Z-3050-2011 | 59329 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith & Nephew, Tendon Stripper, Slotted, 5.0 mm I.D. x12 working length. Non-SterileCatalog Number: 013554Orthopedic Manual Surgical Instrument | Smith & Nephew | 2 | August 22, 2011 | 2011 | Smith & Nephew, Inc. Endoscopy Division | Shaft tip (casting) does not meet the SNE 416specification, which could result in rusting of the Tendon Strippers | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smith & Nephew sent an Urgent - Product Recall Notification letter dated July 7, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the form included in the letter and return to Hoangthi.le@smith-nephew.com or fax to 508-261-3620. Each affiliate must then notify each facility affected by the recall. To return product and obtain credit or order replacement product customers should contact the Smith & Nephew Endoscopy Returns at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to, endo.andreturns@smith-nephew.com. For questions, contact Hoangthi by phone at 508-337-3731. | Contact the recalling firm for information | 570 | 570 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Chile, China, Columbia, Denmark, Dubai, Finland, France, Germany, Greece, Hong Kong, India, Italy, Israel, Japan, Korea, Malaysia, Mexico, Netherland, Norway, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Kingdom, and Venezuela | Open |  |  | 570 | Not_Computer
Z-3045-2011 | 59713 | indicator, physical/chemical sterilization process | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VerifyÂŽ SixCess Indicator Strips, Model #: PCC005, PCC006, PCC007, PCC008, PCC031, PCC033, PCC034, PCC037 & PCC038, Mfr. By: Albert Browne, LTD, Chancery House 190 Waterside Rd. Hamilton Industrial Park, Leicester, Leicester, LE5QZ, United Kingdom.VerifyÂŽ SixCess Indicator Strips are used to confirm that the parameters of sterilization have been reached within packs, pouches and wrapped items processed through a steam sterilization process. Each indicator strip is specific to a sterilization temperature and type (i.e. 250ÂŽF 30 minute Gravity Cycle). | Verify SixCess Indicator Strips | 2 | August 22, 2011 | 2011 | Steris Corporation | During an inspection the firm was notified that language contained in labeling and promotional literature directs Customers that the products can be used to replace biological indicators and/or release loads containing implants. | MISBRANDING: Labeling False and Misleading | STERIS sent a Customer Notification letter dated March 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review and follow the Challenge Packs' revised instructions for use enclosed as Attachment A (270-4) and Attachment B (275F 3).For any questions regarding this recall customers were instructed to contact their STERIS Account Manager or call 440-392-7421. |  | 2500 | 2,500 (one per customer) | Worldwide Distribution - USA (nationwide) and the countries of Lebanon and Qatar. | Terminated | March 19, 2013 | 575 | 2500 | Not_Computer
Z-3049-2011 | 59466 | kit, i.v. start | LRS |  |  | General & Plastic Surgery | Enforcement Discretion | Churchill Medical Latex-Free IV Start KitProduct Code: AMS-623ATLFThe kit includes a tourniquet, alcohol prep pads, a PVP ampule swab, 1 3M" Transpore tape roll, 2 gauze sponges, a 3M" Tegaderm" dressing, and a dressing change labelAMS-623ATLF is a convenient kit that aids in preparing the insertion site for the start of intravenous therapy | Churchill Medical | 2 | August 22, 2011 | 2011 | Churchill Medical Systems, Inc. | Outer Kit labeled expiration date longer than the expiration of the shortest expired product | TRAINING: Employee Error | Churchhill Medical Systems, Inc A Vygon Company issued an URGENT MEDICAL DEVICE RECALL letter was via overnight courier mail on July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The product was requested to be returned. Distributors were provided with instructions to communicate the recall to customers who received the affected product by providing a copy of the notification and a copy of the Recall Acknowledgement and Inventory Return Form. The form should be completed in its entirety, signed and returned to the Distributor. Contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email to customerservice @vygonus.com. |  | 130750 | 130,750 units | Nationwide Distribution | Terminated | March 27, 2013 | 583 | 130750 | Not_Computer
Z-3057-2011 | 59576 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20:HL 20 4-pumps console base: P/N 917741HL 20 5-pumps console base: P/N 917751The HL 20 is indicated for use as an extracorporeal circulation device for perfusion lasting not more then six hours. The HL console is equipped with a rack for accommodating the system's electronic modules. | Maquet HL Console for Heart Lung Machine HL20 | 2 | August 22, 2011 | 2011 | Maquet Inc. | HL 20 Consoles for 4-pump and 5-pump configuration can malfunction if an excessive amount of liquid is used causing circuit board malfunction. | DESIGN: Component Design/Selection | Maquet sent an URGENT PRODUCT RECALL NOTIFICATION letter dated June 17, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge receipt of the notification by completing and sending back the attached response form whether or not they have any affected product, fax to 973-807-9210 no later than June 30, 2011. For any questions customers were instructed to contact thier local MAQUET representative. |  | 229 | 229 units | Nationwide Distribution | Terminated | October 16, 2012 | 421 | 229 | Not_Computer
Z-3048-2011 | 59568 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Gamma X XL Portable Patient Monitor. Draeger Medical Systems, Inc.Portable monitor used to monitor vital signs of neonatal, pediatric, and adult patients for bedside and transport applications. | Gamma X XL Portable Patient Monitor | 2 | August 22, 2011 | 2011 | Draeger Medical Systems, Inc. | Monitor keys may spontaneously become inoperative or active. This may cause a membrane switch panel malfunction, which could result in the monitor discharging a patient automatically. | DESIGN: Device Design | Draeger Medical sent an "Urgent Medical Device Recall Notification" letters dated July 2011 to all affected customers.The letter included product and problem affected. The letter informed the customer a Draeger representative will contact them to upgrade the system. Until the solution is implemented, Draeger recommends the Delta and/or Gamma XXL may be used. If the affected monitors become inoperative or activate without user interaction, Draeger advises removing the monitor from service and contacting a Draeger representative.For questions or for information on this recall please call Draeger Service at (800) 543-5047 (press 1 at the prompt, then 2, then 32349). | No consumer action necessary | 29 | 29 | Worldwide Distribution - (USA) Nationwide distribution including the states of AK, FL, LA, MN, MS, MO, NJ, NC, OH, OK, PA, TN, TX, NC, NH, NJ, NY, WV, and WI; and the countries of Australia, Austria, Argentina, Bahrain, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, France, Germany, Georgia, Greece, Hungary, Italy, Ireland, India, Kazakhstan, Latvia, Macedonia, Malaysia, Mexico, Netherlands, Norway, Russian Fed., Romania, Saudi Arabia, Slovakia, Spain, South Africa, Switzerland, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Terminated | August 16, 2012 | 360 | 61 | Not_Computer
Z-3055-2011 | 59402 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Flexible Probe with Blocking Washer (part of Segmented Cylinder Set)The Segmented Cylinder Applicator Set was developed to treat cancer of the vagina, vaginal stump and rectum. | Flexible Probe with Blocking Washer | 2 | August 22, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Blocking washers on the Flexible Probe with Blocking washers may shift and cause potential for treatment length error. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate old design applicator probes and guiding tubes and to cease use of the affected products. Information in the letter included instructions on how to return the products to Varian BrachyTherapy.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 1019 | 1019 total | Worldwide Distribution | Terminated | June 20, 2012 | 303 | 1019 | Not_Computer
Z-3056-2011 | 59476 | ligator, hemorrhoidal | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The Resolution II Clip, 155 cm size and 235 cm size, Boxes of 1, 10, and 20 clips consists of a pre-loaded, radiopaque single use clip on a flexible delivery system.Endoscopic Clipping Device, Hemorrhoidal Ligator. The device is used for clip placement within the GI tract for hemostatis ( mucosal defects, bleeding ulcer, arteries, polyps); endoscopic marking; and anchoring to affix jejunal feeding tubes to the wall of the small bowel, as a supplementary method for closure of GI tract luminal perforation | Resolution II Clip | 2 | August 22, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific has received reports where, following clip deployment, the alignment of the clip and delivery system does not allow for the withdrawal of the delivery system away from the clip, resulting in tissue trauma, increased bleeding and the clip being pulled off tissue as the delivery system is withdrawn. |  | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL letter dated August 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Medical Device Recall to their customers. Customers were instructed to:Immediately discontinue use of any affected product.Complete and return the Reply Verification Tracking Form by fax to 508-683-5578.Package and return any affected product to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171For questions regarding this recall call 508-683-5578. |  | 93 | 14, 093 boxes of both 155 cm and 235 cm size clips | Worldwide Distribution - including USA (nationwide) | Terminated | June 22, 2012 | 305 | 14093 | Not_Computer
Z-3052-2011 | 59226 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 Multileaf CollimatorHD 120 MLC is an accessory x-ray collimator designed to be mounted on Varian Trilogy Tx and Trilogy linear accelerators and is intended to shape s-ray field perimeter. Field shape can be either static (fixed) or dynamic. The intended use is to assist clinician in the delivery of radiation to defined target volumes while sparing surrounding normal tissue and critical organs from excess radiation. | HighDefinition 120 Multileaf Collimator (HD120 MLC) | 2 | August 22, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | When using the HD120 Multi Leaf Collimator for beam shaping, and under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceeds the expected transmission dose. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes various jaw position recommendations to be followed when planning and treating dynamic fields on the HD120 MLC. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice. |  | 106 | 173***12-20-12 AMENDED TO: 106 units*** | Worldwide Distribution. | Terminated | December 23, 2011 | 123 | 106 | Not_Computer
Z-3059-2011 | 59524 | cpb check valve, retrograde flow, in-line | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Adjustable Vent Valve, TruChoice Alliant Healthcare Products REFERENCE#213-030NS QUANTITY: 100 DESCRIPTION: ADJUSTABLE VENT VALVE, NON-STERILE; and 3T Medical Systems Catalog# 030NS Adjustable One Way Vent Valve, Quantity: 100, NON-STERILEThe TruChoice and 3-T Vacuum Control One-Way Vent Valve is indicated for use during cardiopulmonary bypass surgery to help prevent the buildup of excess vacuum pressure when suctioning either from the heart or the surgical field, and to help prevent inadvertent retrograde flow and air to the heart. | TruChoice and 3T Medical Systems Adjustable Vent Valve | 2 | August 23, 2011 | 2011 | Alliant Enterprises LLC | The firm is initiating this recall due to the receipt of 2 complaints of malfunction of the joint holding the 2 halves of the vent valve together. |  | Alliant Healthcare sent an URGENT: MEDICAL DEVICE RECALL letter on 7/21/2011 and 7/25/2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review the recall notification and quarantine any affected product. If product was further distributed, customers were to forward the letter with the Recall Return Response Form immediately to all affected locations and indicate which locations on the Recall Return Response Form. Customers were also instructed to discontinue use immediately, dispose of the affected product per their normal disposal protocol, and fax the Recall Return Response Form to Alliant Healthcare Products at 269-629-5770.Upon receipt of the Recall Return Response Form, a substitute vent valve (non-adjustable only) will be sent to their location for each affected product for which disposal has been properly documented in writing. For questions call 269-629-0300, ext 226. | Contact the recalling firm for information | 149 | 149 Boxes of 100 | Worldwide Distribution - USA including Michigan, Tennessee, New York, and California and the countries of Chile and Korea | Terminated | April 23, 2012 | 244 | 678 | Not_Computer
Z-3053-2011 | 59693 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Katzen" Infusion Wire, 0.035 in x 180 cm, UPN Product No. M001461930, RED Catalog No. 46-193, Made in USA 8600. Boston Scientific. NW 41 Street, Miami, FL 33166.Intended for the infusion of therapeutic agents (i.e., heparin, saline, thrombolytic agents, etc) in the peripheral vasculature. The wire allows for the delivery of agents in either a "pulse -spray" or " slow weep" technique. | Katzen" Infusion Wire | 2 | August 23, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling one lot/batch of Katzen" Infusion Wires because we have identified that the label on the carton and the inner pouch indicated an incorrect device length. The labels indicate a device length of 180cm while the actual device length is 145cm. | PRODUCTION CONTROLS: Labeling Mix-Ups | Boston Scientific sent an "Urgent Medical Device Recall Removal" letter dated August 1, 2011 to all affected customers.The letter described the problem and product being recalled and advised customers to segregate the affected product immediately and return it to Boston Scientific. Boston Scientific requested the completion and return of the Reply Verification Tracking Form.For additional information on this recall call Boston Scientific Quality Systems at (763) 494-1133. | Contact the recalling firm for information | 33 | 33 | (USA) Nationwide distribution including the states of AZ, AR, CA, CO, FL, GA, IL, IN, NV, NY, NC, OH, PA, RI, and TX. | Terminated | August 10, 2012 | 353 | 33 | Not_Computer
Z-3063-2011 | 59523 | products, red-cell lysing products | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A. The STROMATOLYSER-IM lysing reagent is used to destroy red blood cells to aid in the counting and characterization on immature white blood cells. This reagent is used on the Sysmex XE-Series and SE Series hematology analyzers. | Sysmex StromatolyserIM (SIM220A) | 2 | August 23, 2011 | 2011 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent were found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated April 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax to Sysmex Customer Service at 800-689-8296. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the notification are directed to the Sysmex Technical Assistance Center at 1-888-879-7639. | No consumer action necessary | 1887 | 1,887 cubetainers | Worldwide Distribution--USA (nationwide and Puerto Rico) and the countries of Canada, Colombia, Ecuador, Mexico, Netherland Antilles and Peru. | Terminated | March 14, 2012 | 204 | 1887 | Not_Computer
Z-3054-2011 | 59276 | antigens, all groups, streptococcus spp. | GTY | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | @ Home Strep TestRapid Strep TestProduct Usage: For in-vitro diagnostic use - detection of Strep. | Home Strep Test | 2 | August 23, 2011 | 2011 | Jay Bird Investments Inc | Home Strep Test does not qualify as a Class I device and needs pre-market approval. | PREMARKET APPROVAL: No Marketing Application | Jay Bird Investments, Inc. sent a notification letter dated July 6, 2011 to customers. The letter describes the product, problem and the actions to be taken. The customers were instructed to notify retailers and remove the product from the shelves. For questions regarding disposition of the product please contact Myron Bouwhuis at 801-721-3528. | Contact the recalling firm for information | 7541 | 7,541 | Nationwide Distribution (USA) including the states of OR and UTAH. | Terminated | June 20, 2012 | 302 | 7541 | Not_Computer
Z-3070-2011 | 59588 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Entroy Pool Lift; a patient lift used in conjunction with a chair or stretcher to lift them in/out of a swimming/therapy pool; Arjo Hospital Equipment AB, Eslov, Sweden; Models GAB1000-01-xx, GAB1000-11-xx and GAC100xEntroy is a pool lift that docks to either a chair or a stretcher and aids disabled people getting in and out of a pool. This equipment is intended for disabled persons who need to transfer in and out of a therapy and swimming pool under supervision of skilled and trained personnel in accordance with the Entroy Operating and Product Care Instructions. All other use must be avoided. Active/semi-active persons can and should use the Entroy chair (with or without chassis) while bedridden persons without trunk stability can and should be transferred with the Entroy stretcher with chassis. | Entroy Pool Lift | 2 | August 24, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The Entroy Pool Lift boom may rapidly lower un-commanded to the bottom of its stroke, with a potential for injury to the resident or caregiver. | DESIGN: Device Design | ArjoHuntleigh sent an Urgent Field Correction Notice I.D. 110701AHE to the affected end user via UPS 2nd Day on August 4, 2011. The notice informed the account of the potential for lift boom to rapidly lower un-commanded to the bottom of its stroke. The account was requested to immediately cease use of the affected lift, and disconnect the power to the lift to prevent further use until corrected. The account was requested to complete and sign the Customer Response Form, and return it to ArjoHuntleigh to schedule a service call to correct the lift. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 54485. |  | 1 | 1 unit in the U.S., 311 units in the rest of the world | Worldwide Distribution - USA including CT | Terminated | November 28, 2011 | 96 | 1 | Not_Computer
Z-3073-2011 | 59680 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Habib, 4x Kit, Bipolar Resection Device RITA (Radiofrequency Ablation Devices). AngioDynamics, Inc.,One Horizon Way, Manchester, GA 30816.Intended to be used to assist in coagulation of tissue during intraoperative surgical procedures. | Habib, 4x Kit 4401 | 2 | August 24, 2011 | 2011 | Angiodynamics, INC | The labeling contains errors which result in the products being misbranded | MISBRANDING: Labeling False and Misleading | AngioDynamics, Inc. sent out a "Medical Device Recall" letter dated August 30, 2010 to all affected customers.The letter included; products affected, problem and asked them to return the product back to AngioDynamics. Customers were asked to fill out and return the included Recall Reply Form. For question on this recall contact AngioDynamics at (800) 772-6446. |  | 527 | 527 units | Worldwide Distribution: (USA) Nationwide distribution and the countries of Austria, China, Denmark, Egypt, France Germany, India, Italy, Latvia, Netherlands, Poland, Russian Federation, South Africa, Switzerland and the United Kingdom | Terminated | March 12, 2012 | 201 | 769 | Not_Computer
Z-3069-2011 | 59560 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Applied Medical 15x100 mm Kii Optical Access Systems, Non-Threaded 6/Box, Model C0R39Product Usage: Applied Medical's C0R36, C0R37, C0R38, and C0R39 trocars are diposable single-use devices indicated for use in general, abdominal, gynecological and thoraic minimally invasive surgeries. Trocars establish a path of entry through tissue planes and into internal spaces for purposes of inserting endoscopic instruments that are used to perform surgery on internal organs. | Applied Medical Kii Optical Access System | 2 | August 24, 2011 | 2011 | Applied Medical Resources Corp | The recall was expanded on December 22, 2009 to include four models: C0R36, C0R37, C0R38, and C0R39. Initially, Applied Medical conducted a voluntary recall on June 9, 2009 of two lot numbers, 1084426 and 1085506, of the 15 mm Kii non-threaded optical access system, model C0R36, due to a potential weakness in the cannula tip, possibly resulting in a tip breakage. | DESIGN: Component Design/Selection | Applied Medical Resources Corp. notified customers via an overnight "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2009. The letter informed customers of the recalled products, problem, and actions to be taken. The letter asked that the listed recalled products be returned immediately. The customers were asked to check their inventory, complete the attached Voluntary Recall Notification Confirmation Form and fax as per the instructions. The recall letter provided customers with a contact and telephone number for product return questions (1-949-713-8652) and a separate contact for regulatory questions (1-949-713-8369). | Contact the recalling firm for information | 13791 | 13,791 units total | Worldwide Distribution - Nationwide in the USA and to countries of: Australia, Austria, Belgium, Canada, Chile, Columbia, Cyprus, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Italy, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, and Switzerland. | Terminated | September 06, 2011 | 13 | 13791 | Not_Computer
Z-3075-2011 | 58489 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, DASH 3000, 4000, 5000 Patient MonitorsThe Dash 3000/4000/5000 patient monitor is intended for use under the direct supervision of a licensed healthcare practitioner. The intended use of the system is to monitor physiologic parameter data on adult, pediatric and neonatal patients. The Dash 3000/4000/5000 patient monitor is designed as a bedside, portable, and intra-hospital transport monitor that can operate in all professional medical facilities including but not limited to: emergency department, operating room, post anesthesia recovery, critical care, surgical intensive care, respiratory intensive care, coronary care, medical intensive care, pediatric intensive care, or neonatal intensive care areas located in hospitals, outpatient clinics, free-standing surgical centers, and other alternate care facilities. Physiologic data includes but is not restricted to: electrocardiogram, invasive blood pressure, noninvasive blood pressure, heart rate, temperature, cardiac output, respiration, pulse oximetry, carbon dioxide, bi-spectral index, impedance cardiography, oxygen, and anesthetic agents as summarized in the operator's manual. The Dash 3000/4000/5000 patient monitor is also intended to provide physiologic data over the UNITY NETWORK TM indirectly to clinical information systems (via our Enterprise Gateway) and allow the user to access hospital data at the point-of-care. The information can be displayed, trended, stored, and printed. The Dash 3000/4000/5000 patient monitor was developed to interface with nonproprietary third party peripheral devices that support serial data outputs. | GE Healthcare, DASH 3000, 4000, 5000 Patient Monitors | 2 | August 25, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue with the Dash 3000/4000/5000 patient monitors configured with the 802.11b wireless LAN option that may impact patient safety. The Dash monitor may exhibit reboots on some networks when the 802.11b wireless LAN option is used to connect to the hospital wireless LAN infrastructure. This reboot results in a loss of patient monitoring for approxima | DESIGN: Software Design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. GE recommended actions that should be taken by a qualified service representative. For any questions regarding the medical device correction or identification of affected items contact Technical Support at 1-800-558-7044 or customers should contact their local Service Representative.UPDATE: A second letter dated 12/13/11 was sent on 1/20/12 to consignees. The revised letter provides an update to the safety instructions and product correction. | Contact the recalling firm for information | 8029 | 8029 devices | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI and the countries of UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVENIA, SINGAPORE, SAUDI A ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, KUWAIT, JAPAN, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, ECUADOR, DENMARK, CANDA, BELGIUM AUSTRIA, AUSTRALIA, and ALGERIA. | Terminated | September 11, 2012 | 383 | 8029 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-3043-2011 | 59317 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂŽ II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. | Medtronic SynchroMed II | 1 | August 29, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | Medtronic is updating information regarding the potential for reduced battery performance that can lead to sudden loss of therapy in a small percentage of Medtronic Model 8647 SynchroMed II pumps that was communicated with Healthcare providers in July 2009. The purpose of the current communication is to provide updated information regarding the scope and occurrence of this issue and to emphasize | DESIGN: Component Design/Selection | Medtronic Neuromodulation notified physicians with an Urgent Medical Device Correction letter beginning July 05, 2011. A press was issued July 08, 2011. The letter described the issue, severity, recommendations if they notice issue, and patient management recommendations. UPDATE: in November 2011, MDT began to exchange unused pumps in inventory with pumps that contained the new battery. A letter dated November 2011 was left behind at locations in which the MDT representative swapped out the devices. | Contact the recalling firm for information | 139653 | 139,653 (105,002 US, 34,61 OUS) | Worldwide including USA, Puerto Rico, Algeria, Aruba, Australia, Austria, Belarus, Belgium, Brazil, Chile. Colombia. Costa Rica. Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Iran Iraq, Ireland, Israel Italy, Jordan, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Panama, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | August 24, 2012 | 361 | 139653 | Battery | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3086-2011 | 59469 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic ControlTip Atraumatic Fenestrated Grasper Forceps TipRef: 3642The ReNew Forceps Laparoscopic Surgical Tips are intended to cut, grasp and dissect various abdominal tissue for use in endoscopic, including laparoscopic surgical procedures where instruments are introduced into the body through a cannula | Microline Surgical ReNew Laproscopic ControlTip Atraumatic Fenestrated Grasper Forceps Tip | 2 | August 29, 2011 | 2011 | Microline Surgical, Inc. | Mislabeled: ReNew product number 3262 (Long Fenestrated Grasper) was packaged and labeled asproduct number 3642 (Super-Atrau Fenestrated Grasper). ReNew product number 3642 (Super-Atraumatic Fenestrated Grasper) was packaged andlabeled as 3262 (Long Fenestrated Grasper | PRODUCTION CONTROLS: Packaging Process Control | Microline Surgical sent a RECALL NOTICE letter dated July 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine product subject to the recall. If customers further distributed the product, customers should identify their accounts and notify them immediately of the product recall. Customers were instructed to contact Microline Customer Service to obtain a Return Authorization Number for return of the recalled item. Customers were asked to complete and return the enclosed Recall Return Response Form as soon as possible. For any questions customers were instructed to call the Regulatory Affairs Manager at 978-867-1726. | No consumer action necessary | 300 | 300 units | Worldwide Distribution - USA including FL, MI, MN, TN and WI and the countried of Canada, The Netherlands, Thailand, UK, Norway, and Italy. | Terminated | September 09, 2011 | 11 | 470 | Not_Computer
Z-3079-2011 | 58874 | [6 empty] | Elekta SynergyThe Elekta Synergy System is intended to be used for radiation therapy treatment of malignant neoplastic diseases, as determined by a licensed medical practitioner. | The Elekta Synergy System | 2 | August 29, 2011 | 2011 | Elekta, Inc. | A small number of the Modular Reverse Diodes (D1A), a component in the equipment room of the Elekta Linear Accelerator, have ruptured. | OTHER/UNDETERMINED: Pending | ELEKTA sent an IMPORTANT NOTICE (FCO 20000103022), dated May 6, 2011. The notice identified the product, the problem, and the aciton to be taken by the customer. Elekta recommended that service personnel always wear protective safety glasses, when they do work on the modulator assembly with the covers removed. Elekta also recommended that users make sure that the modulator cover is installed when the digital accelerator is in operation. Elekta will release a Field Change Order when a corrective action is identified. | Discard product or return to place of purchase for a full refund or exchange | 368 | 368 | Nationwide and Puerto Rico | Terminated | October 02, 2012 | 400 | 368 | Not_Computer
Z-2600-2011 | 58804 | analytical measuring and detection laser product | REO |  |  |  |  | Prism Laser 7642Laser System | Prism Laser 7642 | 2 | August 29, 2011 | 2011 | American Stress Technologies, Inc. | It was discovered The Prism 7642 was not certified at the time of delivery. The product lacks the required labeling and the product was sold and delivered without the submission of the Initial Product Report. The product lacks User In formation | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | American Stress Technologies sent a customer notification letter dated April 1, 2011, informing them of the action they have taken to correct this issue. (1) ASTI conducted laser safety tests and checks on May 3,2011 to verify the Prism was compliant with FDA regulations. This was confirmed in a June 17, 2011, electronic mail message.(2) ASTI placed the required labels described in the initial product report on the Prism on May 3, 2011.(3) ASTI delivered updated user manuals to UNO on May 3, 2011, that describe potential laser hazards and the required labeling in the same way as described in the initial product report.(4) ASTI completed the above corrective actions at no cost to UNO as confirmed in a June 20, 2011, electronic mail message.If you have any further questions please call (724) 410-1030. | Contact the recalling firm for information | 1 | 1 | Nationwide Distribution | Open |  |  | 1 | Not_Computer
Z-3083-2011 | 59483 | hemodialysis system for home use | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K@homeThe Fresenius 2008K@home is indicated for acute and chronic dialysis therapy in an acute or chronic facility. The 2008K@home is also indicated for hemodialysis in the home and must be observed by a trained and qualified person as prescribed by their physician. | Fresenius | 2 | August 29, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Upgrade software to implement new on screen graphics and instructions, the addition of a bracket on the side of the machine to temporarily hold the bloodlines during set up in order to conform to approved 510(k) | DESIGN: Device Design | Fresenius Medical NA issued a formal field correction notification dated May 4 ,2011, sent by certified mail with return receipt requested. The letter identified the product the problem and the action needed to be taken by the customer. The customer was notified of the retrofit that would be done to their machines. The patients' facilities were made aware of the changes that would be made to each of their patient's 2008K@home machines on May 4, 2011. On May 13, 2011, all facilities with active 2008K@home machines were sent a notification letter of the retrofit.Please return the receipt for confirmation that you have received this notice.Please call (781) 699-2053, to schedule a certified technician to conduct the machine retrofit. | No consumer action necessary | 705 | 705 units | Nationwide Distribution.Canada - FMCNA is not conducting the recall to Canada, as Health Canada did not require the same modifications to the machines | Open |  |  | 705 | Software |  | Software update | Software Update
Z-3081-2011 | 59462 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Mini MaxLock Extreme Plating System Case & Tray, Model Number(s): MXM-901, Product Usage: The Mini MaxLock Extreme Plating System is intended to stabilize and aid in the repair of fractures, fusions, and osteotomies for small bones and bone fragments. Surgical Case and Tray Delivery System used to house plating and screw orthopedic implant system and instrumentation. | Mini MaxLock Extreme Plating System Case and Tray | 2 | August 29, 2011 | 2011 | Orthohelix Surgical Designs Inc | The surgical tray does not meet our Instruction for Use (IFU) for dry time requirement. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Orthohelix Surgical Designs Inc sent a RECALL NOTIFICATION letter dated July 6, 2011 and an UPDATED URGENT RECALL 2ND NOTIFICATION letter dated July 13, 2011 to all affected customers. The letter identified the recalled product, the problem, and the Field Correction actions to be taken. Customers were instructed to remove the blue rubber mat in the utility bay of the tray base immediately, tag the mat with the tray number, and return it to recalling firm immediately. The letter states once customers have removed the blue rubber mat from the tray base, they may continue to use the tray and system normally. No further actions or precautions are needed. Distributors were also asked to notify their customers of the recall. If you have any questions, please call 330-247-1444 or 330-247-1445. | Contact the recalling firm for information | 61 | 61 trays | USA Nationwide Distribution including the states of: AR, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, WI & WY. | Terminated | September 10, 2012 | 378 | 61 | Not_Computer
Z-3080-2011 | 59400 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 754/754M Disposable MRI Ventilator Circuit Impact Instrumentation Inc., 23 Fairfield Place, West Caldwell, NJ 07006.Accessory for 700 Series 754/754m Ventilators during MRI examinations | Model 754/754M Disposable MRI Ventilator Circuit | 2 | August 29, 2011 | 2011 | Impact Instrumentation, Inc. | Model 754/754 M Disposable MRI Ventilator Circuits are being distributed without market clearence (510k). | MISBRANDING: Labeling False and Misleading | Impact Instrumentation Inc. sent an "Urgent Product Recall" dated May 27, 2011 to all affected customers.The letter informed customers the product is not a cleared product and to return the product back to Impact Inst. for credit. A notification form is included in the letter to be return to the firm.For information on this recall contact the firms Customer Service Department at (800) 969-0750, 8:30 am to 5:00 pm EST. |  | 492 | 492 units | Nationwide Distribution; including the states of CA, FL, IL, NJ, OH, TX, UT, MN, MO, MS, OH, OR, TX, AR, and PA. | Terminated | May 23, 2013 | 633 | 492 | Not_Computer |  |  | Other
Z-3084-2011 | 59577 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20HL 20 4-pumps console base: P/N 917741HL 20 5-pumps console base: P/N 917751Intended for use as an extracorporeal circulation device for perfusion lasting not more than six hours. The HL console is equipped with a rack for accommodating the system's electronic modules. | Maquet HL Console for Heart Lung Machine HL20 | 2 | August 29, 2011 | 2011 | Maquet Inc. | Maquet HL 20 Twin Pump Modules may unintentionally increase speeds for a 1.5-2 second period of time before returning to its original speed. | DESIGN: Software Design | Maquet sent a PRODUCT RECALL NOTIFICATION LETTER/return response forms, dated June 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until a software update was performed by their trained technicians. Customers were also asked to forward the information contained in the notification letter to all users and to all appropriate Biomedical Staff. Customers were instructed to acknowledge receipt of the notification letter by completing and faxing back the attached response form to 973-807-9210 whether or not they have any of the affected product. For questions customers were instructed to contact their local MAQUET representative. For questions regarding this recall call 973-709-7994. |  | 299 | 299 units | Nationwide Distribution | Terminated | October 16, 2012 | 414 | 299 | Software | Device Operation | Software update | Software Update
Z-3078-2011 | 59370 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 2.0 (software version 1.1.23 and 1.1.24)Roche Diagnostics Operations, Inc.An automated instrument used for isolation of nucleic acids (DNA, total RNA, total viral nucleic acids) from different kinds of biological research sample material (whole blood, serum, blood cells, culture cells, tissue, bacteria, fungi) using the specially designed MagNA Pure LC kits, for the purpose of life science research only and automated filling of different kinds of PCR reaction vessels (LightCycler Capillaries, 96-well PCR plates, PCR tube strips) with PCR reaction mixes and template nucleic acid, or other reaction tubes for the making of dilution series, reaction mixes. | MagNA Pure LightCycler | 2 | August 29, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Roche has detected a problem with the MagNA Pure LC 1.0 (software version 3.0.11) and the MagNA Pure LC 2.0 (software versions 1.1.23 and 1.1.24) instruments. Under certain circumstances, only a fraction of each sample is processed which may lead to false negative results, and an impact to patient health if used in a homebrew diagnostic application. | OTHER/UNDETERMINED: Pending | Roche Diagnostics Operations, Inc. sent a letter dated July 21, 2011 entitled "URGENT MEDICAL DEVICE CORRECTION" to all affected customers.The letter included the name product, software versions, and reason for recall and informed them the problem is under investigation. In the meantime, customers are asked to wait 30 seconds after the finalization of the decontamination cycle before initiating the next operation, in order to avoid this issue.For questions on this recall please call Roche at (800) 428-5074. | Contact the recalling firm for information | 148 | 148 | Nationwide Distribution; including the islands of Puerto Rico and Hawaii. | Terminated | July 02, 2012 | 308 | 570 | Computer
Z-3082-2011 | 59399 | prosthesis, elbow, semi-constrained, cemented | JDB | Elbow joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Latitude", Tige Humerale, Humeral Stem, Medium 77 mm, Humeral Stem Right, For Cemented use only, Sterile R, REF 0030402, Tornier In. Edina, MN 55435 USA.The Tornier In. Latitude Elbow Prosthesis is intended for total elbow arthroplasty. Prosthetic replacement with this device may be indicated to relieve severe pain or significant disability following the effects of primary or secondary osteoarthritis and rheumatoid arthritis; correction of functional deformities; revision procedures where other treatments or devices have failed; treatment of fractures that are unmanageable using others techniques. The Tornier In. Latitude Elbow Prosthesis is intended for cemented use only. | Tornier Latitude Elbow Prosthesis | 2 | August 29, 2011 | 2011 | Tornier, Inc | Tornier has discovered that some models of the Tornier latitude Elbow Prosthesis (Small and Medium Humeral stems) were manufactured with a humeral screw that may not function as designed. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Tornier, Inc. sent an "Urgent Product Correction" letter dated June 23, 2011. The letter was addressed to the doctors. The letter described the product and the problem. Doctors were insructed to monitor patients that have received the units for evidence of humeral screw loosening.For questions customers were instructed to contact the Customer Service Department at 1-888-494-7950. | No consumer action necessary | 3 | 3 | Nationwide Distribution including FL and MO | Terminated | July 17, 2012 | 323 | 3 | Not_Computer
Z-3091-2011 | 57806 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Software version 3.3 for Rapid Programmer devices, Model 3832.Software version 3.3 for Rapid Programmer devices, a personal digital assistant (PDA) based programmer supporting implantable pulse generators of the Eon, Genesis, and Renew family. | Rapid Programmer | 2 | August 30, 2011 | 2011 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.All affected devices were returned to the firm as instructed. |  | 14 | 14 units | Distributed in New Hampshire and Puerto Rico. | Terminated | September 27, 2012 | 394 | 14 | Software |  | Software update | Software Update
Z-3092-2011 | 59754 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Evident MWA Generator, REF VTGEN, Output: 60 W / 50 Ohm, Valleylab, a division of Tyco Healthcare Group LP Boulder, CO 80301.Intended for the coagulation of soft tissue. Not intended for use in cardiac procedures. | Evident MWA Generator | 2 | August 30, 2011 | 2011 | Covidien, PLC | Generator may deliver full power when a lower power is selected. | DESIGN: Component Design/Selection | Covidien sales reps began notifying affected customers on February 6, 2009, followed by a letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the affected product to:CovidienEnergy Based Devices5920 Longbow RoadBoulder, Colorado 80301Customers were asked to label the exterior of the carton with their Return Authorization number. Covidien EbD would address the issue and return the unit to them.For any questions customers were asked to call Covidien EbD's Service Department toll-free at 1-800-255-8522, x 4923 or contact their local sales representative for additional information. |  | 357 | 357 units | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Denmark, France, Germany, Hong Kong, Hungary, Israel, Italy, Netherlands, Poland, Portugal, Romania, Singapore, Slovakia, Spain, Switzerland, and the UK. | Terminated | August 30, 2011 | 0 | 357 | Not_Computer
Z-3100-2011 | 59351 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, Interpretation of a RapidArc Treatment Plan as a 3D Conformal Arc Treatment Plan; version v8.6 or 8.8 in combination with treatment planning systems; Model Number: H51; Reference/PSCA Identifier: CP-03063; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CAProduct Usage: 4D Integrated treatment console is designed to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | 4D Integrated Treatment Console (4DITC) | 2 | August 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The firm has identified that in certain configurations, the 4DITC can interpret a RapidArc or VMAT treatment field as 3D Conformal Arc Treatment Plan. If a RapidArc or VMAT treatment field is incorrectly interpreted as Conformal Arc, the field will be delivered with a constant dose rate and gantry speed for the entire duration of the delivery, which can lead to delivered dose being different than | DESIGN: Software Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 27, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that Varian Service will contact customers to arrange to perform update to 4D Integrated Treatment Console (4DITC) software. For further clarification, contact your local Varian Customer Support District or Regional Manager. The letter also provides Varian Oncology Help desk contact information 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  | 645 | 645 | Worldwide Distribution - USA (nationwide) | Terminated | December 05, 2013 | 828 | 645 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3089-2011 | 59662 | unit, phacofragmentation | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Bausch & Lomb Stellaris Optimized Stability Vacuum Pack, Ref: BL5112, packaged 6 units/case, RX, Sterile. Bausch & Lomb Incorporated, Rochester, NYFor use with the Bausch & Lomb Stellaris Vision Enhancement System for the phacoemulsification of an opacified crystalline lens during anterior segment surgery | Bausch & Lomb Stellaris Optimized Stability Vacuum Pack, Ref: BL5112 | 2 | August 30, 2011 | 2011 | Bausch & Lomb Inc | The aspiration channel on the affected vacuum cassettes may be blocked due to a molding malfunction | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Bausch & Lomb Inc. notified all affected customers by direct telephone contact on July 27, 2011 and a follow-up FAX or e-mail dated July 27, 2011The letter included; affected product, problem, and requests customers return the product. An acknowledgement form was included to complete and return with the product.For question on this recall contact Bausch & Lomb at (949) 521-7895. | Contact the recalling firm for information | 8 | 2/6-unit cases | Nationwide Distribution. | Terminated | June 21, 2013 | 661 | 270 | Not_Computer
Z-3097-2011 | 59781 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Encision AEM Suction Irrigation Electrode, Needle Tip, REF ES3885, Encision Inc., Boulder, CO 80301. | Encision AEM Suction Irrigation Electrode, Needle Tip, | 2 | August 30, 2011 | 2011 | Encision, Inc. | Electrode tips have been reported breaking off during use or cleaning. | DESIGN: Device Design | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers.The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory.For information on this recall call Encision at (303) 339-6917. |  | 10 | 10 units | Worldwide Distribution: Nationwide distribution, and the country of New Zealand. | Terminated | September 04, 2012 | 371 | 1078 | Not_Computer
Z-3099-2011 | 59238 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PFC Sigma Round Dome Patella 38 MM 3 PegDevice is indicated for use in knee replacement procedures in conjunction with femoral knee prosthesis components and polymethylmethacrylate (PMMA) bone cement. | PFC Sigma Round Dome Patella | 2 | August 30, 2011 | 2011 | DePuy Orthopaedics, Inc. | DePuy initiated a voluntary recall effective immediately, for one lot of P.F.C. SIGMA RD Dome Patella 3 Peg (38mm) due to packaging labels that did not correctly reflect the size of the enclosed patellar component. In the recalled lot, the packaging is labeled for a P.F.C. SIGMA 38mm 3 peg Round Dome Patella, but the product enclosed in the package is a P.F.C. SIGMA 35mm 3 peg Oval Dome Patella. | TRAINING: Employee Error | The firm, Depuy , sent an email that included a letter entitled "URGENT INFORMATION - RECALL NOTICE" dated June 28, 2011 to DePuy Area customers ( Directors, Distributors, Office Managers, Hospital Representatives, and Doctors). The letter described the product, problem and actions to be taken. The letter also included implications of implanting the recalled device and patient communication information. The customers were instructed to immediately cease further implantation/use or distribution of the product, immediately return any recalled devices in their inventory, notify their customer of the recall and complete, sign and return the Reconciliation Form(s) via fax to DePuy at 574-372-7567 within seven business days.The letter also states that DePuy is not recommending prophylactic revision or additional follow up in the absence of symptoms. Additionally, DePuy recommends that surgeons continue to follow patients as per their standard care.Questions regarding the recall may be directed to the following DePuy representatives--questions about recall information call the Manager of Customer Quality at 574-372-7333; Salesforce questions, call the Group Product Director, 574-372-5046 or Associate Product Director at 574-372-7154 and for Clinical/Surgeon questions, contact DePuy's Scientific Information Office at 1-888-554-2482. | Contact the recalling firm for information | 29 | 29 | Nationwide distribution: CA, MI, NJ and WI. | Terminated | August 28, 2013 | 729 | 29 | Not_Computer
Z-3102-2011 | 57590 | urethrotome | EZO | Urethrotome. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OLYMPUS Pediatric Knives and Resection ElectrodesThe pediatric knives are intended for use in urethrotomy in pediatric applications. The HF resection electrodes are intended for use in transurethral resection in pediatric applications. | OLYMPUS Pediatric Knives and Resection Electrodes | 2 | August 31, 2011 | 2011 | Olympus America Inc. | The device sterile package seal shows evidence of being breached or was breached | PRODUCTION CONTROLS: Packaging Process Control | Olympus America Inc. sent an "URGENT PRODUCT RECALL" letter dated December 22, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected product and to destroy or discard, if found. A questionnaire was attached to the letter for customers to complete and return via fax to Olympus at 484-896-7128.Contact the firm at 484-896-5688 for questions regarding this recall. |  | 1305 | 1305 devices | USA (nationwide) including the states of AL, AR, CA, CO, CT, FL, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OR, TX, VA, VT, and WI. | Terminated | September 28, 2012 | 394 | 1305 | Not_Computer
Z-3106-2011 | 59229 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Mitek Ratchet Handle with Quick Connect Product code: 219970. Instrument used with anchor inserters for insertion of DePuy Mitek Interference Screws (orthopedic). | Depuy Mitek | 2 | August 31, 2011 | 2011 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Quick Disconnect Ratchet Handle may slip up the driver shaft or disengage leaving the shaft in the screw head. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Depuy Mitek issued Original Customer Notification Letter on 6/27/11 and an Updated letter to clarify the lot number information sent 6/29/11. Hospitals requested to immediately check all inventories for affected product and return for repalcement product. Also notify surgeons aware of this potential issue and potential risk of a delay in surgery. questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. |  | 75 | 75 units | Nationwide distribution. | Terminated | January 19, 2012 | 141 | 75 | Not_Computer
Z-3103-2011 | 59217 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrachyVision, versions 8.2, 8.5, 8.6 and 8.9. Varian Medical Systems.3100 Hansen Way, Palo Alto, California 94304-1038.Used to plan radiotherapy treatments for patients with malignant or benign diseases. In addition, the module is specifically indicated for planning proton treatment of neoplasms of the eye. | BrachyVision Software, part of Eclipse TPS | 2 | August 31, 2011 | 2011 | Varian Medical Systems, Inc. | Medical device for brachytherapy treatment may cause users to enter incorrect syntax and lead to potential misinterpretation of patient treatment plan. | DESIGN: Software Design | Varian Medical Systems, Inc. sent a "Medical Device Recall" letter dated August 05, 2010 to all affected customers.The letter included description of product affected and problem. Customers were advised to review template files to ensure that there are no characters immediately after the string in the customized dose quality parameters. Device users were requested to return a response form.For information on this recall contact Varian Technical Support at (800) 360-7909. |  | 1252 | 1252 units | Worldwide and Nationwide Distribution. | Terminated | November 25, 2011 | 86 | 1252 | Software |  | Instructions | Safety Notice/Insructions
Z-3101-2011 | 59262 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager (OBI) 1.3, 1.4, 1.5; and Trilogy Mx, TrueBeam 1.0, Offline Review 1.0 - 2.0;Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CAThe On-Board Imager device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields relation to anatomical and/or fiducial landmarks. Trilogy Mx is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation treatment is indicated. | OnBoard Imager (OBI); Trilogy Mx | 2 | August 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified a software anomaly in the 3D Match environment, regarding a slight mismatch between CT image and Structures in On-Board Imager w/s. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 21, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter recommends several user actions. The firm is developing a correction for the issue and will contact customers when the correction is available.Questions regarding this notice may be addressed by contacting the Varian Oncology Help Desk at 1-888-827-4265. |  | 2268 | 2268 | Worldwide Distribution | Terminated | March 15, 2012 | 197 | 2268 | Software | Display/Image | Correction | Repair
Z-3113-2011 | 59632 | enzyme immunoassay, propoxyphene | JXN | Propoxyphene test system. | Toxicology | Toxicology | 510(k) | DrugCheckÂŽ, Distributed by Express Diagnostics Int'l Inc., 1550 Industrial Drive, Blue Earth, MN 56013. Following cup Model numbers involved in the recall: 60670, 60700, 60713, 60717, 60870, 60923, 61012, 61019, 61023, 61024, 61040, 61042, 61057, 61102, 61105, 61203, 61204, 61206, 61208, 61301, 61303, 61401, 70603, 70809, 71032, 71043, 71055, 71065, 71202, 71204, 71205, 71401, 71402, 60718-3, 60720-6, 60925-A, 60927-3, 60930-A, 60980-3A, 61002-3, 61010-4, 61023-A, 61028-3A, 61039-4, 61040-A, 61055-A, 61056-3, 61057-6, 61100-3, 61102-3,61104-3,61105-5 CRL,61120-3,61202-4,61204-4,61204-A,61206-3A,61207-4,61302-3, 61303-A, 61401-A, 70660-3, 70716-3, 71045-A, 71055-6, 71102-3A, 71204-4, 71206-3A.In Vitro immunoassay screening device for the qualitative detection of multiple drugs of substances in human urine. | DrugCheck | 2 | August 31, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc (EDI), is recalling DrugCheckÂż Drug Screen Cups containing buprenorphine (BUP) and propoxyphene (PPX) because they must be labeled as "forensic use only", meaning they cannot be distributed or used for treatment or diagnostic purposed in the United States. | MISBRANDING: Labeling False and Misleading | The firm, Express Diagnostics Int'l, Inc., sent a "RECALL NOTICE" letter dated October 4, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to notify their customers of the labeling change, and complete and return the DISTRIBUTOR/CUSTOMER CONFIRMATION FORM via fax to Express Diagnostics Int'l, Inc., at 1-507-526-2252. The letter also noted that "Further shipment of product containing BUP and/or PPX cannot continue until Express Diagnostics Int'l, Inc. receives the signed form."If you have any questions, call 507-526-3951. | No consumer action necessary | 196750 | 196,750 | Nationwide distribution:AZ, CA, CT, GA, MD, MA, MN, MO, NH, NJ, OR, SC, and WA. | Terminated | July 06, 2012 | 310 | 196750 | Not_Computer
Z-3104-2011 | 59567 | alcohol dehydrogenase, specific reagent for ethanol enzyme method | DIC | Alcohol test system. | Toxicology | Toxicology | 510(k) | Drug Check, Distributed by Express Diagnostics Intl Inc., 1550 Industrial Drive, Blue Earth, MN 56013.Following Cup and Dip product models involved in the recall: 409004090130500-A30575 -A30603-A30671-A30800-A30802-3A30900-A30950-A31102-A40901C60410-3A60500-A60505-A60524-A60526-A60535-A60600-A60602-4A60630-A60635-4A60640-4A60705-4A60709-A60800-6A60800-A60801-A60820-A60903-A60910-A60921-A60925-A60930-A61020-A61023-A61026-4A61028-3A61030-3A61044-6A61052-4A61100-A61204-A61206-4A61206-6A61303-A61401-A61404-A65500-4A70410-4A70500-3A70500-A70550-A70602-3A70604-6A70630-3A70640-4A70808-6A70910-AIn Vitro immunoassay screening device for the qualitative detection of multiple drugs of substances in human urine. | DrugCheck | 2 | August 31, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Intl'l Inc is recalling DrugCheck devices because they were informed via letter from the alcohol test strip manufacturer on October 27, 2010 that the product was not 510 K cleared for in vitro diagnostic use and could not be labeled as such in the US. | PRODUCTION CONTROLS: Error in Labeling | Express Diagnostics Intl., Inc. sent an Updated Recall Notice dated November 1, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have open orders of the affected product to contact customer service about their order. Customers were asked to complete and fax the form on page 3 of the document to 507-526-2252. Further shipment of products could not continue until Express Diagnostics received the signed form.For questions regarding this recall call 507-526-3951. | No consumer action necessary | 56650 | 56,650 | Nationwide Distribution including AZ, CA,K CT, FL, GA, IL, IA, LA, MD, MA, MN, MS, MO,K MT, NE, NJ, NY, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY. | Terminated | July 06, 2012 | 310 | 56650 | Not_Computer
Z-3108-2011 | 59439 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HiÂŽArt versions 4.0.x, and HD versions 1.0.x, H-0000-0003.The TomoTherapy HI-ART System is intended to be used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. The megavoltage x-ray radiation is delivered in a rotational, non-rotational, modulated (IMRT), or non modulated (non-IMRT/three dimensional conformal) format in accordance with the physician approved plan. | TomoTherapy HiArt System | 2 | August 31, 2011 | 2011 | Accuray Incorporated | As a result of an internal review, TomoTherapy has identified an issue with HiÂżÂżArt versions 4.0.x, and HD versions 1.0.x that we would like to bring to your attention.During DICOM export of plan level images with a non square exported Field of View (FOV), an anomaly in the process of squaring the plan level image may cause the image to shift with respect to ROIs and dose. When the anomaly occu | DESIGN: Software Design | TomoTherapy sent a "Urgent Medical Device Correction" letter dated July 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This issue will be remedied with HiÂŽÂŽArt software version 4.1 and HD software version 1.1 and later.If you have any questions, please contact the TomoTherapy Technical Solutions Center by e-mail "support@tomotherapy.com" . Please refer to the letter for what phone number you should call depending on where you are. | Contact the recalling firm for information | 346 | 346 | Worldwide Distribution-- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and the countries United Arab Emirates United Kingdom Turkey Switzerland Sweden Spain Saudi Arabia Poland Netherlands Italy Germany France Belgium Taiwan Singapore Malaysia Korea Japan India Hong Kong China Australia Mexico El Salvador and Canada | Open |  |  | 346 | Software | Display/Image | Software update | Software Update
Z-3111-2011 | 59297 | display, cathode-ray tube, medical | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips Medical:865047 IntelliVue Clinical Information Portfolio Release D.03865209 IntelliVue Clinical Information Portfolio Release E.00 and E.01865300 IntelliVue Clinical Information Portfolio NICU LinkIntended for use in data collection, storage, and management with independent bedside devices, and ancillary systems that are connected either directly or through networks. It is indicated for use by healthcare providers whenever there is a need for generation of a patient record and computation of drug dosage. | IntelliVue Clinical Information Portfolio | 2 | August 31, 2011 | 2011 | Philips Healthcare Inc. | Under certain circumstances, edits to an order can result in unintended scheduled interventions on the IntelliVue Clinical Information Portfolio. A patient may receive interventions or additional medication doses that are not intended. | DESIGN: Software Design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice. Philips has developed a software upgrade to address this problem in the short term. This upgrade will be available immediately. This software upgrade will be performed by Philips service personnel. While customers await the correction of their devices, customers are instructed to follow the instructions for use for the device, which includes the directions to:1. Always use the medication administration record (MAR) in determining patient care.2. Always review orders and pending administrations in the administration record after any changes to an order are stored.For the longer term, Philips is developing safeguards intended to minimize the potential for similar errors resulting from order changes and upgrade to be available in September 2011.If you need any further information or support concerning this issue, please contact your local Philips representative. | Contact the recalling firm for information | 454 | 454 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open |  |  | 454 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3098-2011 | 59350 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, version 8.1 SP2 to 10.0 SP1; Model Number: H51; Reference/PSCA Identifier: CP-03940; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CAProduct Usage: 4D Integrated treatment console is designed to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | 4D INTEGRATED TREATMENT CONSOLE | 2 | August 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has scheduled for an upgrade of the device from current version to 10.2. There are circumstances in which the 4DITC may be involved in a failure to correctly record treatment history information. | DESIGN: Software Design | Varian Medical Systems, Inc. sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 23, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that the recalling firm will issue a Customer Technical bulletin to inform users of the mandatory upgrade and related information. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. Any questions regarding this Product Notification letter, contact Varian Medical Systems Help desk at 1-888-827-4265 for USA and Canada and +41417498844 for Europe. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) | Terminated | March 15, 2012 | 197 |  | Software | Device Operation | Software update | Software Update
Z-3105-2011 | 59572 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Treatment version 6.6.5052.Varian Medical Systems,Palo Alto, Ca Intended to provide accurate treatment set-ups for each patient by monitoring linear accelerator set-up parameters and by preventing the radiation therapy device form commencing irradiation while any linear accelerator parameter is out of conformance with the treatment plan. | Varian Treatment | 2 | August 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential for Varian Treatment application failing to load the physical wedge filter for a treatment field on a Siemens Mevatron linear accelerator. Delivery of a treatment field without the planned wedge would deliver more dose than intended and would change the shape of the dose distribution for that field. | OTHER/UNDETERMINED: Pending | Varian Medical Systems sent out "Urgent Medical Device Correction/Urgent Field Safety Notice" letters dated July 20, 2011 to all affected customers.The letter identified product, problem and provided recommended user actions when setting up and running the system. It also asked customers to notify appropriate personal provided instructions.For information or questions on this recall please call Varian at (888) VARIAN5 (888- 827-4265). |  | 118 | 118 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, FL, GA, IN, MA, MD, MI, MN, MO, NJ, NY, PA, SC, TN, VA and WV; and the countries of Canada, Austria, Belgium, Brazil, Denmark, Finland, France, Germany, India, Ireland, Italy, Lithuania, Poland, Portugal, Romania, South Africa, South Korea, Spain, Sweden, Switzerland and UK. | Terminated | December 05, 2013 | 827 | 118 | Other | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3116-2011 | 59777 | patient examination glove, specialty | LZC | Patient examination glove. | General Hospital | General Hospital | 510(k) | MICRO-TOUCH(R), L (indicating size large), NITRILE, Powder Free, Medical Examination Gloves, REF 6034303, 200 Pieces, Manufactured for: Ansell Healthcare - Made in Malaysia | MICROTOUCH(R), NITRILE, Powder Free, Medical Examination Gloves | 2 | August 31, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves were inadvertently released for distribution prior to completion of testing required by the U.S. Food and Drug Administration. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Ansell account managers were notified on August 4, 2011 to contact (by phone, fax, e-mail) their respective accounts and notify them of the recall situation. The Ansell account managers were provided with a copy of the recall notification to provide to their accounts. The recall notification includes a request that if these gloves have been further distributed to other customers, please advise them of the recall situation and have them return any outstanding stocks. |  | 250 | 250 cases of 10 | Nationwide distribution: NC, GA, IL, MO, LA, and PA | Terminated | November 27, 2012 | 454 | 1442 | Not_Computer
Z-3107-2011 | 59470 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Li-Ion Docking Station; a lithium ion battery powered docking station for use with GemStar Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; List Numbers 13075-04-01, 13075-04-03, 13075-04-05, 13075-04-07, 13075-24-01, 13075-24-03, 13075-24-05 and 13075-24-07, Note: list numbers with xxxxx-04-xx are for U.S., Canada, Japan, Netherlands and Singapore, while those with xxxxx-24-xx are for international distribution only.The docking station for the GemStar pump is used to provide an alternate power source, acting either as a direct current battery backup, or as an alternating current power source when the docking station is plugged into AC power. The docking station also serves as a stand in which to place the GemStar during pump operation. | Hospira GemStar Docking Station | 2 | August 31, 2011 | 2011 | Hospira Inc. | Hospira has received reports of sparking, smoking, charring and electrical shock when using the GemStar Docking Station due to fluid ingress into the docking station. | DESIGN: Device Design | Hospira Inc. sent an "URGENT DEVICE FIELD CORRECTION" letter dated August 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The requests that customers do not hang or place fluid containers over the docking station, not spray fluid directly on the docking station, and to clean the docking station using a cloth dampened with an apporoved cleaning solution. A reply form was attached to the letter for customers to complete and return via fax at 262-577-6921.Contact Hospira at 1-800-441-4100 for questions concerning this notice.*** Hospira sent an update letter dated September 26, 2012 to the accounts, informing them of the availability of the improved design replacement docking stations, and requesting them to complete and return the enclosed reply via fax or e-mail to Hospira, indicating the number of docking stations to be returned for replacement at no charge if the old docking stations are returned within 45 days of receipt of the replacement stations. ***. |  | 18000 | 18,000 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Japan, Jordan, the Netherlands, Portugal, Saudi Arabia, Singapore, Spain, Sweden and Switzerland. | Open |  |  | 18000 | Not_Computer
Z-3109-2011 | 59366 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 91200, Juggerknot system, Soft Anchors, -Two implants size 1, Biomet Sports Medicine, STERILE. Product Usage: The JuggerKnot Soft Anchors include various size ranges, each consisting of a coreless sleeve and suture(s) assembled on an inserter instrument. The system includes instrumentation allowing for proper preparation and placement of the soft tissue fixation device. | Juggerknot Soft Anchors | 2 | August 31, 2011 | 2011 | Biomet, Inc. | The device packaging was punctured by the device, resulting in loss of implant sterility. | DESIGN: Packaging Design/Selection | Biomet Sports Medicine notified distributor accounts by telephone on July 08, 2011 with instructions to quarantine all lots of the recalled product. Urgent Medical Device Recall Notice letters dated July 21, 2011 were sent to all affected customers. The letter identified the recalled product, problem, and actions to be taken. The customers were instructed to immediately locate and discontinue use of the recalled product. The distributors are advised to communicate this notice to their customers if the recalled products were further distributed. The letter also states that distributors are responsible for the return of the recalled product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Customers were instructed to provide phone confirmation at 800-348-9500, extension 3755 or 3756, when the recall notice is received. Further questions should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8 AM to 5 PM. | Contact the recalling firm for information | 4925 | 4925 | Worldwide Distribution - USA (nationwide) and the countries of: Columbia, Europe and Korea. | Terminated | April 26, 2013 | 604 | 4925 | Not_Computer
Z-3112-2011 | 59345 | drug metabolizing enzyme genotyping systems | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | xTAG CYP2D6 Kit v3 IVD For use with Luminex 100/200 Instrument Luminex Molecular Diagnostics, Inc. 439 University Ave. Toronto, Ontario Canada Catalog # I030B0373 The xTAG CYP2D6 Kit v3 is a device used to simultaneously detect and identify a panel of nucleotide variants found within the highly polymorphic CYP2D6 gene located on chromosome 22 from genomic DNA extracted from EDTA and citrate anticoagulated whole blood samples. | Luminex xTAG | 2 | August 31, 2011 | 2011 | Luminex Corporation | Inaccurate test results, due to possible defective primers, which could lead to incorrect dosage on manufactured drug assays. | OTHER/UNDETERMINED: Pending | Luminex sent a "Customer Advisory" letter dated May 24, 2011, by email and courier to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to destroy any remaining product and to submit a return postcard to confirm that the consignee had been notified of the recall and had destroyed the product.For further questions please call Luminex Technical Support at 1-877-785-2323 ( US ) and 1-512-381-4397 ( Outside US ) Monday through Friday. | Contact the recalling firm for information | 45 | 45 kits | Worldwide Distribution -- USA ( nationwide) and the countries of Belgium, UK and Canada. | Terminated | March 23, 2012 | 205 | 45 | Not_Computer
Z-3110-2011 | 59436 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque antibody CD99 (H036-1.1)Mouse Monoclonal Antibody.Cel Marque, 6600 Sierra College Boulevard, Rocklin, Ca 95677The Cell Marque CD99 Mouse Monoclonal Antibody is intended for qualified laboratories to qualitatively identify by light microscopy, the presence of associated antigens in sections of formalin fixed, paraffin embedded tissue sections using manual or automated IHC test methods. Use of this antibody is indicated as an aid in diagnostic determination of human tumors within the context of the patient's clinical history and other diagnostic tests evaluated by a qualified pathologist of physician. | CD99 (H0361.1) Mouse Monoclonal Antibody. | 3 | August 31, 2011 | 2011 | Cell Marque Corporation | Incorrect label on product bottle, although labeling on outer box is correct. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Cell Marque, notified customers by telephone on July 12, 2011. The customers were instructed to return the product. Cell Marque immediately issued replacement of the product.If you have any further question or require additional documentation, contact the VP of Regulatory and Quality Affairs at (916) 746-8900. | Contact the recalling firm for information | 2 | 2 units | Worldwide distribution: USA (nationwide) include state of : CA and country of: Great Britain. | Terminated | August 31, 2011 | 0 | 2 | Not_Computer
Z-3118-2011 | 59570 | disinfector, medical devices | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance Vision Single Chamber Washer/Disinfector, Model # 1802983,The Vision Single Chamber Washer is intended for use in the cleaning and intermediate level disinfection of soiled reusable utensils, trays, glassware, bedpans and urinals, rubber and plastic goods, simple hard-surfaced rigid surgical instruments such as forceps and clamps, theatre shoes and other similar and related articles found in healthcare facilities. | Reliance Vision Single Chamber Washer/Disinfector | 2 | September 01, 2011 | 2011 | Steris Corporation | STERIS learned that Vision SC Washer users were experiencing nuisance alarms and/or aborted cycles that are adversely affecting the reliability and operation of their unit. |  | The firm, STERIS sent an "URGENT FIELD CORRECTION NOTICE" letter dated May 9, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed that they can continue using their Vision SC Washer(s) prior to installation of the system upgrades. The letter states that STERIS is correcting the issues by implementing a series of quality improvements to the customers Vision SC Washer(s) to ensure satisfaction. Additionally, a STERIS Service Representative will contact the customers facility to arrange for STERIS to upgrade their Vision SC Washer(s).For further information or if you have questions regarding the servicing of your unit, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 404 | 404 Units | Worldwide distribution: USA (nationwide) including states of: L, AR, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI and WY; and countries of: Australia, Canada, Hong Kong and South Korea. | Terminated | March 01, 2013 | 547 | 404 | Not_Computer
Z-3125-2011 | 58705 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis SystemProduct Usage: UniCel DxH 800 Coulter Cellular Analysis System Instructions for Use PN 629743AD The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: 1. Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood. 2. Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF). | Unicel DxH 800 Coulter Cellular Analysis System | 2 | September 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the probe wash collar tubing may become stretched or trapped at the Sample Aspiration Module (SAM) during routine use of single-tube presentation station. When the tension on the tubing is severe, the tubing may not be able to return to its home state and may pull the wash collar out of alignment with the aspirate probe. The misal |  | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated May 11, 2011 with an attached Product Correction Response Form to all affected customers. The letter provides the customers with an explanation of the problem identified, potential impact should the problem occur and actions to be taken. The letter states that this issue will be corrected in a future hardware release. Customers were instructed to share this information with laboratory staff and retain this notification as part of Quality System documentation. If ownership or location of the analyzer(s) has been transferred to another laboratory, customers were instructed to provide a copy of the letter to that party. Customers were also asked to complete and return the enclosed Response Form within ten (10) days, so Beckman can be assured they have received this important communication. Questions concerning this notice, were directed to the Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 680 | 680 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 18, 2012 | 291 | 680 | Not_Computer
Z-3124-2011 | 59752 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Model # P6500, S/N: 400001 through 405005; Mfr. By: STERIS Corporation, Mentor, OH 44060.Product Usage: The SYSTEM 1E Liquid Chemical Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned immersible and reusable critical and semi-critical heat sensitive medical devices, including endoscopes and their accessories. | SYSTEM 1E Liquid Chemical Sterilant Processing System | 2 | September 01, 2011 | 2011 | Steris Corporation | STERIS learned that SYSTEM 1E users may experience aborted cycles due to a failure within the UV light system. | OTHER/UNDETERMINED: Pending | STERIS sent an "Urgent Field Correction Notice" letter dated July 2011 to all affected customers. The letter describes the product, problem and actions to be taken. The letter states that STERIS Service Representative will contact customers to schedule the repair of the recalled product. For further information or questions contact STERIS Field Service Dispact at 1-800-333-8828. |  | 2235 | 2,235 units | Worldwide Distribution - USA (nationwide) and the country of Bahrain. | Terminated | March 19, 2013 | 565 | 2235 | Not_Computer
Z-3122-2011 | 59232 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 3.0T Signa MR750 (Marketing Name: GE Discovery MR750)Product Usage: The GE Sigma MR750 System is a whole body magnetic resonance scanner designed to support high resolution and high signal-to-noise ratio images in short exam times. It is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique anatomical images, spectroscopic images, parametric maps, and/or dynamic images of the structures and organs of the entire body, including, but not limited to head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. The images produced by the GE Sigma MR750 reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and spectra, when interpreted by a trained physician yield information that may assist in diagnosis. | SIGNA 3.0T MR750 SYSTEM | 2 | September 01, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of two potential safety issues associated with MR products (Optima MR450w, Discovery MR450, Discovery MR750, Brivo MR355, Optima MR360) that may impact patient safety.1) Inline post-processed (i.e. automated post-processing) image locations are not correctly annotated under some conditions. 2) Reference images displayed in Functool FiberTrak do not match with | DESIGN: Software Design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated July 27, 2011 to customers with affected systems. The letter described the recalled product, safety issues and recommended actions to be taken. The letter provided immediate mitigation (Safety Instuctions) that address each issue. The letter states that the recalling firm will contact customers to arrange to provide a software upgrade to all affected systems at no cost. For any questions concerning this notification contact your local service representative. |  | 171 | 171 | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, LA,ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WT, WY.and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROAT IA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUKEYand UNITED KINGDOM. | Terminated | August 07, 2013 | 706 | 385 | Software | Display/Image | Software update | Software Update
Z-3119-2011 | 59691 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Operations and Technical Manuals for the firm's WarmAir Model 135 Hyperthermia System (115V, 100V, and 230/240V), 510 (k) #K101148, Serial Numbers: 994-1350001 through 104-1356894. The WarmAir 135 Hyperthermia System is a patient temperature management device which provides forced air warmed by the controller to ... a blanket that is placed over or around adult, pediatric or neonatal patients in order to warm them. The heated air is blown through a connecting flexible hose to a disposable blanket to provide patient therapy by the means of warmed air. The system can be used in long-term care facilities, surgical facilities, hospitals including the Post-anesthesia Care Unit (PACU), Intensive Care Unit (ICU), Surgical Intensive Care Unit (SICU), Emergency Room (ER), Operating Room (OR), medical and surgical floors, or any other department or hospital facility requiring patient temperature management. To prevent hypothermia and/or reduce cold discomfort before, during, and after surgical procedures. Used to raise a patient's temperature and/or maintain a desired patient temperature through convective heat transfer from the controller to a warm-air heated blanket. | WarmAir Model 135 Hyperthermia System and Blankets | 2 | September 01, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Cincinnati Sub-Zero is conducting a field correction due to changes that have been made to the WarmAir Model 135 devices Operation and Technical Manuals in order to comply with FDA device labeling regulations. | MISBRANDING: Labeling False and Misleading | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 27, 2011, to all affected user/customers. The letter includes: description of device; the problem with the suspect manuals, and provides contact information for the customers to use to request new versions of the manuals. Upon receipt of the new manuals, the customers are asked to immediately collect and destroy/discard all of the old manuals at their facility and to replace them with the new versions. After the manuals have been discarded and replaced, the customers are asked to complete and return the attached "Recall Response Form" via fax to (513) 772-9119 or scan and email to WA135manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, OH 45241.For additional information the customers are provided with an e-mail address: WA135-2011@cszinc.com as well as a toll-free telephone number and fax number to reach the firm and obtain additional information as necessary. The contact numbers are: 1 (800) 989-7373, or (513) 772-8810, and/or fax # (513) 772-9119.. |  | 6561 | 6561 units were distributed. | Nationwide distribution: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 05, 2013 | 673 | 6561 | Not_Computer
Z-3128-2011 | 59685 | enzyme immunoassay, inhibin-a | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA, Part Number: DSL-10-28100-4Product Usage: The Inhibin A ELISA kit is intended for the quantitative measurement of dimeric Inhibin A in human serum or plasma. It is intended strictly for in vitro use as an aid in the diagnosis and monitoring of various hormonal reproductive disorders. | Inhibin A ELISA | 3 | September 02, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the microbial contamination in the conjugate diluent bottles included in the Manual Inhibin A Elisa kit (Part Number DSL-10-28100-4), kit lot number 110761. | OTHER/UNDETERMINED: Pending | Beckman Coulter initiated an "URGENT PRODUCT CORRECTION" on June 15, 2011, via a telephone script to the only one customer impacted by the recalled product. The telephone script provided an explanation of the problem identified and the actions to be taken. Beckman Coulter recommended that the customer discontinue use of the recalled product as a preventative measure. Beckman will send a representative to the customer's facility to replace the recalled product in all remaining kits of the affected lot. The customer was instructed to share this information with laboratory staff. For any questions concerning this information, contact Beckman Coulter at 952-368-7874. | Contact the recalling firm for information | 342 | 342 units | (USA) Nationwide Distribution in the state of CA | Terminated | July 11, 2012 | 313 | 342 | Not_Computer
Z-3129-2011 | 59601 | enzyme linked immunoabsorbent assay, (chlamydiae group) | LJC | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | VIDAS Chlamydia test kitsVIDAS Chlamydia is an automated qualitative test for use on the VIDAS family instruments, for the detection of Chlamydia antigen in endocervical or urethral specimens, using the ELFA technique (Enzyme Linked Fluorescent Assay. | VIDAS Chlamydia kits | 2 | September 02, 2011 | 2011 | bioMerieux, Inc. | Complaints were received of high rates of false positive results, leading to unnecessary antibiotic therapy being prescribed to patients. | OTHER/UNDETERMINED: Pending | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" letter dated July 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue using and to destroy any of the affected product. The firm advises customers to confirm all positive results obtained with the affected lot number with another method prior to reporting results. An Acknowledgement and Product Replacement form was attached for customers to complete and return to the firm.Customers should contact their local bioMerieux Customer Service representative for questions regarding this notice. | Contact the recalling firm for information | 283 | 283 units | Nationwide Distribution- including CA, DC, FL, IL, IL, IN, MD, MI, MN, MO, MT, NC, NE, NJ, NY, PA, UT, VA and WI. | Terminated | October 23, 2012 | 417 | 283 | Not_Computer
Z-3130-2011 | 58613 | system, endovascular graft, aortic aneurysm treatment | MIH |  |  | Cardiovascular |  | Cook Medical Zenith Flex AAA Endovascular Graft Main Body, REF TFFB-24-82-ZT, REF G48407, 18 Fr ID/ 7.1 mm OD, Sterile | Zenith Flex AAA | 3 | September 06, 2011 | 2011 | Cook, Inc. | The firm initiated this action in response to information that there may be potential damage in the braided suture which is used to attach the external stents to the graft material. | TRAINING: Employee Error | The firm sent URGENT PRODUCT RECALL notices dated 4/15/2011 to Risk Managers/Recall Administration (Chief of Vascular Surgery and or Interventional Radiology). The letters explained the problem with the devices and the action to be taken. Consignees were instructed to return the recalled devices to Cook Medical for credit or replacement. Customers with questions or concerns can contact the firm at 800-457-2235. | Contact the recalling firm for information | 4 | 4 units | US, Canada, Japan | Terminated | September 13, 2011 | 7 | 71 | Not_Computer
Z-3159-2011 | 59659 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Link Hemoglobin A1c Program Reorder Pack.Calibrator/Diluent Set.IVD.Kit components:Elution Buffer A, 2 x 2500 mLElution Buffer B, 1 x 2100 mLWash/Diluent Solution, 1 x 2500 mLCartridge Set, 2 Cation exchange Analytical Cartridges, 4 Cation Guard Cartridges;CD-ROM -Calibration/Diluent Set, 2 levels, 2 vials each of lyophilized human red blood cell hemolysate with preservative, 1 bottle of diluent, 100 mL, deionized waterWhole Blood Primer, 6 x 1 mLSample Vials, 2 x 100Instruction Manual.Bio-Rad Laboratories Inc. Hercules, CA 94547Intended for the percent determination of hemoglobin A1c in human whole blood using ion-exchange high-performance liquid chromatography (H LPC). | VARIANT II TURBO Link Hemoglobin A1c Program | 3 | September 06, 2011 | 2011 | Bio-Rad Laboratories, Inc. | The Elution Buffer B, lot no. GA 10430, was shipped in a 2.0 L bottle instead of the standard 2.5 L bottle. | TRAINING: Employee Error | The firm, BIO-RAD, sent a "Medical Device Correction" letter dated July 28, 2011 to all consignees/customers via Fed Ex on August 8, 2011. The letter described the product, problem and actions to be taken. The customers were informed that they could continue to use and report results with the affected buffer lot and the stability studies will continue until lot expiration date. The consignees were instructed to distribute this letter to all customers that have received the affected lot and to distribute the MEDICAL DEVICE CORRECTION RESPONSE FORM only to those customers that request replacement Elution Buffer B. In addition, the customers were to complete and return the MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, Attention: BIO-Rad CSD Regulatory Affairs Department.If you have any questions, please contact your regional Bio-Rad office or 510-724-7000. |  | 40 | 40 | Nationwide distribution: MD, MI, MN, NY, OH, SC, VA and WI. | Terminated | November 01, 2011 | 56 | 40 | Not_Computer
Z-3120-2011 | 59692 | tray, start/stop (including contents), dialysis | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom dialysis prep kits labeled as follows: KIDNEYCARE, 125/CS, LOT #A1160, EXP 7-12, BIN 35A, Lee Medical International, Inc., New Orleans, LA 70123MESCALERO-HEMO, 200/CS, LOT #B1113, EXP 8-12, Lee Medical International, Inc., New Orleans, LA 70123DAV-MT-POCONO-H, 200/CS, LOT #B1123, EXP 8-12, BIN 45B, Lee Medical International, Inc., New Orleans, LA 70123ADV-CATH-ISL44, 100/CS, LOT #B1127 EXP 8-12, BIN 34B, Lee Medical International, Inc., New Orleans, LA 70123ADV-HEMO-15G, 100/CS, LOT #B1128, EXP 8-12, BIN 34A, Lee Medical International, Inc., New Orleans, LA 70123NORTHHUDSON-C, 100/CS, LOT #B1134, EXP 8-12, BIN 42B, Lee Medical International, Inc., New Orleans, LA 70123NORTHHUDSON-H, 75/CS, LOT #B1135, EXP 8012, BIN 42B, Lee Medical International, Inc., New Orleans, LA 70123JOHNHAYES-HEMO, SMALL GLOVES, LOT #B1143, B1144, and B1145. EXP 8-12, 100/CS, Lee Medical International, Inc., New Orleans, LA 70123JOHNHAYES-HEMO, MEDIUM GLOVES, LOT #B1146 and B1147, EXP 8-12, Lee Medical International, Inc., New Orleans, LA 70123JOHNHAYES-SUB, SMALL GLOVES, 100/CS, LOT #B1148, EXP 8-12, Lee Medical International, Inc., New Orleans, LA 70123JOHNHAYES-SUB, SMALL GLOVES, LOT #B1149, EXP 8-12, Lee Medical International, Inc., New Orleans, LA 70123SNCH-CATH, 50/CS, LOT #B1152, EXP 8-12, BIN 43A, Lee Medical International, Inc., New Orleans, LA 70123COMMUNITY-HEMO, 125/CS, LOT #B1156, EXP 8-12, Lee Medical International, Inc., New Orleans, LA 70123FLAMS-HEMO, 125/CS. LOT #B1162, EXP 8-12, BIN 38B, Lee Medical International, Inc., New Orleans, LA 70123KINDEYSPA-HEMO, 125/CS, LOT #B1161, EXP 8-12, BIN 35B, Lee Medical International, Inc., New Orleans, LA 70123DAV-POCONO-H, 500/CS, LOT #C1107, EXP 9-12, Lee Medical International, Inc., New Orleans, LA 70123DAV-POCONO-CATH, 75/CS, LOT #C1108, EXP 9-12, BIN 43B, Lee Medical International, Inc., New Orleans, LA 70123DAV-MT-POCONO-H, 200/CS, LOT #C1109, EXP 9-12, BIN 45B, Lee Medical International, Inc., New Orleans, LA 70123MESCALERO-HEMO, 200/CS, LOT #C1110, EXP 9-12, Lee Medical International, Inc., New Orleans, LA 70123KIDNEYTRMT-HEMO, 125/CS, LOT #C1112 and C1123, EXP 9-12, BIN 45A, Lee Medical International, Inc., New Orleans, LA 70123DAV-POCONO-H, 500/CS, LOT #C1126, EXP 9-12, Lee Medical International, Inc., New Orleans, LA 70123METHODIST-HEMO, 150/CS, LOT #C1139, EXP 9-12, Lee Medical International, Inc., New Orleans, LA 70123KINGMAN-HEMO, 150/CS, LOT #C1141, EXP 9-12, Lee Medical International, Inc., New Orleans, LA 70123DAV-PALMERTON-H, 200/CS, LOT #C1145, EXP 9-12, BIN 43B, Lee Medical International, Inc., New Orleans, LA 70123EMANUEL-CATH, 100/CS, LOT #C1146, EXP 9-12, BIN 34B, Lee Medical International, Inc., New Orleans, LA 70123Product Usage: custom dialysis prep kits | Custom dialysis prep kits | 1 | September 07, 2011 | 2011 | Lee Medical International Inc | Firm manufactured custom dialysis kits with Povidone Iodine (PVP) Prep Pads which were recalled by H & P industries. Those pads were subsequently recalled due to possible bacterial contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Lee Medical International Inc., notified their consignees by email and telephone on 03/22/2011. Consignees were instructed to pull product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers.For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 19625 | 19,625 units | Nationwide Distribution (USA) - including states of: AZ, CA, FL, GA, MD, NJ, NM, NY, PA, and TX. | Terminated | January 30, 2012 | 145 | 19625 | Not_Computer
Z-3160-2011 | 59737 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Picis ED PulseCheck-Software, transmission & storage, patient dataService Pack v5.0 L and MThis product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources, such as health care professionals, medical devices connected to the system, and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | Picis ED PulseCheckSoftware | 2 | September 07, 2011 | 2011 | Picis Inc. | Software does not display appropriate allergy interaction warning | DESIGN: Software Design | Picis issued notification on 08/01/11 to affected clients, explaining the potential for the error, with instructions to monitor documented patient allergies closely when prescribing medication orders in ED PulseCheck, using Muitum drug database content. Picis will track the progress on acknowledgement of notification letter receipt through the implementation of software upgrade, for all potentially affected Customers.For questions regarding this recall call 781-557-3000. | Contact the recalling firm for information | 62 | 62 | Nationwide Distribution | Terminated | June 27, 2013 | 659 | 62 | Software | Alarm/Message | Software update | Software Update
Z-3162-2011 | 58473 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Ratcheting Torque Limiting Handle, Catalog Number 52-1012, Orthofix Spinal Implants, Firebird Spinal Fixation System, Firebird Spinal Fixation System DDD Instrument Case.Product is intended for use with the Set Screw Driver (52-1061) and Adjustable Counter Torque Wrench (52-1265) to perform the final tightening of set screws in a spinal implant construct. | Ratcheting Torque Limiting Handle | 2 | September 07, 2011 | 2011 | Orthofix, Inc. | Upon reaching torque limit, there is a possibility the product may jam or exhibit a ratchet direction switch from forward to reverse. | DESIGN: Device Design | Orthofix, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify and remove the affected product from use. Customers were advised to use the combination of the Ratching R Handle and the Set Screw Driver if they experience problems with the Ratching Torque Limiting Handle. A Tracking and Verification form was attached for customers to complete and return to the Orthofix Inc. Regulatory Affairs Department via fax at 214-937-3322.Contact your Account Services Representative at 888-298-5700 for questions regarding the removal and return of the affected product. | Contact the recalling firm for information | 287 | 287 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia and Germany. | Terminated | March 23, 2012 | 198 | 287 | Not_Computer
Z-3161-2011 | 59305 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TiLite ZR and ZRA Series 2 are mechanical wheelchairs. Devices are sold under brand name TiLite.The TiLite ZR had a model catalog number of Z2FS1. The TiLite ZRA has a model catalog number of Z2FS2. Provide mobility for physically impaired individuals. | TiLite ZR Series 2 and ZRA Series 2 Wheelchairs | 2 | September 07, 2011 | 2011 | TiSport, Llc | The Pivot Stud Bolt comprising part of the backrest assembly of TiLite ZR Series 2 and ZRA Series 2 mechanical wheelchairs have shown a tendency to fail due to the head of the bolt being sheared off. | DESIGN: Component Design/Selection | An Urgent Medical Device Correction letter, Product Correction Kit , and the Installation Instructions were sent to domestic suppliers, international suppliers, Veterans Administration hospitals and international distributors by US Priority Mail and by FedEx three days service.The letter identified the affected product and the reason for the recall. Customers were instructed not to send the product to their end user until the correction was made or to identify their end user and notify the end user of the product correction. Replacement parts and instructions to correct the issue were enclosed with the letter. Customers were to destroy the removed parts once they had completed the correction. Customers were to also return the Medical Device Correct Report that was provided, by fax to 509-586-2413, or email to ProdCorrection@tilite.com.Customers can call Technical Support at 800-545-2266, ext. 239 with any questions about recall. |  | 1037 | 1,037 units | Worldwide Distribution | Terminated | July 27, 2012 | 324 | 1037 | Not_Computer
Z-3172-2011 | 59453 | kit, wound dressing | MCY | Elastic bandage. | General Hospital | General & Plastic Surgery | Enforcement Discretion | Wound Dressing Kit - Tier I, 10 MM Flat Drain; product code 0877061.Medela Healthcare, McHenry, IL 60051; A Wound Drain Kit containing a 10 mm Flat Drain, antimicrobial gauze pad, non-adherent wound contact layer, transparent dressing, tubing clamp, barrier film wipe, adhesive tape, sterile saline, and suction extension tubing; 15 individually bagged kits per case; Some of the kits were mislabeled as Invia Wound Therapy Tier III Wound Dressing Kit, 15 Fr Round Channel Drain, product 0877067, lot 370118 | Invia Wound Therapy Tier I Wound Dressing Kit, 10 MM Flat Drain | 2 | September 08, 2011 | 2011 | Medela Inc | Medela determined that some of the individual kits or bags of lot #370118 of "0877061 Wound Dressing Set, Tier I, 10 MM Flat Drain", were labeled incorrectly as "0877067 Wound Dressing Set, Tier III 15 Fr Round Channel Drain". The outside label on the case is correctly labeled as "0877061 Wound Dressing Set, Tier 1,10 MM Flat Drain". | PRODUCTION CONTROLS: Labeling Mix-Ups | Medela Inc. sent an "Urgent Device Recall" letter dated July 26, 2011 to all affected customers.The letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) not to use the affected lot and place them in quarantine 2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form included with this letter 4) complete and return the Product Inventory Form via fax to 800-995-7867 or e-mail suction@medela.com, even if they have no affected product on hand 5)if they have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. 6) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. If you have any questions about this recall please call (877) 735-1626 or e-mail at suction@medela.com. |  | 3000 | 3,000 kits | Nationwide Distribution including the states of California, Florida, Georgia, Illinois, Kansas, Kentucky, Massachusetts, Michigan, Mississippi, New Jersey, New Mexico, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee and Texas. | Terminated | February 16, 2012 | 161 | 3000 | Not_Computer
Z-3164-2011 | 59798 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DUET TRS 45 3.5mm Universal Straight Single Use Loading Unit Surgical Stapler with Implantable StaplesReference Number: DUET4535 The ENDO GIA Universal Staplers with DUET TRS Loading Unit have applications in abdominal, gynecologic, pediatric and thoracic surgery for resection, transection, and creation of anastomosis. They may be used for transection and resection of liver substance, hepatic vasculature and biliary structures. | DUET TRS Loading Units | 2 | September 08, 2011 | 2011 | Covidien LP | Sterility of the device is compromised due to a breach in the sterile packaging | PRODUCTION CONTROLS: Storage | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 15, 2011, to its customers via federal express. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (Note: Units from the affected lots may have been incorporated into Covidien BEST PRACTICE KITS. If you have purchased BEST PRACTICE KITS, please evaluate whether these kits contain units from the affected lots.); complete and return the DUET TRS Universal Loading Unit Recalled Product Return form via fax to (203) 822-6009; and return all inventory from the affected lots (note: All affected units must be returned) with the complete form to: Field Returns Department, located at 195 McDermott Road, North Haven, CT 06473. Questions or concerns, contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 4362 | 4362 units | Worldwide distribution: USA (nationwide) and countries of: Asia, Australia, Canada, Europe, Japan, Latin America, and Pacific. | Open |  |  | 457461 | Not_Computer
Z-3173-2011 | 59830 | [6 empty] | BardPort Titanium Implanted Port.Bard Access Systems.Assembled in Mexico.A totally implantable vascular access device designed to provide long term repeated access to the vascular system | BardPort Titanium Implanted Port. | 3 | September 08, 2011 | 2011 | Bard Access Systems | Implanted port has discrepant product information on package labeling. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Bard Access System, Inc. sent a "Label Discrepancy Notification dated August 30, 2011 to all affected customers.The letter included affected product and informed them of the labeling discrepancy. Customers could either use the device or return it to Bard for replacement. For information on this recall call (801) 595-5568. | Contact the recalling firm for information | 1375 | 1375 units | Worldwide Distribution: Nationwide Distribution including the states of; AL, AR, AZ, CA, CT, DC, FL, GA, HI, IL, KY, LA, MA, MD, MI, MO, NJ, NV, NY, OK, PA, RI, SC, TN, TX, UT, VA, WI, and WV; and the countries of Canada and Belgium. | Terminated | March 30, 2012 | 204 | 1375 | Not_Computer
Z-3175-2011 | 58425 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Kyphon brand KyphoPak Express Tray First Fracture, 15/2 with KyphX Osteo Introducer System; Catalog Number: KPE1003; Product is manufactured and distributed by Medtronic Spine LLC (formerly known as Kyphon Inc), Sunnyvale, CADesigned to reduce kyphosis that results from vertebral body collapse due to osteoporosis. | KyphoPak Express Tray First Fracture, 15/2 with KyphX Osteo Introducer System | 2 | September 09, 2011 | 2011 | Medtronic Spine LLC (formerly Kyphon Inc.) | Device was shipped past the labeled expiration date. | TRAINING: Employee Error | On 3/23/2011, the firm initiated the recall and sent Urgent Notice: Device Recall letters identifying the affected product and explaining the reason for the recall. The letter asks customers not to use the affected product with the lot number 0001653390. The firm will replace any affected product at no cost to the customer. A Medtronic Spine Sales representative will assist with the segregation and return of the affected product. They will also assist the customer with contacting Medtronic Spine Customer Service to return the affected product. Included with the letter was a reconciliation form with instructions regarding the affected product. Questions regarding the letter should be directed to Customer Service at 866-959-7466. |  | 44 | 44 | Nationwide Distribution | Terminated | September 09, 2011 | 0 | 44 | Not_Computer
Z-3182-2011 | 59677 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica Bond Ready-to-Use Primary Antibody Epithelial Membrane Antigen (GP1.4); an in vitro diagnostic immunohistochemistry primary antibody kit; Catalog No. PA0035; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. Epithelial Membrane Antigen (GP1.4) monoclonal antibody is intended to be used for the qualitative identification by light microscopy of human epithelial membrane antigen in formalin-fixed, paraffin-embedded tissue by immunohistochemical staining using the automated Bond-max system. The clinical interpretation of any staining or its absence should be complemented by morphological studies and proper controls and should be evaluated within the context of the patient's clinical history and other diagnostic tests by a qualified pathologist. | Leica Bond ReadytoUse Primary Antibody Epithelial Membrane Antigen (GP1.4) | 3 | September 09, 2011 | 2011 | Leica Microsystems, Inc. | There was an issue with the controlled temperature storage unit where the products were stored prior to shipment to the customers. | TRAINING: Employee Error | Leica Microsystems, Inc. notified all affected customers by telephone on July 26, 2011. The product, problem, and actions to be taken by the customers were discussed. Customers were instructed to discontinue use of the product and to return the product. All affected product has been returned to the firm. | No consumer action necessary | 3 | 3 units | Nationwide Distribution-in the states of Indiana and Kentucky. | Terminated | September 09, 2011 | 0 | 4 | Not_Computer
Z-3176-2011 | 59455 | accessory equipment, cardiopulmonary bypass | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns Level Sensor II Pads.Terumo Cardiovascular Systems Corporation.Used to attach level sensors to the sides of the venous reservoir. | Terumo Sarns Level Detect II Mount Pads | 2 | September 09, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Level Sensor II mounting pads may not stick to the venous reservoir and may detach either at case setup or during a case. | PRODUCTION CONTROLS: Error in Labeling | Terumo Cardiovascular Systems Corp. sent an "Urgent Medical Device Removal" notice dated August 10, 2011 to all affected customers.The notice instructs customers to discard all remaining level sensor pads that do not have an expiration date on the package, and notify Terumo of the number of affected pads via a response form. Terumo will replace recalled product.For questions on this recall call Terumo Customer Service at (800) 521-2818, 8 am-6 pm EST. | Contact the recalling firm for information | 1531440 | 1,531,440 devices | Worldwide Distribution -- Nationwide Distribution, and the island of Puerto Rico. | Terminated | June 20, 2012 | 285 | 1531440 | Not_Computer
Z-3183-2011 | 59421 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | LATITUDEÂŽ Patient Management System Communicator, Model 6482.The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | LATITUDE Patient Management System Communicator | 2 | September 09, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has identified two occurrences where the LATITUDE system did not process a data payload, which is the information uploaded to the LATITUDE system from a remote interrogation of the implanted device. These occurrences resulted in non-displayed payloads because the data is not posted to the website. If the payload is not displayed, the pertinent data will not be available on the | DESIGN: Device Design | Boston Scientific notified all affected customers by telephone on May 25, 2011. The customers were informed that there was an interrogation for one of their patients that was not able to be displayed on the website. Boston Scientific detected that a payload failure had occurred, but could not conclude what information was contained within the payload. Each patient has had previous and subsequent payloads that have processed and displayed on the LATITUDE website. |  | 2 | 2 | Nationwide Distribution-in the states of Florida and Texas. | Terminated | December 04, 2012 | 452 | 2 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-3177-2011 | 59759 | suture, nonabsorbable, synthetic, polyamide | GAR | Nonabsorbable polyamide surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Monomid Non-Absorbable Suture are blue, monofilament synthetic absorbable suture, with different sizes and different needle sizes.The products are labeled in parts.These sutures are intended for use in general soft tissue approximation and/or ligation including use in cardiovascular, ophthalmic, and neurological surgery. | Monomid NonAbsorbable Suture | 2 | September 09, 2011 | 2011 | C P Medical Inc. | Unapproved dye was used in manufacturing the non-absorbable blue nylon suture material. | PRODUCTION CONTROLS: Process Control | CP Medical sent a "RECALL NOTICE" letter dated June 29, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products in the human market and to return all their existing stock. A Recall Checklist was enclosed for customers to complete and return.Contact Customer Service at 1-800-950-2763 for a Return Good Authorization and shipping account number. The firm will pay for shipping cost and replace product at no cost to customers. | Contact the recalling firm for information | 776 | 776 dozens boxes | Worldwide Distribution-USA (nationwide) and the countries of Italy, Romania, South Africa, and Taiwan. | Terminated | May 24, 2012 | 258 | 776 | Not_Computer
Z-3174-2011 | 59564 | collector, ostomy | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Coloplast Corp.-MN 06093, Country of origin: DK ( SenSura 19011 Post-Op, ostomy bag with window, 100 mm , Sterilized using ethylene oxide; SenSura 19010 Post-Op ,ostomy bag with window, 70 mm , Sterilized using ethylene oxide; SenSura 19001 Drainable bag with soft tap 1 piece/Opaque with inspection windows, 10-76 mm Sterilized using ethylene oxide, Manufacturer: Colopast A/S 3050 Denmark.)Coloplast Corp.-MN 06093, Country of origin: HU , No. 128100 1400 (Assura Post-Op 12810 pouches with window single use only, Sterile EO, 1/2" -3 7/8"/13-100mm) ; No. 128 180 1400 (Assura 12818 Post-Op, pouches without window, single use, Sterile EO, 1/2" -2 3/4"/13-70mm , Cut-to fit.; No. 128080 1400 ( Assura Post Op, 12808, pouches with window, single use, Sterile EO, 2 3/4 / 70 mm , Cut- to- fit) , Manufacturer: Colopast A/S 3050 Denmark.These products are typically the first pouches applied in the OR following ostomy surgery. The one-piece drainable minimizes pressure on the abdomen, the soft drainable outlet accommodates the post-operative liquid output, and the inspection window allows a nurse to inspect and / or tend to the new stoma or wound. Thereafter, non-sterile bags are used for ostomy management both prior to and after the patient is discharged from the hospital. | SenSura and Assura Postop Pouches | 2 | September 09, 2011 | 2011 | Coloplast Manufacturing US, LLC | On July 19, 2011 Coloplast became aware of the failure or routine sterility testing for one lot of SenSuera and Assura Post-Op Ostomy Bags. Based on these results it was determined that all affected lots be recalled and quarantined. This affects lots manufactured and distributed from January 2010 until July 19, 2011. | PRODUCTION CONTROLS: Process Control | Coloplast sent a "RECALL OF STERILE SENSURA AND ASSURE POST-OP POUCHES" letter dated July 27, 2011 to all affected customers. An enclosed Field Safety Notice described the product, problem, and actions to be taken by the customers.Contact your Customer Service representative at 800-533-0464 for questions regarding this recall. | Contact the recalling firm for information | 4020 | 4020 | Nationwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, GA, LA, MA, MD, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, RI, TN, TX, VA, WA, and WI. | Terminated | June 05, 2012 | 270 | 4020 | Not_Computer
Z-3179-2011 | 59656 | evacuator, gastro-urology | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Fluid Aspiration Set; a sterile set of silicone tubing and connectors for use with the Olympus SP-2 Fluid Irrigation & Aspiration Pump; 10 sets per box; Manufactured for: Olympus America Inc., Two Corporate Center Drive, Melville, NY 11747-3157; catalog #72-00114-0The fluid aspiration set is used in conjunction with the Olympus SP-2 pump for fluid aspiration during urological endoscopic procedures or laparoscopic procedures. | Olympus Fluid Aspiration Set | 2 | September 09, 2011 | 2011 | Northgate Technologies, Inc. | The Red Tag was assembled to the wrong end of the silicone tubing. The Red Tag indicates how the fluid aspiration tubing set is loaded into the Olympus SP-2 Fluid Irrigation and Aspiration Pump. | TRAINING: Employee Error | The firm, Northgate Technologies, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated August 18, 2011 to the Olympus America customers who ordered the Aspiration Tubing Set, 72-00114-0. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall for subsequent return to Northgate Technologies; if they further distributed this product, they were to identify and notify their customers of the recall, and complete and return the enclosed Return Response Form via fax to (847) 608-9405 or mail to: Northgate Technologies Inc., 1591 Scottsdale Court, Elgin, IL 60123. If you have any questions, please call (224) 856-2222 or e-mail to: ckurek@northgate-tech.com. | Discard product or return to place of purchase for a full refund or exchange | 10680 | 10,680 sets | Nationwide distribution: USA including states of: CA, FL, ID, IL, MI, NY, PA, SD and WV. | Terminated | December 14, 2011 | 96 | 10680 | Not_Computer
Z-3185-2011 | 55319 | stimulator, vestibular acceleration, therapeutic | MHZ |  |  | Neurology | Contact ODE | The 91220 mCare 300 Vital Signs monitor has a 10.4 inch resistive touchscreen display.The Spacelabs Medical, mCare 300, Model 91220, Vital Signs Monitor is a compact monitoring system with a resistive touchscreen display. It is indicated for use in adult, pediatric and neonate patient populations wherever there is a need for the monitoring of ECG, respiration, invasive or noninvasive blood pressures, body temperature, functional arterial oxygen saturation, or expired or minimum inspire CO2. The product is a prescription device intended to be used by healthcare professionals in all areas of a healthcare facility. | Spacelabs Medical Vital Signs Monitor, Model 91220 | 2 | September 12, 2011 | 2011 | Spacelabs Healthcare, Incorporated | There have been reports that the Spacelabs Medical Vital Signs Monitor, Model 91220, will not power up with AC power and will no longer charge the batteries. | DESIGN: Component Design/Selection | SpaceLabs Healthcare sent an URGENT - MEDICAL DEVICE CORRECTION letter dated April 15, 2010, to all affected customers. International consignees were notified by e-mail on April 22, 2010, with the customer notification letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use the device until their monitors can be upgraded. Customers were asked to brief their staff about the possibility of the monitor failure to help prevent an adverse event from happening. |  | 614 | 614 units | Worldwide Distribution - USA (nationwide) and the countries of UNITED ARAB EMIRATES, ARGENTINA, AUSTRALIA, BELGIUM, BAHRAIN, BRAZIL, CANADA, CHINA, COLOMBIA, GERMANY, ECUADOR, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, JORDAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, MALAYSIA, OMAN, POLAND, PUERTO RICO, PARAGUAY, ROMANIA, SAUDI ARABIA, SINGAPORE, SYRIAN ARAB REPUBLIC, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and URUGUAY. | Terminated | September 12, 2011 | 0 | 614 | Not_Computer
Z-3187-2011 | 59764 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ENDOPATHÂŽ ETS Compact Flex (240 mm shaft length) Linear Cutters, Model Number(s): SCW45; CTS45; SCB45; CTS45NK; 6CB45 and SCG45. Mfr. By: Ethicon Endo-Surgery, LLC, Guaynabo, PR 00969The ETS Compact-Flex45 Articulating Linear Cutter is intended for transection, resection, and/or creation of anastomoses. The instruments have application in multiple open or minimally invasive general, gynecologic, urologic, thoracic, and pediatric surgical procedures. They can be used with staple line or tissue buttressing materials. | ENDOPATH ETS Compact Flex (240 mm shaft length) Linear Cutters | 2 | September 12, 2011 | 2011 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENDOPATHÂż ETS Compact Flex 45mm articulating linear cutters because the possibility exists that damage to the blister pack may have compromised the sterility of the device. | PRODUCTION CONTROLS: Packaging | Ethicon Endo-Surgery Inc sent an URGENT DEVICE RECALL letter dated August 11, 2011 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. Fill out the Business Reply Form and fax it to 1-866-853-2802 if they have any of the affected product. To ensure adequate identification of the return product at their facility customers should complete the Customer Purchase Order Field on the Business Reply Form. For any questions regarding this recall call 1-866-708-2330 and reference Event 2471. For clinical or product support call 1-800-USE-ENDO, option 6, or contact your local sales representative. |  | 27582 | 27,582 instruments | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, GREECE, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, POLAND, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, THAILAND and UNITED ARAB EMIRATES | Terminated | February 15, 2013 | 522 | 27582 | Not_Computer
Z-3186-2011 | 59477 | catheter, steerable | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic FlexCath, model 3FC12, Steerable Sheath. The FlexCath Steerable Sheath is a percutaneous catheter introducer fitted with a valve to allow for introduction, withdrawal, and swapping of catheters and wires while preventing air ingress and minimizing blood loss. A side-port with stopcock is integrated to allow continuous drip infusion, injection through the center lumen, flushing, aspiration, blood sampling and pressure monitoring. Rx only. Sterilized using ethylene oxide. Manufacturer: Medtronic CryoCath LP, Kirkland, Quebec H9H 5H3, Canada.Product Usage: The FlexCath Steerable Sheath is intended for percutaneous catheter introduction into the vasculature and into the chambers of the heart. The sheath deflection facilitates catheter positioning. | Medtronic FlexCath 3FC12 | 2 | September 12, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath would like to inform you of an observation related to a potential leak in the hemostatic valve of the Flex Cath 12 Steerable Sheath, Model 3FC12. There wave been reports of blood and saline solution leaking out of, and air ingress into the hemostatic valve. This includes reported instances of the introduction of air bubbles through the valve during aspiration of the side port |  | Medtronic Sales Reps began on July 15, 21011 hand delivering an "IMPORTANT MEDICAL DEVICE iNFORMATION" letter dated July 2011 to consignees. The letter identified the product, the problem, and the actions needed to be taken. The letter stated the rate of occurrence, described the potential clinical implications, and reinforced the importance of following the information found in the instructions for use as consignees continue to use the device. The letter also stated that Medtronic CryoCath is working urgently on a number of fronts to mitigate this observation and will continue to communicate with consignees as appropriate.For further information, please contact Medtronic VP of Quality at (763) 526-0576. | Contact the recalling firm for information | 6712 | 6712 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MI, MO, MN, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI. and the countries of: Australia, Austria, Belgium, Canada, Czech, Denmark, France, Germany, Greece, Hong Kong, Hungary, Ireland, Italy, Netherlands, Norway, Poland, Qatar, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | June 07, 2012 | 269 | 6712 | Not_Computer
Z-3188-2011 | 58478 | dispenser, cement | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Orthovita(R) Cortoss(TM) Delivery GunProduct Usage: Used for the application of Cortoss(TM) Bone Augmentation Material | Orthovita(R) Cortoss(TM) Delivery Gun | 2 | September 12, 2011 | 2011 | Orthovita, Inc. | Recalling firm has discovered through internal testing that the Cortoss(TM) Delivery Gun may abrade its packaging material under certain transportation and handling conditions. This abrasion may cause small holes or tears in the packaging material, which may not be immediately visible to the user. Although t the Cortoss(TM) Delivery Gun is provided in double pouches and there have been no report | PRODUCTION CONTROLS: Packaging | Orthovita sent an "URGENT FIELD SAFETY NOTICE" letter dated March 31, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to Identify, quarantine and discontinue use of the affected product. The letter included an acknowledgement form to be completed and returned as per instructions.. The letter states that the recalling firm will contact customers to arrange for the return or destruction of any affected product. A pre-paid shipping label is provided for product return. The letter also states that the recalling firm will provide replacement product at no cost. For any questions regarding this notice call 484-323-8860. | Contact the recalling firm for information | 3174 | 3174 | Product was distributed within the US to medical facilities in AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, and WI. Government accounts include Bethesda Naval Center, 8901 Rockville Pike, Bethesda, MD; Lackland Air Force Base, 2200 Berquist Drive, Suite 1, Lackland AFB, TX; Hunter Holmes McGuire VA Hospital 1201 Broad Rock Blvd., Richmond, VA; Madigan Army Medical Center, Madigan Army Medical Center, Tacoma, WA; Ralph H. Johnson VA Medical Center, 109 Bee Street, Charleston, SC; VA Medical Center, 1481 W. 10th St., Indianapolis, IN; VA Hospital of Palo Alto, 3801 Miranda Ave., Palo Alto, CA; VA Loma Linda Healthcare System, 11201 Benton Street, Loma Linda, CA; and VA Medical Center/NY Harbor Healthcare, NY, NY. | Terminated | September 25, 2012 | 379 | 3174 | Not_Computer
Z-3178-2011 | 59331 | cardiac ablation percutaneous catheter | LPB |  |  | Cardiovascular |  | STOCKERT 70 Radio Frequency Generator, Catalog Number 57001 Indicated for use with a Biosense Webster, Inc. diagnostic/ablation deflectable tip catheter to deliver radio frequency energy during cardiac ablation procedures.Stockert 70 in foreign countries is also known as: EP-Shuttle/Stockert J50/Stockert J70With software versions: 1.035/1.035J/1.035J70 to 1.037/1.037J70 | EPShuttle/Stockert J50/Stockert J70 | 2 | September 12, 2011 | 2011 | Biosense Webster, Inc. | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator, to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product. | DESIGN: Software Design (Manufacturing Process) | The recall communication was initiated on 7/8/11 and 7/11/11 with Biosense Webster forwarding A Customer Notification Letter, Acknowledgement Form and Software Release Note (Attachments 3, 4 and 1) via FedEx to all U.S. customers who purchased the Stockert 70 Radio Frequency Generator. Two letters and forms were sent to U.S. customers via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively.The Release Note provides tips for the avoidance of potential events when working with the EP-SH UTTLE, STOCKERT 70, STOCKERT J50 or STOCKERT J70 RF generators with SW versions 1.035 / 1.035J /1.035J70 to 1.037/1.037J70. |  | 3397 | 3,397 units total (1,217 in U.S.) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Terminated | October 12, 2012 | 396 | 3397 | Software | Physical Safety Hazards | Notification | Safety Notice/Insructions
Z-3200-2011 | 59175 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips GEMINI TF64 Diagnostic Imaging Systems, Model Number: 882471, 510(k) #K052640, Serial Numbers: 7502, 7504, 7506, 7514, 7083 and 7164. The recalled GEMINI TF64 system units equipped with software versions: 3.5.1.1 and 3.5.2.1 are subject to recall/software correction. The GEMINI TFI6, TF64 and TF Big Bore (BB) are all diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other organ imaging). Attenuation correction is accomplished by CT AC. The device also provides for list mode, dynamic, and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI TF64 Diagnostic Imaging System | 2 | September 13, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | DESIGN: Software Design | The firm , Philips Healthcare, issued a 'Customer Information Letter' dated August 9, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were not required to perform any action. A Philips' Field Service Engineer will visit each customer site and install the necessary software upgrade and Philips will track the progress of the software corrections as they take place.For customers in North America, if you need any further information or support concerning this update, please contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5 for "All Imaging Systems". Enter your site ID number. If you do not know your site ID, please hold and then select option 5 for "Nuclear Medicine" and finally select option 2 for "PET" support. In all other countries, the local Philips Healthcare office should be contacted at 440-483-3000. |  | 6 | 6 units | Worldwide distribution: USA including states of: FL, MA, NE, NC, OH, TX, WA and WV; and countries of: France, Germany, Italy, Kuwait, The Netherlands, Spain and Turkey. | Terminated | February 22, 2013 | 528 | 19 | Computer
Z-3191-2011 | 58985 | motor, drill, electric | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Sterile Bone Cutting Burrs model QD8-4OVD labeled in part:***QD8-4OVD***QTY: 1***4 mm x 8 mm Diamond Oval. Use with***QD8 Attachments.***Use by 2016-03***ANSPACH***Manufacturer***Made in USA***The Anspach Effort, Inc.***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel: (800) 327 6887 / + 1 561 627 1080***EC REP***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe, HP 11 1LA***United Kingdom***Tel: +441494 616126***Rx Only***LOT E173049563***.Cutting and Shaping bone including bones of the spine and cranium. | Anspach Single Use, Sterile Bone Cutting Burrs model QD84OVD | 2 | September 13, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Inc., Palm Beach Gardens, FL, is recalling their Anspach Single Use, Sterile Bone Cutting Burrs, catalog numbers, IMRI-S-8NS-M, M-6P, QD8-4OVD, S-5B-4, and SIL-7-9, lot #'s C323023789, C403026551, C473028567, C403026563, D153034013, D173034385, and C363024789, because the bar code which defines the expiration date may be incorrect. This could cause the bar code reader system to | TRAINING: Employee Error | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated May 27, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and to segregate any affected product found. A Customer Response Form was attached for customers to complete and return to the firm via fax at 1-800-327-6661.Contact Anspach Customer Support at 800-327-6887 for questions regarding this recall. | No consumer action necessary | 17 | 17 | Worldwide Distribution--USA (nationwide) including the states of Alabama, California, and Wisconsin and the countries of Columbia, France, Germany, Great Britain, Italy, Panama, Poland, Saudi Arabia, South Africa, Spain, and Switzerland. | Terminated | February 15, 2012 | 155 | 114 | Not_Computer
Z-3198-2011 | 59690 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101A: AQUILION 16; System, X-Ray, Tomography, ComputedX-Ray Tomography | TSX101A: AQUILION 16; System, XRay, Tomography, Computed | 2 | September 13, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on their TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed,TSX-301A: AQUILION One; System, X-Ray, Tomography, Computed,TSX-20IA: AQUlLlON LB; System, X-Ray, Tomography, Computed, and TSX-1O1A: AQUILION 16; System, X-Ray, Tomography, Computed due to several software issues that cause the scanners to function incorr | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 22, 2011, to all affected customers. Customers were instructed that a revised software to prevent occurrence of the problem would be installed on their system. Their Toshiba Service Representative would contact them for an appointment to install the software to their system when it became available.Customers were requested to take temporary measures until corrective measures were taken.1) In dual scanoscopy, do not press the short key between the completion of the first scanoscopy and the start of the second scanoscopy.2) In dual scanoscopy, do not click the (Scano Skip) button immediately before the completion of the first scanoscopy.3) When the SplitCase function is set to ON, do not select two or more orders simultaneously from the appointment list.4) When an image comment is to be preset by making a selection from the preset memo, do not select any image comment with more than 10 characters. Customers were asked to share the information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at the facility.Customers were asked to complete and return the attached form and fax it to 877-349-3054.For any questions call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 155 | 155 units total | Nationwide Distribution | Open |  |  | 155 | Computer
Z-3184-2011 | 59797 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, version 8.0, used to guide a clinician in the proper implant selection, oral placement and the proper selection of an oral drill.Imaging processing System | Straumann coDiagnostiX | 2 | September 13, 2011 | 2011 | Straumann Manufacturing Inc. | Software version 8.0 has a software anomaly, in which the length is not adjusted accordingly when it is changed after the surgical sleeve position planning step has been completed. | DESIGN: Software Design | Staumann USA notified accounts by an Urgent Field Safety Notice letter dated August 4, 2011 and delivered via Federal Express Delivery Service.The letter identified the affected product and explained the reason for the recall. Customers were asked to review the letter, take appropriate actions, including completing and faxing the enclosed. Customer Confirmation Form. Straumann released a corrected software upgrade (coDiagnostiX 8.0.2.3672), which is available for download when users go online and activate current software. If the user does not have internet access a replacement CD can be ordered from Technical Support at 866 - 531 -7365. Questions should be directed to Technical Support Monday through Friday from 8:00 am to 8:00 pm Eastern Time. | Contact the recalling firm for information | 272 | 272 | Nationwide Distribution | Terminated | May 23, 2012 | 253 | 272 | Software | Device Operation | Software update | Software Update
Z-3076-2011 | 59760 | industrial x-ray | RBS |  |  |  |  | Hamamatsu model L10951-35 Micro Focus x-ray generator.The x-ray generator is incorporated into a cabinet x-ray system that is used for a variety of studies. | Micro Focus XRay generator | 2 | September 13, 2011 | 2011 | Hamamatsu Corporation | It was discovered about a potential issue of X-ray leakage with our Model # LI1 0951-35 Micro Focus X-Ray generator. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Hamamatsu Corporation sent a letter dated July 19, 2011, to the manufacturer and their subcontractor informing them of the affected product and what they propose to do. They will repair the units both in the field and those returned to our plant by installing additional lead shielding in the units. The lead shielding will be fixed to the unit by soldering the addition to the existing lead shield (see attachment 2). We expect to complete the repair of all units in the field by October 1, 2011. All new units to be shipped from our plant will be repaired before shipment.For further questions, please call (908) 231-0960 X 2300. |  | 8 | 8 | Worldwide Distribution -- USA ( nationwide ) the countries of Aichi, India and Japan. | Open |  |  | 8 | Not_Computer
Z-3206-2011 | 57889 | oximeter, ear | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 (GE Healthcare) | DATEXOHMEDA TRUSAT PULSE OXIMETER | 2 | September 14, 2011 | 2011 | GE Healthcare, LLC | The plastic enclosure on the external power supply of the TruSat Pulse Oximeter is prone to breakage, and overheating. The issue may impact patient or user safety. Potential for system or components to start on fire or explode (general case). | DESIGN: Device Design | On 4/8/11 GE Healthcare sent consignees an "Urgent Medical Device Correction" letter dated March 30, 2011. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering & Home Healthcare Provider.The letter listed the Safety Issues, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 31846 | 31,846 | Worldwide distribution: USA, ZAMBIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UGANDA, TURKEY, TUNISIA, TRINIDAD & Tobago, THAILAND, UNITED REPUBLIC OF TANZANIA, PROVINCE OF CHINA TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SUDAN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, NEPAL, MOZAMBIQUE, MEXICO, MALTA, MALAYSIA, MACEDONIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GUADELOUPE, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, KOREA, EL SALVADOR, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE , CANADA, BRITISH INDIAN OCEAN TERRITORY, BRAZIL, BOLIVIA, BENNIN, BELGIUM, BANGLADESH, BAHAMAS, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Open |  |  | 31846 | Not_Computer
Z-3213-2011 | 58820 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Miller Disposable Blades (sizes 0 art no: 50676 and 00 part no 50647Disposable laryngoscope blades | Crystal Metal Miller Disposable Blades | 2 | September 14, 2011 | 2011 | Penlon, Ltd. (Fmly East Healthcare) | Disposable laryngoscope blades incorrectly labeled with wrong size. | MISBRANDING: Labeling False and Misleading | Penlon Ltd ( frmly East Healthcare ) sent an " Urgent Field Safety Notice" dated March 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to check their stock with expiry dates between August 2010 to July 2015, and ensure that the sizes written on the blades, inner plastic bags, and the carton all match.Please complete the reply form below and send to the address provided. If you have any stock where the sizes on the bags, blades, and carton do not match, quarantine the blades and contact Penlon using the information below.For further questions please call 44 1235 547093. |  | 5 | 5 cartons (20 blades each) | Worldwide Distribution -- USA (nationwide) including the states of CA, VA and WI., and the countries of Belgium, Ireland, France, Netherlands, Norway, Portugal, Switzerland, UK, Hong Kong, Greece, Bahrain and China. | Terminated | September 21, 2012 | 373 | 5 | Not_Computer
Z-3211-2011 | 59438 | vehicle, motorized 3-wheeled | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 10, 3 Wheel (SC610)Mobility medical device | Victory 10, 3 Wheel (SC610) | 2 | September 14, 2011 | 2011 | Pride Mobility Products Corp | Contract manufacturer error. Component part(s) (fuses) specifications were not adhered to as outlined in Pride's Engineering drawings. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pride Mobility Products Corp. sent a "RECALL OF VICTORY 3 AND 4 WHEEL UNITS; REPLACEMENT OF FUSES" letter dated June 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers are asked in the letter to locate their affected Victory unit(s) and inform users of the recall. A letter has been provided by the firm to distribute to users alerting them to the recall. Additionally, customers are to set up an appointment with users to replace the fuses. Customers are instructed to contact Pride's Upgrade department at 866-222-6212 to order the kit and for questions regarding this recall. . |  |  |  | Worldwide Distribution-USA (nationwide) including the District of Columbia and Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and countries of Australia, Canada, France, and United Kingdom. | Open |  |  |  | Not_Computer
Z-3202-2011 | 59182 | motor, drill, electric | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Cutting Burr labeled in part:"***S-1504TD QTY:1***Twist Drill for 1.5 mm x 4 mm Screw. Use with SHORT Attachments.***ANSPACH***Manufacturer Made in USA***The Anspach Effort, Inc***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel (800) 327 6887 / +1 561 627 1080***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe HP11 1LA***United Kingdom***Tel: +441494616126***Sterilized using Irradiation***Do Not Re-use***Keep Dry***Caution; consult accompanying documents***Sterile unless damaged or opened***Rx Only***". Cutting and shaping bone including bones of the spine and cranium. | Anspach Single Use, Sterile Bone Cutting Burs | 2 | September 14, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Palm Beach Gardens, FL, is recalling their Single Use, Sterile Bone Cutting Burrs, Catalog #, S-1504TD, Lot #s D363039140. D323038110, D293037349 and Catalog # TURQ-434-4, Lot #s D343038845, D323038131, due to the burr being 2.0 mm longer than stated on the label. As a result, a hole drilled using this burr would be 2 mm deeper than the screw. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Anspach Effort, Palm Beach Gardens, FL sent a Customer letter dated March 31, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to screen their inventory, contact Anspach Customer Support at (800) 327-6887 if any of the recalled product was found, and complete and return the attached reply form. If the product had been further distributed, they were asked that the recall information be forwarded as appropriate. |  | 599 | 599 | Worldwide Distribution -- USA (nationwide) including the states of CA, CO, IN, KY, MD and TX. and the countries of Germany, Japan, Taiwan, United Kingdom and Sweden. | Terminated | October 11, 2011 | 27 | 628 | Not_Computer
Z-3210-2011 | 59826 | massager, powered inflatable tube | IRP | Powered inflatable tube massager. | Physical Medicine | Physical Medicine | 510(k) | Game Ready Professional Therapy System,a cryotherapy device.Intended to treat postsurgical and acute injuries to reduce edema, swelling and pain where indicated. | Game Ready Professional Therapy System | 3 | September 14, 2011 | 2011 | CoolSystems, Inc. | It was discovered that some Game Ready Professional Therapy System units that were to have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. | TRAINING: Employee Error | The firm, CoolSystems, Inc., contacted customers by phone on August 23, 2011. The customers were informed that some Game Ready Professional Therapy System units have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. Instructions were given to immediately cease fulfilling orders from that inventory. Loaners were shipped to the customers overnight. The customers were instructed to replace the shipped device with the loaner that was sent and return it to the company ASAP. If you have any questions, contact the Vice President, Quality Assurance and Regulatory/Clinical Affairs at 510-868-5378. |  | 26 | 26 units | Nationwide distribution: CA, MA, MO, NH, OH, TX, UT, and WA. | Terminated | November 14, 2011 | 61 | 26 | Not_Computer
Z-3205-2011 | 59473 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva Radiological Image Processing System with software R.1.2 Philips Healthcare, Inc.A computed radiography image processing system. | Philips Healthcare PCR Eleva Radiological Image Processing System | 2 | September 14, 2011 | 2011 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be interpreted as a RIGHT marker for side indication if the image is mirrored again within a PACS system. This could potentially lead to a misdiagnosis | DESIGN: Software Design | Philips Healthcare Inc. sent out "Urgent --Medical Device Correction" notices dated June 10, 2011 to all affected customers.The letter described the product and problem. Customers were informed the "burned-in Mirror-Icon" may be misinterpreted as a Right marker. Customers are advised to preferably use lead markers during each X-ray acquisition or alternatively electronic annotations on each image for side indication. The presentation of these markers within each image will clearly show the difference between the intended Left/Right indication and the little mirror ICON. Philips will install improved software that prevents the misinterpretation of the symbol and will schedule the installation of the updated software.For questions on this recall call Phillips Healthcare Inc. at 978-687-1501. |  | 197 | 197 units | Worldwide Distribution: Nationwide distribution; and including the countries of Albania, Angola, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Croatia, Denmark, Egypt, France, Germany, Ghana, Greece, Greece, Haiti, Hungary, Indonesia, Iraq, Israel, Italy, Kazakhstan, Kenya, Latvia, Libya, Liechtenstein, Lithuania, Malawi, Malta, Martinique, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Portugal, Puerto Rico, Romania, Russia, Saudi-Arabia, Serbia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Tanzania, Thailand, Tschech. Rep, Turkey, Turkmenistan, UAE, Uganda, and United Kingdom. | Open |  |  | 197 | Software | Display/Image | Software update | Software Update
Z-3208-2011 | 59681 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, 510 (k) #K060937, Model #728306, Serial Numbers: 100019, 100023, 100087, 100103, and 100170. The recalled system units are equipped with operating software version; 3.2.0.19011.The Brilliance ICT and Brilliance ICT SP are Computed Tomo grapy X-ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance iCT CT XRay System | 2 | September 14, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. Philips Medical's Field Service Engineers are installing software upgrades to certain of their Brilliance iCT and Brilliance ICT SP CT X-ray system units in order to correct non-conformances evidenced in these systems due to software errors. | OTHER/UNDETERMINED: Pending | Philips Medical ( Cleveland) Inc. indicated in the Correction and Removal Report dated June 21, 2011, that a Philips Field Service Engineer will be contacting each site in order to schedule a time to implement the Field Change Order. After which time, the Field Service Engineer will visit each customer site to install the software upgrade and the firm will be tracking the progress of this corrective action.For further questions please call (440) 483-7000. |  | 5 | 5 units | Worldwide Distribution -- USA (nationwide) including the states of DC, IL, IN, OH, PA and VT and the countries of Belgium and Israel. | Terminated | February 22, 2013 | 527 | 8 | Software | Device Operation | Software update | Software Update
Z-3204-2011 | 58474 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Inc., Millennium and HD-120 MultiLeaf Collimators (MLC), a component of the Clinac, All Varian Clinac using the collimator cover P/N 1106900-xx, Model Numbers: H14, H18, H27, H29, HCX; part of the Trilogy Radiotherapy Delivery System. | Millennium and HD120 MultiLeaf Collimators (MLC), Clinac | 2 | September 14, 2011 | 2011 | Varian Medical Systems, Inc. | An improperly installed collimator cover may detach and fall from the machine, potentially striking the patient or nearby personnel. | DESIGN: Component Design/Selection | On 3/17/2011, Varian Medical Systems sent an Urgent Medical Device Correction Letter to all affected users, with a description of the problem and user corrective action steps. |  | 3791 | 3791 units | Worldwide distribution including USA, Australia, Argentina, Italy, Morocco, South Korea, Belgium, Canada, Netherlands, Denmark, Japan, Hong Kong, Malaysia, Hungary, New Zealand, Sweden, Taiwan, Switzerland, Romania, Brazil, Slovakia, India, Colombia, UK, Ireland, Spain, Singapore, Martinique, Reunion, Venezuela, Iceland, Portugal, Turkey, Israel, Bahrain, Macau, Chile, Saudi Arabia, Latvia, China, Finland, Kazakhstan, Thailand, France, Poland, and Germany. | Terminated | June 20, 2012 | 280 | 3791 | Not_Computer
Z-3194-2011 | 59586 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console v 8.6 with ARIA RT Chart v8.5, all of 4DITC v 8.3;Varian Medical Systems, Palo Alto, CADesigned to assist the operator of a radiation therapy device in providing the radiation therapy device from commencing irradiation when any parameter is out of conformance with the treatment plan. | 4D Integrated Treatment Console | 2 | September 14, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The 4D Integrated Treatment Console may erratically fail to load the planned physical wedge accessory without warning or notification when receiving an approved plan from the oncology information system. |  | Urgent Medical Device Correction/Field Safety Notification letters were sent to all US consignees by Certified Mail on July 21, 2011., as well as Varian Sales, Marketing and Service Organizations. Varian International Customer Support managers will notify international customers by appropriate methods. The letter informed customers of the issue and the corrective action steps. Letters request that all appropriate personnel working in the radiotherapy department be notified of the contents of the letter. A new release version is being created to correct the issue. All sites will be notified when it is available.Questions or concerns should be directed to 1-888-827-4265. Contact information for international customers is also provided as well as a list of all affected serial numbers. | Contact the recalling firm for information | 1125 | 1125 | Worldwide Distribution -- USA, China, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guadeloupe, Ireland, Italy, Latvia, Lithuania, Martinique, Netherlands, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Turkey, United Kingdom, Singapore, Algeria, Bangladesh, Belarus, Egypt, Guam, Hong Kong, India, Indonesia, Iraq, Israel, Kenya, Kuwait, Macau, Malaysia, Philippines, Qatar, Russia, South Africa, South Korea; Taiwan, Thailand, Tunisia, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Columbia, Dominican Republic, Ecuador, Mexico, Peru, Venezuela, New Zealand, and Australia. | Terminated | August 01, 2013 | 687 | 1125 | Not_Computer
Z-3214-2011 | 59528 | pump, blood, cardiopulmonary bypass, roller type | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns 8000 Cardioplegia Monitor, Catalog Number 16414; also known as PROG COMP BD AMPRO 420 CPG 8K, Catalog Number 816373The Sarns 8000 Cardioplegia Monitor is indicated to measure extracorporeal line pressure and signal when the pressure exceeds the operator-set limit. | 8K Cardioplegia Monitor | 2 | September 15, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Cardioplegia Monitor for the Sarns Modular Perfusion System 8000 may fail to track cardioplegia delivery volume. | DESIGN: Software Design (Manufacturing Process) | Terumo Cardiovascular Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 29, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter advises customers to make sure all users are aware of the notice. Customers are instructed to complete and return an attached Customer Response Form. Terumo will replace the software on all monitors of the affected product.Customers with questions or concerns should call Terumo CVS Customer Service 1-800-521-2818. | Contact the recalling firm for information | 85 | 85 | Worldwide Distribution--USA (nationwide) including the states of AL, AR, CA, CO, GA, IL, IN, MS, NC, OH, TN, VA, and WI and countries of Belgium, Columbia, Hong Kong, India, Singapore, and Taiwan. | Terminated | March 08, 2013 | 540 | 85 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3207-2011 | 59290 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Synthes T1 Matrix Locking Cap of the Synthes Matrix System, Part Number 04.632.000.Intended for posterior pedicle screw fixation, posterior hook fixation, or anterolateral fixation. | Synthes Matrix System T1 Matrix Locking Cap | 2 | September 15, 2011 | 2011 | Synthes USA (HQ), Inc. | Labeling Correction -- The locking cap was not able to be loosened, resulting in the breakage of the screwdriver or the rod to be cut and the assembly to be removed (screw, cap, and rod). This poses a potential risk of significant prolongation of OR time. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm initiated this action on March 11, 2011 by issuing an Urgent: Medical Device Labeling Correction letter to all consignees. The letter identified the affected product and stated the reason for the correction. It also instructed the consignee to review and attach the supplementary technique page to any MATRIX Spine System Technique Guides that they may have had in their facility. Customers are to complete the Verification Section of the letter and return it to the firm via mail, fax, or e-mail. If customers have questions, they can call 610-719-5063 or their Synthes Spine Sales Consultant. | Contact the recalling firm for information |  |  | Nationwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI. | Open |  |  |  | Not_Computer
Z-3221-2011 | 59519 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101: Aquilion 32/64; System, X-Ray, Tomography, Computed. | TSX101: Aquilion 32/64 | 2 | September 16, 2011 | 2011 | Toshiba American Medical Systems Inc | There is the potential that the current may not be controlled correctly in examinations using the ECG modulation function, which adjusts the x-ray output in synchronization with the ECG waveform. | DESIGN: Software Design | Toshiba America Medical Systems (TAMS) issued an Urgent: Medical Device Correction notification to consignees via USPS return receipt mail. The letter identified the affected product, the problem, and corrective action. The firm plans to install revised firmware to prevent the occurrence of the problem. A Toshiba Service Representative will contact customers to schedule an appointment for installation. The firm asked customers to stop using the Power Reduction option during ECG modulated scanning until the new firmware has been installed. Customers are to share the information in the letter with all users, reviewing radiologist, and the clinical engineering or Biomedical group at their facility. Customers are also to complete and return the attached reply form via fax or e-mail to TAMS as proof of notification. Questions pertaining to the letter should be directed to 800-421-1968 or customers' local Toshiba Representative at 800-521-1968. |  | 528 | 528 units | Nationwide Distribution. | Terminated | March 18, 2013 | 549 | 528 | Software | Output/Calculation | Firmware update | Software Update
Z-3217-2011 | 59359 | system, nucleic acid amplification test, dna, methicillin resistant staphylococcus aureus, direct specimen | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD MAX DNA Extraction Kit DNA-2 (Whole Blood) Catalog Number 437502. Kit contains 24/2D barcode sample preparation reagent tubes *** For use with the BD MAX System ***to extract DNA from Gram-negative bacteria and viruses, which may be present in clinical specimens. The specimen is mixed with BD MAX Sample Preparation Reagent and processed using the BD MAX System. The BD MAX DNA Extraction Kit is used with the BD MAX System to extract DNA from Gram-negative bacteria and viruses, which may be present in clinical specimens. Purified DNA obtained with the BD MAX DNA Extraction Kits may be analyzed using the BD MAX System or another commercially available system for nucleic acid amplification and detection. BD MAX DNA Extraction Kits have not been validated for use with any specific analytical test method. The BD MAX System uses a combination of lytic and extraction reagents to perform cell lysis, DNA extraction and removal of inhibitors. Following cell lysis, by a combination of heat and lytic enzymes, the released nucleic acid is captured by magnetic affinity beads. The beads, with the bound nucleic acids, are washed and the nucleic acids is eluted using -10 ul of release solution. DNA is neutralized for a final volume of -20 ul. Product Quantity Distributed (Int) 21 kits | BD MAX DNA Extraction Kit DNA2 (Whole Blood) | 2 | September 16, 2011 | 2011 | Becton Dickinson & Co. | Sample results maybe switched between time it was scanned with handheld barcode reader and time it was placed in the sample rack, associating sample results to wrong patient. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated May 2011, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customers. The letter advised customers to discontinue use of the Sample Preparation Reagent tubes in the referenced kits and discard any remaining products. Customers were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to (410) 316-4258. For further assistance regarding replacements contact BD Customer Service Dept. 1-800-675-0908. For all other inquiries, contact BD Technical Services Dept. 1-800-638-8663. | Contact the recalling firm for information | 9 | 9 kits | Worldwide Distribution-USA (nationwide) including the states of CA, KY, MD, MI, NC, NY, TX, and WI and the country Belgium. | Terminated | October 12, 2011 | 26 | 392 | Not_Computer
Z-3254-2011 | 59770 | clamp, circumcision | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision trays and packs listed as follows: (1) DeRoyal (R) ST. Circumcision Tray, REF 50-5613, 1 Per Pack, Rx Only STERILE EO, Distributed by: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA; Made in Dominican RepublicLot Numbers: 20854003, 22892644(2) DeRoyal (R) ST. Circumcision PACK 1.3 CLAMP CLAMPMEMORIAL HEALTH, MEMORIAL HOSP TAL WEST, REF 89-5699, 1 Per Pack, Rx Only STERILE EO, Distributed by: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA; Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USALot Number 20012949Circumcision Clamp | Circumcision trays and packs | 2 | September 16, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm distributed surgical kits which contained circumcision clamps that were subsequently recalled by Surgical Design. | OTHER/UNDETERMINED: Pending | DeRoyal Industries, Inc. sent a recall notification letter dated 05/05/2011. A second notice was sent on 07/26/2011 to affected customers who did not respond to the initial notice. Questions regarding this recall may contact the Corporate Attorney at (865) 362-1037. |  | 390 | 390 units | USA (nationwide) including the states of AR, AZ, FL, GA, LA, MI, NJ, NM, and OH. | Terminated | January 30, 2013 | 502 | 490 | Not_Computer
Z-3255-2011 | 59683 | cardiac ablation percutaneous catheter | LPB |  |  | Cardiovascular |  | 7 F Therapy Ablation Catheter, 4 mm Tip Thermistor, Model (Reorder #) 83724, Description: 1304-7-25-M-TH.St Jude Medical CompanyA sterile, single use catheter with one 4mm ablation electrode at the tip and three 2mm diagnostic electrodes. The catheter includes a temperature sensor in the tip electrode for temperature monitoring and the handle is equipped with a steering mechanism to detect the distal tip of the catheter; both uni-directional and bi-directional versions are available. The catheter is available with various distal curve configurations; the curve is indicated on the catheter label. | Therapy Ablation Catheter | 2 | September 16, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed that the Therapy Ablation Catheter, 4 mm Tip Thermistor Quadripolar, were packaged with the incorrect Instructions for Use (IFU). To ensure accuracy of device labeling, St. Jude Medical is requesting that the incorrect IFU be replaced with the enclose, correct version. | PRODUCTION CONTROLS: Process Control | St. Jude Medical sent an "important Medical Device Safety Information" letter dated February 22, 2011 to all affected customers.The letter includes: affected problem, reconciliation form and corrective action. Customers are asked to remove the incorrect Instructions for Use (IFU) and use the IFU provided in the letter, and to fill out and return the reconciliation form.For questions on this recall please contact the firm at (512) 329-1054 or (651) 756-2869. | Discard product or return to place of purchase for a full refund or exchange | 179 | 179 units total | Nationwide Distribution including the states of TN, MS, IN, MO, FL, WI, CA, and MT. | Terminated | April 03, 2012 | 200 | 179 | Computer
Z-3215-2011 | 59743 | biopsy needle kit | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific 25ga (0.52mm) Expect Endoscopic Ultrasound Aspiration Needle REF 5002 Use By 1999-12 LOT 13615651C1 STERILE EOThe Expect Needle is designed to sample targeted submucosal and extramural gastrointestinal lesions through the accessory channel of a curvilinear echoendoscope. | Expect Endoscopic Ultrasound Aspiration Needle | 3 | September 16, 2011 | 2011 | Boston Scientific Corporation | The firm has determined that a single lot of 25 ga Expect EUS/FNA devices were mislabeled. The carton label incorrectly described the devices as 22 ga. The inner pouch label incorrectly describes the device as 25 ga. | OTHER/UNDETERMINED: Pending | Boston Scientific Corporation sent a "URGENT MEDICAL DEVICE RECALL letter dated August 11, 2010, to all affected customers. The letter identified the product, the problem and the action needed to be taken by customer.Customers were instructed to immediately remove the affected lot of Recall products from your inventory ( whether in Labs, Central Supply, Shipping and Receiving or any other location). Also, segregate this product in a secure location for return to Boston Scientific. Complete and return the Reply Verification Tracking Form ( even if you do not have any product to return) following the directions on this page and Reply Verification Tracking Form. Fax the Form to (508) 683-5578. Further questions please call (508) 683-4678. |  | 30 | 30 units | Nationwide Distribution including the states of VT, FL, SC, NC, WA, MA, and MN. | Terminated | November 10, 2011 | 55 | 30 | Not_Computer
Z-3222-2011 | 59704 | cpb check valve, retrograde flow, in-line | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Custom Procedure Kits, REV OAPERFUSION W/4:1 CARDIOPLEGIAItem Number 72990Terumo Cardiovascular Systems Corp.The Cardiovascular Procedure Kits are indicated for use only in the extracorporeal circuit for the cardiopulmonary bypass procedure(s) for which the user designed it. The 3-T L. V. Control Valve is a one-way vent valve used during cardiopulmonary bypass surgery to help prevent the buildup of excess vacuum pressure when suctioning either from the heart or the surgical field, and to help prevent inadvertent retrograde now and air into the heart. | Terumo 3T L. V. Control Valve | 2 | September 16, 2011 | 2011 | Terumo Cardiovascular Systems Corp | 3-T Medical Systems one way vent valves, included in the Custom Terumo Cardiovascular kits, may separate at the bond that holds the two halves | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo Cardiovascular Systems initiated contact on August 12, 2011 by phone, fax, email and Urgent: Medical Device Recall mail to all affected customers.The firm advised customers of the affected product, problem and required recall actions. Customers can remove whole customized packs that are not in urgent medical need and return packs. Or the User may opt to use a sterile LH 130 vent valve from Terumo Cardiovascular Systems. A response form was included to return back to the firm.For questions on this recall contact Terumo Customer Service at (734) 525-9580. | Discard product or return to place of purchase for a full refund or exchange | 34 | 34 units | Nationwide Distribution, including the states of FL and MI. | Terminated | March 20, 2012 | 186 | 4233 | Not_Computer
Z-3259-2011 | 59794 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS Angiographic Catheter, Soft-Vu Cobra (2) BRAIDED, 5 F x 90 cm x .035 in., Catalog No./REF 10709105, STERILE, MANUFACTURED IN USA, AngioDynamics, Inc. 603 Queensbury Avenue, Queensbury, NY 12804Angiographic catheter | SoftVu Angiographic Catheter | 2 | September 16, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The lot of Soft-Vu Catheters was labeled 90 cm in length; however the product was cut to a 65 cm length. | TRAINING: Employee Error | The firm ANGIODYNAMICS, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2011 to the consignees via Certified Mail, Return Receipt Requested. The letter described the product, problem and action to be taken. The customers were instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. The letter also instructed the consignees to (1.) Identify and segregate the recalled lot that is in your possession, (2.) Complete the enclosed Soft-Vu Angiographic Catheter Recall Reply Form and fax it to the attention of the Soft-Vu Angiographic Catheter Recall Coordinator at 518-798-1360. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number XXXXXXXXXX on the outside of the shipping box and include a copy of the Soft-Vu Angiographic Catheter Recall Reply Form with your shipment. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. | Contact the recalling firm for information | 30 | Domestic - 25 units; Foreign - 5 units | Worldwide distribution: USA (nationwide) including states of: California, Tennessee, and Wisconsin, and country of: Spain. | Terminated | October 18, 2012 | 398 | 30 | Not_Computer
Z-3266-2011 | 59820 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS UNI*FUSE INFUSION SYSTEM WITH COOPER WIRE, Catalog No./REF 12401818, CATHETER SIZE 4F, CATHETER LENGTH 90 cm, INFUSION LENGTH 10 cm, OCCLUSION WIRE SIZE .035 in., STERILE, AngioDynamics Inc., 603 Queensbury Avenue, Queensbury, NY 12804 | UNIFUSE INFUSION SYSTEM | 2 | September 19, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of UniFuse Catheters were packaged with the wrong occluding ball guidewires. | PRODUCTION CONTROLS: Packaging Process Control | URGENT MEDICAL DEVICE RECALL LETTERS (dated 1/28/11) and Recall Reply Forms were sent to the consignees via Certified Mail, Return Receipt Requested. The letters identified the scope of the recall, the reason for the recall, the status, and action to be taken. Customers were to identify and segregate the recalled lot in their possession. They were also to complete and fax back the Reply Form. The recalled product should be returned along with a copy of the Reply Form. Replacement product would be shipped upon receipt and confirmation of the returned product. Questions or concerns should be directed towards customers' local sales representative or call Julie Blair, Customer Service Manager at 1-800-772-6446. | Contact the recalling firm for information | 23 | 23 units | Worldwide Distribution -- USA and The Netherlands. | Terminated | October 18, 2012 | 395 | 47 | Not_Computer
Z-3268-2011 | 59467 | glucose dehydrogenase, glucose | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Contour, Contour TS Blood Glucose Test Strips and Bonus PacksIntended for self-testing by people with diabetes and healthcare professionals to monitor glucose concentrations in whole blood | Contour, Contour TS Blood Glucose Test Strips and Bonus Packs | 2 | September 19, 2011 | 2011 | Bayer Healthcare, LLC | All 10- or 25- count glucose test strip vials, either loose vials or found in patient starter kits of Contour, Contour TS and Contour" USB blood glucose meters are being recalled a packaging issue which, under certain conditions, can cause some test strips packaged in some small-count vials to operate outside of US FDA-cleared performance specifications and produce a low-bias reading during blood | PRODUCTION CONTROLS: Packaging | Bayer sent an "URGENT DIABETES TEST STRIP RECALL" letter dated August 12, 2011, to its customers with the name of the test strips, sizes recalled, reason for the recall, potential injury, and instructions to discontinue use and return the test strips to Bayer. The recall letter also included a recall response form that was to be faxed back to Bayer at 1-800-876-2243. For questions call Customer Logistics at 1-800-348-2637. | Contact the recalling firm for information | 58172747 | 3,617,287 vials/ 54,555,460 test strips | Nationwide Distribution and Puerto Rico | Terminated | July 02, 2012 | 287 | 3617287 | Not_Computer
Z-3263-2011 | 59726 | products, red-cell lysing products | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER DxH Cell Lyse, Part number 628018(1L) and 628019 (5L).The COULTERÂŽ DxH Cell Lyse is an erythrocyte lytic agent used for quantitativedetermination of hemoglobin, enumeration of NRBC, and counting and sizingleukocytes on UniCelÂŽ DxH 800 CoulterÂŽ Cellular Analysis System. | COULTER DxH Cell Lyse | 2 | September 19, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified unexpected high particle counts in the lots of DxH Cell Lyse listed. These lots may generate high background counts;could result in a falsely elevated Total Nucleated Count (TNC) of up to 50 cells/ÂżL; or may also result in recovery of TNC values that were higher than expected for Level 1 of the Body Fluid Control. |  | The recall communication was initiated on 08/01/2011 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US customers) to all customers, who purchased Coulter DxH Cell Lyse. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact Beckman Coulter for product replacement. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their Beckman Coulter representative. | Contact the recalling firm for information | 4913 | 4913 units total (2245 in US) | Worldwide Distribution -- Australia, Belgium, Brunei Darussalam, Canada, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Macao, Malaysia, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | June 04, 2012 | 259 | 4913 | Not_Computer
Z-3261-2011 | 59831 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS PROFILER BALLOON DILATATION CATHETER, Catalog No./REF 16801101, BALLOON DIAMETER 4MM, BALLOON LENGTH 2 CM, CATHETER LENGTH 40 CM, STERILE, MANUFACTURED IN USA AngioDynamics 603 Queensbury Avenue, Queensbury, NY 12804 Cardiovascular catheterization | Profiler Balloon Catheter | 2 | September 19, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of Profiler PTA Balloon Catheters were labeled incorrectly. Lot 532109 contains product that is actually 5 mm x 4 cm x 40 cm and Lot 532794 contains product that is actually 4 mm x 2 cm x 40 cm. | TRAINING: Employee Error | Angiodynamics, Inc sent a "Urgent Medical Device Recall letter" dated February 25, 2011 to all affected customers via Certified Mail Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following steps1. Identify and segregate the recalled lot that is in your possession.2. Complete the enclosed Profiler PTA Balloon Catheter Recall Reply Form and fax it to the attention of the Profiler PTA Balloon Catheter Recall Coordinator at 518-798-1360.The form lists the catalog numbers, lot numbers, and quantity our records indicate your facility has received.For further questions please call 1-800-772-6446. | Contact the recalling firm for information | 15 | 15 units | Nationwide Distribution including the states of CA, GA, ID, MI, OH and PA. | Terminated | October 18, 2012 | 395 | 29 | Not_Computer
Z-3265-2011 | 59303 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Powerheart Automated External Defibrillator (AED), G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation, if the victim is breathing, the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs, the device will charge automatically and be able to deliver therapy.Model numbers: Powerheart 9300A, 9300C, 9300D, 9300E, 9300P, 9390A, 9390E, CardioVive 92532, 92533, NK 9200G, 9231, and Responder 2019198, 2023440.Brand names: Powerheart AED, G3, CardioVive AED, GE ResponderAED, Nihon Kohden CardioLife AED. | Powerheart AED, G3, CardioVive AED, GE Responder AED, Nihon Kohden CardioLife AED | 2 | September 19, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) has identified two errors, 0x48 and 0x99, that have repeatedly occurred during a device self-test in the G3 installed base over the last 6 years. These self test errors cause the device to enter into a service required state.CSC remind owners/operators of the Powerheart AED devices to verify the "Rescue Readiness" of their equipment and follow maintenance pro | OTHER/UNDETERMINED: Pending | Cardiac Science began mailing out the recall notification letter on August 1, 2011, to all their US consignees via USPS first class mail. Cardiac Science's distribution partners will receive a notification letter and be provided with a template letter to send to their customers directly.CSC remind owners/operators of the Powerheart AED devices to check the status of of the AED every day to ensure it is Rescue Ready, and perform the regularly schedule maintenance as outlined in the Powerheart AED "Operator and Service Manual" provided with every AED.Customers can call Cardiac Science at 800-426-0337 or e-mail at aed205@cardiacscience.com. |  | 195739 | 195,739 units in US | Worldwide Distribution - US (nationwide)This action applies only to domestic | Terminated | February 13, 2014 | 878 | 195739 | Not_Computer
Z-3270-2011 | 59865 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | VASOVIEW 6 Endoscopic Vessel Harvesting SystemManufactured by Maquet Cardiovascular, LLC; 170 Baytech Drive, San Jose, CA 95134.Product Usage: Endoscopic vessel harvesting systems are indicated for use in minimally invasive surgery allowing access for vessel harvesting and are primarily indicated for patients undergoing endoscopic surgery for arterial bypass. | VASOVIEW 6 Endoscopic Vessel Harvesting System | 2 | September 20, 2011 | 2011 | Maquet Cardiovascular, LLC | During routine production quality monitoring process, it was found that UV curing adhesive was not used in a limited number of devices. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | MAQUET Cardiovascular, LLC, sent an "Urgent Device Removal/Immediate Action Required" letter dated December 17, 2009, to all affected customers. The letter identified the recalled products, the problem, and the actions needed to be taken. The letters asked customers to examine their stocks (inventory) immediately to determine if they had any affected lot numbers and to discontinue dispensing (distributing) the lot. The letter states that customers are to complete the attached Field Action Response form regardless of whether or not the affected products are located in their inventory. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. The letter instructed all users to return all unused affected products to MAQUET. For customer's convenience, all affected products that are returned may be exchanged with other unaffected VASOVIEW Endoscopic Vessel Harvesting System by overnight delivery. Questions should be directed to MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874. | Contact the recalling firm for information | 141 | 141 units | Nationwide Distribution (USA) including the states of: AL, AZ, CA, CT, FL, KS, KY, IL, IN, MI, MO, ND, NY, NV, OH, OK, OR, PA, SC, TN, TX, UT, WA, and WI. | Terminated | September 21, 2011 | 1 | 390 | Not_Computer
Z-3269-2011 | 59664 | autonomous extracorporeal blood leak detector/alarm | ODX | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSetÂŽ True Flow" 2008K@home BloodlinesFresenius Medical Care.Used for transferring blood during the hemodialysis treatment from the patient to the dialyzer (artificial kidney), removing extra water and waste products and then returning the filtered blood back to the patient | Fresenius Medical CombiSet True Flow" | 2 | September 20, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Blood lines coiling tabs #1 and #5 on the arterial line are not in the correct positions | DESIGN: Device Design | Fresenius Medical Care sent out an "Urgent" recall notice dated August 11, 2011 to all affected customers .The letter included description of the product name, problem and asked them to examine their stock immediately and quarantine bloodlines from the recalled lots. Customers were asked to return the product to Fresenius.For information on this recall call the Fresenius Customer Service Team at 1-800-323-5188. |  | 118 | 118 cases | Nationwide Distribution. | Open |  |  | 118 | Not_Computer
Z-3273-2011 | 57865 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Synthes Bone Marrow Aspiration System, Sterile, 11 gauge, 15cm needle (without side holes)Indicated for aspirating bone marrow | Synthes Bone Marrow Aspiration System, Sterile, 11 gauge, 15cm needle (without side holes) | 2 | September 21, 2011 | 2011 | Synthes USA (HQ), Inc. | Firm became aware that some systems were released to stock without going through the sterilization process. |  | Synthes (USA) sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. Affected product removed from inventory will be replaced by Synthes. Contact the Synthes Spine Sale Consultant at 1-800-620-7025 ext.6883 for questions concerning this recall. | Contact the recalling firm for information | 96 | 96 parts | Nationwide Distribution-including the states of AR, CO, FL, GA, ID, LA, MA, MD, MO, MT, NY, NC, OH, SC, SD, TN, TX, and VA. | Open |  |  | 96 | Not_Computer
Z-3274-2011 | 59791 | tourniquet, pneumatic | KCY | Pneumatic tourniquet. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker SmartPump Extension Tubing (Pneumatic tourniquet).Sterile tubing extension 40" and 60", set, luer lock. Rx only.Stryker Instruments Div. of Stryker Corporation.Optional accessories used with the Stryker SmartPump Tourniquet System. The luer locks are intended to be connected to the tourniquet pump - to fill the line - and to the cuff. The extension tubing is an optional accessory used to connect the tourniquet cuff with the pump when a longer distance between the pump and the tourniquet is desired. The extension tubing comes packaged 10 per box. | Stryker SmartPump Extension Tubing | 3 | September 21, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | Three complaints have been reported indicating that the product inside the package did not match the label description. The units had quick connectors instead of luer lock connectors as stated in the label description. |  | Stryker Instruments sent out "Urgent Medical Device Recall" notices dated August 22, 2011 to all affected customers.The letter included the name of the recalled product, lot numbers, and the reason for the recall. Customers were instructed to send back any product to Stryker that is within the affected lots. The letter further stated Stryker will quarantine the product and replace any product that is returned. For information on this recall contact Stryker Instruments at (269) 389-4534. |  | 11848 | 5920-000-003 (30 bx) and 5920-000-005 (30 bx) | Worldwide Distribution: Nationwide Distribution including the states of HAWAII, NORTH DAKOTA, NEBRASKA, PENNSYLVANIA, and IOWA, and the country of FRANCE | Terminated | October 12, 2011 | 21 | 60 | Not_Computer
Z-3277-2011 | 53897 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager Software, transmission & storage, patient data.This product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources, such as health care professionals, medical devices connected to the system, and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | CareSuite" Critical Care Manager, PACU Manager and Anesthesia Manager | 2 | September 23, 2011 | 2011 | Picis Inc. | Electronic health record software applications malfunctioned in a manner which failed to display the discontinued status for medication orders, as entered by the attending physician | DESIGN: Software Design | Picis Inc. sent a customer notification letter dated August 11, 2009 to all affected customers. The letter describes the product, problem, and action to be taken. Customers informed of the potential patient safety error were advised to make the appropriate configuration change, within current software versions 7.1.4 or later. Contact your local Picis support representative for questions regarding this notice. |  | 8576 | 8576 UNITS | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Brussels, Belgium, Denmark, Finland, France, Germany, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and UK. | Terminated | September 27, 2011 | 4 | 8576 | Software | Display/Image | Instructions on changing configurations | Safety Notice/Insructions
Z-3275-2011 | 59908 | microscope, surgical | EPT | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | OPMIÂŽ PenteroÂŽ, OPMIÂŽ PenteroÂŽ C, OPMIÂŽ PenteroÂŽ 900,OPMIÂŽ Vario/S8, OPMIÂŽ Vario/S88, OPMIÂŽ Neuro Multivision/NC4, OPMIÂŽ Vario/NC33, OPMIÂŽ Neuro/NC4, OPMIÂŽ Neuro/NC4 Ceiling Mount.A surgical microscope and accessories is an AC powered device intended for use during surgery to provide a magnified view of the surgical area. | Zeiss Surgical Microscopes,OPMI Pentero, OPMI Vario,OPMI Neuro | 2 | September 23, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Some third party surgical drapes have optically active protective lenses that could affect the ZEISS surgical microscope position information that is passed on to third party navigation systems. |  | The firm, Zeiss, sent an "IMPORTANT SAFETY ALERT FOR ZEISS SURGICAL MICROSCOPES" letter dated September 7, 2011, to all affected customers. The letter described the product, problem and action to be taken. The customers were instructed to use surgical drapes compatible with ZEISS surgical microscopes and always follow the instructions for use of the third party navigation system manufacturer for calibration, calibration verification and operation. In addition, the firm supplied customers with a supplement to their Zeiss surgical microscope user documentation and requested that they read the information and keep a copy with their original instructions also the customers were instructed to complete and return the Acknowledgement of Receipt form within 5 business days via fax to 1-866-809-6044 or USPS to Stericycle, Inc., Attn: Event 2480, 2670 Executive Dr. Suite A, Indianapolis, IN 46241.Should you have any questions about this Safety Alert, please contact your Carl Zeiss Customer Care representative at 1-888-871-7120. | Contact the recalling firm for information | 1677 | 1677 consignees | Nationwide distribution: USA only. | Terminated | February 24, 2012 | 154 | 1677 | Not_Computer
Z-3276-2011 | 59734 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration.Product Usage: The LOGIX OE application is an order entry system intended for use with Baxters LOGIX CM software. The software program aids the pharmacist in creating and authorizing nutritional solution orders. | Logix Order Entry (OE) Software | 2 | September 23, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for an order entry error during use of the LOGIX Order Entry (OE) software. When adding ingredients to a TPN order, the user can accept the pre-populated "per units" of "Bag" or select other "per units". A dosing error may occur if the user does not select the prescribed "per units" for the added ingredient. | DESIGN: Software Design | Baxter Healthcare Corp. sent an "Urgent Device Correction" letters dated August 15, 2011 which included LOGIX OE screen illustrations of a representative neonate and adult patient, via first class mail to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed to follow safe compounding practice, which includes verifying that the ingredient per units entered match the TPN prescription prior to preparation. Baxter requests that customers communicate this issue to all personnel responsible for LOGIX order entry and order authorization at their facility. Customers are instructed to retain a copy of the letter at each computer on which LOGIX OE software is installed. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. | Contact the recalling firm for information | 95 | 95 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the country of Canada. | Open |  |  | 95 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3260-2011 | 59877 | tray, start/stop (including contents), dialysis | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Dialysis prep kits labeled as follows: ADV-CARE-ISL44, 100/CS, LOT #K1018, EXP 5-12, BIN 34B; COMMUNITY-CATH, 150/CS, LOT #K1052, EXP 5-12, BIN 36B; COLUMBUS-CATH, 50/CS, LOT #1080, EXP 05-12, BIN 39B; RENALSG-CATH, 75/CS, LOT #L1046, EXP 6-12, BIN 40B; SUNYDOWN-HEMO, 100/CS, LOT #l1015, EXP 6-12, BIN 28A; KIDNEYCARE 125/CS, LOT #l1070, EXP 6-12; METHODIST-CATH, 75/CS, LOT #L1072, EXP 6-12; METHODIST-HEMO, 150/CS, LOT #L1073, EXP 6-12; MERCYMED-CATH, 75/CS, LOT #L1077, EXP 6-12, BIN 39A and SANTACLA-CATH, 75/CS, LOT #A1118, EXP 7-12, BIN 44A.Product Usage: custom dialysis prep kits | Custom Dialysis prep kits | 1 | September 23, 2011 | 2011 | Lee Medical International Inc | The firm used alcohol triple swabstick and prep pads manufactured by H & P Industries in the manufacture of their custom dialysis kits. Those swabsticks and pads were subsequently recalled due to possible bacterial contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Lee Medical International Inc., notified their consignees by email and telephone on August 17, 2011. Consignees were instructed to remove the affected product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers. For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 4075 | 4,075 units | Nationwide Distribution (USA) including the states of: CA, FL, GA, IA, LA, MD, NY and TX. | Terminated | January 30, 2012 | 129 | 4075 | Not_Computer
Z-3258-2011 | 59825 | low energy direct current thermal ablation system | OAB | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ANGIODYNAMICS NANOKNIFE IRE GENERATOR, Catalog/REF No. 20300101, ANGIODYNAMICS, 603 Queensbury Ave., Queensbury, NY 12804Intended to remove tissue and control bleeding by the use of high- frequency electrical current (surgical ablation of soft tissue). | NANOKNIFE | 2 | September 23, 2011 | 2011 | Angiodynamics, Inc. | Some units of Nanoknife have the potential to deliver a shortened pulse width without indicating an error. The affected units contain a "GOMe" circuit board which may cause the generator to drift from the calibration points set at the time of manufacture. | DESIGN: Component Design/Selection | AngioDynamics, Inc. sent a "NOTIFICATION OF MEDICAL DEVICE CORRECTION" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter instructs customers to review the measurement graphs (displayed on generator) after each delivery of energy, and to contact AngioDynamics Customer Service at 1-800-772-6446 if shortened pulses are detected. Contact your local sales representative or call Customer Service for any additional concerns or questions regarding this notice. | Contact the recalling firm for information | 21 | 21 units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Germany, Israel, Italy, and Spain, | Open |  |  | 21 | Hardware | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-3278-2011 | 59862 | nitrophenylphosphate, alkaline phosphatase or isoenzymes | CJE | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Alkaline Phosphatase, List Number 7D55-21 & 7D55-31.Used for the quantitation of alkaline phosphatase in human serum or plasma. | Abbott Clinical Chemistry Alkaline Phosphatase Assay | 2 | September 26, 2011 | 2011 | Abbott Laboratories, Inc | The reagent has the potential to be contaminated with mold, the Cladosporium fungal species. |  | The firm decided to recall and notified consignees by a Product Recall - Immediate Action Required letter on 08/26/11. The letter identified the affected product, the reason for the recall, the patient impact, and necessary actions to be taken by the customer. Customers are to determine if they are using or have the affected product in their inventory. Customers are to discontinue use and destroy any remaining product of the specified lots in accordance with their facility's policies and procedures. Alternative lots of the product should be ordered and the enclosed Customer Reply form should be completed and returned. The letter also states that customers may want to decontaminate their systems by following the steps in Attachment A. If the product from the affected lots were further distributed, a copy of the letter should be provided to those customers as well. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. | Contact the recalling firm for information | 6269 | 6269 units | Worldwide Distribution | Terminated | August 17, 2012 | 326 | 6269 | Not_Computer
Z-3280-2011 | 58556 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Dual Blue-Release Adult Walkers - Wheeled, and Blue Release Junior Walkers with 5 inch Fixed Wheels, Model Numbers: 6240-5F and 6240-JR5F. The recalled product is packaged four (4) units per shipping box.Product Usage: The walkers are intended for people who need additional support to maintain balance or stability while walking. | wheeled adult and junior walkers | 2 | September 27, 2011 | 2011 | Invacare Corporation | The firm is conducting a recall/correction of their adult and junior wheeled walkers due to a component defect. The firm determined that the caster bolt on the wheeled extension was too short to adequately maintain the Nylex fastener. Consequently, over time, this condition incurs the potential for the caster to loosen and detach, causing the wheels to fall off the walker. This scenario would res | DESIGN: Device Design | The firm, Invacare Corporation notified their customers of this recall/corrective action via an 'URGENT RECALL INFORMATION' letter dated April 7, 2011. The letter was sent to the customers by registered mail or Federal Express. The letter informs the customer(s) of the component defect involved and the inherent possibility that over time, a walker wheel(s) could fall off resulting in instability of the product and potential for serious injury to the user of the device. The letter provides a list of the affected units sold to each facility on an enclosed tracking sheet. It also asks that the customer: 1) Examine their stock for the presence of any of the affected walkers and to quarantine any such products found; 2) Complete an attached recall response form indicating the customers receipt and understanding of the instructions provided in the letter and to fax the completed form to the telephone number listed on the bottom of the form. The letter goes on to state that once the faxed back form is received by Invacare, the company will send the customer replacement wheels kits for all affected walkers purchased; 3) the customer is asked upon receipt of the replacement wheel kits to replace the wheels of any units still in stock and discard all of the old wheel parts removed; 4) Lastly, the customer is asked to contact their sub-account customers and notify them of the recall. If the customer have any questions concerning the instructions, call the recall support line at 1-866-567-7899. |  | 12764 | 12,764 walkers were distributed. | Worldwide Distribution - USA (nationwide: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, including Puerto Rico and countries of Canada, Columbia, and Israel. | Terminated | July 02, 2013 | 644 | 12764 | Not_Computer
Z-3282-2011 | 59527 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTIONGE Medical Systems, SCS.A software module that provides supplemental information to those images extracted from CT and MR temporal datasets. | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTION | 2 | September 27, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that image orientation of reference series displayed in Functool Diffusion Tensor/FiberTrak protocol on Advantage Workstation may not match with the orientation annotations displayed in the viewport. This issue may impact patient safety. | DESIGN: Software Design | GE Healthcare sent "Urgent Medical Device Correction" letters dated August 18, 2011 to all affected customers.The letter included name of product affected and described problem. Recommended actions where provided to customers and they were notified a GE Healthcare representative will contact them for a software system upgrade.For question of this recall contact the GE Service Representative at (800) 437-1171. | Contact the recalling firm for information | 697 | 697 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, CT, DC, FL, ID, IL, KY, LA, MD, MI, MN, NJ, NY, OR, PA, PR, TN, TX, VA, WA, and WI; and the countries of VIETNAM, MEXICO, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, ST HELENA, SPAIN, SLOVENIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂŽDÂŽRATION, ROMANIA, REUNION, REPUBLIC OF CHINA, QATAR, PORTUGAL POLAND, PARAGUAY, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LUXEMBOURG, LIBYA, KUWAIT, KOREA, JAPAN, ITALY, ISRAEL, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, CZECH REPUBLIC, CROATIA, COSTA RICA, COLOMBIA, CHILE, CANADA, BELARUS, BRAZIL, BOSNIA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | August 09, 2012 | 317 | 697 | Software | Display/Image | Software update | Software Update
Z-3281-2011 | 59383 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #78231, Serial Number: 9592.The Brilliance 64 is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at difference angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance CT XRay System | 2 | September 27, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Component defect. Philips Medical decided to conduct a recall field /correction (rework) of the mounting key locking threaded insert of the Detector Module system (DMS) bolt on one (1) Brilliance 64 CT X-ray unit. The firm initiated this action following Philip's receipt of the customer's complaint reporting that the retaining inserts were pulling out of the rotor during the removal of the UDMS | OTHER/UNDETERMINED: Pending | Philips sent a Field Change Order (FCO) 72800519 on June 13, 2011 to the affected customer. A Field Service Engineer visited the customer site and installed a hardware upgrade to the affected Brilliance 64 device system (Serial Number: 9592). The correction action was completed on August 9, 2011.For questions regarding this recall call 1-440-483-7000. |  | 1 | 1 unit | Nationwide Distribution including PA | Terminated | March 15, 2013 | 535 | 1 | Not_Computer
Z-3279-2011 | 59503 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trabecular Metal Modular Acetabular System is a modular acetabular cup system consisting of a shell with a preassembled locking ring, a liner, and optional screws. The Shells are available in a variety of sizes and are designed to be used with Trilogy Acetabular System liners. The shell implants are used to replace the hip socket during total hip arthroplasty.The device is indicated for primary or revision surgery for rehabilitating hips damaged as a result on noninflammatory degenerative joint disease (NIDJD) or its composite diagnoses of osteoarthritis, avascular necrosis,protrusio acetabuli, traumatic arthritis, slipped capital epiphysis, fused hip, fracture of the pelvis and diastrophic variant. The device is intended for either cemented or noncemented use. | Zimmer Trabecular Metal Modular Acetabular System | 2 | September 27, 2011 | 2011 | Zimmer Inc. | Lot # 61132698 and Lot 001122467 (added 9/ 18/2012) may have a manufacturing non conformance that could prevent the devices from functioning as intended. Specifically, devices in the affected lot may have a locking ring groove depth that prevents the corresponding liner from snapping into the shell. The groove provides the locking ring room to displace when inserting a liner into the shell. Th | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 27, 2011 to all affected customers of lot 61132698 and 9/18/2012 to only affected customer of lot 001122467. The letters describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to stop using and to quarantine the affected product. An Inventory Return Certification Form was included with the letter for customers to complete and return to Zimmer, Inc. via fax at (574) 372-4265. Replacement product will be provided for all returned product. Contact Zimmer, Inc. at 1-800-613-6131 for questions regarding this recall. |  | 62255181 | 24 devices lot 61132689; 1 device lot 001122467 | Worldwide Distribution-USA (nationwide) including the states of CA, NC,NY, PA, WA, and WI and countries of Australia, Germany, The Netherlands, and Switzerland. | Terminated | August 29, 2013 | 702 | 25 | Not_Computer
Z-3283-2011 | 59686 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific, LATITUDEÂŽ Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St Paul, MN 55112-5798Boston Scientific, LATITUDEÂŽ RF Communicator & Software, Model 6482.Distributed by Boston Scientific, 4100 Hamline Avenue North, St Paul, MN 55112-5798LATITUDE Communicator REF 6468 (only distributed outside the US)The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | Boston Scientific, LATITUDE Patient Management System Communicator | 3 | September 28, 2011 | 2011 | Boston Scientific CRM Corp | A limited number of Model 6476, 6482, and 6468 (Not available in the US) LATITUDE Communicators are unable to authenticate, and thus cannot connect to the LATITUDE System. Customers are instead receiving Latitude Temporarily Unavailable (LTU) message. | DESIGN: Software Design (Manufacturing Process) | Boston Scientific CRM Corporation sent a " Dear Patient", letter dated August 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new communicator is being shipped separately to you with a corrected security certificate that will allow it to be activated and function normally.1. Please activate this new equipment at your convenience to begin home monitoring. Activation instructions are enclosed with the communicator.2. We also ask that you return the original communicator previously delivered to you:" Place the original equipment in the shipping box used for the new equipment." Pack newspaper or other filler around the equipment to prevent damage during shipping, and seal the box." Enter your address in the FROM section of the prepaid Federal Express label, and apply the return label to the outside of the box directly over the original label." Call Federal Express at 1.800.463.3339 to schedule a pick-up.We apologize for this inconvenience. If you have further questions, please call 1.866.484.3268. |  | 470 | 470 | Worldwide Distribution -- (USA) nationwide including the states of AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE,NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY., and the countries of United Kingdom, France | Terminated | November 06, 2012 | 405 | 470 | Other | Device Operation |  | Other
Z-3286-2011 | 59953 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list 11005.Hospira has received reports of incorrect seating of the regulator closer on the Plum A+ Infusion Pumps. If the regulator closer does not close the flow regulator actuator when the pump door is opened, and if the clinician has not engaged the clamp prior to opening the cassette door, unrestricted flow may occur. | Plum A Hyperbaric Single Channel Infusion Pumps | 2 | September 29, 2011 | 2011 | Hospira Inc. | Hospira has received reports of incorrect seating of the regulator closer on the Plum A+ Infusion Pumps. If the regulator closer does not close the flow regulator actuator when the pump door is opened, and if the clinician has not engaged the clamp prior to opening the cassette door, unrestricted flow may occur. |  | Hospira, Inc. sent an "URGENT DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers. The letter instructs customers on how to prevent occurrences of continuous recycling/rebooting during start-up. A reply form was attached to the letter for customers to complete and return via fax to 1-888-345-5358. Contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance or information regarding this notice. | Contact the recalling firm for information | 120 | 120 pumps | Worldwide Distribution-USA (nationwide) including Puerto Rico and the U. S. Virgin Islands and countries of Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open |  |  | 290217 | Not_Computer
Z-3288-2011 | 59312 | spirometer, monitoring (w/wo alarm) | BZK | Monitoring spirometer. | Anesthesiology | Anesthesiology | 510(k) | FloTrak Elite modulesFloTrak Elite P/N 1044971FloTrak Elite P/N 1051005 | Philips | 2 | September 30, 2011 | 2011 | Respironics Novametrix, LLC. | Pressure lines maybe occluded and display incorrect tidal volumes. | PRODUCTION CONTROLS: Manufacturing Material Removal | Respironics Novametrix, LLC on July 29, 2011 notified accounts via e-mail or tracked receipt, an Urgent - Medical Device Recall letter to notify customers of the issue and to instruct them on conditions for continued use. Monitors will need to be returned for functional performance testing at their facility. Customers can contact Respironics US Customer Service at 1-877-387-3311 for customers residing in the USA or +011 (203)697-6348 for International customers, if they have any questions or need additional information. |  | 41 | 41 units | Worldwide Distribution -- USA, Belgium, Canada, Chile, Columbia, Egypt, India, Italy, Spain, Saudi Arabia Thailand, GB, and Russia. | Terminated | May 23, 2012 | 236 | 295 | Not_Computer
Z-3289-2011 | 59724 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Reverse Humeral Fracture InserterREF 406993 Comprehensive Instrumentation Reverse Humeral Inserter QTY 1.Instrument that is attached to the trial prosthesis for placement in the prepared bone canal to check for correct sizing, version and humeral length. Also used to insert the humeral implant into the canal. | Biomet Reverse Humeral Inserter | 2 | September 30, 2011 | 2011 | Biomet, Inc. | Recall initiated in response to a Biomet Canada report, which noted that while testing the instruments, the pin used to secure the implant on the device did not engage on the effected lots and were distributed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 8/3/11 Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to locate the affected products, discontinue use, return the devices, and a fax a copy of the response form.For information or questions on this recall call Biomet at (574) 372-1570. | Contact the recalling firm for information | 6 | 6 units (internationally) | International Distribution Only: Canada, Australia, and Columbia. | Terminated | October 26, 2011 | 26 | 6 | Not_Computer
Z-0001-2012 | 57858 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Probe, Tactile, Left, ASM, Long, Part #962009S, Medtronic Navigation,Louisville, CO 80027.Trackable hand held instruments used in spine surgeries for preparation of the pedicle. These instruments are used in conjunction with the Medtronic Navigation StealthStation platform and various software applications. | Left Long Tactile Probe | 2 | October 03, 2011 | 2011 | Medtronic Navigation, Inc | Distal end of tactile probe may break off during use. | DESIGN: Labeling Design | Medtronic sent an "Urgent Product Safety Notice" to all affected customers dated Jan 20, 2011.The letter included affected product and updated instructions for use. On-site training was given for the use of the probes by Medtronic Representatives.For information on this recall call Medtronics at (720) 890-3409. | Contact the recalling firm for information | 6 | 6 probes | Worldwide Distribution: Nationwide distribution including the states of CO and WI; and the country of South Korea | Terminated | October 03, 2011 | 0 | 6 | Not_Computer
Z-0007-2012 | 59847 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES Safety Infusion System, Model 621-100ES, 621-200ES, 621-300ES, and 621-400ESB. Braun Medical Inc., 1601 Wallace Drive, Suite 150, Carrollton, TX 75006The Outlook ES pump is intended for use with B. Braun Medical Inc., Horizon Pump IV Sets to regulate the flow of primary and secondary fluids when positive pressure is required. | Outlook ES Safety Infusion System | 2 | October 03, 2011 | 2011 | B Braun Medical, Inc | The Outlook ES Dose Guard Care Area can inadvertently be exited when the pump enters the KVO (Keep Vein Open) state. The same condition can occur when a sequence of HOLD-HOLD key strokes are performed on the pump key panel. | DESIGN: Software Design | The firm, B. Braun Medical Inc., sent an "URGENT: FIELD CORRECTION" letter dated September 8, 2011 to its customers. The letter described the product, problem and actions to be taken. The firm informed the customers that a B. Braun representative will be contact ing them to schedule a field software upgrade of the affected pumps. The customers were instructed to complete and return the FIELD CORRECTION INFORMATION FORM via mail using the enclosed envelop and contact B. Braun Customer Service for shipping instructions at 1-800-627-7867 if product is to be returned.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Technical Product Support at 1-800-627-7867 (800-627-PUMP). | Contact the recalling firm for information | 11525 | 11,525 pumps | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL, GA, IA, IL, KS, KY, LA, MA, MI, MO, NC, NJ, NY, OH, PA, SC, TX, VA, WA, WI, and WV; and country of: Canada. | Open |  |  | 11525 | Software | Device Operation | Software update | Software Update
Z-0003-2012 | 59979 | implant, cochlear | MCM |  |  | Ear Nose & Throat |  | Cochlear Nucleus CI512 Cochlear Implant, REF Z209051, New with Nucleus 5, Sterile EO, Made in Australia, Cochlear Ltd 14 Mars Road, Lane Cove 2066, Australia.The cochlear implant is intended to restore a level of auditory sensation via the electrical stimulation of the auditory nerve. | Cochlear Nucleus CI512 Cochlear Implant | 2 | October 03, 2011 | 2011 | Cochlear Americas Inc. | Cochlear implant may shut down and cease to function. | DESIGN: Component Design/Selection | Cochlear Americas, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and quarantine the affected product. A recall response form was attached to the letter for customers to complete and return. Contact Cochlear Clinic Customer Service at 877-883-3101 for questions concerning this notice. | Contact the recalling firm for information | 60902 | 33,645 units (global); 25,516 (implanted); 1,741 remaining in U.S. Market for implantation. | Worldwide Distribution-USA (nationwide) and countries in Europe and Australia. | Terminated | July 10, 2012 | 281 | 33645 | Not_Computer
Z-0006-2012 | 59139 | stimulator, autonomic nerve, implanted for epilepsy | LYJ |  |  | Neurology |  | Model 104 VNS Therapy Demipulse Duo Generator Cyberonics, Inc. Houston, Texas, Model 104 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page.Product Usage: The VNS Therapy System is indicated for use as an adjunctive therapy in reducing the frequency of seizures in adults and adolescents over 12 years of age with partial onset seizures, which are refractory to antiepileptic medications and for the adjunctive long-term treatment of chronic or recurrent depression for patients 18 years of age or older who are experiencing a major depressive episode and have not had an adequate response to four or more adequate antidepressant treatments. | Cyberonics | 2 | October 04, 2011 | 2011 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator. | DESIGN: Software Design (Manufacturing Process) | The firm, Cyberonics Inc., initiated containment actions on March 17, 2011 to prevent additional affected devices from being shipped. The firm initiated contact with implanting facilities in the United States on May 6, 2011 and international facilities and distributors on May 10, 2011 to remove product from the field that has not yet been implanted. Additionally, a "Field Safety Alert" letter was sent to all affected medical professionals. The letter identifies the affected product, the issue, and the action(s) to be taken by the physicians. Consignees are instructed to complete and return an attached Effectiveness Check Form. If assistance is needed physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com. |  | 13 | 13 units | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, ID, IL, KS, OH, AR, MA, MI, MN, MO, NC, NH, NJ, NM, NY, OH, OR, PA, TN, TX, UT, WA, WI and WV. Products were also distributed to WASHINGTON DC and Puerto Rico and the countries of: Austria, Belgium, Finland, Germany, Ireland, Italy, Netherlands, Norway, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | December 04, 2012 | 427 | 105 | Software
Z-0008-2012 | 59213 | system, facet screw spinal device | MRW |  |  | Orthopedic | 510(k) | OSTEOMED SPINE Facet Fixation System PrimaLOK FF Threaded Implant Driver Shaft OsteoMed L.P. 3885 Arapaho Road, Addison, Texas 75001 | OSTEOMED PrimaLOK FF Threaded Implant Driver Shaft | 2 | October 04, 2011 | 2011 | Osteomed, Lp | Two of the Implant Driver Assembly tips were reported to break during surgery. | DESIGN: Device Design | OsteoMed LP sent an Urgent --Product Recall letter dated June 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately for affected inventory and promptly return any affected inventory to:OsteoMed LPAttan: Rebecca Ellis3885 Arapaho RoadAddison, Texas 75001If the customer further distributed any of the affected product customers were instructed to immediately contact their accounts and advise them of the recall situation and have them return any devices they may have to OsteoMed at the above address.Customers were asked to complete and return the enclosed response immediately in the enclosed, postage paid envelope even if they do not have the device. For any questions please call (800) 456-7779. |  | 20 | 20 units | Nationwide Distribution including PA, TX, CA, IL IA, FL, CT, and MO | Terminated | December 04, 2012 | 427 | 20 | Not_Computer
Z-0004-2012 | 58823 | pump, infusion, pca | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂŽ-Solis ambulatory infusion pumps, Models 2100, REF 21-2101-0200-51 and 2110, REF 21-2111-0100-51, Rx Only. Indicated for intravenous, intra-arterial, subcutaneous, intraperitoneal, in close proximity to nerves, into an intraoperative site (soft tissue, body cavity/ surgical wound site), epidural space, or subarachnoid space infusion. | CADDSolis ambulatory infusion pumps. | 2 | October 04, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of non-delivery/ significant under-delivery of fluids when the CADDÂż Medication Cassette Reservoir is not correctly latched to the pump.EXPANDED: On 7/14/11, Smith Medical expanded this recall to include all consignees with affected pumps. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smiths Medical sent a customer's bulletin notice dated April 11, 2011, only to the customers who reported this issue. Smiths Medical will work directly with the five Consignees who have reported complaints for this issue.The customer was instructed to read the attached Customer Information Bulletin. If you have additional questions on how to correctly latch a cassette to the pump, than please visit the following website: https://www.smiths-medical-education.com.Smiths Medical will also provide replacement pumps, as requested. This also will be considered an extension of Smiths Medical's investigation into this issue and to further validate the effectivity of the corrective and preventative actions in actual use.For further questions please call (651) 628-7280.EXPANDED:Consignees were sent a Smith Medical " Urgent Safety Alert" letter dated June, 30, 2011. The letter was addressed to Risk/ Safety Managers, Clinicians/ Biomedical Engineering, Nursing, Pharmacy, Pain Management and Anesthesia Professionals, Medical Equipment Rental Companies,Distributors and other users of these devices. The letter described the problem and the product. It provided "Advice on Action to be Taken by the User". This included to read the attached Customer information Bulletin and to complete and return the attached Confirmation Form. | No consumer action necessary | 4536 | 4536 (3792 USA, 744 OUS) | Worldwide Distribution -- (USA) nationwide including the states of AR, CA, MO, NY and WI., and the countries of BELGIUM, CANADA, SWITZERLAND, GERMANY, DENMARK, FRANCE, UNITED KINGDOM, IRELAND, ITALY, LUXEMBOURG, MALAYSIA, NETHERLANDS, NEW ZEALAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE. *** EXPANDED*** AL, AK, AZ, CO, DC, FL, GA, ID, IL, IN, KY, LA, MD,MA, JI, MN, MT, NE, NH, NJ, NM, NC, OH, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WY. UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA. | Terminated | July 19, 2012 | 289 | 4536 | Not_Computer
Z-0011-2012 | 59956 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list 11005 Product Usage: The Plum A+ Hyperbaric Infusion System is exclusively designed for hyperbaric therapy in monospace and multiplace chambers. | Plum A Hyperbaric Single Channel Infusion Pumps | 2 | October 04, 2011 | 2011 | Hospira Inc. | Hospira has received reports of continuous recycling and/or rebooting of Plum A+ infusion pumps when the backlight intensity and/or display contrast settings for the LED display has been adjusted from the original default setting. This condition could cause a delay in therapy if the clinician is unable to start the device due to the continuous recycling and/or rebooting. | DESIGN: Software Design | The firm, Hospira sent "URGENT DEVICE RECALL" letters dated September 16, 2011 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to Hospira via fax at 1-888-345-5358. The letter states that a software upgrade to address the issue is currently being developed. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance. *** Follow-up letters dated June 26, 2012 were sent to the accounts listing the device modules with software versions v13.2, v13.3, v13.40 and v13.5 that will require the software upgrade. ***. | No consumer action necessary | 120 | 120 pumps | Worldwide Distribution - USA (nationwide) including Puerto Rico and the US Virgin Islands, and countries of: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open |  |  | 290217 | Software
Z-0013-2012 | 59952 | collector, ostomy | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative /Surgical System (57 mm)Required for application to a patient immediately after ostomy surgery. | Natura Durahesive ConvaTec Moldable Technology PostOperative /Surgical System (57 mm) | 3 | October 05, 2011 | 2011 | Convatec Inc. | Post-operative kits 57 mm may contain 70 mm pouches. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, ConvaTec, sent a "MEDICAL DEVICE RECALL" letter dated June 22, 2011 via UPS 2nd day air to all customers and sales representatives that received product. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop distributing and quarantine the recalled lot; perform a count of recalled product currently in inventory; complete and return the Recall Response Form via fax to (908) 904-2120 or mail to 200 Headquarters Park Drive, Skillman, New Jersey 08558; return the Recall Response Form even if no recalled product is in inventory; call the toll free number 1-800-582-6514 for return instructions and RGA number-(recall product must be returned in a separate container from any other product returns); mark all return shipments "Recalled Product" and provide supplied RGA number, and forward this letter to customers, if they distributed the product.Should you have any questions, call 1-800-582-6514. | Contact the recalling firm for information | 283 | 283 units | Nationwide distribution: AL, AZ, CA, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX and UT. | Terminated | January 11, 2012 | 98 | 283 | Not_Computer
Z-0012-2012 | 59943 | band, elastic, orthodontic | ECI | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Biomet Microfixation outer packaging labeled in part:"REF: SP-2411***RUBBER BAND, LATEX FREE SMALL, HEAVY, 300/PACKAGE 4MM, (3/16" INNER DIAMETER) NON LATEX ELASTICS LOT 307850***DISTRIBUTED BY: BIOMET MICROFIXATION 1520 Tradeport Drive, Jacksonville, FL 32218 USA www.biometmicrofixation.com***."Inner Vendor/Manufacturer Label labeled in part:"H4 Elastics 4mm (3/16") Heavy 4 oz. REF 11-102-04***Manufactured for DENTSPLY GAC INTERNATIONAL DENTSPLY INTERNATIONAL, INC. Phone 631-419-1700 Fax 631-419-1599 355 Knickerbocker Avenue Bohemia NY 11716 USA***MediMark Europe BP2332 F-38033 Grenoble Cedex 2. France***LOT 82936-32 QTY:100 pk use by 2012-07 Made in USA ISSUED DATE 2009-04 GN***"Elastic bands used as elastomeric ties for ligation that facilitate treatment of fracture of the mandible and maxilla. They are commonly used with braces or arch bar wire. | Latex Free Rubber Band | 2 | October 05, 2011 | 2011 | Biomet Microfixation, Inc. | On 09/18/2011, Biomet Microfixation, Jacksonville, FL, recalled their SP-2411, Rubber Band, Lot 307850 because it was labeled and sold as "LATEX FREE" although it contains "LATEX". The firm's HHE identified this as a potential high health risk to the patients or end users with latex allergies. | PRODUCTION CONTROLS: Error in Labeling | Biomet Microfixation, Jacksonville, FL sent a Urgent Medical Device letter dated August 19, 2011, to all affected customers. They were also notified by phone, and followed-up by e-mails starting on 08/18/2011. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed If you have any of the affected products in your facility, please remove it and return it immediately to Biomet Microfixation for credit. In order to return these products, please utilize the Return Goods Authorization (RGA) number provided on the reconciliation form.Please follow the instructions on the form and fill out the attached Inventory Reconciliation sheet and FAX to (904) 741-9425. If you have any further questions, please call 1-800-874-7711 or (904) -741-4400, extension 9468, Monday through Friday, 8am-4pm ET. |  | 22 | 22 packs | Worldwide Distribution-- USA (nationwide) including the states of FL, KY, MN, NJ, SD, and VA., and the country of Australia. | Terminated | February 06, 2012 | 124 | 22 | Not_Computer
Z-0016-2012 | 59854 | catheter introducer kit | OFD | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | DeRoyal (R) Heat/Moist Exchange, REF 83-000042, Rx Only, NON-STERILE, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA: Intended Use: Varies by kit; an anesthesia breathing circuit is a device that is intended to administer medical gases to a patient during anesthesia. It provides both an inhalation and exhalation route and may include a connector, adaptor, and Y-piece. | DeRoyal Surgical Kits and Vital Signs Anesthesia Circuits | 2 | October 06, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm manufactured kits containing a HCH device, manufactured by Vital Signs, that was subsequently recalled because the product may have an occlusion that could prevent proper flow of medical gases or oxygen to the patient. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, DeRoyal, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated September 13, 2011 to its customers. The letter describes product, problem, and actions to be taken. Customers are instructed to identify inventory being recalled and place in quarantine to prevent further use. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Customers were asked to complete the attached Recall Response Spreadsheet and return as per the instructions. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Distributors were instructed to perform a sub-recall to their customers.Customers with questions should call 865-362-1037 between 9am and 5pm EST or email recalls@deroyal.com. |  | 200 | 200 units | Nationawide Distribution USA - including the states of: CA, FL, GA, NJ, and WA. | Terminated | January 30, 2013 | 482 | 1140 | Not_Computer
Z-0020-2012 | 59868 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, 1600 tests, IVD, Made in United States.Kit contains:1. Elution Buffer A 3 x 2500 mL, Bis-Tris Phosphate Buffer2. Elution Buffer B 1 x 1500 mL Bis-Tris Phosphate Buffer3. Cartridge Set, 2 cation exchange analytical cartridges, 4 cation exchange Guard cartridges.4. CD-ROM 1 - CD Rom with program parameters5. Calibration/Diluent set - 2 levels, 2 vial each of lyophilized human red blood cell hemolysate with preservative, 1 bottle diluent, 100 mL deionized water.6. Whole Blood Primer, 6 x 1 mL, lyophilized human red blood cell hemolysate with preservative,7. Sample vials - 2 x 100, 1.5 mL polypropylene vials with pierceable caps.8. Instruction Manual.Bio-Rad Laboratories, Hercules, CA 94547.Percentage determination of hemoglobin A1c in human whole blood using ion-exchange high=performance liquid chromatography (HPLC) | VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, 1600 tests, | 3 | October 07, 2011 | 2011 | Bio-Rad Laboratories, Inc. | A previous recall related to reports of calibration failure and late retention times using Elution Buffer led to an examination of other buffers manufactured during the same time period and a decision was made to field correct these products until a definitive root cause and corrective action have been found and implemented. | PRODUCTION CONTROLS: Process Control | Bio-Rad sent an Urgent Medical Device Correction letter on September 12, 2011, to all affected subsidiaries and customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard all affected product in accordance with all local, state and federal disposal procedures. If customers distributed the product further they should direct their customers to discard all unused product.Customers were asked to please translate (if necessary) the attached "Customer Medical Device Correction Response Form," customers should add their local contact information and distribute to all affected customers in their region by mail or fax.Compile tracking documentation (e.g. Excel spreadsheet) for all customers who received the Medical Device Correction Letter.Complete the "SUBSIDIARY MEDICAL DEVICE CORRECTION RESPONSE FORM" and fax to Bio-Rad Regulatory Affairs at (510) 741-3954, or e-mail PDF to CSD_RA/Hercules/US/BIO-RAD.Customers would be issued credits for the affected products based on their request. For questions call (510) 741-3954. |  | 82 | 82 | Nationwide and to New Zealand, Canada, France, England, China, India, Mexico, Italy, Brazil, Australia, Singapore, Thailand, SriLanka, hong Kong, Sweden, Germany, Peru, Jamaica, Uruguay, Columbia, Argentina, Chile, Paraguay and Korea. | Terminated | May 11, 2012 | 217 | 287668114 | Not_Computer
Z-0030-2012 | 59957 | fibrin split products | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Minutex D-Dimer ; Trinity Biotech plc,Bray,Co. Wicklow,Ireland.Tel: (353) 1 276 9800,Fax: (353) 1 276 9888,Web: www.trinitybiotech.comMinutex D-Dimer is a latex agglutination test for semiquantitative determination of fibrin D-Dimer. | Minutex DDimer | 2 | October 07, 2011 | 2011 | Tcoag Us, Incorporated | There is a performance issue with Minutex D-Dimer Kit which may cause false negative results for patient samples containing levels of greater than 250ng/mL D-Dimer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm,Tcoag US, Incorporated, sent an "Urgent Device Recall" letter dated June 13, 2011 via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products subject to recall; complete and return the attached Product Recall Notification Response Form by June 22, 2011 via fax at 973-775-3085; scrap the products consistent with their laboratory policy and if they have further distributed these products, notify their customers to quarantine and scrap the recalled product.For questions or technical assistance, please call the Tcoag hotline at 1-888-291-0415, option 2 (7.30 am to 6.00 pm Eastern Time). |  | 4409 | 4409 Units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | December 11, 2012 | 431 | 4409 | Not_Computer
Z-0032-2012 | 59853 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | MDC PACS - release R2.3 SP1 .Phillips HealthcareAndover, MassachusettsA software application that is used for receiving, managing, archiving, distributing and recording medical images onto portable digital media (including but not limited to Compact Disk and DVD). | Philips Medical | 2 | October 07, 2011 | 2011 | Philips Healthcare Inc. | Clinical measurements on true size printouts may be inaccurate | DESIGN: Software Design | Philips Medical notified all affected customers with a "Field Safety Notice" dated August 2011.The letter provides instructions to the customer for how to avoid the problem. Customers are informed that until the software upgrade is recieved from Philip, do not use the true-size printing feature if the system is configured for more than one DICOM printer. Customers will be instructed that they will receive a software upgrade when it is available. For questions on this recall contact Philips representative at (800) 722-9377. | Contact the recalling firm for information | 20 | 20 units | Worldwide Distribution: Nationwide distribution including Puerto Rico; and the countries of Argentina, Brazil, China, France, Greece, India, Israel Mexico, and Turkey. | Open |  |  | 20 | Software | Output/Calculation | Software update | Software Update
Z-0029-2012 | 59787 | biopsy needle kit | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OLYMPUS ViziShot EBUS Aspiration NeedleEndobronchial Ultrasound (EBUS) Needle for Fine Needle Aspiration (FNA) - the needle has been designed to be used with ultrasonic endoscopes for ultrasonically guided fine needle aspiration of submucosal and extramural lesions of the tracheobronchial tree and the gastrointestinal tract. | OLYMPUS ViziShot EBUS Aspiration Needle | 2 | October 07, 2011 | 2011 | Olympus America Inc. | Firm has received complaints from users of the ability to wipe material from the needle's stylet and the presence of material in patient cytology samples obtained by the needle following endobronchial ultrasound fine needle aspiration of a lesion. Olympus Medical Systems Corporation (OMSC) in Japan has confirmed that metal particles can occur on the stylet due to friction between the stylet and n | DESIGN: Device Design | OLYMPUS sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A revised Instruction Manual with the new WARNINGS will be sent to the customer's attention.Customers were instructed to examine their inventory to determine if they have the affected product.To share the information on the new WARNINGS with all of the appropriate staff at their facility and any other organization where the product has been transferred.Complete and return the enclosed questionnaire even if they do not have affected product. Customers were instructed to fax the completed questionnaire to 484-896-7128.For any question call 484-896-5688. |  | 22498 | 22,498 | Nationwide Distribution including the District of Columbia and the following States: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | September 28, 2012 | 357 | 22498 | Not_Computer
Z-0031-2012 | 59746 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) Cuvettes The Dimension Vista(R) Cuvettes are for use on the Dimension Vista(R) System. The Dimension Vista(R) integrated system is an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluid. Vista (TM) system chemical and immunochemical applications utilize photometric, turbidimetric, chemiluminescence, nephelometric and integrated ion selective multisensor technology for clinical use. | Dimension Vista(R) Cuvettes | 2 | October 07, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed a low frequency defect in the molding of the Cuvettes that may result in Cuvette Ring and Cuvette Loader jams. It has been determined that the issue is related to a molding defect with the Cuvette flange or "wing" breaking, resulting in an error condition. Therefore, patient results are not impacted. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens Healthcare Diagnostics sent a Urgent Field Safety Notice letter dated August 2011, to all affected customers. The letter identifed the product the problem and the action needed to be taken by customer.Customers were instructed to discard all cuvette cartridges of the identified lots. Customers were also informed that until a replacement lot is available, if an error is encountered, they may contact the Technical Solutions Center for assistance at (800) 441-9245. For replacement product , please call the Siemens Atlanta Customer Service Center at (800) 241-0420. |  | 1786 | 1,786 | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and Puerto Rico., and the countries of Austria, Belgium, France, Germany, Italy, The Netherlands, Portugal, Spain, Switzerland, Australia, Japan, South Korea, and Canada. | Terminated | December 24, 2013 | 809 | 1786 | Not_Computer
Z-0028-2012 | 59535 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2, Part Number: 378600, Serial Numbers: 1001008698, 1001008699, and 10018700.Coulter PrePlus 2 is a microprocessor-controlled pipetting and diluting system, designed for automating sample preparation or assay methods. It is capable of aspirating and dispensing liquid samples. | Coulter PrepPlus 2 | 2 | October 07, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the Coulter PrePlus 2. The reagent level sensing is impaired and will give no warning when a reagent vial is low or out of reagent. Due to the size difference in the vertical mounts, the reagent rack is not aligned optimally for the probe to aspirate or level sense the antibody. This misalignment yields impaired reagen | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Beckman Coulter initiated an URGENT: PRODUCT CORRECTIVE ACTION (PCA) letter dated August 24, 2010, with attached PCA Response Form to all affected customers who purchased the Coulter PrepPlus 2. The letter provided the customers with an explanation of the problem identified and an action to be taken. The customers were instructed to not use the PrepPlus2 until their Beckman Coulter representative corrects the problem. A review of all results generated from samples prepared from the system should be conducted. Any unexpected dim fluorescence and/or decreased number stained cells should be evaluated for re-analysis, including preparing a new sample. Customers were also instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding this PCA, were instructed to call Beckman Coulter Customer service at (800) 526-7694. Outside of USA, customers were instructed to contact their local Beckman Coulter Representative. | No consumer action necessary | 391 | 391 units total (193 in US) | Worldwide distribution. | Terminated | April 04, 2012 | 180 | 391 | Not_Computer
Z-0036-2012 | 59937 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, Synchrony Respiratory Tracking System (Robocouch Patient Positioning System), a Radiation Therapy Device. Model Numbers: 026430, 028186. Product is manufactured and distributed by Accuray, Inc., Sunnyvale, CA | CyberKnife Robotic Radiosurgery System | 2 | October 11, 2011 | 2011 | Accuray Incorporated | A report and investigation determined that the set screws, which were part of the retrofit of a previous recall were incorrectly engaged at one facility which allowed rotation and subsequent descent of the assembly. | PRODUCTION CONTROLS: Process Control | Urgent Device Correction Letters were sent to all sites affected by the retrofit on September 12, 2011 via Fed Ex. The letter identified the affected product and the description of the potential problem. Customers were asked to inspect their Synchrony System by following the instructions provided. If the lower extension feels loose, customers should contact Accuray Customer Support immediately to schedule service of the system. The letter also states that all affected systems will be inspected and corrected, if necessary by Accuray to prevent the potential problem. Accuray Customer Support will contact customers to schedule the inspection. Questions or concerns regarding the issue should be directed to Accuray Customer Support at 1-877-8667 for US customers and 1-408-716-4700 for customers outside of the US, or customersupport@accuray.com. | Contact the recalling firm for information | 236 | 236 total devices | Worldwide Distribution | Terminated | February 27, 2012 | 139 | 236 | Not_Computer
Z-0034-2012 | 59756 | curette, surgical, general use | FZS | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Axle Interspinous Fusion System Torque Driver Handle, Model #: X060-0323, Lot # 525941, S/Ns 001-017, Lot # 525942, S/Ns 001-050, Mfr. By: Bradshaw Medical Inc., Kenosha, WI 53144The Axle Torque Driver Handle is used to lock the set screw into the locking plate. | Axle Interspinous Fusion System Torque Driver Handle | 2 | October 11, 2011 | 2011 | X Spine Systems Inc | The firm learned that the Torque Driver Handles were not actuating at a proper torque. It was determined that the handles were speed determinant and could output a higher torque if the handles were turned at a very slow speed. | DESIGN: Component Design/Selection | X-spine Systems, Inc. sent an URGENT DEVICE RECALL letter dated September 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their stock immediately to determine if they have any affected product. If so, customers were instructed once the set screw is tightened onto the locking plate, and resistance is detected in the handle, it will take less than 2 seconds to hear the handle "click." The "click" indicates the torque limit was reached. If further resistance is felt after 2 seconds, the surgeon should release the handle from tension, and try again.If the product was distributed further, they should contact their accounts, and advise them of this recall.The firm will send replacement handles.For any questions call (937) 847-8400, ext 115. | Contact the recalling firm for information | 39 | 39 Torque Driver Handles | Nationwide Distribution including CA, CO, FL, IA, IN, LA, NY, PA, TN, TX and UT. | Terminated | March 01, 2013 | 507 | 39 | Not_Computer
Z-0033-2012 | 59879 | adaptor, y | FJP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380, Serial Numbers: 5103031102, 5101101101, 5102161007, 5103221007, 5104221007, 5105101008, 5109131010, 5111231011, 5114301012, 5102131101 & 5113091012. Manufactured by: Isopure Corp, Simpsonville, KYSystem is used to mix and distribute bicarbonate as a part of a clinical dialysis system. | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model MD380 | 3 | October 11, 2011 | 2011 | Isopure Corp | Software problem causes transfer line between mixing and storage tank to not be rinsed after disinfection, leaving residual chlorine in the line. | DESIGN: Software Design | Isopure sent a letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Isopure devised an immediate solution which must be implemented at their faclity in conjunction with the disinfection procedure. Customers were asked to inform all appropriate personnel, clinical and technical, who may be involved in operating the system during the normal disinfection schedule.On 5/19/2011, the firm sent letters to their customers via US mail and email. They also telephoned their customers to notify them of the problem. | Contact the recalling firm for information | 11 | 11 units | Nationwide Distribution including AR, CT, GA, MA, ME, NJ, PA, TN and WV. | Terminated | February 19, 2013 | 497 | 11 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-0043-2012 | 58746 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Harleco Giemsa Stain 25 gram glass bottle, Item #228-12 Manufactured for EMD Chemicals Inc., Gibbstown, NJ 08027 1-800-222-0342Made in India.The product is a Class I, IVD dye and chemical stain for use in Hematology. | Harleco Giemsa Stain | 3 | October 12, 2011 | 2011 | EMD Chemicals Inc. | Stains may yield lower then expected absorbance levels resulting in under stained hematology specimens. | TRAINING: Employee Error | EMD Chemicals Inc. sent an "IMPORTANT PRODUCT RECALL" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. An attached response form is included for customers to complete and return via fax to EMD Chemicals at 856-423-6314. Any questions concerning this recall notice should be directed to EMD Chemicals Technical Service Department at 1-800-222-0342. | Contact the recalling firm for information | 36 | 36 units | USA Distribution in the states of PA and NY. | Terminated | October 13, 2011 | 1 | 36 | Not_Computer
Z-0037-2012 | 59898 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | LATITUDEÂŽ Paceart Integration Software, version 1.00 Model 6472.LATITUDE Paceart Integration is software and therefore does not have any serial numbers. However, the software version affected is version 1.00 and is distributed on USB pen drives. These pen drives are lot number controlled and denoted as Model 6493.When used in conjunction with the LATITUDE EMR Integration software transfers data files from the Boston Scientific LATITUDE Patient Management system to a clinic's Paceart system. | LATITUDE Paceart Integration Software | 2 | October 12, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that under certain conditions, LATITUDE Paceart Integration (LPI) softwaremay not properly transfer data resulting in the incorrect summary of programmed parameter values in the ICD Summary tab display and Current Programming Tachy summary section of the Paceart report. | DESIGN: Software Design | Consignees were sent on 8/12/11 a Boston Scientific letter dated August 2011. The letter was addressed to "Dear Doctor". The letter described the product and problem. It also gave Background Information, Description of Potential Data Translation Error, and Additional Information and Action information.A letter, including an HRS form, was also sent to all US LATITUDE Clinics that have active LPI software to notify them of the need to install the corrected LPI Software. | Contact the recalling firm for information | 398 | 398 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT DE, DC, FL, GA, IL,IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OJ, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV, WI, and WY. | Terminated | November 06, 2012 | 391 | 398 | Software | Device Operation | Software update | Software Update
Z-0042-2012 | 59891 | test, urea (breath or blood) | MSQ | Campylobacter fetusserological reagents. | Microbiology | Microbiology | 510(k) | IDkit:HP" Breath Test for Detection of H. pylori, Manufactured by: Exalenz Bioscience Ltd 4 Ha' maayan St. Modiin, Israel 71700, Distributed by: Exalenz Bioscience Inc., 1313 N. Market St., Sute 5100, Wilmington, DE 19808 USAProduct Usage: The IDkit:HP Two Test System is intended for use in the qualitative detection of urease associated with Helicobacter pylori in the human stomach and as an aid in initial diagnosis and post treatment monitoring of H. pylori infection in adult patients. Urea Breath Tests are well established as reliable, non-invasive methods for detecting H. pylori. They are highly endorsed by various medical societies as the test of choice for initial detection as well as for post-treatment testing. | IDkit:HP" Two Test System | 2 | October 12, 2011 | 2011 | Exalenz Bioscience Ltd | The product is in violation of the FD&C; Act in that it was illegally marketed without proper clearance (no 510(k) submitted). |  | The firm, Exalenz Bioscience Ltd. notified consignees located in the United States of the recall by telephone, electronic mail, or written correspondence between June 9, 2011 and June 30, 2011. The letter informed consignees that Exalenz will immediately cease marketing the IDkit:HP TWO test in the United States. Consignees were instructed to return shipment of the IDkit:HP TWO test kits that they have in their possession to the companys US warehouse. For question call the VP of North American Sales at 704-770-6162 or VP of Marketing at 617-794-6936. |  | 1240 | 1240 units | Nationwide Distribution (USA) - including the states of: FL, IL, LA, MA, MS, NY, TN, TX and VA. | Terminated | July 06, 2012 | 268 | 1240 | Not_Computer
Z-0039-2012 | 59316 | microtome, rotary | IDO | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific Microm: Rotary Microtome: REF HM355S-3Used for slicing tissue blocks. | Thermo Scientific Microm | 2 | October 12, 2011 | 2011 | Richard-Allan Scientific Company | The firm has received two complaints of finger lacerations in operators of this device. | DESIGN: Device Design | ThermoFisher Scientific sent "URGENT MEDICAL DEVICE CORRECTION, RE: Thermo Scientific HM355S-3 User Safety Advisory Notice" to all affected customers.The letter states the manufacturer identified a need to enhance the safety instructions and the description of the safety devices in the user manual. It mentions the HM3555S-3 User Manual has been modified to include more specific safe work practices and a new Quick -Reference Card summarizing the manual's safety enhancements and that the changes have been included on a CD that is enclosed with the letter. The firm requests confirmation of receipt of the new User Manual and Quick-Reference Card by October 31, 2011 to be sent by FAX to (269) 372-2674For questions on this recall call the firms Technical Support representatives at 1-800-522-7270 08:00 AM- 4:30 PM. | No consumer action necessary | 786 | 786 | Worldwide Distribution: Nationwide distribution; including the country of Canada. | Terminated | July 25, 2013 | 652 | 786 | Not_Computer
Z-0038-2012 | 59888 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) EXL LOCI Cardiac Troponin I Calibrator (RC621)The LOCI Cardiac TNI Calibrator is an in vitro diagnostic product for the calibration on Cardiac Troponin I method on the Dimension(R) EXL (TM) Integrated chemistry system with LOCI(R) module. | Dimension(R) EXL LOCI Cardiac Troponin I Calibrator (RC621) | 2 | October 12, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Internal testing has confirmed that recovery of levels 2 and 3 of LOCl Cardiac Troponin I Calibrator is lower than expected bottle value. Following calibration with LOCl Cardiac Troponin I Calibrator, lot 1AD045 an upward shift in patient and QC results may be observed. | OTHER/UNDETERMINED: Pending | Siemens sent an Urgent Field Safety Notice dated August 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard any remaining inventory. Replacement orders would be processed and sent free of charge. Customers were also instructed to complete the attached FIELD CORRECTION EFFECTIVENESS CHECK form and fax it to (302) 631- 8457.Customers were instructed to forward the Safety Notice to anyone to whom they may have distributed the affected product. For any technical questions customers were to contact the Siemens Solutions Center at 800-441-9250. |  | 998 | 998 | Worldwide Distribution - USA (nationwide) including AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of GUAM, Austria, Canada, China, Czech Republic, Egypt, France, Germany, Italy, Poland, Portugal, Puerto Rico, Spain and Switzerland. | Terminated | May 20, 2013 | 586 | 998 | Not_Computer
Z-0040-2012 | 59712 | system, applicator, radionuclide, manual | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Prostrate Seeding Applicator Needle. Product code PSS1820ATPacked in boxes of (25) with five Single Packs and four five packs or as singles or five packs.Worldwide Medical Technologies, LLCUsed to introduce radionuclide seeds of 1-125 or Pd-103 into the body for radiation therapy. This device is typically used in the transperineal approach for placing radionuclide seeds in the prostate | Prostate Seeding Needle, Applicator Technique | 2 | October 12, 2011 | 2011 | Worldwide Medical Technologies, LLC | Length of the cannula hub is longer and may not permit a fit with the Applicator | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BrachySciences, a division of Biocompatibles, Inc., sent out "Urgent Recall Device Recall" letters Dated August 12, 2011 to all affected customers.The letter included affected product and problem description. Customers are asked to quarantine and return unused product. A Product Recall & Retrieval Form was included to send back to the firm.For questions on this recall contact the firm at (203) 262-0571, extension 105. | Contact the recalling firm for information | 1872 | 1872 units | Worldwide Distribution: Nationwide distribution including the states of AR, CO, DE. FL, KY, MD, MI, MN, NY, TX, and VA; and the countries of Belgium, and Germany. | Terminated | September 19, 2012 | 343 | 1872 | Not_Computer
Z-0041-2012 | 59418 | electrocardiograph | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ECG Cable Part Number: 01.57.471017 for use with the SE-1 series, SE-3 series and SE-601 series Smart Electrocardiograph ECG cable is an accessory of the following machines: SE-1 series, SE-3 series and SE-601 series electrocardiograph (ECG). The EDAN ECG is a diagnostic and prescription device only intended to be used in hospitals or healthcare facilities by qualified doctors and trained healthcare professionals. The cardiogram recorded by the electrocardiograph can help medical professionals to analyze and diagnose heart disease. However, the interpreted ECG with measurements and interpretive statements is offered to clinicians on an advisory basis only. ECG cable is packed in plastic bag, as an accessory putting together with ECG machine in the package box. | ECG Cable | 2 | October 12, 2011 | 2011 | Edan Instruments, Inc. | The 3 leads of the cable, RL, LA and LL were labeled in wrong sequence. The left arm (LA) is mislabeled as right leg (RL). The left leg (LL) is mislabeled as left arm (LA). The right leg (RL) is mislabeled as left leg (LL). There have been no reports of injury as a result of this issue. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Edan Instruments, Inc. sent an Urgent - Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that wrongly labeled cable may lead to incorrect ECG waveforms during ECG test, customers were reminded to review previous diagnoses made from this cable to mitigate any inappropriate recordings or treatment made.Customers were instructed to replace the affected ECG Cables once they receive the new ECG Cable, and collect all the affected ECG Cables and then destroy and dispose of them according to local regulations. Questions or concerns were directed to the following numbers: Oceania and Asia: +86 (755) 2689 8322North America, South America:+86 (755) 2689 9914Europe: +86 (755) 2669 4684. | Contact the recalling firm for information | 615 | 615 Ccables | Worldwide Distribution - USA (nationwide) and the countries of Dominica, Ecuador, Colombia, Honduras, Peru, Mexico, El Salvador, Guatemala, Venezuela, Uruguay, Chile, Panama, Nicaragua, Ecuador, Australia, Philippines, Malaysia, Indonesia, Chinese Taipei, & Spain | Terminated | October 13, 2011 | 1 | 615 | Not_Computer
Z-0045-2012 | 58461 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS System, Single Use; CATALOG #'s 9045-01 & 9045-02. Interventional Spine, Inc.A sterile kit used to stabilize the spine as an aid to fusion through immobilization of the facet joints. | PERPOS PLS System Single Use | 2 | October 13, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to drivers that may contain a burr in the distal tip preventing full seating of the driver onto the head of the facet screw, resulting in difficulty screwing the facet screw into the bone.Use of this product may result in a delay of the surgical procedure | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Interventional Spine, Inc. sent an "Urgent Recall" letter dated March 15, 2011 to all affected customers.The letter included; affected product, reason for recall and recommendations.Customers were asked to immediately examine inventory, quarantine affected product and to notify International Spine, Inc if any product remain.For information on this recall contact the firm at (949) 472-0006. |  | 2620 | 2,620 units | Word wide Distribution: (USA) Nationwide including the states of CO, CT, FL, KY, LA, MA, NC, NY, OK, OR, TN, TX, VA, and WI; and the countries of Italy, Spain, Turkey, and the U.K. | Terminated | February 05, 2013 | 481 | 2692 | Not_Computer
Z-0055-2012 | 59526 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Safe Sheath Introducer, Catalog number: V18; AngioDynamics, INC.The Valved Introducer sheath is utilized to surgically place the catheter during the port implantation procedure. | Safe Sheath ULTRALITE 8F Safe Sheath Introducer. | 2 | October 13, 2011 | 2011 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | AngioDynamics, Inc contacted affected customers initially via telephone and a follow up "Notification of Medical Device Correction" letter was sent dated August 1, 2011. The letter informed them of the problem and provided instructions. Customers were instructed to utilize the 8F Introducer Sheath provided within the 8F Smart Port CT Implantable Port during port implantation. The 8F Valved Sheath (V18) may continue to be utilized during the placement of the 7.5F Smart Port Implantable Port (for example: CT75STSD; CT75STSD-NF). For question on this recall contact AngioDynamics, Inc at (800) 772-6446. | Contact the recalling firm for information | 1095 | 1095 total units | (USA) Nationwide distribution including the states of CA, FL, GA, IA, IL, IN, KY, MD, MI, MO, NC, NE, NJ, NY, OH, OK, OR, TN, UT, WA and WI. | Terminated | March 07, 2012 | 146 | 1095 | Not_Computer
Z-0048-2012 | 59768 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott m2000rt System software; an automated system for performing fluorescence-based PCR that results in quantitative and qualitative detection of nucleic acid sequences; Abbott Molecular Inc., Des Plaines, IL 60018 USA; catalog #09K15-01The Abbott m2000rt System provides for real-time measurement of the stages of the polymerase chain reaction (PCR). Real-time PCR measures PCR amplification as it occurs, cycle-by-cycle, allowing quantitative measurements to be made during the highly reproducible exponential phase of PCR. | Abbott m2000rt System software | 2 | October 14, 2011 | 2011 | Abbott Molecular | The Abbott m2000rt System software has software bug which can affect customers who manually enter a Test Order onto the m2000rt and sorts a column on the calibration or control grids of the Test Order screen by double clicking the column header. This causes a mismatch between specific concentration values and the designated calibrators and controls in the final results, which could cause an inval | DESIGN: Software Design | Abbott Molecular Inc. sent a Field Correction Recall letter dated September 8, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to your m2000rt software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2012.Please complete the enclosed reply sheet and fax it to (847) 775-6728 at your earliest convenience to acknowledge receipt. We appreciate your assistance in this matter. |  | 926 | 926 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Belgium, Benin, Bolivia, Bosnia and Herzegovia, Brazil, Bulgaria, Burkina Faso, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. of the Congo, Denmark, Egypt, Estonia, Ethiopia, Finland, France, Gabon, Gaza & Jericho, Germany, Greece, Guatemala, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Libya, Lithuania, Malawi, Malaysia, Mali, Mauritius, Mexico, Moldova, Morocco, Mozambique, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Venezuela, Vietnam, Zambia and Zimbabwe | Open |  |  | 926 | Software | Output/Calculation | Software update | Software Update
Z-0044-2012 | 60008 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA Toric (enfilcon A) TORIC soft (Hydrophilic) Contact Lenses, 6 SOFT CONTACT LENSES FOR ASTIGMATISM, CooperVision, Scottsville, NY 14546 USA; Juana Diaz, PR 00795 USA; Hamble SO31, 4RF, UK; Device Listing # D033736.Product Usage: Indicated for the correction of ametropia (myopia or hyperopia with astigmatism) in aphakic and non-aphakic persons with non-diseased eyes in powers from -20.00- to +20.00 diopters and astigmatic corrections from -0.25 to -10.00 diopters. The AVAIRA Soft Contact Lens is indicated for single - use disposable wear and daily wear. When prescribed for frequent replacement, the lens may be disinfected using a chemical or hydrogen peroxide disinfecting system. | AVAIRA | 1 | October 14, 2011 | 2011 | CooperVision Inc. | This recall was initiated because of the unintended presence of a residue (silicone oil) on lenses from certain lots of Avaira Toric contact lenses. The company received complaints of temporary hazy or blurry vision and discomfort, most likely caused by transient or reversible swelling of the surface layer of the cornea, which is known as corneal epithelial edema. Corneal epithelial edema results | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, CooperVision sent an "Urgent: Company Initiated Voluntary Medical Device Recall - Contact Lenses" notification letter to all affected consignees. The notice identified the product, the problem, and the actions to be taken. The letter includes an attached "Recall Response Form" that consignees are required to complete and fax to (585) 889-5688. Consignees were instructed to immediately examine their inventory, stop any further distribution, quarantine and return product subject to recall. In addition, consignees were told if they have further distributed the affected product to notify customers at once of this product recall. The letter states that accounts will either be credited or exchanged for an unaffected product. If you have any further questions, call 1-800-341-2020. |  |  | At least ( 778,301 lenses: Domestic - 659,847 lenses; Foreign - 118,454 lenses; including EXPANSION: Domestic - 69,047 lenses) ( to be updated) | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Europe, Germany, Hong Kong, Italy, New Zealand, and Spain. | Open |  |  |  | Not_Computer
Z-0049-2012 | 59978 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ArthroCare brand MiniMagnum Knotless Fixation Implant Device with Inserter Handle Plus SpeedStitch Magnum Wire Suture Cartridge; Catalog Number: OM-6007;Product is manufactured and distributed by ArthroCare Corporation, Sunnyvale, CAThe MiniMagnum Knotless Fixation Device is indicated for use in fixation of soft tissue to bone. Examples of such procedures include: Shoulder: Bankart Repair, SLAP lesion repair, acromio-clavicular separation, rotator cuff repair, capsule shift/capsule-Iabral reconstruction, biceps tenodesis and deltoid repair. Ankle: Lateral instability, medial instability, Achilles tendon repair/reconstruction and midfoot reconstruction Foot: Hallux valgus reconstruction Elbow: Tennis elbow repair, biceps tendon attachment Knee: Extra-capsular repairs; reattachment of: medial collateral ligament, posterior oblique ligament or joint capsule closure to anterior proximal tibia; extra capsular reconstruction, ITB tenodesis; patellar ligament and tendon avulsions. | MiniMagnum Knotless Fixation Device | 2 | October 14, 2011 | 2011 | ArthroCare Corporation | ArthroCare initiated the recall of one lot of the MiniMagnum Knotless Implant Fixation Device where the incorrect diver block was packed with the implant. | PRODUCTION CONTROLS: Packaging Process Control | The firm, ArthroCare Corporation, sent a letter dated May 10, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were asked to immediately return the product; check their inventory, refer to the copy of the Mini Magnum outer label to help identify the product, and if the affected product lot is in their inventory, notify ArthoCare Product Support at (800) 797-6520, press 2 and return the product immediately. The letter also stated, "In addition, if you do not have any product from this lot remaining in your inventory, we request that you contact ArthroCare at the number above, or by e-mail at ra@arthrocare.com, to inform us of the lack of remaining product. This will ensure that you will not be contacted again and that we have accounted for product shipments from this lot." If you have any questions, please call the Chief Regulatory Officer and Senior Vice President, Quality Systems and Assurance at (408) 735-6229. |  | 32 | 32 units | Nationwide distribution: USA including states of: CA, FL, NY, OH and OK. | Terminated | October 24, 2011 | 10 | 32 | Not_Computer
Z-0050-2012 | 53046 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Total ÂŽhCG.The Access Total ÂŽhCG assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of total ÂŽhCG levels in human serum and plasma using the Access Immunoassay Systems. | Access Immunoassay Systems TOTAL hCG | 2 | October 17, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that certain Access Total ÂżhCG reagent packs show decreased stability from the date of manufacture, which may lead to increased imprecision. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recall communication was initiated on March 18, 2009 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter dated March 13, 2009 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems total ÂŽhCG. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the reagent lot numbers 824542 and 825652, and discard any unused reagent packs.Customers were asked to review their historical records for low level quality control to ensure that the assay performance for the referenced lot numbers has been within their laboratory specifications. They were also asked to share the information with their laboratory staff and retain the letter for their Quality System documentation, and then complete and return the enclosed response form within 10 days so that they may be assured that consignees received the notification.Finally they were told to contact Customer Technical Support Center at 1-800-854-3633 in the US and Canada, or their local Beckman Coulter Representative with any questions. |  | 14764 | 14,764 | Worldwide Distribution -- US, Algeria, Angola, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, canada, Chile, China, Colombia, Czech Republic, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kenya, Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Russian Federation, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Zambia. | Terminated | April 27, 2012 | 193 | 14764 | Not_Computer
Z-0054-2012 | 59912 | sucker, cardiotomy return, cardiopulmonary bypass | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(k) | SARNS INTRACARDIAC SUCKER, BULK, NON-STERILE | Sarns Adult Rigid Intracardiac Sucker; bulk, nonsterile | 2 | October 17, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) recently discovered three instances of metal fragments inside the stainless steel tube of the Sarns" Adult Rigid Intracardiac Sucker. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 9/13/2011 Terumo sent a letter entitled "URGENT MEDICAL DEVICE REMOVAL" to its customers that included the name and description of the device, lot numbers, the potential hazard, correction, the reason for the recall and instructions. Customers are to review the notice, ensure that all users are aware of the issue, and a recall response form to be completed and returned. Customers are to call Terumo CVS Customer Service to obtain a Return Goods Authorization number. The letter also included MedWatch reporting information and stated that Customers with questions should call 1-800-521-2818. | Contact the recalling firm for information | 5700 | 5700 | Worldwide Distribution -- USA, Belgium, Australia, Mexico, and Japan. | Terminated | December 10, 2012 | 420 | 15470 | Not_Computer
Z-0052-2012 | 59961 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Impac Medical Systems SequencerProduct Usage: The intended us of Sequencer is to assists in the process of patient treatment on a radiotherapy treatment machine. Sequencer provides the capability to notify the user of actions that need to take place prior to treatment, to display reference images for set up purposed and to auto setup the machine to predefined settings. The actual settings are read from the treatment machine, via the machines communication interface, and compared to predefined values. if a mismatch occurs between the planned values and the actual machine settings, treatment is prohibited. At appropriate points during the treatment, the actual delivered values are recorded to provide treatment tracking. | SEQUENCER | 2 | October 17, 2011 | 2011 | Impac Medical Systems Inc | Varian reported a problem related to certain collimator rotation values which are incorrect on TrueBeam compared to planned values in Mosaiq. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | "Impac Medical Systems Inc., sent an USER NOTICE letter to all affected consignees on September 20, 2011. The letters identified the source characterization error, clinical impact, and final resolution. The letter states that Elekta Support will follow up with each consignee and remotely correct the source characterization. Customers were instructed to sign, date and return the attached Confirmation of Receipt Form by one of the three ways indicated. The letter instructs consignees to distribute this notice to any and all users of IMPAC software that may potentially be affected by the issue. For questions, concerns, and requests for upgrade, email support@impac.com or call. The letter provided consignees with a complete list of contact information.. |  | 11 | 11 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, NJ, NY, NV, TX, VA and WA. | Terminated | March 06, 2012 | 141 | 11 | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0051-2012 | 59848 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES, 621-400ES, 621-400ESR.To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | Outlook ES Safety Infusion System | 1 | October 18, 2011 | 2011 | B Braun Medical, Inc | B Braun Outlook ES Safety Infusion System "Run" Light Emitting Diodes (LEDs) pumps stop infusing and a backup alarm sounds, but the "Run" LEDs advance as if the pumps were infusing. | DESIGN: Software Design | B. Braun Medical, Inc. sent an "URGENT: FIELD CORRECTION" letter dated September 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A firm representative will contact all affected customers in order to schedule the field correction. The correction consists of revising the Door and Main Processor software, inspecting the main board, and replacing the board, if warranted. A Field Correction Information Form was included in the letter for customers to complete and return. Contact Technical Product Support at 1-800-627-7867 for questions regarding this notice. |  | 6321 | 6321 | Nationwide Distribution | Open |  |  | 6321 | Software | Alarm/Message | Revise software/insepct and replace board if needed | Software Update
Z-0057-2012 | 59442 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx ultrasound system and M-Turbo ultrasound system is Medical Diagnostic Ultrasound.General purpose ultrasound system intended for use by a qualified physician for evaluation by ultrasound imaging or fluid flow analysis of the human body. | MicroMaxx ultrasound system and MTurbo ultrasound system | 2 | October 18, 2011 | 2011 | Sonosite, Inc. | This is a software anomaly. In some circumstances, if the Systolic/Diastolic (S/D) ratio value is interpreted directly from the display, the clinician could potentially use an incorrect value of the S/D ratio in his/her assessment. | DESIGN: Software Design | The firm SonoSite, Inc., sent an "Urgent - Medical Device Correction" letter dated September 29, 2011 to its foreign and domestic customers. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the three (3) steps provided in the letter to verify if the extended OB/GYN Calculations (EOBC) is enabled on their system. Additionally, SonoSite will perform a software upgrade to M-Turbo and/or MicroMax ultrasound system in order resolve the issue. The customers were instructed to contact Sonosite Technical Support via website: www.sonosite.com/sdratio; email: regulatory@sonosite.com (include contact information and address), or call in U.S. and Canada: 1-877-590-4982. If you have any questions or concerns regarding this matter, do not hesitate to contact in US and Canada at 1-877-657-8118, Monday-Friday, 6 am-5 pm PST. Customer can also send an email to Technical Support at service@sonosite.com. | No consumer action necessary | 4236 | 4236 units (US) and 4089 units (internationally) | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Afghanistan, Algeria, American Samoa, Argentina, Australia, Bolivia, Botswana, Brazil, Canada, China, Costa Rica, Cyprus, Dominican Republic, Estonia, France, Gambia, Germany, Great Britain, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Tanzania, Uganda, United Arab Emirates, Viet Nam, and Zambia. | Terminated | September 14, 2012 | 332 | 4236 | Software | Output/Calculation | Software update | Software Update
Z-0060-2012 | 59998 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANCHORLOK(R) LEADING EDGE(R) SOFT TISSUE ANCHOR SYSTEM, 2.5mm, 2.0 SUTURE, REF 5100-006, 1 each, STERILE R, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002Product Usage: The instrument is used as a soft tissue anchor | ANCHORLOK | 2 | October 18, 2011 | 2011 | Wright Medical Technology Inc | The product is labeled with an expiration date without adequate justification. | EXPIRATION DATING: Incorrect or No Expiration Date | The firm, Wright Medical Technology, Inc., sent an "URGENT: PRODUCT RECALL NOTIFICATION" letter dated September 20, 2011 to all affected consignees. The letters identified the products, the problem, and the actions to be taken. The letter instructed consignees to identify affected inventory. If recalled products are located consignees are instructed to contact customer service at 1-800-238-7117, for specific instructions on how to return the recalled products. The letter asked that this communication be forward to any associated sites if the recalled products were further distributed. Consignees are to complete and return the enclosed notification form by fax to 901-867-7401 or by email to: cathy.park@wmt.com. For questions regarding this recall contact 901-867-4324. |  | 596 | 596 units | Worldwide Distribution - USA (nationwide) and countries of: Argentina, Canada, Denmark, France, Great Britain, Ireland, Japan, and Spain. | Terminated | July 17, 2012 | 273 | 11960 | Not_Computer
Z-0071-2012 | 58302 | system, esophageal pacing | LPA |  |  | Cardiovascular |  | ***REF #200***TAPSCOPE 550. Each pouch contains one (1) 18 French esophageal stethoscope with bipolar pacing and recording electrodes with a YSI 400 series male connector thermistor. LOT xxxxxxx 400095-0200-C 100902. www.cardiocommand.com. ***REF #200-0010***TAPSCOPE 550 CONTAINS: Ten (10) 18 French esophageal stethoscopes with bipolar pacing and recording electrodes with YSI 400 series male connector thermistors. NON-STERILE. SINGLE USE ONLY. DO NOT REUSE. CAUTION: Federal (U.S.A.) law restricts this device to sale, distribution and use by or on the order of a licensed physician. Manufacturer: CARDIOCOMMAND, INC. Tampa, FL USA. (800) 231-6370 (813) 289-5555. www.cardiocommand.com. LOT xxxxxxx 400155-B 091101.***REF #200-0025***TAPSCOPE 550 Contains: Twenty five (25) 18 French esophageal stethoscopes with bipolar pacing and recording electrodes with YSI 400 series male connector thermistors. ***LOT xxxxxxx***FOR SINGLE USE ONLY. DO NOT REUSE. NON-STERILE. CAUTION: Federal (U.S.A.) law restricts this device to sale, distribution and use by or on the order of a licensed physician. Read Model 2A operator's manual prior to initiating pacing procedures. Manufacturer: Tampa, FL USA 33607. (800) 231-6370 (813) 289-5555. www.cardiocommand.com. 400215-A 100902. | TAPSCOPE 550 | 2 | October 19, 2011 | 2011 | Cardiocommand Inc. | CardioCommand, Inc. is recalling their TAPSCOPE 550 esophageal pacing/recording catheters because there maybe electrical discontinuity, which can result in the inability to reliably pace and/or record ECG. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On October 25, 2010, an Urgent Medical Device Recall Notice was sent by certified mail to all facilities advising them of the recall. The letter identified the affected product and reason for the recall. Users were asked to discontinue use and return the affected product to the firm. If all product was already used, then this should be indicated on the form provided. The firm will send replacement product in quantities equal to the returned product. The attached form should be completed and returned via fax. Customers were instructed to report any suspected incidents related to this issue to CardioCommand Customer Service at 1-800-231-6370. Also questions or comments should be directed to CardioCommand Customer Service. |  | 567 | 567 total | Nationwide Distribution -- GA, OR, TX, CO, KY, MT, MA, AL, CA, FL & TN. | Terminated | August 28, 2012 | 314 | 567 | Hardware | Device Operation | Replace | Remove or Replace
Z-0068-2012 | 60000 | drape, adhesive, aerosol | KGT | Drape adhesive. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | STRIP T's STERILE ADHESIVE ORGANIZER, Model #SC5000-4R, Lot #090710, 510(k) exempt. One (1) adhesive organizer is packed in a Tyvek pouch and 60 pouches are placed into a 'Shelf Box'. The shelf box is placed into a shipping container. The Sterile Adhesive Organizer is used to hold tubing or instruments to the drape in surgery. | STRIP T'S Sterile Adhesive Organizer | 3 | October 19, 2011 | 2011 | Kapp Surgical Instrument Inc | A portion of the labeled expiration date is missing. The expiration date on approximately half of the pouch labels distributed are missing the last digit of the year of expiration. However, the shelf box containing the pouches is labeled correctly. | EXPIRATION DATING: Incorrect or No Expiration Date | Surgical Concepts, Inc ( A subsidiary of Kapp Surgical Instrument Inc) sent a notification letter dated March 5, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm ceased distribution of the suspect product on December 5, 2007. The letter informs the firm's customers (both distributors and medical user/customers) of the expiration date labeling omission of the last digit of the year of expiration. It goes on to explain that the incomplete expiration date is present on the individual packages only and that the shelf boxes are labeled correctly with the complete/correct expiration date. The letter provides the customers with a contact person and a toll free telephone number 1-(800) 282-5277, to call in order to make arrangements for receive replacement labels for any misbranded packages received. |  | 59 | 59 boxes (3540 units) | Nationwide Distribution-- including the states of AK, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MO, MT, NC, NM, OH, PA, TN, TX, UT, and VA | Terminated | February 19, 2013 | 489 | 59 | Not_Computer
Z-0070-2012 | 60007 | catheter, intravascular occluding, temporary | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre 9F Pruitt F3 Outlying Carotid Shunt with T-portThe Pruitt F3 Carotid Shunt (the Shunt) is a multi-lumen device with balloons at both the distal (internal carotid) and proximal (common carotid) ends of the shunt. The balloons, when inflated independently, act as a stabilization mechanism to maintain the position of the Shunt when it is placed within the common and internal carotid arteries. An external safety balloon located on the inflation arm leading to the distal (internal carotid) balloon acts as a mechanism to relieve pressure on the internal carotid balloon in the event it inflates above optimal size and pressure. The external safety balloon feature reduces the possibility of balloon over-inflation and resultant vessel damage. The Pruitt F3 Carotid Shunt has features to aid the user during shunt insertion and balloon inflation. The inflation path of the proximal (common carotid) balloon is color coded. Sterile saline is injected from the blue stopcock, through the blue lumen and into the blue common carotid balloon. The sleeve of the external safety balloon is yellow, to increase its visibility. Depth markings on the shunt body are for reference during insertion | 9F Pruitt F3 Outlying Carotid Shunt with Tport | 2 | October 19, 2011 | 2011 | Lemaitre Vascular, Inc. | Extended Expiration on the outer box label 2016-05, the correct expiration date is on the tray label: 2014-05 | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, LeMaitre Vascular, Inc., sent an "URGENT: 9F PRUITT F3 Carotid Shunt Device Field Safety Notice" dated October 4, 2011 to its customers. The notice describes the product, problem and action to be taken. The customers were instructed to identify their inventory; return the affected products for exchange with a properly labeled device; contact Le Matire customer service for replacement at 800-628-9470, and complete and return the RECALL NOTICE RECONCILIATION FORM via fax to 781-425-5049, even if no product is on hand. If you have any questions concerning this safety notice, please contact the Quality Assurance Engineer at 781-221-2266 ext. 183. | Contact the recalling firm for information | 144 | 144 units | Nationwide distribution: USA including states of: CA, FL, IA, IL, MA, MI, NE, NJ, NY, OH, SC, TX, and WI. | Terminated | March 14, 2012 | 147 | 144 | Not_Computer
Z-0069-2012 | 59889 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc. Dimension(R) EXL Cardiac Troponin I Flex(R) Reagent Cartridges | Dimension(R) EXL Cardiac Troponin I Flex(R) Reagent Cartridges | 2 | October 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed customer complaints for low recovery of QC and patient samples. | PRODUCTION CONTROLS: Process Control | Siemens Healthcare Diagnostics, Inc. Urgent Field Safety Notice was sent to customers on August 31, 2011 regarding the issue. The letter instructs customers to immediately discontinue use of the affected lots and discard all remaining inventory. |  | 6072 | 6,072 cartons | Worldwide distribution: USA, Puerto Rico, Guam, Austria, Canada, Croatia, France, Greece, Germany, Italy, Japan, Mexico, Portugal, Spain and Switzerland. | Terminated | January 14, 2014 | 818 | 6072 | Not_Computer
Z-0076-2012 | 59911 | guide, surgical, instrument | FZX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | MICROPLASTY TAPERLOC 7.5 MM RESECTION GUIDE, STAINLESS STEEL, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE PO BOX 587 WARSAW IN 46582 USA.This device is used to mark the femoral neck resection level | BIOMET MICROPLASTY TAPERLOC 7.5 MM RESECTION GUIDE | 2 | October 20, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the blueprint calls forengraved measurements on both sides on the Microplasty Taperloc 7.5MM Resection Guide. Theabove lots, as supplied to Biomet by a vendor, are engraved on only one side. To date, Biomet hasreceived 2 reports involving 2 units where the instrument was missing the engraving on one side. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Arrangements will be made for customers to return the affected product to Biomet. To confirm receipt of this notice call Biomet at 800-348-9500 ext. 3755 or 3756. Questions related to this notice should be directed to 574-371-3755 or 3756, Monday through Friday, 8 AM to 5 PM. |  | 22 | 22 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, IN, MI, OK, and VA and countries of Australia and Chile. | Terminated | March 07, 2013 | 504 | 22 | Not_Computer
Z-0075-2012 | 59968 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Total Hip Joint ReplacementAccolade TMZF Plus 127 Neck Angle V40 Hip Stem #3Accolade TMZF Plus 127 Neck Angle V40 Hip Stem #5Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 A subsidiary of Stryker Corp.Stryker FranceZAC Satolos Green PusignanAv de Staolas Green 69881 MetziaeCedex FranceThe TMZF Press Fit HA Stem and TMZF Press Fit Plus HA Stem are intended for the reconstruction of the head and neck of the femoral joint. The device is intended for primary reconstruction of the proximal femur or revision of a previous total hip arthroplasty. These stems can be used with any currently available Howmedica Osteonics acetabular components and V40 Femoral Heads that can be mated with a 5 40 BG taper. | Total Hip Joint Replacement | 2 | October 20, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that a size 3 Accolade stem was in a box labeled as a size 5 Accolade stem. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Stryker Orthopaedics sent an "URGENT: PRODUCT RECALL" letter dated July 14, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgment Form via fax to 201-831-6069. Additionally, information is provided for customers to return the affected products. Please contact the Divisional Regulatory Compliance Manager at 201-972-2100 for questions regarding this notice. |  | 21 | 21 units | Nationwide Distribution | Terminated | November 15, 2013 | 757 | 21 | Not_Computer
Z-0079-2012 | 59948 | gauge, pressure, coronary, cardiopulmonary bypass | DXS | Cardiopulmonary bypass coronary pressure gauge. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack, which are designed as specified by the user. The following models contain a Pressure Display Set. Sterilized by Ethylene Oxide. Non-pyrogenic. Manufacturer: Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604. Product Usage: This product is intended for use in monitoring catheter, cannula, or line pressures associated with cardiopulmonary bypass equipment and/or related products. | Medtronic Disposable Pressure Display Set | 2 | October 20, 2011 | 2011 | Medtronic Perfusion Systems | Pressure Display Set (models 61000, 61006, 62000, 62006) DFU will be updated to clarify that the intended use of the disposables is only for cardiopulmonary bypass surgery (procedures up to 6 hours). Use beyond 6 hours may result in air leakage in the system resulting in a pressure reading that is not representative of the actual line pressure. | DESIGN: Software Design | Medtronic Perfusion Systems, sent an "Urgent Medical Device Notification letter dated September 9, 2011, via UPS 2-day delivery to all affected customers. The letter identified the affected devices, described the issues, and the actions needed to be taken. The letter indicates when approved, the updated Directions for Use for the Pressure Display Box and Pressure Display Sets will be sent to customers. Customers were asked to complete the Customer Confirmation Certificate and return to Medtronic per the instructions on the form. The letter states that if the recalled products have been further distributed, forward this letter immediately to all customers who received the affected products, and gather proof of notification. For further questions concerning this notification, contact Medtronic Cardiovascular Technical Support at 877-526-7890 or Medtronic Sales Representative. | Contact the recalling firm for information | 74804 | 74,804 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and Virgin Islands (U.S.) and the countries of: Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, El Salvador, France, Georgia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Yemen. | Terminated | October 11, 2012 | 357 | 270088 | Software
Z-0083-2012 | 52105 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager -Software Versions 8.05, 8.06, 8.07, 8.08, and 8.09.It is a stand-alone software package that provides connectivity between laboratory instruments and Laboratory Information Systems, aiding the flow of laboratory information, including test orders and results, specimen information, and patient information. Instrument Manager also provides data management functions that aid in the interpretation and management of laboratory results. The Instrument Manager software is put on a CD which is distributed in a box. User documentation consists of user manuals, PDF copies of which are on the CD Intended to be used by trained laboratory personnel as an aid in laboratory instrument interfacing and laboratory data management. | Instrument Manager software | 2 | October 21, 2011 | 2011 | Data Innovations, Inc. | The "previous result" retrieved from the Specimen Management Database may not be the most recent "previous result". | DESIGN: Software Design | Data Innovations sent an "IMPORTANT PRODUCT NOTICE" dated May 11, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter offers temporary workarounds for customers to use until corrected versions of the software is complete. Versions of the corrections will be provided on a staggered schedule. Customers will be notified by e-mail. Questions regarding this notice are directed to Support at 802-658-1955. |  | 188 | 188 | USA (nationwide) and the country of Canada. | Terminated | October 25, 2011 | 4 | 188 | Software | Output/Calculation | Software update | Software Update
Z-0084-2012 | 59892 | tray, catheterization, sterile urethral, with or without catheter (kit) | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | O'Neil Sterile Field Intermittent Urinary Catheter Kit labeled in mart:"***ORIGINAL CLOSED SYSTEM BY DR O'NEIL O'NEIL STERILE FIELD INTERMITTENT URINARY CATHETER KIT WITH IODINE SWABSTICKS (3)*** REF: ONSM 14KPES SIZE:14Fr Product No: AAM 8014KPES***SINGLE USE ONLY***Original design by: Dr. George O'Neil 200 Churchill Ave, Subiaco, WA 6008, Australia Tel.: +61 (0)8 9388 1700 Fax: +61 (0)8 9382 4952 European Representative: Advena Ltd Hereford, HR4 9DQ, UK DISTRIBUTED IN THE USA BY: American Australian Medical, Inc. 1-888-880-8871***"Case of Catheter Kits labeled in part:"***The O'Neil Sterile Field Urinary Catheter Kit REF: ONSM 14KPES Product No: AAM 8014KPES***CAUTION: Federal law restrict sale of this device to or on the order of a physician. STERILE Single use only***Designed by: GO MEDICAL INDUSTRIES PTY LTD200 Churchill Avenue, Subiaco, WA 6008, Australia Tel.: +61 (0)8 9388 1700 Fax: +61 (0)8 9382 4952 email: gomedic@linet.net.au Website: www.gomedical.com.au***Distributed in the U.S.A. by: AMERICAN AUSTRALIAN MEDICAL INC. BUFFALO, NY 14209 USA TEL: 1-888-880-8871***Country of Origin: ROMANIA***QUANTITY 50***"This kit includes a Urinary Catheter and accessories including gloves, tissue, underpads, and Triad Povidone Iodine Swabsticks (3). These accessories are included in the kit as optional and are not required for urinary catheterization. | O'Neil Sterile Field Intermittent Urinary Catheter Kit | 2 | October 21, 2011 | 2011 | Go Medical Industries Pty., Ltd. | O'Neil Sterile Field Intermittent Urinary Catheter Kits With Povidone Iodine Swabsticks are being recalled due to a recall on the Triad Povidone Iodine Swabsticks contained within the kit. H&P; Industries are recalling the Triad Povidone Iodine Swabsticks because it was determined that the swabsticks were manufactured without having a system for microbial testing in place at the time of release. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Go Medical Industries Pty Ltd, sent an e-mail on September 3, 2011 and follow-up with an "URGENT: MEDICAL DEVICE RECALL" letter dated September 07, 2011 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and not sell the product in the market place; immediately perform a sub-recall and notify their customers and complete and return the RETURN RESPONSE FORM via fax at 716-883-8099 or mail to: American Australian Medical Inc, 1083 Delaware Avenue, Buffalo, NY 14209.For further details you can contact QA Manager at Go Medical Industries at +61 89 381 3226 or pmoretta@gomedical.com.au. |  | 203743 | 203,743 (189,043 English; 14,700 Spanish) | Nationwide distribution. | Terminated | January 30, 2012 | 101 | 235243 | Not_Computer
Z-0104-2012 | 60004 | intervertebral fusion device with bone graft, cervical | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine Zero-P Instruments and Implants. Zero-profile anterior cervical interbody fusion (ACIF) device. | Synthes Spine ZeroP Instruments and Implants | 2 | October 21, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION - Following review of labeling and promotional materials associated with the Zero-P Implant system, it was found that the materials reference a study where PEEK-CR is utilized as a cervical interbody spacer. PEEK-CR is not currently indicated for use as a cervical interbody spacer. | MISBRANDING: Labeling False and Misleading | Synthes (USA) sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated September 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to review the updated supplemental labeling page included with the letter and replace/attach the technique guide page in any Zero-P technique guide that customers may have. The Verification Section of the letter is to be completed and returned to Synthes via fax at 610-251-9005. For questions regarding this notice call 1-800-620-7025 ext. 5403. | Contact the recalling firm for information |  |  | USA (nationwide) including Puerto Rico, Washington, DC and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Open |  |  |  | Not_Computer
Z-0100-2012 | 59837 | burr, orthopedic | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***REF A-CRN***1.85 mm x 16 mm Fluted Router. Use with CRANI-A Attachments.***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only ***Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327-6887 / + 1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126***"Lot numbers: 309600 and 315857.Product Usage: Cutting and shaping bone including bones of the spine and cranium. | Anspach Single Use, Sterile Bone Cutting Burrs | 2 | October 21, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Inc. initiated a recall on May 6, 2008 for their Sterile Bone Cutting Burrs. All affected customers were notified of the recall by letter dated 5/29/2008. | TRAINING: Employee Error | The firm, Anspach sent an "URGENT PRODUCT REMOVAL" letter on May 29, 2008, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Screen their inventory immediately for all products listed on the attached "Product Replacement Form." Upon visual inspection if the pouch is sealed on both ends it is sterile. If one end of the pouch is not sealed it is non-sterile and should be returned for credit or replacement. 2) Complete the "Product Replacement Form" by indicating the quantity returned, and whether the customer would prefer to have their returned devices replaced, or their account credited. 3) Place products to be returned in a box and mark the box with the Return Authorization Number on the "Product Replacement Form." 4) Use the Anspach FedEx account number listed on the form to return the unused devices with the completed "Product Replacement Form." Following receipt of the form, Anspach will process the request immediately. For questions related to the product replacement, please contact Customer Service at 800-327-6887. For questions related to the Voluntary Removal, please contact Anspach Regulatory Affairs and Quality Assurance at 800-327-6887. | Contact the recalling firm for information | 40 | 40 | Worldwide Distribution - USA (nationwide) including the states of: CA, CO, FL, GA, IA, ID, SC, MA, MD, MI, MO, MN, MS, NY, OH, OR, PA, TN, TX, and WI, and the countries of: Africa, Netherlands, Panama, South Korea, Spain, and United Kingdom. | Terminated | October 24, 2011 | 3 | 905 | Not_Computer
Z-0103-2012 | 60047 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hamilton brand Disposable Syringe Pack for use with Microlab F.A.M.E; Part Number: 147692 (case containing 96 syringes); Manufactured by: Treff AG, Degersheim, Switzerland for Hamilton Bonaduz AG, Bonaduz, Switzerland (Manufacturer of record); Product is distributed by Hamilton Company, Reno, NVThe F.A.M.E Syringes are a disposable product that is used on our Microlab F.A.M.E. series instruments for diluting/dispensing. Microlab F.A.M.E was released under 510(k) K943493/S1 under product codes JJE (862.2160 Discrete Photometric Chemistry Analyzer for Clinical Use), JJQ (862.2300 Colorimeter, photometer, or spectrophotometer for Clinical Use), and JTC (866.2500 Microtiter Diluting/Dispensing Device), all of which are currently Class I exempt devices. The syringes are also used on Ortho Summit Processor instruments distributed by Ortho-Clinical Diagnostics. | Disposable Syringe Pack for use with Microlab F.A.M.E | 3 | October 21, 2011 | 2011 | Hamilton Co | Certain lots of syringes have caused an increase in error message "DIS541", which indicates "improper dispensation". | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Hamilton, sent an "Urgent: Product Recall Notification" letter dated September 26, 2011 to its customers. The letter describes the product, problem and action to be taken. The customers were instructed to return the syringes with a completed Hamilton Company Product Recall Questionnaire form via fax to: (775) 856-7259 or with an emailed signed pdf copy to wayne.mcauliffe@hamiltoncompany.com. The firm's Customer Service Department will contact customers with complete Returned Material Authorization information. If you have any questions, please do not hesitate to contact the Director of Quality at (775) 858-3000. |  | 1096 | 1096 units | Worldwide distribution: USA(nationwide) including states of: NJ, NV, TX and WA; and countries of: Brazil, P.R. China, France and Germany. | Terminated | March 06, 2012 | 137 | 1096 | Not_Computer
Z-0082-2012 | 60054 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Smith & Nephew, HD1200 Autoclavable Camera HeadCatalog Number: 72203360Product Usage: Endoscope & Accessories | ENDOSCOPIC Camera | 2 | October 21, 2011 | 2011 | Smith & Nephew, Inc. Endoscopy Division | HD1200 camera heads may cause image issues |  | Smith Nephew sent an "URGENT  PRODUCT RECALL NOTIFICATION" letter dated sent September 23, 2011 to all affected consignees. The letter identified the product, problem and actions to be taken. Consignees were instructed to complete requested information, return affected product and order replacement product. To return product and order replacement product, consignees were instructed to contact the Smith & Nephew Endoscopy Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. For questions call 405-917-8556. |  | 15 | 15 units | Nationwide Distribution (USA) including the states of: MI, NH, OK. | Open |  |  | 15 | I/O | Display/Image | Replace | Remove or Replace
Z-0080-2012 | 59984 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | "***Vipat REF H0020-100***Lot 100533***2015-04***STERILE EO Elastomeric pump***2mL/Hr***100mL***Medpro International (Thailand) Ltd. 360, Moo 6, Chonburi Industrial Estate, Tambon Bowin, Amphur Sriracha, Chonburi 20230, Thailand***Victus Inc. 4918 SW 74 Court 33155 44400***"Packed 24 units per case.The devices are intended for patients requiring intravenous, percutaneous, subcutaneous, intra-operative sites or epidural administration of medications. | Elastomeric Infusion Pump | 1 | October 24, 2011 | 2011 | Victus, Inc. | Victus, Inc. Miami, FL is recalling Vipat Brand Elastomeric Infusion Pump, Model H0020-100, Lot 100533. The product may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to patients. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Victus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 15, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use, remove the products from inventory and quarantine the products. Additionally, the firm indicated that their customer service department would contact customers to coordinate the return of affected products. Customers are directed to call 305-663-2129 ext. 110 for questions regarding this notice. | Contact the recalling firm for information | 63 | 63 units | USA (nationwide) including Puerto Rico and the states of FL and PA. | Terminated | January 12, 2012 | 80 | 63 | Not_Computer
Z-0081-2012 | 59981 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro AccuFlo Elastomeric Infusion Device, Model #CT-0020-270C, packaged 12/box. The box label shows the product is Manufactured By: Medpro International (Thailand) Ltd. for Medpro Corporation Pte. Ltd., Singapore.The devices are intended for patients requiring intravenous, percutaneous, subcutaneous, intra-operative sites or epidural administration of medications. | Medpro AccuFlo Elastomeric Infusion Device | 1 | October 24, 2011 | 2011 | Progressive Medical Inc | The product may have a higher flow rate than specified resulting in over administration of drug solutions to the patients. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Progressive Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 28, 2011 via email and regular mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue their use; remove the products from inventory and quarantine them; inform their customers to return unused products to Progressive; return the product, and complete and return the RECALL RETURN RESPONSE FORM via fax to 314-961-5786. Progressive customer service will contact the customers to coordinate the return of the goods in question.Should you have any questions, please call 800-969-6331 or email to dsullivan@progressivemedinc.com. | Contact the recalling firm for information | 180 | 180 units | Nationwide distribution: USA including states of: FL, KY, PA, VA and WY. | Terminated | December 19, 2011 | 56 | 180 | Not_Computer
Z-0105-2012 | 60017 | microscope, surgical | EPT | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | OPMI VARIO 700 Surgical Microscope.Manufactured by: Carl Zeiss Meditec AG,Oberkochen, Germany.Provide a magnified view of the surgical field. | OPMI VARIO 700 Surgical Microscope | 2 | October 24, 2011 | 2011 | Carl Zeiss Meditec, Inc. | An incorrect nut and bolt set was used in the axis three (3) spring assembly of the OPMI VARIO 700 surgical microscope distributed to three US locations. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carl Zeiss Meditec had Field Service Engineers directly contact all affected customers. The Field Service Engineers will replace the defective component. Questions regarding this Field Corrective Action should be directed to the Director of Clinical/Regulatory Affairs at 925-557-4616. | Contact the recalling firm for information | 3 | 3 | USA (nationwide) including the states of FL, GA, and LA . | Terminated | November 04, 2011 | 11 | 3 | Not_Computer
Z-0106-2012 | 58984 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | "***Assembled for Invivo Corporation 12501 Research Parkway Orlando FL 32826***Made in U.S.A.***MRI Patient Monitoring System Model 865214***Use only with Model 989803169201 Power Adapter***REF 453564155341***SN US11401652***".Note: REF 453564155341, REF 453564181201, and 453564180091 are subject to recall. REF is the firms Service Number which identifies the device. Product Usage The Expression MRI Patient Monitoring System is intended for use by healthcare professionals to monitor vital signs for patients undergoing MRI procedures. The Patient Monitoring System also provides signals for synchronization of the MR scanner. | MRI Patient Monitoring System | 2 | October 25, 2011 | 2011 | Invivo Corporation | Philips Invivo Corporation of Orlando, FL is recalling their Expression MRI Patient Monitoring System. During transport of the system's cart with a docked Display Controller Unit (DCU), a front caster (wheel) may become loose and fall off causing the cart to become unbalanced and fall. | PRODUCTION CONTROLS: Process Control | Philips Invivo Corporation issued a Field Safety Notice "URGENT - MEDICAL DEVICE RECALL" letter to all affected customers notifying them that the Expression MRI Patient Monitoring System could pose a risk for patients and users. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the user in order to prevent risks for patients or users and the actions planned by Philips Invivo to correct the problem. This notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. Customers were instructed to contact Philips Invivo Customer Support at (877) 468-4861 (option 1, then option 3) for questions or support concerning this issue. |  | 285 | 285 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI., SC, SD, TN, TX, VA, WA, and WI and the countries of: Australia, Austria, Canada, Denmark, Germany, Israel, Malaysia, Netherlands, Norway, Singapore, Spain, Switzerland and, United Kingdom. | Terminated | November 15, 2012 | 387 | 285 | Not_Computer
Z-0089-2012 | 59969 | pump, infusion, enteral | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | CORFLO 300 Enteral Pump; a volumetric rotary peristaltic intermittent enteral feeding pump; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #20-0500.Intended to deliver formula to patients requiring tube feedings. | CORFLO 300 Enteral Pump | 2 | October 25, 2011 | 2011 | CORPAK MedSystems | Corpak is removing the Corflo 300 Enteral Pump from the market due to the possibility that with fluid ingress or other damage, the keypad may not function, and the possibility of motor mount failure. | PRODUCTION CONTROLS: Process Control | Corpak MedSystems sent Recall Notification letters dated September 19, 2011 to all CorFlo 300 Enteral Pump customers on the same date. The customers were informed that there is a possibility that with fluid ingress the keypad may become inoperable, and that there is a possibility of motor mount failure. Either event could possibly affect the rate at which formula is delivered to the patient, and fluid ingress could cause a potential of audible alarm failure. Corpak determined that the removal of the pumps from the market was the best option for their customers, and requested the return of all CorFlo 300 pumps to Corpak by no later than December 1, 2011. The customers were requested to contact Corpak at 1-800-323-6305 to arrange for the return of the pumps, and to complete and return the enclosed acknowledgement form, indicating the number of pumps on hand, via fax at 847-541-9526 or e-mail to questions@corpak.com. Dealers/distributors were requested to notify their customers of the recall. | Contact the recalling firm for information | 13000 | 13,000 pumps | Nationwide Distribution | Open |  |  | 13000 | Not_Computer
Z-0115-2012 | 59326 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology |  | GE SENOGRAPHE DS FULL FIELD DIGITAL ESSENTIAL MAMMOGRAPHY SYSTEM.The Senographe DS system generates digital mammographic images, which can be used for screening and diagnosis of breast cancer. The Senographe DS is intended to be used in the same clinical applications as traditional mammographic film/screen systems. | GE SENOGRAPHE DS FULL FIELD DIGITAL ESSENTIAL MAMMOGRAPHY SYSTEM. | 2 | October 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurement on GE magnified images provided by Senographe DS and Essential Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | DESIGN: Software Design | GE Healthcare sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated September 21, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers. The letter provides safety instructions and advises customers to complete and return a Customer Reply Form via fax to +33(0)1 30 70 39 30. A GE Healthcare service representative will contact customers if the updated software version is required. Contact the Call Center at 800-437-1171. |  | 1344 | 1344 | Worldwide Distribution-USA (nationwide) including Puerto Rico and Washington, D.C. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA,GERMANY, GREECE, GUADELOUPE, GUYANA, HONDURAS, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF KOREA, RUSSIAN FEDERATION SAINT PIERRE AND MIQUELON, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE FORMER YUGOSLAV REPUBLIC OF MECEDONIA, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VENEZUELA. | Terminated | October 16, 2012 | 356 | 3457 | Software
Z-0112-2012 | 58454 | mask, oxygen, low concentration, venturi | BYF | Venturi mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mask adult, elastic strap style, Percent-O2 Lock air entrainment system, w/7' tube (Linde) - 50/cs;Salter Labs, Arvin, CA 93203Intended usage: respiratory therapy. | Mask adult, elastic strap style, PercentO2 Lock air entrainment system | 2 | October 26, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | A small number of oxygen mask components were found to have 100% obstructed airways rendering them incapable of delivering oxygen to patients undergoing respiration therapy. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Salter Labs, called all their customers on March 25, 2011 and asked that they hold all affected product. In addition, the firm sent an "URGENT PRODUCT RECALL" letter dated April 4, 2011 and a follow-up letter dated April 7, 2011 via e-mail and certified mail to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory, return any affected lots listed, contact Salter Customer Service at 1-800-235-4203 to arrange for return of affected product and complete and return the Customer Reply Form via 661-854-6816 or toll free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 800 | 800 all varieties | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Colombia and Iceland. | Terminated | October 26, 2011 | 0 | 800 | Not_Computer
Z-0107-2012 | 59522 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE HEALTHCARE, MSK 1.5T EXTREME MR SCANNER, MODEL AA5000.The MSK Extreme TM MR Scanner is intended for use as a diagnostic imaging device to produce axial, sagittal, coronal and oblique images of the internal structures of the leg (excluding the thigh), knee, ankle, foot, forearm, elbow, wrist and hand. | GE Healthcare, MSK 1.5T Extreme MR Scanner Model AA5000 | 2 | October 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware that an exposed high voltage gradient wire internal to the MSK 1.5T Extreme system equipment may cause electrical shock, which may impact service personnel safety. This issue does not impact patient or operator safety. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare LLC sent a "Urgent Medical Device Correction letter" dated August 29, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to please ensure that all potential service personnel at your facility are made aware of this safety notification and the recommended actions.Contact your local service representative if you have any questions concerning this notification. For further questions please call ( 262 ) 513-4122. | Contact the recalling firm for information | 52 | 52 | Worldwide Distribution -- USA (nationwide) including the states of AL, CA, FL, ID, NY, NH, NJ, NV, OH, PA, TX, VA and WI., and countries of UNITED ARAB EMIRATES, UKRAINE, UNITED KINGDOM, SWITZERLAND | Terminated | May 29, 2012 | 216 | 52 | Not_Computer
Z-0074-2012 | 59933 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator, Model 19250-001. Manufactured by: CareFusion, 22745 Savi Ranch, Yorba Linda, CA 92887 USA. EnVe is a portable critical care ventilator. It is suitable for a wide range of life support and critical care applications. The ventilator normally operates from external DC power. When the ventilator is portable, a Removable Battery Pack powers the unit. Devices are designed for use on patients who require respiratory support or mechanical ventilation and weigh a minimum of 5 kg (11 lbs). | EnVe Ventilator | 1 | October 26, 2011 | 2011 | CareFusion 203, Inc. | CareFusion has identified potential risks associated with the EnVe ventilator manufactured between December 2010 and May 2011. There are 3 issues that can result in ventilation delivery to the patient being interrupted. CareFusion will be contacting facilities to coordinate implementation of the necessary corrective actions. | DESIGN: Device Design | The firm, CareFusion, sent an "URGENT PRODUCT RECALL" letter dated September 12, 2011 to its customers.The letter identified affected product, described the 3 issues providing the description of the event, potential risk, and provided required action for Users. The letter also stated that CareFusion would be contacting the facility to coordinate the corrective action.If you have any questions and/or need support, please contact CareFusion Technical Support at 1-800-554-8933 from Mon. - Fri. 8:00am -5:00pm Pacific Time. |  | 128 | 128 units | Nationwide distribution: USA including states of: CA, CO, FL, IL, MI, MN, MT, ND, NJ, NM, NY, OK, PA, TX, UT, WI, and WV. | Terminated | November 22, 2012 | 393 | 128 | Not_Computer
Z-0124-2012 | 59147 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON CX4 Clinical System Part #: 758300Product Usage: The Beckman SYNCHRON CX4 DELTA, CX7 DELTA, CX4CE, and CX7 systems are fully automated and computer controlled instruments designed for the in vitro diagnostic quantitation of biological fluid components and therapeutic drugs as well as the qualitative determination of drugs of abuse in urine | SYNCHRON CX4 | 2 | October 27, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed software limitations which affect only manually-loaded reagent cartridges on the CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. | DESIGN: Software Design | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 20, 2010 with an attached Fax Back Response Form (via US Postal Service for US Customers) to all customers who purchased the SynchronÂŽ CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. The letter provides the customers wtih an explanation of the problem identified and actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected product listed above to another laboratory, they were asked to provide a copy of this letter to them. Questions regarding this Product Corrective Action letter, were directed to Customer Support Center (Call Center) on the Beckman Coulter website athttp://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. Outside the United States and Canada, they were directed to their local Beckman Coulter Representative. |  | 1070 | 1070 systems total in US | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | June 28, 2012 | 245 | 1070 | Software |  | Instructions | Safety Notice/Insructions
Z-0119-2012 | 60041 | first aid kit with drug | LRR |  |  | General & Plastic Surgery | Enforcement Discretion | Base Camp First Aid Kit, Part Number 4150, UPC 0-91966-04150-3, contains 171 pieces.The kit is labeled in parts: "***BASE CAMP FIRST AID KIT***171 PIECE***6 IODINE PREP PADS***Lifeline First Aid, LLC***Wilsonville, OR 97070, USA***www.lifelinefirstaid.com***".First aid kit. | Base Camp First Aid Kit | 2 | October 27, 2011 | 2011 | Lifeline First Aid LLC | The first aid kits are recalled because the kits contain Iodine Prep Pad single use wipes which were recalled by Triad. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Lifeline First Aid, LLC sent a " PRODUCT RECALL NOTICE-LIFELINE FIRST AID KITS WHICH INCLUDE H & P [TRIAD] PVP IODINE WIPES" letter dated July 7, 2011 to all affected customers. The letter states that Lifeline First Aid, LLC has become aware of a recall of Povidone Iodine Wipes manufactured by H&P; Industries Inc. (TRIAD GROUP) and marketed under various brand names, including Triad and Swift. The recall of the Povidone Iodine Wipe is due to potential microbial contamination of the product. The letter instructs customers to collect and quarantine all affected products and to complete and return an attached Customer Response Form via fax to (503) 783-3721.Affected products will be replaced without charge and acceptable Lifeline first aid kits will be re-worked and returned using Lifeline's standard RGA process. Customers may call Lifeline customer service at 1 (877) 377-8243 for further instructions. | Contact the recalling firm for information | 6072 | 6072 kits | Nationwide Distribution | Terminated | March 29, 2012 | 154 | 26150 | Not_Computer
Z-0117-2012 | 60064 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Endura Fluoroscopic Image Intensified X-ray Catalog Number: 718074 Product Usage: Fluoroscopic x-ray system, image-intensified Fluoroscopic x-ray system. | Philips BV Endura | 2 | October 27, 2011 | 2011 | Philips Healthcare Inc. | Front plate broke in to two pieces causing the C-Arm to suspend on the single central bolt | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare Inc., sent an "URGENT - Field Safety Notice" dated September 30, 2011 to all affected customers. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the customer/user in order to prevent risks for patients or users and the actions planned by Philips Medical Systems to correct the problem. The notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. The notice stated that Philips will replace the front plates for systems on the Unit Affected List by the end of June 2012. The notice instructed customers to contact Philips Healthcare representative for further information or support. |  | 12 | 12 units in the US | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Austria, Belgium, Bharain, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Netherlands, Norway, Portugal, Russian Fed., Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Open |  |  | 266 | Not_Computer
Z-0113-2012 | 60025 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Radio used in the Precess MRI Patient Monitoring System. 3160 (Precess) and HE97B, 94044, 94045 (radio and radio kits).Label located on the back of the WPU. Product is labeled in part: "***Invivo Corporation 12601 Research Parkway, Orlando, FL 32826***Made in U.S.A.***REF 3160***SN***CAUTION: Law in U.S.A. Canada and EU restricts this device to sale by or on the order of of a licensed medical practitioner only.***" Label located on the bottom of the DCU. Product is labeled in part: ***Invivo Corporation 12601 Research Parkway Orlando, FL 32826 Made in U.S.A.***REF 3160DCU***SN***CAUTION: Law in U.S.A. Canada and EU restricts this device to sale by or on the order of of a licensed medical practitioner only.***" Display Controller Unit (DCU) is labeled in part: "***Invivo***PRECESS***". Wireless Processing Unit (WPU) is labeled in part: "***Invivo***". NOTE: WPU AND DCU MAY HAVE THE SAME OR DIFFERENT SERIAL NUMBERS.The radio is responsible for wireless communication between the Wireless Processing Unit and the Display Controller Unit. Precess is intended for use by healthcare professionals to monitor vital signs for patients undergoing MRI procedures. | Precess MRI Patient Monitoring System | 2 | October 27, 2011 | 2011 | Invivo Corporation | Philips Invivo has determined that it is possible for the system's alarm limits and other device settings to spontaneously change without user intervention. This problem is associated with intermittent and frequent loss of communication between the wireless Processing Unit and the Display Controller Unit. | DESIGN: Device Design | A Field Safety Notice entitled Urgent - Medical Device Correction letter, dated September 2011, with attached Field Safety Notice for customers were sent to distributors. The letter described the reason for the recall and the intention of the Field Safety Notice. Also, the hazards involved and instructions on how to identify units that are affected by the recall were provided . The actions to be taken by the user and the actions planned by the Philips Invivo were explained. Users were to follow the instructions provided on how to safely use the device until the unit can be serviced by Philips Invivo. For further information or support concerning this issue, customers are to contact Philips Invivo Customer Support at (877)-468-4861 or their local Philips Invivo representative. |  | 875 | 875 | Worldwide Distribution -- Australia, Austria, Brazil, Canada, China, France, Germany, Malaysia, Norway, Portugal, Spain, Sweden, The Netherlands, United Arab Emirates, United Kingdom, and US. | Terminated | August 07, 2013 | 650 | 875 | I/O | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0073-2012 | 59536 | table, surgical with orthopedic accessories, ac-powered | JEA | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Mizuho OSI Modular Table System;Product Usage: For use in orthopaedic surgery to provide support to a patient. | Modular Table System | 1 | October 27, 2011 | 2011 | Mizuho Orthopedic Systems Inc | Reports of injury related incidents while using the Mizuho OSI Modular Table System. |  | Mizuho Orthopedic Systems Inc sent a "Field Advisory Notice (Recall)" letters dated July 29, 2011 to all customers using the Modular Table System via Federal Express. The Field Advisory Notice reference the devices involved and explain that there is no need to return any product or to discontinue the use. The letter provides WARNINGS and RECOMMENDATIONS for all users to consider when using the devices and request that the response form be completed as instructed in the advisory communication. The letters also recommend that awareness, training, inspection prior to use, and proper maintenance be established and practiced. The letter recommend that the Field Advisory Notices be posted in prominent areas and communicated to all staff and potential users of this device. Additional support and training can be obtained by contacting Mizuho OSI's Customer Resource Group at 1-800-777-4674. For further questions concerning this notification, contact the Director of Customer Resource Group for assistance (1-800-777-4674, outside the USA 00+1-510-476-8199).Revised letters were approved by CDRH and sent out in January 2012. | Contact the recalling firm for information | 15011 | 15,011 units | Worldwide Distribution - Nationwide USA | Terminated | August 27, 2013 | 670 | 15011 | Not_Computer
Z-0135-2012 | 60175 | system, gastrointestinal motility (electrical) | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AR ManoShield, MSS-3599, 10 pack.FGS: MSS-3599,Individual package label: LBL-110275-03,Box label: LBL-110282-03,Package Inserts: DOC-3682,User Manual: DOC-1332-12.Product Usage: The MVS disposable sheath (a.k.a. ManoShield) for the high-resolution MVS catheter is designed to provide a biocompatible barrier between the catheter and patient. | Sierra Scientific Motility Visualization System (MVS) | 2 | October 28, 2011 | 2011 | Sierra Scientific Instruments Inc | The recall was initiated after Given Imaging received a test report during a Canadian licensing process where the AR ManoShield exhibited pinholes and did not pass the test, the Field Safety Corrective Action (FSCA) Committee was convened. As a result of the data discussed during this meeting, a decision was made to voluntarily recall this product, and cease its further commercialization or distr | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Given Imaging sent an "Urgent Product Recall 1st Notice" letter with an attached Device Recall Confirmation Form (via Fed Ex with signature for tracking purposes) to all customers who purchased the AR ManoShield, part numbers MSS-3599. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to locate the affected product and removed from use immediately. The letter stated that if the affected product have been further distributed, forward this letter immediately to all customers who received the affected product. Customers were asked to complete and return the enclosed confirmation form as per the instructions. Customers with any questions were instructed to contact Karla Guerrero at (310) 641-8492 ext. 2138 or email at karla.guerrero@givenimaging.com. |  | 292 | 292 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Chile, China, Columbia, Ecuador, France, Germany, Israel, Mexico, Poland, Portugal, Sweden, and Switzerland. | Terminated | August 30, 2012 | 307 | 292 | Not_Computer
Z-0133-2012 | 59521 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, InnovaÂŽ 3100IQ (Optima Edition) Fluoroscopic Vascular X-ray System.Indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational imaging-procedures. | Innova 3100IQ (Optima Edition) Fluoroscopic Vascular Xray System | 2 | October 28, 2011 | 2011 | GE Healthcare, LLC | A potential weakness on the monitor suspension end-stop device may impact user safety. The end-stops may break if the monitor-suspension is repeatedly pushed with excessive force on the end rails. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A GE Healthcare "Urgent Medical Device Correction" letter dated August 22, 2011 was sent out 8/24/11. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology & Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were told to operate the suspension without excessive force. If one of the end-stops breaks, then the system should not be used and their local GE Healthcare Service Representative should be contacted. GE will contact customers to replace all end-stops at no cost to the customer. Questions or concerns should be directed to 800-437-1171 in the US and 0120-055-919 in Japan. For other countries, customers can contact their local GE Healthcare Service Representative. | Contact the recalling firm for information | 15 | 15 | Worldwide Distribution -- USA, including the states of TX & NC and the countries of TURKEY, POLAND, MEXICO, CHINA, CAMBODIA, BRAZIL, BOLIVIA, and SAUDI ARABIA. | Terminated | May 29, 2012 | 214 | 15 | Not_Computer
Z-0136-2012 | 59866 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | SwitchPoint Infinity Control SystemStryker Communications, 1410 Lakeside Parkway, # 100, Flower Mound, TX 75208The SwitchPoint Infinity (SPI) Control System is an integrated voice, video, data router and teleconferencing interface for operating rooms. The SwitchPoint Infinity is used for direct control of room equipment, audio/video equipment, and indirect control of surgical equipment in the operating room. | SwitchPoint Infinity Control System | 2 | October 28, 2011 | 2011 | Stryker Communications Corp | Installation records for the SwitchPoint Infinity, Booms, Lights, and Flat Panel systems are incomplete or inadequate. | PRODUCTION CONTROLS: Process Control | Stryker sent an URGENT FIELD SAFETY NOTICE: RA2010-245 dated August 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Customer Response Form within 5 business days to Stryker. A Stryker representative would contact them at their convenience to schedule a re-inspection for all of the units identified and still in use. In the event the inspected equipment does not meet Stryker's established specifications, their representative will visit their facility and hopefully resolve the issue. If a subsequent visit is required the date will be scheduled in full consideration of their services.For questions regarding this recall call 972-410-7310. | Contact the recalling firm for information |  |  | Worldwide Distribution including Puerto Rico, Brazil, Ecuador, Venezuela, Costa Rico, Bolivia, Peru, Mexico, China, India, Korea, Philippines, Singapore, Malaysia, Taiwan, Middle East, Russia, Africa, Poland, South Africa, Australia, New Zealand, Canada, Japan, Denmark, Finland, Norway, Sweden, Belgium, The Netherlands, France, Germany, Austria, Italy, Portugal, Spain, Romania, Switzerland, and United Kingdom. | Terminated | August 20, 2012 | 297 |  | Not_Computer
Z-0134-2012 | 60168 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6301 Dual Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8064.The Flo-Gard Volumetric Infusion Pumps are an electromechanical pump used for intravenous infusion of fluids at user-selected rates. | Baxter FloGard 6301 Volumetric Infusion Pump | 2 | October 28, 2011 | 2011 | Baxter Healthcare Corp. | The Flo-Gard infusion Pump was released to the customer with OOS values for the Air Sensor Calibrations with the Air in Tubing test. | TRAINING: Employee Error | A Baxter Healthcare technician initially notified the sole customer on August 26, 2011 of the defective product and requested that the product be quarantined. On September 14, 2011 a technician visited the site to correct the problem observed on the Flo-Gard Infusion Pump. All required tests were performed and the results were found acceptable. |  | 1 | 1 unit | Distributed only in Puerto Rico. | Terminated | October 31, 2011 | 3 | 1 | Not_Computer
Z-0140-2012 | 59909 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act:a) Split Dressing - 2"x 2", Yellow, component C-033-0022, vendor part number S1072; b) Tubing Extension Sets and IV Sets component C-FG711297CP, vendor pn FG711297CP;component C-FG711093CP, vendor pn FG711093CP; component C-FG711298CP, vendor pn FG711298CP; component C-FG711092CP, vendor pn FG711092CP; c) Cast Padding component C-9046, vendor pn 9046; component C-9044, vendor pn 9044;component C-9043, vendor pn 9043;component C-9062, vendor pn 9062;d) Scalpel Holder component C-DD1628-6, vendor pn DD1628-6; e) Latex Tubing component C-806-12, vendor pn 806-12;component C-808R, vendor pn 808R;component C-606R, vendor pn 606R;component C-122BA10, vendor pn 122BA10;component C-122BA12, vendor pn 122BA12;component C-122BA07, vendor pn 122BA07;component C-102BA12, vendor pn 102BA12;component C-099BA10, vendor pn 099BA10;component C-102BA10, vendor pn 102BA10;component C-103BA10, vendor pn 103BA10;component C-124BA10, vendor pn 124BA10;component C-140BA10, vendor pn 140BA10;f) Clinic Dropper component C-12548, vendor pn 10-871-00; g) Tongue Blade (depressor) component C-11820-010, vendor pn 07-920-00; h) Catheter Clamp component C-043-0497, vendor pn 033403-887;i) Gillette Good News Razorcomponent C-P4417A, vendor pn 4417GN; j) Prepsicle Round Short Blue Swabstickcomponent C-051-0401, vendor pn SBR20035;k) Tubingcomponent C-SFM3-3050, vendor pn SFM3-3050;component C-SFM3-3650, vendor pn SFM3-3650;l) Specimen Containercomponent C-B9021YN, vendor pn B9021YN;component C-C8846-18, vendor pn 15706-806;The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then sterilized and are ready to be used by the customer for the specific procedure. | Presource Custom Sterile Surgical and Procedure Kits | 2 | October 31, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturer(s) does not have all necessary FDA registrations or filings needed for this product. | DESIGN: Component Design/Selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 2, 2011 to all affected consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. If the consignee does not wish to accept overlabeled product or to overlabel the product in their possession, they were asked to contact Presource Sales Operations at 800-766-0706 or their Sales Representative for further instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact the following: Hospitals - Presource Sales Operations at 800-766-0706, Distributors - Distributor Management Group at 800-635-6021, and Federal Government facilities - Government Service at 800-444-1186.For questions regarding this recall call 847-887-6412. |  |  |  | Nationwide Distribution | Open |  |  |  | Not_Computer
Z-0139-2012 | 60006 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Picis Anesthesia Manager, PACU Manager,Critical Care Manager-8.2 Service Pack 4 or higher (Software, transmission & storage, patient data)This product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources, such as health care professionals, medical devices connected to the system, and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | Picis Anesthesia Manager | 2 | October 31, 2011 | 2011 | Picis Inc. | Periop/Critical Care Manager EHR Software Application displayed one patient's demographic and health information in place of another patient that was in current view. | DESIGN: Software Design | Picis Inc. sent a "Safety Notification letter" dated August 12, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed until this service pack is applied, we recommend: * For patients who have been transferred from other facilities within the same VISN, you should discontinue the use of these Picis applications. Manually check the patient's name with that displayed in the Picis patient blue banner to ensure you catch all of these cases. * For all other patients, you may safely continue the use of these Picis applications.For further questions please call (781) 557-3000. | Contact the recalling firm for information | 10 | 10 | Nationwide Distribution -- AZ, DE, FL, IL, NY, OH, PA, WI and WV. | Terminated | March 23, 2012 | 144 | 10 | Software | Output/Calculation | Software update | Software Update
Z-0142-2012 | 59582 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 including all Catalog Numbers to include: Catalog # 500AVHCT with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. Catalog Number 500AHCT, CDI 500 Monitor with Arterial blood parameter module and Hematocrit/Saturation probe. Catalog Number 500AV CDI 500 Monitor with Arterial and Venous blood parameter modules; and Catalog 500A, CDI 500 Monitor with Arterial blood parameter module.Intended for use during cardiopulmonary bypass procedures when continuous monitoring of arterial and /or venous pH, PCO2, K+, or oxygen saturation, hematocrit, hemoglobin, and temperature are desired. | CDI 500 Monitor | 2 | October 31, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS identified a defective circuit board, which could potentially cause the CDI monitor to not function after a power loss. Once power is lost, the monitor will not power up again, even if AC or battery power are restored. | DESIGN: Component Design/Selection | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" notice dated September 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides information on what to do in the event of a failure of the device. Additionally, a Customer Response Form was attached to the letter for customers to complete and return. Terumo CVS will schedule the return of affected products for inspection and replacement of defective parts. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. | No consumer action necessary | 432 | 432 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Chile, Japan, Philippines, and Thailand. | Terminated | April 11, 2013 | 528 | 432 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0143-2012 | 55537 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | Bioplate Titanium Fixation System, Sterile Kit.The sterile plate and screw kit configurations for use with the Bioplate Titanium Fixation System are intended for use on non-load bearing fixation, including cranial bone fixation and brow fixation and the treatment of fractures and reconstructive procedures of the craniomaxillofacial skeleton. Each device is intended for single use only and only in conjunction with other titanium and titanium alloy implant. | Bioplate Titanium Fixation System, Sterile Kit. | 2 | November 01, 2011 | 2011 | Bioplate Inc | There is potential defect in seal of the sterile barrier containing the device. | DESIGN: Process Design | BIOPLATE sent an URGENT - PRODUCT RECALL NOTIFICATION letter dated March 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to return any recalled product they have at their siteand if distributed further they should forward the information contained in the letter to all customers. If a device from one of the lots noted above has already been used, Bioplate would like to be advised of that use.They asked that consignees complete the enclosed Fax form and transmit the requested information to them as soon as possible at (310) 815-2126.Returned recalled product was to be sent to: Bioplate3643 Lenawee AvenueLos Angeles, CA 900163643 Lenawee Avenue, Los Angeles, CA 90016For any questions regarding this recall call (310) 815-2100. |  | 18246 | 16,020 Individual Kits; 2,226 Boxes. | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, Japan, Taiwan, Middle East and Australia. | Terminated | April 04, 2012 | 155 | 18246 | Not_Computer
Z-0144-2012 | 60057 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BardÂŽ Ventralex" ST Medium Circle with Strap Hernia Patch Diameter 6.4cm /2.5", Catalog Number: 5950008.Intended for use in the reconstruction of soft tissue, where weakness exists, in procedures involving soft tissue repair, including repair hernias and deficiencies caused by trocars. | Bard Ventralex" ST Medium Circle with Strap Hernia Patch | 2 | November 01, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Ventralex" ST Medium Circle was packaged in a pre-printed carton identified for a BardÂżVentralex" ST Large Circle. | PRODUCTION CONTROLS: Packaging Process Control | Davol Inc. issued an Urgent; Medical Device Recall letter on 10/4/11 for the mislabeled Bard Ventralex ST Medium Circle. The letter identified the affected product and explained the reason for recall. Customers are to immediately examine their inventory and quarantine any affected product. In addition, if any product was further distributed then those customers should be notified of the recall as well. Accounts are to contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. The attached recall effectiveness check form should be completed and returned. Questions should be directed towards Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services & Support Department at 1-800-562-0027. |  | 96 | 96 units | Nationwide Distribution | Terminated | September 20, 2013 | 689 | 96 | Not_Computer
Z-0146-2012 | 58162 | orthopedic stereotaxic instrument | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317.RIO System - The Tactile Guidance System v2.0 is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The Tactile Guidance System v2.0 is indicated for use in surgical knee procedures, in which the use of stereotactic surgery may be approriate, and where reference to rigid anatomical bony structures can be identified relative to a CT base model of the anatomy. These procedures include unicondylar knee replacement and/or patellofemoral knee replacement. | RIO Robot Arm Interactive Orthopedic System | 2 | November 01, 2011 | 2011 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. | Contact the recalling firm for information | 70 | 70 | Worldwide distribution: USA including states of: AZ, CA, CO, FL, GA, IA, IL, LA, MI, MO, MS, NC, NJ, NV, NY, OK, OH, OR, PA, RI, TX, UT, VA, WA, WI, and WV; and countries of: Scotland, Italy and Korea. | Terminated | December 13, 2011 | 42 | 70 | Software | Physical Safety Hazards | Software update | Software Update
Z-0145-2012 | 60046 | implant, dermal, for aesthetic use | LMH |  |  | General & Plastic Surgery |  | MERZ AESTHETICS, RadiesseÂŽ Volume Advantage 1.5 CC, Injectable Implant, Part Number 8071M0H1 and 8071M0K1, Sterile, Rx ONLY.Part Number 8071M4K1 distributed in CanadaPart Number 8071M5 distributed in Europe/Middle East/AfricaPart Number 8071M9 distributed in Hong Kong, Korea, Malaysia. | Radiesse | 2 | November 01, 2011 | 2011 | Merz Aesthetics, Inc. | There is a potential leakage of product that may occur with Merz Aesthetics 1.5cc syringes of Radiesse Dermal Filler. When using the syringe, Health Care Professionals should be alerted to the potential for product that my leak pass the plunger at the proximal end of the syring. The defect may result in infection requiring medical intervention to prevent impairment or damage. To date Merz has n | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 8/18/11, consignees were sent a Merz Aesthetics "Urgent Medical Device Recall" letter dated August 15, 2011. The letter was addressed to "Dear Health Care Professional". The letter described the product and the problem and listed the details regarding this issue. | Contact the recalling firm for information | 311582 | 311,582 | Worldwide distribution including USA, Puerto Rico, Dominican Republic, Hong Kong, Israel, Netherlands, Germany, South Korea, and Canada. | Terminated | May 01, 2012 | 182 | 311582 | Not_Computer
Z-0147-2012 | 60240 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad(R) Sterile Disposable Stellant Syringe Kit, Catalog # SDS-CTP-SPK, Lot # 112865.Angiographic syringe intended to be used for the specific purpose of injecting intravenous contrast media into humans for diagnostic studies in computed tomagraphy. | Medrad(R) Sterile Disposable Stellant Syringe Kit | 2 | November 02, 2011 | 2011 | Medrad Inc | Packaging may exhibit open seals - sterility compromised | PRODUCTION CONTROLS: Packaging Process Control | Medrad sent a " Urgent Medical Device Recall letter" dated October 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed if they have any of the affected product lot, please contact MEDRAD customer support to make arrangements for the return of the product. Call 1 -800-633-7231, select option 5, then select option 1. Indicate that you are calling about the affected syringe kit recall.Complete and sign the attached response form and fax it back to MEDRAD at (412) 406-0942.Please complete and fax this form even if you do not have any of the referenced product so we can maintain proper accountability for all syring kits that have been shipped. |  | 17380 | 17,380 | Nationwide Distribution -- including the States of : AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, VS, WA, WV, and WI. | Open |  |  | 17380 | Not_Computer
Z-0149-2012 | 60034 | catheter, percutaneous, cardiac ablation, for treatment of atrial fibrillation | OAE |  |  | Cardiovascular |  | Medtronic Arctic Front Catheter, Sterile EO, Medtronic CryoCath LP, Kirkland, Quebec, H9H 5 H3, Canada, Model 2AF282 and 2AF231 ( not distributed in US)Product Usage: The Medtronic CryoCath System consists of a Cryoablation console, cardiac Cryoablation catheters, connection components and accessories. The Arctic Front Cardiac Cryoablation Catheter is an over the wire Balloon catheter. The inside of the Balloon at the distal end of the Cryoablation Catheter reaches temperatures of -50ÂŽC or colder when refrigerant is injected from the console to the inside of the balloon. Arctic Front Cardiac Cryoablation catheters are introduced into the vasculature by traditional minimally invasive techniques. The CryoConsole houses the electrical and mechanical components as well as proprietary software for controlling and recording a cryotherapy procedure with Arctic Front and/or the Freezor family of catheters. It stores and controls the delivery of the liquid refrigerant through the coaxial umbilical to the catheter, recovers the refrigerant vapor from the catheter under constant vacuum, and disposes of the refrigerant through the hospital scavenging system. Multiple features are built into the catheter and the CryoConsole system to ensure safety. | Medtronic CryoCath Arctic Front CryoAblation Catheter | 2 | November 02, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath has identified specific lots of Arctic Front CryoAblation Catheters that were distributedwith incorrect programming that may result in false expired catheter notifications. Impacted lots are 21571,25375, 25376, 25377, and 25381 of the Arctic Front Catheter Model Number 2AF282. Medtronic CryoCathis communicating this information to the FDA.On affected catheters, the incor | TRAINING: Employee Error | Medtronic sent an "Important Device Recall" letter dated August 2011, hand delivered by Medtronic Field Representatives on August 29, 2011 to all affected customers. The letter described the product affected, the problem, and the actions needed to be taken. The customers were instructed to return all unused catheters from affected lots for credit and replacement. The letter states that Medtronic AF Solution representative will assist customers with the return of affected product. For further questions contact your local Medtronic AF Solutions representative. | Contact the recalling firm for information | 263 | 263 | Worldwide Distribution - USA (nationwide) including the states of: AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MN, MS, MO, NE, NY, NC, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI. and the countries of: Austria, United Kingdom. | Terminated | June 07, 2012 | 218 | 263 | Software | Alarm/Message | Return | Remove or Replace
Z-0150-2012 | 59991 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RSP Humeral Socket Reamer (Size 32 SML):DJO Surgical Instruments and Instrument CasesDJO surgicalEncore Medical, L.P. 9800 Metric Blvd, Austin, TX 78758Contact information: DJO Surgical, ATTN: Customer Service, 9800 Metric Blvd, Austin, TX 78758, 1-800-456-8696.The product is intended to be used to ream prepare the humeral bone for accepting the implant. It has a spring loaded alignment pin on the proximal (reaming) end used in both alignment of the reamer and as an indication that the reamer has been seated fully. | RSP Humeral Socket Reamer (Size 32 SML) | 2 | November 02, 2011 | 2011 | Encore Medical, Lp | DJO Surgical has determined that the alignment pins cannot be consistently cleaned using the technique provided in the cleaning/sterilization section of the instrument IFU. | DESIGN: Component Design/Selection | DJO Surgical sent an URGENT FIELD SAFETY NOTICE letter dated September 28, 2011, to all affected customers. The letter identified the product. the problem, and the action to be taken by the customer. Customers were instructed to return the RSP Humeral Socket Reamers and RSP Baseplate Rim Planers. The RSP Humeral Socket Reamers will be reworked and returned to the customer. The Baseplast Rim Planer will be replaced when a suitable design is available. During the interim until instruments are returned, the use of FMP acetabular reamers size 42mm, 43mm, 45mm (DJO PIN 803-05-226, 227, 200) for the humeral preparation was validated to be effective to prepare the humerus for the implant.Customers were instructed to immediately notify the hospitals and surgeons using these devices and document it on the Field Safety Notice Response Form. For questions regarding this recall call 512-832-9500. | Contact the recalling firm for information | 399 | 399 Reamers | Worldwide Distribution - USA including AR, AZ, CA, CO, FL,GA, ID, IN, IL, LA, MA, MD, ME, MI, MO, NJ, NM, NY, OH, OK, PA, PR, SC, TN, TX, UT, VA, WA, and WI and Internationally toGermany, South America, Italy, Saudi Arabia, United Kingdom, and Canada. | Terminated | March 29, 2013 | 513 | 532 | Not_Computer
Z-0158-2012 | 60254 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | "***BIOMET 3i***4555 Riverside Drive Palm Beach Gardens FL 33410 USA***REF AC4425***Conical 25 ANGLED ABUTMENT***4.1mm(D) X 4mm(H)***Abutment; Pilier: Abutment; Abutment; Pilastro; Pilar***RX Only***Sterile***BIOMET 3i Dental Iberica S.L. WTC Almeda Park. Ed. 1. Planta 1'***".Attach an implant to an implant restoration. | Conical Angled Abutment | 2 | November 03, 2011 | 2011 | Biomet 3i, LLC | On 09/28/2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Conical Angled Abutment, Model # AC4425 Lot 887628-5 and their Standard Conical Abutment Cylinder Model # SWCA62 Lot 848478. The packaging for the referenced products may not have been completely sealed prior to shipment. | TRAINING: Employee Error | Biomet 3i, LLC sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to inspect their inventory and return any affected product for replacement.A Product Recall Response Form was attached for customers to complete and return via fax to 561-514-6316. Customers are directed to call 1-800-342-5454 for questions regarding this recall. | Contact the recalling firm for information | 122 | 122 | Worldwide Distribution-USA (nationwide) and the countries of Chile, Italy, Japan, Korea, Spain, and Sweden. | Terminated | March 09, 2012 | 127 | 130 | Not_Computer
Z-0152-2012 | 60100 | enzyme immunoassay, amphetamine | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ProScreen PSCup-5M Waived, Catalog Number: PSCup-5M WaivedProduct Usage: The ProScreen Drugs of Abuse Test Cup is a one-step immunoassay for the qualitative detection of multiple drugs in human urine (refer to PN: 42141-PS-W). The ProScreen Drugs of Abuse Test Cup is used to obtain a visual, qualitative result and is intended for professional use only. The ProScreen Drugs of Abuse Cup is not for legal or medical diagnostic purposes. This device is not intended for workplace testing. For diagnostic and treatment purposes, consult with a healthcare or substance abuse professional. This assay provides only a preliminary result. Clinical consideration and professional judgment must be applied to any drug of abuse test result, particularly in evaluating a preliminary positive result. In order to obtain a confirmed analytical result, a more specific alternate chemical method is needed. Gas Chromatography/Mass Spectroscopy (GC/MS) is the preferred confirmation method. | Us Diagnostic ProScreen Drugs of Abuse Cup | 2 | November 03, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech, Inc. received a complaint from US Diagnostics (USD) that First Check Drug Screen Cups were packaged in the ProScreen Cup boxes. Some of US Diagnostics' customers received ProScreen 5-panel cups that contain First Check THC cups. The recalled products are mislabeled. If the USD customer receives the FirstCheck tes, there will be no instructions for use | PRODUCTION CONTROLS: Packaging Process Control | The recall communication was initiated on June 9, 2011 with Ameditech Inc. forwarding an Urgent Medical Device Recall notification with an attached Verification Form (via phone and written notification) to the affected distributor. The letter provides the distributor with an explanation of the problem identified and actions to be taken. The distributor was required to: cease distribution of the affected lot and provide remaining inventory count to Ameditech; destroy any remaining inventory of these lots and document destruction per site requirements; and ensure that all user of the device have received a copy of the Urgent Medical Device Recall. In addition, the distributor was instructed to complete and return the enclosed Verification Form with in 10 days.For any questions about the information contained in this notice, contact Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 2000 | 2000 test cups (800 kits) | Nationwide Distribution (USA) including the state of Alabama. | Terminated | January 18, 2012 | 76 | 2000 | Not_Computer
Z-0154-2012 | 60114 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | 37% Etching Gel-1ml syringe 20 pk, part # T06AThe intended use of these devices is to etch tooth dentin/enamel in order to prepare the tooth surface prior to the application of a bonding agent. | 37 Etching Gel1ml syringe | 3 | November 03, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron Clinical is voluntarily recalling all lots of products labeled as 37% Phosphoric Acid Etching Gel produced since April 2007 because the product continues to extrude out of the syringe even after pressure is no longer applied, and it contains a higher percentage of phosphoric acid than that indicated on the label. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Pentron Clinical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 31, 2011, with attached Return Form (via USPS 1st class mail) to all their consignees/customers. The letter describes the product, problem and action to be taken.All consignees/customers were instructed to complete the Recall Return Form via fax to: 1-877-677-8844 and to return any affected product in their inventory. Affected products will have expiration dates between April 2011 and April 2015.Consignees/customers with questions can contact Pentron Clinical Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement. |  | 41744 | 41,744 units | Worldwide distribution: USA (nationwide) and countries of: Armenia, Austria, Australia, Antigua & Barbuda, Barbados, Belarus, Bulgaria, Canada, Czech Republic, Croatia, Cyprus, Ecuador, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kenya, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Poland, Romania, Russia, St. Vincent & Grenadine, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, UAE, Ukraine, UK, and Vietnam. | Terminated | August 09, 2012 | 280 | 41744 | Not_Computer
Z-0160-2012 | 57234 | gutta-percha | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Microseal Master Cone Part Number: 815-9033Microseal Master Cone is intended to obturate root canals during endodontic therapy procedures. | Microseal Master Cone Part Number: 8159033 | 3 | November 04, 2011 | 2011 | Ormco Corporation | The recall was initiated because SybronEndo has confirmed that the Microseal Master Cones has been mislabeled. | PRODUCTION CONTROLS: Error in Labeling | Sybron Dental Specialties, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return the affected lots of product.. Products, would be replaced at no charge or credited to their account. Customers were instructed to contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. The affected product should be returned to SybronEndo at the following address. Please label returned product "RECALLED PRODUCT Attention: Customer Returns"SybronEndo 1332 South Lone Hik AvenueGlendora, CA 91740ADDITIONALLY, THEY KINDLY REQUEST CONSIGNEE COOPERATION IN COMPLETING AND FAXING BACK THE ENCLOSED ACKNOWLEDGEMENT/RETURN FORM. THIS FORM MUST BE COMPLETED AND RETURNED WHETHER OR NOT YOU HAVE ANY PRODUCT TO RETURN.Customers were also asked to identify and recover the affected product lots listed above that may have been shipped to their customers.For questions regarding this recall call 909-962-5600. |  | 200 | 200 units | Worldwide Distribution - USA including AR, CA, FL, PA, GA, IL, TX, WA, NY, IN, AL & MT; Internationally to Canada, CIS (Commonwealth of Independent States, Israel, United Kingdom, Sweden, France, Germany, and Italy. | Terminated | November 21, 2011 | 17 | 200 | Not_Computer
Z-0161-2012 | 60003 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine T-PAL Spacer System. Transforaminal posterior atraumatic lumbar spacer system.Transforaminal posterior atraumatic lumbar spacer system. Synthes T-PAL Spacer is indicated for use in patients with degenerative disc disease (DDD) at one or two contiguous levels from L2 to S1 whose condition requires the use of interbody fusion combined with supplemental fixation. The interior of the T-PAL Spacer should be packed with autogeneous bone graft (i.e. autograft). DDD is defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies. These patients should be skeletally mature and have had six months of non-operative treatment. | Synthes Spine TPAL Spacer System. | 2 | November 04, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION for Medical Device - previous labeling and promotional materials associated with the T-PAL Spacer system reference the use of T-PAL spacer with allograft material. T-PAL Spacer is currently indicated for use with autograft material. All accounts with record of receiving T-PAL Implant will receive notification of the labeling correction and a copy of the updated brochure. All | PRODUCTION CONTROLS: Error in Labeling | SYNTHES sent an URGENT MEDICAL DEVICE LABELING CORRECTION letter dated September 23, 2011, to all affected customers. The labeling has been revised. Updated information packets have been sent to all accounts which included the final labeling and communications to accounts and sales force. Customers were instructed to review the updated supplemental labeling provided in the packageCustomers were instructed to attach the supplementary technique page to an T-PAL Spacer Brochure that may be in their facility. They can access an electronic version of the Brochure by accesssng Product Literature at the following link:http://us.synthes.com. The new version can be identified by code J10120 located at the bottom right corner of the back cover.To confirm receipt and understanding of the information, customers were requested to complete the Verification Section of the letter and return it to Synthes by either of the following methods:Mail provided postage paid envelopeFax 610-251-9005Scan/e-mail to FieldAction@synthes.comFor any questions call 1-800-420-7025 x5403 or contact their Synthes Spine Sales Consultant. |  |  | N/A - Labeling Correction | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, HI, KS, KY, LA, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. and Puerto Rico and Washington, D.C; Internationally to Canada and Japan | Open |  |  |  | Not_Computer
Z-0162-2012 | 60094 | cyclosporine | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | ARCHITECT Cyclosporine; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cyclosporine in human whole blood Distributed by Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1L75-25.The ARCHITECT Cyclosporine assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cyclosporine in human whole blood on the ARCHITECT i System. The ARCHITECT Cyclosporine assay is used as an aid in the management of heart, liver and kidney transplant patients receiving cyclosporine therapy. | ARCHITECT Cyclosporine | 2 | November 04, 2011 | 2011 | Abbott Laboratories | The ARCHITECT Cyclosporine assay has the potential to generate falsely elevated or depressed results with patient samples due to precision exceeding the Package Insert claim of | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated October 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to only run ARCHITECT Cyclosporine using Reaction Vessel (RV) lots manufactured with the same resin as follows: a) RV lots beginning with lot 06083P100 and higher can be used together or, b) RVs lots lower than 06083P100 can be used together. Any RV lot number that does not begin with a zero (e.g. 90032P100) should also be considered lower than 06083P100. Further instructions were included for the accounts to follow when running ARCHITECT Cyclosporine assays. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051. Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. A second letter was sent to the accounts on 11/23/11 to simplify the instructions for identifying the Architect Reaction Vessel lots that can be used together for use with the Architect Cyclosporine assay. |  | 8913 | 8,913 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Terminated | November 07, 2012 | 369 | 8913 | Not_Computer
Z-0165-2012 | 60241 | keratome, ac-powered | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Abbott brand IntraLase Femtosecond Laser System, Model 2; Catalog Numbers: 20003D/J/K, R20003D/J/K, 20005D/J/K; Product is manufactured and distributed by AMO Manufacturing USA, LLC, Milpitas, CA Product Usage: Intended Use for: IntraLase Femtosecond (Model 2 and 3) and iFS Advanced Femtosecond Laser Systems. FS Laser Systems are ophthalmic surgical lasers used in certain patient populations: (1) In patients undergoing surgery or other treatment requiring initial lamellar resection of the cornea; (2) In patients undergoing surgery or other treatment requiring initial lamellar resection of the cornea to create tunnels for placement of corneal ring segments; (3) In lamellar IEK [lntraLase Enabled Keratoplasty] and corneal harvesting ; (4) In the creation of a corneal flap in patients undergoing LASIK surgery or other treatment requiring initial lamellar resection of the cornea; (5) In the creation of a lamellar cut/resection of the cornea for lamellar IEK for the creation of a penetrating cut/incision for penetrating IEK. | INTRALASE FS Laser | 2 | November 04, 2011 | 2011 | AMO Manufacturing USA, LLC | The firm has identified a discrepancy in cut depth that may affect deep corneal incisions (e.g. 300 Âżm or deeper) created using the Laser System. | DESIGN: Process Design | Abbott Medical Optics Inc. (AMO) issued an "ADVISORY NOTICE" letter dated October 12, 2011 to all affected customers. The letter identifies the affected products, the problem, and the actions needed to be taken. Customers were instructed to exercise caution when making cuts that approach the corneal endothelium. The letter includes a 1-page Operator's Manual addendum that clarifies the recommendations. The letter states that AMO will conduct a field correction to adjust the EOF on the affected FS Laser Systems. The letter includes a response form to be completed and returned as per the instructions. Customers were instructed to notify all appropriate staff/personnel of the Advisory Notice. For any further questions regarding the information or recommendations, contact an AMO representative at 1-800-266-3375. | Contact the recalling firm for information | 1226 | 1226 total laser systems | Worldwide Distribution - USA (nationwide) | Terminated | May 01, 2013 | 544 | 1226 | Not_Computer
Z-0164-2012 | 60143 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BioPro Femoral Head 28 mm +18, Item 19010.Orthopedic femoral head for implantation | BioPro Femoral Head 28 mm 18 | 2 | November 04, 2011 | 2011 | Biopro, Inc. | BioPro Femoral Head Outer Packaging was found lacking proper seal ,compromising sterility. | PRODUCTION CONTROLS: Packaging | Biopro, Inc notified all affected customers via telephone on September 21, 2011 with instructions to quarantine and return affected product to the manufacturer. The firm sent a follow up "RECALL NOTICE" dated September 29, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter gives instructions on how to return affected product. Contact Customer Service at 1-800-252-7707 for questions regarding this recall. |  | 8 | 8 units | USA (nationwide) including the states of FL, MI, and TX. | Terminated | February 15, 2013 | 469 | 12 | Not_Computer
Z-0174-2012 | 60255 | introducer, syringe needle | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | 1.0mL 28 G x 1/2" blister packaged insulin syringe Catalog Number 329424; Lot Number 0039920Intended use: Subcutaneous injection of insulin. | Becton Dickinson 1.0mL 28 G x 1/2" blister packaged insulin syringe | 2 | November 07, 2011 | 2011 | Becton Dickinson & Company | Experiment labels may be on shelf cartons of 1.0ml 28G x 1/2 blister packaged insulin syringes. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, BD, sent an "PRODUCT RECALL NOTIFICATION" letter and response form dated June 15, 2011 to their consignee/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately review their inventory for the affected product, remove the product from inventory, quarantine the product and return it to BD; if the product was further distributed, they should identify their customers and notify them at once, and complete and return the attached Tracking/Verification Form via fax to: 201-847-4853.If you have any questions, contact Customer Service at 1-800-237-2762, Monday -Thursday, 8:00am-6:30pm and on Friday, 8:00-5:30pm Eastern Standard Time. | Contact the recalling firm for information | 10230 | 10,230 shelf cartons (1,0230,000 syringes) | Nationwide Distribution. | Open |  |  | 10230 | Not_Computer
Z-0173-2012 | 60183 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter System 1000 Tina Single Patient Hemodialysis Delivery System; Baxter Healthcare Corporation, Renal Division, McGaw Park, IL 60085; product code S1000L3P. Hemodialysis and Continuous Renal Replacement Therapy devices for acute and chronic hemodialysis | Baxter System 1000 Tina Hemodialysis Instrument | 2 | November 07, 2011 | 2011 | Baxter Healthcare Renal Div | The Tina System was released to the customer with OOS values for the Pressure Displays Accuracy and Leakage Rate Test. | TRAINING: Employee Error | Baxter Healthcare Corporation sent a service technician to the lone customer site on September 19, 2011 to correct the affected product. The technician executed all tests as established in the Service Manual, with acceptable results. | Contact the recalling firm for information | 1 | 1 unit | Distributed in Puerto Rico only. | Terminated | November 07, 2011 | 0 | 1 | Not_Computer
Z-0169-2012 | 60069 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act: a) 24" Infant Tape Measure, component 30942P; b) Silicone Tubing, component SFM3-4450; c) Triple Dye Drug, 10 unit dose, component PM1022; d) Assembly Swabstick, component PS004; e) 18x26 Poly Lined Towel, component 8580Z; f) Disposable Vein Stripper, component 63-0431;g) Cotton Glove, Large, White, PR, NS, component 1051; h) Cotton Glove, Medium, White, PR, NS, component 1051M; i) 6 Inch Metric-Inch Ruler, component 18C-Thru; The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then sterilized and are ready to be used by the customer for the specific procedure. | Presource Custom Sterile Surgical and Procedure Kits | 2 | November 07, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturers do not have all necessary FDA registrations or filings needed for these components. | DESIGN: Component Design/Selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 30, 2011 to the affected wholesale consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit and shipping carton in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact Distributor Management Group at 800-635-6021. This wholesale recall was expanded to include additional kits via letter dated October 5, 2011, with the same instructions as listed above.For questions regarding this recall call 800-292-9332. |  |  |  | Nationwide Distribution including California, Florida, Illinois, Kentucky, Minnesota, Missouri, Ohio, Tennessee, Washington, West Virginia and Wisconsin | Open |  |  |  | Not_Computer
Z-0171-2012 | 58766 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | ***REF IOSS511***OSSEOTITE Certain Implant 5 x 11.5mm***LOT 910695***Sterile using Radiation***Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n 08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment. The Thread form titanium screws (cylindrical dental implants) designed to be placed into mandibular or maxillary jaws in order to affix a dental prosthesis (edentulous or partially edentulous segments) that replaces 1 or more missing teeth. | OSSEOTITE Certain Implant | 2 | November 08, 2011 | 2011 | Biomet 3i, LLC | The product may contain a manufacturing defect, that if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | PRODUCTION CONTROLS: Process Control | The firm sent an Urgent: Medical Device Recall letter, dated November 2010, to notify customers that received the Osseotite Certain Implant (IOSS511) of the problem. The customers were instructed to return the unopened device(s) to Biomet 3i for replacement. If customers notice an issue, they were to contact Customer Service at 1-800-342-5454 for assistance. Customers were to respond via fax and return the implant(s) to BIOMET 3i using the address and incident number provided. Customers with questions or concerns were instructed to contact Mark Mashburn at 561-776-6906. |  | 241 | 241 | Worldwide Distribution -- USA , BR, JP, TW, CA, Germany, and Guam. | Terminated | November 09, 2011 | 1 | 241 | Not_Computer
Z-0176-2012 | 59684 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Microtainer MAP Microtube for Automated Process with K2EDTA. For the collection, transport, and automated processing of skin punctures samples from infants, children, geriatric, oncology, and critical care patients. | BD Microtainer MAP Microtube for Automated Process. | 3 | November 08, 2011 | 2011 | Becton Dickinson & Company | BD Microtainer MAP Microtube is being recalled due to the potential clotting of blood specimens resulting from a reduced amount of EDTA additive. | DESIGN: Device Design | A Becton Dickinson Sales Consultant conducted a "on-site" visit to the single Customer account, Florida Hospital, on February 11, 2011. All units were successfully recovered. All product recovered was quarantined and destroyed. Becton Dickinson believes that all reasonable efforts to communicate the issue to the affected consignee have been completed. If you have any questions or request additional information please do not hesitate to call (201) 847-5636. | Contact the recalling firm for information | 10800 | 10,800 units | Distributed only to Florida. | Terminated | November 10, 2011 | 2 | 10800 | Not_Computer
Z-0175-2012 | 59996 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter Assembly Manufactured by: Stryker Spine SAS Z. I. Marticot - 33610 - Cestas France; Distributed in the USA by: Stryker Spine 2 Pearl Court. Allendale, NJ 07401-1677 USA Stryker Mantis Rod Inserter Inner Shaft Manufactured by: Stryker Spine SAS Z. I. Marticot - 33610 - Cestas France; Distributed in the USA by: Stryker Spine 2 Pearl Court. Allendale, NJ 07401-1677 USA IS2097MAN Mantis Rod Inserter ShaftUsed to facilitate insertion and delivery of a Mantis rod. | Mantis Rod Inserter | 2 | November 08, 2011 | 2011 | Stryker Spine | There have been reports of difficulty in rod rotation using the newly designed Mantis Rod Inserter. | DESIGN: Device Design | Stryker notified branches/agencies, surgeons and hospitals beginning June 24, 2011 with an Urgent Product Recall letter, dated June 24, 2011, via Fed Ex. The letter identified the affected product, the issue, potential hazards, risk mitigations, and the return process. Customers were asked to examine their inventory and reconcile any inserters and shafts with the catalog numbers provided using the Customer Response Form. Customers are to fax the form to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should be retrieved and returned. Questions should be directed to 201-760-8298. |  | 471 | 471 units US, 667 units international | Nationwide Distribution | Terminated | November 13, 2012 | 371 | 471 | Not_Computer
Z-0177-2012 | 59425 | calibrator, dose, radionuclide | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Capintec CRC 25 PET, item number 5130-3217 with Software Revisions 2.01 and 2.02PET dose calibrator | Capintec CRC 25 PET | 2 | November 09, 2011 | 2011 | Capintec Inc | Software error in Rev. 2.01 and Rev 2.02 for Capintec CRC 25PET | DESIGN: Software Design | Capintec, Inc. sent a Software Error Alert letter dated July 8, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contac their Capintec Customer Support Representative to receive a software upgrade. For questions call Capintec Customer Support at 1-800-631-3826. |  | 69 | 69 | Worldwide distribution - United States including CA, FL, IL, MA, NJ, OK, PA, and VA; and internationally to Brazil, China, India, Indonesia, Italy, Japan, Korea, South Korea, and Turkey. | Terminated | December 10, 2013 | 762 | 69 | Software |  | Software update | Software Update
Z-0181-2012 | 60022 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Clinac, Trilogy, Trilogy Tx and Novalis linear accelerators.Manufactured by Varian Medical System, Palo Alto, CA.Varian High Energy Linear Accelerator is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | Clinac, Trilogy, Trilogy Tx and Novalis linear accelerators | 2 | November 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has received reports in which a user has remotely rotated the gantry into contact with the couch or with the patient, in both the manual mode and the automate mode. | DESIGN: Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated September 9, 2011 to all affected customers via Federal Express with attached instructions. The letter provides the description of problem; references, details, recommended user actions, and varian actions.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***In addition, the letter also states "THIS LETTER IS NOT A SUBSTITUTE FOR THE CAREFUL REVIEW OF ALL USER DOCUMENTATION ASSOCIATED WITH YOUR PARTICULAR EQUIPMENT AND SOFTWARE."If you have any questions, contact Varian Oncology Help Desk at USA and Canada: 1.888.VARIANS (888.827.4265); Europe +41 41 749 8844 and/or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 5823 | 5823 devices | Worldwide Distribution. | Terminated | July 24, 2012 | 258 | 5823 | Not_Computer
Z-0179-2012 | 59918 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Quick Shield with Snappy Safety Holder, Greiner bio-one.The VACUETTE QUICKSHIELD with SNAPPY Tube Holder is intended to be used only with the VACUETTE Blood Collection Needles and VACUETTE Blood Collection Tubes as a system in routine venipuncture procedures. These devices are to be used by properly trained healthcare professionals only in accordance with these instructions. | Quick Shield with Snappy Safety Holder | 2 | November 09, 2011 | 2011 | Greiner Bio-One North America, Inc. | Reports of the Vacuette QuickShield Safety Holders breaking during use due to a crack at the holder hub. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Greiner Bio-One sent a "QUICKSHIELD WITH SAFETY HOLDER" notification letter dated August 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to stop distributing and using the affected products. A Product Disposition Form was attached for customers to complete and return for replacements. Contact the firm at 888-266-3883 for questions concerning this recall. | Contact the recalling firm for information | 443950 | 443,950 pieces | Nationwide Distribution-USA (nationwide) including the states of CO, IL, KY, NY, PA and VA. | Terminated | March 20, 2012 | 132 | 443950 | Not_Computer
Z-0178-2012 | 60104 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | LifeCell Strattice Reconstructive Tissue Matrix for Stoma ReinforcementCatalog #0606008 (6 x 6 cm), 0808008 (8 cm x 8 cm), (0610008 ( 6 cm x 10 cm).Intended for use as a soft tissue patch to reinforce soft tissue where weakness exists and for the surgical repair of damaged or ruptured soft tissue membranes. | Strattice Reconstructive Tissue Matrix for Stoma Reinforcement | 2 | November 09, 2011 | 2011 | LifeCell Corporation | The use of Strattice for stoma reinforcement at the time of stoma creation is not within the product's cleared indications for use in the US market. | PREMARKET APPROVAL: No Marketing Application | LifeCell began notifying their customers via phone on 7/13/2011. Recall Notice Letters and return response forms, dated July 12, 2011, were sent out on the same day. The letter identified the affected product and informed customers of the reason for the recall. Customers were asked to examine their inventory and quarantine the affected product. If customers have affected product, they were to contact Customer Solutions at 1-866-423-2433 to arrange for return of the product to LifeCell. Also, the attached Recall Notification form should have been completed and returned no later than August 17, 2011. Questions should be directed to Customer Solutions at 1-866-423-2433. |  | 2651 | 2651 sublots (of 42 lots) | Nationwide Distribution | Terminated | March 04, 2013 | 481 | 2651 | Not_Computer
Z-0172-2012 | 59829 | massager, therapeutic, electric | ISA | Therapeutic massager. | Physical Medicine | Physical Medicine | 510(K) Exempt | Shoulder Flex Massage DeviceDevice is an electric personal massage device labeled in part: "ShoulderFlex Deep-kneading Shiatsu* Model #SH61 *Read all instructions before using the massager* **WARNING NEVER wear a necklace or similar items while using this product Made in China."Personal use massage device designed for home use to provide deep tissue massage to the neck, shoulder and back while user lies on a flat surface. | Shoulder Flex Massage Device | 1 | November 10, 2011 | 2011 | King International | The covered rotating mechanism can catch on user's clothing, jewelry (e.g. necklace), and hair. There have been reports of one death and a strangulation associated with using the device. | DESIGN: Component Design/Selection | An Urgent Product Recall letter was sent via regular mail to customers. The letter identified the affected product and described the reason for the recall. Customers were asked to stop using Shoulderflex massagers immediately and to safely dispose of them in the trash. The device components should be disposed of separately so that the device cannot be reassembled and used. This means that the power supply should be separated from the massager unit, and the massage fingers should also be removed and disposed of separately. Customers are to complete and return the enclosed the postage-paid card to King International in order to confirm receipt of the recall letter. If additional information is needed regarding this recall, customers are to contact King International LLC at 503-524-7046 or visit their website at www.shoulderflex.com or via mail to PO Box 2384, Beaverton, Oregon 97075.On 8/25/2011 FDA published a health alert on the FDA website. On 8/31/2011 the firm issued a nationwide press release. |  | 11934 | 11934 | Nationwide Distribution | Open |  |  | 11934 | Not_Computer
Z-0180-2012 | 60048 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat cTnI cartridgesIs an in vitro diagnostic test for the quantitative measurement of cardiac troponin I in heparinized whole blood or plasma samples. | iStat cTnI cartridges | 2 | November 10, 2011 | 2011 | Abbott Point Of Care Inc. | i-Stat cTnI cartridges beginning with letter "T" may exhibit a higher than expected variability in reported results. This increased variability may result in falsely elevated or falsely depressed results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | APOC issued Urgent Recall Notices/Business Reply Forms dated September 2011 to consignees. The letter identified the affected product and explained the background of the problem. Customers are to discontinue use of the cartridges from the lots identified. All unused product will be credited and should be returned as instructed on the attached Business Reply Card. If the affected product was further distributed, a copy of the Recall Notice should be provided to those customers as well. APOC shipped replacement product in order to minimize any disruption. If replacement product has not been received then customers should contact Customer Service at 1-800-366-8020, option 5. Questions regarding this information should be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1 or customers' local APOC representative. | Contact the recalling firm for information | 1688755 | 1,688,755 cartridges | Worldwide Distribution | Terminated | March 11, 2013 | 487 | 1688755 | Not_Computer
Z-0182-2012 | 60036 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation-NM (EBW-NM) with software version 1.0P, 1.1.1A, 1.5H, 1.5.1A and 2.0Q -Field Test PeriodA nuclear medicine image display and processing application suite that provided software applications used to process analyze and display medical images/data. | Extended Brilliance WorkstationNM (EBWNM) | 2 | November 10, 2011 | 2011 | Philips Medical Systems | A problem related to a software error has been detected in the Philips Multiple Gated Radionuclide Angiocardiography (MUGA) application. A false negative interpretation due to a falsely elevated ejection fraction or a false positive interpretation due to a false low ejection fraction could occur. | DESIGN: Software Design | Philips sent all affected sites an Urgent Medical Device Correction Letter on September 27, 2011 for the MUGA issue. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy of the communication with the equipment instruction for use. Additionally, a field change order was released on September 26, 2011, customers will be receiving release notes as part of their correction informing them of the changes and corrections to the software. The new software will be installed on affected units by Philips Field Service Engineers.For customers in North America contact Customer Care Solutions Center at 1-800-722-9377. Select option 5. In all other countries contact the local Philips Healthcare office. |  | 1181 | 1181 systems | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 30, 2013 | 447 | 1181 | Software | Output/Calculation | Software update | Software Update
Z-0187-2012 | 59461 | bed, therapeutic, ac-powered, adjustable home-use | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare 600 lb. Bariatric Bed, Head motor extension cable used on component numbers BAR5490 and BAR5490IVC. These components are part of the following Bed Model Number(s): BAR600, BARPKGCA, BAR600IVC, BARPKGIVC-1633, and BARPKG. | Invacare bariatric bed | 2 | November 14, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that miskeying or disconnection of the head motor extension cable during bed set-up may result in product damage, or in rare instances, the potential for fire. | DESIGN: Labeling Design | On 9/19/11, Invacare sent Recall Notificaiton Letters to their customers. They instructed their customers to quarantine the product, complete and return a response form. | Contact the recalling firm for information | 35747 | 35,747 | Distribution Nationwide and to United Arab Emirates, Dominican Republic, Egypt, Guam, Mexico, & Virgin Islands. | Terminated | July 02, 2013 | 596 | 35747 | Not_Computer
Z-0186-2012 | 60242 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, Part Number: N97A,Product Usage: The intended use of this device is for cementation of porcelain, resin, and metalbased inlays, onlays, crowns, bridges and posts. The material is not recommended for use on veneer restorations. | Breeze SelfAdhesive Resin Cement | 2 | November 14, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron is voluntarily recalling the Breeze Self Adhesive Resin Cement because it was discovered that there may be a curing problem that may cause cemented crowns to fall off after use of the product. |  | Pentron Clinical sent an "Urgent: Medical Device Recall" letter dated July 18, 2011 with an attached Recall Return Form (via USPS First Clas mail) to all affected customers. The letter identified the affected product, explanation of the problem, and actions to be taken. Customers were instructed to discontinue use and return any affected product in their inventory. The letter asks that each customer complete and return the enclosed Recall Return Form via fax to 714-288-4609, in order to confirm receipt of the customer notification.Customers were instructed to contact Pentron Clinical Customer service at (800) 551-0283 (option 1) directly to handle the arrangements of a quick return and replacement. |  | 336 | 336 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Cypress, Greece, Germany, Italy, Jordon, Poland, Romania, Turkey and Vietnam. | Terminated | June 28, 2013 | 592 | 336 | Not_Computer
Z-0183-2012 | 60122 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Pulstar(R) DVT PREVENTION SYSTEM, Anti-Embolism Wrap System, albahealth(R), A Subsidiary of Encompass Group, LLC.Prevention of DVT. | Pulstar(R) DVT PREVENTION SYSTEM | 2 | November 14, 2011 | 2011 | Albahealth LLC | The wrap portion of the device was tearing at the seam. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Albahealth, L.L.C. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory and to quarantine all affected product. A response form was enclosed for customers to complete and return. The firm provide information on how to return the affected product and receive replacement product. Questions concerning this recall are directed to call 865-354-5229 for assistance. | Contact the recalling firm for information | 2582 | 2,582 pairs | Nationwide Distribution- including the states of CA, CO, FL, GA, IN, KY, LA, NY, OH, PA, TX, and TN. | Terminated | March 30, 2012 | 137 | 2582 | Not_Computer
Z-0184-2012 | 59919 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Greiner bio-one, 4ml K2E EDTA K2 Vacuette tube. Used for venous blood collection. Item number 454209. | Vacuette 4ml EDTA DK2 tube | 2 | November 14, 2011 | 2011 | Greiner Bio-One North America, Inc. | Vacuette 4ml EDTA DK2 tubes have been reported for short fills to no fills. | PRODUCTION CONTROLS: Process Control | Consignees were notified by Greiner bio-one letter on August 11, 2011. They were instructed to: Stop distributing the material, complete the product Disposition form attachÂŽ and to fax it to Greiner bio-one and to destroy or return any of the affected products for replacement or credit. Distributors were asked to provide a list of their customers that have received the item, preferably in an Excel file, so that they could be notified of the recall. |  | 504000 | 504,000 pieces | Nationwide. | Terminated | March 12, 2012 | 119 | 504000 | Not_Computer
Z-0188-2012 | 60370 | pad, alcohol, device disinfectant | LKB |  |  | General Hospital | 510(k) | PDI Alcohol Prep Pads packaged in the INRatio kit. PDI Alcohol Prep Pads, 100/box.Alcohol Prep Pads, packaged in the Alere INRatioÂŽ 2 PT/INR Home Monitoring Kit; or Hemosense INRatioÂŽ 2 Prothrombin Time/INR Testing Kit, Self-Test System; Alcohol Prep Pads, packaged in the Hemosense INRatioÂŽ Prothrombin Time/INR Testing Kit, Self-Test System;Alcohol Prep Pads, packaged in the INRatio Starter Kit ProductModel Number: 0200086, 0200432, 0100007, 0100072The Alere INRatioÂŽ2 PT/INR Home Monitoring System is used for the quantitative measurement of Prothrombin Time (PT) in fresh, capillary whole blood. The Alere INRatioÂŽ2 PT/INR Home Monitoring System is intended for use outside the body (in vitro diagnostic use) by people taking warfarin and other oral anticoagulant (blood thinning) therapy who need to monitor the clotting time of their blood. The Alere INRatioÂŽ2 PT/ INR Home Monitoring System is not intended to be used for screening purposes. The Alere INRatioÂŽ2 PT/INR Home Monitoring System should only be used with a prescription from your physician. The INRatio2 PT Monitoring system is used for the quantitative measurement of Prothrombin Time (PT) in fresh, capillary whole blood. The INRatio2 PT Monitoring system is intended for use outside the body (in vitro diagnostic use) by people taking warfarin and other oral anticoagulant (blood thinning) therapy who need to monitor the clotting time of their blood. The INRatio2 PT Monitoring System is not intended to be used for screening purposes. The INRatio2 PT Monitoring system should only be used with a prescription from your physician. The INRatioÂŽ system performs Prothrombin Time (PT) testing on fresh, capillary whole blood. The PT test is the principle assay used to monitor warfarin therapy. Patients taking warfarin and other oral anticoagulants should consult with their healthcare provider before adjusting their dosage. The thromboplastins used in performing PT tests can vary substantially causing differences in results between methods. The International Normalized Ratio (INR) is a mathematical correction of the PT result that adjusts for differences in thromboplastins. Consult your physician for the appropriate therapeutic range for you. | PDI Alcohol Prep Pads packaged in the INRatio kit. | 2 | November 16, 2011 | 2011 | Alere San Diego | Alere has initiated a recall of the Professional Disposables International (PDI) alcohol prep products due to the potential presence of low levels of the microorganism Bacillus cereus. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Alere sent a recall notification letter dated October 6, 2011, and Verification Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use an alternate prep pad that is not involved with the PDI recall, or use a sterile gauze pad with isopropyl alcohol. If customers purchase alcohol pads from a store or pharmacy, customers can confirm with their pharmacist that the pads they are purchasing are not associated with the PDI recall. INRatio/INRatio2 kits at Alere are currently being reworked to include a sterile PDI alcohol pad.Customers were instructed to complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of the notice and to indicate the number of discarded/required replacement alcohol prep pads.Customers with questions about INRatioÂŽ/ INRatioÂŽ 2 were instructed to consult their healthcare provider, or call Alere Technical Service at 877-866-5313 (24 hours a day, 7 days a week). | Contact the recalling firm for information | 14248 | Estimated 14,248 Total units | Nationwide Distribution | Terminated | February 29, 2012 | 105 | 14248 | Not_Computer
Z-0190-2012 | 60366 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Step 2, 3000-22000 GPD (Sodium Hydroxide 2.0N) cleaning solution. 4 oz plastic bottle. Lot # 11IC002, Serial #'s: 11IC00210 - 11IC002294.The Step 2, 3000-2000 GPD chemical is used in the cleaning procedures for the Isopure MD400 Series Reverse Osmosis System. Step | Step 2, 300022000 GPD (Sodium Hydroxide 2.0N) | 3 | November 16, 2011 | 2011 | Isopure Corp | The cleaning solution Step 2, 3000-22000 GPD labeled bottles (NaOH) were filled with Step 1, 3000-22000 GPD chemical (HCl). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Isopure sent a Product Recall Notice dated October 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue the use of any Step 2 chemical from the affected product, place any remaining chemical into the original box, seal it with tape, and segregate the product to prevent it from being placed back into stock. The firm will pick up the product for return on a call tag. Customers would receive replacement product for all affected cases of the Step 2 chemical the firm received.For questions call 502-722-1000. |  | 107 | 91cs/12/4 oz plastic bottles | Nationwide Distribution including the following states: AL, AZ, CA, FL, GA, IN, KS, LA, MA, MO, MS, NC, NH, NJ, NM, NV, OH, OR, PA, SC, TN, TX, and VA | Terminated | February 24, 2012 | 100 | 91 | Not_Computer
Z-0192-2012 | 60097 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Modular Dual Mobility (MDM) Liner and X3 Acetabular Insert; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; Stryker France ZAC Satolos Green Pusignan Av de Satolas Green 69881 Mey Sieu Cedex FranceThe MDM Liner and X3 Acetabular Inserts is a sterile single-use device intended for use in primary and revision total hip arthroplasty to alleviate pain and restore function. This device is intended to be used only with any currently available Howmedica osteonics 22.2 mm and 28 mm diameter femoral heads. | Modular Dual Mobility (MDM) Liner and X3 Acetabular Insert; | 2 | November 16, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become ware that there is the potential for interpretation of the product labeling which may lead to an incorrect implant being used. | DESIGN: Labeling Design | The firm, Stryker, sent an "URGENT PRODUCT CORRECTION" letter and Product Correction Bulletin dated June 3, 2011 via Fed Ex to its customers. The letter and bulletin describes the issue, potential hazards, risk mitigation and product correction instructions. The customers were instructed to complete and return the enclosed PRODUCT CORRECTION ACKNOWLEDGMENT FORM within 5 days via fax to: 201-831-6069.If you have any questions, feel free to call (201) 972-2100 or Hip Marketing Team at 201-831-5280. |  | 1951 | 1951 units | Nationwide distribution. | Terminated | November 13, 2013 | 728 | 16891 | Not_Computer
Z-0195-2012 | 60145 | driver, wire, and bone drill, manual | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | IBO Blade, Right, for use with Stryker TPS Reciprocating Saw, REF 5100-37-901, Stryker Instruments, Kalamazoo, MI.Product Usage: The intended use of these devices is to cut bone and hard tissue in maxillofacial procedures and have been specifically designed for osteotomies of the mandibular ramus. | Stryker | 2 | November 16, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | There is a potential for these blades to break where the blade is welded on to the arbor. This hazard can result in the following harms: tissue damage due to the broken blade being aspirated or swallowed, additional surgery to remove a broken blade, infection due to an unretrieved device component, complications associated with increased time under anesthesia to retrieve a broken blade, osteomyeli | DESIGN: Device Design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated October 18, 2011 to Risk Managers, Office Managers, and Oral Surgeons. The letter described the products under recall, the problem and a series of actions to be taken. Customers were asked to immediately locate recalled products and quarantine. Products were to be shipped back to Stryker. The letter included a Business Reply Form to be faxed back to Stryker Instruments Regulatory Department, 866-521-2762. The letter instructs to forward this notification to all affected locations if the recalled products were further distributed. Customers with questions can call Stryker at 1-800-253-3210. |  | 5691 | 5,691 | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Canada, China, Germany, Greece, Hong Kong, India, Jordan, Miramar, Saudi Arabia, Singapore, Spain, Switzerland, United Kingdom. | Terminated | July 12, 2013 | 604 | 12631 | Not_Computer
Z-0189-2012 | 60059 | urease and glutamic dehydrogenase, urea nitrogen | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Urea Nitrogen List Number 7D75-21 and 7D75-31 used for the Quantitation of urea nitrogen in human serum, plasma, and urine. | Abbott Clinical Chemistry Urea Nitrogen | 2 | November 16, 2011 | 2011 | Abbott Laboratories, Inc | The reagents are being recalled due to particulate matter in some of the reagent cartridges. The particulate matter has been identified as Aureobasidium fungus species. | PRODUCTION CONTROLS: Process Control | The firm decided to recall and they sent a Product Recall Letter to their customers on 10/03/2011. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. Customers should determine if you have any remaining inventory. Please discontinue and destroy any remaining inventory of these lots. Order alternative lots of Clinical Chemistry Urea Nitrogen. Complete and return the enclosed Customer Reply form. |  | 5704 | 5704 units of Urea Nitrogen | Worldwide Distribution. | Terminated | May 11, 2012 | 177 | 5704 | Not_Computer
Z-0191-2012 | 60174 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, Caridian BCT, Lakewood, CO 80215.A blood component separator used for therapeutic plasma exchanges. | Spectra Optia Apheresis System | 2 | November 16, 2011 | 2011 | Caridian BCT, Incorporated | Possible higher than expected platelet loss following therapeutic plasma exchange or plasmapheresis. | DESIGN: Software Design | CaridianBCT sent a "Safety Alert" letter dated September 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm will be sending a recall letter to consignees in November 2011, stating they are voluntarily recalling the software and will provide an upgrade once the software has been approved.If you have any further questions please contact your CaridianBCT Representative, the CaridianBCT Support Center at 1-877-3-FYU-BCT ( US Toll free 1-877-299-4228 ) or (303) 231-HELP ) ( 1 (303) 231-4357, or your local CaridianBCT Customer Service office. |  | 852 | 852 units (US: 333; Foreign: 519) | Worldwide Distribution -- USA ( nationwide ) and the countries of Australia, Belgium, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Pakistan, Republic of Korea, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand and Vietnam. | Terminated | February 01, 2013 | 443 | 852 | Software |  | Software update | Software Update
Z-0200-2012 | 60118 | bracket, plastic, orthodontic | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Brackets, part #494-1210.The brand name of the device is Spirit MB Brackets, an orthodontic bracket. Spirit MB Brackets have been assigned the product code 76 DYW (Plastic Orthodontic Bracket) by the FDA and is classified as a Class II Medical Device. Spirit MB Brackets are to be fixed attachment on a tooth which holds an archwire during orthodontic treatment. | Spirit MB Brackets | 3 | November 16, 2011 | 2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error. | DESIGN: Device Design | Ormco Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product. If any of the affected product is found, customers are directed to contact Ormco Customer Care at 1-800-854-1741 to receive a RMA number. The RMA number will allow a quick return and replacement or credit. Additionally, a Recall Return/ Acknowledgement Form was attached for customers to complete and return via fax to 909-962-5605. | Contact the recalling firm for information | 10 | 10 packs | Worldwide Distribution-USA (nationwide) including the states of FL, MN, TX, and VA, and the countries of Australia, France, Germany, Italy, Japan, Mexico, Morocco, Netherlands, New Zealand, South Africa, Turkey, and Ukraine. | Terminated | April 24, 2012 | 160 | 145 | Not_Computer
Z-0185-2012 | 60202 | dc-defibrillator, low-energy, (including paddles) | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 20 defibrillator/monitor is an AC powered external defibrillator with battery backup used by healthcare providers in hospital and clinic settings.Affected Product Part Numbers: 3202487-000 thru -012, 3202487-014 thru -027, 3203487-029 thru -039, 3202488-000 thru -019, U3202487-000, U3202487-015, U3202488-000. | LIFEPAK 20 defibrillator | 2 | November 16, 2011 | 2011 | Physio Control, Inc. | When users do not disconnect the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform. | TRAINING: Employee Error | Physio Control sent a " Urgent Medial Device Safety Alert" letter dated January 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The letter instructed the customers to follow the Operating Instructions for AED and Manual Defibrillation mode (Section 4-Therapy, pages 58 and 61) that instructs users to "Connect therapy electrodes to the therapy cable and confirm cable connection to the defibrillator" to prevent users from mistakenly leaving the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform.Recommendations to users who use a Therapy Cable:-Disconnect the QUIK-COMBO test plug prior to patient use-Confirm therapy electrodes - to - therapy cable connection-Confirm therapy cable - to - defibrillator connection-Confirm therapy electrodes - to - patient connectionAdding 2005 American Heart Association software that includes audible and visual prompts to LIFEPAK 20 defibrillators manufactured prior to January 2007 is optional and not required. If the customers would like to have this optional kit installed in your defibrillators, please contact Physio Control, Inc.Customers can call Technical Support at 1-800-442-1142 for any questions about this recall or or visit our website at www.physio-controlnotices.com/LP20testplug. | Contact the recalling firm for information | 23467 | 23,467 units worldwide | Worldwide Distribution--USA ( nationwide ) and the countries of Bahrain, Lebanon , Cyprus, Netherlands, Egypt, Pakistan, Germany, Saudi Arabia, India, South Africa, Iran, Syria, Israel, Turkey, Jordan, UAE,Kuwait,Yemen, Asia Pacific, Australia, New Caledonia, Guam, New Zealand, Hong Kong, Latin America, Anguilla, Jamaica, Argentina, Mexico, Bahamas, Nicaragua, Brazil, Panama, Chile, Peru, Colombia, Trinidad, Costa Rica,Uruguay, El Salvador, Venezuela, Guatemala, Virgin Islands and Canada. | Terminated | November 21, 2011 | 5 | 23467 | Other |  |  | Other
Z-0209-2012 | 60420 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable, trained, and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical, gynecological, medical cytological, and autopsy case specimens within the Pathology labs, for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SoftPath GUI | 2 | November 17, 2011 | 2011 | SCC Soft Computer | Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. SoftPath architecture does not correctly handle AP macros and data fields inside report sections that are stored in the database. Improper information can be displayed in all places where the diagnosis from the database is read and displayed. | DESIGN: Software Design | The firm, SCC Soft Computer, sent a "CORRECTION COMMUNICATION" notice on July 13, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to restrict security access to the template set-up files and to acknowledge receipt of the notice; grant SCC personnel permission to check their report templates and correct any that are affected; indicate if they wish to schedule the utility to be run and if they would like to have the SoftWeb parameter PathPDFLinkOnly set to "Y". SCC will review the templates on the customers systems; discuss any irregularities requiring change and move or remove AP macros and data elements as needed.If you have any questions, contact the Director of Quality Management at (727)789-0100. |  | 142 | 142 | Worldwide distribution: USA (nationwide) including states of: AR, AL, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI, WV including Puerto Rico; country of Canada. | Terminated | July 17, 2012 | 243 | 142 | Software |  |  | Other
Z-0207-2012 | 59990 | transmitters and receivers, electrocardiograph, telephone | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | LifeStar ACT III Platinum Sensor Ambulatory Cardiac Telemetry; an automatically activated system with a high performance triple channel EKG, office hookup and auto send/auto detect that requires no patient intervention to either capture or transmit an arrhythmia when it occurs; Manufactured by Card Guard, 2 Pekeris St, Rehovot, 76100, Israel; Model FG-00084LifeStar ACT provides critical ECG information by capturing the onset and termination/escape of an arrhythmia to assist in correctly identifying and treating the patient. Patients are monitored for up to 30 days and all of the data is available at the request of the physician during the service period and for 7 days following the end of service. The ACT III will provide more sensitive data for initial or early detection of arrhythmia in patients that have minimal or atypical symptoms. | LifeStar ACT III Platinum Sensor | 2 | November 17, 2011 | 2011 | Lifewatch Services Inc | LifeWatch received complaints of minor shocks and burns during use of the LifeStar ACT systems with shorts in the leads. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, LifeWatch, sent an " Urgent Device Recall" letter dated September 27, 2011 to the affected customers. The letter describes the product, problem and action to be taken. The customers were instructed to immediately refrain from providing these devices to their customers; stop using the affected sensors, return their entire kit in the pre-paid mailer, and use the new device kit included with the recall letter. In addition, the letter informed the users that the replacement device has upgraded software and firmware that allows for detection of shorts in the cable, and if detected, cuts off current flow to the patient to prevent minor shocks and burns.If you have any questions, please call (847) 813-4258. |  | 442 | 442 sensors | Nationwide distribution. | Terminated | May 15, 2012 | 180 | 1094 | Computer
Z-0202-2012 | 59960 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed paradigm mio Infusion Sets MMT-921The Paradigm mio infusion sets are intended for the subcutaneous infusion of insulin from a Medtronic MiniMed Paradigm insulin infusion pum. | Medtronic MiniMed paradigm mio Infusion Sets | 2 | November 17, 2011 | 2011 | Medtronic Minimed | The consignee received four (4) boxes of mio infusion sets (MMT-921) that were shipped to them in error. These infusion sets were shipped in a package that included a MiniMed Paradigm REAL-TIME Insulin Pump among other items.The mio infusion sets were not authorized to be shipped to the consignee based on Medtronics internal product handling procedures and they may not function properly.T | TRAINING: Employee Error | The firm, Medtronic, sent an "IMPORTANT RECALL INFORMATION" letter dated May 16, 2011 to consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use right away; use the replacement mio infusion sets provided to them in this package; return all unused Lot 3700255 mio infusion sets using the emailed pre paid return label, and drop the package off at any UPS location or call UPS at 800.742.5877 to schedule a pick up."Please note that only the Lot 3700255 mio infusion sets were shipped in error. The other products included in the package were not shipped in error and do not need to be returned."If you have any additional questions relating to this notification or any questions in general call 800.999.9859 ext.63704. | Discard product or return to place of purchase for a full refund or exchange | 4 | 4 boxes | Nationwide distribution: AL only. | Terminated | December 14, 2011 | 27 | 4 | Not_Computer
Z-0203-2012 | 60124 | enzyme linked immunoabsorbent assay, rubella | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | EHEC Kit Lot# 608096 Containing: Sample Diluent 18.8 ml Lot# 8602EXP.1-BUF WASH 20X II 100.0ml.Lot# 2607EXP.2-EHEC Detection Antibody 10.0ml Lot# 8603EXP.3-EHEC-Enzyme Conjugate 10.0ml Lot# 8604EXP.4-Substrate II 12.5ml Lot# 1601EXP.5-Stop-Solution II 10.0 ml Lot# 1601EXP.6-EHEC Positive Control 3.4 ml Lot# 8601EXP.7-EHEC Negative Control 3.4 ml Lot# 1603EXP.8-96 transfer pipettes-1 Microwell holder-3 Microwell strip sealersThe Premier EHEC test kit is a rapid in vitro microwell EIA for the detection of Shiga Toxins I and II ( verotoxins) in stool specimens, broth cultures, from individual colonies or colony sweeps of agar plates. The Premier EHEC test is intended for use as an aid in the diagnosis of enterohemorrhagic E. Coli (EHEC) infections. | Premier EHEC Kit | 2 | November 17, 2011 | 2011 | Meridian Bioscience Inc | Internal testing has shown that the indicated kits produce false negative results when tested with some direct stool specimens. These false negative reactions are not observed for all devices tested, however, the rate of occurrence will likely impact clinical sensitivity as defined by product registration and associated instructions for use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Meridian Bioscience, Inc. sent an Urgent Medical Device Recall letter dated October 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, evaluate testing performed with the affected product., and segregate and destroy remaining inventory.Customers were asked to contact Meridian's Technical Service Department by phone or facsimile as necessary. For replacement product phone 800-343-3858, outside North America 513-271-3700, Fax 513-272-5432.Customers were also asked to complete and return the attached Customer Response Form. In the event that they have no remaining inventory, customers were instructed to indicate such on the Customer Response Form and return to Meridian Technical Service. For questions phone 800-343-3858, outside North America 513-271-3700. |  | 3236 | 3236 Kits | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NC, NH, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI and WV and the countries of AUSTRALIA, CANADA, CHILE, COSTA RICA, ITALY, JAPAN and NEW ZEALAND. | Terminated | February 19, 2012 | 94 | 3236 | Not_Computer
Z-0204-2012 | 60392 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SUPER TORQUE MB Angiographic Catheter product is labeled in part: "***5F (1.65 mm)***110 cm .038" (0.97 mm)***PIG***Max. 1200 psi (8274 kPa)***Markers***Open End 6 Sideholes***Angiographic Catheter*** REF Cat No.***Lot***Use By***Assembled in Mexico Cordis Corporation, 14201 NW 60th Ave. Miami Lakes, Florida 33014, USA***EC***REP*** Cordis Cashel, Cahir Road Cashel, Co Tipperary, Ireland***Cordis***a Johnson&Johnson; company***"UPN H739532598A3, Cat No 532-598A, No. MB 10, French Size 5, Shape (PIG), Length cm 110, Flow Rate, ml/sec, 18, Sideholes 6. UPN H739532598B3, Cat No 532-598B, No. MB 20, French Size 5, Shape (PIG), Length cm 110, Flow Rate, ml/sec, 18, Sideholes 6. UPN H73953298C3, Cat No 532-598C, No MB 20 French Size 5, Shape (PIG), Length, cm 65, Flow Rate, ml/sec, 30, Side Holes 6.UPN H73953298D3, Cat No 532-598D, No MB 2 French Size 5, Shape (PIG), Length , cm 70, Flow Rate, ml/sec, 30, Side Holes 10.Cordis SUPER TORQUE MB angiographic catheters with Market Bands are designed to provide angiographic visualization and linear measurement of the vasculature when combined with the delivery of radiopaque contrast media to selected sites in the vascular system. | SUPER TORQUE MB Angiographic Catheter | 2 | November 17, 2011 | 2011 | Cordis Corporation | Cordis Corporation is issuing a Correction for their SUPER TORQUE MB Angiographic Catheter after receiving a report of marker bands dislodged during a procedure. Model Numbers 532598A, 532598B, 532598C, 532598D, SRD5724MB and SRD5727MB are affected by the Cordis Correction. | PRODUCTION CONTROLS: Error in Labeling | Cordis Corporation sent a Correction letter dated November 21, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers are instructed to read the "Description" and "Recommendations" sections, sign and return an enclosed Acknowledgement Form, and to maintain awareness of the communication until the information has been incorporated into the Cordis Super Torque MB Catheter labeling. Contact Cordis at 800-781-0282 for questions regarding this notice. |  | 118690 | 118,690 units | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | July 23, 2013 | 614 | 119000 | Not_Computer
Z-0214-2012 | 57585 | catheter care tray (kit) | OHR | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | Enforcement Discretion | Angiotech***HSG Procedure Tray 5Fr***1 - 18ga x 1-1/2" Needle; 1 - Fenestrated Drape; 1- 20cc Syringe L/L; 1 - Povidone-Iodine Ointment (1g); 2 - Povidone-Iodine Solution (1oz); 1 - Sanitary Napkin; 2 - Sponge Sticks (6"); 1 - Sound Dilator; 1 - Autobag; 1 - Lubricating Jelly; 2 - Cotton Tipped Swab (8"); 1 Extension Tubing (30"); 2 - 3"x3" Sponges (12-Ply); 1 - Vaginal Speculum; 1 - HSG Catheter (5Fr); 1 - Fixed Volume Syringe - 2cc; 1 - CSR Wrap (30"x30")***Catalog Number: 660005000***Sterile - For single patient use. STERILE EO. Caution: Federal (USA) Law restricts this device to sale by or on the order of a physician.MEDICAL DEVICE TECHNOLOGIES, INC. 3600 SW Ave. Gainesville, FL 32808 U.S.A. Domestic: 877-991-1110, International: 732-626-6466, customer_service@angio.com, www.angiotech.com.HSG Catheters are intended for use in the injection of contrast material in the examination of the uterus and fallopian tubes. | HSG Procedure Tray 5Fr | 2 | November 18, 2011 | 2011 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc. dba Angiotech, is recalling HSG Procedure Trays and Bone Biopsy Trays distributed between December 2007 and December 14, 2010 for package integrity. | PRODUCTION CONTROLS: Packaging | The firm, Angiotech, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 22, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review their current inventory, segregate the affected lot numbers and complete and return the attached Customer Acknowledgement Form via fax to our Quality Assurance department at 1.352.338.0662 or 1.800.333.0440 to arrange for return of the identified product.If you have any questions concerning this notification, call (352) 338.0440 ext. 355 or ext 358. |  | 2902 | 2902 boxes | Worldwide distribution: US (nationwide) and countries of: Canada, Columbia, Costa Rica, Japan, Mexico, Peru and Saudi Arabia. | Terminated | November 30, 2011 | 12 | 32491 | Not_Computer
Z-0213-2012 | 60101 | densitometer, bone | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DICOM Option for the Norland Illuminatus Software Revs 4.2.0 though 4.3.1. Models Norland XR-600, XR-800, XR-46, EXCELL, XR-36 and Eclipse DXA Bone Densitometers.The XR 46 performs DXA scans of the AP Spine, Hip, Forearm, Lateral Spine, and Whole Body, as well as other user selectable sites. It provides BMD (g/cm2), Area (cm2), and BMC (g) values. It compares these values to gender and ethnic matched reference populations and provides T-Score and % young Reference, Z- Score and % age matched, and long term and short term change values. This includes sBMD (mg/cm2). The XR 46 performs soft tissue assessment and provides lean mass, fat mass, percent fat, and total soft tissue values for all scan sites, including Whole Body. The XR 46 includes a Report writer that allows the operator to customize the appearance of their report stand to automate the tedious part of the report. It also allows the user to enter ranges of patient values and the specific statements they want to be printed on their report for each of these ranges. The bone density measurements from the XR 46 can be used as an aid to physicians in determining fracture risk. | DICOM Option for the Norland Illuminatus Software Revs | 2 | November 18, 2011 | 2011 | Cooper Surgical, Inc. | The DICOM Software, Norland Illuminatus DXA Software with the DICOM Option (Part Number 434A133) version 4.2.0 through 4.3.1, which is installed on the XR-600, XR-800, XR-46 and EXCELL DXA Bone Densitometers, could lead to improper patient information getting into a previous patient's record. | DESIGN: Software Design | The firm, CooperSurgical, sent a "Notice of RECALL-URGENT" letter dated September 23, 1011 to its customers. The letter described the product, "Problem/Situation", and "When can this problem occur?". CooperSurgical will send a new revision of Illuminatus software to the customers once available.If you have any concerns about your existing database of patients, please contact our Norland Technical Support team at (203) 601-5200 Extension 9888 or 9889. |  | 184 | 184 | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, FL, GA, IN, KS, KY, LA, MD, MA, MI, MO, NJ, NY, NC, OK, PA, TN, TX, UT, and WI ; and countries of: ARGENTINA, AUSTRALIA, BELGIUM SPAIN, CANADA, CHINA, DENMARK, EMIRATES, ENGLAND, FRANCE, GERMANY, GREECE, IRELAND, MALTA, POLAND, SAUDI ARABIA, SOUTH KOREA, and TAIWAN. | Terminated | April 09, 2013 | 508 | 184 | Software | Output/Calculation | Software update | Software Update
Z-0210-2012 | 59818 | stimulator, autonomic nerve, implanted for epilepsy | LYJ |  |  | Neurology |  | VNS Therapy Aspire HC Generator and VNS, Model 105Cyberonics, Inc. 100 Cyberonics Blvd., Houston, TX 77058For implantation in humans to treat Epilepsy and Depression. | VNS Therapy Aspire HC Generator and VNS | 2 | November 18, 2011 | 2011 | Cyberonics, Inc | The devices are being recalled because the output current delivered to the vagus nerve is less than the design intent and there is a potential charge imbalance at the lead cathode and generator-can during stimulation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cyberonics sent a Field Safety Alert on August 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm decided to recall and initiated contact with implanting facilities in the United States on 08/15/2011. A Field Safety Alert was delivered on 09/19/2011 to implanting facilities. Customers were instructed not to prescribe AspireHC pulse generators for implant until Cyberonics resumes shipment of AspireHC pulse generators that have the issue remedied. Cyberonics recommends that physicians continue their current practice of titrating programming parameters for patients implanted with an AspireHC pulse generator to achieve appropriate clinical outcomes, paying particular attention to report of:Lack or loss of efficacyPain, inflammation, swelling/edema in the neck area; and/orPatient perception of stimulation or muscle twitching in the area surrounding the implant. Cyberonics recommended routine systems diagnostic testing to aid in device and patient monitoring. Physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com.For questions regarding this recall call 281-228-7323. |  | 340 | 340 units | Worldwide Distribution - USA including TN, TX, IL, OH, FL, CA, CO, OR, AZ, NY, PA, MO, NJ, WI, OK, MI, ID, KY, IA, MA, NE, GA, LA, NY, and Norway. | Terminated | August 20, 2012 | 276 | 340 | Not_Computer
Z-0211-2012 | 60154 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Results Management (RM) Software Versions RM 2.04.37.04 to RM 7.8 SU2The IMPAX Cardiovascular suite is a cardiovascular information system providing image archiving and image display and modality study-specific structured reporting. | Results Management (RM) Reporting Module | 2 | November 18, 2011 | 2011 | AGFA Corp. | Improper merging of IMPAX Cardiovascular Results Management (RM) reports causes the signed report to become unavailable. | DESIGN: Software Design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 27, 2011, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to carry out a merge post finalization of an RM reports. If it is necessary to modify a finalized report, AGFA recommends that they make the modification through one of the following actions:Amend the report through the RM interface, orCall AGFA Service to change the status of the report to "preliminary" and then proceed with the finalization of the report.AGFA issued an Acknowledgment, via FAX- Back that the information was received and understood to affected customers. For any questions customers were instructed to contact their AGFA Healthcare Service at 877-777-2432. |  | 323 | 323 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 02, 2013 | 592 | 323 | Not_Computer
Z-0212-2012 | 60165 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) The IMPAX Cardiovascular suite is a cardiovascular Information system providing image archiving and image display and modality study-specific structured reporting. | IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) | 2 | November 18, 2011 | 2011 | AGFA Corp. | Incorrect entries in mapping files could lead to inaccurate measurement display in adult and pediatric echo reports when using the IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR). |  | The firm, AGFA HealthCare" sent an 'URGENT FIELD SAFETY NOTICE" dated May 27, 2011 via FED-EX by June 6, 2011 to all customers using the affected product. The letter describes the product, problem and actions to be taken. The customers were instructed to distribute this information within their facility to all those who need to be aware and to complete and return the attached FEEDBACK FORM via fax to 864-421-1668 or email to debbie.norris@agfa.com. AGFA will provide software corrections to all affected sites and will contact the customers to schedule the analysis and provide any corrections if required.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. | Contact the recalling firm for information | 280 | 280 | Worldwide distribution: AK, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WV; and country of: Canada. | Open |  |  | 280 | Software | Output/Calculation | Software update | Software Update
Z-0247-2012 | 60021 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Gamma Camera System;SKYLIGHT 8FT, 3/8" SPECTManufactured by Philips Medical Systems (Cleveland), Inc.595 Miner Road, Cleveland, OH 44143.The SKYLight Gamma Camera System is intended to product images depicting the anatomical distributions of single photon and positron emitting radioisotopes within the human body for interpretation by medical personnel. | SKYLight Gamma Camera System | 2 | November 19, 2011 | 2011 | Philips Medical Systems | A mechanical malfunction of the button on the hand controller on the Philips Precedence and SKYLight SPECT systems may get stuck and cause unintended motion of a detector. | DESIGN: Device Design | Philips Medical Systems sent a "Customer Information" letter dated September 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter lists additional precautions the customers should use while operating the device. Philips will notify customers for replacing the affected product through an Action for Performance-Proactive Field Change Order to resolve the issue. Contact Philips Customer Care Solutions at 1-800-722-9377, option 5, for questions concerning this notice. | Contact the recalling firm for information | 658 | 658 - all units in recall | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | December 21, 2012 | 398 | 658 | Not_Computer
Z-0259-2012 | 60173 | suture, surgical, absorbable, polydioxanone | NEW | Absorbable polydioxanone surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Product is a synthetic absorbable suture with needle and labeled as sterile. The product is sealed inside an inner foil pouch with an outside Tyvek pouch. The outer sterile barrier seal is intact. The product is boxed and shrink wrapped for protection against contamination. The product is a Violet, monofilament synthetic absorbable suture with needle, size 4/0 (1.5 Metric), NFS-2 Needle (Product code M397-Polydioxanone) or size 3/0 (2.0 Metric) (Product code M398 {Butler code 029254}).Product is labeled in part: "***POLYDIOANONE Violet (PDO) Monofilament Absorbable Suture***".Intended for use in general soft tissue approximation and/or ligation including use in pediatric cardiovascular tissue where growth is expected to occur, and ophthalmic surgery. The suture is not recommended for adult cardiac tissue, microsurgery or neural tissue. The absorbable sutures are useful where extended wound support is desirable. | MonoDox Synthetic Absorbable Suture | 2 | November 19, 2011 | 2011 | C P Medical Inc. | The seal integrity of the outer product pouch may be compromised due to partial delamination of the two layers that make up part of the pouch. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CP Medical, Inc. sent a "RECALL NOTICE" letter dated January 5, 2009 to all affected customers.The letter describes the product, problem, and the actions to be taken by the customers. Customers are instructed to identify affected product and to return the product to the firm. Customers should call 1-503-232-1555 for a return goods authorization number and shipping account number. Contact your customer service representative at 1-800-950-2763 for questions concerning this recall. | Contact the recalling firm for information | 13632 | 13632 sutures | Nationwide Distribution | Terminated | December 13, 2011 | 24 | 13632 | Not_Computer
Z-0261-2012 | 60405 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. SoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable, trained, and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities, for the purpose of storing, managing, querying, and reporting laboratory data. | SoftLab | 2 | November 21, 2011 | 2011 | SCC Soft Computer | On 08/08/2011 SCC Soft Computer, Clearwater, FL initiated a correction on following SoftLab used with: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. A client reported a flagging error where the result was flagged as critical in the Instrument Menu but the printed report did not show the corresponding flag. All affected clients were notified of the issue on 0 | DESIGN: Software Design | SCC Soft Computer notified all affected customers on August 8, 2011, of the recall via the SCC Soft Computer proprietary Task management system to manage client communication. Clients use this tool to log in requests for services and receive and provide updates to any issue. Each client user has a unique password and every action is date and time-stamped. Each Task has a unique identifier and a full audit trail.SCC recommended to the customer until Live on the correction avoid making edits in the instrument menu.For further questions please call (727) 789-0100. |  | 255 | 255 | Worldwide Distribution - USA ( nationwide) and Canada. | Terminated | November 22, 2013 | 732 | 255 | Not_Computer
Z-0265-2012 | 60301 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Stryker Endoscopy 3.55 MM Super 90 S-SERFAS Energy Probe; electrosurgical device. Model number: 279-351-300; | SERFAS | 2 | November 21, 2011 | 2011 | Stryker Endoscopy | An increased number of complaints for tip breaks were observed for the 3.5mm Super 90-S SERFAS Energy Suction Probes manufactured from June to August 2011. | DESIGN: Component Design/Selection | Stryker Endoscopy Urgent Medical Device Removal notifications datyed October 25, 2011, were sent to all consignees via traceable mail. Sales representatives and regional managers were notified by e-mail. Product is to be returned to Stryker Endoscopy using the pre-paid shipping label. Replacement product may be obtained by calling Stryker at 1-800-624-4422. The enclosed Acknowledgement receipt form is to be completed and returned. Questions or comments may be directed to Stryker at 1-800-624-4422 or e-mail at 90S@stryker.com. | Contact the recalling firm for information | 6307 | 6307 | Worldwide distribution including USA, Australia, Canada, China, Greece, Hong Kong, India, Italy, Japan, France, Netherlands, Poland, South Africa, Spain, Sweden, Switzerland and the United Kingdom | Terminated | February 29, 2012 | 100 | 6307 | Not_Computer
Z-0262-2012 | 60205 | degreaser, skin, surgical | KOY | Surgical skin degreaser or adhesive tape solvent. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | PDI Electrode Prep Pads, Skin Prep B339 are non-sterile alcohol prep pads. They are included in the Spacelabs Healthcare Hook-Up kits. Each kit contains 2-4 pieces of PDI Electrode Prep Pads . Hook-up kits that contain the affected alcohol prep pads are as follows:Aria, Kit Part number 23189-010, Kit, Basic hookup, AA;CardioCall, Kit Part number 23189-020, Kit, Hookup, CardioCall;Aria, Kit Part number 23189-106, Kit, Aria Hookup;Aria, Kit Part number 23189-110, Kit, Aria Hookup;Flashcorder, Kit Part number 23189-111, Kit, Flashcorder Hookup;Aria, Kit Part number 23189-113, Holter Hookup Kit 7 electrode;LifeCard 12, Kit Part number 23189-114, Kit, Hookup, LC12, Multilingual;Aria, Kit Part number 23189-115, Kit, Basic Hookup, AAA;LifeCard, Kit Part number 23189-116, Kit, Holter Hookup, LifeCard;483, Kit Part number 23189-117, Kit, Holter Hookup 483;Aria, Kit Part number 23189-153, Holter Hookup Kit, 7 electrode;Aria VLC, Kit Part number 23189-207, Kit, Aria Hookup, VLC;EVO, Kit Part number 50-0048, 4 electrode hookup kit without pouch;EVO, Kit Part number 50-0049, 4 electrode hookup kit with pouch;EVO, Kit Part number SK10456, Sample Kit, Disposables.A surgical skin degreaser or an adhesive tape solvent is a device that consists of a liquid such as 1,1,2-trichloro-1,2,2-trifluoroethane; 1,1,1-trichloroethane; and 1,1,1-trichloroethane with mineral spirits intended to be used to dissolve surface skin oil or adhesive tape. | Spacelabs Healthcare HookUp kits | 2 | November 21, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | The non-sterile alcohol prep pads (B339) are included in Spacelabs Healthcare hook-up kits may have the presence of a bacterium, Bacillus cereus. These alcohol prep pads were manufactured and recalled by Professional Disposables International (PDI). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | SpaceLabs Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the affected product and dispose of them per their hospital or clinic protocol. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2for Technical Support. |  | 25015 | 25,000 pieces of PDI alcohol prep pads-15 different Spacelabs hook-up kits | Worldwide Distribution - USA (natiowide) and the countries of South Africa, Turkey, Bangladesh, Dominican Republic, Peru, Belgium, Ecuador, Romania, Austria, Honduras, Guatemala, Denmark, Republic of Korea, Argentina, Thailand, Italy, Austria, Chile, Bolivia, Mexico, Sweden, Netherlands, Switzerland, India, Serbia, Australia, Morocco, Turkey, ISpain, Czech Republic, Slovakia, Poland, China, Columbia, Portugal, Canada, Germany, United Kingdom, Singapore, France, Russian Federation, Bahrain, and Kuwait | Terminated | February 23, 2012 | 94 | 25015 | Not_Computer
Z-0264-2012 | 60365 | dilator, vessel, for percutaneous catheterization | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Super Sheath and Super Sheath R/O Introducer Sheaths, Sterile EO, Manufactured by Togo Medikit Co., Ltd. 17148-6 Aza Kamekawa, Oaza Hichiya, Hyuga City, Miyazaki Prefecture 883-0062, Japan, Distributed by Boston scientific Corporation One Boston Scientific Place, Natick, MA 01760-1537 USA.Super Sheath 8F x 11 cm ,038 with Guide Wire BX/10 UPN H7491603508B1, Catalog # 16035-08B. Batch/Lot 11B15B6, 11B23B7, 11B28B9, 11C09BA, 11C16B9, 11C30B9.Super Sheath 8F x 11 cm ,038 without Guide Wire BX/10 , U PN M00115724B1, Catalog # 15-724B1. Batch/Lot 11C07B8.Super Sheath R/O 8 F x 11 cm without Guide Wire BX/10, M00159664B1, Catalog # 15-964B, Batch/Lot 11B01B4, 11B25B6.The device is indicated for use in the introduction of diagnostic and interventional devices inserted into the human vasculature. | Boston Scientific Super Sheath and Super Sheath R/O Introducer Sheaths | 2 | November 21, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating a medical device recall removal regarding nine lots/batches of Super Sheath and Super Sheath R/O Introducer Sheaths. Boston Scientific was informed through product complaints that the 0.035" ID dilator may be labeled incorrectly as a 0.038" ID dilator. | MISBRANDING: Labeling False and Misleading | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED letter dated November 21, 2011. The letter described the product and the problem. Customers were instructed to segregate the affected product immediately and return it to Boston Scientific. Distributors were advised that the recall removal depth is to the hospital level and the recall removal notification should be forwarded to their customers.For questions regarding this recall call 763-494-1133. |  | 2660 | 2660 | Worldwide Distribution - USA including AZ, AR, CA, CO, FL, GA, IL, KY, LA, MD, MA, MN, MO, NE, NJ, NM, NY, OH, OK PA, RI, TN, TX, UT, VA, WY, WI, and WV and the countries of BELGIUM, BRAZIL, CZECH REPUBLIC, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, ITALY, NETHERLANDS, RUSSIAN FEDERATION, SOUTH KOREA, SPAIN, SWEDEN, and UKRAINE. | Terminated | January 15, 2013 | 421 | 2660 | Not_Computer
Z-0270-2012 | 60278 | chair, adjustable, mechanical | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | "***UPLIFT TECHNOLOGIES INC.***Bringing quality to life.***Model No.***CA200***Description***Lifting commode chair***Weight Range***80-300+ lb-136+ kg***"Labeling on product: "***1C0101 and K0061014***".A lifting commode chair that is self-powered (non-electric) and provides lifting assistance for those who have reduced mobility due to arthritis, post-operative conditions, degenerative joint disease, Parkinson's disease or other neuromuscular disorders. | Uplift Commode Assist | 2 | November 22, 2011 | 2011 | Uplift Technologies Inc. | On June 26, 2011 Uplift Technologies Inc. Nova Scotia, Canada initiated a recall of the Uplift Commode Assist Model No. CA200, Lot No. 1C0101. The plastic seat does not meet manufacturing specifications. The seat may not fit the unit properly and may interfere with the unit's normal locking mechanism. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Uplift Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 27, 2011 to all affected customers The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and to quarantine the affected product. Customers are to call Uplift's Customer Service Department at 1-800-387-0896 to request a Return Materials Authorization number in order to return the product.. | Contact the recalling firm for information | 219 | 219 chairs | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 191 | 219 | Not_Computer
Z-0266-2012 | 60005 | source, brachytherapy, radionuclide | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | Rapid Strand Rx; the RAPID Strand Rx kit consists of custom loaded absorbable seeding spacers and I-125 radionuclide brachytherapy sources (seeds) spaced at prescribed distance and configuration within a sleeve (tube) made of absorbable suture material, stiffened, loaded into prostate seeding needles, packaged and then sterilized by Gamma sterilization method, accompanied by a calibration strand, a sealed source calibration certificate (100% assay report), a radiograph of the loaded needles in the needle-holding card, and a source certification technical data sheet.; Products Provided by: Oncura, 3350 North Ridge Avenue, Arlington Heights, IL 60004 USA, Prescription Filled by: AnazaoHealth, 5710 Hoover Blvd., Tampa, FL 33634 USA; product identifier: I125RSRX, model 7000.RAPID Strand Rx is indicated for permanent interstitial implantation of selected localized tumors that are of low to moderate radio sensitivity. It may be used either as primary treatment (such as prostate cancer or unresectable tumors) or for the treatment of residual disease after excision of the primary tumor. RAPID Strand Rx may be indicated for use concurrent with or at the completion of other treatment modalities such as external beam radiation therapy. | Rapid Strand Rx | 2 | November 22, 2011 | 2011 | Medi-Physics Inc. dba GE Healthcare | There is a potential brachytherapy over or under dose with the RAPID Strand Rx kit that may impact patient therapy. | PRODUCTION CONTROLS: Process Control | GE Healthcare/Oncura sent a "Urgent Medical Device Correction" letter dated September 20, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were advised to assess the pre and post plan dosimetry for their patient, taking into account the possible variation of the seed strength. GE Healthcare/Oncura will contact the customers in the near future to confirm their receipt of the letter and to discuss any questions they may have.Please call if you have any further questions (847) 398-8400 and 011-44-1494-737251 (Amersham, United Kingdom, GMT). | No consumer action necessary | 23 | 23 orders | Worldwide Distribution -- USA (nationwide) including the States of Arizona, Florida, Louisiana, Mississippi, New Jersey, New York, Pennsylvania, Puerto Rico, Virginia and Washington, and the countries of Canada and the United Kingdom | Terminated | January 20, 2012 | 59 | 23 | Not_Computer
Z-0271-2012 | 60374 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TUM-E-VAC REF 2075, Rx ONLY, NON STERILE --- Ethox International, Buffalo, NY 14204, USA --- Intended use: gastric lavage | TUMEVAC | 3 | November 22, 2011 | 2011 | Ethox International, Inc. | Ethox Tum-E-Vac Gastric Lavage Kit #2075, Lot Number 031123255 was labeled with expiration date symbol and date of 2010/10 on the unit product label. The carton was labeled with a manufacturing date symbol and date of 2010/10. The Ethox Tum-E-Vac Gastric Lavage Kit #2075 does not have an expiration date. | TRAINING: Employee Error | Ethox International, Inc. sent an "URGENT DEVICE RECALL" letter dated March 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine if they have the affected product. Customers are to return the affected product to the firm by using UPS account 128959. Customers will be reimbursed for the returned goods. Additionally, a response form was enclosed for customers to complete and return. Contact the firm at (716) 842-4000 for questions regarding this recall. | Contact the recalling firm for information | 130 | 130 units | Nationwide Distribution-including the states of California, Connecticut, Illinois, Ohio, and Virginia. | Terminated | November 28, 2011 | 6 | 130 | Not_Computer
Z-0263-2012 | 60155 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Physiological Patient Monitoring SolutionNetworked solution/system used to monitor patients' vital signs and therapy, control alarms, review Web-based diagnostic images, access patient records, and scale up functionality based on patient acuity level. | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) | 1 | November 22, 2011 | 2011 | Draeger Medical, Inc. | During product training on the Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) it was discovered that the weight-based drug dosage calculation may indicate incorrect recommended values. In addition, a misalignment of ECG and ART waveforms delivered from an Infinity(RO Acute Care System (IACS) Monitoring Solution (M540) was observed on the Infinity Central Station. |  | The firm initiated their recall of this device on October 17, 2011 by issuing a letter to the consignee instructing customers to follow the Instructions for Use and use the Infinity Central Station only for remote assessment of a patient's status. They recommend that you do not consider the waveforms on Infinity Central Station for primary diagnosis purposes. In case of a high priority alarm situation, "Check the patient and treat if necessary".If the misalignment issue is observed, it can be cleared by undocking and redocking the M540 or by changing the order of waveforms displayed on the Cockpit screen (or any other change to the auto display screen layout). |  | 476 | 476 | Nationwide distribution - IL | Terminated | April 01, 2013 | 496 | 476 | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0268-2012 | 60316 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Wolf Medical Supply, Inc., WOLF-PAK REDI-FLO ELASTOMERIC INFUSION PUMP, FILL VOLUME 100 mL, FLOW RATE 2mL/Hr. Model # CT-0020-100 Item # RFC100020. Manufactured by Medpro International (Thailand) Ltd. | WOLFPAK REDIFLO ELASTOMERIC INFUSION PUMP | 1 | November 23, 2011 | 2011 | Wolf Medical Supply Inc. | Redi-Flo Elastomeric Infusion Pumps may have a higher flow rate than specified (more than +15%). | PRODUCTION CONTROLS: Process Control | On 10/20/2011 Wolf Medical Supply Inc., customers were notified by telephone and email. Product was distributed to end users. |  | 8 | 8 cases (192 units) | Nationwide distribution, including Puerto Rico. | Terminated | April 25, 2012 | 154 | 40 | Not_Computer
Z-0272-2012 | 60164 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CardioVascular (CV) Admin ToolThe IMPAX Cardiovascular suite is a cardiovascular information system providing image archiving and image display and modality study-specific structured reporting. | AdminTool2 | 2 | November 23, 2011 | 2011 | AGFA Corp. | Potential for misattribution of data and for a report to become unavailable when the user requests for it to be moved to another study that already has a report associated with it. | DESIGN: Software Design | An 'URGENT FIELD SAFETY NOTICE" was sent via FED-EX by June 6, 2011 to all the sites using the affected product. The letter described the issue and mitigation. The letter recommends that users not move reports to studies where a report already exists. If a report is being moved because the existing report is irrelevant to the study, then the existing report should be deleted prior to moving the new report to that study. If the report is being moved because the user wants to combine the content of the reports, then the "merge" function should be utilized. The letter also states that a software enhancement is currently available. Acknowledgment via FAX-Back was requested from the sites. Questions regarding the issue should be directed to customers' local Agfa HealthCare Service at 877-777-2432. The firm has identified 185 additional sites and an "URGENT FIELD SAFETY NOTICE" was sent on October 12, 2012. |  | 323 | 323 units | Worldwide Distribution -- USA, Virgin Islands and Canada | Terminated | June 26, 2012 | 216 | 323 | Software | Output/Calculation | Software update | Software Update
Z-0273-2012 | 60123 | sucker, cardiotomy return, cardiopulmonary bypass | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(k) | IBC INTERNATIONAL BIOPHYSICS CORP. 2101-2 EAST ST. ELMO, SUITE 275, AUSTIN, TX 78744 SUCTION HANDLE (STERILE) PART NUMBER: 1990S LOT NUMBER 041811-2065 EXPIRATION DATE: 18/APR/2015For use in open heart procedures, Autologus blood recovery or blood suction. Product is a component of Blood Recovery Systems. | IBC Suction Wand | 2 | November 23, 2011 | 2011 | International Biophysics Corp | Potential for the sterility to be compromised. | PRODUCTION CONTROLS: Packaging | International Biophysics Corp. sent a" Medical Device Recall Notice" dated October 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Use of the product in the referenced lot number should cease immediately. Customers were asked to follow the instructions in the letter to return the affected product.If you have any further questions please call (512) 814-0070. | Contact the recalling firm for information | 1200 | 1200 units | Nationwide Distribution including the states of AZ, CA, CO, DE, MS, OR, PA, TN, and TX | Terminated | March 23, 2012 | 121 | 1200 | Not_Computer
Z-0282-2012 | 59913 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Flexible Probe with Blocking Water, GM11002420, (Part of the Segmented Cylinder Set, GM11004150) | Flexible Probe with Blocking Washer | 2 | November 23, 2011 | 2011 | Varian Medical Systems, Inc. | The Instructions For Use ( IFU) distributed with product prior to 2004 lacks the correct limit to the sterilization cycles and checks regarding the firm seating of the blocking washer. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Varian Medical Systems sent a Medical Device Recall (PNL-GM-42697-2) in May 2010 for the blocking washer on the flexible probe, Varian part number GM11002420, a component of the Segmented Cylinder Applicator Set. The letter was to inform users that a flexible probe with an unstable or shifted blocking washer must not be used due to the potential for treatment length error. The blocking washer may move from rough handling or sterilization. Users were told to discard any instances of the obsolete Instructions for Use (IFU) in use and replace it with the IFU (PGM11004150 4 March 24, 2010.pdf). A Receipt Verification Card was included requesting consignees to fill out, fax or e-mail immediately to Varian BrachyTherapy. | Contact the recalling firm for information | 81 | 81 parts | Worldwide Distribution: USA, Germany, Belgium, Canada, Mexico City, Spain, Brazil, Switzerland, Uzbekistan, India | Terminated | December 05, 2011 | 12 | 81 | Not_Computer
Z-0278-2012 | 60430 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. | Softlab | 2 | November 23, 2011 | 2011 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. |  | 12 | 12 | Sold throughout US and Canada. | Terminated | October 29, 2012 | 341 | 12 | Software | Output/Calculation | Software update | Software Update
Z-0279-2012 | 60305 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085SRT Surgical Tables, Catalog Numbers: 5085- ST01-410-1; 5085SRT- ST0-1420-6. | STERIS 5085 and 5085SRT surgical tables | 2 | November 23, 2011 | 2011 | Steris Corporation | Some hydraulic column cylinders installed in certain 5085 and 5085SRT Surgical Tables were assembled by their supplier with incorrect snap rings and an anomaly was present in the tilt cylinder. This error may affect the user's ability to move the table top out of the full right tilt position. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On October 17, 2011, Steris sent 'Urgent Voluntary Field Correction Notices' to their customers via Fedex with tracking numbers for delivery confirmation. The letter identifies the affected product and notifies the user/customers of the problem with the surgical tables. The notification letter goes on to inform the customers that a STERIS Service Representative will be contacting them to arrange for the field correction of the affected unit(s). This involves the replacement of the hydraulic column cylinder of each surgical table(s). |  | 10504 | 167 tables were distributed - 5085=108 and 5085SRT=59 | Nationwide distribution. | Terminated | March 19, 2013 | 482 | 167 | Not_Computer
Z-0281-2012 | 60269 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity(R) Delta, Delta XL, Gamma XXL and Kappa Monitors are used to monitor, physiologically, the vital signs of patients (with Arrhythmia Detection or Alarms). | Draeger Infinity(R) Delta, Delta XL, Gamma XXL and Kappa Monitors | 2 | November 23, 2011 | 2011 | Draeger Medical, Inc. | Firm received 2 complaints reporting Ventricular Fibrillation (Vfib) arrhythmias were not recognized by the monitor and did not alarm when the Masimo SET pod is used for SpO2 measurements. | DESIGN: Software Design | Firm initiated their recall of this device on October 25, 2011 by sending an "Urgent Medical Device Recall" letter to consignees. A Technical Service Bulletin will be generated and distributed globally that instructs the affected International Draeger subsidiaries to inform customers in their countries. The software anomaly has been corrected in software version 7.3. Draeger Service Representatives will schedule time with customers to upgrade the software in affected monitors. The software upgrade will be provided free of charge. |  | 14894 | 14,894 | Worldwide distribution- USA including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SD, TN, TX, UT, VA, WI and WV and International distribution including the countries of Algeria, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Botswana, Brasilien, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Deutschland, Dominican Republic, Ecuador, Egypt, El Salvador, Equatorial Guinea, Estonia, Finland, France, Frankreich, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indien, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Jordanien, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Madagascar, Malaysia, Malta, Mauritius, Mayotte, Mexico, Moldavia, Moldova, Morocco, Netherlands, Norway, Norway, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Rumania, Russian Fed., Saudi Arabia, Serbia, Seychelles, Singapore, Slovenia, Slowenien, South Africa, South Korea, Spain, Sri Lanka, SÂŽdafrika, Sweden, Switzerland, Syria, Taiwan, Thailand, Tschechische Republic, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and White Russia. | Terminated | July 01, 2013 | 586 | 14894 | Software | Alarm/Message | Software update | Software Update
Z-0276-2012 | 59980 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic |  | Sensar Foldable IOL with OptiEdge. Model number AR40e 21.5 Diopter.The Sensar Foldable IOLs with OptiEdge Design are ultraviolet-absorbing posterior chamber intraocular lenses and are indicated for the visual correction of aphakia in adult patients in whom a cataractous lens has been removed by extracapsular cataract extraction. The lens are intended to be placed in a capsular bag. | Sensar Foldable IOL with OptiEdge | 2 | November 23, 2011 | 2011 | Abbott Medical Optics Inc (AMO) | The firm is initiating this voluntary recall because of the possibility that certain Sensar IOLs are labeled with an incorrect diopter power. | TRAINING: Employee Error | Abbott Medical Optics, Inc. sent an URGENT FIELD SAFETY NOTICE dated September 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and remove the affected product from their inventory, complete and return the included facsimile form to Abbott Medical Optics and indicate the status of each affected serial number on the form. Customers were asked to return the form even if they have no inventory affected by the recall. For any questions customers were instructed to contact their Abbott Medical Optics Customer Service Representative.For any questions regarding this recall call 714-247-8656. |  | 96 | 96 | Worldwide Distribution including China, India, Indonesia, Korea, Malasyia, Taiwan, Thailand, Austria, Finland, France, Germany, Great Britain, Hungary, Italy, Netherlands, Portugal, and Saudy Arabia | Terminated | January 08, 2014 | 777 | 179 | Not_Computer
Z-0274-2012 | 60328 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSDV), sterile EO, Rx Only, Synovis Surgical Innovations, A Division of Synovis Life Technologies, Inc, 2575 University Ave W., St Paul, MN 55114-1024.The device is intended to be used as a staple line buttress | PeriStrips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSDV) | 3 | November 23, 2011 | 2011 | Synovis Surgical Inovation Div. of | Synovis Surgical Innovations is voluntarily recalling one lot of Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) product because some of the devices were mislabeled. | PRODUCTION CONTROLS: Labeling Mix-Ups | Synovis Surgical Inovation sent a "Recall Notification" letter dated October 26, 2011. The letter described the problem and product affected. Customers were instructed to identify and quarantine the affected product and to complete and fax the Product Return/Replacement Form. | Contact the recalling firm for information | 132 | 132 units | Nationwide Distribution including CT, GA, NJ, NY, OH, TX, VA. | Terminated | April 24, 2012 | 153 | 132 | Not_Computer
Z-0284-2012 | 60142 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System. | Symbia T Series SPECTCT System | 2 | November 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Errors could potentially impact image interpretation. The acquisition software is exhibiting a defect in the energy window position when an isotope is peaked. There is a potential patient safety issue when running syngo MI Applications version VA60A with Foresight detectors firmware version v1.5.9.12 on the Symbia systems. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated October 14, 2011 to all affected customers. The letters informed customers of the acquisition software defect in the energy window position when an isotope is peaked, which could potentially impact image interpretation, and that the AEB firmware version on their imaging system has been changed to a previously released version until acquisition software service pack is available. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at (800) 888-7436. | Contact the recalling firm for information | 3 | 3 units | Florida, Illinois and Texas | Terminated | October 16, 2012 | 328 | 5 | Software | Physical Safety Hazards | Software update | Software Update
Z-0280-2012 | 60219 | antinuclear antibody, indirect immunofluorescent, antigen, control | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostic Inc., NOVA Lite ANA KSL KitPart Number: 708390. in vitro diagnostic. | NOVA Lite ANA KSL Kit | 3 | November 23, 2011 | 2011 | Inova Diagnostics Incorporated | The NOVA Lite ANA KSL kit label has the wrong fill volume for the IFA System Negative Control (P/N 508186) with 1mL. The component in the kit is correctly labeled and filled with the 0.5mL fill volume. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | On 9/26/11, Inova Diagnostics forwarded an Urgent Advisory Notice and Product Advisory Notice Fax Back Form to all their customers who purchased the Nova Lite Ana KSL, Product Code: 708390. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete the attached fax back form indicating how many of these kits they have in their inventory, quantity destroyed, method of destruction and the quantity of kits that need to be replaced. Customers were also instructed to notify their customers of the advisory notice. Customers were instructed to contact Tana Keivens with any questions or comments at (858) 586-9900 ext 351. |  | 500 | 500 kits | Worldwide distribution to: USA, Mexico, Kuwait, Uraguay, Romania, Italy, Philippines, India, Malaysia, Algeria, Czechoslovakia, Brazil, Colombia, Ireland, Korea, and Saudi Arabia. | Terminated | April 04, 2012 | 133 | 500 | Not_Computer
Z-0277-2012 | 60436 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. SoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable, trained, and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities, for the purpose of storing, managing, querying, and report laboratory data. | SoftLab GUI versions 4.0.2.84.0.2.10; 4.0.3.124.0.3.13; and 4.0.4.14.0.4.4. | 2 | November 23, 2011 | 2011 | SCC Soft Computer | On 05/03/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the following SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. A client reported the PT and INR results did not get flagged. The PROT should have been flagged as HIGH and the INR should have been flagged as PANIC. Neither test was flagged. All affected clients were notified of the issue on 05/03 | DESIGN: Software Design | SCC Soft Computer initiated a Field correction on May 3, 2011. All customers with affected software have been notified of this issue. Clients were instructed to have a "valcheck" utility run to determine if there are any tests in their system affected by the issue. They were supposed to indicate whether they wanted to schedule to receive the patch or hot fix.If you have any further questions please call ( 727 ) 789-0100. |  | 69 | 69 | Worldwide Distribution-- USA ( nationwide) and Canada. | Terminated | August 07, 2013 | 623 | 69 | Software | Output/Calculation | Software update | Software Update
Z-0285-2012 | 60166 | film, radiographic | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Curix Opthos H, size 35x43 cmA radiographic medical x-ray film designed to optimize diagnostic chest imaging by providing complete diagnostic information including superbly visualized lung detail along with a high level of retro-cardiac, sub-diaphragmatic and mediastinal information. | Curix Opthos H | 3 | November 25, 2011 | 2011 | AGFA Corp. | The film notch was located in the wrong position. | PRODUCTION CONTROLS: Process Control | An "URGENT SAFETY NOTICE" letter was sent via email to the dealers on October 7, 2011. The letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the dealer sites. Included in each dealer letter, Agfa also provided a customer notification letter to be sent to their end-use customers. Agfa has requested dealers and end-use customers to file a complaint at 1-877-777-2432 for any identified product still in their inventories. Agfa will replace product and arrange for return shipment. |  | 125 | 125 | Worldwide Distribution -- USA, including the states of IN, KS, MI, MN, MO, NC, NE, NY, OH, and TN and the country of Canada. | Terminated | December 22, 2011 | 27 | 125 | Not_Computer
Z-0292-2012 | 60413 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 150 mL Burette Set, with Calibrated Burette and 2 Clave Ports, Convertible Pin, 77 inch , Soluset Microdrip with Option-Lok; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 19208-01Intended use: for the administration of fluids | Hospira LifeShield LatexFree 150 mL Burette Set | 2 | November 25, 2011 | 2011 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | TRAINING: Employee Error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated November 10, 2011 to the customers of record via UPS on the same date. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-866-913-0667 or email to: hospira2530@stericycle.com; return affected product to Stericycle using the label provided with this letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. | Contact the recalling firm for information | 12580 | 12,580 units | Worldwide distribution: USA (nationwide) including states of: Alabama, Alaska, California, Colorado, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kentucky, Massachusetts, Michigan, Minnesota, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oregon, Pennsylvania, Rhode Island, Tennessee, Texas, Vermont, Virginia, Washington and Wisconsin, and internationally to Canada | Open |  |  | 12580 | Not_Computer
Z-0289-2012 | 60355 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | ImmuneTech brand MyAllergyTest, Specimen Collection Kit, Part No. 8015, Screening Pack: 50 test (with lancet),Product is manufactured distributed by ImmuneTech, Inc., Foster City, CA; but the non-sterile alcohol prep pads are manufactured by H & P Industries, Inc., Triad Group, Hartland, WIThe MyAIIergyTest is for the collection and transport of a capillary blood sample to a clinical laboratory for allergen specific lgE profile testing. | MyAllergyTest, Specimen Collection Kit | 2 | November 25, 2011 | 2011 | Immunetech Inc | ImmuneTech is recalling certain Specimen Collection Kits that contain Triad brand non-sterile alcohol prep pads which have been recalled by Triad Group, Hartland, WI due to Bacillus cereus, microbial contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, IMMUNETECH, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated November 4, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory and quarantine all affected kits; if they have further distributed the affected kits, contact Immunetech to accurately identify and immediately notify their customers and facilitate the return and exchange of all remaining kits. If you have any questions, call (650) 312-1066, Monday-Friday, 9:00am -5:00pm (pacific) or email: customersupport@immunetech.com. |  | 50 | 50 | Worldwide distribution. | Terminated | June 14, 2012 | 202 | 3340 | Not_Computer
Z-0294-2012 | 60031 | scale, stand-on, patient | FRI | Stand-on patient scale. | General Hospital | General Hospital | 510(K) Exempt | Health o Meter Professional Remote Display Digital Physician Scale, Model 498KL; digital patient weighing scale; Pelstar LLC, 11800 South Austin Avenue, Unit B, Alsip, IL 60803, U.S.A., Made n China; Model 498KLModel 498KL scale with remote display head uses sophisticated microprocessor technology. Each instrument is designed to provide accurate, reliable and repeatable weight measurements and features that make the weighing process simple, fast and convenient. The scale is set up to determine the weight of a stationary patient. The weight can be displayed in pounds or kilograms. The unit can be operated by an AC adaptor (optional) or by 6-AA cell batteries (included). | Health o Meter Professional | 2 | November 28, 2011 | 2011 | Pelstar, Llc | There is a potential for an electrical short when installing or replacing the batteries in the remote display head of the scales. This short may cause the batteries to smoke, burn or burst. | DESIGN: Device Design | Pelstar LLC. sent a Urgent Device Recall letter dated October 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informed them of the potential of an electrical short caused when installing or replacing the batteries which may result in minor injuries to the person installing the batteries. The customers were requested to immediately examine their inventory and quarantine all product subject to the recall. In addition, if they distributed this product to end users, they were requested to identify all customers and notify them of this product recall at once in writing, using the enclosed recall notification letter. The customers were also requested to complete and return the enclosed mandatory response form as soon as possible. Any questions were directed to Health O Meter Professional at 800-815-6615 or 708-377-0600. |  | 2498 | 2,498 units | Worldwide Distribution -- USA (nationwide) including the countries of Canada, New Zealand, Israel and Lebanon. | Open |  |  | 2498 | Hardware | Physical Safety Hazards | Remove | Remove or Replace
Z-0293-2012 | 60304 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 1.5T and 3.0T MRI Systems, Model Numbers: 781396 and 781377, 510(k) #K110151. System units equipped with software version: 4.1.1 are subject to recall/field correction. This MRI system is indicated for use as a diagnostic device. It can produce cross-sectional images, spectroscopic images and/or spectra in any orientation of the internal structure of the head, body, or extremities. Image appearance is determined by the spatial distribution and flow as well as by many different NMR related properties of the tissue and anatomy studied and the MRI scan technique applied. These images and/or spectra, when interpreted by a trained physician, provide information that may assist the diagnosis. | Ingenia 1.5T and 3.0T MRI Systems | 2 | November 28, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm determined that during scanning procedures, for some larger patients, some reddening of the skin of can occur in certain instances due to RF heating when the torso of the patient is positioned within the body coil. | DESIGN: Software Design | Philips Medical Systems ( Cleveland) Inc. sent a Urgent Field Safety Notice dated August 3, 2011, to all affected customers. Philips Field Service Engineers were dispatched to the affected customer sites to deliver the recall notification and to perform the necessary software corrections. The letter lists the affected MRI systems; describes the problem; the actions to be taken by the customer/user; it also informs the customer of the upcoming software correction which will be applied to each unit by the Philips Field Service Engineers in order to address this issue. If any additional information is required, the customer is instructed to contact their local Philips Field Service Engineer. For further questions please call (440) 483-7000. |  | 26 | 26 units - 12 units distributed in U.S. & 16 distributed to foreign customers | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, GA, OR, MI, NJ and TX., and the countries of AU, BE, CH, DE, FR, JP, NL and SE. | Terminated | February 22, 2013 | 452 | 26 | Software | Physical Safety Hazards | Software update | Software Update
Z-0297-2012 | 60233 | calibrators, drug specific | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Multigent Vancomycin Calibrators;Manufactured by Microgenics Corp, Fremont, CA.Distributed by: Abbott Diagnostics, Abbott Park, IL Product Usage:The Multigent Vancomycin calibrators are intended for use in the calibration of the multiagent vancomycinVancomycin is a commonly used antibiotic to treat staphylococcus and other gram-positive bacterial infections. | Multigent Vancomycin Calibrators | 2 | November 28, 2011 | 2011 | Microgenics Corp | Complaint from the sole distributor that the third party control material is recovering high, and outside the published range. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Microgenics sent a Product Removal notification letter dated June 15, 2011 to all affected customers. The letter identified the affected product, the problem, and the necessary actions to be taken. Customers were instructed to discontinue use and dispose of the affected product per local waste ordinances, review all patient results reported with the use of this lot, and retain a copy of this notification for their records. The letter states that if the affected product has been forwarded to another laboratory, a copy of the letter is to be provided to them. An alternative calibrator is available to customers impacted. It can be ordered by calling Thermo Fischer Scientific Customer Service at 1-800-232-3342.Questions regarding the information should be forwarded to Abbott Customer Service at 1-877-4ABBOTT. |  | 118839963 | 809 kits lot 59446762; 651 kits lot 59391741 | Nationwide Distribution (USA) in the state of ILLINOIS | Terminated | February 17, 2012 | 81 | 1460 | Not_Computer
Z-0305-2012 | 60121 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5Product Usage: Radiation therapy treatment of malignant neoplastic diseases, as determined by a licensed medical practitioner. | Medical linear acelerator (with patient imaging) | 2 | November 29, 2011 | 2011 | Elekta, Inc. | If "Confirm" or "Unconfirm" settings are clicked during the "Terminated Checking" phase of treatment delivery, Desktop Pro cancels the loaded beam and sets the Beam MUs to zero. Desktop Pro then sends the delivered MUs to MOSAIQ a a zero value. | DESIGN: Software Design | Elekta sent an IMPORTANT NOTICE letter dated March 9, 2011 to affected customers. The notice identified the product, the problem, and the corrective action to be taken by the customer. The notice advise all customers to follow any instructions or recommendations covered in the Notice. For questions call +44(0)1293 654200. |  | 56 | 56 | Nationwide Distribution (USA) including the states of: AZ, CA, GA, IL, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, TN, VA, WI and Puerto Rico | Terminated | August 09, 2012 | 254 | 56 | Software |  |  | Other
Z-0303-2012 | 59479 | pump, blood, cardiopulmonary bypass, roller type | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery Module for Sarns Modular Perfusion System 8000 Product Usage: The Sarns Modular Perfusion System 8000 is indicated for up to 6 hours use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist. | Sarns Modular Perfusion System 8000, Circuit Board for MBattery Module | 2 | November 29, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems is aware on one instance in which an incorrectly labeled printed circuit board for the Sarns Modular Perfusion System 8000 was shipped to a single customer. The affected board mislabeled two connection between the board and the transformer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo conducted a Correction/Removal - Customer Communicated Script on November 7, 2011 to the affected customer. The phone script communicated the correction and reason for removal, potential hazards, affected population, and confirmation of possession of affected module. A replacement date was set for November 7, 2011. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818, hours are Monday to Friday. 8 AM -6 PM EST. | Contact the recalling firm for information | 10 | 10 units | Nationwide Distribution (USA) including the state of San Antonio, TX | Terminated | February 17, 2012 | 80 | 10 | Not_Computer
Z-0306-2012 | 60192 | dressing, wound, drug | FRO |  |  | General & Plastic Surgery | 510(k) | ON-Q Silver Dressing is a Antimicrobial wound dressing containing silver. It is individually wrapped in a foil pouch. Each pouch contains 1 single use 4" x 4.5" pad. There are 5 pouches per box. Ref: SD4X4.5; Part Number: IF-151023 and IF-151026.The product is labeled in parts: "***5 DRESSINGS***4 X 4.5 INCHES***REF SD4X4.5***PART NO. 5001826***ON-Q SilverDressing is a silver containing adhesive transparent polyurethane film for fixation and securement of devices to the skin as a primary contact sound dressing***MANUFACTURED BY ACRYMED***DISTRIBUTED BY I-FLOW***AcryMed, Inc Beaverton, OR 97008***".Intended for use in securing devices such as catheters to the skin. Also intended for use as a primary or secondary cover wound dressing on surgical incision sites. | ONQ Silver Dressing | 2 | November 29, 2011 | 2011 | Acrymed Incorporated | The Silver Dressing, lot H0705J and H0706J may contain dressings that are not properly sealed in the foil pouch, which could affect sterility of the dressing. | PRODUCTION CONTROLS: Process Control | On June 30 2008, the firm e-mailed all affected customers with an Important Product Recall Notification letter.The customers were notified that two lots of SilverDressing may contain dressings that are not properly seated in the foil pouch.The customers were instructed to inspect the inventory, ceased distribution of the affected lots and quaratine the remaining inventory and notify AcryMed.Customers can call AcryMed at 503-624-9830 for any questions about this recall. | Discard product or return to place of purchase for a full refund or exchange | 41061 | 14,125 units from lot # H0706J and 25,525 units from lot # H0705J | Nationwide distribution -- DE, FL, KY, OR, IN, TX, NV, MT, CA and NY. | Terminated | December 20, 2011 | 21 | 41061 | Not_Computer
Z-0298-2012 | 59581 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | The Terumo Advanced Perfusion System 1, Catalog Number 801763, System 1 Base, 110 VProduct Usage: The Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures. | Advanced Perfusion System 1 | 2 | November 29, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | 1 US customer was operating the Advanced Perfusion System 1 with an unapproved, modified cable, which was in product development, to connect the system an an occluder for the Sarns Modular Perfusion System 8000 . Five(5)of twelve(12) cables from Terumo Japan, which were in product development, were given to a US customer. | DESIGN: Device Design | Terumo sent an URGENT MEDICAL DEVICE CORRECTION: Recall Notification dated November 8, 2011 to all affected customers. The letter described the reason for the recall, known/potential hazards and request immediate location and correction of product. The letter instructs customers to stop using the modified cables and return to Terumo for replacement and destruction. Customers were asked to complete the attached response form and fax to the number indicated on the form. All affected customers will be receiving a safety advisory to advise them that Terumo do not recommend the alternative configuration of a System 1 Base with an 8K Occluder. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818 or 1-800-292-6551 (fax). |  | 12 | 12 units | Worldwide Distribution - (USA) Nationwide Distribution | Terminated | August 22, 2012 | 267 | 12 | Not_Computer
Z-0302-2012 | 60289 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1- 4" Roller Pump for Terumo System 1Product Usage: The Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | November 29, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS discovered that our In house service center has incorrectly installed a felt washer into the main bearing of nine roller pumps that had been returned for repair or preventative maintenance unrelated to the felt washer. The felt washer was added as a validated running change to production to all roller pumps manufactured since February 2009. The felt washer should not have been installed | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo sent an URGENT MEDICAL DEVICE REMOVAL notification letter dated November 10, 2011 to all affected customers. The letter described the affected products, problem, potential hazards, and important clinical information. Replacement roller pumps will be configured at the customer site by a Terumo Field representative with the affected roller pumps being returned via normal returned goods process. Customers were asked to complete the attached response form and fax to the number indicated on the form. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818. |  | 3 | 3 units | Worldwide Distribution - USA (nationwide) including the states of: CA, KY, MO, NE, PA and TX and the countries of Canada and Europe. | Terminated | September 06, 2012 | 282 | 9 | Not_Computer
Z-0304-2012 | 60204 | system, measurement, blood-pressure, non-invasive | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Ambulatory Blood Pressure (ABP) Monitor, Model 90217A.The Spacelabs Model 90217A Ambulatory Blood Pressure (ABP) Monitor is a small, lightweight unit designed to take blood pressure and heart rate measurements for a 24 hour, 48 hour, or longer period. | Ambulatory Blood Pressure (ABP) Monitor | 2 | November 29, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | If external pressure is applied to the cuff of the Ambulatory Blood Pressure, Model 90217A, during deflation, the monitor will stop working and display an EC75 error code. | DESIGN: Device Design | On October 21, 2011, a Medial Device Correction customer letter was mailed to all US customers. A return receipt was requested.The letter identified the affected devices and described the problem. If the device displays an EC75 error code, customers can cycle the power off and then on to restore normal functionality. Please advise patients of this mitigation to avoid a delay in obtaining results. The firm is in the process of upgrading their software to avoid the issue and will contact customers to arrange for an update to their monitors.Customers can call Spacelabs Healthcare at 1-800-522-7025, select 2 for Technical support for any question about this recall. |  | 785 | 785 units worldwide (83 in US and 702 outside US) | Worldwide Distribution | Terminated | May 30, 2012 | 183 | 785 | Software | Device Operation | Software update | Software Update
Z-0311-2012 | 60146 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Cobas c 111 AnalyzerThe Roche Cobas c 111 analyzer is an in-vitro diagnostic analyzer capable of performing clinical chemistry, specific protein and electrolyte tests. Analytes are measured photometrically or turbidimetrically; the analyzer also has an optional ISE module for measuring sodium, potassium and chloride. | Roche Cobas c111 Analyzer | 2 | November 30, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Internal investigations have revealed a potential printer malfunction on the Cobas c 111 analyzer, where one or more numbers or characters can be missing on the hard copy result printout, or the printer may skip leaving blanks which may result in an incorrect print layout. The missing number of characters can be in any position, therefore it could be in the middle of a result, (e.g. 48.53 mmol in | DESIGN: Software Design | Roche Diagnostics Operations, Inc sent a Urgent Medical Device Correction dated October 18, 2011, via UPS Ground (receipt required) to all affected customers and one distribution center, with instructions to disable the Auto Print Result feature until new software is available. All printed results will need to be confirmed against the results indicated on the display screen. Questions are to be directed to Roche Diagnostics Technical Support at 1-800 428-2336,. |  | 135 | 135 | Worldwide Distribution - USA including West Virginia and Texas. | Terminated | February 21, 2013 | 449 | 135 | Software | Output/Calculation | Software update | Software Update
Z-0314-2012 | 60287 | fastener, fixation, biodegradable, soft tissue | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek Panalock Loop, Lupine Anchor Catalog #210704The LUPINE BR Anchor System is indicated for use in soft tissue to bone fixation in association with adequate postoperative immobilization. | Panalock Loop, Lupine Anchor | 2 | November 30, 2011 | 2011 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Incorrect suture configuration | TRAINING: Employee Error | The firm, DePuy Mitek Inc., sent an "URGENT VOLUNTARY PRODUCT RECALL"letter dated October 31, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to not use or sell any of the identified problem products included in this recall; complete the enclosed business reply form and fax to: 1-508-828-3750 or email to: mitekcomplaints@its.jnj.com (note: It is important that we receive this business reply form, even if you have no product remaining in your inventory from the affected lots); isolate all inventory and call 1-877-379-4871 for return goods authorization and replacement product; and return all identified affected product to: (USA)- DePuy Mitek, ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or (foreign)-GMED Healthcare EDC Quality Dept Rue de Luxembourg 5, ZI Trazegnies, BE-6180 Courcelles, Belgium. If you have any questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-4871. | Contact the recalling firm for information | 298 | 298 units | Worldwide distribution: USA (nationwide) and countries of: Austria, Australia, Brazil, Canada, Estonia,Switzerland, China, Czechoslovakia, Denmark, Spain, Finland, France, Great Britain Israel, India, Japan, Korea, Mexico, Martinique, Norway, Portugal, Sweden, Singapore, Solvonia, and Slovakia. | Open |  |  | 1768 | Not_Computer
Z-0309-2012 | 60314 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Arrow International StimuCath(R) Continuous Nerve Block Kit, ASK-19608-SFH, Lot # MF0127039Permits placement of catheters next to nerves and nerve plexuses for continuous nerve block anesthesia of analgesia techniques for periods not exceeding 72 hours | Arrow International StimuCath(R) Continuous Nerve Block Kit, ASK19608SFH | 2 | November 30, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging of the ASK-19608-SFH Continuous Nerve Block Kits, lot # MF0127039 were found to have an insufficient seal. Therefore, the packaging may have been compromised, and sterility of the product cannot be guaranteed. | PRODUCTION CONTROLS: Packaging | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 20, 2011, to its customer. The letter included a stock response form to be completed and returned to the firm. The letter also described the product, problem and action to be taken. The customer was instructed to check their stock for the products included in the scope of the recall; cease use and distribution; quarantine all affected product immediately; complete and return the enclosed Recall Acknowledgement & Stock Status Form to their sales representative, and return any affected product freight collect to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, attn: Asheboro Quality Assurance Manager. Arrow International can provide replacement with a similar product or credit their account.If you have questions or concerns, please call this toll free number 800-233-3187. | Contact the recalling firm for information | 17 | 17 kits | Nationwide distribution: CT | Terminated | October 01, 2012 | 306 | 17 | Not_Computer
Z-0319-2012 | 60150 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular CollimatorProduct Usage: This device is intended to hold a patient's head in a fixed position and to localize and center the output of a linear accelerator (UNAC) to allow radiotherapy of brain tumors and other types of cerebral lesions. | Stereotactic Circular Collimator | 2 | November 30, 2011 | 2011 | Elekta, Inc. | Recent newspaper articles outlined improper use of SRS Cone Collimator accessories that injured patients on Brainlab and Varian systems. |  | Elekta sent an IMPORTANT NOTICE letter dated March 25, 2011 to all affected customers. The letter identified the affected product, the problem, key safety procedures and information to users for safe operation of the Stereotactic Circular Collimator. Elekta will issue a Mandatory Action Field Change Order which will include an updated Circular Collimator Instructions for Use - 102022201 and new labels. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. Also, Elekta is offering some free WebEx sessions which will include the procedures and recommended operation of these types of accessories. Customers are instructed to register at www.elekta.com/SafeAccessories. |  | 89 | 89 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN, and WV | Terminated | October 02, 2012 | 307 | 89 | Not_Computer
Z-0310-2012 | 60282 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas c 311 Analyzer, Part Number: 04826876001In-vitro diagnostic analyzer capable of performing clinical chemistry, specific protein and electrolyte tests. Analytes are measured photometrically or turbidimetrically; the analyzer also has an optional ISE module for measuring sodium, potassium and chloride. | cobas c 311 Analyzer | 2 | November 30, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The cobas c 311 analyzer software may incorrectly calculate the volume of the ISE Reference Electrolyte solution remaining on the analyzer. In some cases the incorrect results obtained will not be flagged with any data alarms. | DESIGN: Software Design | October 26, 2011 Urgent Medical Device Corrections were sent to all direct accounts in the USA that have installed the cobas c 311 analyzer and included a faxback form to be completed and faxed to 1-877-766-7452. Consignees were instructed to temporarily replace the ISE Reference Electrolyte bottle based upon the frequency calculated from the maximum number of ISE samples run per day and activate the yellow threshold alarm on the system. Customers with question and concerns are directed to call Roche Diagnostics Technical Support at 1-800-428-2366, available 24-hours, 7 days a week. |  | 940 | 940 worldwide | Worldwide Distribution -- USA including the states of Maryland and Texas. | Terminated | July 09, 2012 | 222 | 940 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0313-2012 | 60323 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Rel 3, a radiographic system, digitalIntended use: Radiographic system, digital | Digital Diagnost Rel 3 | 2 | November 30, 2011 | 2011 | Philips Healthcare Inc. | Missing a small, metal protective cover on the bottom side of the control handle for the ceiling suspension. exposes printed circuit | PRODUCTION CONTROLS: Process Control | The firm, Philips Healthcare, sent a "URGENT - Field Safety Notice" letter dated October 25, 2011, to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed: Do not place your hand in the open slot below the Control Handle and ensure that no patient places his/her hands the opening; review the following information with all members of their staff who need to be aware of the contents of this communication, and retain a copy with the equipment instruction for use. Philips has issued a Field Change order instructing Service Engineers to check affected systems and install the cover where it is missing free of charge.If you need any further information or support concerning this issue call 978-659-4519. | Contact the recalling firm for information | 28 | 28 (7 US; 2 Canada; 19 OUS) | Worldwide distribution: USA (nationwide) including states of: NC, SC, TX, and WV; and countries of: Australia, Austria,Canada, Germany, Italy, Netherlands, Sweden, and Switzerland. | Open |  |  | 28 | Not_Computer
Z-0295-2012 | 59856 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator Probe, Part GM11002560 - 3.2 mm diameter braced PVDF. Varian Medical Systems, Palo Alto, CA, Mfg. by Varian Medical Systems, Haan GmbH. Designed for brachytherapy intracavity treatments in areas such as cancer of the vagina, vaginal stump or rectum. | GammaMed Flexible Applicator Probe | 2 | November 30, 2011 | 2011 | Varian Medical Systems, Inc. | GammaMed Flexible Applicator Probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | PRODUCTION CONTROLS: Process Control | Varian Medical Systems, Inc. sent a Customer Technical Bulletin CTB-GM-603A on August 5, 2008 to all affected customers. The letter included description of device and problem. The bulletin also advised that if they find that the applicator and source guide tube combination is not within the indicated range on the length gauge, to call VBT technical support using the contact information on page 5 of the bulletin to obtain a replacement applicator probe. Customers are to complete and return the attached verification card. If support is needed the letter provides customer support contact information. | Contact the recalling firm for information | 653 | 653 parts | Worldwide Distribution -- USA (nationwide): Arkansas,California, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Minnesota, Montana, Michigan, Missouri, Nebraska, Mississippi, New Jersey, Nevada, New York, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington, West Virginia, and Wisconsin, and the countries of Austria, Azerbaijan, Brazil, British Columbia, Canada, Chile, Germany, Greece, India, India, Ireland, Israel, Mexico, Morocco, Pakistan, Poland, Republic of Macedonia, Russia, Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Ukraine, and Venezuela. | Terminated | December 07, 2011 | 7 | 653 | Not_Computer
Z-0312-2012 | 60128 | system, radiation therapy, radionuclide | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Knife C 1.2, 4 and 4CProduct Usage: Leksell Gamma Knife is a teletherapy device intended for stereotactic irradiation of head structures. | Radionuclie radiation therapy system | 2 | November 30, 2011 | 2011 | Elekta, Inc. | Several of the LMR03 actuators with bronze drive nut have failed unexpectedly, creating a potential safety hazard for the operator and patient. |  | Elekta sent an IMPORTANT NOTICE: Urgent Field Safety Notice letter dated May 12, 2011 to affected customers. The notice identified the affected product, problem, and actions to be taken by the user. The letter describes special care to be used in operating the helmet changer actuator LMR03. The letter states that Elekta will replace all actuators LMRO3. For more details contact your local Elekta representative.. |  | 44 | 44 | Nationwide Distribution (USA) - including the states of: CA, CO, DAL, FL, GA, IL, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OK, OR, PA, TN, TX, VA, WA, WI and Puerto Rico | Open |  |  | 44 | Not_Computer
Z-0307-2012 | 60125 | system, radiation therapy, radionuclide | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma PlanProduct Usage: Gamma Plan is a computer based system designed for Gamma Knife treatment planning. | Elekta Leksell Gamma | 2 | November 30, 2011 | 2011 | Elekta, Inc. | There is a risk that the user may interpret the instructions in the manual too literally, and adjust the TMR 10 dose description to obtain an ETCPD that matches previous clinical prescriptions, which in turn could lead to a prescription dose that deviates a few percent from the intended dose. | MISBRANDING: Labeling False and Misleading | Elekta sent an IMPORTANT NOTICE letter dated June 23, 2011 to affected customers. The notice identified the product, problem, and corrective actions to be taken by the user. The letter advise all customers to follow any instructions or recommendations covered in the Notice. The letter states that Elekta is working on a new software version and customers will receive the future version when released. For questions contact your local Elekta representative. |  | 2 | 2 | Nationwide Distribution (USA) - including the states of: NM and WA | Terminated | January 31, 2012 | 62 | 2 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0326-2012 | 59779 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0 mm diameter, length 113 mm. Remote controlled radionuclide applicator system. Varian Medical Systems, Inc. Mfg by Varian Medical Systems Haan GmbH, Haan, Germany.The interstitial plastic needles with 2 mm diameter are designed for interstitial treatments in areas such as head and neck, gynecology, breast or prostate. In order to avoid bending and to facilitate the introduction of the needles, each needle is supplied with a mandrin. | GammaMed Plastic Needle with Mandrin | 2 | December 01, 2011 | 2011 | Varian Medical Systems, Inc. | Plastic needle tip could possibly crack or separate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian Medical Systems sent a "CUSTOMER TECHNICAL BULLETIN" letter dated September 30, 2008 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed users to inspect needle tips before and after every clinical application to prevent needle tip separation or cracking. | Contact the recalling firm for information | 1716 | 1716 parts | Worldwide Distribution - USA (nationwide) including the states of Arizonia, California, Florida, Illinois, Indiana, Michigan, North Carolina, and Vermont, and the countries of Argentina, Austria, Belgium, Canada, Estonia, France, Germany, India, Norway, Pakistan, Russia, Slovenia, So. Africia, Spain, Switzerland, United Kingdom, and Ukraine. | Terminated | December 07, 2011 | 6 | 1716 | Not_Computer
Z-0339-2012 | 59767 | incubator, neonatal | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Rabee Incu i (Atom Infant Incubator Model 102) Atom Medical Corporation 3-18-15 Hongo, Bunkyo-Ku, Tokyo, Japan 113-0033Usage: An incubator for infants. | Rabee Incu i | 2 | December 01, 2011 | 2011 | Atom Medical Corporation | The caster with lock on the incubator may become loose or break. | DESIGN: Device Design | The firm, Atom Medical, sent an "Urgent: Field Safety Notice" letter dated August 18, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all devices in their inventory and quarantine them until the rework is complete; identify any customers to whom they have further distributed the product; contact their customers at once and arrange a time in which the field service technician can perform the rework (the parts needed to complete the rework will be sent from Atom Medical's export department ); complete rework of all affected devices and document rework on the Distributor FSCA/Recall Response Form and return via Fax to: +81 3 3812-3199 or email to kira-utc@atomed.co.jp; t-sugino@atomed.co.jp or kenichi.shimada@atomed.co.jp (all forms should be completed and returned by January 31, 2012.) If you have any questions, contact the Manager Quality Control or Manager, Regulatory Affairs at email addresses above, phone: +81-3-3815-3632 or fax +81-3-3812-3199. |  | 32 | 32 units | Worldwide distribution: USA (nationwide) including state of: PA and countries of: Algeria, Austria, Australia, China, Egypt, Germany, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Russia, Saudi Arabia, Spain, Switzerland, Thailand, The Netherlands, and UAE. | Terminated | September 09, 2013 | 648 | 270 | Not_Computer
Z-0331-2012 | 60229 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Model CAT-850B Catheterization Table, a component of the Toshiba INFX-8000 series X-Ray Interventional System.These systems are single or dual plane systems that employ x-rays to obtain fluoroscopy or radiography images of the body. The system is comprised of a computer system, support arms (that contain the tube and solid state detector), x-ray generator and a patient table. This device is a digital x-ray system configured for use as a diagnostic x-ray angiography system. This x-ray angiography system is used for diagnostic and interventional procedures for cardiac blood vessels, cerebral blood vessels, abdominal blood vessels and lower extremities. | Model CAT850B Catheterization Table | 2 | December 01, 2011 | 2011 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that during patient transfer from a gurney to the catherization table, the tabletop may rotate unexpectedly when sufficient force is applied to the tabletop. The result of the tabletop rotating unexpectedly is the patient being tranferred may fall from the tabletop. | DESIGN: Device Design | Toshiba America Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers were instructed to follow the user instructions by positioning the tabletop at the foot-end position, and ensuringsufficient assistance is available, when transferring a patient from the gurney to the CAT-850B catheterization table. Customers were also instructed to follow the instructions for transferring a patient to the CAT-850B that are included in the operation manual. A Customer Reply Form was included with the letter for customers to complete and return via fax to (877) 349-3054. The customer reply form can also be sent via e-mail to raffairs@tams.com. Questions regarding the recall letter can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. | Contact the recalling firm for information | 300 | 300 units | Nationwide Distribution. | Terminated | March 18, 2013 | 473 | 300 | Not_Computer
Z-0329-2012 | 60231 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian brand GammaMed 12i Accessory: Clamping adapters for bronchial catheters, 1.8mm diameter, GammaMed 12i; Ref: BT-00264, Model 1-H63; Product is distributed by Varian Medical Systems Inc., Palo Alto, CA and Manufactured by Varian Medical Systems Haan GmbH,Bergische Strasse 16, D-42781 Haan, Germany.The Varian Medical System GammaMed Afterloader system is a remotely controlled intelligent robotic device intended to be used to deliver high dose-rate radiation therapy, in close proximity to and from within diseased tissues. Palliative care and treatments with curative intent to both benign and non-benign disease using interstitial, intracavitary, intraluminal, intra and inter-operative brachytherapy or combinations of any these treatment types are possible using this device. | GammaMed 12i Accessory | 2 | December 01, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Product Models 1 & 2 have a design flaw problem where the clamping adapter canobstruct movement of the source cable in the event the bronchial catheter is dislodgedfrom the connector. | DESIGN: Component Design/Selection | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION LETTER/FIELD SAFETY NOTICE" letter dated October 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Contact the Varian Help Desk at 1-888-827-4265 for questions regarding this notice. | Contact the recalling firm for information |  |  | Worldwide Distribution--USA (nationwide). | Terminated | October 22, 2012 | 326 |  | Not_Computer
Z-0325-2012 | 59885 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VariSource iX, Software Version 1.1, Part Number AL20020000 VariSource iX Afterloader.The VariSource iX Series High Dose Rate Afterloader system is a computer controlled remote electra/mechanical system used for medical purposes, for placing a NiTiNol wire incorporating an irradiated iridium seed internally or close by, a malignant tumor or tumor bed in a practice known as brachytherapy. The device has up to 20 channels for treatment. | Brachytherapy System, Remote Afterloader | 2 | December 01, 2011 | 2011 | Varian Medical Systems, Inc. | VariSource iX console software may cause incorrect dwell times to be delivered to patient and may result in a mistreatment. | DESIGN: Software Design | Varian Medical Systems sent a "MEDICAL DEVICE RECALL" letter dated April 1, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to conduct a simple test to establish if their machine is affected by the problem. Device users were requested to complete an attached receipt verification form and return it to Varian BrachyTherapy via fax or e-mail. Call the Contact Service Dispatch number at 1-800-864-1672 for questions regarding this detail. |  | 72 | 72 worldwide/50 in U.S. | Worldwide Distribution-USA (nationwide) | Terminated | May 30, 2013 | 546 | 72 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0336-2012 | 60286 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 10mm x 2.7mm RibLoc Screw, a component of the RibLoc Rib Fracture Plating System, Model RBP1202.The product is labeled in parts: "***Ribloc Rib Fracture Plating System***MATERIAL: Titanium***Rx ONLY***NON STERILE***SCREW***ACUTE Innovations LLC 21421 NW Jacobson Road, Suite 700 Hillsboro, OR 97124 USA www.acuteinnovations.com*** MediMark Europe Sarl*** 11 rue Emile ZOLA BP 2332****38003 GRENOBLE CEDEX 2 FRANCE"The Acumed Rib Congruent Bone Plate System is indicated for use in providing fixation during fractures, fusions, and osteotomies for the rib. The Acumed Rib Congruent Bone Plate System includes plates and screws designed specifically for the rib. | RibLoc System, Standard Screw | 2 | December 01, 2011 | 2011 | Acute Innovations Llc | RibLoc screw from lot L1008007 was out of specification. It would not engage with the hex tip properly. | PRODUCTION CONTROLS: Process Control | The firm, ACUTE Innovation, sent an "URGENT NOTICE: DEVICE RECALL" letter dated February 21, 2011, to all consignees/customers via FED-EX on February 28, 2011. The letter describes the product, problem and actions to be taken. The customers were instructed to identify product having the lot code(s) and quarantine them; return all product from this lot to Acute Innovations immediately using shipping labels provided or return to: Acute Innovations, LLC, 21421 NW Jacobson Rd, Suite 700, Hillsboro, OR, 97124, and if they further distributed this product, notify those customers of this recall and request return of product(s).Customers can call ACUTE Innovation at 1-866-623-4137 for any questions about this recall. | Discard product or return to place of purchase for a full refund or exchange | 299 | 299 units (267 in US and 32 to South Africa) | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, GA, IL, MN, NC, NY, UT, and WA; and countries of: South Africa. | Terminated | December 06, 2011 | 5 | 299 | Not_Computer
Z-0343-2012 | 60302 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips GEMINI TF 16 Diagnostic CT X-ray/PET Imaging System, 510(k) #K052640. System units equipped with software version: 3.5.2 are subject to recall/field correction. The GEMINI TF I6, and TF 64 are diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other organ imaging). Attenuation correction is accomplished by CT AC. The device also provides for list mode, dynamic, and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI TF 16 PET/CT Xray Diagnostic Imaging System | 2 | December 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | In response to complaints received by Philips Medical concerning certain identified software errors reported by their customers, the firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | DESIGN: Software Design | Philips Healthcare sent a "CPE CUSTOMER LETTER" dated October 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Additionally, device operation information is provided to customers. A Philips Field Service Engineer will be visiting each customer site to install the necessary software upgrades. Contact the Customer Care Solutions Center at 1-800-722-9377 for questions concerning this notice. |  | 9 | 9 units | Nationwide Distribution including the states of CO, LA, MD, MI, MN, PA, TX, VA, VT, and WI. | Terminated | February 22, 2013 | 449 | 19 | Software
Z-0345-2012 | 59934 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | OdysseyA radiation treatment planning system | Odyssey | 2 | December 01, 2011 | 2011 | PerMedics, Inc. | Anomaly 1 - Invalid Export of Beam Table and or Collimator Angle for DICOM RT Ion, DICOM Image and RTP Link export files:During PerMedics' internal code reviews for lEC 61217 compliance verification, an issue was identified by PerMedics engineers in three export files: D1COM RT Ion Plan, D1COM RT Image and RTP Link. In some cases, this issue can result in table and/or collimator angle values t | DESIGN: Device Design | PerMedics sent an Odyssey TPS URGENT NOTICE OF CORRECTION letter dated November 4, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Anomaly 1 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations were provided if they planned to use Odyssey RT Ion Plan, DICOM RT Image, and/or RTP Link export:Workaround for the invalid exported table and/or collimator angle errors: If, prior to the software correction, any of these affected file types were exported from Odyssey, manually verify that the table and collimator angles are correct. Additionally, apply appropriate changes to any angles deemed incorrect that were obtained from Odyssey's DICOM RT Ion Plan, DICOM RT Image, and/or RTP Link files.Only those machines with table and/or collimator angles that increase in the clockwise direction will have the incorrect angles.Anomaly 2 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations are provided if the user- Changes the patients scheduled start date prior to the software correction:Workaround for Incorrect Monitor Unit Values when Changing Treatment Start Date: If, prior to the software correction, the user changes the patient's treatment start date on any fully-computed IMRT plan, IMAT plan or plan containing a "Region DVH" prescription, the user should force a full recommendation oft hose patient plans prior to fixing and exporting the plan.The following procedure can be used to force a full re-computation of the plan:- Open the patient's plan worksheet- Change the algorithm type of the plan- Change the algorithm type back to its original value- Close t. | No consumer action necessary | 15 | 15 | Worldwide Distribution - USA including NY, GA, CA and the countries of Hungary, Brazil, China, India, Italy, and Mexico | Terminated | March 02, 2012 | 92 | 15 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0323-2012 | 59814 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZMR Crossover Instrumentation, utilized for the following ZMR devices: Porous Stem,Nitrited Porous Revision, Slotted Porous Bowed Stems with and without Compression Nut, Splined Straight Stem, Taper Stem and Bodies, Calcar Bodies, and Spout Bodies.Product Usage: The ZMR Hip Prosthesis is indicated for cementless total hip arthroplasty in patients whose bone stock is of poor quality or inadequate for other reconstruction techniques as indicated by deficiencies of the femoral head, neck or portions of the proximal femur. | ZMR Porous Revision Hip Prosthesis and ZMR Revision Taper Hip Prosthesis | 2 | December 01, 2011 | 2011 | Zimmer Inc. | This recall was initiated because it may not be clear to the surgeon whether the device can be used based on the instructions for use. | PRODUCTION CONTROLS: Error in Labeling | The firm Zimmer, sent an "URGENT DEVICE CORRECTION" letter dated August 26, 2011 to Marketing Personnel, Distributors, Sales Representatives, Distribution Operation Managers, and Surgeons. The letters identified the reason for the recall, the product, and required actions to be taken. The letter also included Medwatch reporting information and a Certificate of Acknowledgement form to fax back to the recalling firm to confirm receipt of the recall notice. For questions contact Zimmer at 1-877-946-2761. | Contact the recalling firm for information | 119284 | 119,284 total devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dubai/UAE, Eschbach, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Japan, Jordan, Israel, Korea, Lebanon, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, S. Africa, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Taiwan, Thailand, Turkey, UK, Ukraine, and Venezuela, | Terminated | August 29, 2013 | 637 | 119284 | Not_Computer
Z-0320-2012 | 60250 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R.4.5 and R4.6ProducT Usage: To be used as part of radiation therapy treatment process. | XVI R4.5 Xray Volume Imaging System | 2 | December 01, 2011 | 2011 | Elekta, Inc. | The software did not operate as expected. | DESIGN: Software Design | Elekta sent an IMPORTANT NOTICE letter dated October 14, 2011 to all affected consignees. The letter identified the affected product, the problem, and the actions to be taken. This notice contained precautionary measures that are required for the operation of the equipment. The notice advise all customers to follow any instructions or recommendations covered in the Notice. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. For questions contact your local Elekta representative. |  | 77 | 77 | Nationwide Distribution (USA) - including the states of: AZ, CA, CT, GA, ID, IL, LA, MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI and Puerto Rico | Open |  |  | 77 | Software |  | Instructions | Safety Notice/Insructions
Z-0333-2012 | 60244 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | LATITUDEÂŽ Patient Management System, Models 6441*, 6442*, 6465*, 6488 (USA).There are two software applications: the programmer software and the LATITUDE remote monitoring system software. The Model 2868 Software Application is loaded onto the Model 3120 programmers. These programmers are used during in-office follow-ups. The software version affected domestically is version 1.05 for COGNIS and TELIGEN and internationally is version 2.04 for COGNIS, TELIGEN, and PUNCTUA, ENERGEN, and INCEPTA. The PUNCTUA, ENERGEN, and INCEPTA are not currently approved in the USThe LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. The LATITUDE Patient Management System software is web-based and therefore at any point in time, there is only one (1) released version for each model number. The current version is 7.1.Both the programmer software and the LATITUDE remote monitoring system software similarly display the shock lead impedance measurements. | LATITUDE Patient Management System | 2 | December 01, 2011 | 2011 | Boston Scientific CRM Corp | Current Boston Scientific defibrillator systems measure shock lead impedances up to 200 ohms, and export results to the programmer and LATITUDE monitoring system via the Daily Measurements feature. A measured shock lead impedance that is greater than 125 ohms or less than 20 ohms will generate an alert message on programmer screens and a Red Alert within LATITUDE (if activated). However, the mea | DESIGN: Software Design | Boston Scientific sent an "IMPORTANT DIAGNOSTIC INFORMATION" letter dated October 2011 to all affected customers. The letter described the product, problem and actions to be taken by the customers. | Contact the recalling firm for information | 25300 | 25,300 (14,700 US; 10,600 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, CANADA, AUSTRALIA, NEW ZEALAND, CYPRUS, CZECH REPUBLIC, GREECE, GUADELOUPE, HUNGARY, BANGLADESH, CHINA, COCOS ISLAND, HONG KONG, INDIA, INDONESIA, KOREA, MALAYSIA, PHILIPPINES, SINGAPORE, SOUTH KOREA, SRI LANKA, THAILAND, RUSSIA, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, NEW CALEDONIA, POLAND, PORTUGAL, QATAR, REUNION,IRAN, ISRAEL, JORDAN, KUWAIT, LATVIA, LEBANON, LUXEMBOURG, MONACO, MOROCCO, JAPAN, ARGENTINA, ARUBA, BAHAMAS, BARBADOS, BELIZE, BERMUDA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, FALKLAND ISLANDS, HAITI, HONDURAS, JAMAICA, MEXICO, PANAMA, PERU, TRINIDAD & Tobago, URUGUAY, VENEZUELA, AND NETHERLANDS ANTILLES. | Terminated | December 04, 2012 | 369 | 25300 | Not_Computer
Z-0332-2012 | 60238 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Head Ring Posts with part number 970.280 - re-usable components of the Frame Array Module (of the Optical Guidance Platform and Floorstand devices.Varian Medical Systems, Palo Alto, CA 94304.The Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment for stereotactic radiosurgery or radiotherapy treatments on cranial extracranial lesions. | Head Ring Posts; of the Optical Guidance Platform and Floorstand devices | 2 | December 01, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the Head Ring posts used by both the Optical Guidance Platform FrameArray module and the Floorstand where the head ring posts may be damaged due to excessive mechanical stress resulting in possible failure during usage. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated October 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers are advised to verify that they are using the current design of the head ring posts (part # 970.284) and to contact Varian if replacements are needed. Customers may contact the Varian Help Desk for questions regarding this notice at 1-888-827-4265. | No consumer action necessary | 205 | 205 kits ( Each kit contains 4 Head Post Rings) 820 HPR's | Worldwide Distribution. | Terminated | July 25, 2012 | 237 | 205 | Not_Computer
Z-0335-2012 | 60228 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, 510(k) #K033326, Model Number: 728231, Serial Numbers: 9745, 90169, and 95414. The recalled device systems are equipped with Version 3.5 Beta 1 software. | Brilliance CT XRay System | 2 | December 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were found for the Brilliance Big Bore and Brilliance 64 Computed Tomography X-ray Systems equipped with software version; 3.5 Beta1. The non-conformances stem from software errors identified in the operating software. | DESIGN: Software Design | Philips Healthcare released a 'Field Change Order' (FCO #72800531) on September 12, 2011. Philips sent an - URGENT Medical Device Correction letter dated October 17, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to this situation. The letter informed customers on how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update is installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. Customers were instructed to contact their Philips representative or local Philips Healthcare Office for further information or support for this issue. For North America and Canada customers were instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 3 | three (3) units were distributed. | Worldwide Distribution - USA including Texas and the countries of Italy, India and China. | Terminated | February 22, 2013 | 449 | 5 | Software |  | Softwarae update | Software Update
Z-0341-2012 | 60153 | monitor, carbon-dioxide, cutaneous | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TC pCO2 Electrodes, Model/Part #945-655-E5260, 510(k) #K043003. Product Usage: The device is intended for used with the firm's TCM series monitor. The E260mpCO2 electrode is a non-invasive device, used with the TCM series monitor to measure transcutaneous pCO2 (through the skin). | transcutaneous pCO2 electrodes | 2 | December 01, 2011 | 2011 | Radiometer America Inc | Transcutaneous pCO2-electrodes were wrongly coded as: pO2 electrode on the hardware plug for the firm's TCM series monitors. This causes the monitor to interpret the electrode wrongly, causing it to fail all calibrations and never get ready for use. | TRAINING: Employee Error | Radiometer contacted the affected U.S. customer and informed them of the recall via telephone on September 27, 2011. Radiometer requested the return of the affected product received from the firm. For questions call (440) 871-0463. |  | 4 | 4 units were distributed. | Worldwide Distribution - USA (nationwide) in the state of San Diego, CA. | Terminated | February 19, 2012 | 80 | 4 | Not_Computer
Z-0342-2012 | 60497 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI X-ray Volume Imaging SystemTo be used as part of radiation therapy treatment process. | Synergy XVI kV Generator | 2 | December 01, 2011 | 2011 | Elekta, Inc. | A small number of sites have experienced situations with kV generator transformers overheating due to a diode failure on the interface board. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Elekta, issued a "FIELD CHANGE ORDER MANDATORY ACTION" notice dated October 20, 2011 to its customers to add fuse protection for the kV transformers. The notice describes the product, problem, provides instructions and procedures and actions to be taken. The customers were instructed to follow any instructions or recommendations covered in the notice and to file the document in the Important Notice section of the appropriate User Manual. Elekta will contact the customer if the FCO is applicable to their kV generator.If you have any questions, call 770-300-9725. |  | 123 | 123 | Nationwide distribution: USA including states of: AL, AK, AZ, CA, CT, FL, GA, IL, IA, KY, MD, MS, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TX, VT, VA, WA, WV, WI, Puerto Rico, & DC. | Terminated | August 09, 2012 | 252 | 123 | Hardware | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-0363-2012 | 60458 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12.SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable, trained, and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical, gynecological, medical cytological, and autopsy case specimens within the Pathology labs, for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SoftPath GUI | 2 | December 02, 2011 | 2011 | SCC Soft Computer | On 07/20/2011 SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. A client reported that the HPV results on a patient report were incorrect. The HPV results were from the previous order on the patient and not the current order. All affected clients were notified about the issue on 07/20/2011. | DESIGN: Software Design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" dated July 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. SCC personnel will review the customized Lab Results AP macro and/or the Micro Results AP macro on customer systems and will either discuss with them the irregularities requiring change or reconfigure macros as necessary. |  | 97 | 97 | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 21, 2013 | 536 | 97 | Software | Output/Calculation | Upgrade | Repair
Z-0355-2012 | 60451 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable, trained, and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical, gynecological, medical cytological, and autopsy case specimens within the Pathology labs, for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. | 2 | December 02, 2011 | 2011 | SCC Soft Computer | On 06/13/2011, SCC Soft Computer initiated a correction on the following SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. A client reported when they verified the contents of the fax reports, they found one of the specimen diagnoses to be missing. Parts A through D were present as well as Part F, however, Part E diagnosis was missing for one of the patient reports. All affected clients wer | DESIGN: Device Design | SCC Soft Computer initiated a Field Correction on June 13, 2011. All customers with affected software have been notified of this issue.SCC recommends until the corrected patch is loaded to temporarily use one of the following alternative working solutions: 1. Use MS Word Editor functionality within SoftPath application for creating patient reports. In order to ensure Text Control editor is not inadvertently activated, the templates for Text Control editor will be deleted and reinstalled at the clients request by their SCC representative after receipt of the correction. 2. In Report Template set-up, the bottom margin should be made slightly larger than the size of the footer plus its distance from the bottom page border. For further explanation, please consult your SCC representative for further details, or call (727) 789-0100. | No consumer action necessary | 14 | 14 | Class II Recall - Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 181 | 14 | Software | Output/Calculation | Software update | Software Update
Z-0359-2012 | 60409 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15 Consult instruction for use. RaySearch Laboratories AB Sveavagen 25, SE-111 34 Stockholm Sweden. support@raysearchlabs.comRayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment, and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system. | RayStation | 2 | December 02, 2011 | 2011 | RAYSEARCH LABORATORIES AB | The firm has found an issue with DICOM RT Image representation of DRR images from RayStation 2.0 which needs to be clarified in the user manual. The symptoms involve the Pixel Spacing and Image Position attributes in the RT images being incorrect, thus enabling the user to make measurements that are erroneous and possibly misalign the patient. | DESIGN: Device Design | The firm, RaySearch Laboratories, sent an email with an "Urgent Field Safety Notice Medical Device Correction" notice dated October 19, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to pass on the information in the field safety notice to everyone involved and create routines to compensate for the incorrect DRR images; if the DICOM DRR images are not used to make any form of absolute measurements, no action needs to be taken. Raysearch laboratories will provide a script to update these DICOM files to have the correct values upon request.For further information please contact RaySearch Laboratories AB Quality and Regulatory Affairs Manager at +46 722 366 110 or david.hedfors@raysearchlabs.com; RaySearch Americas Director of Customer Support at +1 877 778 3489 or freddie.cardell@raysearchlabs.com. | Contact the recalling firm for information | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: MA, NJ, and NY; and country of: Netherlands | Terminated | May 11, 2012 | 161 | 19 | Software | Display/Image | Software update | Software Update
Z-0346-2012 | 60335 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part - Pouch label: "***ORTHOFIX***Manufactured By: Nanotherapeutics, Inc. 13859 Progress Blvd, Suite 300 Alachua, FL 32615 (386) 462-9663***Description: Origen***DBM Paste***Size***REF:***ID***Expiration***STERILE R***Donated Human Tissue Allograft***SINGLE PATIENT USE ONLY***Rx Only***" Product is labeled in part - Carton label: "***Origen DBM with Bioactive Glass***REF***ID***Expiration***Manufactured for Blackstone Medical, Inc by: Nanotherapeutics, Inc. 13859 Progress Blvd. suite 300 Alachua, FL 32615***Distributed By: Orthofix Inc 1720 Bray Central Drive McKinney, TX 75069***Rx Only***ONE EACH***"Origen DBM with Bioactive Glass Catalog #:22-2002 (2cc size) 22-2005 (5cc size)22-2010 (10cc size)"Bone void filler comprised of a gelatin carrier, bioactive glass (45s5), and human demineralized bone matrix (DBM) intended for transplantation."Indicated to be gently placed into bony voids or gaps of the skeletal system that are not intrinsic to the stability of the bony structure (e.g. the extremities and pelvis). | Origen DBM with Bioactive Glass | 2 | December 02, 2011 | 2011 | Nanotherapeutics, Inc | On 03/10/2011 Nanotherapeutics Inc, Alachua, FL initiated a recall of their Origen DBM with Bioactive Glass Catalog # 22-2002 (2cc), 22-2005 (5cc) and 22-2010 (10cc). The same product is also distributed as NanoFUSE DBM Catalog # NAN109-02 (2cc) NAN109-05 (5cc) and NAN109-10(10cc). The products were manufactured without an approved 510k. Direct accounts were notified of the recall on 04/08/2011 |  | Nanotherapeutics sent an Urgent: Medical Device Recall letter dated April 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove any unimplanted/unsold product shipped to them. Customers were asked to notify their customers and instruct them to make notification until the end user (hospital/physician) is reached. Recall Instructions:Immediately examine inventory and quarantine product subject to recall.Return all quarantined product subject to recall to:Nanotherapeutics, Inc.Attn: Receiving Department13859 Progress Blvd, Suite 300Alachua, Florida 32615For questions call 386-462-9663. |  | 6276 | 6276 units | Nationwide Distribuiton including TX, FL, AZ, CA, NY, and MA. | Terminated | April 03, 2012 | 123 | 6407 | Not_Computer
Z-0348-2012 | 60237 | instrument, surgical, orthopedic, ac-powered motor and accessory/attachment | HWE | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Oscillating Saw Attachment/Large w/Key for Small Battery Drive.This surgical instrument motor and accessories are AC-powered, battery-powered or air-powered devices intended for use during surgical procedures to provide power to operate various accessories or attachments to cut hard tissue or bone and soft tissue. Accessories or attachments may include a bur, chisel (osteotome), dermabrasion brush, dermatome, drill bit, hammerhead, pin driver and saw blade. | Synthes Oscillating Saw Attachment/Large w/Key for Small Battery Drive | 2 | December 02, 2011 | 2011 | Synthes USA (HQ), Inc. | The potential exists for the Large Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece, there is a potential for injury to the patient and/or user. |  | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers return their affected device(s) for servicing. Instructions are provided for customers to follow. Customers having questions regarding this notice should contact the firm at 610-719-5450. | Contact the recalling firm for information | 5 | 5 | Nationwide Distribution including the states of CO, LA, NY, and OK. | Open |  |  | 5 | Not_Computer
Z-0349-2012 | 60245 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Tibial Broach P/N 801-01-013Foundation Tibial Broach to be used in conjunction with the Tibial Sizing Template and Tibial Broach Guide to prepare the resected tibial plateau for acceptance of a keeled tibial Baseplate stem. P/N 801-01-013 | Foundation Tibial Broach | 2 | December 02, 2011 | 2011 | Encore Medical, Lp | The strike plate may dislodge from the broach stem handle during impaction of the device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DJO Global sent an URGENT FIELD SAFETY NOTICE letter dated October 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information to all those who need to be aware within their organization or to any organization where the affected devices have been transferred. Customers should contact Customer Service to place a replacement order and receive an RMA number and return all affected devices using the RMA number. For any questions call (512) 834-6255. |  | 30 | 30 units | Worldwide Distribution - USA including AR, AZ, CA, LA, MD, NY, TX, UT and Puerto Rico and the countries of United Kingdom and Italy | Terminated | September 05, 2012 | 278 | 30 | Not_Computer
Z-0351-2012 | 60347 | instrument, biopsy, suction | FCK | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The EnCorÂŽ biopsy probe, designed for use with EnCorÂŽ and EnCorÂŽ Enspire Vacuum Assisted Biopsy (VAB) driver systems.The EnCorÂŽ Biopsy Device is indicated for use in acquiring tissue for diagnosing breast abnormalities and is supplied sterile and intended for single use. | The EnCor biopsy probe | 2 | December 02, 2011 | 2011 | Bard Peripheral Vascular Inc | This recall was voluntarily initiated by the firm due to the product 's damaged packaging trays; cracked EnCor (Trademark) packaging trays, plastic PETG tray the probe is packaged in was cracked. | PRODUCTION CONTROLS: Packaging | The firm sent the customer notification letter on 11/07/11, via FedEx with proof of delivery notice. The letter notified customers of the reason for recall that units from the affected lot may have packaging damage resulting from transportation. The letter instructed customers to check all inventory locations , remove any identified product, to not use the affected product, and return the affected product to Bard Peripheral Vascular, Inc., 1415 W. 3rd Street, Tempe, Arizona 85281. The letter also instructed customers to fill out the recall and effectiveness check form and fax it to 1-800-894-6772. Customers were instructed to call 1-800-321-4254, Option #2 Ext 2727 for a return authorization number to facilitate the expedient return of the product.For questions regarding this recall call 480-894-9515. | No consumer action necessary | 64570 | 64570 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, LAPAC, and Australia | Terminated | March 29, 2012 | 118 | 64570 | Not_Computer
Z-0362-2012 | 60416 | medical computers and software | LNX |  |  | General Hospital | 510(k) | SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. SoftReports is a web-based report launching tool to be used in a hospital or laboratory environment by knowledgeable, trained, and experienced personnel. SoftReports enables the user to schedule and generate reports for the purpose of viewing and printing data stored in SCC product databases. | SoftReports | 2 | December 02, 2011 | 2011 | SCC Soft Computer | On 10/25/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. When SoftReports Designer opened a layout, the system did not load the layout with the previously saved value in the NewPage property. This caused multiple patient records to be printed on the same page. All affected clients were notified on the issue on | DESIGN: Software Design | SCC Soft Computer sent a "Correction Notification" letter dated October 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The hot fixes that were previously loaded which introduced this issue were corrected and delivered to all affected customers as a mandatory correction in hot fixes 1.7199.8 and 1.7200.8 and higher. | No consumer action necessary | 5 | 5 | Nationwide Distribution. | Terminated | March 16, 2012 | 105 | 5 | Software | Output/Calculation | Software update | Software Update
Z-0364-2012 | 60411 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and TrueBeam STx, Model Number: H19, Ref/FSCA identifier: CP-06381 are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation treatment is indicated. | TrueBeam and TrueBeam STx, | 2 | December 02, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the TrueBeam and TrueBeam STx systems where, under certain tuning conditions, the electron beam emerging from the bend magnet may have an elongated spot shape. | DESIGN: Device Design | On 11/04/2011, an Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps. Varian is developing a corrective action for this issue. A customer service representative will contact you when the correction is available to schedule installation on your system. Questions regarding the letter should be directed to the Varian Oncology Help Desk Contact Information at USA and Canada: 1-888-827- 4265 or Europe +41-41-749-8844. | No consumer action necessary |  |  | Worldwide distribution including the US and Internationally. | Terminated | July 25, 2012 | 236 |  | Other |  | Software update | Software Update
Z-0360-2012 | 59886 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #728231, 510(k) #K 033326, Serial Numbers: 95504, 91003, 300030, 95130, and 95157.The Brilliance 64 is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at difference angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance CT XRay System | 2 | December 02, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64 and Ingenuity CT software version 4.0 Beta 2. The non-conformances are due to software errors identified in these device systems. | DESIGN: Software Design | Philips Healthcare released a 'Field Change Order' (FCO #72800542) on September 15, 2011. Philips sent an URGENT - Medical Device Correction letter dated October 10, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to the situation. The letter was mailed by Philips to their customers on October 27, 2011. The letter provided information telling the customer how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update was installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. The letter closes by providing information which all customers in Canada and North America can use to contact the firm through a toll-free telephone number 1-800-722-9377, option 5; enter side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 5 | 5 units | Worldwide Distribution - USA including Ohio and the countries of Australia, China, Belgium, Israel, and Sweden | Terminated | March 15, 2013 | 469 | 7 | Software |  | Software update | Software Update
Z-0353-2012 | 60338 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Servo-S ventilator systemDevice Part number 66 40 440Intended for treatment of and monitoring of patients in range of neonates, infants and adults with respiratory failure of respiratory insufficiency; used only by healthcare providers in hospitals or healthcare facilities and for in hospital transport. | ServoS ventilator system | 2 | December 02, 2011 | 2011 | Maquet Inc. | Servo-I and Servo--S ventilator systems are equipped with a specific revision of a printed circuit board (PC1772 revision 22) that includes a component that has generated a number of nuisance alarms due to technical error message 24 (TE 24). |  | The firm, Maquet Getinge Group, sent an " PRODUCT RECALL URGENT - MEDICAL DEVICE CORRECTIVE ACTION" letter dated September 19, 2011, to its customers. The letter describes the product, problem and actions to be taken. Maquet service team will contact the customers within two weeks of receiving this letter to schedule an on-site service of the affected equipment. If you have any questions, call technical support at our toll free number at 1-888-627-8383. | Contact the recalling firm for information | 26 | 26 units | Nationwide distribution. | Open |  |  | 202 | Not_Computer
Z-0356-2012 | 60300 | light source, fiberoptic, routine | FCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Integra Surgical Luxtec MLX Light Source Catalog Numbers: 00MLXProduct Usage: The Integra Luxtec MLX Light Source is designed to supply high intensity white light through a fiber optic cable for illumination of a surgical field during a surgical and or medical procedure. | Integra LifeSciences | 2 | December 02, 2011 | 2011 | Integra LifeSciences Corp. | A loose screw may result in shifting of components and improper function of the turret or actuator assembly. | PRODUCTION CONTROLS: Process Control | Integra sent an URGENT: Field Corrective Action letter dated October 14, 2011 to all affected customers via FedEx. The letter identified the affected product, the problem, and the actions to be taken with instructions to immediately inspect their inventory for potentially loose screws within the affected unit(s). If the unit(s) were observed with loose screws, the unit(s) were to be returned to Integra Surgical for repair and servicing. Customers were instructed to complete the Acknowledgement and Return Form per the instructions on the form. For questions call Integra Surgical Customer Service at 1-800-431-1123. |  | 66 | 66 units | Worldwide Distribution - USA (nationwide) including the states of: Arizona, California, Colorado, Florida, Indiana, Mississippi, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Montana, Washington, Wyoming, and Puerto Rico, and the countries of: Australia, Germany, Italy, Japan, Netherlands, New Zealand, Pakistan, Sweden, Switzerland, and Turkey. | Terminated | March 08, 2012 | 97 | 97 | Not_Computer
Z-0388-2012 | 60418 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea (available in various configuration as Vitrea Enterprise Suite). Vitrea Enterprise Suite 1.2, Vitrea Enterprise Suite 1.3, Vitrea enterprise Suite 6.0 and Vitrea Enterprise Suite 6.1.A medical diagnostic system that allows the processing, review, analysis, communication and media interchange of multi-dimensional digital images acquired from variety of imaging devices. | Vitrea | 2 | January 11, 2012 | 2012 | Vital Images, Inc. | Vital Images has found two issues that may display incorrect Standard Update Values (SUV) when viewing certain PET/CT images. This issue affects the VitreaCore PET/CT feature for Vitrea Enterprise Suite 1.2, 1.3, 6.0, 6.1, including all updates to those versions. | DESIGN: Software Design | Vital Images, Inc. sent an URGENT DEVICE RECALL letter dated November 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to inform all potential users of the issues. Customers were to be contacted by their Service Account Manager to discuss and schedule resolutions.For questions customers were to contact Vital Images Customer Support with any questions toll free at 1-800-208-3005, in Europe +31 (0) 70 4135801, in Asia t +86 10 8588 1800, or online at support@vitalimages.com. |  | 338 | 338 (233 US) | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of IRELAND, FRANCE, UNITED KINGDOM, SWITZERLAND, AUSTRALIA, SLOVENIA, GERMANY, NETHERLANDS, BELGIUM, PAKISTAN, BRAZIL, ARGENTINA, SWEDEN, QATAR, CHINA, FINLAND, ITALY, DENMARK, PORTUGAL, REPUBLIC OF KOREA, SPAIN, and CANADA | Terminated | December 03, 2013 | 692 | 338 | Computer
Z-0522-2012 | 60426 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plazaImage processing system | Siemens syngo.plaza | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Potential malfunction when using syngo.plaza. On syngo.plaza workplaces large studies can be loaded into the viewer if transferred from a DICOM node or imported even while image transfer is not yet completed. In these situations more images of that study may exist which are not shown in the syngo.plaza viewer. | DESIGN: Software Design | Siemens Medical Solutions USA sent a Customer Safety Advisory Notice dated March 10, 2011 to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the customer of the potential issues and provides additional information on the implemented resolution.For further questions please call (610) 219-4834. |  | 1 | 1 | Nationwide Distribution New York | Terminated | January 31, 2014 | 751 | 1 | Computer
Z-0444-2012 | 60455 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab used with SA INST 4.0.2.0 to 4.0.2.51, 4.0.3.0 to 4.0.3.13, 4.5.0.0 to 4.5.0.9, 4.5.2.0 to 4.5.2.8 and SA INST 3.1.3.0 to 3.1.3.29, 3.1.6.0 to 3.1.6.17, 4.0.0.0, 4.0.1.0 to 4.0.1.32.SoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable, trained, and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities, for the purpose of storing, managing, querying, and reporting laboratory data. | SoftLab | 3 | January 11, 2012 | 2012 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL, initiated a correction for the SoftLab used with SA INST. A test performed on the Roche Cobas Integra series of instruments when used with the GenInst/Dispatch listening program could cause incorrect flagging for a test result posted to the lab database. | DESIGN: Software Design | The firm, SCC Soft Computer, sent a "Correction Communication" letter on July 26, 2011 to customers/consignees. The letter describes the product problem and actions to be taken. The firm will correct the software to recognize and support all known formats of the resulting date and time sent by an instrument. The software is available via a recommended patch or hot fix to the client's current release. The customers were asked to acknowledge receipt of the task and grant permission for SCC to correct the affected configuration files. If you have any questions, call (727) 789-0100. |  | 265 | 265 | Worldwide Distribution - USA (Nationwide) and Country of: Canada. | Terminated | May 21, 2013 | 496 | 265 | Computer
Z-0691-2012 | 60829 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC) TrueBeam Linear AcceleratorVarian Medical Systems, Palo Alto, CA.The 4DITC is designed to interface to the Varian Clinac Treatment Control System for the purpose of providing verification of treatment parameters, imaging and providing a record of treatment.. The True Beam is a medical linear accelerator which delivers X-ray and electron radiation for sterotactic radiosurgery and therapeutic radiation. | 4D Integrated Treatment Console (4DITC) TrueBeam Linear Accelerator | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has discovered that on 4D Integrated Treatment Console and TrueBeam, the ad-hoc CIAO(Complete Irradiated Area Outline) calculation results in the CIAO generation that is different from the CIAO in RT Chart. The 4DITC algorithm adds dose for the control point at field edge, which should not be added according to the system specifications. | DESIGN: Device Design | Varian Medical Systems sent a Urgent Medical Device Notice dated December 22, 2011, to all affected customers via certified mail. The letter identified the product the problem and the action needed to be taken by the customer. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the recall.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Canadian account letters were sent by Fed Ex. International Customer Support Managers were sent letters for translation (when necessary) prior to sending to accounts.All of the information will be posted on the Varian web site and all appropriate personnel in the radiotherapy department should be advised of the content of this letter. A complete list of affected code numbers is included as well as contact information for Varian Oncology Help Desk. USA and Canada - 888-827-4265. |  | 4875 | 4875 units | Worldwide Distribution -- USA (nationwide) | Terminated | July 26, 2012 | 197 | 4875 | Computer
Z-0488-2012 | 60452 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8.Product Usage: SoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable, trained, and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities, for the purpose of storing, managing, querying, and reporting laboratory data. | SoftLab | 2 | January 11, 2012 | 2012 | SCC Soft Computer | On 05/03/2011, SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. A client reported that results were not flagged correctly. This can occur when a test is ordered in Order Entry, then collected | DESIGN: Software Design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" notification letter to all affected customers. The letter describes the product, Risk-to-Health notification task, potential hazard, and corrective action. The customers were asked to acknowledge receipt of the task and grant permission for SCC to make the required and recommended changes. For any questions, call (727) 789-0100. |  | 220 | 220 | Worldwide Distribution - USA (nationwide) and the country of: Canada. | Terminated | August 07, 2013 | 574 | 220 | Computer
Z-0462-2012 | 60461 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS.This correction applies to customers who use the syngo RT Therapist v4.2 or the combination of syngo RT Therapist v4.2 with syngo RT Oncologist v4.2 from Siemens.12/7/2011: Update: Syngo RTT4.2.108 SW contains the same SW content when installed on the syngo Oncologist Workspace (license key controlled). Currently the syngo Oncologist SW version is 4.2.92 P04. By installing this update, the Oncologist will match the RTT version 4.2.108.The intended use of the Siemens branded Artiste, Oncor and Primus family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | syngo RT Therapist and syngo RT Therapist | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | The software release will bring current Installed Base system to the forward production version. It is a result of the final system testing addressing several safety issues that are not critical for the safe use of the device and multiple performance defects. | DESIGN: Software Design | Siemens sent Customer Information letters and Update Instructions TH024/11/S for the installation of software release 4.2.108 beginning on November 2, 2011 to all affected consignees. Sites with active devices for the Oncology system were sent the current Update Instructions TH014/11/S for the installation of software release 4.2.108 and Customer Information letter beginning November 22, 2011.Consignees were asked to include the Customer Information in their syngo RT Oncologist System Owner Manual chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 925-602-8157. |  | 147 | 85 affected customers RT Therapa[ist, 62 customers Oncology | Worlwide Distribution - USA (nationwide) including UT, OH, TN, FL, WI, IN, KY, CA, PA, SD, NY and the countries of India, Germany, United Kingdom, Australia, Poland, Japan, Croatia, India, Columbia, Belgium, Puerto Rico, China, Sweden, South Africa and Czech Republic. | Terminated | March 21, 2012 | 70 | 147 | Computer
Z-0370-2012 | 60441 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Release 10 with RT Chart Release 10 (Client Build 10.0.28 or Build 10.0.34), An Eclipse Treatment Planning System used with ARIA Radiation Oncology (RT Chartmodule), Reference/FSCA Identifier: CP-05277, Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA.The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with Varian Eclipse and RT Chart where adding a reference pointin RT Chart and editing the field dose contributions may change the Prescribed Dose Per Fractionin Eclipse. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated October 31, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Varian has corrected the issue and will contact customers to schedule installation of the corrected version of the system. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice.. | Contact the recalling firm for information |  |  | Worldwide Distribution. | Open |  |  |  | Computer
Z-0466-2012 | 60459 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7.SoftMic is a laboratory information system to be used in medical research, clinical, and public health laboratories by knowledgeable, trained, and experienced personnel. SoftMic enables the user to handle data processing and laboratory activities for microbiology and serology testing, for the purpose of storing, managing, querying, and reporting laboratory data. | SoftMic GUI versions: 4.0.3.04.0.3.15, 4.0.4.14.0.4.7. | 2 | January 11, 2012 | 2012 | SCC Soft Computer | On 08/25/2011, SCC Soft Computer initiated a correction on the following SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. For clients using SoftMic Result Entry, when a certain series of keystrokes on the test grid are performed while entering results and statusing a test, the test status may not be maintained upon saving. The status may be removed or copied from another test on the order | DESIGN: Software Design | SCC Soft Computer initiated a Field Correction dated August 25, 2011. to all affected customers.Customers with the affected software have been notified and are scheduling receipt of the hot fix..Further questions call (727) 789-0100. |  | 144 | 144 | Worldwide Distribution -- USA (nationwide) and Canada. | Open |  |  | 144 | Computer
Z-0520-2012 | 60592 | ventilatory effort recorder | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Lifescreen Apnea is Lifescreen ECG Screening System, Version 3.31. Lifescreen Apnea is indicated for screening adults for the probability that they suffer from obstructive sleep apnea, mixed apnea or hypopnea, in order to evaluate the necessity for a polysomonographic examination. The ECG recording may be obtained at any location specified by a physician including home, hospital or clinic. Subjects screened for sleep apnea should have periods of sleep of at least 4 hours duration during which the ECG is predominantly sinus rhythm in nature. | Lifescreen Apnea | 2 | January 11, 2012 | 2012 | Del Mar Reynolds Medical, Ltd. | When Lifescreen is left open and a new recording is loaded from Sentinel Cardiology InformationManagement System, customers can get corruption in the data or a wrong ECG after the first32 hours. Both recordings have to be longer than 32 hours for this to happen. | DESIGN: Software Design | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated November 11, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to close Lifescreen every time before loading another recording from the Sentinel database. Spacelabs is in the process of upgrading software and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 523 | total 523 units and 171 kits (34 units and 7 kits in US; 489 units and 164 kits outside US) | Worldwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, KY, LA, MA, MD, MI, MN, NC, NY, OH, TN, TX, UT, WV, and PUERTO RICO and the countries of AUSTRIA, BELGIUM, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, KUWAIT, MALTA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, Slovakia, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM | Terminated | June 05, 2012 | 146 | 523 | Computer
Z-0456-2012 | 60636 | hyperthermia monitor | NZE | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ThermoSuit Manufactured for Life Recovery Systems. HD, LLC, 170 Kinnelon Rd. Kinnelon, NJ 07405 USATel.: (973) 283-2800FAx.: (973) 283-2910Temperature reduction in patients where clinically indicated; Monitoring of patient temperature | Thermosuit | 2 | January 11, 2012 | 2012 | Life Recovery Systems HD, LLC | ThermoSuit System Model TSP-50 Pump/Controller may have corrosion on the temperature circuit board which may result in incorrect patient temperature reading. | PRODUCTION CONTROLS: Process Control | Life Recovery Systems sent recall letters/return response forms dated 10/24/2011 to all consignees. Life Recovery Systems has expanded their recall to include additional units and to provide additional service to products out in the field. A new recall letter was sent to consignees on 4/27/2012. | Contact the recalling firm for information | 28 | 28 units | Nationwide Distribution including NY, NC, LA, NJ, AL. PA TX, KY, MO, MA, and TN | Terminated | January 24, 2013 | 379 | 28 | Computer
Z-0515-2012 | 60424 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, versions VA70A and VA70B Client/Server InstallationsRadiological image processing system | Siemens syngo Imaging XS, versions VA70A and VA70B Client/Server Installations | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, versions VA70A and VA70B Client/Serve installations. Loading additional images/series from Patient Browser using the thumbnails can cause an unintended deletion of images/series. | DESIGN: Software Design | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory letter dated August 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistenceis in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of CA, IN, IA, NV, NE, NY, OH, TN and TX. | Terminated | January 24, 2014 | 744 | 12 | Computer
Z-0507-2012 | 56545 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga Unified Intelligence System (UIS), This is a software-only device, WHIE Implementation v1.1i. A calculator/data processing module, for clinical use.Intended to provide the clinical user with the ability to view data from multiple, different data systems. | Microsoft Amalga Unified Intelligence System (UIS) | 2 | January 11, 2012 | 2012 | Microsoft Corporation Health Solutions Group | The database may incorrectly match patient data resulting in the incorrect patient data being displayed. This could potentially present a risk to patient safety. | DESIGN: Software Design | Microsoft notified the sole consignee by FAX on 7/22/2010 with an Urgent Medical Device Safety Information. The letter identified the affected product, the description of the problem, and actions to be taken to correct the problem, along with some additional information for the consignee. In the interim, the firm proposes to block access to patient records identified as mismatched. The firm also stated that the issue has been addressed in version 2.0. The customer should report any incidents to Michael Stokes, Director of Policy Compliance at 208-938-5397 or via e-mail at Michael.Stokes@microsoft.com. | Contact the recalling firm for information | 1 | 1 | Nationwide Distribution -- Wisconsin. | Terminated | December 20, 2011 | -22 | 1 | Computer
Z-0705-2012 | 60773 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ E9 Diagnostic Ultrasound Systems running software versions R3.1.0 and R3.1.1.The device is intended for use by a qualified physician for ultrasound evaluation of Fetal; Abdominal; Pediatric; Small Organ (breast, testes, thyroid); Neonatal Cephalic; Adult Cephalic; Cardiac (adult and pediatric); Peripheral Vascular; Musculo-skeletal Conventional and Superficial; Urology (including prostate); Transrectal; Transvaginal; Transesophageal and Intraoperative (abdominal, thoracic, vascular and neurosurgical). | GE LOGIQ E9 Diagnostic Ultrasound Systems running software versions R3.1.0 and R3.1.1. | 2 | January 11, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with the data of your LOGIQ E9 system that may impact patient safety.1 - User may notice an image from a different patient, or an image from the current patient with title information from a different patient.2  Incorrect measurements/calculations on images, when viewed on the off-line review station. As a re | DESIGN: Software Design | The firm, GE Healthcare" sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 20, 2011 to its customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, Contact Information and actions to be taken. The customers were instructed to ensure that all potential users in their facility are made aware of this safety notification and recommended actions and to follow the safety instructions noted in the letter.If you have any questions or concerns, contact your local GEHC service representative at 1-800-321-7937 (USA/Canada). |  | 894 | 894 | Worldwide distribution: USA (nationwide) including states of: AK, AZ, ALR, CA, CO, CT, FL, GA, IL, IN, IA, LA, LME, MD, MA, MI, MN, MS, MO, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TIN, TX, UT, VA, WA, WV, WI, and WY; and countries of: AUSTRALIA, AUSTRIA, BELGIUM CANADA, CHILE, CHINA, CROATIA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PHILIPPINES, PORTUGAL, REPUBLIC OF KOREA, ROMANIA, RUSSIAN FEDERATION,SAUDI ARABIA, SINGAPORE, SLOVAKIA, SPAIN,SWEDEN, SWITZERLAND, TAIWAN, and UNITED KINGDOM.. | Terminated | January 10, 2013 | 365 | 894 | Computer
Z-0713-2012 | 60830 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand ARIA Radiation Oncology, ARIA versions 10.0.39 and below, Reference/FSCA Identifier: CP-06070, Model Number: HIT, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CAProduct Usage:The ARIA Radiation Oncology product is a treatment plan and image management application. It enables the authorized user to enter, access, modify, store and archive treatment plan and image data from diagnostic studies, treatment planning, simulation, plan verification and treatment. ARIA Radiation Oncology also stores the treatment histories including dose delivered to defined sites and provides tools to verify performed treatments. ARIA also includes optional tools to manage the oncology departments clinical and business information. | ARIA Radiation Oncology | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian received customer complaints involving ARIA Oncology Information Systems,where the Planned Total Dose did not calculate correctly in Physicians Intent. It was identified that incorrect values were being stored for Planned No. Fractions, leading toan unexpected calculation. |  | Varian sent an "Urgent Medical Device Correction - Urgent Field Safety Notice" letter dated December 16, 2011 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the product, problem and recommended user actions to be taken. Customers were instructed to notify all appropriate personnel of the correction notice. The notice instructs customers to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems. For questions contact your Varian Customer Support District or Reginal Manager. |  | 2235 | 2235 units***12-12-12 AMENDED TO: 2984 units*** | Worldwide Distribution - USA (nationwide) | Open |  |  | 2235 | Computer
Z-0693-2012 | 60741 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity CT (Computed Tomography) System, 510 (k) #K033326, Model #728326, Serial Numbers: 300005, 300010, 300011, 300012, 300013, 300014, 300015, 300016, 300017, 300018, 300019, 300020 and 300021.Product Usage:The Ingenuity CT is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of X-Ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Ingenuity CT System | 2 | January 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Component defect. When the user is operating the control panel, he/she may inadvertently activate the wrong switch on the control panel which could, in turn, cause the patient couch to move in an unwanted direction. | DESIGN: Device Design | Philips sent an 'URGENT - Medical Device Correction Ingenuity CT' letter dated November 15, 2011. The letter describes the affected product, problem, hazard involved and actions to be taken by the customer and user. In this case, the customers are advised that extra care should be taken when using the 8-way Gantry Panel switch so as to avoid inadvertently activating an unintended switch resulting in unwanted (couch) movement. The letter notifies the customers that a Philips representative will be visiting each customer site and installing a new control panel through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, option 5; enter site ID or follow prompts). |  | 13 | 13 system units were distributed (5 in U.S. & 8 overseas) | Worldwide Distribution - USA (nationwide) in the states of: LA, NY, OH, WV and the foreign countries of: Australia, Germany, Korea, India, Malaysiaa and New Zealand. | Terminated | February 22, 2013 | 408 | 13 | Computer
Z-0715-2012 | 60826 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) v10.2.3, Reference/FSCA Identifier: CP-06873,Model Number: H51, distributed by and manufactured by:Varian Medical Systems Inc., Palo Alto, CAThe 4D Integrated Treatment Console [4DITC] is designed to interface to the Varian Clinac Treatment Control System [TCS] for the purpose of providing verification of treatment parameters, imaging and providing a record of treatment. The general function of the 4D Integrated Treatment Console is to allow treatment plans and images to be retrieved from the Oncology Information System [OIS] and sent to the TCS, planned treatment plan parameters to be verified against the TCS delivery parameters for accuracy and treatment history to be recorded in the OIS for use and display. The Barcode Conical Collimator Verification [BCCV] device assists operators of radiation therapy devices by preventing irradiation until the conical collimator required by the treatment plan is in place. | 4D Integrated Treatment Console (4DITC) | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | When multiple conical collimator treatment plans are scheduled to be treated in a session, thePlan Label should be displayed on the BCCV dialog box should match the Plan Label on the4DITC. However, the name of the plan is not being updated correctly. | DESIGN: Software Design | The firm, VARIAN Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" dated December 12, 2011, to all affected customers. The notice described the product, problem and user corrective action steps. The customers were provided to do the following: Do not use BCCV as a primary reference for identifying the correct plan to be treated; use caution when selecting an isocenter in a positioning system that is not interfaced to the 4DITC; any discrepancies in planned versus actual parameters should be brought to the attention of the medical physicist for investigation before treatment begins, and one workaround is to schedule one plan per session. Special instructions for customers outside the USA and Canada: complete the attached Proof of Notification or Receipt Verification Card once you have read this document and return it to Varian Medical Systems.Varian is developing a technical correction for this issue. The customers will be contacted by a customer representative when this technical correction is available to schedule installation on their system.If you require further clarification, please feel free to call your local Varian Customer Support District or Regional Manager: USA and Canada: 1-888-827-4265; Europe: +41 41 749 8844; Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com and Latin America: soporte.al@varian.com; and Internet: www.myvarian.com or www.varian.com.Additional recall letters were sent on 11/5/12 to additional consignees that Varian determined needed to be notified of the recall. |  | 55 | 55 | Worldwide distribution:USA (nationwide) and countries of: Canada, Australia, Austria, Brazil, China, France, Georgia, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Latvia, Lithuania, Macau. Malaysia, Mexico, Netherlands, Singapore, Slovenia, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Ukraine, United Kingdom.. | Terminated | December 05, 2013 | 694 | 64 | Computer
Z-0486-2012 | 60357 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare AW Server version 2.0 with release 2.0 - 4.0 A medical software system that allows multiple users to remotely access AWapplications from compatible computers on a network. | GE Healthcare AW Server | 2 | January 11, 2012 | 2012 | GE Healthcare, LLC | 2 Unrelated Software Issues: An incorrect patient name maybe displayed when switching between two Volume Viewer tabs with different patients. Also, if a second patient exam is selected while the application is starting, the second patient's data will be displayed with the first patient's name display on the applications tab. | DESIGN: Software Design | Consignees were sent on 11/17/11 a GE "Urgent Medical Device Correction" letters dated November, 16 2011 The letter was addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letter described the Safety Issue, Safety Instructions Affected Product Details, Product Correction and Contact Information. For the first issue, customers were asked to contact their field engineer to install the correct configuration file. Also, customers should not apply the recommended changes from the Customer Release Note as those changes cause the issue to occur. With the second issue, customers are to wait until the application has fully initialized before making another patient exam selection. GE Healthcare will correct all affected AW server installations via a software upgrade at no cost to the customer. A representative will contact customers to arrange for the correction. Questions should be directed to customers' local GE Healthcare Service Representative or at the numbers provided in the letter. | No consumer action necessary | 308 | 308 | Worldwide Distribution -- USA, including states of AZ, CA, CO, CT, DE, DC, FL, HI, IL, IN, LA, MD, MI, MN, MS,MO, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI and countries of UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, REUNION, QATAR, POLAND, NORWAY, NEW ZEALAND, LEBANON, SOUTH KOREA, JAPAN, ITALY, ISRAEL, HONG KONG, GUADELOUPE, GREAT BRITAIN, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CANADA, BRAZIL, BOSNIA AND HERZEGOVINA, BELGIUM, and AUSTRALIA. | Terminated | July 03, 2012 | 174 | 308 | Computer
Z-0372-2012 | 60307 | stethoscope, esophageal, with electrical conductors | BZT | Esophageal stethoscope with electrical conductors. | Anesthesiology | Anesthesiology | 510(k) | Cincinnati Sub Zero, Steri-Probe Temperature Probe, Esophageal/Rectal Temperature Probe, Part Number: 39622, Catalog #491B. TA disposable temperature probe used in conjunction with a measuring/monitoring device to continuously measure/monitor a patient's core body temperature. | SteriProbe Temperature Probe | 2 | January 11, 2012 | 2012 | Cincinnati Sub-Zero Products Inc | Reports that the Rectal/ Esophageal Temperature Probes are not reading correctly. This situation causes the Blanketrol device, which the probes connect to, to shut down operation and display a 'check probe' error. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cincinnati Sub-Zero Products, Inc., issued an "Urgent Medical Device Field Action" letter on November 1, 2011 to all affected customers. The letter includes; description of product and problem, and a list of the affected lot numbers. The letter instructs the customers to discontinue use of the device. All unused product is to be returned to the Biomedical/Clinical Engineering and/or Service Department for an replacement and all used probes are to be disposed of in accordance with hospital policy. The customers are also instructed to complete and return an attached Recall Response Form ASAP. For additional information, customers may contact the Beatrice Washington at 1-800-989-7373 or (513) 772-8810. |  | 369680 | 369,680 | Nationwide distribution and Canada. | Terminated | August 13, 2013 | 580 | 369680 | Computer
Z-0672-2012 | 60151 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500Product Usage: The CDl Blood Parameter Monitoring System 500 is intended for use during cardiopulmonary bypass procedures when continuous monitoring of arterial and/or venous pH, PC02 , P02, K+, oxygen saturation, hematocrit, hemoglobin, and temperature are desired. | CDI 500 Monitor | 2 | January 11, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | In a recent review of in-house Service Procedures for the CDI Blood Parameter Monitoring System 500, Terumo Cardiovascular Systems (Terumo CVS) found that one of the procedures it had been performing on older systems should have been considered a recall at the time it was implemented:Terumo CVS began installing a protective shield (plastic grommet) in the system chassis during manufacture of all | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Terumo sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 6, 2011. The letter identified the affected products, the problem, and the actions to be taken by customer/user. The letter states that Terumo CVS will update all affected CDl Monitors and no additional actions are necessary for those units previously returned that received the correction as part of the service process. For questions or concerns contact Terumo CVS Customer Service 1-800-521-2818. |  | 451 | 451 units total | Worldwide Distribution - USA (nationwide in the states of: AR, AZ, CA, FL, KS, IL, IN, LA, MA, MO, MI, MS, NJ, NV, NC, OH, PA, TN, TX, VA, WA, WI, WV including Puerto Rico and the countries of: AE, Australia, BR, Belgium, Brazil, Canada, China, Dominican Republic, EC, Germany, Guatemala, Japan, Korea, Mexico, SA, Taiwan and UY. | Terminated | March 01, 2013 | 415 | 451 | Computer
Z-0591-2012 | 60625 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Computed Radiography Systems with DX-G Digitizer.DX-G Digitizers haved the same intended use as the predicate devices. they are intended to use in providing diagnostic images to aid the physician with diagnosis. DX-G Digitizers provide the ability to use both standand needle phosphor image plates and can be used with either Musica or Musica 2. In the US, Agfa's Computed Radiography Systems with DX-G Digitizers are not intended for use in mammography. | DXG Digitizer | 2 | January 11, 2012 | 2012 | AGFA Corp. | Occasionally images can be lost or sent to an incorrect folder or the digitizer may stop either during startup or during a cassette scan cycle. | DESIGN: Software Design | Agfa HealthCare sent a 'URGENT SAFETY NOTICE" letter dated December 6, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Agfa Service will contact your facility and schedule a software upgrade if this has not already been performed. Please complete the attached feedback form as soon as possible and fax it to 864-421-1662.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 125 | 125 | Worldwide Distribution-- USA (nationwide) including the states of AZ, CA, CO, FL, GA, HI, IL, IN, KS, LA, ME, MI, MN, MO, NE, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI and WV, and Canada. | Terminated | April 11, 2012 | 91 | 125 | Computer
Z-0583-2012 | 60258 | apparatus, nitric oxide delivery | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DS (Delivery System) Model 10003 & 10007.The INOMAX DS Delivery System delivers INOMAX (nitric oxide for inhalation) therapy gas into the inspiratory limb of the patient breathing circuit in a way that provides a constant concentration of nitric oxide (NO), as set by the user, to the patient throughout the inspired breath. It uses a specially designed injector module, which enables tracking of the ventilator waveforms and the delivery of a synchronized and proportional dose of NO. It may be used with most ventilators. The INOMAX DS provides continuous integrated monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. | INOmax DS (Delivery System) | 1 | January 11, 2012 | 2012 | Ino Therapeutics | Through the review and monitoring of its complaint data, Ikaria identified a product anomaly with the INOmax DS Model 1003. Ikaria concluded the anomaly risk to the patient safety was low and non reportable (4/23/10). During a subsequent Level 3 inspection by FDA, it was concluded that the action should have been reported under 21 CFR & 806.10. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ikaria sent a "RECALL- CUSTOMER PRODUCT REMEDIATION" notice dated December 22, 2011 to all affected customers. The letter describes the product, problem, actions taken by the firm. No action was required by customers. The notice served as correspondence to inform customers of the issue. Customers may direct their questions to Ikaria Customer Care at 1-877-566-9466. | Contact the recalling firm for information | 2446 | 2446 | Worldwide Distribution-USA (nationwide) and the country of MALAYSIA. | Terminated | May 29, 2012 | 139 | 2446 | Computer
Z-0598-2012 | 60521 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Product Code: 2M8161. | Colleague CX Infusion Pump | 2 | January 11, 2012 | 2012 | Baxter Healthcare Corp. | There was no documentation to show the required tests were performed following the replacement of the battery and battery harness. | PRODUCTION CONTROLS: Process Control | Baxter sent a service technician to the consignee location on 05/08/2009 to perform the battery check and discharge test. The pump was found to be within specification. | Contact the recalling firm for information | 1 | one pump | Puerto Rico | Terminated | January 03, 2012 | -8 | 1 | Computer
Z-0594-2012 | 60549 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning System, software version 9.0.Model # Description453560446041Pinnacle3 version 9.0 English453560446051Pinnacle3 version 9.0 Simplified Chinese453560446061Pinnacle3 version 9.0 Traditional Chinese453560446091 Pinnacle3 version 9.0 Dutch453560446101Pinnacle3 version 9.0 French453560446111Pinnacle3 version 9.0 German453560446141Pinnacle3 version 9.0 Italian453560446161Pinnacle3 version 9.0 Japanese453560446181Pinnacle3 version 9.0 Spanish453560446201Pinnacle3 version 9.0 TurkishIntended to provide support for radiation therapy treatment planning for the treatment of benign or malignant disease processes. | Pinnacle3 Radiation Therapy Planning System | 2 | January 11, 2012 | 2012 | Philips Medical Systems (Cleveland), Inc. | Phillips Medical Systems have recently determined that a software (Pinnacle3 Software version 9.0) nonconformance can cause incorrect beam geometry. | DESIGN: Software Design | Consignees were sent on 11/18/11 a Philips "Urgent Medical Device Correction" letter dated 2011 Nov 16. The letter identified the affected product and described the Problem, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer/User, Actions Planned by Philips and Further information and Support. If customers observed the issue, they were to manually make the suggested change. In addition, customers were to place the letter in the Instructions for Use until otherwise notified. The firm will deliver 9.2 software release free of charge. Customers in North America are to contact Customer Care Solutions Center at 1-800-722-9377, if they have questions or need support concerning the recall. |  | 715 | 715 (463 US) | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and countries of AUSTRALIA, AUSTRIA, BELGIUM, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, MALAYSIA, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, and CANADA. | Terminated | February 28, 2014 | 779 | 715 | Computer
Z-0516-2012 | 60684 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Portal Vision, R-Arm with 4D Integrated Treatment Console;Varian Medical Systems. Palo Alto, CA.The Portal Vision device is used for verification of the treatment field and shielding block in relation to anatomical landmarks in radiotherapy treatment. Portal Vision will also allow for verification of the exit dose in radiotherapy treatment. | Portal Vision (Electronic Portal Image Device) RArm with 4D Integrated Treatment Console | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was discovered when using the 4DITC in combination with an R-Arm to acquire double exposure MV images. | DESIGN: Software Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated December 1, 2011 to all affected sites via Certified mail, return receipt. Canadian accounts were sent by Fed Ex and internationally to Varian International Customer Support Managers for translation and notification of customers. The notice described the product, problem and actions to be taken. The customers were instructed to always acquire both the treatment port and the open port images BEFORE displaying the digital graticule, when using a standard double exposure template; and to advise the appropriate personnel working in their radiotherapy department of the content of this letter.Special instructions for customers outside the USA and Canada: these customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems, once they have read this document. Varian is developing a correction for this issue and a Customer Service rep will contact the customers when the correction is available to schedule its installation on their systems.If you have any questions, contact USA and Canada at 1-888-827-4265; Europe: +41 41 749 8844; Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: China.apps.helpdesk@varian.com; Japan: Japan-Apps.Helpdesk@varian.com and Latin America: soporte.al@varian.com or Internet: Oncology systems customer site: www.myvarian.com and Varian Medical Systems public site: www.varian.com. | Contact the recalling firm for information | 1486 | 1486 devices | Worldwide distribution: USA (nationwide) and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France,, Georgia, Germany,Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Lativia, Macedonia, Malaysia, Martinique, Mexico, Montenegro, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Open |  |  | 1486 | Computer
Z-0399-2012 | 60230 | calibrator for hemoglobin and hematocrit measurement | KRZ | Calibrator for hemoglobin or hematocrit measurement. | Hematology | Hematology | 510(k) | VARIANT II Link Hemoglobin A1c Program Reorder Pack, 1600 tests, IVD, For the determination of Hemoglobin A1c in human whole blood.Contents:Elution Buffer A , 1 x 2500 mLElution Buffer B, 1 x 2100 mLWash/Diluent Solution, 1 x 2500 mLCartridge SetCD ROM - 1Calibrator/Diluent SetWhole Blood Primer, 6 x 1 mLSample vials, 2 x 100Instruction ManualBio-Rad Laboratories, Inc., 400 Alfred nobel Drive, Hercules, CA 94547-1803.Intended for the percent determination of hemoglobin A1c in human whole blood using ion-exchange high-performance liquid chromatography (HLCP). | VARIANT II Link Hemoglobin A1c Program Reorder Pack | 3 | January 11, 2012 | 2012 | Bio-Rad Laboratories, Inc. | It was discovered that the Variant II Turbo Link Calibrator Values card contains incorrect values for IFCC units. The numbers are missing a decimal point. | PRODUCTION CONTROLS: Process Control | Bio-Rad Laboratories, Inc. sent a Medical Device Correction Notification letter dated October 25, 2011, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The correction was initiated by the firm, customers were instructed to contact their regional Bio-Rad office for any questions.For questions regarding this recall call 510-741-6114. |  | 63 | 63 | Nationwide Distribution including MI, SC, OH, IL, MN, MD, VA, NY and WI. | Terminated | February 02, 2012 | 22 | 63 | Computer
Z-0429-2012 | 60134 | implantable pacemaker pulse-generator | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular |  | Implantable Pacemaker Accent DR models PM2110 and PM2210St. Jude Medical pacemakers monitor and regulate a patient's heart rate with pacing therapy. | Accent DR | 2 | January 11, 2012 | 2012 | St Jude Medical CRMD | St. Jude Medical has received reports of out of range pacing lead impedance (PLI) measurements.These measurements are low and out of range ( | DESIGN: Software Design | The firm, St. Jude Medical, sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated September 22, 2011 to its customers by FedEx priority overnight service for planned receipt on September 23, 2011. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed that in order to prevent a false reading, a new programmer software version (12.1.0.1) is available. This software when used to interrogate an Accent DR or Anthem CRT-P pacemaker will eliminate the potential for this anomaly to occur. The onetime upgrade is performed automatically on affected devices and will not change the operation of the implanted device. A St. Jude Medical Sales Representative will assist the customers in loading the new programmer software onto their Merlin programmer. Customers with questions or concerns, can contact their local St. Jude Medical representative or their Technical services department at (800) 722-3774. | Contact the recalling firm for information | 152000 | 152,000 units | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Asia Pacific, Belgium, Bulgaria, Canada, China, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, India, Iran, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latin America, Lebanon, Malaysia, Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, UK, and U.A.E. | Terminated | April 10, 2012 | 90 | 166000 | Computer
Z-0514-2012 | 60425 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, all versionsRadiological image processing system | Siemens syngo Imaging XS, all versions | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, all versions. The "Rename" functionality can cause an unintended patient merge if the renamed patient is sent to another system. | DESIGN: Software Design | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory Notice letter dated August 3, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistencies in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 38 | 38 | Nationwide Distribution including the states of CA, FL, GA, IN, IA, KS, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Open |  |  | 38 | Computer
Z-0638-2012 | 60631 | instrument, measuring, corneal radius | HJB | Corneal radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Octopus 900 Perimeter, EyeSuite Perimetry version i4.000 (Flash v. 2.1.4 & 2.1.5), Model Number: Octopus 900, Distributed by: Haag-Streit USA, Inc., 3535 Kings Mills Rd., Mason, OH 45040 Product Usage: The OCTOPUS 900 Perimeter is designed for the examination, analysis and documentation of the field of sight, especially the light difference sensitivity and other functions of the human eye. | Octopus 900 | 2 | January 11, 2012 | 2012 | Haag-Streit USA Inc | Currently there is a possibility, at the start of the perimetry examination, for the background illumination of the cupola not to turn on. If no illumination of cupola occurs, data obtained from the examination could provide the doctor with results that would appear to be better than actual. | DESIGN: Software Design | Haag-Streit USA Inc sent an Urgent Field Safety Notice letter dated November 11, 2011 to their customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to check their unit and if needed upgrade their unit with the new software provided. Customers were advised to complete and return the Software Update Verification Form included. Customers who had distributed the affected product outside of their facility were instructed to immediately notify their customers of the recall. | Contact the recalling firm for information | 158 | 158 Units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA and WA and the country of Mexico. | Open |  |  | 158 | Computer
Z-0737-2012 | 60797 | medical device data system | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | NicVue Software version 2.9.2 and 3.0.1. Medical Device Data System (MDDS) Intended use: The NicVue software is intended to be the patient database for various Nicolet applications. The NicVue software is intended to aid the hospital in tracking patient data from collection to review to archive. The NicVue Software includes the following features and intended uses: Launch Applications, Enter patient information, Import/Export patient information from the Hospital Information System, Move data files, Archive data files on DVD or CDROM. | NicVue Software version 2.9.2 and 3.0.1. | 2 | January 13, 2012 | 2012 | CareFusion 209 Inc. | Under certain work flow processes, when selecting a new patient name in NicVue, the associated exams list may not properly refresh, and may continue to display the previous patient's exams. | DESIGN: Software Design | The firm, CareFusion, sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated December 16, 2011 to its customers. The letter included a "Description of the issue" and "Actions you must take to correct this issue". The customers were instructed to follow the instructions included with the CD to update all of your NicVue systems; complete and return the SOFTWARE UPDATE VERIFICATION FORM to CareFusion-NeuroCare via fax to: +1 608 829 8517 ; mail to ATTN: Regulatory Affairs, CareFusion-NeuroCare, 1850 Deming Way, Middleton, WI USA 53562 and/or email (PDF) to: Susan.Niesen@CareFusion.com, and if you have transferred any of these products to another location, send those users a copy of this Field Correction notification and notify CareFusion-NeuroCare at +1 800 356 0007 or +1 608 829 8500 ext. 5129, or via email at susan.niesen@carefusion.com of this transfer.If assistance is needed in completing these actions, contact CareFusion NeuroCare Call Center at +1 800 356 0007 or +1 608 829 8500 for further directions; press option 2 for Technical Support. | Contact the recalling firm for information | 1062 | 1,676 ( 1062 OUS, 614 US) | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, DE, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA,WV, and WI; and countries of: ARGENTINA, AUSTRALIA, BULGARIA, BANGLADESH, BELGIUM, CHILE COLOMBIA, CYPRUS, CZECH REPUBLIC, CANADA, CHINA, CROATIA, CURACO, DENMARK, DOMINICANA, EGYPT, FRANCE,GERMANY, HUNGARY, HONG KONG, IRELAND, INDIA, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NORWAY, NETHERLANDS, OMAN, PARAGUAY, PERU, POLAND, PORTUGAL, PALESTINIAN, PAKISTAN, PHILIPPINES,QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SERBIA, SLOVENIA, SPAIN, SLOVAKIA, SWITZERLAND, SRI LANKA, RUSSIA, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, URUGUAY, UNITED ARB EMIRATES, and VIET NAM. | Terminated | October 11, 2012 | 272 | 1676 | Computer
Z-0736-2012 | 60837 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Treatment Planning System, 7.3, 8.0, 8.1, 8.2, 8.6, 8.9, 10.0; Model Number: H48; Reference/FSCA Identifier: CP-07007, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithim is specifically indicated for planning proton treatment of neoplasms of the eye. | Varian brand Eclipse Treatment Planning System | 2 | January 13, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the Eclipse Treatment Planning System when planning for the Siemens Multileaf Collimator (MLC) where a virtual jaw can be positioned inside the MLC aperture and then used as the beam limiting device for the dose calculation. This will result in an incorrect dose distribution within Eclipse. | DESIGN: Software Design | Varian Medical Systems sent a Urgent Medical Device Correction letter dated December 12, 2011, to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to pay careful attention to any warnings displayed regarding the Virtual Collimator Jaws when the plan is calculated. If the warning indicates that the Virtual Collimator Jaws are inside the MLC aperture, do not proceed with the calculation.Varian is continuing to evaluate possible technical solutions for this issue.Special Instructions for customers outside the USA and Canada: In order to satisfy regulatory requirements, we request that you complete the attached Proof of Notification or Receipt Verification Card once you have read this document and return it to Varian Medical Systems."We sincerely apologize for any inconvenience and thank you in advance for your cooperation."Please call for USA and Canada: 1.888.VARIAN5 (888.827.4265)Europe: +41 41 749 8844. |  | 11807 | 11807 units, ***12-12-12 Amended to: 13, 616 units*** | Class II Recall - Worldwide Distribution -- USA (nationwide) including the countries of Canada and Europe. | Open |  |  | 11807 | Computer
Z-0744-2012 | 60520 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015. Product Code: 2M8161. Product Usage: Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. Electronic infusion pumps indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration, such as intravenous (IV), intra-arterial (IA), subcutaneous, epidural or irrigation of fluid spaces applications. | Colleague CX Volumetric Infusion Pump | 2 | January 13, 2012 | 2012 | Baxter Healthcare Corp. | Baxter's Service Center did not perform the ground impedance and leakage current test which verify if any power overflow occurs and that it is properly discharged through the device electrical system. | TRAINING: Employee Error | Baxter sent service technicians to the consignees location on July 14, 2010 and October 6, 2010 and conducted the ground impedance and leakage current test on the affected pumps. The pumps were found with specifications. |  | 2 | 2 units | Nationwide distribution: Puerto Rico. | Terminated | January 13, 2012 | 0 | 3 | Computer
Z-0779-2012 | 59887 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity CT (Computed Tomography) System, 510 (k) #K033326, Model #728326, Serial Numbers: 300003, 300004, and 300005. The recalled CT systems are equipped with Software Version 3.5 Beta 2.Product Usage: The Brilliance Big Bore is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at difference angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Ingenuity CT system | 2 | January 17, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64, Brilliance Big Bore, and Ingenuity CT systems, equipped with software version 3.5 Beta 2. The non-conformances are due to software errors identified in these device systems. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips sent an "URGENT - Medical Device Correction" notification letters to all affected customers via Certified mail on October 19, 2011. The letter describes the referenced software issues; provides instructions to the customers as to what they can do in order to address these issues; and informs the customers of the firm's upcoming software field/correction. The letter states that a Philips representative will be contacting all affected customers to install the necessary software upgrade through a Field Change Order. This will be implemented free of charge.For information or support contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solution Center 1-800-722-9377, option 5. |  | 3 | 3 units | Worldwide Distribution - Nationwide (USA) including the states of: IN, MI, MO, NY, OH, and OR and the countries of Germany, Israel and Netherlands. | Terminated | February 22, 2013 | 402 | 3 | Computer
Z-0776-2012 | 60764 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Ultraview SL Command Module, Model 91496.The Spacelabs Multi-parameter Module is intended for use with the Patient Care Management System (PCMS) to acquire, monitor, and process various clinical parameters from an adult or neonate/infant populations in any type of clinical environment other than home use. Physiological parameters that may be monitored include ECG with arrhythmia detection, respiration, invasive and noninvasive blood pressure, temperature, oxygen saturation (SpO2) and cardiac output. Acquired data may then be communicated to an information network for display, recording, editing and analysis. | Spacelabs Medical Ultraview SL Command Module, Model 91496 | 2 | January 17, 2012 | 2012 | Spacelabs Healthcare, Llc | One Spacelabs Medical Ultraview SL Command Module, Model 91496, was removed from customer due to software error. | DESIGN: Software Design | Spacelabs removed one affected unit of Medical Ultraview SL Command Module from a hospital in Troy, MI. during a site visit on July 13, 2011. This recall pertains to one single unit at one location. Contact Spacelabs at 425-657-7200 for questions regarding this recall. | No consumer action necessary | 1 | one unit | One device distributed in Troy, MI. | Terminated | January 20, 2012 | 3 | 1 | Computer
Z-0805-2012 | 60933 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF 16 PET/CT System (with software version 3.5.1 and 3.5.1.1), Model # 882470 & GEMINI TF 64 PET/CT System (with software version 3.5.1 and 3.5.1.1), Model #882471 Philips Medical Systems, Highland Heights, OHThe device is a diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. | Gemini TF 16 and Gemini TF 64 PET/CT Systems Software Versions 3.5.1 and 3.5.1.1 | 2 | January 18, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips received notification by a customer of an issue relating to unwanted auto delete occurrences resulting in raw data being deleted and, in some cases, raw data being deleted immediately after scan acquisition but before images are reconstructed. | DESIGN: Software Design | Philips Medical System sent a Urgent Medical Device Correction letter dated December 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The following are recommended actions:1. Please ask your local Philips Service Engineer to ensure that the Auto Delete function is configured to run automatically.2. Please share this information with all Philips GEMINI TF system operators at your site.3. Retain this document with your system's instructions for Use Manual.For further information or support concerning this issue please contact your local Philips representative in the United States and Canada at 1-800-722-9377, and follow the recorded menu options to reach Customer Solutions Engineer; in all other countries please dial your local Philips Healthcare office. |  | 197 | 197 units | Worldwide Distribution -- USA (nationwide) including the states: AL, AR, AZ, CA, CO, DC, FL, GA, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NY, OH, OR, PA, PR, TN, TX, WA and WV. and the countries of: Argentina, Australia, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Denmark, Egypt, France, Germany, Hungary, India, Israel, Italy, Japan, Netherlands, Pakistan, Philippines, Poland, Russian Fed, Saudi, Arabia, Singapore, South Korea, Spain, Sweden, Taiwan, Thailand, Turkey, Ukraine and United Kingdom. | Terminated | February 22, 2013 | 401 | 197 | Computer
Z-0827-2012 | 60761 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza VA20A Server Farm setup integrated to a RIS, model number 10592457. | syngo.plaza VA20A, model number 10592457 | 2 | January 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | The firm has become aware of the potential for an unintended merge of multiple patients when using syngo.plaza version VA20 with Server Farm setup integrated to a RIS. Sporadically it can happen that certain Patient IDs in the database are deleted and sent to NULL. This can cause incorrect merging with patients that already had a NULL Patient ID. | DESIGN: Software Design | The firm initiated their product recall in May 2011 by issuing a "Customer Safety Advisory Notice" informing all customers of the malfunction and providing instructions on how to avoid the error. The issue will be solved with the next version release, planned for June 2011. |  | 3 | 3 | Nationwide Distribution - Including the states of Florida, Massachusetts and Wisconsin. | Terminated | January 24, 2014 | 735 | 3 | Computer
Z-0834-2012 | 60781 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ P6 Ultrasound System.The device is intended for use by a qualified physician for the evaluation of soft tissue and blood flow in the clinical applications of: Fetal; Abdominal; Pediatric; Small Organ (breast, testes, thyroid); Neonatal Cephalic; Adult Cephalic; Cardiac (adult and pediatric); Peripheral Vascular; Musculo-skeletal Conventional and Superficial; Urology (including prostate); Transesophageal (TE); Transrectal (TR); Transvaginal (TV); and intraoperative (abdominal, thoracic, vascular and neurological). | GE LOGIQ P6 Ultrasound System | 2 | January 21, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a safety issue associated with the Probe of your LOGIC P6 system that may impact patient safety. GE Logic P6 probe head has a possibility to cause burn injuries to patients or operators. | DESIGN: Software Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 23, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the affected product. Contact the GE Support Center at 1-800-321-7937 or your local GE Healthcare service representative for questions regarding this issue. |  | 2270 | 2270 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KY, LA, ME, MA, MI, MN, MS, MO, MT, NE, NV, NH, NY, NC, ND, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, and WI. and countries of VENEZUELA, UNITED KINGDOM TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARAIA, RUSSIAN FEDERATION, ROMANIA, REPUBLIC OF KOREA, QATAR, PROTUGAL, POLAND, PHILIPPINES, PERU, OMAN, NORWAY, MEXICO, MALAYSIA, LITHUANIA, LEBANON, KOSOVO, KENYA, KAZAKHSTAN, JAPAN, JAMAICA, ITALY, ONDONESIA, INDIA, HUNGRY, HONG KONG, GUATEMALA, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BAHRAIN, AUSTRIA, and ALGERIA. | Terminated | January 21, 2014 | 731 | 2270 | Computer
Z-0840-2012 | 60463 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Roller Pump, 6 Inch DiameterThe Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | The Terumo Advanced Perfusion System 1 | 2 | January 24, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) discovered and rejected a single digital encoder cable assembly during incoming inspection because it had a loose wire. The assembly is used in the 4" and the 6" roller pumps for the Terumo Advanced Perfusion System 1. | PRODUCTION CONTROLS: Process Control | TERUMO sent an Urgent Medical Device Correction letter dated January 13, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to review the Medical Device Correction notice. Assure that all users are aware of the notice. Confirm receipt of the communication by faxing the attached Customer Response Form to 1-900-292-6551. Roller pumps with affected cables will have the cables assembly replaced.For questions regarding this recall call 1-800-521-2818. |  | 44 | 44 | Worldwide Distribution - USA including MD, WI, and IA and the countries of Canada, Mexico, and Vietnam | Terminated | December 03, 2012 | 314 | 44 | Computer
Z-0845-2012 | 60873 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray Systems (Brilliance 64, Model Number: 728231; Brilliance Big Bore, Model Number: 728243; Brilliance iCT, Model Number: 728306; Brilliance iCT SP, Model Number: 728311; and Ingenuity CT, Model Number: 728326); Mfr. by Philips Medical Systems; Cleveland, OH. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Brilliance 64, Brilliance Big Bore, Brilliance iCT, Brilliance iCT SP, Ingenuity CT | 2 | January 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The Common Image Reconstruction System (CIRS) on the affected Brilliance Systems encounters an error and requires reboot of the system. The unit failed to initialize axial scan due to timeout that occurred waiting for "READY" state | DESIGN: Software Design | On 10/10/2011 the firm sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected products, the problem, the hazard involved, the action to be taken by the customer, and actions planned by Philips. Customers were to reboot their systems if the two possible scenarios occurred. Philips Healthcare plans on installing the update through a Field Change Order free of charge. For further information or support concerning the issue, contact a local Philips representative or local Philips Healthcare office. In North America and Canada, contact Customer Care Solutions Center at 1-800-722-9377, option 5. | Contact the recalling firm for information | 548 | 548 units | Worldwide Distribution -- USA, including the states of: AL, AZ, CA, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NM, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA & WV and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BANGLADESH, BELGIUM, Brazil, CHINA, COLUMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KOREA, LUXEMBOURG, MALAYSIA, MALI, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PHILIPPINES, POLAND, QATAR, RUSSIA, SINGAPORE, SLOVENIA, SOUTH KOREA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN & VIETNAM. | Terminated | June 18, 2013 | 511 | 548 | Computer
Z-0869-2012 | 60742 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | STERIS SYSTEM 1E Liquid Chemical Sterilant Processing System. Model P6500, 510 (k) Numbers: K090036 Cleared 4/05/2010; K101409 cleared 8/2/2010; K102462, cleared 9/2/2010; add to file 2/24/2011.The SYSTEM 1E Liquid Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned immersible and reuseable critical and semi-critical heat sensitive medical devices, including endoscopes and their accessories. | SYSTEM 1E Liquid Chemical Sterilant Processing Systm | 2 | January 25, 2012 | 2012 | Steris Corporation Hopkins Facility | Software deficiency. STERIS identified that water specifications were not being met at some of their customer sites. Critical parameters for initiating cycles are: water quality, water pressure and water temperature. Failure to meet these criteriae may lead to cycle aborts by the SYSTEM 1E Liquid Chemical Sterilant Processing System. The firm intends to perform a software field correction which | DESIGN: Software Design | The firm, STERIS Corp. prepared an "URGENT VOLUNTARY FIELD CORRECTION NOTICE" letter dated December 8, 2011. The firm intends to send the letter to its customers upon STERIS' receipt of 510(k) clearance from FDA. The letter includes a description of the Product-SYSTEM ÂŽÂŽ Liquid Chemical Sterilant Processing System, problem and actions taken. The customers were instructed to in case of an incomplete or cancelled cycle, reprocess devices in the cycle following the directions in the Operator manual. The letter also states that STERIS Corporation will install updated software on the affected system, and will be providing the customers with a new Operator Manual reflecting the revised wording resulting from the installation of the new software. For further information or if they have any questions regarding the firms visit to their facility, contact STERIS Field Service Dispatch at 1-800-333-8828. If you have questions regarding this matter, contact Director Low Temperature Sterilization, at 440-392-7455. | No consumer action necessary | 3882 | 3,882 units were distributed | Worldwide distribution: USA (nationwide) and countries of: Bahrain, Hong Kong, Japan and South Korea. | Open |  |  | 3882 | Computer
Z-0802-2012 | 60372 | electrode, cutaneous | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | BIS Bilateral Sensors, part number 186-0212, packaged in boxes of 10. For use with BIS Monitoring System, Labeled as either Covidien Llc, or Aspect Medical Systems, Inc.Product Usage:To enable recordings of electrophysiological (such as EEG) signals. The Sensor is a single patient use, disposable, pre-gelled 6 electrode array that is applied directly to the patients forehead. BIS technology is used as an adjunct to clinical judgement and training. Reliance on BIS technology alone for intraoperative anesthetic management is not recommended. | BIS BILATERAL SENSOR | 1 | January 26, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | During manufacture, circuit traces were reversed which causes inaccurate display of information. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Finished Device Change Control | Covidien sent an Urgent Medical Device Recall letter dated November 8, 2011, to all affected customers. The letter identified the product, problem, and actions to be taken. Customers were instructed to segregate and return any product with the affected lot numbers according to the directions provided. If product was further distributed, the recall was to be continued to those customers. Further information is available at 800-635-5267. |  | 8510 | 8510 units (851 boxes of 10) | Nationwide Distribution (USA) including the states of: AZ, CA, CO, FL, GA, IL, KY, KS, LA, MA, MO, MS, NC, NM, NY, OH, PA, SD, TN, TX, UT, VA, WA, WI, WV and WY. | Terminated | March 13, 2013 | 412 | 8510 | Computer
Z-0801-2012 | 60223 | electrode, cortical | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Nicolet C64 Stimulus Switching Unit (SSU Amplifier), Catalog/ Part Numbers: 382-610800, 688-624200, 688-624300, 842-687900, 672-610800F.Product Usage:The SSU Amplifier is used with the Cortical Stimulator Control unit which is used for functional brain mapping procedures during treatment of patients with seizure disorder and brain tumors, providing stimulation via electrode pairs or a hand held bipolar probe. | Nicolet C64 Stimulus Switching Unit (SSU Amplifier), | 1 | January 26, 2012 | 2012 | CareFusion 209 Inc. | CareFusion is recalling the NicoletOne Software v5.60.2-5.71.4 used with the Cortical Stimulator Control Unit. It is possible that during cortical stimulation the electrode annotation label may be incorrect on the screen and in the report when the current is correctly delivered through the selected electrodes. Also, it has been determined that the stimulus cable connection in the Cortical S | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CareFusion sent an "Urgent: Field Safety Corrective Action Notification" letter dated October 5, 2011 to all affected consignees. The letter described the product, the problem and urgent actions to be taken. Consignees were advised to immediately provide their customers with the attached customer notification letter. Consignees were instructed to complete and return the Verification Form enclosed in the notice. For questions on this matter call 1-800-356-0007 ext. 5129. | Contact the recalling firm for information | 80 | 80 units | Worldwide Distribution - USA (nationwide) including the states of: AR, CA, FL, GA, HI, KY, MI, MO, NY, TX and WA and the countries of: CANADA, CHILE, CHINA, COATIA, DENMARK, FRANCE, GERMANY, INDIA, ISRAEL, KOREA, NORWAY, PORTUGAL, SINGAPORE, SPAIN, SWITZERLAND, TAIWAN, RUSSIA and UNITED KINGDOM. | Terminated | May 21, 2012 | 116 | 167 | Computer
Z-0927-2012 | 60989 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis and AXIOM Artis zee modular angiographic systemsProduct Usage: Angiographic x-ray systems | AXIOM Artis and AXIOM Artis zee modular angiographic systems | 2 | January 27, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential misalignment in the connection of the remote foot switch and the stationary unit. This potential malfunction may occur if either the wireless foot switch or the stationary unit is equipped with improper firmware installed in the bluetooth interface performing the communication between foot switch and system. This is not a general issue but sporadic as a vendor unsystematical | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Siemens sent an Urgent Field Safety Notice letter on January 16, 2016, to all affected customers The letter identified the product, problem, potential risk, and actions to be taken. Customers were instructed to immediately pass this information to all those who need to be aware and instruct personnel accordingly. Siemens ask that if the device/equipment is no longer in the customers possession to forward this safety notice to the new owner of the device/equipment. |  | 45 | 45 | Nationwide Distribution (USA) including the states of: CA, CT, FL, GA, HI, IL, IN, KY, ME, MD, MI, MS, MO, NH, NJ, NY, NC, OH, OR, TN, TX, UT, VA, WI and Puerto Rico. | Terminated | January 31, 2014 | 735 | 45 | Computer
Z-0924-2012 | 60847 | fluorometer, for clinical use | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂŽ ECiQ Immunodiagnostic System, Catalog No. 1922814 --- COMMON/USUAL NAME: VITROSÂŽ ECiQ System --- Ortho-Clinical Diagnostics, 100 Indigo Creek Drive, Rochester, NY 14626 --- All systems that have Software Version 3.8 or below installed and on which the VITROSÂŽ Total B-hCG II Reagent Pack (human chorionic gonadotropin) is processed.Product Usage:For in vitro diagnostic use only. To perform random access, batch and STAT immunodiagnostic assays on human fluid specimens | VITROS ECiQ Immunodiagnostic System | 2 | January 27, 2012 | 2012 | Ortho-Clinical Diagnostics | Software defect; Ortho Clinical Diagnostics received complaints of observed results from diluted samples that were reported as | DESIGN: Software Design | Ortho Clinical Diagnostics (OCD) sent an URGENT Product Correction Notification letter dated August 24, 2011 via FedEx overnight letter to all US consignees with VITROSÂŽ ECi/ECiQ Systems. Foreign Affiliates were also notified of the problem on August 24, 2011. The letters provided information regarding the affected product, the problem, the actions to be followed prior to reporting results from their laboratory and a Question and Answer Section. Customers were instructed to 1) Post the enclosed Procedure for Reviewing VITROSÂŽ Total B-hCG II Results near the VITROSÂŽ ECi/ECiQ System to advise operators of this issue and assist them in reviewing results from diluted samples; 2) Users should consult with their Laboratory Medical Director and request a physician to resolve any concerns users may have regarding previously reported results from Lots 0484 and above of VITROSÂŽ Total B-hCG II Reagent Packs; 3) Complete and return the attached Confirmation of Receipt form upon receipt of this notification; and 4) Forward the information in this notification if you have distributed this product outside of your facility. For additional questions contact OCD Customer Technical Services at 1-800-421-331. |  | 1294 | Domestic: 906, Foreign: 1294 | Worldwide Distribution - USA (nationwide) and including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open |  |  | 2407 | Computer
Z-0935-2012 | 60842 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | FloWire Doppler Guide wires.Volcano Corporation,Rancho Cardova, Ca.The Volcano FloWire Doppler guide wire is a single use guide wire intended for use in all blood vessels including both coronary and peripheral arteries to measure blood flow velocities during diagnostics angiography and/or interventional procedures. | Volcano FloWire Doppler Guidewire | 2 | February 01, 2012 | 2012 | Volcano Corporation | Reports of no signal displayed on the monitor when FloWire Doppler Guide wires were connected to the Combomap. | PRODUCTION CONTROLS: Process Control | Volcano Corporation sent recall letters dated October 6, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and fax the attached Customer Reply Letter so that an RMA could be issued for the return of the unused affected product, and any replacement or credit could be arranged. Customers were instructed to return the Customer Reply letter even if they no longer have the affected product in their inventory. For questions regarding the use of the product. customers were instructed to contact their Sales Representative.For questions regarding this recall call 916-281-2790. | Contact the recalling firm for information | 14 | 14 in US; 42 to EU and 15 to Japan | Worldwide Distribution - USA including MI, AZ and FL and the countries of Japan and Europe | Terminated | February 01, 2012 | 0 | 14 | Computer
Z-0942-2012 | 52807 | computer, diagnostic, pre-programmed, single-function | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Vigilance II Monitor Model: VIG2/VIG2E Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69,00.71, 00.73.Product Usage:The Vigilance II CCO/SvO2/CEDV Monitor is indicated for use in patients requiring monitoring of hemodynamic parameters, including cardiac output, oximetry and right ventricular, ejection, fraction, and end diastolic volume. | Vigilance II Monitor | 2 | February 02, 2012 | 2012 | Edwards Lifesciences, Llc | Edwards Lifesciences has made the decision to initiate a field corrective action for the Edwards Vigilance II Model: Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, 00.71, 00.73 due to the potential of inaccurate Injectate Cardiac Output (ICO) measurements. | DESIGN: Software Design | Edwards initiated two Urgent Product Notification - Action Required letters to all affected customers. The letters identified the affected product, problem, and actions to be taken. The letters instructed users to immediately discontinue use of in-line temperature probes for Intermittent Cardiac Output (ICO) measurements on the Vigilance II instrument and recommends use of the bath probe only (model 9850A) to measure the temperature of the injectate solution for ICO calculations, until software update have been completed. Customers were instructed to complete and return a software upgrade form to Edwards as record of completion. For questions call Edwards Technical Support department at 800-822-9837. |  | 6934 | 6934 instruments | Worldwide Distribution - Nationwide Distribution (USA) - and the countries of: Argentina, Australia/ New Zealand, Brazil, Canada, China, Columbia, Ecuador, Europe, Indonesia, Japan, Korea, Macao, Malaysia, Mexico, Middle East, Peru, Puerto Rico, Singapore, Thailand, Taiwan and Venezuela | Terminated | February 08, 2012 | 6 | 6934 | Computer
Z-0936-2012 | 60936 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD Software release: PBL 6.1.1Product Code 708037Product Usage:The Philips MultiDiagnost Eleva system is intended to be used as a multifunctional/universal system. General R/F, Fluoroscopy, Radiography and Angiography can be performed along with more specialized interventional applications. This includes the following general areas: Digestive system, Skeletal system, Urinary system, Reproductive system, Respiratory system, Circulatory system, and Various: Arthrograms, Myelograms, Facet joint injections, Discography, Sialography. | Philips MultiDiagnost Eleva FD (MDFD) | 2 | February 02, 2012 | 2012 | Philips Healthcare Inc. | The skin dose and skin dose rate values are calculated using an incorrect point of reference. Therefore, the AKR as displayed and stored in patient records is too low, although the dose as received by the patient is not affected. | DESIGN: Software Design | Philips Healthcare issued on January 9, 2012 a mandatory Field Safety Notice to all affected customers. The letter identified the affected product, problem description, hazard involved and actions that should be taken by the customer/user. The letter contains important information for the continued safe and proper use of the equipment. Customers were instructed to review the following information with all members of there staff and retain a copy with the equipment "Instruction for Use". For information or support concerning this notice contact your local Philips representative 866-767-2822. | Contact the recalling firm for information | 28 | 28 units (US) | Worldwide Distribution - Nationwide Distribution (USA) including the states of AR, AZ, CA, CO, FL, IA, KS, LA, MI, MO, NC, NY, OH, PA, TX, VA, VT, WA and WV | Open |  |  | 28 | Computer
Z-0941-2012 | 61010 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DRAD-3000E: Radrex-i X-Ray is a general purpose X-ray System. Labeling reads in part "TOSHIBA***DIGITAL RADIOGRAPHY SYSTEM RADREX-i DRAD-3000E Systems Data No. MSDXR0144EA***" | DRAD3000E Radrexi Xray System | 2 | February 02, 2012 | 2012 | Toshiba American Medical Systems Inc | It has been found that some of the starter cables used for x-ray tube high speed rotation are not sufficiently durable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recall notification titled "URGENT: MEDICAL DEVICE CORRECTION", was sent on 12/29/11 to customers via certified mail. The letter informed the customer of an issue with the starter cable used for high speed rotation on the Radrex X-Ray System. The letter also informs customers of the corrective measures: a replacement cable will be installed on your system to prevent this issue from occurring. An 800 number was provided for any questions regarding the letter. | Contact the recalling firm for information | 39 | 39 | Class II Recall - Nationwide Distribution (USA) - Including the states of VA, IL, CA, PA, TN, OH, LA, TX, OK, AR, AZ, MS, WV, TN, WI, NJ, MN, NY, IA, FL, and Washington D.C. | Open |  |  | 39 | Computer
Z-0944-2012 | 60836 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra PAC1 Mod 2 Cables used with the Integra MPM, CAM01 Monitors | MPM, CAM01 PAC 1 | 3 | February 03, 2012 | 2012 | Integra LifeSciences Corp | Some of the connector cables associated with Integra's intracerebral pressure (ICP) monitors may cause temporary interference with the ICP waveform on the monitor display. |  | Integra sent a Urgent Product Recall Notification letter dated October 20, 2011, via FedEx, email, and fax to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to examine their inventory at your earliest convenience. If you have an effected PAC1 cable, please complete the attached Recall Acknowledgement Form and return it to Integra Customer Service per the instructions on the form. An Integra Sales Representative will contact you to assist in the return and replacement of new cables at no cost. In the meanwhile, you may continue to use the cables until you have been contacted and replacement product has been provided. Distortion of the ICP wave form can be stopped or prevented by keeping the PACI catheter receptacle out of human contact.Please contact Integra Customer Service at 1-800-654-2873. |  | 221 | 221 Units | Worldwide Distribution-- USA (nationwide ) . | Terminated | June 18, 2013 | 501 | 221 | Computer
Z-0946-2012 | 60977 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance Big Bore Model #728243 and #728244, Philips Medical Systems.Computed Tomography X-Ray System intended to produce of x-ray transmission data taken at different angles and planes. | Brilliance Big Bore | 2 | February 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The Block and Multi Leaf Collimator (MLC) shape is not propagated to the beam that is modified by the Oppose operation. The contours can appear with the incorrect color in the Beam's Eye View Viewer, resulting in organs being incorrectly placed or incorrectly shaped. | DESIGN: Software Design | On 12/13/2011, the firm sent Field Safety Notices, entitled URGENT - Medical Device Correction, to their customers. The letter identified the affected product, the problem, the hazard involved, and the actions to be taken by customers and the firm. Customers were to follow the instructions provided in order to address or avoid the issues. Philips will be installing a software update to address the identified issues free of charge. If customers need further information or support in regards to these issues, they are to contact their local Philips representative or local Philips Healthcare office. In North America and Canada, customers can contact the Customer Care Solutions Center at 1-800-722-9377. |  | 142 | 142 Units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, FL, IA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NH, NM, NY, OH, OR, PA, PR, TN, TX, VA, & WA and countries of AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, DOMINICAN REPUBLIC, FRANCE, GABON, GERMANY, HUNGARY, ISRAEL, ITALY, JORDAN, LUXEMBOURG, MALI, MEXICO, NETHERLANDS, OMAN, SLOVAKIA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UKRAINE, UNITED ARAB EMIRATES, & UNITED KINGDOM. | Terminated | February 22, 2013 | 385 | 142 | Computer
Z-0945-2012 | 61012 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Clinac Accelerators with Mark, Millennium MLC 120 (Multileaf Collimator) and BrainLAB m3 MCL, Reference/FSCA CP-05591, Model Number: H53, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA Intended use: The Varian Millennium 120 Collimator (MLC) is an x-ray collimator designed to be mounted on a C or EX Series Varian Clinac radiation therapy linear accelerator and is intended to shape the x-ray field perimeter. Field shape can either be static (fixed) or dynamic. Dynamic field shapes are controlled as a function of either Clinac dose fraction or gantry angle. The intended use is to assist the radiation oncologist in the delivery of radiation to defined target volumes while sparing surrounding normal tissue and critical organs from excess radiation. | Millennium MLC 120 (Multileaf Collimator) and BrainLAB m3 MCL | 2 | February 03, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Failure to retract or park the Varian MLC leaves prior to use of the demountable BrainLAB m3 could result in the Varian MLC leaves interfering with the treatment field defined by the BrainLAB m3 without an interlock being asserted. | DESIGN: Process Design | The firm, Varian Medical Systems, Inc., sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated November 10, 2011 to all affected users. The letter describes the product, description of problem, details, recommended actions and Varian's actions. Varian Service will contact customers to arrange installation, once modifications are available.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter also included special instructions for customers outside the USA and Canada. The customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems once they have read the document.If you require further clarification, feel fee to contact your local Varian Customer Support District or Regional Manager at: USA and Canada: 1-888-VARIANS (1-888-827-4265); Europe: +41 41 749 8844; or Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com; Latin America: soporte.al@varian.com or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 205 | 205 units***12-12-12 AMENDED TO: 213 units*** | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Algeria, Austria, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Slovenia, South Korea, Spain, Sri Lanka, Sweden, Taiwan, Turkey, United Kingdom and Venezuela. | Open |  |  | 205 | Computer
Z-0973-2012 | 60867 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano pcFM Software Kit ;Manufactured by: Volcano Corporation, 2870 Kilgore Road, Rancho Cardova, CA 95624.The volcano pcFM Software Kit is a research only off line tool intended for review and printout of physiological measurements recorded by ComboMap System. It is not intended for analysis or manipulation of the original measurement data. | pcFM Software Kit | 2 | February 07, 2012 | 2012 | Volcano Corporation | Internal review of validation of pcMF Software kit found that instructions directs user to remove the rear panel, causing the potential for electric shock. | DESIGN: Software Design (Manufacturing Process) | The firm, Volcano Corporation, sent an "IMPORTANT PRODUCT INFORMATION" letter dated June 17, 2010 to all customers. The letter describes the product, problem and actions to be taken. The customers were instructed to return the pcFM Software kit to: Volcano Corporation at 2870 Kilgore Road, Rancho Cordova, CA 95670 and contact (800) 228-4728 for instructions on returning the product. The customers were also instructed to complete and return the attached Customer Reply Letter form to Volcano even if they no longer use the pcFM Software Kit via fax to Volcano Customer Support at (916) 638-8812.If you have any questions, contact your Sales Representative or Customer Service at (800) 228-4728. | Contact the recalling firm for information | 8 | 8 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, TX, and WI; and countries of: The Netherlands, Sweden and United Kingdom. | Terminated | February 07, 2012 | 0 | 8 | Computer
Z-0970-2012 | 60721 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Hemotherm 48ÂŽC Dual Reservoir Heater/Cooler System, Model Number: 400 and Model Number: 400MR. The 48 o ( Degree) C Hemotherm Dual Reservoir Heater/Cooler I provides a flow of temperature-controlled water to a heat exchanger. The device, in conjunction with a heat exchanger, is used to warm/cool the fluid that enters and exits a heat exchanger. The conductive heat/cool transfer is accomplished through any type of blood oxygenator/heat exchanger unit. The operator of the unit regulates the desired fluid temperature. Frequent observation of the actual blood/fluid temperature and body temperature is required. The Hemotherm does not contact the patient of the fluids; the water being circulated by the Hemotherm is separated from the fluid by the heat exchanger. To active temperatures above 42 Degrees C, the user must follow explicit instructions in to the owner's manual in order to activate the higher setting. Without following the explicit instructions, the device can only operate within the normal maximum temperature range of: 42 degrees C. The Hemotherm can be used with adult and pediatric patients for extracorporeal circulation of the patient's limb, entire body, or a chemo medication lavage of the abdominal area. The Hemotherm 48ÂŽC device system is used to lower/raise the blood/fluid temperature and/or maintain blood/fluid temperature, as required, through conductive heat transfer. The conductive heat transfer is accomplished through any type of Blood Oxygenator/Heat Exchanger unit. This device is intended to hyperthermic treatment of AIDS and cancer patients and associated research. The device can be set to reach an upper temperature of: 48ÂŽC . | HyperHypothermia System | 2 | February 07, 2012 | 2012 | Cincinnati Sub-Zero Products Inc | The suspect hyper-hypothermia device systems sold to physicians for research and/or the treatment of AIDS and Cancer patients, were not appropriately cleared by the FDA prior to distribution. | DESIGN: Software Design | CSZ sent an "URGENT: FIELD ACTION NOTICE" letter dated January 20, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to become familiar with the amended manual page and to get approvals from the Institutional Review Board before further use of the device. Additionally, a response form was enclosed with the letter for customers to complete and return via fax to (513)772-9119. Contact CSZ Service Department at 1-800-989-7373 for questions regarding this notice. | No consumer action necessary | 20 | A total of 20 devices were distributed in the U.S. & 3 units were distributed internationally. | Nationwide Distribution-including the District of Columbia and states of CA, GA, MA, MD, MI, NJ, NY, OH, and UT. | Terminated | February 26, 2013 | 385 | 36 | Computer
Z-0985-2012 | 60898 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva IITV (MD-EII) with Software Release: PBL 6.1.1Product Code 708036Product Usage:The Philips MultiDiagnost Eleva system is intended to be used as a multifunctional/universal system. General R/F, Fluoroscopy, Radiography and Angiography can be performed along with more specialized interventional applications. This includes the following general areas: Digestive system, Skeletal system, Urinary system, Reproductive system, Respiratory system, Circulatory system, and Various: Arthrograms, Myelograms, Facet joint injections, Discography, Sialography. | Philips MultiDiagnost Eleva IITV | 2 | February 08, 2012 | 2012 | Philips Healthcare Inc. | Incorrect Skin Dose calculation. AKR (Air Kerma Rate) displayed and stored in patient records is too low | DESIGN: Software Design | Philips Healthcare issued a mandatory Urgent Field Safety Noticedated October 3, 2011 to all affected customers. The letter identifies the affected product, problem description, hazard involved, action to be taken by customer/user and actions planned by Philips to correct the problem. The letter provide important information for the continued safe and proper use of the equipment. The letter instructs customers to review the following information with all members of their staff who need to be aware of the contents of this communication and retain a copy with the equipment Instruction for Use. For further information or support concerning this issue contact your local Philips representative at 866-767-2822. |  | 10 | 10 units (US) | Worldwide Distribution - USA (nationwide) including the states of CA, IL, OH, SD, TX, WY and the countries of: Australia, Austria, Canada, Egypt, Germany, mexico, Turkey and UK | Open |  |  | 10 | Computer
Z-1004-2012 | 60943 | system, immunomagnetic, circulating cancer cell, enumeration | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CellTracks Analyzer II.A semi-automated fluorescence microscope used to enumerate fluorescently labeled cells that are immuno-magnetically selected and aligned. The system is for in-vitro diagnostic use when used in tandem with specimen preparation equipment and reagents that are legally marketed for in vitro diagnostic use with this device. | CellTracks Analyzer II | 2 | February 13, 2012 | 2012 | Veridex, LLC | Control cell images may i intermix with Patient cell images while changing between the Control review screen and Patient review screen on the Cell tracks Analyzer II. | DESIGN: Software Design | Veridex, LLC sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated August 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides procedure sheets with step by step directions for customers to follow. Additionally, a Confirmation of Receipt form was enclosed for customers to complete and return to the firm. Contact Customer Technical Services at 1-877-837-4339 for questions regarding this notice. |  | 37 | 37 Units | Worldwide Distribution. | Open |  |  | 37 | Computer
Z-1036-2012 | 60946 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion 100-BGM TeleHealth Monitor, Model No. 09377512, Product Code/Part No. 81842478, Bayer HealthCare LLC 555 White Plains Road, Tarrytown, NY 10591-5097 USA --- Manufactured in Japan for Bayer HealthCare LLC -- Common Name: Viterion100 TeleHealth Monitor -- Classification Name: Monitor, Physiological, Patient (Without Arrhythmia Detection or Alarms) [ref. 21CFR 870.2300]Product Usage:The Viterion 100 TeleHealth Monitor and related Data Server are physiological monitoring systems. The systems collect, accumulate and periodically transmit patient vital signs and other physiological data from a patient who may be remote from the healthcare practitioner to the practitioner, and provides communication between the patient and the practitioner. The physiological monitoring instrument, selected by the healthcare practitioner, operates in conjunction with the Patient Terminal located in the patients home, a hospital room, nursing home, or other healthcare facility. The Patient Terminal connects to the server via a conventional or cellular telephone line. The system is not a real-time monitoring system, and there is no requirement that a health care professional review the information at any specific interval or provide immediate feedback to patients. Rather, the system is intended to aid the physicians longer term management of patients. Patients are expected to take the same short term actions in response to physiological data as they ordinarily would in the absence of the Tele Health Monitor. | Viterion 100 TeleHealth Monitor | 2 | February 15, 2012 | 2012 | Viterion TeleHealthcare Llc | Software defect; Specifically, blood glucose (BG) results from the Roche Aviva meter that are tagged by the user with either the meter hypoglycemic or general "asterisk" flags (*) will not be automatically transmitted through the Viterion 100-BGM monitor containing software version 3.2.0 or displayed to the health care provider through the data server. | DESIGN: Software Design | On May 12, 2011, Bayer Healthcare (aka Viterion) telephoned their consignees regarding the issue with the Viterion 100-BGM and Roche ACCU-CHEK Aviva Flagged Data and followed-up with a letter (dated May 27, 2011) via e-mail. The letter notified customers of the issues and provided information on how to work around the problems. The letter provided the following recommendations: (1) The information concerning the general asterisk (*) flag be communicated within their organization as an amendment to the current patient Safety Alert Notice that was distributed and posted in April 2011. (2) For patients who need to use the Aviva Hypo and/or the asterisk flag, please assign Manual Data Entry for blood glucose on the V100-BGM. The letter informs customers that Bayer has completed the software development to rectify the aforementioned data transmission issue and will upgrade all affected products. Bayer request that customers confirm their receipt of this memo in writing and for any questions concerning this situation, contact Thuy Nguyen, Manager Quality Assurance and Regulatory Affairs at 914-333-6821. |  | 3905 | 3905 units domestically | (USA) Nationwide Distribution | Open |  |  | 3905 | Computer
Z-1034-2012 | 61111 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE; System, X-Ray, Tomography, ComputedLast four of Serial Number:2005, 2008, 2010, 2023, 2026, 2034, 2036, 2038, 2048, 2049, 2051, 2052, 2053, 2061, 2063, 2067, 2070, 2071, 2072, 2077, 2080, 2092, 2115, 2116, 2123, 2131, 2133, 2139, 2143, 2148, 2151, 2158, 2161, 2162, 2165, 2167, 2168, 2169, 2170, 2174, 2208, 2211, 2212, 2219, 2220, 2222, 2241, 2253, 2257, 2265, 2275, 2283, 2284, 2291, 2294, 2308, 2311, 2316, 2324, 2328, 2332, 2333, 2006, 2007, 2008, 2011, 2014, 2015, 2016, 2017, 2018, 2019, 2020, 2021, 2022, 2025, 2027, 2030, 2031, 2032, 2033, 2035, 2002, 2004.Aquilion One is a multislice CT system that supports whole body scanning. The system supports simultaneous data aquisition of 320 slices using the Selectable Slice-thickness Multi-row Detector (SSMD). In addition, the high speed rotation mechanism and the reconstruction unit of the system allow quick image acquisition to further improve throughput in CT examinations. Aquilion ONE, Toshiba's top of the line X-ray CT system, is provided with a 7.5 MHU large capacity X-ray tube unit and detector and permits scanning at the high speed per rotation. | TSX301A: AQUILION ONE; System, XRay, Tomography, Computed | 2 | February 15, 2012 | 2012 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that their is a potential software issue for the TSX-301A Aquilon One and the TSX-301B Aquilon One. | DESIGN: Software Design | Toshiba sent an Urgent Medical Device Correction letter to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax the attached form to (877) 349-3054. The form can also be sent via email to raffairs@tams.com.Customers with any questions regarding this letter were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 82 | 82 units total | Nationwide Distribution - USA (nationwide) and Puerto Rico | Terminated | April 08, 2013 | 418 | 82 | Computer
Z-1056-2012 | 61110 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TruFlight Select PET/CT System equipped with software version 3.5.3.10, Model #459800111661.A diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. | TruFlight Select PET/CT Scanning System | 2 | February 17, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances have been noted in the TruFlight Select PET/CT System. A number of these non-conformances have been corrected to improve the overall quality and functionality of the suspect system units. | DESIGN: Software Design | Philips sent an 'URGENT - Medical Device Correction TruFlight Select PET/CT Serial Numbers: 17002, 17003 letters. The letter was dated December 13, 2011. The letter describes the affected product, the nature of the problem along with the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. In this case, the customers (operator/users) are provided with separate recommended actions. The letter goes on to notify the customers that a Philips representative will be visiting each customer site and installing a software update (3.5.3) in order to address several of the non-conformance issues through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, Customers in all other countries are advised to dial their local Philips Healthcare office. |  | 2 | two (2) units were distributed | Nationwide Distribution -- OH and MD. | Terminated | February 22, 2013 | 371 | 2 | Computer
Z-1064-2012 | 61071 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS. Radiological image processing system. | Siemens syngo Imaging XS | 2 | February 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In September 2011 firm issued two Customer Safety Advisory Notices for syngo Imaging XS. Firm became aware of an unintended behavior when using syngo Imaging XS, version VA70A and VA70B with Client/Server Installations. Loading additional images/series from the Patient Browser using the thumbnails may cause unintended deletion of images/series. In response, firm issued a Customer Safety Advisor | DESIGN: Software Design | Siemens Medical issued an "UPDATE INSTRUCTIONS SY055/11/S/UPDATE PACKAGE VA70B_0211" to all affected customers. The package described the product, problem, and instructions for customers to follow. |  | 128 | 128 | Nationwide Distribution-including the states of CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | September 29, 2012 | 222 | 128 | Computer
Z-1063-2012 | 60058 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology |  | GE SENOGRAPHE 2000D FULL FIELD DIGITAL MAMMOGRAPHY SYSTEM ( SENO 2000D).The Senographe 2000D system generates digital mammographic images that can be used for screening and in the diagnosis of breast cancer. The Senographe 2000D is intended to be used in the same clinical applications as traditional mammographic film/screen systems. | GE SENOGRAPHE 2000D FULL FIELD DIGITAL MAMMOGRAPHY SYSTEM | 2 | February 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurements on GE magnified images provided by Senographe 2000D Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | DESIGN: Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-437-1171 for questions regarding this notice. |  | 1060 | 1060 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA,CT, DC,FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, VA, WV,and WI and countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, PORTUGAL, POLAND, NETHERLANDS, MOROCCO, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, JAM,AICA, ITALY, IRELAND, HUNGARY, HONG KONG, HONDURAS, GREECE, GERMANY, GEORGIA, FRANCE, EGYPT, DOMINICAN REPUBLIC, COLUMBIA, CHINA, CHILE, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Open |  |  | 1060 | Computer
Z-1066-2012 | 61050 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 and R4.6To be used as part of radiation therapy treatment process. | XVI R4.5 Xray Volume Imaging System | 2 | February 21, 2012 | 2012 | Elekta, Inc. | If a user, using an on-line correction protocol, clicks the accept correction button, an approval dialog box appears. If cancel is selected, the table correction values return to their initial values, but the image panes continue to display the changed position. | DESIGN: Software Design | ELEKTA sent an Important Field Safety Notice (20001507042) dated January 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 770-670-2548. |  | 96 | 96 units | Nationwide Distribution including AR, AZ, CA, CT, FL, GA, ID, IL, LA MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI, and Puerto Rico | Terminated | April 17, 2012 | 56 | 96 | Computer
Z-1072-2012 | 60881 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂŽ 3600 Immunodiagnostics Systems, Catalog # 6802783; COMMON/USUAL NAME: VITROSÂŽ 3600 --- NOTE: VITROSÂŽ 3600 that are e-Connected are not affected by this issue.Product Usage:For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROSÂŽ Immunodiagnostic Products Reagents. | VITROS 3600 | 3 | February 21, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints regarding the loss of calibration data following the installation of a development Assay Data Disk (ADD) that was available in the on-screen mailbox on the Main Menu System Status screen of their VITROSÂż 5600 Integrated System. OCD's investigation determined that ADD, Data Release Version (DRV) 9241 or 9251 was electronically transferr | DESIGN: Software Design | Ortho-Clinical Diagnostics (OCD) sent an Important Product Correction Notification letters dated April 13, 2011 via FedEx overnight to customers with an affected VITROSÂŽ 3600 or 5600 System (those not e-Connected). Foreign affiliate consignees were informed of this action via e-mail on April 13, 2011. The letters described the affected products, problem and actions to be followed. The letters states that an OCD Field Engineer will contact customers who has the affected products and perform a Modification (MOD) procedure that will remove ADD DRV 9241 or 9251 from the on-screen mailbox of their system(s). Customers were instructed to complete and return the accompanying Confirmation of Receipt Form and fax to1-888-557-3759 or 1-585-453-4110. The letter asks that distributors forward a copy of the enclosed customer communication to all customers who were shipped an affected VITROSÂŽ 3600 or 5600 System. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. | Contact the recalling firm for information | 96 | Domestic - 15 units, foreign - 81 units | Worldwide Distribution - USA (nationwide) including the countries of: Brazil, Canada, Chile, China, England, France, Germany, India, Italy, Japan, Mexico, Singapore, Spain and Venezuela. | Open |  |  | 228 | Computer
Z-1082-2012 | 60890 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga - previously known as Microsoft Amalga Unified Intelligence System (UIS) 2009Product Usage:Microsoft Amalga UIS 2009 enables healthcare providers to create a comprehensive continually updated data ecosystem to aggregate, store, display, and share disparate data either coming from many source systems or directly entered by end users for rapid self-service data exploration and analysis by front line managers, clinicians and ancillary staff. Furthermore, Amalga UIS allows healthcare providers to rapidly create new applications to respond to evolving business and clinical needs. | Microsoft Amalga Unified Intelligence System | 2 | February 22, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Healthcare providers reported that when sharing a single browser session on Amalga Web, patient A's data can be displayed under patient B's name. | DESIGN: Software Design | Microsoft Corporation sent an "Urgent Field Safety Notice" letter dated January 5, 2012 by email to all affected customers. The letter described the product, problem, and actions to be taken by the customers. If you have any questions about this corrective action program, please contact Microsoft Corporation at 425-882-8080. |  | 9 | 9 units | Worldwide Distribution - USA (nationwide) including the states of California, China, Kentucky, North Carolina, Wisconsin and country of Switzerland. | Terminated | December 12, 2012 | 294 | 9 | Computer
Z-1093-2012 | 60932 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 DIAGNOSTIC ULTRASOUND SYSTEM with R1.1.0 Software. Intended for use by a qualified physician for ultrasound evaluation of Fetal; Abdominal; Pediatric; Small Organ; Peripheral Vascular, etc. | LOGIQ S8 DIAGNOSTIC ULTRASOUND SYSTEM | 2 | February 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where an image from a different patient, or an image from the current patient with title information from a different patient maybe noticed. | DESIGN: Software Design | Consignees were sent on 2/7/12 a GE Healthcare "Urgent Medical Device Correction" letter dated February 6, 2012. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Head of General Imagining Ultrasound Department. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were instructed not to press the PRINT key during the specified view in order to avoid seeing the issue. GE will provide a software update at not cost to customers in order to remove the malfunction. Customers are to contact their local GEHC service representative or call 1-800-321-7937 if they have questions or concerns. | No consumer action necessary | 161 | 161 | Worldwide Distribution -- US, including the states of NJ, PA, and WI and countries of UNITED KINGDOM, THAILAND, SOUTH KOREA, SLOVENIA, SINGAPORE, ROMANIA, REPUBLIC OF MOLDOVA, POLAND, MALAYSIA, JAPAN, ITALY, INDIA, GERMANY, FRANCE, EUROPE, CZECH REPUBLIC, AUSTRIA, AUSTRALIA, and ARAB EMIRATES. | Terminated | July 06, 2012 | 134 | 161 | Computer
Z-1098-2012 | 61045 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with 4D Planning procedure option of MultiPlan Treatment Planning System, versions 3.0, 3.1, 3.5.1, 3.5.2 and 3.5.3 with Ray-Tracing dose calculation.Accuray Incorporated, Sunnyvale, CA.Treatment planning and image guided stereotactic radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | CyberKnife Robotic Radiosurgery System | 2 | February 23, 2012 | 2012 | Accuray Incorporated | An anomaly was discovered during internal regression testing, where the dose information is displayed incorrectly during treatment planning during a specific workflow using the optional 4D Ray Tracing dose calculation algorithm. As a result, it is possible that the dose calculation will display a lower dose than the intended dose prescribed for treatment delivery. | DESIGN: Software Design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides a work around to prevent the issue from happening. Contact Accuray Customer Support at 1-877-668-8667for questions regarding this notice. |  | 207 | 207 | Worldwide Distribution-USA (nationwide) and the countries of Korea, Hong Kong, China, Canada, Turkey, India, Russia, Italy, France, Belgium, Poland, Germany, Netherlands, United Kingdom, Spain, Greece, Saudi Arabia, Switzerland, Ukraine, Czech Republic, and Japan. | Terminated | September 28, 2012 | 218 | 207 | Computer
Z-1096-2012 | 60945 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | LATITUDE Patient management System, Model 6488, software version 7.2. It is indicated for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | Latitude Patient Management System | 3 | February 23, 2012 | 2012 | Boston Scientific CRM Corp | Boston Scientific is conducting a recall involving the LATITUDE Patient Management System - Model 6488, software version 7.2 because patients are not able to perform setup. There is no risk to health associated with this issue. | DESIGN: Software Design | Beginning on December 19, 2011, for the patients that have attempted to set-up their communicators, Boston Scientific called the patients and their associated LATITUDE clinics to let them know that they can now complete the set-up. | Contact the recalling firm for information | 1 | 1 | Nationwide Distribution, USA - including the states CA, MA, MI, NC, OK, PA, RI, TX, and WA. | Terminated | October 25, 2012 | 245 | 1 | Computer
Z-1092-2012 | 61121 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR and PRIMUS Linear Accelerators with RT Therapist Workspace, version 4.1. Manufactured by Siemens AG, Kemnath, Germany. The indication is to deliver X-ray, photon and electron radiation for the therapeutic treatment of cancer. | Syngo RT | 2 | February 23, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Software upgrade to installed base to the latest release of the RT Therapist version 4.1.144 which corrects multiple safety issues that could lead to patient mistreatment. | DESIGN: Software Design | Siemens customers with active devices will receive the software update through the distribution of the "Update Instructions" with a release date of January 22, 2012. The software update will be delivered and installed by a qualified Field Service Engineer utilizing a SW kit and instruction distributed by the Update instruction. Consignee notification of the Update Instructions will be delivered by Customer Service or by certified mail. |  | 182 | 182 active devices | Worldwide Distribution - USA, including the states of WI, UT, PA, TN, OH, LA, DE, MI, NJ, FL, NE, WV, MO and the countries of Germany, Spain, Poland, Australia, Malaysia, New Zealand, India, Republic of Korea, Canada, Italy, Egypt, Philippines, France, Belgium, Saudi Arabia, South Africa, Norway, Turkey , Hungary, Russian Federation, P.R. China, Lebanon, Iran, and the United Kingdom. | Terminated | November 07, 2012 | 258 | 182 | Computer
Z-1102-2012 | 60976 | monitor, cardiac (incl. cardiotachometer & rate alarm) | DRT | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series MonitorMindray DS USA, Inc.800 MacArthur Boulevard.Mahwah, NJ 07430The Indications for Use for the V Series include the monitoring of the following human physiological parameters: ECG waveform derived from 3, 5, 6 and 12 lead measurements Heart Rate Pulse Oximetry (Sp02) ST Segment Analysis Arrhythmia Detection Non Invasive Blood Pressure (NIBP) Invasive Blood Pressure (IBP) Cardiac Output (CO) Respiratory Gasses Respiration Rate Temperature The V Series Monitor has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices | V Series Monitor | 2 | February 24, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | A system database corruption may occur that causes the V Series Monitor to reset. | PRODUCTION CONTROLS: Process Control | The firm, Mindray, sent an "URGENT PRODUCT CORRECTIVE ACTION" letter dated October 14, 2011 on October 17, 2011 via certified mail, return receipt requested, signature required to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to contact their Mindray Service Representative at 1-800-288-2121, to arrange for software upgrade of the V Series Monitor(s) in their facility.If you have any questions, contact Manager, Regulatory Affairs and Clinical Affairs at 201-995-8407 or email: d.arpino@mindray.com. |  | 542 | 542 units US ; 153 units International | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, IN, MO, NE, PR, and SD; and countries of: Australia, Chile, Columbia, United Kingdom, and Venezula. | Terminated | June 11, 2013 | 473 | 542 | Computer
Z-1057-2012 | 60992 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Powerheart G3 Automated External Defibrillators (Non-Wearable) and include the following models: Powerheart 9300A, 9300E, 9300P, 9390A, 9390E; CardioVive 92532, 92533; Nihon-Kohden 9200G, 9231; and GE Responder 2019198, 2023440 AEDs.The Powerheart AED G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation, if the victim is breathing, the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs, the device will charge automatically and be able to deliver therapy. The device is intended to be used by personnel who have been trained it is operation. The user should be qualified by training in basic life support or other physician-authorized medical response. | Powerheart AED, CardioVive AED, GE Responder AED, Nihon Kohden CardioLife AED | 1 | February 27, 2012 | 2012 | Cardiac Science Corporation | The affected Automated External Defibrillators (AEDs) contains a circuit board manufactured with a component that may fail unexpectedly due to a supplier manufacturing defect. If the component were to fail during a rescue attempt, the AED may not deliver defibrillation therapy. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cardiac Science sent an "URGENT-Voluntary Medical Device Recall" letter dated January 13, 2012 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to immediately return the device to the manufacturing location for repair. The affected devices will receive a hardware correction, and the same serial number device returned to the consignee in most cases.Customers in the US can call 1-888-402-2484 and customers outside the US can call 1-425-402-2482 or email at aed210@cardiacscience.com to arrange delivery of the shipping materials.Customers can also visit www.cardiacscience.com/aed210 for more information about AEDs affected by this recall. |  | 4877 | 11,019 units (6142 units in the US and 4877 units international) | Worldwide distribution: USA (nationwide) and Puerto Rico; and to the following countries: Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Malaysia, Maldives, Malta, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Qatar, Romania, Russia, Saint Lucia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, Ukraine, and United Kingdom. | Terminated | September 27, 2013 | 578 | 11019 | Computer
Z-1127-2012 | 61141 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza with software version VA20B_HF04Product Usage:Radiological image processing system | Siemens syngo.plaza with software version VA20B_HF04 | 2 | February 29, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo.plaza VA20B_HF04 with USB dongle and archive license. In syngo.plaza VA20B_HF04, there is a scenario where data in Short-Term Storage (STS) could reaming unprotected and not archived. this can happen when an archive server with a USB dongle-based license is enabled with auto archiving. If the disk fill level is reached which triggers auto deleti | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notice letter dated January 20, 2012 to all affected consignees. The letter identified the affected product, problem, steps to take to avoid this issue, and actions to be taken. Customers were instructed to contact their local Siemens Service to plan for their software update and ensure that this safety advisory is placed in the system's instruction for use. |  | 7 | 7 | Nationwide Distribution (USA) - CA, FL, GA, MO, NY and WA. | Terminated | September 29, 2012 | 213 | 7 | Computer
Z-1125-2012 | 61043 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips Intellispace Portal, Release 4.0.0. For storing, processing and management of electronic information/dataModel: 881001.Indicated for use as a diagnostic device. | Philips Intellispace Portal | 2 | February 29, 2012 | 2012 | Philips Medical Systems Nederlands | There maybe incorrect values in the diffusion/perfusion maps. | DESIGN: Software Design | Philips issued URGENT - Medical Device Correction Notifications, dated January 20, 2012. The letter identified the affected product, all Intellispace Portal with software release 4.0.0 Installed, and gave a description of the problem. Users were provided with the actions to be taken prior to the software upgrade( to version 4.0.1). Users are to follow the instructions provided in order to generate maps with the correct values and to make sure that the setting is correct before using the generated data. Philips will contact customers in order to install the update free of charge. Customers should contact their local Philips representative at 866-767-2822. | Contact the recalling firm for information | 25 | 25 units | Worldwide Distribution -- US, including states of IN and TX, and countries of Australia, Austria, Czech, Denmark, France, Germany, Indonesia, Netherlands, Poland, Spain, and Switzerland. | Open |  |  | 25 | Computer
Z-1135-2012 | 61107 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific, LATITUDEÂŽ Patient Management System Communicator, Models 6476.The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. | LATITUDE Patient Management System Communicator | 2 | March 01, 2012 | 2012 | Boston Scientific CRM Corp | One (1) Model 6476 LATITUDE Communicator is unable to interrogate the patient's pulse generator. Boston Scientific has determined that this is due to a patient receiving an incorrect communicator. Therefore the communicator did not find and was unable to interrogate the assigned patients PG. Additionally when the patient attempted to set-up this incorrect communicator, the system reflected | PRODUCTION CONTROLS: Packaging Process Control | On January 19, 2012 Boston Scientific telephoned the affected patient and their clinic. The patient was notified that the communicator they received will not be able to activate successfully and that their communicator and sensors would be replaced. The clinician was informed of the patient's communicator was sent incorrectly to the other patient who attempted to activate it, the monitored status of the website was incorrect and that a new communicator would be shipped to the patient that will work with their device. | Contact the recalling firm for information | 1 | 1 | Nationwide distribution: UT only. | Terminated | July 05, 2012 | 126 | 1 | Computer
Z-1134-2012 | 61127 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Patient Monitors: The monitor is intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates in health care facilities. The monitor is to be used by trained health care professionals. Physiological Monitoring, Patient Monitor , Arrhythmia detector and alarm (including ST -segment measurement and alarm. | Philips IntelliVue Patient Monitors | 2 | March 01, 2012 | 2012 | Philips Healthcare Inc. | Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor | DESIGN: Software Design | Philips Healthcare sent a Urgent Medical Device Correction letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips has recently received reports that, under certain circumstances, alarms announced at the patient monitor are not announced (either visual or audible) at the central station. This issue affects the following IntelliVue patient monitor models with SW rev. H.xx.xxMP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), Intelligent Display D80 (M8016A), MX600 (865242), MX700 (865241) and MX800 (865240)If this issue occurs, the primary alarm function at the bedside monitor is not affected.Philips is conducting this voluntary correction to upgrade software on affected devices. Please refer to the following pages, which provide instructions for actions to be taken. Follow the Action to be taken by Customer/User section of the instructions. This issue has been reported to the appropriate regulatory agencies.Ensuring that you have the highest quality medical devices, accessories and supporting documentation is our top priority. Your satisfaction with Philips products is very important to us.Review this information with all staff members who interface with the central station and need to be aware of the contents of this communication.The Product Number and Serial Number is contained on the devices product label, located on the front of the device. The SW revision can be accessed via the Revision Screen at the bedside monitor. charge.A Philips Healthcare representative will contact customers with affected devices contact Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this correction to arrange an upgrade of the Intellivue software. | Contact the recalling firm for information | 83534 | 83,534 units | Worldwide Distribution -- USA (nationwide) and the countries of ALBANIA, ANGOLA, ANTILLES, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BYELORUSSIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, GHANA, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LEBANON, LIBYAN, ARAB, JAMAHIRI, LITHUANIA, LUXEMBOURG,MACAU, MALAYSIA, MOROCCO, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NEPAL, NETHERLANDS, ANTILLES, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PALESTINE, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIET NAM, YEMEN, ARAB REP and ZAMBIA | Open |  |  | 83534 | Computer
Z-1153-2012 | 61208 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5/C5 Control panel for 2 mast roller pumps 85, Part No.: 28-95-85, Sorin Group Deutschland GMBH, Lindberghstrasse 25, 80939 Munchen, Germany, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004, USAThe S5 System is used to perform, control and monitor extracorporeal blood circulation during an operation for up to six hours. | S5 Control Panel | 2 | March 05, 2012 | 2012 | Sorin Group USA, Inc. | Touch screen may be unresponsive, inhibiting user input. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | All affected US customers and distributors were notified of the Sorin S5 HLM Touch Screen Field Correction by certified mail on 02/15/12. The letter informed users of the potential for unresponsive S5 HLM touch screens and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to replace affected touch screens on site.Customers and distributors were also provided with a Response Form to confirm they had received, read and understood the Field Correction Notice. They were instructed to fax the completed form to (303) 467-6502.For questions regarding this recall call 303-467-6306. |  | 2 | 2 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Bulgaria, China, Denmark, Germany, Ecuador, Egypt, Finland, France, Georgia, Greece, United Kingdom, Hong Kong, Indonesia, Iraq, Iran, Israel, Italy, Japan, Jordan, Canada, British Columbia, South Korea, Kuwait, La Reunion, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Morocco, Mexico, Monaco, New Zealand, Netherlands, Norway, Oman, Austria, Phillippines, Poland, Portugal, Puerto Rico, Qatar, Serbia, Russia, Saudi Arabia, Sweden, Switzerland, Singapore, Spain, South Africa, Syria, Thailand, Czechoslavakia, Tunisia, Turkey, Hungary, Vietnam, and the United Arab Ermirates | Terminated | November 09, 2012 | 249 | 187 | Computer
Z-1159-2012 | 61251 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio, AXIOM Luminos dRF, AXIOM Luminos Agile/TF, Uroskop OmniaSolid state x-ray imager (flat panel/digital imager) | Siemens Ysio, AXIOM Luminos dRF, AXIOM Luminos Agile/TF, Uroskop Omnia | 2 | March 05, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of unintended behavior when using Ysio, AXIOM dRF, AXIOM Luminos Agile/TF and Uroskop Omnia with RAD Fluoro Uro function using software version VC10A. Due to an error in the image pipeline, images may display dark and not suitable for diagnostic. Due to this software error customers have to repeat x-ray exposure to receive diagnostic images. | DESIGN: Software Design | Siemens sent Update Instruction XP057/11/S by mail on February 15, 2012, which releases software update package VC10D to remedy the described issues.For questions regarding this recall call 610-219-4834. |  | 58 | 58 | Nationwide Distribution including AK, CA, CO, FL, GA, IL, IN, IA, KY, MD, MA, MO, NE, NY, NC, OH, OR, PA, SC, TX, and UT. | Terminated | January 14, 2014 | 680 | 58 | Computer
Z-1175-2012 | 61196 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and True Beam STx versions 1.0 through 1.5, Model number H19; Varian Medical Systems, Palo, Alto, CA 94304. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | TrueBeam Radiotherapy Delivery System | 2 | March 06, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the respiratory gating software of TrueBeam. When importing breath-hold gating protocols, the gating thresholds can be reset to default values, rather than retaining the thresholds established during planning. | DESIGN: Device Design | Urgent Medical Device Correction/Urgent Field Safety Notices were sent on January 30, 2012 by certified mail to US and Canadian consignees.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Varian Sales, Marketing and Service Organizations also were informed via letter the same day. The Field Safety Notice with customer distribution list was e-mailed to Varian International Customer Support Managers for translation (if necessary) to be sent to affected sites with return proof of notification to be sent to Varian. The notice informed users of the recommended actions. Questions were directed to the Varian Help Desk 1-888-827-4265 (US and Canada). | Contact the recalling firm for information | 182 | 182 units ***12-12-12 AMENDED TO: 200 units*** | Nationwide Distribution (USA) and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, New Zealand, Norway, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and the United Kingdom. | Open |  |  | 182 | Computer
Z-1161-2012 | 61198 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, versions 7.3, 8.0, 8.1, 8.2, 8.6, 8.9 and 10.0;Model H48Varian Medical Systems, Palo Alto, CA.Product Usage:The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithim is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | March 06, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has been identified an anomaly with the Eclipse Treatment Planning System where the calculated dose distribution is not updated or invalidated when a recalculation or reload is performed while the compensator or fluence editor is open. | DESIGN: Device Design | Varian sent an URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE letter dated January 30, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, provide recommended user actions, and special instructions for customers outside the USA and Canada. The letter stated that this document will be posted on the customer support website: http://www.MyVarian.com. Customers were instructed to complete the attached Proof of Notification or Receipt Verification Card and return to Varian Medical Systems. For further clarification and questions contact your Varian Customer Support District or Regional Manager. | Contact the recalling firm for information | 25604 | 12,077 units ***AMENDED TO 13,527 units*** | Worldwide Distribution - (USA) Nationwide. | Open |  |  | 13527 | Computer
Z-1186-2012 | 60321 | radioimmunoassay, parathyroid hormone | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH (iPTH) Assay -assay (lot 138)Cat. Nos.100 Test kit 04385983 (129501)500 Test kit 06587575 (129502).For in vitro diagnostic use in the quantitative determination of intact parathyroid hormone (PTH) in EDTA plasma or serum using the ADVIA Centaur systems. This assay is intended to be used to aid in the differential diagnosis of hyperparathyroidism, hypoparathyroidism, or hypercalcemia of malignancy. | ADVIA Centaur Intact PTH (iPTH) Assay | 2 | March 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Poor sensitivity and imprecision on ADVIA CentaurÂż systems; Serum samples low bias; EDTA plasma samples low bias | PRODUCTION CONTROLS: Process Control | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated October 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructed customer to immediately discontinue the use of ADVIA Centaur Intact PTH (iPTH) assay kits that contain intact PTH (iPTH) reagent lot 138. It also recommended that for serum samples analyzed on the ADVIA Centaur CP system, the content of the Urgent Field Safety Notice be discussed with their laboratory director to determine the need to review previous test results, conduct patient follow up, and repeat testing for iPTH results that were less than the upper limit of the reference range, and had a companion calcium result that was elevated or at the upper limit of normal. Additionally, a Completion Notification form was attached for customers to complete and return. Customers should contact their local technical support provider for questions regarding this recall. | Contact the recalling firm for information | 8146 | 100 Test =5768; 500 Test =1778 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Chile, Columbia, Egypt, Finland, France, Hong Kong, India, Israel, Japan, Malaysia, Mexico, Netherlands, New Zealand, Poland, South Korea, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, Uruguay, and Venezuela. | Terminated | March 08, 2012 | 1 | 7546 | Computer
Z-1184-2012 | 61237 | block, beam-shaping, radiation therapy | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Siemens mMLC (ModuLeaf) safety bar and trolley label.Siemens Medical Solutions USA, Inc.Concord, CA.The intended use of the Siemens branded mMLC is an accessory used with the Linear Accelerator systems to deliver x-ray photon and electron radiation for the therapeutic treatment of cancer. This accessory shapes the radiation beam, per the leaf parameters, to minimize the amount of dose to outlying tissue. | Siemens mMLC (ModuLeaf) | 2 | March 07, 2012 | 2012 | Siemens Medical Solutions USA, Inc | This correction is being performed to update the safety knob of the Siemens brand mMLC (ModuLeaf) to conform to tie IEC 60601-1 safety standard. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, SIEMENS, issued a "Customer Information" letter to its customers. The letter described the product, problem and actions to be taken. Siemens developed a new steel knob that complies with the latest standards; Siemens checked all ModuLeaf mMLC trolleys and if necessary refitted with the safety label and also arranged for an inspection and modification of the customers system- a Release of the safety update UI TH 026/11/S to distribute and hardware update commenced on February 7, 2007 by certified mail.If you have any questions, contact the Regulatory Specialist II at 925-602-8175. |  | 46 | 46 active devices | Worldwide distribution: USA (nationwide) including states of: CO, IA, KY, NJ, OH and TX; and countries of: Algeria, Austria, Belgium, Brazil, China, Egypt, France, Germany, Hungary, India, Italy, Japan, Kuwait, Malaysia, Mexico, P.R., Russian Fed., Saudi Arabia, South Africa, Spain, Turkey, and United Kingdom. | Terminated | August 14, 2012 | 160 | 46 | Computer
Z-1185-2012 | 61238 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON G5 ventilators with software version V2.0x.Hamilton Medical AG, SwitzerlandImported and distributed by Hamilton Medical, 4990 Energy Way, Reno, NV 89502.Product Usage:Hamilton-G5 Ventilators are designed for intensive care ventilation of adult and pediatric patients, and optionally infant and neonatal patients. The devices are intended for use in the hospital and institutional environment with trained healthcare professionals. | HamiltonG5 Ventilator | 2 | March 07, 2012 | 2012 | Hamilton Medical Inc | Hamilton-G5 ventilators with software version V2.0x may experience an unintentional change of the displayed ventilation mode under certain conditions. | DESIGN: Software Design | Hamilton Medical sent a Medical Device Safety Alert and Correction Action letter dated February 17, 2011 to all affected customers via a traceable method. The letter identified the affected product, problem and immediate actions required by operators, distributor and manufacturer. The letter states that a Hamilton Medical Account manager will contact customers to schedule a convenient time for the Technical Support Department to install the software update. Questions or concerns may be directed to Technical Support Manager at 800-426-6331, Ext 125. | Contact the recalling firm for information | 456 | 456 devices in US. | (USA) Nationwide Distribution including the states of: CA, CO, FL, GA, MA, MI, MO, NC, NV, NY, PA, TN, TX, VA and WV. | Terminated | October 23, 2013 | 595 | 456 | Computer
Z-1191-2012 | 61021 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 is Microsoft Amalga Unified Intelligence System (UIS), Calculator/data processing module, for clinical use.Microsoft Amalga UIS 2009 enables healthcare providers to create a comprehensive continually updated data ecosystem to aggregate, store, display, and share disparate data either coming from many source systems or directly entered by end users for rapid self-service data exploration and analysis by front line managers, clinicians and ancillary staff. Furthermore, Amalga UIS allows healthcare providers to rapidly create new applications to respond to evolving business and clinical needs. | Microsoft Amalga UIS 2009 | 2 | March 08, 2012 | 2012 | Microsoft Corporation Health Solutions Group | The problem occurs when a user filters the data in Amalga datagrid to exclude specific values. If the field specified in the filter criteria contains NULL values, the filtering will exclude those records as well. | DESIGN: Software Design | Microsoft Corporation Health Solutions Group sent a "Urgent Field Safety Notice" email dated January 19, 2012, to all affected customers. Notification of customers in China is still pending due to the Chinese New Year Holiday. Customers were instructed to :1. If the user notices varying row counts or inconsistent values in a given column and they are using the exclude filter, they may have to look for alternative ways to filter the data that do not use the above mentioned filter conditions.2. To avoid the indicated issue in your implementation there are two steps to be taken:a. Ensure that all parsers are inserting empty strings () rather than NULLwhen ingesting data into the system andb. Identify all Baseview definitions (i.e. the sql view that a baseview isaccessing) that include an outer join and update those definitions toinclude the ISNULL replacement function [e.g. ISNULL( column_name, )]on any columns which may return NULL due to a non-matching row in theouter join tabAs of January 23, 2012 The firm's members of Services Team have spoken by phone with representative of every customer.Customers should call 425-882-8080 for any questions about this recall. | Contact the recalling firm for information | 28 | 28 units | Worldwide Distribution -- USA (nationwide) including the states of AR, CA, KY, HI, FL, MA, MD, MN, NC, NJ, NY, WA, WI and countries China, Switzerland, and United Kingdom. | Terminated | December 12, 2012 | 279 | 28 | Computer
Z-1197-2012 | 61267 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD, Software R5.1.xModel: 708034Philips MultiDiagnost Eleva system is intended to be used as a multifunctional I universal system. General R/F. Fluoroscopy, Radiography and Angiography can be performed along with more specialized interventional applications. | Philips MultiDiagnost Eleva FD | 2 | March 09, 2012 | 2012 | Philips Medical Systems Nederlands | Displayed skin dose rate is calculated using an incorrect point of reference, resulting in too low displayed values. | DESIGN: Software Design | Philips Healthcare issued Urgent Field Safety Notices dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. No action is necessary by the customer/user. A mandatory Field Safety Correction (FCO) Action will be issued to solve the problem. The FCO will consists of a Software update. For further support customers were instructed to contact their Phillips representative.For questions regarding this recall call 978-687-1501. |  | 45 | 45 US; 20 Canada;142 International=207 units | Worldwide Distribution - USA (nationwide) and the following countries:Canada, AustraliaAustriaBelgiumBahrainCzech RepublicDenmarkEstoniaFranceFinlandGermanyHong KongIsraelItalyLatviaLuxemburgMalaysiaMexicoNetherlandsNew ZealandNorwayRussian FedSAUDI ARABIASerbiaSlovakiaSolveniaSouth AfricaSpainSwitzerlandSwedenSyriaThailandTurkeyUnited Kingdom | Terminated | March 26, 2013 | 382 | 59 | Computer
Z-1206-2012 | 61273 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂŽ 3500 Syringe Pumps, Software version 5.0.0, Rx Only. These pumps are indicated for use by healthcare professionals in the delivery of IV solutions, drug solutions, blood or blood products, and enteral feedings in continuous, intermittent, and body weight delivery modes. | Medfusion 3500 V5 Syringe Pumps | 2 | March 14, 2012 | 2012 | Smiths Medical ASD, Inc. | Pumps rendered inoperable at power up (displaying a blank screen and alarm). | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Smiths Medical ASD, Inc. contacted two consignees by telephone in November 2010 and January 2011. Smiths Medical worked with one of the two customers to correct 540 of the 595 devices (91%) while on-site in January 2011. The remaining devices were corrected by hospital personnel.The other customer was also made aware of the issue, root cause and corrective action. This customer reported to Smiths Medical that they understood the issue, but did not want Smiths Medical to perform the correction on their pumps at the time (January 2011). They preferred to send their pumps to Smiths Medical for repair, as and when needed. As of December 2011, The hospital has returned 67 devices to Smiths Medical.For questions regarding this recall call 651-628-7477. |  | 595 | 595 (WI )+ 864 (DC) = 1459 | Nationwide Distribution including WI and DC | Terminated | July 25, 2013 | 498 | 595 | Computer
Z-1216-2012 | 61204 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips EasyDiagnost Eleva DRF, Fluoroscopic x-ray systemSoftware Release 3.0 and 3.1Model: 706034Product Usage: Radiological image processing system (computed radiography x-ray system) | Philips | 2 | March 15, 2012 | 2012 | Philips Healthcare Inc. | Misinterpretation of Mirror Icon-Mirror icon on the bottom of each image could be misinterpreted as aRIGHT lead marker for side indication if the image is mirrored within a PACS system | DESIGN: Software Design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 8, 2012 to all affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter contains important information for continue safe and proper use of the equipment. Philips will contact customers to schedule the installation of the updated software. For further information or support contact your local Philips representative at the Technical Support Line 866-767-2822. | Contact the recalling firm for information | 60 | 60 units | Worldwide Distribution - USA Nationwide | Open |  |  | 60 | Computer
Z-1218-2012 | 60934 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Elecsys 2010 Analyzer Software version 06-07The Elecsys 2010 analyzer is intended for the in vitro quantitative and qualitative analysis of analytes in body fluids. | Elecsys 2010 Analyzer Software version 0607 | 2 | March 16, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Elecsys 2010 customers have reported getting no value and flag 48 results below expected value range after performing automatic dilution for a sample that previously recovered above the measuring range. Incorrect but believable results may occur for the CEA, Rubella IgG or SHBG assays. The other affected assays will generate either no value or a negative result, both of which are not bel | DESIGN: Software Design | Roche sent an Urgent Medical Device Correction letter dated January 31, 2012, via UPS Ground (receipt required) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: use one of the workarounds outlined in this Urgent Medical Device Correction (UMDC) if you are running an automatic dilution on an Elecsys 2010 analyzer with software version 06-07; complete the attached faxback form (4600-00-0112) and fax it to 1-888-912-8460; and file this UMDC for future reference.For questions regarding this recall call 1-800-428-2336. | Contact the recalling firm for information | 2010 | Software version 06-07 was installed on Elecsys 2010 analyzers at 13 accounts. | Worldwide Distribution - USA including NY, GA, DE, SD, WI, ME, OR, IA, AND CA and the countries of Austria, Cyprus, Czech Republic, Estonia, France, Germany,Greece, Hungary, Iceland, Italy, Latvia, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, and Switzerland | Terminated | September 06, 2012 | 174 | 13 | Computer
Z-1217-2012 | 61193 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FastPlan Treatment Planning System, Model HZ3, Varian Medical Systems, Palo Alto, CA. Varian FastPlan Treatment Planning System is a radiation treatment planning system composed of a high resolution UNIX graphics computer, a printer, a digitizing tablet, and radiation treatment planning application software. | FastPlan Treatment Planning System | 2 | March 16, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the CT Localization module of the FastPlan treatment planning system. If the module reports a significant difference in tilt or spin on a slice and the use elects to "Discard" that slice, an incorrect calculation of the isocenter location may result. | DESIGN: Device Design | "Urgent Medical Device Correction/Urgent Field Safety Notification" letters were sent on January 30, 2012 by certified mail to consignees in the US and by Federal Express for Canada. International sites will be notified by Varian International Customer Support Managers, who will translate letters if necessary. The notice provided a description of the issue, the actions that can be taken to avoid or mitigate the issue, and the steps Varian is taking to address the issue. Customer service number for US and Canada is: (888) 827-4265. | Contact the recalling firm for information | 201 | 201 | Worldwide Distribution - Nationwide (USA) and the countries of Canada, China, Germany, Ireland, Spain, Singapore, Georgia, Hong Kong, India, Macau, Malaysia, Russia, Sri Lanka, Taiwan, Japan, Chile, Columbia, Venezuela and Australia. | Terminated | October 21, 2013 | 584 | 201 | Computer
Z-1219-2012 | 61250 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Brilliance Workspace Portal is a software component of the Computed Tomography X-Ray System. The Brilliance Workspace Portal is a software component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance Workspace Portal | 2 | March 16, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect Standardized Uptake Value ("SUV") values are displayed in two instances:1. After loading PET data (with a Slope value different than 1) into the Quick Review (QR) application from a remote device; and2. During the loading the above mentioned type of PET data into the CT viewer application from a remote device. | DESIGN: Software Design | The Philips Healthcare, sent an "URGENT-Medical Device Correction" Field Safety Notice dated February 08, 2012 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of their staff and to retain a copy with the equipment Instruction for Use. The letter notifies the customers that a Field Service Engineer will be contacting them to schedule a visit to implement the corrective action by installing the software update version 2.6.1.If you need further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 822 | 822 | Worldwide distributions: USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, VA, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WI, and WV; and countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech, Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Poland, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uzbekistan, and Venezuela. | Terminated | February 22, 2013 | 343 | 822 | Computer
Z-1124-2012 | 59944 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova X-ray Imaging Systems.The Digital Fluoroscopic Imaging System is indicated for use in diagnostic and interventional angiographic procedures of human anatomy. It is intended to replace image intensifier fluoroscopic systems in all diagnostic or interventional procedures. This device is not intended for mammography applications. | GE Healthcare Innova | 2 | March 19, 2012 | 2012 | GE Healthcare, LLC | It was discovered that when the Diamentor M4-KDK is installed on an Innova system the dose display is disabled on the certified Innova monitor and enabled on the Diamentor M4-KDK which is not certified as required by FDA regulation (Subpart J). | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will alert affected Innova users with a Diamentor M4-KDK unit of this issue and the workaround to mitigate this issue. GE Healthcare will bring affected Innova units in the field into compliance by enabling the Innova certified display and providing affected customers with an Operators Manual Addendum and Service Manual Addendum which will address this topic.GE Healthcare will correct all affected systems by updating the system's Operator and Service Correction manuals and enabling the innova dose display on the Live and Reference monitors and the DLconsole. The Diamentor M4-KDK can still be used to display additional informative dose values. GE Healthcare will, without charge, remedy the defect or bring the product into compliance; the details of which will be included in a subsequent communication to you or through a GE field engineer site visit. GE Healthcare updated the Innova User Instructions, as part of the recall, to indicate to meet regulatory requirements based on the international safety standards IEC 60601-2-43:2000 and/or 21 CFR 1020.30 (k), use as reference the exposure dose displayed on the DL console, live or reference monitor. If you have any further questions or concerns please call 800-437-1171. | Contact the recalling firm for information | 19 | 19 | USA Nationwide Distribution | Terminated | March 29, 2012 | 10 | 19 | Computer
Z-1241-2012 | 61255 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 780 Analyzer, part number: LH780 System Help 1B1.103272, PN 723585.SW Version: 1B1, 1B2. Serial Number: all.The COULTER LH 780 Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The COULTER LH 780 Hematology Analyzer also provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids. | COULTER LH 780 Analyzer, | 2 | March 21, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Auto Stop and Auto Transmit features for Quality Control become disabled after scanning assay values from Coulter 5C and Retic-C Cell Controls assay sheets using the 2D Barcode scanner. | DESIGN: Software Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 30, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers who use the Auto-Stop or Auto-Transmit to ensure the settings are enabled in Control Setup after scanning the assay sheets. Additionally, a Response Form was included for customers to complete and return. Contact Beckman Coulter Customer Service at (800) 526-7694 for questions regarding this notice. |  | 8922 | 8,922 units total (4,073 units in US) | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the countries of Algeria Andorra Angola Antigua and Barbuda Argentina Australia Bahrain Bangladesh Belgium Bermuda Brazil Bulgaria Burundi Cameroon Canada Chile China Colombia Croatia Czech Republic Denmark Ecuador Egypt Ethiopia France Germany Greece Guadeloupe Guam Guatemala Honduras Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Jordan Korea, Republic of Kuwait Latvia Lebanon Libyan Arab Jamahiriya Lithuania Macao Malaysia Mayotte Mexico Monaco Morocco Myanmar Namibia Netherlands New Zealand Norway Oman Pakistan Panama Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Saudi Arabia Serbia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Taiwan Thailand Trinidad and Tobago Tunisia Turkey United Arab Emirates United Kingdom Uruguay Venezuela Viet Nam Yemen Zimbabwe | Open |  |  | 8922 | Computer
Z-1270-2012 | 61308 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ System;customers with Siemens branded linear accelerators with a COHERENCE RT Therapist workspace version 2.1a and the 3rd party MOSAIQ Oncology Information System.MOSAIQ Oncology information system is owned and marketed by IMPAC Medical System, Inc. MOSAIQ is distributed by Siemens Medical Solutions USA, Inc, Radiation Oncology under a distributor contract with IMPAC Medical.Distributed by Siemens Healthcare, Concord, Ca.Product Usage: Medical charged particle radiation therapy system. The intended use of the SIEMENS branded PRIMUS, ONCOR and ARTISTE Solution family of linear accelerators systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer.P | MOSAIQ System; | 2 | March 21, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Images used in treatment planning may belong to another patient. Risk of treatment to wrong area. Software update to upgrade Siemens installed base to the latest release of MOSAIQ version 2.00X3 SP5, for RT Therapist version 2.1A | DESIGN: Device Design | Siemens sent a Customer Information letters to all affected customers. The Customer information letters inform customers of the release of MOSAIQ 2.00X3 SP5 which specifically addresses a safety issue that is only found whenMOSAIQ is used with COHERENCE Therapist 2.1 MQ (also called RTT2.1A ). This letter is to be kept with Customer information documentation. Update Instruction TH004/12S are being issued with the update. The Software update is delivered and installed by qualified Field Service Engineers. For questions, contact your responsible SIEMENS Service Engineer or Application Specialist. |  | 44 | 44 affected Siemens Customers | Worldwide Distribution - USA (nationwide) including the states of KY, PA, TN, WI, and P.R. and the countries of Australia, Belgium, Bosnia Herzeg. China, Colombia, Croatia, Czech Republic, Germany, India, Ireland, Japan, Lebanon, Malaysia, Mexico, New Zealand, Philippines, Poland, Russian Federation, South Africa, Saudi Arabia, Thailand, Turkey, and United Kingdom. | Terminated | September 21, 2012 | 184 | 44 | Computer
Z-1268-2012 | 61236 | system, test, vitamin d | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Vitamin D Diluent 2-pack,Reference 10494100in vitro diagnostic use in the quantitative determination of total25 (OH) vitamin D in human serum and plasma (EDTA, lithium-heparin, sodium heparin) using the ADVIA Centaur and ADVIA Centaur XP systemsThe ADVIA Centaur Vitamin D Total (Vit D) assay is for in vitro diagnostic use in the quantitative determination of total 25 (OH) vitamin D in human serum and plasma (EDTA, lithium-heparin, sodium heparin) using the ADVIA Centaur and ADVIA Centaur XP systems. The ADVIA Centaur Vit Assay is intended as an aid in the determination of vitamin D insufficiency. The ADVIA Centaur Vitamin D Diluent 2-pack is used on board the ADVIA Centaur systems to perform automatic and system-performed dilutions fro Vitamin D assay. | ADVIA Centaur Vitamin D Diluent 2pack | 2 | March 21, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Automatic and system-performed dilution calculation factor for Vitamin D dilutions is incorrect and causesdiluted patient samples to under recover by approximately 50%. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Siemens Healthcare Diagnostics, sent an "Urgent Device Recall Notice" on February 16, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue using the automatic and system-performed dilution features for Vitamin D; to discontinue the use of ADVIA Centaur VitD Diluent, 2-pack (REF 10494100); to manually dilute samples with values greater than the assay range of 150 ng/mL (375 nmoI/L), using the ADVIA Centaur VitD Diluent, 1 bottle (REF 10632114) as stated in the Vitamin D assay Instructions For Use (IFU); to review the contents of the recall notice with their Laboratory Director and retest the previous sample results generated using system- performed dilutions, and complete and return the confirmation fax-back form via fax to TECHNICAL SOLUTIONS CENTER at (302) 631-7597.If you have any questions or need additional information, please contact your local Technical Support Provider or Distributor or call (508) 668-5000. | Contact the recalling firm for information | 442 | 442 units | Worldwide distribution: USA (nationwide) and countries of: Columbia, China, Denmark, Egypt, and Hong Kong. | Terminated | October 17, 2013 | 575 | 442 | Computer
Z-1283-2012 | 61365 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Diagnost Eleva DRF, Software release 4.0Fluoroscopic image-intensified Stationary X-Ray System. | Philips Easy Diagnost Eleva DRF | 2 | March 23, 2012 | 2012 | Philips Medical Systems Gmbh, DMC | Underexposed image due to lag in switching betweenauxiliaries | DESIGN: Software Design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated February 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm provided a work around for customers to use until a software update can be installed. Customers should contact their local Philips representative for questions regarding this notice. | Contact the recalling firm for information | 88 | 88 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Germany, Indonesia, Malaysia, Saudi-Arabia, Singapore, and Thailand. | Terminated | May 23, 2013 | 426 | 88 | Computer
Z-1285-2012 | 61363 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH | Brilliance 64 and Ingenuity CT | 2 | March 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368). | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 17, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of the contents in the communication. Customers were asked to retain a copy of the communication with the equipment instruction for Use. For further information or support concerning this issue, customers should contact their local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 7 | 7 units | Worldwide Distribution - USA including OH and the following countries: Australia, China, Belgium, Israel & Sweden. | Terminated | February 22, 2013 | 336 | 7 | Computer
Z-1318-2012 | 60647 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Anesthesia Machines (digital display assembly); A gas machine for anesthesia is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. | Penlon Prima Anaesthesia Machines (SP2 and SP3) with Electronic Flowmeter Display | 2 | March 26, 2012 | 2012 | Penlon, Ltd. | Penlon, Ltd has identified the potential for the premature failure of a component in some Electronic Flowmeter sensor display unit PCB assemblies. | DESIGN: Component Design/Selection | Penlon Limited sent a Technical Bulletin dated November 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. As a precaution a field upgrade part was released that required all existing PCB assemblies to be removed and the updated item installed. The old PCB must be returned to Penlon with the serial number of the affected machine. The required replacement parts were included with the technical bulletin.For further information customers should contact the Technical Support Team at techsupport@penlon.com. |  | 160 | 160 units | Worldwide Distribution - USA (nationwide) and the countries of Peru, Uganda, Kenya, Malaysia, Poland, Saudi Arabia, United Arab Emirates, Philippines, Equador, and Thailand | Terminated | July 03, 2012 | 99 | 160 | Computer
Z-1317-2012 | 61269 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15, Radiation Therapy Treatment Planning System. RaySearch Laboratories AB Sveavagen 25, SE-111 34 Stockholm Sweden. support@raysearchlabs.comDesigned for treatment planning and analysis of radiation therapy. | RayStation | 2 | March 26, 2012 | 2012 | RAYSEARCH LABORATORIES AB | An issue was found with the dose calculation in the RayStation 2.0, that could result in a potential overestimation of the dose. This could lead to the potential under-dosage of a patient. | DESIGN: Device Design | The firm decided to recall and sent Urgent Field Safety Notice Medical Device Correction letters to consignees on 12/06/2011. The letters described the issue, identified the affected product, actions to be taken by the user, a solution, and contact information. Users are to follow the two part workaround, and educate their staff and all users on the workarounds. The letter states that the issue will be resolved in the next version of RayStation. If further information is needed, users are to contact either Freddie Cardel, Director of Customer Support, RaySearch Americas, at 1-877-778-3849 or freddie.cardel@raysearchlabs.com; or Niclas Borglund, Director of Service, RaySearch Laboratories AB, at 46-8-5450-6130 or niclas.borglund@raysearchlabs.com. | Contact the recalling firm for information | 21 | 21 units (16 domestically and 5 internationally) | Worldwide Distribution -- US, including the states of FL, MA, NJ, and NY and the country of the Netherlands. | Terminated | May 11, 2012 | 46 | 21 | Computer
Z-1178-2012 | 61311 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare EasyDiagnost Eleva DRF 4.0Stationary X-Ray System with DRF as an option for digital radiography. | Philips Healthcare EasyDiagnost Eleva DRF 4.0 | 2 | March 27, 2012 | 2012 | Philips Healthcare Inc. | It was discovered in some instances the detector did not reset as quickly as the user prompted a change form one image receptor (auxiliary) to the next, and then back again. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter, dated February 28, 2012, was sent to customers. The letter identified the affected product, a description of the issue, the hazard involved, along with the actions to be taken and planned. Customers were to use one of the three work-a-rounds provided. The issue will be corrected with Field Change Order 70600064, which consists of a Software Update 4.0.2SP Service Pack installed on the system. If further information or support is needed, customers can contact Dominic Siewko at 978-659-7936. | Contact the recalling firm for information | 96 | 96 | Worldwide Distribution - Including USA. | Open |  |  | 96 | Computer
Z-1224-2012 | 60991 | industrial x-ray | RBS |  |  |  |  | Hamamatsu Corporation, Bridgewater NJ 08807 Microfocus X-ray SourceL8601 series, L9121 series (90kV type MICROFOCUS X-RAY SOURCE),L7901 series, and L8031 series (100kV type MICROFOCUS x-ray source | Hamamatsu Microfocus Xray Source | 2 | March 29, 2012 | 2012 | Hamamatsu Corporation | It was discovered that the Micro Focus X-ray Source could be experiencing software related problems that would allow the unit to exceed the maximum power setting that was set by the operator. | DESIGN: Software Design | Hamamtsu Corporation sent a Warning letter dated December 22, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. CDRH approves your CAP under the following conditions:1.Hamamatsu is able to work with your manufacturer customers to assure the revised manual and new label are provided to your customers' end users.2.Hamamatsu informs CDRH if you receive information which indicates this control defect has occurred. Relevant information includes direct communication from customers or through observation of excessive wear on returned x-ray tubes.For questions regarding this recall call 908-231-0960, ext 2300. | No consumer action necessary | 100 | 100 | Worldwide Distribution-- USA (nationwide) and the country of Canada. | Open |  |  | 100 | Computer
Z-1212-2012 | 61028 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MD Eleva system.Multi-functional fluoroscopy, radiography and angiography systems that can perform more specialized interventional applications | Philips Healthcare MD Eleva system | 2 | March 29, 2012 | 2012 | Philips Healthcare Inc. | It was discovered there was a software error where the air karma rate is not displayed correctly on the user interface. | DESIGN: Software Design | A field correction action will be undertaken to correct the software error in affected systems worldwide. Your proposed action to bring the product into compliance with regulatory requirements is to institute a worldwide field correction to replace the software error in affected systems by updating the current software to version R6.1.2.The timeframe for completion of this field action will be 6 months post field action publication. Preventive actions: The software has been corrected moving forward in subsequent releases for these products.For further questions please call (978) 787-1501. | No consumer action necessary | 92 | 92 ( 24 Eleva 11s, and 68 Eleva FDs ) | Worldwide Distribution - USA (nationwide) | Open |  |  | 92 | Computer
Z-1337-2012 | 61367 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal, M/N 881001, IntelliSpace Portal (ISP) is a multimodality (CT, NM, and MR) thin- client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. IntelliSpace Portal (ISP) is a medical software system that allows multiple users to remotely access IntelliSpace Workspace Portal (ISP) from compatible computers on a network. | IntelliSpace Portal, | 2 | March 30, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were logged for IntelliSpace Portal running software version 4.0 during in-house random verification testing. When loading a study to Review, after failing to load it through Add to Running Application the Analysis application is not being closed and different (1st loaded study) patient appears in analysis (CCA). | DESIGN: Software Design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy with the equipment instruction for use. For any further information or support customers should contact their local Philips representative or local Philips Healthcare office. Customers in North America and Canada should contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 90 | 90 units | Worldwide Distribution - USA including AL, AR, FL, GA, IO, IN, MI, NJ, NY, OH, OK, OR, PA, SC, TN, TX & WA and the countries of Australia, Canada, China, Czech Republic, France, Germany, India, Indonesia, Israel, Italy, Poland, Spain, Sweden, Switzerland & United Kingdom | Terminated | February 22, 2013 | 329 | 90 | Computer
Z-1340-2012 | 61373 | tomography, optical coherence | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | RTVue CAM with Corneal Power Measurement, Cornea Power Kit (software) | Cornea Power Kit (software) | 2 | March 30, 2012 | 2012 | Optovue Inc. | A few doctors in the US were inadvertently sent the Cornea Power software version kit for international users instead of the domestic version. | TRAINING: Employee Error | Field Correction letters were sent to all doctors affected by the issue on February 23, 2012. The letter identified the affected product and the problem. The letter also stated that customers may continue to use the RTVue until the correction was made. Once the correction has been made, customers are to sign and return the attached acknowledgement form. If customers had questions, they are to contact Optovue at 1-866-344-8948 or via mail at 45531 Northport Loop W, Fremont, CA 94538. | Contact the recalling firm for information | 8 | 8 | Nationwide Distribution -- CA and WA only | Terminated | April 02, 2012 | 3 | 8 | Computer
Z-1338-2012 | 61391 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital | PMA | Medtronic, Model 8870, Application Software Card is part of the SynchroMedÂŽ II Drug Infusion System. The System includes the SynchroMedÂŽ II implantable drug infusion pump 8637-20, 8637-40, N'Vision clinical programmer 8840 and the Software application card 8870. The SynchroMed II Infusion system contains and administers prescribed drugs to a specific site inside patient body. The SynchroMed II Pump is supplied in 20 ml and 40 ml reservoir sizes. In the United States, the SynchroMed II Pump is specifically approved for (indicated): ÂŽ The chronic epidural/intrathecal infusion of Infumorph (preservative-free morphine sulfate sterile solution) in the treatment of chronic intractable pain. ÂŽ The chronic intrathecal infusion of Lioresal Intrathecal (baclofen injection) in the management of severe spasticity. ÂŽ The chronic intrathecal infusion of Prialt (preservative-free ziconotide sterile solution) for the management of severe chronic pain. ÂŽ The chronic intravascular infusion of floxuridine (FUDR) for the treatment of primary or metastatic cancer. ÂŽ The chronic intravascular infusion of methotrexate for the treatment of primary or metastatic cancer. Outside of the United States, the SynchroMed II Pump is indicated for the chronic infusion of drugs or fluids tested as compatible and listed in the product labeling for the specific geographies. | Medtronic, Model 8870, Application Software Card | 2 | March 30, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has confirmed that an algorithm used in the Model 8870 application card software has resultedin nine (9) occurrences of an incorrectly displayed Schedule to replace the pump by date. It is estimated that there are more than 140,000 SynchroMed II pump implants worldwide. A patient with a pump reaching EOS prior to replacement may experience the return of underlying symptoms and/orwit | DESIGN: Software Design | Medtronic mailed an "Urgent Medical Device Correction " letter addressed to Dear Healthcare Professional on March 12, 2012. The letter described the problem, Background, Clinical Manifestations, Recommendations, and Additional Information.To ensure there is no interruption in therapy, pump replacements should be scheduled before the ERI alarm sounds and the pump reaches its 7-year end of service.Customers were advised to forward the Urgent Medical Device Correction letter to patients whose name appears on their Pump Replacement List.Reimbursement SupportFor assistance with prior authorization, contact Medtronic Reimbursement at the numbers below:Therapy for Severe Spasticity: (800) 292-2903Pain Therapies: (866) 962-9909For questions regarding the letter call 800-328-0810. | Contact the recalling firm for information | 6338 | 6338 | Nationwide Distribution | Open |  |  | 6338 | Computer
Z-1358-2012 | 61295 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex eXpert DC Intraoral Dental X-Ray SystemDevice is to be used as an extra-oral source of X-Rays in Dental Radiography | Gendex eXpert DC Intraoral Dental XRay System | 2 | April 02, 2012 | 2012 | Gendex Dental Systems | Improper hardware was used during installation of the eXpert DC Intraoral X-Ray system units causing drifting of the tubehead assembly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Gendex Dental Systems sent a Urgent Medical Device Correction letter dated December 29, 2011, to both affected customers. The Customers were provided background information and details on how the nonconforming installations of the devices would be corrected. Gendex Dental Systems apologizes for any inconvenience caused by this issue, and request your corporation as we correct this problem. The customers were instructed to complete and return the Acknowledgement Form upon receipt of the recall notification letter and fax to (215) 997-5665.For further questions please call (215) 954-0365. | No consumer action necessary | 52 | 52 | USA including the states of Iowa and Illinois. | Terminated | June 01, 2012 | 60 | 52 | Computer
Z-1359-2012 | 61272 | glucose dehydrogenase, glucose | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch Verio IQ Blood Glucose Monitoring System Kit;One Touch Verio IQ Blood Glucose IQ Blood Glucose Monitoring Starter Kit;One Touch Verio IQ Blood Glucose Monitoring Warranty Kit.Distributed by : LifeScan, Inc.Milpitas, CA 95035for: LifeScan EuropeDivision of Cilag GmBH International6300 ZugSwitzerlandMeter made in China.Intended to be used for the quantitative measurement of glucose in fresh capillary whole blood. | OneTouch Verio IQ Blood Glucose Monitoring System Kit; | 2 | April 02, 2012 | 2012 | Lifescan Inc | An error was discovered in the OneTouch IQ meter software so that the meter turns itself off when a user attempts to view results in the "Results Log" when the log has 256 or a multiple of 256 items to display. | DESIGN: Software Design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1. Identify and sequester all lots of OneTouchÂŽ Verio"IQ Meters they have in inventory.UPC No. 353885007702 LifeScan Part No. 022267 NDC No. 53885026701Note: Only OneTouchÂŽ Verio"IQ Meters are subject to this notice. Pleasecontinue to sell all other OneTouchÂŽ Brand products including OneTouchÂŽ Verio" TestStrips and all OneTouchÂŽ UltraÂŽ Brand Meters, Test Strips and Ancillary Products.2. Communicate this replacement program to their distribution centers and retail stores that have received OneTouchÂŽ Verio"IQ Meters. Request that they return only OneTouchÂŽ Verio"IQ Meters per their normal return procedures.3. Once you received all of the OneTouchÂŽ Verio"IQ Meters to be returned, call Inmar Corp. at 336-631-7744 for a Returned Goods Authorization (RGA) number and product return instructions.When product has been returned, a credit memo for an amount equivalent to the invoiced price of the returned product. No deductions will be allowed.If you have any questions about this action, please contact your LifeScan Sales Representative or call the LifeScan Customer Experience Line at: 1-800-600-7226, Option 1. Customers were advised that they can continue use of their affected product until an updated meter is provided free of charge.For questions regarding this recall call 408-942-5937. | Contact the recalling firm for information | 101882 | 101,841 units. 2-27-12 - this figure is an estimate and will be updated when more accurate number is received. | Worldwide Distribution - USA (nationwide) and Canada | Terminated | February 05, 2013 | 309 | 101882 | Computer
Z-1361-2012 | 61445 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Systems. Brilliance 64 Model# 728231, Ingenuity CT Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH.The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance 64 and Ingenuity CT Computed Tomography XRay Systems | 2 | April 02, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare is performing an upgrade to the field for issues related to Philips Brilliance 64 and Ingenuity CT scanner having Software Version 3.5.1. When the user plans a scan with tilt applied, the scan increment changes from the user defined values (noncontiguous) to a contiguous scan. | DESIGN: Software Design | . | Contact the recalling firm for information | 440 | 440 units | Worldwide Distribution-USA (nationwide) including the states of AK, AR, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MT, ND, NE, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA and WV, and the countries of Australia, Bahamas, Bangladesh, Belgium, Bolivia, Brazil, Chile, China, Colombia, Croatia, Denmark, Finland, France, Georgia, Germany, Great Britain, India, Indonesia, Iraq, Israel, Italy, Japan, Kazakhstan, Korea, Laos, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Russia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 18, 2013 | 442 | 440 | Computer
Z-1365-2012 | 61392 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Version 3.10.00.Planning of radiation therapy. | Monaco | 2 | April 03, 2012 | 2012 | Computerized Medical Systems Inc | An unnecessary dose refresh can occur in certain situations which causes inconsistencies in the Beam Doses and the Total Dose. | DESIGN: Software Design | An "IMPORTANT SAFETY NOTICE" was issued between 3/5-9/12 via e-mail, FAX, or traditional mail. The Notice described the correct functioning of the system, the problem, when and why the issue occurs, the clinical impact, and the workaround. Patch 3.10.01 is now available. Customers can download their patch at the link provided in the Notice. The Notice also provided a list of offices where customers can contact if they have questions or further support is needed. | Discard product or return to place of purchase for a full refund or exchange | 44 | 44 | Worldwide Distribution -- USA, including the states of OH, IN, WI, CA, FL, IL, LA, NJ, MA, NC, and MO and the countries of Canada, Australia, Belgium, China, France, Germany, Japan, The Netherlands, New Zealand, Portugal, Switzerland, Taiwan, and the United Kingdom. | Open |  |  | 44 | Computer
Z-1367-2012 | 61359 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation Computed Tomography X-Ray Systems, M/N 728260, Distributed by: Philips Medical System, Cleveland, OHProduct Usage: The Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a fused (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Extended Brilliance Workstation Computed Tomography XRay Systems | 2 | April 05, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list. | DESIGN: Software Design (Manufacturing Process) | Philips sent an Urgent - Medical Device Correction letters to all affected customers. The letter identified the product, what the problem is and under what circumstances it can occur, important information for the continued safe and proper use of the affected equipment, and actions to be taken to minimize the effect of the problem. Philips instructed their customers to put a copy of the letter with their "Instructions for Use" manual. For question contact your local Philips representative or local philips healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site Id or follow the prompts). |  | 49 | 49 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV & WY. Product was also shipped to the following countries: Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakstan, Kenya, Kuwait, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom & Venezuela. | Terminated | February 22, 2013 | 323 | 1110 | Computer
Z-1370-2012 | 61316 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Hologics SecurView DX workstation, software version 7.3.0 intended for viewing mammography images and other imaging modalities.SecurView DX is intended for selection, display, manipulation, filing and media interchange of multimodality images from a variety of different modality systems. It also interfaces to various image storage and printing devices, using DICOM or similar interface standards. The device used with FDA-cleared monitors may be used by a trained physician for display, manipulation, and interpretation of lossless compressed or non-compressed mammographic images for screening and diagnostic mammography, as well as any other DICOM multimodality image. | Hologics SecurView DX workstation | 2 | April 05, 2012 | 2012 | Hologic, Inc. | SecurView DX with software version 7.3.0 defects when used to view mammography tomosynthesis images | DESIGN: Software Design | Hologic Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides instructions on how to re-review the stored images and correct any annotation defect. A SecurView DX 7.3.0 Confirmation of Notification form was enclosed for customers to complete and return to the firm via fax at 866-652-8674. Contact Hologic's Help Desk at 1-877-371-4372 for questions regarding this notice. | Contact the recalling firm for information | 7 | 7 units | Worldwide Distribution-Distributed in the state of Texas and the countries Belgium, Brazil, and Canada. | Terminated | May 07, 2012 | 32 | 7 | Computer
Z-1382-2012 | 60152 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500, 500V.Intended for use during cardiopulmonary bypass procedures. | CDI 500 Monitor | 2 | April 06, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Replacement of the SBC batteries (system batteries) should be managed as "routine maintenance" instead of as an in-house Service Procedure because some systems may be relying on batteries beyond their useful life. | DESIGN: Labeling Design | On 2/3/12 all consignees received an URGENT MEDICAL DEVICE RECALL letter via Federal Express. The letter described the reason for the recall, identified the affected product population, what to do in the event of a failure, the potential hazard, and correction. Terumo CVS will update the Operator's Manual for the CDI 500 with instructions to replace the battery within its useful life. Users are to review the notice and ensure that all users are made aware of the issue. In addition, the Customer Response Form should be completed and returned. Users can contact their local Terumo Cardiovascular Systems representative or call 1-800-521-2818 with questions or concerns, to check the last time system batteries were replaced, or to schedule maintenance for systems approaching 10 years. |  | 3829 | 3,829 total units distributed | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, WV, WA, and WI and the countries of Japan, Germany, Barbados, Saudi Arabia, Columbia, Uruguay, Ecuador, Trinidad, Costa Rica, Paraguay, Panama, Belgium, Jordan, Libya, Mexico, Brazil, Honduras, Canada, Korea, Dominica Republic, El Salvador, Guatemala, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Taiwan, Philippines, Mexico, Thailand, Vietnam, and South Africa. | Terminated | August 29, 2013 | 510 | 3829 | Computer
Z-1377-2012 | 61084 | lipoprotein, low-density, antigen, antiserum, control | DFC | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | Wako Autokit Lp(a) Calibrator, model code# 998-41491, packaged in 1 ml bottles, 2 bottles/kit, and labeled in part ***Manufactured by: Wako Pure Chemical Industries, Ltd. 1-2, Doshomachi 3-Chome, Chuo-Ku Osaka 540-8605, Japan***Distributed by: Wako Chemicals USA, Inc. /Wako Diagnostics1600 Bellwood Road Richmond, VA 23237***Product Usage: The Lp(a) calibrator is intended to be used with the Wako Autokit Lp(a) test to establish the points of reference that are used in the determination of Lp(a) in serum or plasma. | Wako Autokit Lp(a) Calibrator | 3 | April 06, 2012 | 2012 | Wako Chemicals USA, Inc. | In-vitro diagnostic kit calibrator is defective and could impact patient profile for cardiovascular disease. | DESIGN: Device Design | Wako Diagnostics notified its customers by telephone on January 26, 2012 and sent an Urgent Medical Device Recall letter dated January 30, 2012 to inform them of the product recall. The notifications identified the affected product, problem and actions to be taken. Consignees were requested to: 1) Complete and sign the attached Return Materials Authorization (RMA) Form to receive replacement material and 2) Package the recalled Product to be returned and send to the address listed on the RMA form. For questions call Technical Support Team at 1-877-714-1924. |  | 10 | 10 kits | The product had been sold to 2 customers in the US. | Terminated | May 02, 2012 | 26 | 10 | Computer
Z-1376-2012 | 61436 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OneTouch(R) Ping(R) Insulin Pump glucose management system. The device is intended to deliver insulin therapy. | OneTouch(R) Ping(R) Insulin Pump glucose management system | 2 | April 06, 2012 | 2012 | Animas Corporation | There is an error in the pump software where the OneTouch(R) Ping(R) insulin pump did not allow the user to save changes to the time and date on the pump if those changes were made on February 29, 2012 (Leap Day). Users who did not attempt to change the time or date on their pump on February 29, 2012 are not affected by this issue. | DESIGN: Software Design | The firm initiated their recall on March 1, 2012 and sent Urgent Notification letters dated March 6, 2012 to all consignees by UPS with delivery confirmation. Customers were notified of the issue and the appropriate steps to take in order to mitigate the issue. A Web Letter for users was also posted to the Animas web site. Questions were directed to (855) 230-7577. |  | 91239 | 91,239 | Worldwide Distribution, (USA) including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of Canada and the territory of Puerto Rico. | Terminated | December 24, 2013 | 627 | 91239 | Computer
Z-1389-2012 | 52024 | system/device, pharmacy compounding | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Pharmacy compounding system . It is a twenty-four source compounder for multi-source fluid mixing of both macro and micro ingredient down to 0.2 ml quantities.Product Usage: Pharmacy compounding System for parenteral macro and micro ingredients | ExactaMix 2400(EM2400) Compounder Operating Ststem | 2 | April 09, 2012 | 2012 | Baxa Corporation | Software issue with compounding pump resulting in over delivery of component ingredients during routine operations. | DESIGN: Software Design | Baxa sent a Safety Alert letter dated March 23, 2009 to all affected consignees. The letter identified the affected product, problem and recommended actions with specific instructions related to the issue. Additionally, the firm will be providing a warning label for the affected product as soon as possible and is currently working on a product upgrade to correct the issue. For assistance contact Baxa Technical Support at 800-678-2292. The firm is also providing a 24 hour a day, 7 days a week Technical Support for the affected product. | Contact the recalling firm for information | 1322 | 1322 units | Worldwide Distribution - USA Nationwide including Canada. | Terminated | April 10, 2012 | 1 | 1322 | Computer
Z-1385-2012 | 61364 | source, brachytherapy, radionuclide | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | VariSeed 7.1, software application, part number H62, manufactured by Varian Medical Systems. VariSeed 7.1 is a computer based software application for planning and evaluating prostate brachytherapy procedures. | Variseed 7.1 | 2 | April 10, 2012 | 2012 | Varian Medical Systems, Inc. | Brachytherapy device requires the user to enter radioactive source activity, but software may specify unintended units for activity and lead to overdose in patient treatment. | DESIGN: Labeling Design | Varian Medical Systems notified direct account healthcare practitioners/ medical physicists with "Urgent Medical Device Correction/Urgent Field Safety Notice" letters by mail on 01/19/2012. Varian's advisory described how to mitigate user error and the firm's intentions to revise Instructions for Use. Technical Support phone numbers were provided for US and international customers. |  | 1377 | 1377 units | Worldwide Distribution | Terminated | November 07, 2012 | 211 | 1377 | Computer
Z-1400-2012 | 61370 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry, Multi-parameter Ambulatory Telemetry, Model #TS01. Dist. By: ScottCare Corporation, Cleveland, OH.The TeleSentry device is intended for diagnostic evaluation of patients who experience transient symptoms or asymptomatic events that may suggest cardiac arrhythmia. The device continuously monitors and records the data, automatically records events triggered by an arrhythmia detection algorithm or manually by the patient, and automatically transmits the recorded event activity associated with these symptoms for review by a licensed physician. | TeleSentry, Multiparameter Ambulatory Telemetry | 2 | April 13, 2012 | 2012 | Scottcare Corporation | Certain TeleSentry Devices contain a Bad SD Card. When the SD Card goes bad, the TeleSentry Device cannot read or write data to the SD Card. The TeleSentry Device then stops functioning. | DESIGN: Component Design/Selection | ScottCare sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 19, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected products. A Recall Acknowledgement Form was enclosed for customers to complete and return via fax to (216) 257-6129. Contact the firm at (216) 362-0550 ext. 113 for questions regarding this recall. | Contact the recalling firm for information | 513 | 513 units | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, MI, NJ, NV, NY, OH, PA and TX, and the country of Canada. | Terminated | September 27, 2013 | 532 | 513 | Computer
Z-1445-2012 | 61390 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific LATITUDEÂŽ Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798, Made in Malaysia. The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. | Latitude Patient Management System Communicator | 2 | April 20, 2012 | 2012 | Boston Scientific CRM Corp | Boston Scientific LATITUDEÂż Patient Management System Communicators were shipped to the incorrect patients. |  | The patients were notified via phone starting on 2/24/12 that their communicator will not be able to activate successfully and that their communicator and sensors will be replaced. If the patient could not be reached via phone or voice mail, a Boston Scientific letter dated 2/29/12 was sent to the patient. The letter described the problem, explained What is a LATITUDE Communicator and What should you do. Questions were directed to Patient Services at (866) 484-3268. | Contact the recalling firm for information | 8 | 8 | Nationwide Distribution, including the states of: AL, CA, FL, IL, IN, MO, NJ, NY, OH, OR, TX VA, and WA. | Terminated | August 02, 2012 | 104 | 8 | Computer
Z-1451-2012 | 60774 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sarns Modular Perfusion System 8000 Catalog number 16413 Arterial MonitorThe 8K Arterial Monitor is intended for use with the 8000 Modular Perfusion System to display one average pressure and three average temperatures and also provide two digital timers. The pressure channel can have alert and alarm levels set by the user to indicate overpressure conditions. The alert and alarm level settings are retained by the system and can be recalled after the next power up. Alert and alarm conditions are indicated by flashing LED bar with single or dual audio tones; an alarm will also stop the arterial and cardioplegia pumps. | Sarns Modular Perfusion System 8000 | 2 | April 23, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo received multiple reports of malfunctions of the Cardioplegia and arterial monitors for the Sarns Modular Perfusion System 8000. Upon review of complaint investigation, the malfunctions are caused by or could have been caused by soldering anomalies on the circuit board of the monitor. The compromised solder joints could cause intermittent or loss of display, or could cause monitor reset, po | DESIGN: Process Design | TERUMO sent an URGENT MEDICAL DEVICE RECALL- SAFETY ADVISORY to all consignees on March 21, 2012 via Federal Express. Terumo issued the safety advisory to remind users how to intervene if they experienced intermittent or complete loss of pressure display on the arterial monitor or cartiplegia monitor, or a related pump stop. Customers were asked to review the Medical Device Correction notice, assure that all users are aware of the notice, confirm receipt of the communication by faxing the attached Customer Response Form to 1-800-292-6551.For questions regarding this recall call 1-800-521-2818. |  | 3022 | 3,022 | Worldwide Distribution - USA including AK, AL , AR, AZ , CA, CO , CT, DC, DE , FL, GA, HI, IA , ID , IL, IN , KS, KY , LA, MA , MD, MI, MN , MO, MS , MT, NC , ND, NE, NH, NJ, NM , NY, OH , OR, PA , RI, SC , TN, TX , VA, WA , WI, WV, WY, and Puerto Rico and the countries of PERU, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tujunga, United Arab Emirates, Uruguay, Venezuela, Vietnam, Wooddale, and Yemen. | Terminated | July 11, 2013 | 444 | 3022 | Computer
Z-1454-2012 | 61482 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4The Dimension Vista(R) System is an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista(R) chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric and integrated ion-selective multisensory technology for clinical use. | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 | 2 | April 25, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Siemens sent an Urgent Field Safety Notice dated December 19, 2011 to all affected customers alerting them to Invalid Below Manufacturer Assay Range Flags on Dimension Vista(R) Systems with Software Version 3.4. The UFSN included instructions on how to correct the situation and also included an effectiveness check form to be faxed to the firm at 302-631-8467. On January 16, 2012, the firm issued an additional UFSN alerting consignees to Invalid Above and Below Manufacturer Assay Range Flags. This notice also included directions and a Field Correction Effectiveness Check form to be completed and faxed to 302-631-8467.For technical questions customers were instruced to call 800-441-9250.For questions regarding this recall call 302-631-6299. |  | 1683 | 1,683 | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland & Baltics, France, Germany, Italy, Japan, South Korea, Malaysia, Netherlands, New Zealand, Norway, Portugal, Puerto Rico, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland and United Kingdom. | Open |  |  | 1683 | Computer
Z-1463-2012 | 61472 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Bolus Cord; a remote bolus cord and switch for use with GemStar 7 Therapy and Pain Management Pumps, list 13000, 13150, 13087 and 13088; Hospira, Inc., Lake Forest, IL 60045; List Number 13027; An accessory for the infusion pump that allows bolus requests to be made by a patient or clinician up to 6 feet (1.8 m) away from the pump. | GemStar Bolus Cord | 2 | April 26, 2012 | 2012 | Hospira Inc. | Hospira has received customer reports of bolus delivery failures when using the GemStar bolus cord. Possible failures may be a result of cracked or damaged bolus cords which may result in an open or short circuit of the Bolus Cord. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hospira, sent an "URGENT DEVICE FIELD CORRECTION RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to take care when removing the bolus cord from the base of the GemStar unit, inspect the cord for damage, discontinue using any damaged cords, and to call Hospira at 1-800-241-4002 for replacement. A Reply Form was enclosed for customers to complete and return via fax to 1-866-912-7085 or e-mail to: Hospira2627@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for technical assistance. |  | 87095 | 47,127 cords - U.S. and 39,968 cords - ROW | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nola, Norway, Oman, Philippines, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay and Venezuela. | Open |  |  | 87095 | Computer
Z-1464-2012 | 61528 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Systems with Eleva software version 2.x and stitching optionIntended use: Stationary radiographic system, Radiographic system, digital | Digital Diagnost | 2 | April 26, 2012 | 2012 | Philips Healthcare Inc. | Potential misdiagnosis due to improperly automatically stitched images, improper image ruler handling. | DESIGN: Labeling Design | The firm, Philips Healthcare Imaging, issued an "URGENT-FIELD SAFETY NOTICE" dated March 30, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions carefully and to not use the Philips Image Ruler for measurements. The firm plans to supply an Addendum to the Instructions For Use and a Philips Service Engineer will contact the customer as soon as the Field Action Kit is ready to be implemented.If you need any further information or support concerning this issue, please contact your local Philips representative or call 1-866-767-2822. | Contact the recalling firm for information | 606 | 606 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech. Rep, Denmark, Finland, France, Germany, Guadeloupe, India, Indonesia, Iran, Iraq, Italy, Japan, Malta, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Salkowski, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ver.Arab.Emir. | Open |  |  | 606 | Computer
Z-1465-2012 | 61510 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Anatomical Pathology Software Client Releases-Client Service Releases 5.6, 5.63, 5.64, 5.65, 6.0, 6.04, 6.05, 6.06, 6.1, 6.12 and 6.13Product Usage:Meditech Anatomical Pathology Software records, stores, codes, searches and provides reports for pathology cases. It enables pathology staff and other personnel to access related findings in real time. | MEDITECH Anatomical Pathology Software | 2 | April 26, 2012 | 2012 | Medical Information Technology, Inc. | Meditech Anatomical Pathology software is being recalled because data on a pathology specimen may be inadvertently deleted. | DESIGN: Software Design | MEDITECH sent a notification beginning on January 19, 2012 to customers who are LIVE with MEDITECH Client Server Anatomical Pathology 5.6 and higher. This notification was made via e-mailed task updates that informed customers of the software issues and the code corrections that are available. |  | 280 | 280 | Worldwide Distribution - USA Nationwide and the countries of Bahamas, Canada and United Kingdom. | Open |  |  | 280 | Computer
Z-1475-2012 | 60610 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | UMP Personal Sentry Deluxe Pull String Monitor, Model #91650. The firm on the label is Universal Medical Products, Lincoln, NE. Designed to sound an alarm when a patient exceeds his or her safe range of movement from a wheelchair, chair, or bed. | UMP Personal Sentry Deluxe Pull String Monitor, Model 91650. | 2 | April 27, 2012 | 2012 | Stanley Security Solutions, Inc. | Monitor gives false alarms | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stanley Healthcare Solutions sent an "IMPORTANT RECALL NOTICE" dated December 8, 2011. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to return affected product through Customer Service and to have the appropriate authority sign the Acknowledgement Form. Customers with questions can call UMP Customer Service at 1-800-206-1045. | Contact the recalling firm for information | 592 | 592 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | September 24, 2013 | 515 | 592 | Computer
Z-1474-2012 | 61328 | system,network and communication,physiological monitors | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. Product Usage: The Philips eCareManager System software is intended for use in data collection, storage, and clinical information management with independent bedside devices and ancillary systems that are connected either directly or through networks. The eCareManager system provides patient information and surveillance of monitored patients at the point of care location and at a remote, supplementary care location through wide area networking technology and dedicated phone lines and is solely intended for use in a hospital environment. It is not intended to be used in a home environment. | Philips eCareManager System | 2 | April 27, 2012 | 2012 | Visicu, Inc. | Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File. |  | Philips notified all affected customers via a preliminary email message advising them of the issue. An URGENT - Field Safety Notice follow-up letter was mailed to all customers on February 22, 2012. This letter identified the affected product, described the problem and circumstances in which it can occur; recommend actions the customer should take to prevent risks for patients, and list the actions planned by Philips to correct the problem. For questions contact Philips eICU Customer Support 877-374-2872. |  | 12 | 12 | Nationwide Distribution (USA) including states of MA, MO, WI, AR, AZ, PA, SD, VA, and IN. | Terminated | July 06, 2012 | 70 | 12 | Computer
Z-1479-2012 | 61604 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Smith & Nephew 560P High Definition Camera Control UnitCatalog Number: 72201919The 560P camera system is configured to be connected as a stand-alone system on an endoscopic cart or tower. It is designed for use with the Smith & Nephew 560H 3-CCD High Definition Camera Head ( 560H camera head ). The 560P camera system is designed to be used in conjunction with various peripheral devices, including the 560H camera head, Smith & Nephews 300 watt xenon light source, various flat panel displays, image capture devices, and print and video recording devices. | Smith & Nephew 560P High Definition Camera Control Unit | 2 | May 01, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | May experience intermittent loss of videodue to a defective board component | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smith & Nephew Endoscopy Division sent an "URGENT PRODUCT RECALL 1st NOTIFICATION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory and to complete and return the bottom portion of the letter via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this notice. | Contact the recalling firm for information | 239 | 239 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, and UK Asia-pacific: Australia, China, Hong Kong, India, Japan, Singapore, South Korea, and Thailand, Mexico, Puerto Rico, South Africa, Turkey, and Venezuela. | Open |  |  | 239 | Computer
Z-1486-2012 | 61545 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded linear accelerator with a 550 TxT Treatment Table with a serial number | ARTISTE, ONCOR and PRIMUS Linear Accelerator systems | 2 | May 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | SIEMENS has received reports where a dropping down of the 550 TxT" Treatment Tablewas observed while the table was being moved downwards during operation by means ofthe hand control. | DESIGN: Software Design | Siemens sent Customer Information letters to all affected customers beginning April 3, 2012 via Siemens Field Service Engineers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter included an Update Instructions TH006/12/S - TxT Table Software Update VD00A. SIEMENS has arranged for a modification of the affected system. A SIEMENS Service Representative will ensure that the new software version VD00A is installed on affected system. The letter instructs customer to please include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. |  | 624 | 624 active devices | Worldwide Distribution - USA (nationwide) including P.R and the countries of: Africa, Angola, Argentina, Australia, Austria, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russian Federation, Saudi Arabia, Serbia, South Thailand, Sri Lanka, Spain, Switzerland, Trinidad, Tobago, Turkey, Ukraine and United Kingdom, . | Terminated | March 14, 2013 | 315 | 624 | Computer
Z-1487-2012 | 53340 | system, gastrointestinal motility (electrical) | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Insight rev H software. The InSIGHT Gastrointestinal Motility System is intended for use by gastroenterologists, surgeons, and medically trained personnel for esophageal, biliary, and anorectal studies as an aid in documenting and diagnosing digestive motility disorders. | Insight | 2 | May 07, 2012 | 2012 | Sandhill Scientific, Inc | Software error may produce incorrect esophageal contraction readings. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Sandhill Scientific sent a Urgent Product Recall letter dated February 23, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were informed All Suite revision H, BioVIEW Analysis Software rev 5.5.1 are affected. Replacement of the Suite revision H, BioVIEW Analysis Software rev 5.5.1, will be provided at no cost to you. Sandhill Scientific will pay for shipping/handling charges associated with replacement of this software.We anticipate that by the first week of March, we will be able to provide you with the Software that resolves this problem.For further questions please call toll free 800-468-4556. | Contact the recalling firm for information | 60 | 60 programs | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, FL, IA, IL, MA, MD, MI, MN, NC, ND, NE, NM, NY, OH, OR, PA, SC, SD, TX, UT, VA and WA and the countries of Belgium, Canada, Malaysia, Mexico, New Zealand, South-Korea, Switzerland and United Kingdom. | Terminated | May 07, 2012 | 0 | 60 | Computer
Z-1493-2012 | 61580 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are: ARTISTE MV System; ONCOR Avant Garde ONCOR Impression ONCOR Impression Plus ONCOR Expression. Manufactured by Siemens AG, Medical Solutions, Kemnath, Germany; Registered by Siemens Medical Solutions, Concord, CA. ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer.Product Model with Part Number:ARTISTE MV System, 8139789;ONCOR Avant Garde, 5863472;ONCOR Impression, 5857920;ONCOR Impression, Plus 5857912;ONCOR Expression, 7360204;ONCOR Expression, 7360717.ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC | 2 | May 07, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment (issue 1). Additionally, there may be a potential for collision if the operator does not ensure clearance of the extended flat panel positioner for all gantry rotation angles (issue 2). | DESIGN: Labeling Design | Siemens sent a Safety Advisory Notice with Update to all affected customers. The notice informed the customers of the risk of pressing the Motion Stop button repeatedly during treatment. Work Steps were provided for customers to follow. |  |  |  | Worldwide Distribution - USA including the states of CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV | Terminated | July 25, 2012 | 79 |  | Computer
Z-1495-2012 | 61693 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | True Beam;Versions 1.0, 1.5.12.0 and below, 1.6.9.5 and below.Varian Medical Systems,Palo Alto, CA.Product Usage:TrueBeam Radiotherapy delivery system is intended to provide sterotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | True Beam | 2 | May 08, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified in TrueBeam when using respiratory gating with patients in the prone position. The gating curve is inverted when a prone patient plan is loaded into TrueBeam. | DESIGN: Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice dated April 13, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, details and recommended user actions. Varian will provide a software version that will mitigate the issue described in the letter and Varian Customer Support will notify each affected facility when the software is available. The letter instructs customers to retained a copy of this notice along with the most current product labeling. For future reference, this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. For questions contact Varian Oncology Help Desk at 1-888-827-4265 US and Canada. |  | 207 | 207 devices ***12-12-12 AMENDED TO: 216 UNITS*** | Worldwide Distribution - USA Nationwide and the countries ofAustralia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom | Open |  |  | 207 | Computer
Z-1494-2012 | 61578 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded ONCOR" or PRIMUS" linear accelerator with an OPTIVUE" or MVision" imaging system; Model Numbers: ONCOR Expression 7360717, ONCOR Impression Plus 5857912, ONCOR Avant Garde 5863472, PRIMUS HI4504200;ARTISTE/ONCOR/PRIMUS systems are manufactured by: SIEMENS AG, MEDICAL SOLUTIONS,Roentgenstrasse 19-21, Kemnath, GERMANY 95478 Product Usage:Medical charged-particle radiation therapy system. The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens | 2 | May 08, 2012 | 2012 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Siemens-branded ONCOR or PRIMUS linear accelerator with an OPTIVUE or MVision imaging system that has the potential collision problem of the imaging panel if the panel does not retract before the gantry rotates; this may result in patient injury. | PRODUCTION CONTROLS: Process Control | Siemens sent an Update Instruction containing the Customer Information letters beginning April 4, 2012 delivered by Siemens customer Service Group or by certified mail to all affected customers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter instructs customer to include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. | Contact the recalling firm for information | 73 | 73 units | Worldwide Distribution - USA Nationwide | Terminated | September 17, 2012 | 132 | 73 | Computer
Z-1541-2012 | 53196 | injector, fluid, non-electrically powered | KZE | Nonelectrically powered fluid injector. | General Hospital | General Hospital | 510(k) | Trima Accel Automated Blood Collection System, catalog number 917000000, Caridian BCT, Lakewood, CO 80215 USA.Product Usage:Automated Blood Collection System for use in humans. | Trima Accel | 2 | May 11, 2012 | 2012 | CaridianBCT, Inc. | Labeling clarification and software updates to include fixes for end of run summaries, increasing boot reliability, corrections to the double red blood cell count collection TBV limit, and to be in compliance with the most current FDA guidance for post donation platelet count targets. | DESIGN: Software Design | CaridianBCT sent a Letter notification dated March 2011 to consignees The letter identified the upcoming software updates to version 6.0.1 as well as operations manual changes. For questions contact your CaridianBCT Representative, the CaridianBCT Support Center at +1-877-339-4228 or 1-303-231-4357 or your local CaridianBCT Customer Service office. |  | 2278 | 2278 units | USA Nationwide Distribution | Terminated | July 25, 2012 | 75 | 2278 | Computer
Z-1615-2012 | 61719 | system, environmental control, powered | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04; QC200405 (USA) and QCE200403, v.02; QCE200403, v.03 (UK). Manufacturer Living Independently Group, Inc., New York, NY. QuietCare is intended for use in monitoring the environmental conditions and activity (motion) of an individual living in a senior housing community. | QuietCare | 3 | May 16, 2012 | 2012 | Intel-GE Care Innovations LLC | The QuietCare production system experienced a partial outage on 7/29/11. As a result of the partial outage, some monitoring data about residents' movements was not processed in the production system, and alerts were therefore not sent to caregivers. The problem was identified and corrected and the unprocessed data was recovered and processed. However, because of the time lag in processing, a fe | TRAINING: Employee Error | Initial notification to the affected facilities was done by telephone or email on 7/30/11. This was followed up with Urgent Field Notification letters to the customers on 8/12/11 via e-mail. For questions, customers were instructed to contact Customer Support at 1 (866) 772-8243. |  | 82 | 82 units affected by the outage: Domestic - 80 units, Foreign - 2 units | Worldwide distribution - USA and the United Kingdom. | Terminated | May 21, 2012 | 5 | 82 | Computer
Z-1616-2012 | 61803 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S1000 ultrasound systems running software version 1.0, Siemens Medical Solutions, USA, Inc. Mountain View, CA. The product is indicated for use as an ultrasound imaging system. | Acuson S1000 | 2 | May 16, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | DESIGN: Software Design | A field corrective action was initiated on 04/25/2012 under Siemens update program US019/12/S in which a "Customer Safety Advisory Notice" was sent to all customers notifying them of the problem and the recommended actions to avoid potential risk. A follow-up field corrective action will be initiated in which the affected software will be updated in the field. | Contact the recalling firm for information | 45 | 45 | Worldwide Distribution -- USA, including the states of TX, NJ, GA and the countries of Germany, United Kingdom, Republic of Korea, Australia, Japan, Poland, Israel, Singapore, France and Saudi Arabia. | Terminated | October 19, 2012 | 156 | 45 | Computer
Z-1646-2012 | 61809 | tomographic imager combining emission computed tomography with nuclear magnetic resonance | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MRI System v.3.7.1, 510(k) #K103483, Model #882380. The device is a diagnostic imaging system for fixed installations that combines Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI). The system does not expose the patient to any ionizing radiation. The MRI Subsystem produces cross-sectional images, spectroscopy images and/or spectra in any orientation of the internal structures of the whole body. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient 's body (specific pharmaceuticals are used for whole body, brain, and other organ imaging). The PET and MRI portion of the system can be used either as an integrated system or as a stand-alone MRI or PET system. The MRI subsystem provides data suitable for use in attenuation correction of the PET acquired data. | Ingenuity TF PET/MRI System | 2 | May 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were identified and corrected by Philips Medical for their Ingenuity TF PET/MRI Systems running software version 3.7.1. The firm states that these non-conformances stemmed from software and/or documentation errors, and that software corrections are being performed in order to 'improve the overall quality and functionality of the Ingenuity TF CT/MRI S | DESIGN: Software Design | The firm, Philips Medical, sent an 'URGENT - Medical Device Correction' notification letter dated April 13, 2012 to all affected customers via Certified mail. The letter informs the customers of the firm's upcoming software update. It tells the customers that the software for the Ingenuity TF PET/MR system unit will be upgraded from version 3.7.1 to version 3.7.2 and that the update/upgrade will provide a number of improvements to the overall quality and functionality of the Ingenuity TF PET/MR System. The notification document states that Philips will be installing the system update to version 3.7.2 software through a Field Change Order. This will be implemented free-of-charge and will include complementary PET/MR application training. For further information, call toll-free telephone number which they can call 1-800-722-9377 and follow the menu options to reach a Customer Solutions Engineer. |  | 6 | 8units (2 units in U.S. & 6 units in other countries) | Worldwide distribution: USA (nationwide) including states of: OH and NY; and countries of: China, Finland, Germany, South Korea, Spain, and Switzerland. | Terminated | August 05, 2013 | 440 | 8 | Computer
Z-1644-2012 | 61746 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, Versions 4.30.00-4.64.00.Used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. The system will calculate and display, both on-screen and in hard-copy, either two- or three-dimensional radiation dose distributions within a patient for a given treatment plan set-up. | XiO Radiation Treatment Planning System | 2 | May 22, 2012 | 2012 | Computerized Medical Systems Inc | XiO will underestimate the dose delivered when the closed leaf gap is set to 0mm on XiO and a value other than 0mm is set on the accelerator. | DESIGN: Software Design | Computerized Medical Systems, Inc. sent an Important Safety Notice the week of April 30, 2012, by regional support managers via e-mail, fax, or regular mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. North American customers were instructed to go to www.elekta.com, select the SupportPlus Login at the top of the page, enter their portal, select Downloads/Updates to download the latest patches.For Customers outside North America, go to www.elekta.com,select the SOFTWARE button on the left, Treatment PlanningSoftware, then Contact TPS Support tab, Download TPSSoftware Patches..............................................................Customers were also instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue.For questions regarding this recall call 408-380-8023. | Contact the recalling firm for information | 1652 | 1,652 sites received the versions | Worldwide Distribution - USA (nationwide) and the country of Canada | Open |  |  | 1652 | Computer
Z-1656-2012 | 61754 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heaters Model 333W, 115V - Part #86135, and 230V & 240V - Part #86136, Cardiopulmonary by-pass temperature controllers. The ECMO Heaters are used in conjunction with a heat exchanger to warm the fluid temperature. The device is composed of a heater, a circulating pump, fan, safety high limit, and a microprocessor board. Water is heated and pumped from the unit to the heat exchanger. The device can be set so that the operator can monitor the actual water temperature or, with the use of a 400 series probe, the fluid (external) temperature. This device is intended to provide temperature-controlled warm water flow to a blood heat exchanger in order to maintain a patient's blood temperature. It can be used during extracorporeal circulation procedures. It is intended to use with commercially available heat exchangers. | Cardiopulmonary bypass temperature controllers | 2 | May 24, 2012 | 2012 | Cincinnati Sub-Zero Products Inc | Design error. The probe jack contacts (terminals on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with one another. If this situation occurs, a Service Technician, Operator/User, and/or Patient could be exposed to electric shock during device servicing or use. This situation was ori | DESIGN: Device Design | Cincinnati Sub-Zero (CSZ) issued an 'URGENT Medical Device Field Action Letter to their customers dated January 31, 2012. The letter was mailed out starting on February 3, 2012. The letter includes: a description of the product and problem, describes the affected Serial Number Range subject to recall/correction and instructs the customers to discontinue using the affected serial-numbered units.The customers are instructed to return the product to their Service Department or Biomedical Engineering Dept. for correction and to complete and mail to CSZ the attached 'Field Action Form.' For additional information, the firm can be reached on a toll-free # 1-800-989-7373. | No consumer action necessary | 298 | 298 units (115V=295 and 230V/240V=3) | AL, AR. AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL,K IN,, KS, LA, MA, MD, MI, MN, MO, NC, NE, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI | Terminated | February 26, 2013 | 278 | 298 | Computer
Z-1657-2012 | 61692 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Becton Dickinson Phoenix System Software, catalog # 448035/448046, component of BD automated microbiology system, and used with BD Gram Positive MIC or MIC/ID panels. Product Usage:Automated Microbiology System for the quantitative determination of antimicrobial susceptibility by minimal inhibitory concentration (MIC) of Gram Positive bacteria from pure culture. | BD Phoenix Automated Microbiology System Software for Automated Microbiological System | 2 | May 24, 2012 | 2012 | Becton Dickinson & Co. | BD identified an increase in vancomycin-resistant strains that may exhibit erroneous results on all PMIC or PMIC/ID panels and could lead to ineffective treatment regimen in affected patients. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Becton Dickinson (BD) sent an Urgent Field Correction Action letter dated April 2012 and an Update letter dated April 18, 2012 to all affected customers via e-mail and UPS overnight delivery. The letter identified the affected products, problem and actions to be taken. The letter provided specific workaround instructions for device users pending software upgrades. A BD Service Engineer will contact affected customers to arrange for the installation of the updated software when it becomes available. The letter included an Acknowledgement Form to be completed. For questions contact Technical Services at 800-638-8663. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 3454993 | 3,454,993 total products (U.S.) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT ,DC ,FL , GA, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, AND WY. and the countries of Canada, Australia, Philippines, Thailand, Malaysia, India, New Zealand, Belgium, Singapore, Japan, Taiwan, China, Chile, Colombia, Korea, Brazil and Venezuela. | Terminated | August 08, 2013 | 441 | 3454993 | Computer
Z-1436-2012 | 52013 | medical computers and software | LNX |  |  | General Hospital | 510(k) | Abacus TPN Calculating SoftwareProduct Usage:LIMS Windows-based program for order entry and TPN calculation. | Abacus TPN Calculating Software | 1 | May 25, 2012 | 2012 | Baxa Corporation | Ordering salt-based ingredients on an ion-based template resulted in medication errors when calculating amounts of ingredients required. | DESIGN: Labeling Design | Baxa Corporation sent a "Safety Alert" letter dated January 19, 2009 via FedEx to all affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to ensure that users create either a salt-based template OR an ion-based template for ordering, and that both methods of ordering are not used in one template. Customers were provided a list of safe practices to follow on the enclosed ISMP Safety Alert. Customers were advised upon receipt of the Safety Alert that Baxa Technical Support will review and address any issues with the customer on their configuration setup for calcium and magnesium. For questions contact Baxa Technical Support at 1-800-678-2292. | Contact the recalling firm for information | 667 | 667 units | Worldwide Distribution - US (Nationwide) and the countries of: Canada, China, Australia, Panama and Columbia. | Terminated | September 12, 2012 | 110 | 667 | Computer
Z-1684-2012 | 61808 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatialcompounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A); Model #: 10041461 - ACUSON S2000 Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CAThe ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, | 2 | May 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Siemens initiated a Customer Safety Advisory Notification letter Field correction on April 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers' local Customer Service Engineer will install an update to their system as soon as it becomes available. For questions customers were instructed to contact their local service support person. Customers were also asked to pass the notification leter to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-694-5747. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Austria, Australia, Belgium, Canada, Switzerland, Chile, China, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Israel, India, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Myanmar, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, El Salvador, and Thailand. | Terminated | July 16, 2013 | 412 |  | Computer
Z-1689-2012 | 61779 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Practix Convenio Mobile X-ray SystemModel: 70455The Practix Convenio is a mains independent motor driven mobile x-ray system, to x-ray body parts of patients to create images for medical diagnostic purposes. The Practix Convenio is intended to x-ray in rooms for medical use. | Philips Imaging | 2 | May 30, 2012 | 2012 | Philips Healthcare Inc. | Unexpected Low-level X-ray pulse | DESIGN: Device Design | Philips sent a Urgent Field Safety Notice dated April 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of the communication. Please retain a copy with the equipment instruction for Use. A Philips Service Engineer will contact account as soon as the Field Action Kit is ready to be implemented. Communicate with Philips with regard to this program, please reference Field Change Order 70400049.Further questions please contact your local Philips representative at 1-800-722-9377. |  | 10 | 10 US; 66 internattional | Worldwide Distribution-- USA ( nationwide ) including the states of IA, LA, LA, ND, NY, OH and TX and the countries of Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Germany, Hong Kong, Indonesia, Italy, Malaysia, Mexico, Netherlands, Philippines, Qatar, Russia, Saudi Arabia, Singapore, South Africa,Thailand and Ukraine. | Open |  |  | 76 | Computer
Z-1699-2012 | 61790 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System Software Version 1.6 & BelowProduct Usage:For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Products Slides, VITROS Chemistry Products MicroTip Reagents and VITROS Immunodiagnostic Products Reagents. | VITROS 5600 Integrated System | 2 | June 01, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling VITROS 5600 Integrated System Software Version 1.6 and below because of an anomaly that may inadvertently add a surfactant to a CurveTip position. | DESIGN: Software Design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated April 26, 2012, via FedEx overnight courier to US consignees with VITROSÂŽ 5600 Systems. The letter identified the affected product, problem and actions to be taken. Foreign affiliates were informed by email on April 26, 2012 of the issue and instructed to notify their consignees of the issue and actions. OCD recommend customers to group all DAT assays into a single run as described in the letter, until the next Software Version 2.0 is installed on their VITROSÂŽ 5600 System. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. | No consumer action necessary | 1357 | 1357 units | Worldwide Distribution - US (nationwide) | Open |  |  | 1357 | Computer
Z-1702-2012 | 61673 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 (previously known as Microsoft Amalga Unified Intelligence System), Calculator/data processing module, for clinical use. Microsoft Amalga UIS 2009 enables healthcare providers to create a comprehensive continually updated data ecosystem to aggregate, store, display, and share disparate data either coming from many source systems or directly entered by end users for rapid self-service data exploration and analysis by front line managers, clinicians and ancillary staff. Furthermore, Amalga UIS allows healthcare providers to rapidly create new applications to respond to evolving business and clinical needs. | Microsoft Amalga UIS 2009 | 3 | June 04, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Microsoft Amalga UIS 2009 has an issue with the display of lab results in the Lab Details window of the lab result viewer. The Lab Details view does not freeze/pin the row header when the user scroll horizontally. | DESIGN: Software Design | Microsoft Corporation Health Solutions Group sent an "URGENT FIELD SAFETY NOTICE" dated April 12, 2012 to all affected customers via e-mail. The letter identifies the product, problem, and actions to be taken by the customers. The firm plans to revise its labeling for Amalga to clarify that the row headers will not be visible as the user scrolls horizontally. Customers were asked to complete and return to the firm a response form via fax at 425-706-7329 or e-mail to vlpigin@microsoft.com. |  | 13 | 13 units | Worldwide Distribution-USA (nationwide) including the states of AR, MA, MD, NC, NY, and WA and the country of China. | Terminated | December 12, 2012 | 191 | 13 | Computer
Z-1703-2012 | 61834 | system, environmental control, powered | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Brand Name: QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04 -- Classification Name: System, Environmental Control, Powered. -- Firm on label: Manufacturer Living Independently Group, Inc., A wholly owned subsidiary of General Electric Company, PO Box 1466, New York, NY 10116 USA, 1-866-772-8243 --- The system is a passive behavior-monitoring device. QuietCare uses strategically placed, unobtrusive, passive wireless infrared pyroelectric sensors in the senior's residence to detect their movement. Data from these sensors are sent to a base station that periodically transmits the information to a centralized secure file server through standard phone lines over a toll-free number. Sophisticated algorithms analyze the data, learning each individual's normal patterns of behavior. When the system detects significant deviations from these baselines it automatically alerts designated caregivers. QuietCare also generates alerts of late wake-ups, potential bathroom falls and if household temperatures become dangerously low or high. The behavioral changes that the system tracks can often serve as subtle indicators of the potential emerging health problems. These changes include: nighttime bathroom usage, entering or exiting the bedroom at a different time than usual, change in the number of times the refrigerator door is opened/closed change in the number of times the medicine cabinet door is opened/closed. The intended users of the QuietCare device are the caregivers at the Assisted Living Facilities and Independent Living Facilities. The monitored individuals (seniors) do not actively operate the device and the device function is not dependent upon the senior resident's ability to use the device. --- Class 2 device, 510(k) exempt, Device Listing Number D117156. | QuietCare | 3 | June 05, 2012 | 2012 | Intel-GE Care Innovations LLC | On 7/29/11, the QuietCare production system experienced a partial outage from about 15:00 hours until 22:30. During the engineering investigation into the cause of the service disruption, it was discovered that some residents did not have the wander and night motion features properly activated, and alerts were not being transmitted. Failure to generate alerts in a timely manner could result in a | TRAINING: Employee Error | Intel-GE Care Innovations sent an "URGENT FIELD NOTIFICATION" e-mail dated August 12, 2011 to all affected customers. The notification identifies the product and problem. No action is required by the customers. Contact Customer Service at (866) 772-8243 for questions regarding this notice. | No consumer action necessary | 61 | 61 units were identified as affected | Nationwide Distribution-USA (nationwide) including the states of Arizona, California, Connecticut, Idaho, Illinois, Minnesota, New Jersey, New York, Ohio, Pennsylvania, Texas, and Utah. | Terminated | June 05, 2012 | 0 | 61 | Computer
Z-1705-2012 | 61919 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393.The Xprezzon Bedside Monitor is intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other devices via Flexport interfaces. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. It can also function as a generic display or computer terminal. It is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. It passively displays data generated by Spacelabs parameter modules, Flexport interfaces, and other Spacelabs SDLC based products as waveform and numeric displays, trends and alarms. Key monitored parameters available on this model, when employing the Spacelabs Command Module, consist of ECG, respiration, invasive and noninvasive blood pressure, SpO2, temperature and cardiac output. | Model 91393, Xprezzon Bedside Monitor | 2 | June 05, 2012 | 2012 | Spacelabs Healthcare, Llc | The screw used to connect the external power supply to the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, may damage an internal capacitor. The damage may result in an inability of the monitor to power on, power off, or involuntarily shut down. |  | The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 22, 2012 via return receipt requested, to US customers. On May 29, 2012 a translated customer letter was emailed to all international subsidiaries and distributors of record. The letter describes the product, problem and actions to be taken. The customers were instructed to notify staff of the potential failure and to only tighten the power supply screw with their fingers. In addition, the firm will schedule a time to upgrade the monitors at the customers site.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1898 | 1898 monitors (Domestic: 1188) (International: 710) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Chile, China, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, India, Italy, Kuwait, Malaysia, Netherlands, Panama, Poland, Portugal, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | May 07, 2013 | 336 | 1898 | Computer
Z-1791-2012 | 61866 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories.The Brilliance iCT and iCT SP are used to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | The Brilliance iCT and iCT SP | 2 | June 08, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update (version 3.2.4) to the Brilliance iCT and iCT SP systems currently running software version 3.2.0, 3.2.1 or 3.2.3, addressing a number of non-conformances. A review was conducted of all nonconformances that are fixed by the software upgrade. None of these nonconformances result in unacceptable hazards.The nonconformance documented below had the highest sc | DESIGN: Software Design | Philips sent an Urgent - Medical Device Correction letter dated May 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Customers were instructed to review the Release Document provided them with information about the software update, new feature and enhancements. For North America and Canada customers should contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 241 | 241 | Worldwide Distribution - USA including AL, CA, DC, DE, FL, GA, IA, IN, MA, MD, MN, MS, NC, NY, OH, OR, PA, TN, TX, VT, WA & WV and the countries of Argentina, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, EGYPT, France, Germany, Greece, Hungary, ICELAND, India, Indonesia, Israel, ISRAEL, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, PHILIPPINES, POLAND, Portugal, Qatar, Russia, SINGAPORE, South Korea, SPAIN, SWEDEN, Switzerland, Taiwan, THAILAND, Turkey, UNITED KINGDOM & Uzbekistan. | Terminated | July 25, 2013 | 412 | 241 | Computer
Z-1538-2012 | 61830 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Practix Convenio Mobile X-ray systemWorldwide Distribution -- Nationwide | Philips Healthcare Practix Convenio | 2 | June 13, 2012 | 2012 | Philips Healthcare Inc. | When the system is switched off or when the tube arm is parked in the down position (X-ray generator switch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. .When the system is switched off or when the tube arm is parked in the down position (X-ray generatorswitch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. The failure happened whencomponent supplie | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Philips Healthcare sent a proprosed corrective action letter dated May 3, 2012, to FDA. The proposed corrective action plan is to bring the product into compliance with regulatory requirements by issuing a field corrective action which should be completed by Philips by August 1, 2012.For questions regarding this recall call 978-687-1501. | Contact the recalling firm for information | 11 | 11 - US | Nationwide Distribution | Open |  |  | 11 | Computer
Z-1818-2012 | 61926 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror iconStationary Fluoroscopic system | Philips Healthcare Imaging | 2 | June 15, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Philips Healthcare issued a Field Safety Notice on 5/22/12 advising users of the error. A software update (Rel 2.1.2) will issue to correct the problem. Contact your local Philips representative. 1-800-722-9377.For questions regarding this recall call 978-687-1501. | Contact the recalling firm for information | 752 | 752 units | Worldwide Distribution - USA (nationwide) and the countries of:NationwideForeign: CanadaAFGHANISTANALBANIAALGERIAANGOLAARGENTINAARMENIAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHAMASBAHRAINBANGLADESHBARBADOSBELARUSBELGIUMBENINBHUTANBOLIVIABOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIABURUNDICAMBODIACAMEROONCANARY ISLANDSCHADCHILECHINACOLOMBIACONGOCONGO,DEMOCRATICCOSTA RICACOTE D'IVOIRECROATIACUBACYPRUSCZECH REPUBLICDENMARKDJIBOUTIDOMINICAN REPUBLICECUADOREGYPTEL SALVADOREQUATORIAL GUINEAERITREAESTONIAETHIOPIAfake test ssdFINLANDFRANCEFRENCH GUIANAGABONGAMBIAGEORGIEGERMANYGHANAGREECEGUADELOUPEGUAMGUATEMALAGUINEAGUYANAHONDURASHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAMAICAJAPANJORDANKAZAKSTANKENYAKUWAITKYRGYZSTANLAO PEOPLE'S DEM.REPLATVIALEBANONLIBERIALIBYAN ARAB JAMAHIRILIECHTENSTEINLITHUANIALUXEMBOURGMACAUMACEDONIAMALAWIMALAYSIAMALIMALTAMARTINIQUEMAURITANIAMAURITIUSMETRACOMMEXICOMOLDOVAMONGOLIAMOROCCOMOZAMBIQUENAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORTH KOREANORWAYOMANPAKISTANPALESTINEPANAMAPARAGUAYPERUPHC BESTPHILIPPINESPMCC BU CanadaPMCC BU USAPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIARWANDASAUDI ARABIASENEGALSERBIA&MONTE; NEGROSIERRA LEONESINGAPORESLOVAKIASLOVENIASOMALIASOUTH AFRICASOUTH KOREASPAINSRI LANKASUDANSURINAMESWEDENSWITZERLANDSYRIAN ARAB REPUBLICTAIWANTAJIKISTANTANZANIATHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUnidentified CountryUNITED ARAB EMIRATESUNITED KINGDOM URUGUAYUZBEKISTANVENEZUELAVIET NAMYEMENZAMBIAZIMBABWE | Open |  |  | 752 | Computer
Z-1831-2012 | 61913 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The product is Pathfinder SL, version 1.6.0.3057 software used with the Spacelabs Healthcare Pathfinder SL Holter Analyzer, version 1.6.0.Product Usage:The Holter Analyzer analyzes information collected from ambulatory electrocardiogram recorders worn by patients and used to assist physicians in diagnosis and patient monitoring. In addition, incorrect QTc values are reported after moving the cursor. QTc values are timing events between labeled points of interest in the ECG signal that the physician can mark with the cursor. The software is used to edit a family of events and upon conclusion the software crashes with an error message. | Spacelabs Healthcare Pathfinder SL Holter Analyzer | 2 | June 20, 2012 | 2012 | Spacelabs Healthcare, Llc | The firm received reports of the Pathfinder LS software crashing after editing a batch of Holter Analyzer records. The software will stop with an error message requiring a restart. In addition, QTc values will not update after moving the marker. | DESIGN: Software Design | Spacelabs Healthcare, Llc sent a "MEDICAL DEVICE CORRECTION" letter dated May 11, 2012 via mailed, return receipt requested, to US customers. The customer letter (translated as necessary) was emailed to all international subsidiaries and distributors of record. The letter identified the affected product, problem and actions to be taken. Consignees are advised to turn the power off and then on to restore normal functionality. Updated software will be provided upon completion of validation. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 3 for Technical Support. | Contact the recalling firm for information | 114 | 114 | Worldwide Distribution - US (nationwide) including the states of Nevada and Vermont, and the countries of AUSTRALIA, AUSTRIA, CANADA, CYPRUS, GERMANY, ICELAND, ISRAEL, POLAND, ROMANIA, SOUTH AFRICA, SWITZERLAND, THAILAND, TURKEY and UNITED KINGDOM. | Terminated | October 22, 2012 | 124 | 114 | Computer
Z-1843-2012 | 61946 | cap,cooling (infants) | MXM |  |  | Neurology |  | Olympic Cool-Cap System is intended to provide treatment forneonatal hypoxic-ischemic encephalopathy (HIE). The Cool-Cap is indicated for use in full term infants (more than 36 week gestational age) with clinical evidence of moderate to severe HIE. The device provides selective head cooling with mild systemic hypothermia to prevent or reduce the severity of neurologic injury associated with HIE.It is a selective head cooling system for treatment of neonatal hypoxic-ischemic encephalopathy (HIE) in infants. A cooling water cap is placed on the infant's head and the body is warmed using radiant warmers. The Cool-Cap provides 72 hours of monitored cooling.The device is labeled in part: "***OLYMPIC COOL-CAP Cooling Unit***Model No.***Made in USA***Software in this product is Natus Medical Incorporated***" | Olympic CoolCap System | 2 | June 20, 2012 | 2012 | Natus Medical Incorporated | Olympic Cool-Cap Cooling Module's power supply failed during treatment. The power supply failure can happen without prior indication and results in the unavailability of the system to begin or continue cooling treatment. | DESIGN: Component Design/Selection | Natus Medical Incorporated sent a URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTIONS Olympic Cool-Cap System letter dated May 21, 2012, to all affected customers. Fed-Ex to all U.S. customers, specifically to Biomedical department, NICU, with affected devices. The letter identified the product the problem and the action needed to be taken by the customer.In addition, Natus also sent the Urgent Field Safety Notice-Olympic Cool-Cap System (DOC-007459) dated May 2012 to all US customers and International Distribution Partners who were affected by both Frozen Screen and Power Supply field corrective action.Customers were instructed to: " During all routine checks of infants undergoing treatment, look at the clockdisplayed on the screen to verify that the seconds counter on the clock isadvancing. The time displayed should continue to advance if the system isoperating correctly." If the system experiences a screen-freeze during a cooling session, the Olympic Cool-Cap system can be rebooted and cooling resumed." Ensure that all the personnel in your facility that use the Olympic Cool-Capsystem are advised of this Field Safety Notice (see attached).The U.S. customers are advised of the issue specifically stated in Action #2 in the letter and they are also informed that Natus will contact them within the next 30 to 120 days to schedule the power supply replacement in their affected system(s).Natus requires that the Distribution Partners to notify their customers with copies of the Natus recall letter and the Field Safety Notice (DOC-007459). Specifically these two documents should be addressed to both the Biomedical Department and the Neonatal Intensive Care Unit (NICU).Customers with questions about this recall should call Natus Technical Service at 888-496-2887 or email to TService@natus.com. |  | 181 | 181 units | Worldwide Distribution -- USA (nationwide) including the states of : AK; AR; AZ; CA; CO; DE; FL; GA; HI; ID; IL; IN; KY; MA; MI; MN; MS; ND; NE; NJ; NY; OH; PA; SC; SD; TN; TX; and WA. and the countries of : Argentina; Australia; Austria; Canada; Egypt; Germany; Hong Kong; Indonesia; Iran; Italy; Kuwait; Poland; Russia; Saudi Arabia; Singapore; Slovenia; South Africa; South Korea; Spain; Taiwan; Thailand; Turkey; United Arab Emirates; and United Kingdom. | Terminated | February 21, 2014 | 611 | 181 | Computer
Z-1849-2012 | 61741 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm Imaging System with Software Version 3.1.1, Catalog Numbers: BI-700-00027-120R and BI-700-00027-120. The 0-armÂŽ Imaging System is a mobile x-ray system designed for 2D fluoroscopic and 3D imaging. | OARM Imaging System | 2 | June 21, 2012 | 2012 | Medtronic Navigation, Inc. | Software Version 3.1.1 does not meet navigational accuracy. | DESIGN: Software Design | Medtronic Navigation contacted customers by telephone on 10/12/2010 and with "Urgent Field Safety Notice" follow up letters dated 10/14 or 10/15/2010. Sites were visited by Medtronic Service Representative to complete navigational accuracy testing and either replace or correct the unit. | Contact the recalling firm for information | 3 | 3 units | Nationwide Distribution including the states of CO, IN, and MN. | Terminated | June 21, 2012 | 0 | 3 | Computer
Z-1854-2012 | 61761 | system, automated platelet aggregation | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow System, part #85005-6HProduct Usage:The VerifyNow P2Y12 assay device is a whole blood assay used in the laboratory or point of care setting to measure the level of platelet P2Y12 receptor blockade. The VerifyNow Systems intended for use with human whole blood and verify now test devices. The system should be operated by healthcare professionals trained on use of the system and in accordance with institution and policies and procedures. Accumetrics respresentatives will assist your institution in the installation of the system and operations of the operator. | VERIFYNOWP2Y12 ASSAY | 2 | June 21, 2012 | 2012 | Accumetrics Inc | The recall was initiated by Accumetrics because of a field correction to the VerifyNow System; specifically the results reported when running the VerifyNow P2Y12 Test. Currently the VerifyNow P2Y12 Test reports three values: PRU, BASE and % Inhibition. The change implemented is the elimination of the BASE and % Inhibition results. There will be no change to the reporting of the PRU result whic |  | Accumetrics, Inc. sent an Urgent Field Correction notification letter and an Ackknowledgement - fax back form to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter informed customers that an Accumetrics representative will becontacting them to schedule an appointment to update their VerifyNowinstrument software. The letter stated that Accumetrics realizes that procedures and report format may need to be revised as software changes is implemented, and will provide assistance to help in making this transition. For any questions or concerns, call at 1-800-643-1640 ext. 2; oremail at: support@accumetrics.com. |  | 1167 | 1,167 units total | US Nationwide Distribution | Terminated | November 27, 2012 | 159 | 27242 | Computer
Z-1865-2012 | 62016 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Uroview 2800, Model number A349855. Designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical, and interventional procedures. | OEC Uroview 2800 | 2 | June 22, 2012 | 2012 | GE OEC Medical Systems, Inc | Recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mode" and "digital spot mode", and the system does not reboot. | DESIGN: Software Design | On 11/20/06 two written notifications were sent to consignees addressed to the Facility Administrator and the Radiology Administrator. The notifications advised them that GE Healthcare had identified several intermittent potential safety issues that may occur with the product. Interim solutions were provided for each issue. The letter stated that GEHC OEC would contact customers and remedy the issues free of charge when a permanent solution became available. An Urgent Recall Notice, dated April 13, 2007, was also sent to customers. Customers were to contact the service team at 800-874-7378 for further information, if there were any questions or concerns. | Contact the recalling firm for information | 717 | 717 units | Worldwide Distribution. | Terminated | June 25, 2012 | 3 | 717 | Computer
Z-1896-2012 | 62333 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx monitor/defibrillatorsModels: M3535A, M3536A, M3536M4, M3536M5, and M3536M6Product Usage:The Heartstart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips HeartStart MRx | 2 | June 27, 2012 | 2012 | Philips Healthcare Inc. | HeartStart MRx Unexpected Pads/Paddles ECG Failure |  | Philips issued an Urgent Voluntary Medical Device Correction/Field Safety Notice letter on June 20, 2012. The letter informs customers of the issue, identifies details (including models and software versions applicable) of units affected, gives instructions on actions to be taken until the correction is available, and identifies what action Philips plans to eliminate or remedy the issue. The letter states that a Philips Healthcare representative will contact customers to arrange for the replacement of the Power PCA in all the affected devices.Philips will complete the installation free of charge. For questions or support concerning this issue contact your local Philips representative at 1-800-722-9377. |  | 2914 | 2914 units | Worldwide Distribution - US (nationwide) and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, EL SALVADOR, ETHIOPIA, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTILLES, NICARAGUA, NORWAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Open |  |  | 2914 | Computer
Z-1900-2012 | 61943 | heater, breathing system w/wo controller (not humidifier or nebulizer | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | AirLife Isothermal Breathing Circuit, Infant Respiratory Circuit, Non-Heated; an Rx device for single patient use; CareFusion, 22745 Savi Ranch Parkway, Yorba Linda, CA 92887 USA, Made in MexicoRespiratory breathing circuits are used with a ventilator when mechanical ventilation is administered to a patient. | AirLife Isothermal Breathing Circuit, Infant Respiratory Circuit, NonHeated | 1 | June 29, 2012 | 2012 | Carefusion 2200 Inc | The Y adapter within the breathing circuit may spontaneously crack, causing a leak in the closed ventilation system. This leak could lead to a decrease in the tidal volume delivered to the ventilated patient. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, CareFusion, sent an "URGENT RECALL NOTICE" dated May 29, 2012 to all of the affected customers on the same date via registered return receipt mail. The notice described the product, problem and actions to be taken. The customers were instructed to immediately destroy any affected product in-stock at their facility and to NOT further distribute these products. NOTE: The notice stated "Carefusion is removing from the the market all lot numbers less than lot #0000414257 or begin with the letter 'Y' for all Infant Breathing Circuit product codes listed in the provided attachment".The customers were also instructed to not return the product to Carefusion. Return the product directly to: Owens and Minor Health Care Logistics at 1651 California Street, Suite C, Redlands, CA 92374, Attn: Recall Coordinator, for immediate replacement.NOTE: The notice the stated "Failure to follow these important measures may result in a potential patient safety risk".In addition, the customers were instructed to forward the recall notice to any customers to whom these products were sold, notifying them of the potential risk, advising them to discard or return any CareFusion AirLife Infant Breathing Circuits with the noted product/lot numbers for replacement; and to complete and return a enclosed reply form indicating the amount of inventory in stock, amount destroyed and amount returned. The distributors were requested to complete and return the attached Acknowledgement Form indicating that they have received the recall notice and that they communicated the Recall Notice to all customers who have purchased the affected breathing circuits via Fax to (847) 473-7774 or email to David.Wehrheim@Carefusion.com. They were also instructed to contact CareFusion Channel Management at (866) 802-6320 for further details on return and/or replacement of stock.If you have any questions related to this information, please call (847) 473-7495 or email David.Wehrheim@Caref. | Contact the recalling firm for information |  |  | Worldwide distribution: USA (nationwide) including Puerto Rico and internationally to: Canada, Egypt, France, Germany, Great Britain, Italy, Kuwait, Mexico, Portugal, Saudi Arabia, Spain, Taiwan and United Arab Emirates. | Open |  |  |  | Computer
Z-1942-2012 | 62270 | medical device data system | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips Xcelera Connect, Software R2.1 L 1 SP2, an interface engine for data exchange.Philips Xcelera Connect R2.x is a generic interface and data mapping engine between a Hospital Information System (HIS), Imaging Modalities, Xcelera PACS and Xcelera Cath Lab Manager (CLM). This interface engine simplifies the connection by serving as a central point for data exchange. The data consists only of demographic patient information, schedules, textual information and text reports. | Philips Xcelera Connect | 2 | July 05, 2012 | 2012 | Philips Healthcare Inc. | Xcelera Connect R2.1 L 1 SP2 , incomplete information arriving from unformatted reports interface | DESIGN: Software Design | Philips Healthcare sent an "URGENT-FIELD SAFETY NOTICE" letter dated June 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Philips representative will contact customers when the upgrade becomes available. Contact your local Philips Representative for questions concerning this issue. | No consumer action necessary | 226 | 226 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Chile, Denmark , Estonia, Finland, France, Germany ,Hong Kong, Ireland, Italy, Malaysia, Netherlands, Poland, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Taiwan, United Kingdom, and Utd. Arab. Emir. | Open |  |  | 226 | Computer
Z-1957-2012 | 60341 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Central Control Module for System 1: System 1 Base, 100/120VThe Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 06, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of the Central Control Monitor (CCM) for Terumo Advanced Perfusion System 1 losing partial or full functionality. Symptoms that have been associated with a malfunctioning CCM include: full or partial loss of display, loss of control functions; inability to distinguish various status, alert, alarm conditions, the inability to operate the touch screen, or the CCM | DESIGN: Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE RECALL" notice dated June 22, 2012 and addendum to the Operator's Manual via Federal Express to its customers. The letter described the product, problem and actions to be taken. The notice states to the customers "Please note that we are not recommending that you stop using your Central Control Monitor". The customers were instructed to review the Medical Device Recall notice; assure that all users are aware of this notice; place the Addendum at the beginning of the Central Control Monitor Section (page 4-1) of the Operator's Manual and complete and return the attached Customer Response Form to the fax number indicated on the form. If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours are from Monday-Friday, 8 AM - 6 PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries including: Argentina, Belgium, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates (UAE), and Vietnam. | Terminated | March 06, 2013 | 243 | 1640 | Computer
Z-1950-2012 | 62470 | unit, neonatal phototherapy | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Drager PT 4000 Phototherapy System | Drager PT 4000 Phototherapy System | 2 | July 06, 2012 | 2012 | Draeger Medical Systems, Inc. | A routine internal quality check found that the Drager PT 4000 Phototherapy System is fitted with a 16 Amp rated fuse instead of the specified 1.6 Amp rated fuse. This may mean that the internal fuse would not trip promptly in the event of an internal short-circuit. The possible consequence is the wall outlet circuit breaker would trip first on the device may severely overheat. There have not b | OTHER/UNDETERMINED: Pending | Drager Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to bee taken by the customers. In addition, a Customer Reply Card was included for customers to complete and return. Contact the firm at 800-543-5047 for questions regarding this recall. |  | 65 | 65 | Nationwide Distribution-USA (nationwide) and Puerto Rico. | Terminated | July 01, 2013 | 360 | 65 | Computer
Z-1951-2012 | 61529 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | AEM Disposable Electrodes, Device, Electrosurgical, Cutting and Coagulation and Accessories. AEM Disposable Electrodes used as electrosurgical accessories intended for ablation, removal, resection and coagulation of soft tissue in open, endoscopic and laparoscopic surgical procedures channeling monopolar high-frequency electrical current from compatible electrosurgical generators. The device is model FC03XX series.AEM Disposable Electrodes used as electrosurgical accessories intended for ablation, removal, resection and coagulation of soft tissue in open, endoscopic and laparoscopic surgical procedures channeling monopolar high-frequency electrical current from compatible electrosurgical generators. | AEM Disposable Electrodes | 2 | July 06, 2012 | 2012 | Encision, Inc. | Encision is performing a field correction to their AEM Disposable Electrodes due to two consumer complaints reporting unanticipated alarms from the AEM monitor while using the device. This issue could result delays or complications during surgical procedures. | DESIGN: Device Design | Encision sent an Urgent Medical Device Recall letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to lubricate the contact on the instrument cord with a conductive gel (to be supplied by Encision) prior to making the connection of the cord to the instrument. Instructions for applying the gel were attached. This corrective action was temporary; allowing customers to use their inventory of devices and for new product in the short term, until a design correction can resolve the problem. Customers would be contacted by their Encision Sales Representative to review this information for their specific situation and train their staff on its correction. For questions customers should call 303-339-6917 or Customer Service at 800-998-0986.For questions regarding this recall call 303-339-6917. | Contact the recalling firm for information | 48 | 48 boxes | Worldwide Distribution - USA including CA, TX, GA, NY, NC, OH), and New Zealand. | Terminated | August 07, 2012 | 32 | 48 | Computer
Z-1925-2012 | 62118 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | FLOW-I Anesthesia SystemThe system is intended for use in administrating anesthesia while controlling the entire ventilation for patients with no ability to breathe, as well as in supporting patients with a limited ability to breathe. | FLOWI Anesthesia System | 1 | July 06, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | A technical alarm may be generated on the FLOW-i system when using the MAN/AUTO switch to change ventilation modes from manual to automatic or from automatic to manual if the switch is not fully engaged in the "on" or "off" position but remains in the "in-between position. | DESIGN: Device Design | The firm, Maquet Getinge Group, sent an "URGENT DEVICE FIELD CORRECTION" letter dated February 20, 2012 to its customers. The letter describes the product, problem and action to be taken. The letter state that a Maquet Service Representative will upgrade the customers software in the FLOW-i Anesthesia System to the current revision and also provide the customers with the corresponding user's manual. Note: This upgrade includes other minor improvements. The customers were instructed to sign a document to verify satisfactory completion of the work once the software upgrade was completed.If you have questions or require additional information, please contact your local MAQUET representative or Technical Support at 888-627-8383.A press release was issued on 7/6/2012. | Contact the recalling firm for information | 284 | 284 units | Worldwide distribution: USA (nationwide) state of GA only and countries of: Argentina, Australia, Austria, Belgium, China, Czech Republic, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Serbia, Singapore, South Africa, Spain, Sweden incl Maquet Critical Care, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Open |  |  | 284 | Computer
Z-1944-2012 | 62302 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information Systems interfaced to Elekta Treatment machines.An image enabled electronic medical record system used for oncology workflow management.Manufactured by:Impac Medical Systems100 Mathilda Place, 5th floorSunnyvale, Ca 94086 | MOSAIQ Oncology Information Systems | 2 | July 06, 2012 | 2012 | Impac Medical Systems Inc | Impac Medical became aware of a problem with the recording of the Wedge Motor Units in MOSAIQ when it resumes after a machine terminates/faults, the remaining Wedge MU was being erroneously set to zero. | DESIGN: Software Design | A User Notice began being distributed on June 8, 2012. This will be followed by contact at each site by IMPAC support to determine if the site is affected by this issue, and if so, the upgraded version will be suggested. The Notice identified the affected device, the problem, the clinical impact, along with the interim mitigations and the final resolution. Users are encouraged to monitor daily delivery and recordings of treatments to be sure that the Wedge Motor Units treatments are properly delivered. The letter also recommends that treatment charts be checked weekly and that users not override any parameters in MOSAIQ until the machine setup has been completed. Users are to contact Elekta Support for the upgrade. | Contact the recalling firm for information | 11 | 11 affected sites | Worldwide Distribution -- USA, including the states of VA, CT, GA, NJ and NY and the countries of Belgium, Denmark, France, and the United Kingdom. | Terminated | June 19, 2013 | 348 | 11 | Computer
Z-1971-2012 | 62373 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 220/240V.The Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump malfunctions of the Terumo Advanced Perfusion System 1 that could result in the inability to use the pump and would then require replacement with back-up equipment. Description of malfunctions:" Unplanned pump stop or pause" Changes in pump speedIf it is necessary to replace a roller pump during cardiopu | DESIGN: Device Design | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Customer Response Form to the firm via fax at 1-800-292-6551. Contact the firm at 1-800-521-2818 for questions regarding this recall. |  | 648 | 648 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 01, 2013 | 235 | 1640 | Computer
Z-1968-2012 | 62367 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 100/120VThe Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump occlusion malfunctions that could result in the inability to use the pump and would then require replacement with back-up equipment. Occlusion malfunctions includestiff/difficult to adjust occlusion, loss of occlusion setting, unequal roller-to-roller occlusion and inability to adjust the occlusion.If it is | DESIGN: Device Design | The firm, Terumo, sent an "URGENT MEDICAL DEVICE RECALL" notification dated June 25, 2012 and an update to the Operator's Manual for the Terumo Advanced Perfusion System 1 via Federal Express to its customers. The notification describes the product, problem and actions to be taken. The customers were instructed to review this Medical Device Recall notice; assure that all users are aware of this notice; place the addendum at the beginning of the Roller Pumps Section (page 5.1) of the Operator's Manual, and complete and return the attached Customer Response Form to the fax number/email address indicated on the form. Note: Terumo is continuing its investigation into the root cause of the malfunctions, and may initiate further action pending results of the investigation.If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours: Monday-Friday, 8 AM - 6 PM (ET). | Contact the recalling firm for information | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries including: Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Open |  |  | 1640 | Computer
Z-1960-2012 | 61591 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Advantage CT/PET Fusion for Advantage Windows Workstation with Volume View CT Colonography/Navigator 2 software. (Version 4.0, 4.1,or 4.2)The Advantage CT/PET Fusion clinical application provides easy comparison of three dimensional (3D) anatomical images from Computed Tomography (CT) and functional images from Positron Emission Tomography (PET). | Image Processing System | 2 | July 09, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to incorrect alignment of fused CT and PET images associated with the CT/PET fusion capabilities of your Volume Viewer and CT/PET Fusion application. Volume Viewer allows two studies (one PET and one CT) that are registered in the same spatial domain to be viewed simultaneously in fused mode. The two images should be aligned to th | DESIGN: Software Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides safety instructions for customers to use until a software upgrade is available. Contact the firm at 800-437-1171 for questions regarding this recall. | Contact the recalling firm for information | 70 | 70 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IL, MA, NJ, NY, SC, and TX and the countries of TURKEY, TUNISIA, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, KENYA, JAPAN, ITALY, IRELAND, INDIA, GREAT BRITAIN, GERMANY, FRANCE, EGYPT, CHILE, CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, and ARGENTINA. | Terminated | December 06, 2012 | 150 | 70 | Computer
Z-1966-2012 | 61480 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 220/240V The Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. The System 1 base includes an integrated backup battery in the event of power loss. | The Terumo Advanced Perfusion System 1 | 2 | July 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Advanced Perfusion System 1 battery charge is being depleted faster than expected during periods of system shutdown. The System 1 may falsely indicate a longer battery life than actual battery time remaining if not subjected to regular use, such as would occur with a backup system. This may result in unexpected system shutdown. Prolonged battery depletion may cause permanent battery damage. | DESIGN: Device Design | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to review the Medical Device Recall Notice, assure that all users are aware of the notice. Consignees should place the Addendum at the beginning of the System Base Section (page 3-1) of the Operator's manual, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number indicated on the form. All System 1 product is on quality hold. Terumo CVS will release product including batteries as needed only in the event that customers require service. All affected System 1 customers will receive a safety notification and an addendum to the Operator's Manual. There will be a field correction but no product returned for this recall.Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. | Contact the recalling firm for information | 632 | approx. 632 units | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV.Internationally to: Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam, | Terminated | November 08, 2013 | 487 | 1625 | Computer
Z-1983-2012 | 61151 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Eon Mini Neurostimulation (IPG) System (Model 3788);The product is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. | Eon Mini Neurostimulation (IPG) System | 2 | July 10, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | The firm has received 112 complaints of the Eon Mini IPGs that lost the ability to communicate or recharge due to a workmanship issue resulting in loss of pain relief and subsequent explant. The firm has determined the cause of these reports to be related to process variances in the positioning of the internal battery and printed circuit board, causing a short, and therefore, prematurely depletin | DESIGN: Process Design | St. Jude Medical sent an Important Medical Device Recall letter dated December 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions:For unimplated inventory of affected product customers should contact their SJM Representative to have the device returned to St. Jude Medical. A replacement device will be provided at no additional cost to them.For implanted affected product it is recommended that customers do not unnecessarily explant the devices associated with this advisory if the IPGs are functioning as intended. If there is a sudden loss of power or if the duration between recharges becomes significantly shorter, customers should contact their St. Jude Medical Representative. If device replacement is required due to IPG failure related to this voluntary recall notice, St. Jude Medical will provide a replacement IPG at no charge.For questions regarding this recall call 972-309-2154. | Contact the recalling firm for information | 280 | 25, 255 devices | Worldwide Distribution - USA (nationwide) and internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Open |  |  | 280 | Computer
Z-1988-2012 | 62107 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal software.Intellispace Portal is a multimodality thin client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. | Intellispace Portal | 2 | July 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update to the Intellispace Portal systems to address a number of non-conformances. The most frequent and severe non-conformance identified involves the portal setting of "PAC Integration". This setting has a feature that allows the user to launch images of multiple patient studies into an application, but does not display a warning message informing the user that diff | DESIGN: Software Design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Philips representative or the Customer Care Solutions Center at 1-800-722-9377, option 5, for questions regarding this recall. |  | 96 | 96 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IA, IN, MI, NC, NJ, NY, OK, PA, SC, TN, TX, and WA and the countries of Australia, Austria, Canada, China, Czech Republic, Denmark, and France. | Terminated | March 28, 2013 | 260 | 96 | Computer
Z-1987-2012 | 62005 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal software and Extended Brillance Workspace software Intellispace Portal is a multimodality thin-client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. The Philips Extended Brillance Workspace is intended for use as a diagnostic viewing and processing workstation. | Intellispace Portal and Extended Brillance Workspace | 2 | July 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The Intellispace Portal software and Extended Brillance Workspace software may have significantly different results in the colors and size of the affected area when the customer uses the "Time Sensitive" and "Time Insensitive" methods. | DESIGN: Software Design | Philips Healthcare sent a Urgent Medical Device Correction letter dated May 31, 2012, to all affected customers. The letter was sent via certified mail to each customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office at (1-800-722-9377 Option 5 : Entersite ID or follow the prompts).Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. The user was informed that a field service representative will visit the facility to install the new software.Philips apologizes for any inconveniences caused by this problem. | Contact the recalling firm for information | 38 | 38 | Worldwide Distribution -- Nationwide Distribution including the states of AZ, OR, PA and TX and the countries of Australia, Belgium, Canada, Denmark, Finland, France, Germany, India, Indonesia, Israel, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | September 30, 2013 | 446 | 38 | Computer
Z-1986-2012 | 62342 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx M3536A Defibrillator/Monitor with Software R.03.03The Heartstart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips Healthcare | 2 | July 11, 2012 | 2012 | Philips Healthcare Inc. | Defibrillator/monitor cannot analyze an ECG report during acquisition of 12 lead ECG | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare issued an Urgent Voluntary Medical Device Correction letter on June 20, 2012, to all affected customers. The Notification informs customers of the issue, identifies details (including models and software versions applicable) of units affected, instructions on actions to take while awaiting the correction, and identifies what action Philips plans to eliminate or remedy the issueA Philips Healthcare representative will contact customers to arrange for the updated R.03.03 software installation in all the affected devices.For questions customers were advised to call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 35 | 35 units | Worldwide Distribution - USA including FL, IN, MA, M, NY, OH, TN, TX, and VA. Internationally to AUSTRALIA, HONG KONG, and INDIA | Terminated | August 14, 2013 | 399 | 35 | Computer
Z-2000-2012 | 62128 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal softwareProduct Usage: Intellispace Portal is a multimodality thin-client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. Intellispace Portal is a medical software system that allows multiple users to remotely access the system from compatible computers on a network. | Intellispace Portal | 2 | July 12, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | In the Multi-Modality Tumor Tracking application, when changing the contours of the lesion using the editing mode "edit contour" tool, the measurements are not recalculated causing discrepancies with the lesion measurements. | DESIGN: Software Design | Philips Healthcare sent an Urgent-Medical Device Correction letter dated May 23, 2012, via certified mail to each consignee. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. The letter informed affected customers that a Philips Service Engineer will contact them for implementation of the software update on the affected systems free of charge. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 44 | 44 | Worldwide Distribution - US (nationwide) including the states of: AR, LA, MI, NC, NY, OH, OK and PA and the countries of: Australia, Canada, Denmark, France, Germany, Israel, Netherlands, Spain and Switzerland. | Terminated | March 28, 2013 | 259 | 44 | Computer
Z-2007-2012 | 62434 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MonacoProduct Usage:Planning of radiation therapy | Monaco | 2 | July 12, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | DESIGN: Software Design | Elekta sent an Important Safety Notice to affected customers. The notice identified the affected product, problem, clinical impact and actions to be taken. The notice inform customers of how the issue will be resolved and instructions to follow when it is available. |  | 38 | 38 | US Nationwide Distribution - including the states of: CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI | Open |  |  | 38 | Computer
Z-2002-2012 | 62433 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal SIMPlanning of radiation therapy | Focal SIM | 2 | July 12, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | DESIGN: Software Design | IMPAC Medical Systems, Inc. sent an Important Safety Notice LINFFM0001/1.0 on June 21, 2012, to consignees on concerning a patch to the software. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the DICOM coordinates option for Reports when printing the Setup Reference Diaglog. Customers were instructed go to www.elekta.com, select the SupportPlus Login at the tope of the page, enter their portal, select Downloads/Updates to download the latest patches.For questions regarding this recall call 770-670-2548. |  | 38 | 38 units | Nationwide Distribution including CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI. | Open |  |  | 38 | Computer
Z-1985-2012 | 62112 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrak with Multiple Dataset Navigation, 892.2050 System, Image Processing System, Model Number IT3500 Plus.Product Usage: Usage:The system is an aid to locate anatomical structures during open or percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MR, or X-ray. | InstaTrak with Multiple Dataset Navigation | 1 | July 12, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On October 11, 2006, GE Healthcare recalled GE OEC InstaTrak 3500 Plus System with Software version 5.2, the surgical Navigation and Visualization Application due to software related issues. | DESIGN: Software Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notice dated October 11, 2006 to all affected customers. The notice identified the affected product, problem and actions to be taken. Customers were informed that a GEHC OEC would be contacting their facility to verify the current software version installed at their site. If the issue affects their facility a software upgradewill be provided at no cost. Customers were advised to excercise caution when using the affected product until the 5.2.1software upgrade is completed. For question contact GE Healthcare Customer Service at 1-800-874-7378, option 1 or 1-801-536-4688, option 1.For questions regarding this recall call 801-536-4516. | Contact the recalling firm for information | 86 | 86 units | Worldwide Distribution - US (nationwide) | Terminated | August 21, 2012 | 40 | 86 | Computer
Z-2001-2012 | 62381 | orotrachael intubation guide kit | OGK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | Enforcement Discretion | superDimension Systems AAS00016-xx with Software Version 4.0-4.9.Indicated for use for displaying images of the tracheobronchial tree to aid the physician in guiding endoscopic tools or catheters in the pulmonary tract and to enable marker placement within soft lung tissue. It does not make a diagnosis and is not an endoscopic tool. Not for pediatric use. | superDimension Systems | 2 | July 12, 2012 | 2012 | superDimension, Inc | superDimension, Inc., is conducting a recall of certain superDimension Systems operating Software Version 4.0 - 4.9. The affected systems are sold under the product code AAS00016-xx.The products being recalled have been found to malfunction when connected to the superDimension Locatable Guide (LG). Specifically, the Systems using Software Version 4.0-4.9 will not recognize the Locatable Guide | DESIGN: Device Design | superDimension sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advised customers to discontinue use of the system and requested consignees to complete and return the acknowledgement receipt form via fax to 763-210-4098. Contact the firm at 763-210-4057 for questions regarding this recall. |  | 99 | 99 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alabama, Alaska, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Vermont, Virginia, Washington, West Virginia, and Wisconsin and the countries of Canada, Australia, Austria, Belgium, Czech Republic, Denmark, France, Germany, Hungary, Italy, Netherlands, South Africa, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | February 28, 2013 | 231 | 99 | Computer
Z-2010-2012 | 61763 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1System 1 Base, 220/240V.The Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 13, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of several issues that manifest themselves as communications or power failures in the system network for Terumo Advanced Perfusion System 1. Issues that affect a single pump or module are:" The appearance of a red X on a single pump or module icon on the CCM " The appearance or a single '?' symbol on a pump or module icon on the CCM" The fai | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. CUSTOMER INSTRUCTIONS1. Review this Medical Device Recall notice.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/ email address indicated on the form.Terumo CVS Customer Service-1-800-521-2818Fax-1-800-292-6551.Customer Service Hours:Monday - Friday, 8 AM - 6 PM EST. |  | 648 | 648 units | Worldwide Distribution-USA (nationwide) including the states of , AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | June 10, 2013 | 332 | 1640 | Computer
Z-2013-2012 | 62329 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Molecular Bar Code Scanner User's Guide; instructions for setting up the Bar Code Scanner LN 6L89-01, a component of the Abbott m2000 Instrument System; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list number 06L88-02.Language configuration bar codes - Intended Use: The bar code scanner is a hand-held device that provides a convenient means of scanning sample bar codes to allow positive sample identification and is supplementary to the computer keyboard. Both the scanner and the keyboard transmit data to a System Control Center (SCC) through a single wedge (cable with branched connectors). The bar code scanner must be configured to match the keyboard language being used. In its default configuration set by the manufacturer (Jadak), the scanner is in English format. If a keyboard other than the English keyboard is being used, the user is required to use the scanner to scan an appropriate language configuration bar code provided in the User's Guide to match the language of the keyboard. Users in the U.S have the ability to convert to another language. | Abbott Molecular Bar Code Scanner User's Guide | 2 | July 14, 2012 | 2012 | Abbott Molecular | In the Abbott Molecular Bar Code Scanner User's Guide, List Number 6L88-02/09, the section Modifying the Keyboard Format for Bar Code Scanner LN 6L89-01 contains incorrect bar codes that can prevent the correct configuration of the bar code scanner language format.The Abbott m2000 System is intended for use in performing nucleic acid testing in clinical laboratories. It is comprised of the Ab | PRODUCTION CONTROLS: Error in Labeling | Abbott Molecular Inc. sent an Urgent Field Safety Notice letter dated January 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to replace the affected pages of the Bar Code Scanner User's Guide with the replacement pages provided with the letter, and immediately perform the Modifying the Keyboard Format section with these replacement pages. The accounts were requested to complete the enclosed reply form and fax it to (847) 775-6728 to acknowledge receipt, understanding and implementation of the information provided in the notice. |  | 100 | 100 units | California, Florida, Maryland, Oregon and Virginia, and internationally to Belgium, Brazil, China, France, Germany, Hong Kong, Japan, Malaysia, Mexico, Qatar, Switzerland and United Kingdom | Terminated | December 04, 2012 | 143 | 100 | Computer
Z-2006-2012 | 62349 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe VentilatorDesigned for use on patients who require respiratory support or mechanical ventilation and weigh a minimum of 5 kg (11 lbs). It is suitable for service in hospital and transport environments as a source of continuous or intermittent positive pressure ventilatory support, delivered invasively or non-invasively. It is not intended for homecare use. | EnVe Ventilator | 1 | July 16, 2012 | 2012 | CareFusion 203, Inc. | CareFusion identified that the EnVe Ventilator may not hold the set Positive End Expiratory Pressure (PEEP) value either intermittently or continuously or it may alarm for a Patient Circuit Fault. The unit will activate either a Low PEEP or Patient Circuit Fault Alarm both audibly and Visually to alert the healthcare professional. Ventilation delivery to the patient may be compromised. | DESIGN: Device Design | CareFusion sent an Urgent Product Recall letter with a return receipt dated June 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required Action for Users: If the ventilator resumes normal ventilation (intermittent leak) the audible alarm stops. To clear the alarm indicator on the ventilator display, enter Alarm Messages tab, push alarm Reset to clear display. If the ventilator does not resume normal ventilation (continuous leak) provde an alternative method of ventilation to the patient. CareFusion will follow-up with a phone call to coordinate the correction of the devices. For questions customers should call 1-800-554-8933. |  | 266 | 266 | Nationwide Distribution including AZ, CA, CO, FL, IL, KY, ME, MI, MO, MN, MT, ND, NJ, NM, NV, NY, OK, PA, SC, SD, TX , UT, WI, and WV. | Terminated | February 11, 2013 | 210 | 266 | Computer
Z-2015-2012 | 62123 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Invasive Blood Pressure Care Cables.Product ID Number: 2016995-001, 2016995-002, 2016995-003, 2016995-004. | GE Healthcare, Invasive Blood Pressure Care Cables. | 2 | July 16, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware through complaints of a potential safety issue with its Invasive Blood Pressure Cable made for use with a Spectramed / BD invasive blood pressure transducer. Due to use of an incorrect O-ring, the mating of the cable with the transducer is faulty. This situation inhibits blood pressure measurement from being monitored correctly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent a Urgent Medical Device Correction letter dated June 12, 2012, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Inforamtion. Advised customers to discontinue use, further distribution and to contact their customers. Please return any defective Invasive Blood Pressure Cables by isolating all affected product and completing the attached Confirmation form/fax back per its instructions. This step is required to confirm receipt of communications with all customers.If you have any questions or concerns regarding this notification, please contact Customer Service at +1-800-932-0760 (Domestic & International). Hours of Operation: 8:00 am EST to 6:00 pm EST. . |  | 279 | 279 | Worldwide Distribution -- USA (nationwide) including the states of AK, CA, IL, MS, NJ, PR, SD, TN, TX, and VA. and the countries of ARGENTIAN, AUSTRALIA, BELGIUM, CHINA, GERMANY, ALGERIA, ECUADOR, EGYPT, FINLAND, UNITED KINGDOM, JORDAN, JAPAN, MOROCCO, MAURITUS, NETHERLANDS, OMAN, PERU, SEBIA, SUDI ARABIA, SINGAPORE, EL SALVADOR, TURKEY, URUGUAY and SOUTH AFRICA. | Terminated | December 05, 2013 | 507 | 279 | Computer
Z-2019-2012 | 62437 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, installations with Option C. It is a patient monitor with arrhythmia detection or alarms.The Spacelabs Healthcare Xprezzon Bedside Monitor is intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other devices via Flexport interfaces. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. The Spacelabs Healthcare Xprezzon Bedside Monitor may also function as a generic display or computer terminal. As a generic display or terminal, the patient monitor allows network based applications to open windows and display information the Xprezzon and other networked monitors. The Spacelabs Healthcare Xprezzon Bedside Monitor is designed to communicate with a variety of external devices such as displays, network devices, serial devices, user input devices, audio systems, and local/remote recorders. The Xprezzon Bedside Monitors is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. | Spacelabs Medical Xprezzon Bedside Monitor, Model 91393 | 2 | July 16, 2012 | 2012 | Spacelabs Healthcare, Llc | There were reports that a certain combination of key presses will cause the module to disconnect from the Xprezzon bedside monitor, model 91393, and the patient parameters will drop from the display. The parameters return after power cycling the monitor. | DESIGN: Software Design | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated June 29 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to advise their staff of the potential failure of the XPREZZON Bedside Monitor. Spacelabs will contact customers to schedule a convenient time for their hospital to upgrade their affected monitors, at no cost. On July 6 2012, a customer letter (translated as necessary) URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare XPREZZON Bedside Monitor will be emailed to all international subsidiaries and distributors of record. Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025, ext 2. | Contact the recalling firm for information | 21 | 21 units | Worldwide Distribution - USA including Florida, Iowa, Oklahoma, South Caroline and the countries of India, Germany and Panama. | Terminated | November 28, 2012 | 135 | 21 | Computer
Z-2020-2012 | 61753 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 100/120V.The Terumo Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 17, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | TerumoÂż Advanced Perfusion System 1 has displayed a Red X error message on multiple module icons indicating a loss or fluctuation of internal power on one side of the system's base. Terumo Cardiovascular Systems' initial investigation has determined that there may be more than one cause for the malfunction and is investigating further to determine the root cause.Potential risks with the loss or | PRODUCTION CONTROLS: Process Control | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to place the addendum in the Operator's Manual for Terumo System 1, at the beginning of the Central Control Module Section. Additionally, a Customer Response Form was attached for customers to complete and return to the firm. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. | Contact the recalling firm for information | 992 | 992 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 26, 2013 | 252 | 1640 | Computer
Z-2030-2012 | 61691 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMASpectrum Infusion Pump Master Drug Library (MDL) Editor software.Product Usage:A customized in-house library of all IV and epidural drugs, along with their safe delivery parameters. | SIGMASpectrum | 3 | July 18, 2012 | 2012 | Sigma | Sigma is a safety alert for the Master Drug Library (MDL) Editor software associated with the SIGMA Spectrum Infusion Pump System because it may result in a facility drug library that does not represent common clinical practices. | DESIGN: Software Design | Sigma sent an Urgent: Customer Safety Alert dated April 23, 2012 via UPS delivery to all consignees requesting both acknowledgement of notification and confirmation that recommended actions have been taken. Customers are asked to follow the instructions contained in the recall letter. Customers with questions may contact the firm at 800-356-3454 ext 300. | Contact the recalling firm for information | 2048 | 2,048 CD's | Worldwide Distribution - US (nationwide) including Puerto Rico, and Canada. | Open |  |  | 2048 | Computer
Z-2042-2012 | 61479 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1. System 1 base 100/120V.The Terumo Advanced Perfusion System I is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only When used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | July 19, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received several reports of malfunctions of the Air Bubble Detection System used with its heart-lung machines over the last three years. The malfunctions were expressed in various failure modes (oversensitivity, false alarms and were found to have various root causes. The malfunctions occurred in systems that were excluded from prior corrective actions or in syste | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to complete and return a Customer Response Form via fax to 1-800-292-6551. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 992 | 992 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, BRISBANE, CANADA, CHILE, China, COLOMBIA, COLOMBIA, Costa Rica, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, GERMANY, Guatemala, HOLLAND, Honduras, Hong Kong, India, INDIA, Indonesia, Israel, ITALY, Jordan, KOREA, Kuwait, Lebanon, LEICESTERSHIRE, Malaysia, Mexico, NETHERLANDS, New Zealand, Nicaragua, NORWAY, Pakistan, Panama, PERU, Philippines, Russia, Saudi Arabia, SINGAPORE , SOUTH AFRICA, South Korea, SPAIN, Sri Lanka, SWEDEN, SWITZERLAND, Taiwan, THAILAND, Turkey, UNITED ARAB EMIRATES (UAE), Uruguay, Venezuela, Vietnam, WEST GERMANY and Yemen. | Terminated | April 12, 2013 | 267 | 4689 | Computer
Z-2036-2012 | 62284 | humidifier, respiratory gas, (direct patient interface) | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI Concha Therm Neptune Humidifier, Rx Only, Product Usage: Usage:Device that is intended to add moisture to, and sometime to warm, the breathing gases for administration to a patient. | Hudson RCI Concha Therm Neptune Humidifier | 2 | July 19, 2012 | 2012 | Teleflex Medical | Speaker Failure. If the speaker fails and an event occurs which would generate an audio alarm, the alarm will not sound and could cause a potential delay in patient treatment. | DESIGN: Component Design/Selection | Teleflex sent Urgent Medical Device Notification letters dated June 12, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to complete the enclosed Acknowledgement Form and return per the instructions provided. Once the form is received, Teleflex Medial will schedule for a technician to come to their facility and exchange the affected product. For questions call your local sales representative or Customer Service at 1-866-246-6990. |  | 14794 | 14,794 each | Worldwide Distribution - U.S. (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Chile, China, Europe, India, Japan, Mexico, South Korea, Taiwan and UK. | Terminated | January 28, 2014 | 558 | 14794 | Computer
Z-2041-2012 | 62420 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Eclipse Treatment Planning System for radiotherapy; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | July 19, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning (for radiotherapy treatments) where changing the prescription in a calculated VMAT or RapidArc plan may lead to Monitor Units which do not reflect the new prescription. If the dose prescription of a VMAT or RapidArc plan is changed after the initial Dose Volume Optimization, the Monitor Units and dose distribution may not reflect | DESIGN: Software Design (Manufacturing Process) | The firm, VARIAN Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION; URGENT FIELD SAFETY NOTICE" dated June 18, 2012, to all affected users. The notice described the product, problem and actions to be taken. Varian is developing a technical correction for this issue. The customers will be contacted by a Varian Customer Service representative when this correction is available to schedule its installation.If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager.In USA and Canada: 1.888.VARIAN5 (888.827.4265)Europe: +41 41 749 8844; email: support-emea@varian.comEmail: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com and Latin America: soporte.al@varian.com. Internet: Oncology systems customer site - www.myvarian.com and Varian Medical Systems public site - www.varian.com. |  | 7113 | 7113 units | Worldwide distribution: USA (nationwide) and countries including of: Austria, Australia, Bahrain, Belgium, Brazil, Canada, China, Chile, Colombia, Denmark, Egypt, Finland, France,Germany, Greece, Hungary, Hong Kong, India, Indonesia, Israel, Ireland, Italy, Japan, Latvia, Malaysia, Mexico, Morocco, Nepal, Netherlands, Norway, Panama, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, United Kingdom, Taiwan, Turkey, Ukraine, United Arab Emirates and Venezuela. | Open |  |  | 7113 | Computer
Z-2037-2012 | 62372 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria CT System SoftwareProduct Usage: The Scenaria system is indicated to acquire axial volumes of the whole body including the head. The images can be acquired in either axial, helical, gated or dynamic modes. The volume datasets acquired by the Scenaria can be post processed by the Scenaria to provide additional information. Post processing capabilities included in the Scenaria software include CT angiography (CTA), Multi-planar reconstruction (MPR) and volume rendering. | Hitachi Scenaria CT System Software | 2 | July 19, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error in the Reconstruction Status function that creates a potential risk that a patient study will not be fully reconstructed. When multiple reconstructions are queued for processing, if a queue is suspended and the task order changed, the second queue may not reconstruct the full range of slices. This may result in missing diagnostic data. | DESIGN: Software Design | Hitachi sent a "Device Correction Notice" dated May 4, 2012 to their customers. The notice identified the affected product, problem and actions to be taken. Customers were instructed to "not use the Reconstruction Status" function to suspend active reconstructions until the software problem can be corrected. They also informed their customers that a representative will visit their facility to install the software update to permanently resolve the issue at no cost to the facility. For further information regarding this notice call 800-800-3106 x3720. | Contact the recalling firm for information | 6 | Six (6) units | US Nationwide Distribution to the following states of: MD, MI, MX, ND, NY, OH and TX. | Terminated | February 19, 2013 | 215 | 6 | Computer
Z-1993-2012 | 62527 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima Mobile X-ray System. The GE Automatic Mobile X-Ray (AMX ) Series are intended to take exposures utilizing film or computed radiography ( CR ) | GE Healthcare Optima Mobile Xray System | 2 | July 19, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR200amx, ond Optima XR220amx related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication tothe user or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the firmware on the system to address the issue.Your proposed corrective action plan (CAP) (Field Modification Instructions 10865) describes the rework plan that GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. | Contact the recalling firm for information | 254 | 254 units have been shipped to/installed in the U.S. | Worldwide Distribution (USA) nationwide | Terminated | January 10, 2013 | 175 | 254 | Computer
Z-2066-2012 | 61910 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS System Software For use in the in-vitro quantitative, semi quantitative, and qualitative measurement of variety of analytes of clinical interest. | VITROS 5,1 FS System Software | 2 | July 23, 2012 | 2012 | Ortho-Clinical Diagnostics | software automatically switches the lot when a vancomycin (VANC) or valproic acid (VALP) reagent is depleted or the operator manually changes the lot in use causing lack of verification of proper quality assessment. | DESIGN: Software Design | Ortho Clinical Diagnostics notified its foreign direct account consignees via email on May 21, 2012 and their direct account US consignees were sent an Urgent Distributor Safety Notice on May 22, 2012 via US Postal Service Priority Letter. The recall letter notified the consignees of the situation and that Ortho will replace their remaining inventory of non-expired GENs 14 or 15 for VITROS VALP Reagent only. Until then the consignees can continue using the VITROS VALP Reagent, GENs 14 or 15 until the replacement arrives. And to note that VITROS VANC Reagent, GENs 19 &20 are now expired. --- The letter also requested the consignees to complete and return the enclosed Confirmation of Receipt, no later than May 31, 2012. And to review the quality control beginning with the Calibration date of the VITROS VANC Reagent ( GEN 21 or above) or VITROS VALP (GEN & above) so that the QC results can be verified to be within acceptable ranges. --- For customers using VITROS VALP Reagent GEN 14 or 15, they were to respond as to the quantity of the non-expired GENs 14 and 15 remaining inventory; to load and use only one GEN of VITROS VALP Reagent at a time on the system until the replacement arrives. This notification letter should be posted by each VITROS System using the VITROS VALP Reagent or VITROS VANC Reagent or with the user documentation. And in addition, if these products were distributed then this information should be forwarded as well.For questions regarding this recall call 585-453-4110. | Contact the recalling firm for information | 2044 | 2044 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Brazil, Canada, Chile, Argentina, Paraquay, Ecuador, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, UK, Northern Europe, France, Germany, Italy, Spain and Portugal. | Open |  |  | 3508 | Computer
Z-2073-2012 | 62519 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Digital Linear Accelerators with syngo RT Therapist version 4.2/MOSAIQ OIS Systems and Multiple-X, Manufactured by SIEMENS AG. The ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ARTISTE or ONCOR Digital Linear Accelerators with syngo RT Therapist | 2 | July 24, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety issue when using the "IMMEDIATE RESUMPTION" feature of the syngo RT Therapist 4.2 software after a Control Console interlock during the delivery of flattening-filter-free (Multiple-X) beams or segments; in a very rare scenario, the RT Therapist will send again the previously completed beam or segment for delivery to the control console, thus repeating th | DESIGN: Software Design | The firm issued "Urgent: Medical Device Correction" letters delivered by the Siemens Customer Service Group or by certified mail to all its direct consignees, informing them of the affected product and providing instructions on the recall. The firm stated that release of this UI TH018/12/S to distribute the notifications began on June 25, 2012. | Contact the recalling firm for information | 18 | 18 | Worldwide Distribution, including the state of Florida (USA) and the countries of Germany, India, Poland, Thailand, and the Republic of Korea. | Terminated | October 18, 2012 | 86 | 18 | Computer
Z-2069-2012 | 62534 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV ReportingThe Results Management (RM)/IMPAX CV reporting component facilitates the quick, effective creating of digital structured reports for : adult catheterization and echocardiography, pediatric echocardiography, congenital heat disease, nuclear cardiology and non-invasive vascular disease management. This webb-based application allows cardiology departments to view, edit and sign the structured reports. | IMPAX CV Reporting NonInvasive Vascular (NIV)/Results Management (RM) NIV | 2 | July 24, 2012 | 2012 | AGFA Corp. | When users selected "Left stenosis" in the "Graft Duplex Conclusion" section in IMPAX CV Reporting Non-Invasive(NIV) module, "Right stenosis" appeared in the sentence display within the printed representation of the report. | PRODUCTION CONTROLS: Error in Labeling | AGFA Healthcare sent an Urgent Field Safety Notice dated July 10, 2012, to all affected consignees indicating that a representative from the IMPAX CV product management or professional services team would be contacting them to perform a review of their system. For questions customers were asked to call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 6 | 6 | Nationwide Distribution including IN, SC, OH, TX | Terminated | October 29, 2013 | 462 | 6 | Computer
Z-2040-2012 | 62382 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 6800 Miniview, Image-intensified fluoroscopic x-ray system used during diagnostic, surgical and interventional procedures.Designed to provide the physician with general fluoroscopic visualization of the patient including but not limited to surgical orthopedic and extremity imagine. The device is not intended for whole-body pediatric imaging. | OEC 6800 Miniview | 1 | July 24, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 6800 Miniview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lock-up, incorrect image display, incorrect dose area calculated air kerma in "film mo | DESIGN: Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On 11/20/2006, two written notifications were sent to 1513 identified consignees, addressed to the Facility Administrator and the Radiology Administrator for a combined total of 3026 letters. The Urgent Recall Notice, dated 11/8/2006, advised customers that GE Healthcare had identified several intermittent potential safety issues that may occur with the identified affected product. The notification also provided interim or permanent solutions for each of the problems described. For the issues that did not have permanent solutions at the time, the firm was to contact customers and remedy the issues free of charge. For the issues that did have a permanent fix at the time, the GEHC OEC service team was to contact customers to install the correction. If customers had questions or concerns, they were to contact the service team for further information at 800-874-7378 option 8. On 04/23, 2007, a second written notification was sent via certified mail to the same 1513 consignees providing further clarification on the issue of "incorrect Image Display" and a temporary remedy to the issue. |  | 1933 | 1,933 units | Worldwide Distribution | Terminated | May 13, 2013 | 293 | 1933 | Computer
Z-2065-2012 | 62387 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 8800 Flexview, 892.1650 Image-intensified fluoroscopic x-ray system Product Usage: The 8800 Flexview is designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical, and interventional procedures. Clinical applications may include cholangiography, urologic, endoscopic, urologic, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedures. The system may be used for other imaging applications at the physicians discretion. | OEC 8800 Flexview, | 1 | July 24, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 8800 Flexview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mod | DESIGN: Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent two Urgent Recall Notice letters to the Facility Administrator and the Radiology Administrator. The letters identified the affected product, problem and actions to be taken. The letters advised consignees that GE Healthcare is actively working on solutions that will permanently resolve each of the issues. For questions or concerns contact GE Healthcare service team at 800-874-7378 option 8. |  | 1257 | 1257 units (472 domestic systems) | Worldwide Distribution - US Nationwide | Terminated | May 13, 2013 | 293 | 1257 | Computer
Z-2078-2012 | 62462 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza. Radiological image processing system | Siemens syngo.plaza | 2 | July 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential malfunction when using syngo.plaza with software version VA20C_HF01. For datasets with distance measurements it can happen, that all images are not loaded to Viewer but no error message is displayed to user. | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notice letter dated February 24, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to ensure that the prefix defined in the viewer settings for "Content Creator's Initials" is not empty. Then all images will be display correctly. This has to be checked for every user. Further questions please call 610-219-6300. |  | 2 | 2 | (USA) nationwide including the states of CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 14, 2014 | 537 | 2 | Computer
Z-2076-2012 | 62007 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015Product Usage: The device is labeled for prescription use only. | Alaris PC unit | 1 | July 26, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because the Alaris PC unit model 8015 has a component on the PC unit power supply board that is causing an error code(120.4630) "System Error" or "Missing Battery Error" at start up. The error code is accompanied by both an audible alarm and visual error messages on the PC unit screen. | DESIGN: Component Design/Selection | Carefusion sent an Urgent: Medical Device Recall Notification letter dated June 27, 2012 to all their customers who purchased the Alaris PC unit (model 8015). The letter identified the affected units, explanation of the problem, potential risk and actions to be taken. Customers were instructed to not return their devices. If customers observe either of the error messages identified in the letter, then they were instructed to remove the PC unit from use and contact the Carefusion Support Center. Carefusion will contact their facility by phone within 60 days of receiving this notification to schedule a visit to replace their affected power supply board on their PC unit. Customers were instructed to promptly complete and return the enclosed customer response card to expedite the corrective action process. For questions and support contact:CareFusion Support Center at 888-562-6018 for (Recall Related Questions), Customer Advocacy at 800-854-7128, Option 1, Option 1, Option 3 or email customerfeedback@carefusion.com for (Adverse Event Reports), Technical Support at 888-812-3229 for (Technical Questions Regarding the Alaris System). |  | 112912 | 112,912 units total (106,888 units in US) | US Nationwide Distribution | Open |  |  | 112912 | Computer
Z-2107-2012 | 61247 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂŽ Syringe Infusion Pump, Model 4000, Software Version V1.0 & V1.1, Reorder Numbers 4000-0100-50, 4000-0100-95, 4000-0101-50, 4000-0101-51, and 4000-0101-78.Designed for use in care areas for the administration of fluids requiring precisely controlled infusion rates. | Medfusion 4000 V1.0 &V1.1; Syringe Infusion Pump | 1 | August 01, 2012 | 2012 | Smiths Medical ASD, Inc. | Smiths Medical has identified a software anomaly in the Medfusion 4000 Syringe Infusion Pump that causes a device history log corruption and triggers a Watchdog Fail-Safe alarm. When the device exhibits this failure mode, visual and audible alarms will sound and the device ceases operation. | DESIGN: Software Design | Consignees of MedfusionÂŽ 4000 pumps with software version 1.0 & 1.1 were visited by a Smiths Medical representative on 2/14/12 informing them that their pumps require a software upgrade. A record of this software upgrade was documented as part of the records for this action. | Contact the recalling firm for information | 607 | 607 pumps (257 USA, 350 Canada) | Worldwide Distribution -- USA, including the state of CA, and country of Canada. | Terminated | November 14, 2012 | 105 | 607 | Computer
Z-2143-2012 | 62023 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. Product UsageSoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable, trained, and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities, for the purpose of storing, managing, querying, and reporting laboratory data. | Soft Lab | 2 | August 03, 2012 | 2012 | SCC Soft Computer | On 10/17/2011 SCC Soft Computer, Clearwater, FL initiated a correction on SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. A client reported an issue which caused two different sets of results from the Sysmex series of Hematology instruments to display under the same order on the instrument menu. | DESIGN: Software Design | SCC Soft Computer's proprietary Task Management System sent a correction communication letter to all consignees with affected software. The letter identified the affected product, problem and corrective actions to be taken. The letter also recommended alternative working solution if the corrective actions cannot be used. Consignees with affected software are scheduling receipt of the software correction. For questions call 727-789-0100. |  | 88 | 88 | US Nationwide Distribution | Open |  |  | 88 | Computer
Z-2139-2012 | 62558 | system/device, pharmacy compounding | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | AUTOMIX High Speed Compounder System, product 2M8077; automated nutrition compounders; Product Usage: The Automix Compounder Systems are electromechanical positive displacement fluid compounding devices used in the admixture of large volume parenteral nutrition formulations via gravimetric weighing. | AUTOMIX | 1 | August 08, 2012 | 2012 | Baxter Healthcare Corp. | After fluid ingress into the keypad of the AUTOMIX Compounder's (AUTOMIX) control module, the AUTOMIX may generate an incorrect device response to an AUTOMIX operator's key press. Also, the AUTOMIX system may experience an intermittent electrical failure, in which unexpected pump motor operations will occur. Also, the AUTOMIX system may intermittently experience an electrical failure where thec | DESIGN: Device Design | Baxter Healthcare Corp. sent an Urgent Product Recall letters dated July 13, 2012, via first class mail to all affected customers, informing personnel that Baxter was executing a voluntary controlled product withdrawal of the Automix system from the U.S. market by June 1, 2013. Until their units can be replaced, customers were advised to exercise care when entering data, and verify the accuracy of the keypad response; to inspect the umbilical cable of the Automix for any visible cuts, nicks deformation, or physical damage to the cable, discontinuing use of the device if such damage is found; and to return the device to Baxter if the device stops compounding without an alarm and without a visible green complete light, discarding the solution bag. The customers were requested to convey this information to all users of the Automix compounders, and a wall chart was provided to hang over the device to remind users of the mitigations to take. The customers were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667. |  | 270 | 270 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Open |  |  | 5166 | Computer
Z-2162-2012 | 62503 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The ABL software version 6.12 for the ABL800 analyzer. The analyzers are intended for the in Vitro testing of: -Samples of whole blood for the parameters pH, pO2, pCO2,cK+, cNa+, cCa2+, cCl, cGlu, cLac, cCrea, ctBil, and co-oximetry parameters (ctHb, sO2, and the hemoglobin fractions FO2Hb, FCOHb, FMetHb, FHHb and FHbF)-Samples of expired air for the parameters pO2 and pCO2-Pleura samples for the pH parameter.The ABL software version 6.12 for the ABL800 and ABL700 analyzer. The analyzers are intended for the in Vitro testing of: -Samples of whole blood for the parameters pH, pO2, pCO2, cK+, cNa+, cCa2+, cCl-, cGlu, cLac, cCrea, ctBil, and co-oximetry parameters (ctHb, sO2, and the hemoglobin fractions FO2Hb, FCOHb, FMetHb, FHHb and FHbF) -Samples of expired air for the parameters pO2 and pCO2 -Pleura samples for the pH parameter. | ABL 800 FLEX blood gas analyzer | 2 | August 08, 2012 | 2012 | Radiometer America Inc | A customer has reported an error on an ABL8xx configured to a Hospital Information System using the data transfer protocol HL7 version 2.5. If the instrument transmits results to a HIS system and the connection is broken the results are queued for later transmission. If the ABL and HIS system are configured with HL7 2.5, any results in the queue will be mixed up when the instrument transmits t | DESIGN: Software Design | The firm, Radiometer, sent two "Priority Level: Urgent" letters one dated July 9, 2012 and the other dated July 24, 2012 to its affected customers via FedEx with tracking. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the attached Recall Response Fax Form via fax to 800-736-0601. The letter informed the customers that a local service representative will be contacting them shortly to install a software update that will prevent this from occuring (in the event that you utilize HL7 version 2.5 or intend to at some future date.) If you have any questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt. 4. |  | 11937 | 11929 ABL 8xx units | Worldwide distribution: USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech republic, Denmark, Ecuador, Egypt, Eritrea, Estonia, Finland, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Malta, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Qatar, Republic of Korea, Romania, Russia, Saudi Arabia, Switzerland, Singapore, Slovenia, South Africa, Spain, Sweden Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Yugoslavia, and Zimbabwe. | Terminated | August 20, 2013 | 377 | 20225 | Computer
Z-2153-2012 | 62562 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian ARIA Oncology Information System, Import Exportapplication, v 11.0; Reference/FSCA Identifier: CP-08847; Model Number: HIT; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The ARIA Radiation Oncology product is a treatment plan and image management application. It enables the authorized user to enter, access, modify, store and archive treatment plan and image data from diagnostic studies, treatment planning, simulation, plan verification and treatment. ARIA Radiation Oncology also stores the treatment histories including dose delivered to defined sites and provides tools to verify performed treatments. ARIA also includes optional tools to manage the oncology department's clinical and business information. | Varian | 2 | August 08, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Import Export Application version 11.0 whereMLC leaf positions for plans using BrainLab m3 and mMLC are incorrectly imported. | DESIGN: Software Design | Varian Medical Systems sent a " Urgent Medical Device Correction Urgent Field Safety Notice" dated July 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed until a correction is deployed all treatment plans using the BrainLab m3 mMLC must be imported using the Import Wizard only.Varian Medical Systems is developing a technical correction for this issue. You will be contacted by a Customer Service Representative when this correction is available to schedule its installation on your system.For further questions please call 1.888.827.4265. |  | 2 | 2 Units. | Nationwide Distribution including the states of New York and Ohio. | Terminated | July 01, 2013 | 327 | 2 | Computer
Z-2156-2012 | 62551 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva with Flat Detector (MD Eleva FD)Model: 708034Product Usage:Philips X-ray System | Philips | 2 | August 08, 2012 | 2012 | Philips Healthcare Inc. | When doing a 3 Dimensional Rotational Angiography (3DRA) the 3D reconstruction does not always start, even when all images have been acquired. | DESIGN: Software Design | Philips issued an Urgent Field Safety Notice dated July 2, 2012 to all affected customers. The letter identified the affected product, problem description, actions to be taken by the customer/user in order to prevent risks for patients or users and actions planned by Philips to correct the problem. For question or support concerning this issue contact your local Philips representative. | Contact the recalling firm for information | 41 | 41 units | Worldwide Distribution - US (nationwide) including the states of CA, CT, MI, MN, MN, NC, NY, NY, VA and the countries of Belgium, Columbia, Czech Republic and Denmark.FranceGermanyItalyLatviaLuxembourgNetherlandsRussian Fed.SpainSwedenSwitzerlandThailandUnited Kingdom | Terminated | March 26, 2013 | 230 | 41 | Computer
Z-2136-2012 | 62414 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Radiological Image Processing SystemThe system is an aid to locate anatomical structures during open or percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CS, MR, or X-ray. | FluoroTrak Spinal Navigation System | 1 | August 08, 2012 | 2012 | GE OEC Medical Systems, Inc | The FluoroTrak Spinal Navigation Application on the OEC 9900 EliteNAV could result in an incorrect position of the navigated instrument(s) versus the displayed reference image. | DESIGN: Software Design | GE Healthcare sent an Urgent Recall Notice dated April 2, 2007, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were asked to call 800-874-7378. For questions regarding this recall call 801-536-4952,. |  | 17 | 17 | Worldwide Distribution - USA including MO, IL, LA, FL, NH, TX, IN, PA, NJ and the countries of France and China | Terminated | September 28, 2012 | 51 | 17 | Computer
Z-2172-2012 | 62448 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | EasyLink(TM) Informatics System, Part Number 1000034805. EasyLink(TM) Informatics System is classified as a data calculator and processing module for use with In-Vitro diagnostics devices. | EasyLink Informatics System | 2 | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | There is a potential for the EasyLink system to display and transmit to the Laboratory Information System networks an incorrect result following the EasyLink systems application of the configured result number format. The problem is encountered when the result number format configured on the Sysmex CA-7000 or CA-1500 coagulation analyzer does not match the EasyLink system settings on the following | DESIGN: Software Design | The firm initiated this recall on June 7, 2012 by issuing an "Urgent Field Safety Notice" to all customers who are known to have the EasyLink system and who also own and operate Sysmex CA-1500 and CA-7000 instruments. The Notice described the problem and provided recommended actions. Customers can contact 800-441-9250 for questions. |  | 237 | 237 | Worldwide distribution, including the countries of Australia, Austria, Belgium, Canada, Czech Republic, France, Germany, Italy, Netherlands, New Zealand, Slovakia, Spain, Switzerland and the United States, with nationwide distribution. Center Recall Depth is Retail. | Open |  |  | 237 | Computer
Z-2177-2012 | 62444 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 base 100/120VThe Terumo Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | August 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems sent a Safety Advisory (AA-2010-007-C) on December 2, 2010 (Recall# Z-0882/0883-2011 RES 57416) regarding malfunctions of the Occluder and Air Bubble Detection (ABD) Modules to a confined number of Terumo Advanced Perfusion System 1 users. The investigation of the malfunction found a faulty component on the modules' application board caused the failures. Terumo CVS re | DESIGN: Component Design/Selection | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 26, 2012 to its customers. The letter described the product, problem and actions to be taken. To customers were instructed to review this Medical Device Recall notice and addendum; assure that all users are aware of this notice and addendum; place the addendum in the front of Section 7: Modules (p 7.1) and confirm receipt of this communication by faxing the attached Customer Response Form to the fax number/email address indicated on form. Terumo CVS is issuing a Safety Advisory to alert all Terumo System 1 users of the potential for failure and providing an addendum for the Operator Manual with instruction on how to respond to the failure.If you have any question or concerns contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are Monday-Friday, 8AM - 6PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Japan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | November 16, 2012 | 99 | 1641 | Computer
Z-2173-2012 | 62403 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Siemens syngo(R) Lab Data Manager.The syngo(R) Lab Data Manager is classified as a data calculator and processing module for use with In-vitro diagnostic devices. It also provides work flow management as well as remote instrument control to multiple interfaced instruments | Siemens syngo(R) Lab Data Manager | 2 | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The recalling firm has discovered that when using a Sysmex(R) CA-7000 or CA-1500 coagulation analyzer with the syngo Lab Data Manager (LDM), there is a potential for the syngo LDM system, to display and transmit to the LIS an incorrect result following the syngo LDM's application of the configured result number format. The problem is encountered when the result number format configured on the Sys | DESIGN: Software Design (Manufacturing Process) | Siemens Healtcare Diagnostics sent a Urgent Field Safety Notice dated June 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to discuss the contents of this letter with your Laboratory Director regarding the need to review any previous test results.If you have any technical questions or concerns, please contact the Seimens Technical Solutions Center at 1-800-441-9250, for further assistance. |  | 15 | 15 | USA (nationwide) including the states of AZ, CA, MS, MO, NJ, NC, OH, OR, PA, and TX. | Open |  |  | 15 | Computer
Z-2184-2012 | 62646 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular (CV) ReportingReporting tool used for structured reporting and structured report outputs (including XML style sheet customization and PDF) | IMPAX Cardiovascular (CV) Reporting | 2 | August 10, 2012 | 2012 | AGFA Corp. | Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE dated July 27, 2012, to all affected customers. The letter described the safety alert and mitigation. Customer were asked to complete the Acknowledgment form and fax back to 864-421-1664 that the information was received and understood.For questions customers were instructed tol call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 280 | 280 | Worldwide Distribution - USA (nationwide) and Canada | Open |  |  | 280 | Computer
Z-2202-2012 | 62310 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-1 System Battery --- Made in U.S.A. for Carestream Health, Inc., Carestream Health, Inc., 150 Verona St, Rochester, NY 14608. --- The battery may be sold with the DRX-1 imaging device or as a stand-alone accessory. | Carestream DRX1 System battery | 2 | August 13, 2012 | 2012 | Carestream Health, Inc. | Carestream DRX-1 System batteries manufactured before June 2012 due to two occurrences of lithium-ion battery overheating, which may cause burns. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Firm sent consignees "Urgent: Customer Notification and Field Corrective Action" letters on June 8, 2012. The letters described the issue and provided recommended actions. Service dealers will contact all affected consignees to evaluate the battery. U.S. customers with questions can contact Carestream Customer Care at 800-328-2910. Customers outside the United States should contact their local Carestream service number. |  | 9971 | 9971 units | Worldwide Distribution, including Nationwide (USA). | Open |  |  | 9971 | Computer
Z-2203-2012 | 61614 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Advantage Workstations, Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Workstation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA), Model 2378698-2Product Usage:Advantage Workstation version 4.2: The Advantage Workstation 4.2 is a review station, which allows easy selection, review, processing, filming and medial interchange of multi-modality images from a variety of diagnostic imaging systems. When interpreted by a trained physician, filmed images may be used as an element for diagnosis. Advantage Workstation version 4.3/4.4 : The Advantage Workstation 4.3 is a review station, which allows easy selection, review, processing, filming and medial interchange of multi-modality DIACOM images from a variety of diagnostic imaging systems. When interpreted by a trained physician, filmed images of the AW monitor may be used as a basis for diagnosis, except in the case of mammography. | GE Healthcare Advantage Workstations | 2 | August 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is informing you of incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed they are asking you not to use the aorta analysis protocol or the customized protocols for an aneurysm.A second issue was also discovered internally. When loading PET data into the Volume Viewer, ifslices are missing in a PET serie | DESIGN: Software Design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 23, 2012 to consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers & Radiologists. The letter described the affected product details, safety issue, safety instructions and product correction. For more information on this topic, contact a GE Application Specialist or your GE Field Service representative.For questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 1055 | 1055 | Worldwide Distribution - US (nationwide) including the states of : AL, AZ, AR, CA, CO, DE, DC, FL, GA, GU, HI, IK, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO,MT, NE,NH,NJ, NM, NY, NC, ND, OH, OK, PA, PR, SC, SD, TN, TX, VA, WA, WV, WI and the countries of YUGOSLAVIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAÂŽDE ARABIA, RUSSIAN FEDERATION, ROMANIA, REPUBLIC OF MALTA, PORTUGAL, POLAND, PHILIPPINES, NORWAY, NIGERIA, NEW ZEALAND, MEXICO, KOREA, KENYA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HUNGRY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | December 11, 2012 | 120 | 1055 | Computer
Z-2208-2012 | 62568 | table, radiographic, tilting | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy Diagnost Classic-Stationary x-ray systemCatalog Numbers: 70642, 70643, 706030, 706031, 706032Stationary x-ray system, Universal RF system for general use. | Philips Easy Diagnost ClassicStationary xray system | 2 | August 15, 2012 | 2012 | Philips Healthcare Inc. | Inclination switch in the stand may have been mounted incorrectly, as a result, the switch could not stop the tilting movement of the table base due tol hardware or software errors. | PRODUCTION CONTROLS: Equipment Maintenance | Philips issued an Urgent Field Safety Notice dated July 9, 2012, to notify users that the inclination swtich may be mounted incorrectly. A Field Safety Engineer will be dispatched to verify if a switch was installed properly, and in the event it was not installed properly they will re-install it properly. The Field Service Engineer will add a label as described in the spare part replacement manual, which is part of the FCO kit.. Customers were instructed review the information contained in the Field Safety Notice and to retain a copy of the notice with the equipment Instruction for Use. For further questions customers were instructed to call their Phillips representataive Technical Support Line at 1-800-722-9377. For questions regarding this recall call 978-687-1501. |  | 12 | 12 units | Worldwide Distribution - USA including IN, MA, MI, MO, NC, NY PA, SC, and WI and to the countries of Canada, HongKong, New Zealand. | Open |  |  | 12 | Computer
Z-2219-2012 | 62581 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Version 9.5 | syngo Dynamics Version 9.5 | 2 | August 17, 2012 | 2012 | Siemens Medical Solutions, USA, Inc | If multiple uncompressed images/clips are sent to the syngo Dynamics server, around the same time, it is possible for an image from a patient to be placed in another patient's study. |  | On 07/13/2012, Siemens issued a "Customer Safety Advisory Notice" to affected consignees according to Update Instruction SY062/121S. The notice informs the customers of the potential issue with the identified affected product and provides details regarding the resolution. A software update to relieve the symptom of the issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY063/121S. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY064/121S. |  | 30 | 30 | Nationwide Distribution -- MA, GA, MO, IA, CA, PA, OR, CO, NY, NJ, WA, UT, OH, WI, MI, KY, and NC. | Terminated | July 15, 2013 | 332 | 30 | Computer
Z-2223-2012 | 62641 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are whole body Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | The Brilliance 64 and Ingenuity CT Computed Tomography Scanners | 2 | August 17, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | 1. The current system design of software version 4.0.0xx379 does not prevent the system from scanning if the dose level exceeds a set upper dose limit for CT Brain Perfusion scanning, which previous versions did. However, the current version does have the Dose Check (NEMA XR-25) feature, which provides dose notifications and alerts, but it is possible for the user to disable Dose Check after ente | DESIGN: Software Design | Philips Healthcare sent an URGENT - Medical Device Correction letter date July 18, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Phillips will be installing a software update version 4.0.0xx516. A Philips Service Engineer will contact the customer for implementation of the software update on the affected systems free of charge.For further information or support concerning this issue, customers may contact their local Philips representatives or Philips Healthcare office. For North America and Canada, they may also contact the Customer Care Solutions Center (1-800-722-9377, option 5)For questions regarding this recall call 440-483-7000. |  | 8 | 8 | Worldwide Distribution - USA including IN and OH, also, the countries of Australia, Belgium, China, Israel, and Sweden | Terminated | February 22, 2013 | 189 | 8 | Computer
Z-2230-2012 | 62684 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Serial interface module RS-232Product Usage:The System 1 provides perfusion case data transfer through the use of the data transfer module. This module is a general purpose serial interface and is the interface between the TLink Data Management ystem or Sarns CCPro Data Perfusion oftware and the system. The data collected and exported by the data transfer module includes physiological data and device events to record the progress and events during a procedure. | System 1 | 2 | August 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of malfunctions involving the loss of data transfer capabilities using the data transfer capabilities using modules on the Terumo Advanced Perfusion System 1. Some reports stated that the odor of smoke was as associated with the failures. The malfunctions occurred during set-up, during cardiopulmonary bypass, and during routine testing of the modu | DESIGN: Device Design | Terumo sent an URGENT MEDICAL DEVICE Recall letter dated May 4, 2012 to all affected customers. The letter identified the affected product, reason for correction, potential hazard and actions to be taken. The letter stated that Terumo CVS will replace all affected modules. Customers were informed that a Terumo VS Representative will contact the user to schedule a service call for the replacement upon receipt of the Response Form attached. For questions contact Terumo CVS Customer Service at 800-521-2818. | No consumer action necessary | 579 | 579 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV and the countries of: AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Germany, Hong Kong, Indonesia, Japan, Jordan, Malaysia, Singapore, South Africa, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam | Open |  |  | 1427 | Computer
Z-2229-2012 | 62452 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART & Sunquest Laboratory SMART SelectProduct Usage: for clinical use | Sunquest Laboratory SMART and SMART Select | 3 | August 21, 2012 | 2012 | Sunquest Information Systems, Inc. | In GUI inquiry, searching by ID (CID) may return the wrong patient demographic data if the accession number associated to the CID is no longer inactive. | DESIGN: Process Design | Sunquest sent an Urgent - Product Safety Notice letter dated August 12, 2011 to all affected customers. The letter identified the affected product, summary of the problem, actions to eliminate the problem and actions to be taken by clients. Sunquest will distribute a new maintenance release.Client Support can be reached at 1-877-239-0337 for U.S and Canadian sites and at +44 (0) 101-335-0562 for international sites. |  | 47 | 47 sites | US Nationwide Distribution | Terminated | September 04, 2012 | 14 | 47 | Computer
Z-2239-2012 | 62892 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Positron Emission Tomography (PET) and X-ray Computed Tomography, GEMINI TF 16 PET/CT, GEMINI GXL 16 PET/CT, GEMINI TF Big Bore PET/CT, Philips Medical Systems, Cleveland, OH. The device is a diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other organ imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode, dynamic, and gated acquisitions. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a "fused" (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: " The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. " The detection, localization, and staging of tumors and diagnosing cancer patients. " Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of with regional cerebral activity from PET images. | Positron Emission Tomography (PET) and Xray Computed Tomography | 2 | August 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS battery cabinet that appeared to have burned a hole through the casing on one of the batteries. Supplier of the UPS, Chloride/Emerson, determined the batteries, housed within the UPS battery cabinet, entered thermal runaway causing the battery to overheat. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated August 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Field Service Engineers will be implementing a Field Change Order (FCO) to replace the batteries of the affected UPS units. Contact the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 for questions regarding this notice.. | No consumer action necessary | 7 | 7 units | Nationwide Distribution-including the states of CO, FL, MD, MN, NE and PA. | Open |  |  | 7 | Computer
Z-2240-2012 | 62691 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | 892.1650 Image-intensified fluoroscopic x-ray system; the device is designed to provide fluoroscopic and spot film imaging of the patient during diagnostic surgical and interventional procedures. The systems include features specifically designed for use in diagnostic and interventional cardiac imaging procedures, and are also intended for cholangiography, endoscopic, urologic, orthopedic, neurologic, vascular, critical care and emergency room procedures. The systems may be used for other imaging applications at the physician's discretion. | GE OEC 9800 | 2 | August 22, 2012 | 2012 | GE OEC Medical Systems, Inc | GE OEC is recalling certain 9800 Image Intensified Flouroscopic x-ray systems due to customer complaint analysis showing that a replacement Hitachi 160 GB Hard Drive, when installed on certain models is susceptible of causing unanticipated system shut downs, no boots, data loss or data mix, or unexpected hard drive corruption. | DESIGN: Device Design | GE OEC sent a Urgent Recall Notice dated November 6, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification advised them that GE Healthcare had identified a potential safety issue associated with a "Significant increase on Data Loss, Data Mix, and System file corruption of the Hard Drives" on the 9800 Systems.The field correction consisted of installing a "Smart Power Switch" between the power control PCB and the system power switch on 9800 Systems with a Celeron single board computer and software version 29. This will remedy the potential for data loss, data mix, and system file corruption of the hard drives.For further questions please call 800-874-7378 Option 8. |  | 3595 | 3,595 units | Worldwide Distribution -- USA (nationwide) | Terminated | January 28, 2013 | 159 | 3595 | Computer
Z-2244-2012 | 62571 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-armÂŽ Imaging System. Product Numbers: BI-700-00027-100, BI-700-00027-100R, BI-700-00027-120, BI-700-00027-120R, BI-700-00027-230, BI-700-00027-230R. | Medtronic Oarm Imaging System | 2 | August 22, 2012 | 2012 | Medtronic Navigation, Inc. | A fault in a high voltage relay in the 0-armÂż Imaging System might cause the relay to overheat and burn. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Medtronic Navigation issued an "Urgent Field Safety Notice" to consignees dated 7/11/12. The problem was described and recommended actions were provided. For each affected O-armÂŽ Imaging System, a Medtronic service representative will visit the site and replace the 500VDC relay on the system and test the O-armÂŽ Imaging System. | No consumer action necessary | 17 | 17 units | Worldwide Distribution - USA, including the states of AL, CO, FL, IA, MI, IA, NC, NE, NY and the countries of South Africa, Italy, and Spain. | Open |  |  | 17 | Computer
Z-2238-2012 | 62567 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging system.Product Usage:Horizon Medical Imaging is a medical image and information management application that is intended to receive, transmit store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from various medical image systems. The medical modalities of these medical imaging systems include, but are not limited to, all modalities supported by ACR/NEMA DICOM 3.0 (specifically including mammographic images). Horizon Medical Imaging is intended to function with a variety of storage systems and printers via DICOM and other computer industry standard interfaces and protocols. Horizon Medical Imaging is indicated for use by trained medical professionals including, but not limited to, radiologist, physicians, and medical technologists. Horizon Medical Imaging is also indicated for use in soft copy diagnostic interpretation of medical images and video by physicians trained in such practice (specifically including soft copy diagnostic interpretation of mammographic images). | Horizon Medical Imaging system. | 2 | August 22, 2012 | 2012 | Mckesson Medical Imaging Group | There is a potential to save reports in RadReport with an Incorrect Final Status. | DESIGN: Software Design | McKesson sent an Urgent Field Safety Notice letter dated May 29, 2012 to all affected consignees. The letter identified the affected product, circumstances under which the issue occurs and immediate actions to be taken. Customers were advised to notify the Customer Support department that they have read and understood this Advisory Notice at 1-800-663-2533. For questions contact the Support Manager at the Customer Support department. |  | 516 | 516 units | Worldwide Distribution - U.S. Nationwide and the countries of: Canada, France, Ireland and United Kingdom. | Terminated | September 12, 2013 | 386 | 516 | Computer
Z-2243-2012 | 62751 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest EncompassThe reporting site is using the Encompass software in an acute patient care setting. | Sunquest Encompass | 3 | August 22, 2012 | 2012 | Sunquest Information Systems, Inc. | Sunquest is recalling the Sunquest Encompass versions 2.4 and later because the specimen collect time is displaying and printing one hour later than the actual collect time only when Daylight Saving Time is active. | DESIGN: Software Design | The firm, Sunquest, sent a "Urgent-Product Safety Notice: PSN-09-01" letter dated April 13, 2009 all their customers who purchase the Sunquest Encompass versions 2.4 and later. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to request software correction and to create a service request via client support webpage at www.sunquestinfo.com/supportweb. Then choose Submit a Service Request. Select Issue Type and select Software Request. When Software Request is selected, the Work Queue field automatically fills in with "Software Request." Promptly advise appropriate personnel.Customers with questions can contact Client Support and Services at (877) 239-6337 for US and Canadian sites and at +44(0)0800-960-972 for international sites. | Contact the recalling firm for information | 23 | 23 clients | Worldwide distribution: USA (nationwide) and country of: Bahamas. | Terminated | August 22, 2012 | 0 | 23 | Computer
Z-2252-2012 | 62796 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AGFA IMPAC CV DICOMStore with Media Purge Daemon (MPD) and IMPAX CV DICOMStore with Cardiovascular Purge Services (CPS), Medical Image Storage Device | AGFA DICOMStore | 2 | August 22, 2012 | 2012 | AGFA Corp. | Loss of patient data can occur under certain circumstances due to misconfiguration of DICOMstore configured with Agfa's Medical Image Storage Devices - Media Purge Daemon (MPD) and Cardiovascular Purge Service (CPS). | DESIGN: Software Design | On 8/9/2012 AGFA Healthcare issued an "URGENT FIELD SAFETY NOTICE" letter to consignees. This letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood was also required from the consignees. | No consumer action necessary | 535 | 535 | USA Nationwide, Virgin Islands and Canada | Open |  |  | 535 | Computer
Z-2242-2012 | 62460 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatic System is included with the Dimension Vista System. The Dimension Vista Integrated System is an in vitro diagnostic device intended tot duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids.The EasyLink Informatics System is classified as a data calculator and processing module for use with in vitro diagnostic devices. The EasyLink is included with the Dimension Vista System and may be used with other Siemens analyzers as a communications and connectivity workstation for integration with laboratory information system (LIS) networks. | EasyLink(TM) Informatics System included with the Dimension Vista System | 2 | August 22, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the EasyLink Quality Control Package the system may not perform as intended for Multi-Rule QC violations [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions send with the Urgent Field Safety Notice. When this situation occurs, QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient | DESIGN: Process Design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was attached for customers to complete and return via fax to 302-631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 637 | 637 | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, Germany, Spain, Portugal, Italy, Japan, Norway, Australia, Austria, New Zealand, Belgium, Switzerland, Netherlands, Malaysia, Slovenia, South Korea, Denmark, Saudi Arabia, Slovakia, Taiwan, Philippines, Czech Republic, and Singapore. | Open |  |  | 637 | Computer
Z-2257-2012 | 62706 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxI 800 Access Immunoassay Systems, Part Numbers: Dxl 800: 937100; Dxl 800 Remanufactured: A25288;Dxl 800 with spot B: A71456;Dxl 800 Reconditioned: A25285.The DxI Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | UniCel DxI and DXC Immunoassay Systems | 2 | August 23, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received seven reports of a software anomaly for the UniCel Dxl Immunoassay Systems. If undetected, this anomaly may cause erroneous test results. | DESIGN: Software Design | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 with attached PCA Response Form to all customers who purchased the UniCel Dxl Immunoassay Systems and DxC Clinical Systems which includes all software versions. The letter describes the product, problem and actions to be taken. The customers were instructed verify the system temperatures immediately following each initialization; do not begin processing samples until this verification is complete; refer to the appendix included with the letter for detailed instructions about verifying system temperatures; review their quality control (QC) performance history; and share this information with their staff and retain this information as part of their quality system documentation, and to complete and return the enclosed response form within 10 days to Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318; Attn: Regulatory Affairs or fax to: 1-786-639-4000. Customers with any questions regarding the notice, were instructed to contact Customer Support Center via our website, http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representatives. | Contact the recalling firm for information | 2442 | 2442 units total (784 units in US) | Worldwide distribution: USA (nationwide) and countries of: Algeria, Angola, Argentina, Australia, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea,Republic of; Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | August 23, 2012 | 0 | 2442 | Computer
Z-2254-2012 | 62870 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon MRI System/Oasis MRI, Software. These are software controlled MRI devices.Product Usage:The Echelon and Oasis MR systems are imaging devices, intended to provide the physician with physiological and clinical information that can be useful in diagnosis determination. | Echelon and Oasis | 2 | August 23, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error with a feature called Multi-Planar Reconstruction(MPR). MPR takes an original volume dataset reconstructed in any plane view and creates a new image set from another view (i.e. axial to coronal, coronal to sagittal, etc.). The error causes the first image in the dataset to have a right/left reversed orientation. The software error can occur randomly when execu | DESIGN: Software Design | Hitachi sent a Device Correction Notice letter dated June 18, 2012 to all affected customers. All letters are recorded by tracking number and date received by the consignee. The letter identified the affect product, problem and actions to be taken. All sites will be visited by Hitachi Service to install the software update to permanently resolve the issue at no cost to the facility. Software correction is targeted to start by September 1, 2012. For question contact the Manager of Regulatory Affairs at 800-800-3106 x3720. | Contact the recalling firm for information | 165 | 165 units | US Nationwide Distribution - including the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY and Mexico. | Terminated | February 19, 2013 | 180 | 165 | Computer
Z-2266-2012 | 62851 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba CT systems, Asteion and Aquilion, that are equipped with a PC based computer, or to those older systems (SGI based) that received an improved table by an upgrade purchase. This device is designed to produce cross-sectional images of a human body by reconstruction of xray transmission data from the same axial plane taken at different angles. These images have been proven to be clinically useful in the diagnosis of spine and head injuries, intracranial tumors, blood clots in the brain, eye trauma, soft tissue lesions in the extremities, gastrointestinal lesions, abdominal and pelvic malignancies, and hepatic metastases. CT is also used to evaluate intestinal obstruction!;, assess intra-abdominal abnormalities and to examine musculoskeletal degeneration. This device employs no intended uses that are not in cleared devices already found in the marketplace. | Scanner, Computed Tomography, XRay | 2 | August 24, 2012 | 2012 | Toshiba American Medical Systems Inc | The systems all have a similar table control board. It hasbeen found that this board may cause the table position to be incorrect if the table cradle is movedmanually. This is referred to as "Free" mode in the operation manuals. If this error does occur it ispossible that the patient could be scanned at an incorrect position. | DESIGN: Process Design | Toshiba sent a "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to please avoid moving the table in the "Free" or manual mode as much as possible. If it is necessary to move the table in this manner please verify the table position prior to initiating the scan sequence.If it is noted that the position is not correct, exit the eXam plan and start a new one beginning with the scanoscopy. Please share this information with all users of the equipment and your clinical or biomedical engineering departments.Please find the attached customer reply form. The purpose of this form is for FDA tracking. Please take a moment to complete the form and fax it to the toll free number on the form.". The contact information for this recall is:"(800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968.". |  | 757 | 757 | USA ( Nationwide Distribution ) | Terminated | September 04, 2012 | 11 | 757 | Computer
Z-2268-2012 | 62881 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Mini-Infuser Model 150XL Single Speed Infusion Pump, an Rx battery operated syringe infusion pump; Product Usage: These product codes includes refurbished units. Intended for the controlled rate delivery of small volume parental fluids as prescribed by a physician. Usage: The Mini-Infuser System is designed for controlled intermittent administration of intravenous medication. | MiniInfuser Model 150XL Infusion Pump | 3 | August 24, 2012 | 2012 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the Mini-Infuser will not operate. The reason being that the Mini-Infuser Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Baxter sent a Safety Alert letter on January 11, 2011 for the InfusO.R. Pump. Baxter expanded the scope of the recall to include the Mini-Infusers System and sent an Urgent Device Correction letter dated March 4, 2011, via First Class mail to all affected customers. The letters informed the customers that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The customers were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future reference. For Technical questions call Baxter Global Technical Services at 1-800-843-7867. For question regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 26485 | 26,485 units | 45,865 units | Terminated | August 27, 2012 | 3 | 45865 | Computer
Z-2263-2012 | 62770 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Cisco brand HealthPresence, version 2.1; Model/Catalog Number: CHP-ENTPRSYR2.I-K9; Product is manufactured and distributed by Cisco Systems, Inc. San Jose, CACisco HealthPresence, version 2 is intended for remote consultation of patients with a healthcare provider. Cisco HealthPresence enables remote collaboration between a physician and a patient by providing: high-definition video and advanced audio; transmission and display of third-party medical device data; integration with third party EMR: scheduling: and other software applications for remote consultations and exams. Cisco Health Presence is not intended for use in emergency situations. Cisco HealthPresence is not intended for use in active patient monitoring or alarming. | Cisco | 2 | August 24, 2012 | 2012 | Cisco Systems Inc | Software anomaly may result in incorrect values and interpretations. The initial vital signs reading will display accurately on the Cisco Health Presence Vitals tab and the OnePlace tab. However, subsequent vital signs readings will only update on the Cisco HealthPresence Vitals tab and within the OnePlace database, but will not update under the OnePlace tab, and thus, will not provide current vit | DESIGN: Software Design | Cisco notified their one customer by phone on July 25, 2012, with a follow up email and instructed the customer to stop using the OnePlace vitals tab and to use the CHP only. On July 27th, 2012, the vendor of the OnePlace application, fixed the bug remotely.On July 31, 2012, Cisco's staff physician contacted the customer to confirm that the customer was aware of the fix. Validation of the correction was completed by Cisco, along with customer on August 9, 2012. Until validation was satisfactorily completed, the integrated system lie dormant at customer's site. |  | 1 | 1 unit | Nationwide Distribution including the state of TN. | Terminated | August 24, 2012 | 0 | 1 | Computer
Z-2264-2012 | 62678 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS.Radiological image processing system. | syngo Imaging XS | 2 | August 24, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo Imaging XS, version VA70A or higher. In some situations it can happen that a miscalculation is done for the grey scale values in the functions "Regions of Interest", "Pixel Lens", "Edge Enhancement" and "Histograms". | DESIGN: Software Design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated July 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. For questions concerning this notice contact your local Siemens Service Rep.. |  | 52 | 52 | Nationwide Distribution-CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Open |  |  | 52 | Computer
Z-2275-2012 | 62679 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza, Model number 10592457.Intended use: radiological image processing system. | syngo.plaza | 2 | August 27, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in which a miscalculation occurs for images with pixel depth values greater that 12 bit in certain functions. | DESIGN: Software Design | The firm initiated their recall/correction of this device on July 18, 2012 by sending a "Customer Safety Advisory Notice" letter, dated July 12, 2012, to their consignees. The letter identified the affected product along with the malfunction. It also discussed the steps the user can take to avoid the issue. Users were asked to not use the mentioned functions for images from those modalities until the issue was resolved and the system was updated with the software version containing the solution. Users were to contact their local Siemens Service in order to plan for the updated software version that will resolve the issue. Additionally, users should place the safety advisory notice in the system's instructions for use, and personnel should be instructed accordingly and maintain awareness. If the device was sold and is no longer in the users' possession, the notice should be forwarded to the new owner and the firm should be notified of the new owner. |  | 14 | 14 | Nationwide Distribution -- CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 24, 2014 | 515 | 14 | Computer
Z-2279-2012 | 62880 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series Patient Monitors; Mindray DS USA, Inc.Product Usage:The V Series Monitor is intended for intra hospital use under the direct supervision of a licensed healthcare practitioner. The indications for Use for the V Series include the monitoring of the following human physiological parameters: ECG waveform derived from 3, 5, 6 and 12 lead measurements; Heart Rate; Pulse Oximetry (SpO2); ST Segment Analysis; Arrhythmia Detection; Non Invasive Blood Pressure (NIBP); Invasive Blood Pressure (IBP); Cardiac Output (CO); Respiratory Gasses; Respiration Rate; and Temperature. The V Series Monitor has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. The target populations are adult, pediatric and neonate with the exception of: Arrhythmia detection and ST Segment Analysis, for which the target populations are adult and pediatric only; IV Drug Calculations for which the target population is adult only; and Cardiac Output for which the target population is adult and pediatric only. | V Series Monitor | 2 | August 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified two software anomalies contained in the V Series Monitoring System. There have been no reports of injuries associated with these anomalies. These anomalies affect V Series Systems distributed between September 31, 2010 and January 9, 2012. | DESIGN: Software Design (Manufacturing Process) | Mindray sent an Urgent: V Series Monitoring System Correction letters dated May 7, 2012 via certified mail to all affect customer. The letter identified the affected product, problem and actions to be taken. Customers are directed that VPS modules should not be used between systems until a software upgrade has been completed. This work will be done at no cost to the customer. For questions contact your Mindray Serviice Representative at 1-800-288-2121. | Contact the recalling firm for information | 979 | 979 units | Worldwide Distribution - US Nationwide including the states of: NY, SD, NE, IL, MO, CA, OH, OK, WA, IN, WY, CO, TX, FL, NC, AL and Puerto Rico. and the countries of: Australia, Canada, Columbia, New Zealand, Pakistan, Russia, Saudi Arabia, United Kingdom and Venezuela. | Terminated | July 17, 2013 | 323 | 979 | Computer
Z-2282-2012 | 53055 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole Body X--ray ScannerThis device is indicated as a whole body volume CT system. The system produces volumes of up to 320 axial slices in a single rotation. | Aquilion One CT System, TSX301A/2 | 2 | August 28, 2012 | 2012 | Toshiba American Medical Systems Inc | 1. It has been found that in raw data processing for Go&Return; helical scanning, when the raw data acquired in the first scan (Go) and that acquired in the second scan (Return) are selected and reconstructed at one time, incorrect couch positions and scan times are displayed on the reconstructed images for the data acquired by Return scanning, while all related information is displayed on the reco | DESIGN: Software Design | The firm, Toshiba America Medical Systems, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 30, 2009 to its Consignees/customers. The letter described the product, problem and actions to be taken. The letter informed them of the above mentioned reason for recall and telling them that a corrective measure will be installed to prevent the occurrence of the above problems. They were told that their Toshiba service representative will contact them to install the measures.Furthermore Toshiba instructed them to adopt the following temporary measures to avoid these phenomena until corrective measures have been taken.1. When performing raw data processing for Go&Return; helical scan, select the raw data acquired for Go and that acquired for Return separately in order to perform reconstruction2. When using the Sure Exposure 3D function, confirm that the error message shown in the preceding page is not displayed on the monitor at the time of setup (until the Scan Start button lights after the Confirm button is clicked).3. When executing an eXam Plan in which Sure Exposure 3D is selected as the actual scan for Sure Start, confirm that the error message show in the preceding page is not displayed on the monitor when monitoring scan is switched to the actual scan. If the error message is displayed, click [OK] in the error message dialog to continue processing.The letter should be shared with facility's Biomedical or Clinical Engineering Department. Also the customers were instructed to complete and return the Customer Reply Form via fax to: Regulatory Affairs at 877-349-3054.If you have any questions regarding this letter, contact Director, Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968 with any questions regarding the letter. |  | 24 | 24 Systems | Nationwide distribution: USA including states of: AL, AR, CA, CO, FL, GA, KY, MA, MD, MT, ND, NJ, NV, NY, OH, PA and Puerto Rico. | Terminated | September 04, 2012 | 7 | 24 | Computer
Z-2278-2012 | 62853 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | EXCELART Vantage, Atlas, and Titan MRI Systems, MRT-1503 and MRT-1504.Imaging of the Whole Body, Fluid Visualization, 2D/3D Imaging, MR Angiography/MR Vascular Imaging, Blood Oxygenation Level Dependent Imaging, Perfusion I Diffusion Imaging, and Proton Spectroscopy. | EXCELART Vantage, Atlas, and Titan MRl Systems MRT.1503 and MRT1504 | 2 | August 28, 2012 | 2012 | Toshiba American Medical Systems Inc | The status display in the sequence queue window may not change from "Current" to "Done". | DESIGN: Process Design | A recall notification letter titled: "URGENT: MEDICAL DEVICE CORRECTION," was sent to customers. The letter stated the problem, corrective action, and (instructions) request to customers. Customers were instructed to perform the temporary recovery procedures if the problem occurs. Additionally, the acknowledgement sheet should be completed and faxed back. The information should be shared with all users, including the bio-medical or clinical engineering departments. Questions pertaining to the information for this recall were to be directed to (800) 421-1968. | Contact the recalling firm for information | 54 | 54 | Nationwide Distribution | Terminated | September 04, 2012 | 7 | 54 | Computer
Z-2305-2012 | 62709 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo(R) Lab Data Manager and syngo(R) Lab Process Manager, all versions.This device is classified as a data calculator and processing module for use with In-vitro diagnostic devices. It also provides work flow management as well as remote instrument control to multiple interfaced instruments. | syngo(R) Lab Data Manager and syngo(R) Lab Process Manager, all versions. | 2 | August 31, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the syngo Quality Control package the system may not perform as intended for Multi-Rule QC violation [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions on page 2 of the recall notification. QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient results will not be held as expecte | PRODUCTION CONTROLS: Packaging Process Control | Siemens sent a Urgent Field Safety Notice letter dated May, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customers to re-set the Multi-Rule option and instructions are provided on how to do so. This recall notification also included a Field Safety Effectiveness check form to please fax to the Siemens Technical Solutions Center at 302-631-8467 BY June 15th. | No consumer action necessary | 44 | 44 | USA ( nationwide ) including the states of AZ, CA, FL, IL, KY, MD, MA, MN, MS, MO, NE, NH, NJ, NY, NC, OH, PA, TX, VA and WV. | Open |  |  | 44 | Computer
Z-2312-2012 | 62982 | probe, uptake, nuclear | IZD | Nuclear uptake probe. | Radiology | Radiology | 510(K) Exempt | Capintec Captus 3000 Thyroid Uptake System Nuclear uptake probe intended to be used by trained Nuclear Medicine Technologists to perform thyroid uptake procedures, which consist of the measurement of organ uptake and circulating radioactivity. | Capintec Captus 3000 Thyroid Uptake System | 2 | September 05, 2012 | 2012 | Capintec Inc | Customers who received a Capintec Captus 3000 Thyroid Uptake System with Software Revisions 1.22, 1.23, 1.24, 1.25, and 1.26 may encounter an error while using the thyroid uptake protocol ?Measure Liquid and Subtract Residual." This protocol was a new feature added to Software Revision 1.22 in January 2011. The error occurs only if the user does NOT measure the residual dose immediately after me | DESIGN: Software Design | The firm, CAPINTEC, INC., sent a "SOFTWARE ERROR ALERT" letter dated July 20, 2012 to all affected consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the four measuring steps outlined in the letter; DO NOT PRINT, EXIT THE THYROID UPTAKE MODULE, OR MEASURE ANOTHER PATIENT BEFORE YOU HAVE COMPLETED THESE FOUR MEASUREMENT STEPS; and follow the attached addendum for step-by step instructions. Capintec will send corrected software to all affected customers in approximately 2-3 weeks. The customers are instructed to contact Capintec Customer Support Representative if you have any questions or concerns regarding this notification, or if you do not receive your replacement software. Capintec Customer Support can be reached by phone at 1-800-631-3826 or by email at getinfo@capintec.com or by visiting our website at www.capintec.com. |  | 256 | 256 | Worldwide distribution: USA (nationwide) and countries of: Argentina, Belgium, Brazil, Canada, China, Greece, Hong Kong, India, Israel, Malaysia, Philippines, Spain, South Korea, Thailand, and Turkey. | Terminated | December 10, 2013 | 461 | 256 | Computer
Z-2321-2012 | 62891 | subsystem, water purification | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AmeriWater MRO Dialysis RO+ System, AmeriWater 1303 Stanley Avenue, Dayton, OH 45404 (800) 535-5585 www.AMERIWATER.com.The AmeriWater MRO Portable Reverse Osmosis Systems are water treatment systems intended for use in hemodialysis applications. They are designed to pre-treat and purify potable water for use in making dialysate for hemodialysis and to meet current AAMI and Federal (U.S.) standards. The AmeriWater Portable MROS model is intended for use in a hospital, clinic, dialysis center, or for home care for single patient use. The AmeriWater Portable MRO1 model is for treatment of up to two patients in a hospital, clinic, or dialysis centers. | AmeriWater MRO Dialysis RO System | 2 | September 06, 2012 | 2012 | AmeriWater Inc | The firm was notified that a malfunction of the device was caused by an incorrect fuse 32 volt 15 amp fuses instead of 250 volt 15 amp fuses, that was installed in the water system. The incorrect fuse may result in excessive heat in the fuse holder eventually resulting in failure of the device to operate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, AmeriWater, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2012, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; remove affected devices from service until the correction is completed to prevent malfunction of the device during patient treatment; and complete and return the attached response form to AmeriWater within 30 days via fax to 937/461-1988 Attn: Quality Assurance. AmeriWater dealer will be contacting their customers to schedule replacement of the fuses.If you have any questions, contact your local AmeriWater dealer or AmeriWater customer service at 1-800-535-5585. |  | 80 | 80 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, CO, CT, FL, GA, IL, IN, KS, MA, MD, MO, MS, NC, NY, OH, OK, SC, TN, TX and VA; and country of: Canada. | Open |  |  | 80 | Computer
Z-2313-2012 | 62702 | test,natriuretic peptide | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Multi-Diluent 1 - Used for automated dilutions for the ADVIA Centaur BNP and TSH3 Ultra assays when on-board the ADVIA Centaur or ADVIA Centaur XP systems.Catalog Numbers: 2 Pack REF 07907174 (110312) 6 Pack REF 07293184 (110313).Used for automated dilutions for the ADVIA Centaur BNP and TSH3 Ultra assays when on-board the ADVIA Centaur or ADVIA Centaur XP systems. | Siemens ADVIA Centaur MultiDiluent 1 | 2 | September 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Decrease in onboard dilution recovery when using Multi-Diluent 1 that has been stored on board the ADVIA Centaur and ADVIA Centaur XP systems for for BNP and TSH3 Ultra assays | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL NOTICE" dated July 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Response Form was included for customers to complete and return via fax to (302) 631-7597. Contact your local Sales Representative for questions regarding this notice. |  | 23301 | 2 Packs = 15879 and 6 Packs = 7414 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Australia, Brazil, Chile, China, Croatia, Denmark, Ecuduar, Egypt,Guatemala, Hong Kong,India,Ireland, Israel,India, Japan,Korea, Mexico, Malyasia, Peru, Pakistan, Saudi Arabia,Singapore, Thailand,Uruguay, Viet Nam, and South Africa. | Open |  |  | 23301 | Computer
Z-2323-2012 | 62757 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | VersaTREK Windows Software, Version 5.4.3 is a component of the VersaTrek System. The Software CD will be individually packaged using a standard CD plastic case and placed into a standard CD mailerThe VersaTREK System is for cultivating and recovering microorganisms, especially bacteria and yeasts, from blood and other normally sterile body fluids. | VersaTREK Windows Software, Version 5.4.3 | 3 | September 06, 2012 | 2012 | Trek Diagnostic Systems | Review of instrument software databases and investigation of two customer reports determined that on rare occasions (less than 0.3%), following the removal of a positive signaling bottle from the VersaTREK instrument, the positive test result inadvertently reverted to a negative result in the instrument database. | DESIGN: Software Design | ThermoFisher Scientific sent an Important Medical Device Product Correction Notice dated July 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact the Technical Service Department at 800-642-7029 to schedule their upgrade and to complete and return the Correction Notice Customer Acknowledgement by fax or in the enclosed return envelope. Customers were also asked to return the checklist. For questions customers were instructed to call 800-642-7029.For questions regarding this recall call 913-985-4185. | Contact the recalling firm for information | 333 | 333 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO , CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SD, TN, TX, VA, VT, WA, WI, WV, and WY. Internationally to the countries of Bahamas, Bangladesh, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, India, Ireland, Italy, Japan, Jordan, Kenya, Kuwait, Malaysia, Mexico, Pakistan, Peru, Philippines, Portugal, Puerto Rico, Russian Federation, South Africa, Spain, Thailand, Turkey, and United Kingdom | Terminated | June 18, 2013 | 285 | 333 | Computer
Z-2324-2012 | 62967 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Brand name: RayStation. Generic name: Treatment Planning System. Software version and build numbers are 2.5.0.144 and 2.5.1.89. Product lifetime is 5 years.RayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment, and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system. | Radiation Therapy Treatment Planning System | 2 | September 06, 2012 | 2012 | RAYSEARCH LABORATORIES AB | The firm decided to recall when an issue was found with the dose calculation in RayStation 2.5, Raystation 2.5.1. The problem has not caused any patient mistreatment or other incidents. However, the user must be aware of this information to avoid incorrect dose calculations during treatment planning. The dose calculation is not correctly normalized in the unusual case of preventing the optimizatio | DESIGN: Device Design | The firm, RaySearch Laboratories, sent an "Urgent Field Safety Notice, Medical Device Correction" letter dated May 9, 2012 to its customers. The letter describes the product, problem and actions to be taken. The letter informed the customers that the "workaround is to not use segment MU optimization" and that there "is in general no need to manually adjust the segment shapes to such an extent there is need for a separate optimization of the segment MU." Additionally, the customers were instructed -Do not uncheck the "Segment shapes" check box; educate planning staff and all users about this workaround; inspect your product and identify all installed units with the noted software version numbers; and complete and return the Reply Form via fax to: RaySearch Americas Inc. at 888-501-7195; email: freddie.cardel@raysearchlabs.com and/or email to: Raysearch Laboratories AB at niclas.borglund@raysearchlabs.com.For further information and return of the Reply Form contact For the Americas: Director of Customer Support, RaySearch Americas at +1877 778 3849 or freddie.cardel@raysearchlabs.com; and For Europe, Asia and rest of the world: Director of Service, RaySearch Laboratories AB at +46 8 5450 6130 or niclas.borglund@raysearchlabs.com. | No consumer action necessary | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: FL, MA, NJ, and NY; and countries of: Canada Germany, Netherlands and Switzerland. | Terminated | January 25, 2013 | 141 | 19 | Computer
Z-2332-2012 | 62768 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest LaboratorySunquest Laboratory is intended for use by professionals working in a clinical laboratory. Sunquest Laboratory is intended for use only by professionals who have received extensive training in the use of the software. Sunquest Laboratory is controlled via userdefined security levels, which allow the laboratories to manage data security and integrity. The system is marketed as an automated solution for managing laboratory processes, including: " Patient registration " Order entry and order modification " Specimen collection, verification, suitability, and distribution " Quality assurance checking " Workload recording " Billing charge capture " Standards and controls recording " Test order result inquiry and reporting Organism susceptibility and epidemiology records " Microbiology culture direct examination and observation recording " Outreach clients, specimens, and results " Client-specific reporting capabilities " Inbound client service request tracking " Historical data Clients that use the Sunquest Blood Bank system should refer to the intended use statement found in the Blood Bank System Overview Users Guide. | Sunquest Laboratory | 2 | September 07, 2012 | 2012 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1 because when the Windows-based merge application is used and the "Event merge across patients" option is used, a receiving system (for example a HIS or CoPath system) may interpret the event merge as a full patient merge. This occurs because a trigger from the Laboratory system for the event | DESIGN: Software Design | The firm, Sunquest, sent an "Urgent- Product Safety Notice: PSN-12-02" dated March 15, 2012 to all their customers who purchased the Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1. The notice described the product, the problems and the actions to be taken. Customers were instructed to request the software correction by creating a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb. Customers with questions were instructed to contact Client support at (877) 239-6337 for US and Canadian Sites and at +44 (0) 161 335 0562 for international sites. |  | 126 | 126 sites | Worldwide distribution: USA (nationwide) and countries of: England, United Kingdom and United Arab Emirates. | Terminated | October 02, 2012 | 25 | 126 | Computer
Z-2347-2012 | 62328 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx monitor/defibrillatorsModels: M3535A and M3536A, M3536J, M3536M, M3536MC, M3536M1, M3536M2, M3536M3, M3536M4, M3536M5, M3536M6The Heartstart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips HeartStart MRx monitor/defibrillators | 2 | September 10, 2012 | 2012 | Philips Healthcare Inc. | Therapy Cable Connection wear may lead to a malfunction in detection of defibrillation pads/paddles therapy Cables | DESIGN: Device Design | Philips notified customers on June 20, 2012, by an Urgent Medical Device Correction notification/Field Safety Notice and a revised Instructions for Use addendum for each HeartStart MRx unit. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of units affected and and identifies the action. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Carefully read the attached HeartStart MRx Instructions for Use Addendum. Immediately implement ongoing therapy connection inspection on all of your MRx devices to detect wear from higher than expected levels of stress. The HeartStart MRx Instructions for Use Addendum describes how a user can identify wear. If wear is detected, remove affected devices immediately from use and contact Philips to arrange for service.WARNING: The service life of your therapy cables/external paddles is up to three years. To maintain reliable performance and reduce the possibility of failure during patient use, replace them every three years from the time they were initially placed into service or if they fail the inspection criteria in the attached Instructions for Use AddendumThe addendum provides the customer with complete user information regarding ongoing inspections of port. In addition, service support will be provided for Therapy Port and Cables requirement replacements.Further information or support concerning this issue,contact your local Philips representative at 1-800-722-9377.For questions regarding this recall calll 978-687-1501. | No consumer action necessary | 58701 | 58,701 units | Worldwide Distribution - USA (nationwide) and the countries of : AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Open |  |  | 58701 | Computer
Z-2344-2012 | 63008 | stimulator, electrical, evoked response | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Interfaces subject to recall: Medtronic NIM (Nerve Integrity Monitor) 3.0 Patient Interfaces REF 8253210 (Pulse).LABELING: Box Label: REF 8253210 x1***NIM-Response 3.0 Patient Interface***Medtronic Xomed, Inc., 6743 Southpoint Drive North Jacksonville, FL 32216-0980 USA***Rx Only***Lot***SN:***"Top of unit is labeled in part: "***NIM-Response 3.0***Medtronic***" Bottom/Underside of unit is labeled in part: "***5X20mm FA.L., 32mA. 250VAC***REF***MEDTRONIC XOMED INC. JACKSONVILLE, FL USA***" Cord is labeled in part: "***Ref.***Mfg Date***Sn***"The NIM 3.0 Patient Interfaces [REF #'s 8253210 (Pulse), 8253200 (Response), 8253410 (Neuro)] is an EMG Nerve Integrity Monitor intended for locating and identifying cranial and peripheral motor and mixed motor-sensory nerves during surgery, including spinal cord and spinal nerve roots. | NIM (Nerve Integrity Monitor) 3.0 Patient Interfaces Pulse;Response;Neuro | 2 | September 10, 2012 | 2012 | Medtronic Xomed, Inc. | In January 2010, Medtronic issued a recall of Medtronic NIM 3.0 Patient Interfaces after Medtronic identified a trend of open channel conditions (signal loss from the electrode). When the NIM 3.0 detects the signal loss from the electrode (i.e. open channel), it fails the electrode check and provides an audible and visual warning to the user which could also create a delay in the surgical procedu | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Medtronic USA, Inc., sent an "URGENT Product Recall Notification" letter dated September 27, 2010, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify their receipt of this notice and complete and sign the Product Replacement checklist form indicating that their sales representative has completed the replacement of their devices and fax to 1-651-367-8055.U.S. Customer Units - replacement kits with their contents will be provided to each Area Sales Manager (ASM) for each customer along with documentation of serial number units to be replaced. Each ASM will be responsible for delivering a replacement unit(s) and picking up and returning each of their customer's units using the pre-labeled replacement kit boxes.International Customer Units- International units, depending on the geography will either be replaced or repaired and returned. If replacement is planned, the International Business Manager (IBM) will receive replacement kits as described for the U.S. In those geographies where devices will be repaired and returned, kits with the following contents will be sent to IBM's along with documentation of serial number units to be repaired.Foreign distribution: In the larger territories (i.e. Australia, Japan, Canada, EMEA-Europe, Middle East, Africa), repairs/rework were handled at the Medtronic facilities in those geographic areas. In the smaller geographic areas (Singapore, Thailand, Hong Kong, Korea, Latin America ), where there were limited sales, devices were returned to the U.S. In all cases (large and small geographic areas), devices were returned via a Medtronic entity (sales office or distribution center).If you have any questions regarding the subject Actron or the content of this letter, contact the Distinguished Regulatory Affairs Advisor at 904-279-7532. |  | 7 | 7 | Worldwide distribution: USA (nationwide) and countries of: Australia, Austria, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Lebanon, Libya, Luxembourg, Netherlands, New Zealand, Norway, Portugal, Saudia Arabia, Singapore, Slovak, Spain, SRI Lanka, Sweden, Turkey, United Arab Emirates, and United Kingdom. | Terminated | November 26, 2013 | 442 | 524 | Computer
Z-2345-2012 | 63102 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano brand Intravascular Ultrasound, Volcano s5/s5i Family Software Version 3.2.1 or 3.2.2, Models/Part Numbers: 435-0607.01, 435-0607.04, 435-0602.01, 435-0602.02; Product is manufactured and distributed by Volcano Corporation, Rancho Cordova, CA.Medical device for use in imaging of histology. | Volcano brand Intravascular Ultrasound | 2 | September 10, 2012 | 2012 | Volcano Corporation | The display on the system has the potential to change. In house testing eventually identified the unique sequence of events that must occur in order for this issue to manifest itself. | DESIGN: Software Design | Volcano Corporation sent an "IMPORTANT FIELD CORRECTIVE ACTION" letter dated August 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Volcano sales representative or Volcano Customer Service for questions regarding this recall. | Contact the recalling firm for information | 1020 | 1020 units | Worldwide Distribution. | Open |  |  | 1020 | Computer
Z-2376-2012 | 62760 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2121IQ, Innova 3131IQ Cardiovascular X-ray imaging systems.K060259: The Innova 3131IQ (3131-IQ) and 2121 IQ (2121-IQ) biplane systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational angiography procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic, and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. These devices are not intended for mammography applications. K061163: For Innova 2121IQ, Innova 3131IQ, Innova 4100, Innova 4100IQ, Innova 3100, Innova 3100IQ, Innova 2100IQ devices with Innova IVUS option: The Innova systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational imaging procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic, and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. Those devices are not intended for mammography applications. Innova IVUS Option: The Innova IVUS software option simplifies the clinical workflow associated with the use of Volcano IVUS products by: (1) automatically synchronizing the patient demographic information (patient name, date of birth, DICOM attributes etc.) from Innova system with an IVUS imaging system, (2) providing a remote access to commonly used IVUS functions from the Innova table side user interface, (3) displaying the IVUS images on the multi-monitor display of the Innova cathlab system. K091658: The Innova systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally, rotational imaging procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic, and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. These devices are not intended for mammography applications. | GE Healthcare, Innova 2121IQ, Innova 3131IQ Cardiovascular Xray imaging systems. | 2 | September 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Innova 2121IQ & Innova 3131IQ Systems. While performing a fluoroscopic acquisition with the Innova 2121IQ or Innova 3131IQ systems, there is a potential for loss of usable imaging mode which requires the system to restart in order to recover. | DESIGN: Software Design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 16, 2012, to all affected consignees. The letter described the Safety Issue, safety Instructions, Affected Product Details, Product Correction, and Contact Information. Consignees were instructed that hospital staff should wait for Photo or Image to appear on the display (when store photo function is performed) before start of the next short duration fluoroscopy acquisition. For questions customers were instructed to call:United States (800) 437-1171Japan: 0120 - 055 - 919For other countries consignees should contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. | Contact the recalling firm for information | 240 | 240 (133 US and 107 OUS) | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, HI, IK, IL IN, IA, KS, LA, MD, MA, MI, MN,MO, NV, NH, NH, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI,.and internationally to:AUSTRALIA, AZERBAIJIN, BAHRAIN, BELARUS, BULGARIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, JAPAN, KUWAIT, LEBANON, LIBYA, MALAYSIA, MEXICO, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAÂŽDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, TAIWAN, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM, and YEMEN. | Open |  |  | 240 | Computer
Z-2256-2012 | 62946 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Optima XR200amx. Mobile general purpose radiographic imaging of the human head and body. | Optima | 2 | September 13, 2012 | 2012 | GE Healthcare, LLC | Software issue: a violation of 21 CFR 1020.31(a). After an image is taken, when the user selects and clears an additionally displayed screen the system inhibit is also cleared, allowing the user to take an exposure without the technique factors being displayed. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | The firm will issue Important Electronic Product Radiation Warning letters that identify the affected product along with a description of the defect and related hazards. Users are instructed to ensure that the acquisition mode is always selected prior to attempting an exposure to confirm the techniques selected. The firm will remedy the issue or bring the device into compliance free of charge. A GE Healthcare Service Representative will update the software on the system to address the issue. Field Modification Instructions (FMI) 10865(2) describes the rework plan the GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. Users are to contact 800-437-1171 in the US, if they have any questions or concerns. |  | 254 | 254 total in the US | Worldwide Distribution -- USA (nationwide) Distribution | Terminated | January 10, 2013 | 119 | 254 | Computer
Z-2377-2012 | 62844 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron Docking Station, a component in the DataCaptor Connectivity SystemCatalog Number DC-NU-DS-CKThe Capsule Neuron Docking Station is a component in the DataCaptor Connectivity System. The Capsule Neuron Docking Station provides connection for the Capsule Neuron UMPC which is a bedside device that enables automatic collection of vital signs data that can connect to the hospital network through a wired Ethernet connection or via a standard 802.11 b/g/n wireless network. | Capsule Neuron Docking Station | 2 | September 13, 2012 | 2012 | Capsule Tech Inc. | Capsule Neuron Docking Station may experience a lost connection, resulting in loss of data or loss of power to the Caspsule Neuron | DESIGN: Device Design | The firm, Capsule Tech, sent a recall letter dated August 10,2012 to its customers. The letter described the product, problem and actions to be taken. Capsule is sending field correction kits to customers with instructions. Capsule's Professional Services team will contact the customers to schedule shipment of their kit. Final shipments of the field correction kits will be completed by Oct 2012. If you have any questions or concerns, contact your professional services representative and/or call in US: 978-482-2340 or France: +33 153-34-14-00. | Contact the recalling firm for information | 12448 | 12,448 units | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Denmark, France, Germany, Netherlands and UK. | Open |  |  | 12448 | Computer
Z-2382-2012 | 63007 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL 91493 Integrated Module Housing.PCIS patient monitors is intended to sound various alarm tones that are generated by the specific PCIS patient parameter modules. Each patient parameter module determines (a) when an alarm condition is violated; (b) the alarm priority (i.e. high, medium or low); (c) alarm limits; and (d) when to generate and terminate alarms. PCIS patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Medical PCIS patient monitors are designed to communicate with a variety of external devices such as display devices, network devices, serial devices, user input devices, audio system, local/remote recorders and Spacelabs Medical patient parameter modules. In all applications, Spacelabs Medical PCIS patient monitors provide access to that data on the Spacelabs Medical network either via hardwire or wireless transmission. Through this network, a patient's data may be displayed at the bedside monitor, central station or clinical information system. | Integrated Module Housing | 2 | September 14, 2012 | 2012 | Spacelabs Healthcare, Llc | There is a potential failure of the Integrated Module Housing causing the patient module parameters to drop off the monitor display or the patient module will fail to sign on when inserted into the housing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Spacelabs Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-522-7025 and select 2 for assistance regarding this recall. |  | 527 | 527 units total; 350 US and 177 International | Worldwide Distribution. | Terminated | April 15, 2013 | 213 | 527 | Computer
Z-2378-2012 | 63110 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The product affected by this notice is sold under the trade name TomoTherapy SharePlan 1.1 . To determine if the version you are using is affected, bring up the About box in the SharePlan application and check if the build number reported there is 1.3.1.10 .Intended use: t-RayAutoplan is a software that, based on the planned dose distribution from a TomoPlan treatment planning system, generates a family of individually optimized IMRT plans and presents their characteristics to the user in a GUI. Among the generated plans, the user selects and clinically approves a treatment plan and exports it electronically to DICOM, for subsequent treatment of the patient. The intended users of t-RayAutoplan shall be clinically qualified staff, such as medical physicists, medical doctors or dosimetrists | RayAutoplan, aka tRayAutoplan, commercial name (TomoTherapy) SharePlan | 2 | September 14, 2012 | 2012 | RAYSEARCH LABORATORIES AB | This field notice concerns two behaviors of the SharePlan (Multiple Fraction Groups and Depth Offset) that may be unexpected and are not clearly described in the user manual. None of these behaviors have caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivere | MISBRANDING: Labeling False and Misleading | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction dated May 30, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consigness were instructed to completely avoid affecting the erroneous behaviour, verify in TomoTherapy Hi-Art that the plans sent to TomoTherapy SharePlan 1.1 only have one single fraction group with a single fraction size,used for all fractions, and do not attempt to import treatment plans having more than one fraction size into the product. To completed avoid affecting the erroneous behaviour the user can either offset the measured curves before importing them into beam commissioning or add the depth offset tothe calibration depth.For further information consigness should contact RaySearch Laboratories AB Quality and Regulatory Affairs at +46 722 366 110 or the Tomo Therapy support at support@tomotherapy.com.North America 1 866 368 4807Belgium 0800 38783France 0805 631 565Germany 0800 000 +541Italy 800-986 399Netherlands 0800 0201364Spain 800 3000049Switzerland 0800 001927UK 0868 238 6035China/Northern 10 800 712 1701China/Southern 10800 120 1701Hong Kong 800 967912Japan 0044 22 132374Singapore 800 1204683South Korea 0079 81 4800 7204All other locations + 1 608 824 2900 or +32 2 40 44 44. |  | 15 | 15 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, and OR. Internationally to Canada, Belgium, United Kingdom, Australia, France, Germany, and Poland. | Terminated | November 16, 2012 | 63 | 15 | Computer
Z-2379-2012 | 62867 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | HALO Flex Energy GeneratorsManufactured by BÂŽRRX Medical (now Covidien)540 Oakmead ParkwaySunnyvale, CA 95085.The HALO Flex Energy Generator is indicated for use for the coagulation of soft tissue. The HALO Flex System is indicated for use in the coagulation of bleeding and non-bleeding sites in the gastrointestinal tract including but not limited to, the esophagus, including Barrett's Esophagus. | HALO Flex Energy Generators | 2 | September 14, 2012 | 2012 | Covidien LLC | Potential failure of HALO Energy Generator to enter "stand by" mode when initially powered on; and therefore can not be used for patient care. |  | COVIDIEN sent an Urgent Medical Device Correction letter dated August 3, 2012, to all affected customers via Federal Express. Additionally, Field Service Representatives will be contacting their accounts to notify them of the correction.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps prior to initiating patient treatment and specifically, prior to administering conscious sedation and performing upper endoscopy:1. Power-on the HALO Flex Energy Generator and confirm that the generator passes self-test, enters standby mode and displays "Ready, Connect Catheter."2. Confirm that the HALO Flex Energy Generator is functioning properly at power-up and enters the standby mode then the procedure can begin.3. Do not use the HALO Flex Energy Generator if the generator does not pass the power on self-test. In this situation, please call Covidien, GI Solutions customer service at (888) 662-2779 or your local territory manager (sales representative) so that the firm can provide them with a replacemenent generator.4. Customers were asked to reply to Covidien, GI Solutions (formerly BARRX Medical) by signing and returning the attached Product Correction Form via fax to the number provided on the form. Customers with questions were instructed to call 888-662-2779.For questions regarding this recall call 408-328-7300. | Contact the recalling firm for information | 617 | 617 generators | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, Colombia, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Israel, Jordan, Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Venezuela and UK. | Terminated | September 30, 2013 | 381 | 617 | Computer
Z-2380-2012 | 62939 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Ultraview DM3 MonitorModel Number: 91330-NTThe Ultraview DM3 Monitor is indicated for use by health care professionals for monitoring patient vital signs. The Ultraview DM3 Monitor is intended for monitoring, recording, and alarming basic vital signs on adult and pediatric patients. Monitored parameters include SpO2, pulse rate, NIBP, and temperature. | Ultraview DM3 Monitor | 2 | September 14, 2012 | 2012 | Zoe Medical Incorporated | Unit fails to power up, resulting in an equipment alarm | DESIGN: Software Design | The firm, Zoe Medical, coordinated with the Distributor via email on June 7, 2012. The email described the product, problem and actions to be taken. Zoe will remove product the site and return the units for correction.If you have any questions, contact Manager, Global Technical Support at 425-657-7200 ext 5508. | No consumer action necessary | 35 | 35 units | Nationwide distribution: WA only. | Terminated | September 18, 2012 | 4 | 35 | Computer
Z-2401-2012 | 63004 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Zoll X Series Defibrillator/Pacemaker/MonitorProduct Usage:The product is used by qualified medical personnel for the purpose of converting ventricular fibrillation to sinus rhythm or other cardiac rhythms capable of producing hemodynamically significant heart beats. In addition, the product may be used in synchronized mode to terminate certain atrial and ventricular tachycardias and other arrhythmias resistant to drug therapy. The X Series may be used for cardiac pacing in conscious or unconscious patients for up to a few hours duration as an alternative to endocardial stimulation. The X Series may be used for the monitoring a patients electrocardiogram (ECG). | Zoll Medical | 2 | September 17, 2012 | 2012 | ZOLL Medical Corporation | Shipped with incorrect software | PRODUCTION CONTROLS: Process Control | Zoll Medical telephoned affected customers on August 16, 2012. Customers were informed of the affected product, problem and actions to be taken. Customers were instructed to remove and return the affected product for replacement. For question contact ZOLL's Technical Support. | Contact the recalling firm for information | 11 | 11 units | Worldwide Distribution - US Nationwide including the states of: LA, MA, ME, NC, ND, NE, OH and the country of Germany. | Terminated | September 19, 2012 | 2 | 11 | Computer
Z-2403-2012 | 62738 | blood establishment computer software, user developed | OYT |  |  | Hematology | 510(K) Exempt | Amplilink (AL) Software 3.2.3 for use with COBAS AMPLICOR, CAP/CA, COBAS TaqMan,and COBAS AmpliPrep/COBAS TaqMan. The software can be used with a Laboratory Information System. | Amplilink Software v3.2 | 2 | September 18, 2012 | 2012 | Roche Molecular Systems, Inc. | When using a Laboratory Information system (LIS) with automated systems utilizing AMPLILINK software v3.2 series (AL v3.2) there is a potential for wrong results to be assigned to an order. | DESIGN: Device Design | Roche Molecular Systems Inc. sent "Urgent Field Safety Notice" letters and Faxback forms on 4/23/2012 via UPS to all affected customers. The issue was described and recommended actions were provided. A Product Advisory Notice and an Updated Field Safety Notice were issued as well. |  | 1292 | 1292 instruments | Nationwide (USA) Distribution | Terminated | February 12, 2014 | 512 | 1292 | Computer
Z-2404-2012 | 63006 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT 16 SystemProduct Usage:The Brilliance 16 is a Whole Body Computed Tomography X-ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. | Brilliance 16 Slice (Air) | 3 | September 19, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or hard shutdown with the power button. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Philips sent an Urgent - Medical Device Correction letter dated August 24, 2012 via certified mail to all US consignees. Philips Key Market representatives distributed the letter internationally. The letter identified the affected product, problem and actions to be taken by customer/user. Philips is implementing a software update to correct the issue free of charge. A Philips representative will contact affect sites to schedule the software installation. For further information or support contact your Philips representative or local Philips Healthcare office. |  | 93 | 93 units | Worldwide Distribution - US Nationwide including the states of: AL, CA, GA, MA, MD, NC, OH, PA, TX, and WA; and to the following countries of: Algeria, Oman, Australia, Brazil, China, Saudi Arabia, United Arab Emirates, and India. | Terminated | December 24, 2013 | 461 | 93 | Computer
Z-2444-2012 | 63186 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SharePlan 1.1, aka t-RayAutoplan version 1.2.1 (build number 1.3.1.10). Manufacturing name RayAutoplan, aka t-RayAutoplan, commercial name (TomoTherapy) SharePlanIntended use: t-RayAutoplan is a software that, based on the planned dose distribution from a TomoPlan treatment planning system, generates a family of individually optimized IMRT plans and presents their characteristics to the user in a GUI. Among the generated plans, the user selects and clinically approves a treatment plan and exports it electronically to DICOM, for subsequent treatment of the patient. The intended users of t-RayAutoplan shall be clinically qualified staff, such as medical physicists, medical doctors or dosimetrists. | SharePlan 1.1 | 2 | September 24, 2012 | 2012 | RAYSEARCH LABORATORIES AB | This notice concerns behavior of the SharePlan that may be unexpected and not clearly described in the user manual. This behavior has not caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivered treatment. The behavior appears, when importing RFA-files (.asc) | DESIGN: Software Design | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction letter dated October 10, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. To completely avoid effects of the erroneous behaviour, RFA-file has to be modified to have the crossline profiles sgtored in the X column of the file and the inline profiles in the Y column. Upon request, RaySearch Laboratories will provide help to the customers by transferring the RFA-files to the correct format through a scrip. Customers were instructed to pass on the notice to all those who need to be aware within their organization and to maintain awareness on this notice until the new labeling version has been received to ensure effectiveness of the corrective action.For questions regarding this recall call +46 8 54 50 61 42. | No consumer action necessary | 20 | 20 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, OR, SD, and WI. Internationalyl to Canada, Sweden, France, Germany, Australia, Poland, Belgium, and United Kingdom. | Terminated | January 25, 2013 | 123 | 20 | Computer
Z-2443-2012 | 62313 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 ASSEMBLY FLOW PODProduct Usage: The Terumo Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | September 24, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received two reports of the false back flow alarms for the TerumoÂż Advanced Perfusion System 1 when there was no tubing in the flow sensor. Terumo CVS' investigation found that the voltage level on a circuit board was improperly set in a population of flow modules for Terumo System 1, the sensor operates correctly if the tubing is installed and it is highly unli | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Terumo CVS issued an URGENT MEDICAL DEVICE Recall letter dated September 10, 2012 to alert all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Terumo CVS representative will contact all affected US account.to schedule an appointment to replace the affected product. The affected product will be return to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction. For questions call the Terumo CVS Customer Service at 1-800-521-2818. |  | 150 | 150 units | Worldwide and US Nationwide Distribution | Terminated | January 13, 2014 | 476 | 150 | Computer
Z-2449-2012 | 62322 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Flow Moduleindicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Advanced Perfusion System 1 Flow Module | 2 | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received one report of a flow module for the Terumo Advanced Perfusion System 1 not displaying blood flow information on the Central Control Monitor or the Centrifugal Control Unit. The tricolor LED on the flow module was not illuminated. Terumo CVS' investigation found that the malfunction was caused by an improper component on a circuit board used in a populat | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | TERUMO sent an Urgent Medical Device Recall letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; and confirm receipt of the communication by faxing the attached Customerr Resonse Form to 1-734-741-6149. Terumo CVS will send a representative to each affected US account to replace the affected devices and return the affected devices to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction.For question customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 72 | 72 units | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, Singapore, Jordan, Japan, Vietnam, Germany, Russia, Saudi Arabia, BELGIUM, Mexico, and CANADA. | Terminated | January 13, 2014 | 475 | 72 | Computer
Z-2453-2012 | 63180 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hologic Fluoroscan Mini C-Arm Model with InSight- FD Software Version 5.0.1and 5.0.2The device is a Mini C-arm Fluoroscopic Imaging System designed to provide -physicians with general fluoroscopic visualization of a patient, including but not limited to surgical orthopedic and podiatry use, critical and emergency procedures, and light anatomy imaging situation. | Hologic Fluoroscan Mini CArm Model with InSight FD | 2 | September 25, 2012 | 2012 | Hologic, Inc. | Fluoroscan Mini C-Arm models -InSight2, version 4.0 and 4.0.1 - DICOM images stored in PACS may be flipped 180Âżaround vertical axis.InSight FD, version 5.0.1 and 5.0.2 - DICOM images stored in PACS may be lipped 180Âż around vertical axis as well as 180Âż around a horizontal axis. | DESIGN: Software Design | Hologic issued an Urgent: Medical Device recall letter dated September 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Halogen corrected the anomaly and they provided the customer with the revised version of the software. Customers were asked to update their systems in accordance with the instructions provided with the software included with the letter. Customers were also asked to complete the attached acknowledgement form and fax or mail it to the firm at their earliest convenience. For any questions customers were instructed to call Hologic's Help Desk at 1-800-321-4659.For questions regarding this recall call 781-999-7300. |  | 46 | 46 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia,Belgium, Canada, China, Germany, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Italy, Japan Korea, Lithuvania, Neherlands, qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Thailand, Taiwan, South Africa, and Switzerland. | Terminated | January 10, 2014 | 472 | 460 | Computer
Z-2462-2012 | 63093 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Treatment Planning System, Model Number: H48; Eclipse 8.9 and 10.0Reference/FSCA Identifier: CP-08922; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithim is specifically indicated for planning proton treatment of neoplasms of the eyes. | Varian brand Eclipse Treatment Planning System | 2 | September 26, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning where an outdated assigned Hounsfield Unit [HU] value can be used for dose calculation when the clock on the Eclipse Client Workstation is not synchronized with the Database/System Server. Treatment of the patient using these values can therefore lead to under or over-dose. | DESIGN: Software Design | VARIAN Medical Systems sent an Urgent Medical Device Correction Urgent Field Safety Notice on August 17, 2012, to all affected customers with a description of the problem and user corrective action steps. Customers were instructed to ensure that the workstation clock be synchronized with the DatabaseSystem server clock. Customers were also instructed to contact Varian service for assistance with time synchronization if required. Customers were to complete the attached Proof of Notification Form once they have read the document and return it to Varian Medical Systems. For questions customers should call the following numbers:USA and Canada 1-888-827-4265.Europe +41 41 749 8844.For questions regarding this recall call 650-424-5731. | Contact the recalling firm for information | 6087 | 6087 | Worldwide Distribution - USA (nationwide) and Internationally... | Open |  |  | 6087 | Computer
Z-2463-2012 | 62830 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DEXIS Imaging Suite - catalog number 1.008.7932, version numbers 10.0.0 through 10.0.3Radiological Image Processing System | DEXIS Imaging Suite | 2 | September 26, 2012 | 2012 | Imaging Sciences International, LLC | A condition can occur during renumbering of the patient database using the DEXCopy feature. When renumbering is being performed using the DEXCopy feature and multiple patients with an identical name are present, the software will assign all subsequent records to the first patient record, resulting in mixed patient records. In order for this problem to occur several conditions need to be met: 1) | DESIGN: Software Design | DEXIS sent a Medical Device Corrective Action letter dated July 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. An Acknowledgement Form was included with the notification that is to be returned by the consignee to the recalling firm. Customers were instructed to call DEXIS Technical Support at 888-883-3947 as soon as possible to have Technical Support review and potentially correct any issues that may have resulted from its use. To request a free-of-charge upgrade kit, customers should email a request to orders@dexis.com.For questions regarding this recall call 215-997-5666,. | Contact the recalling firm for information | 366 | 366 | Worldwide Distribution - USA including Washington, DC, AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and the country of Canada. | Open |  |  | 366 | Computer
Z-2468-2012 | 63166 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A/2 Aquilion One CT System. The TSX-30IAl2 is a whole body CT scanner. This device captures cross sectional volume data sets. The device consists of a gantry, patient couch (table) and peripheral cabinets used for data processing and display.This device is intended to produce cross sectional volume sets of the anatomy. Additionally the system, when used by a trained professional, is capable of proving data for dynamic and perfusion studies of organs and extremities. | TSX301A/2 Aquilion One CT System | 2 | September 27, 2012 | 2012 | Toshiba American Medical Systems Inc | The SureStart function may not operate in the usual manner, and you may obtainabnormal analysis results when using the cerebral blood flow analysis function. | DESIGN: Software Design | Toshiba American Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.customers: The firm advised customers to take the following temporary measures to avoid issues until a corrective action has been taken:I. In SureStart:" Perform scanning with Observation Direction set to VFF when the PatientInsertion Direction and Observation direction are FF." Perform scanning with the Observation direction set to VFH when the PatientInsertion Direction and Observation Direction are HF.2. When acquiring images for 4D perfusion analysis:" Perform scanning with the Scan Direction set to OUT when the Patient Insertion Direction and Scan Direction are HF.Perform scanning with the scan direction set to IN when the Patient InsertionDirection and Scan Direction are FF.Contact the firm at (800) 421-1968 for questions regarding this issue. |  | 3 | 3 units | Nationwide distribution: USA including states of: MA, MD, and NV. | Terminated | September 27, 2012 | 0 | 3 | Computer
Z-2466-2012 | 62979 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | BariAir Therapy SystemThe BariAir Therapy System is a risk management treatment system designed for large and / or difficult-to-manage patients. The BariAir Therapy System provides low air-loss pressure management therapy, pulsation, percussion, Turn Assist. flexible patient positioning capabilities and built-in scales for patients weighing 300 Ib* to 850 Ib ([136*-386 kg] including accessories). The BariAir Therapy System is indicated to aid in: " Patients whose body weight and size pose a significant risk or care management issue to the patient or staff during the performance of routine nursing care. " Large patients weighing between 300 Ib* and 850 Ib (136*-386 kg) including accessories. " Large patients who are difficult to turn. " Preventing and treating pressure ulcers. " Large patients requiring percussion therapy. | BariAir Therapy System | 2 | September 27, 2012 | 2012 | KCI USA, Inc. | KCI has issued a medical device correction for the BariAir Therapy System Hand Pendent. The BariAir Therapy System is not designed to allow the patient to place the bed into the patient exit position. KCI has received complaints indicating that patients used the hand pendent to inadvertently place the BariAir Therapy System into a pseudo-exit position. In a small number of instances, this allow | DESIGN: Device Design | KCI sent an Urgent - Voluntary Medical Device Correction letter dated September 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers would be contacted by KCI to coordinate a shipment of a replacement product. For questions customers would instructed to call 1-800-275-4524, Select Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. | Contact the recalling firm for information | 575 | 575 | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | July 29, 2013 | 305 | 575 | Computer
Z-2492-2012 | 61796 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Voice + Monitor, with wire bracket, Model 25223. Stanley Security Solutions, Inc., Lincoln, NE.Designed to help reduce the risk of falls by alerting caregivers when a resident is leaving a bed, chair, room or wheelchair. | Tabs Professional Voice Monitor, Model 25223 | 2 | September 28, 2012 | 2012 | Stanley Security Solutions, Inc. | The monitor may fail to sound an alarm | DESIGN: Device Design | Stanley Security Solutions, Inc. sent an Important Recall Notice dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease using any of the affected product. The customers were informed, however, that they may continue to use the units as long as they were powered with a 9-volt alkaline battery in accordance with current instructions because the issue only occurs when using the 12V DC adaptor. The letter also informs the consignee to expect a follow-up notice no later than 6/1/12 that will provide additional instructions.The recalling firm issued additional letters dated 6/5/12 to distributors and end users with updated instructions. The letter informed them further testing revealed the possibility that a subset of affected monitors may also fail to alarm when powered by a 9V battery if weight is rapidly applied and removed from the pressure pad. Therefore, the recalling firm decided to recall and repair this subset of affected monitors and provided the serial number ranges that must be returned for repair. The letter also provided a list of serial number ranges that could be corrected in the field by permanently blocking access to the 12V DC jack on the monitor. Instructions were enclosed to show the customer how to disable to the jack. In addition the distributor letters request the customer to provide the recall information to their downstream customers. A response form was included to report the amount of affected product in the customers' inventory.For questions customers were instructed to call 888-224-2460.For questions regarding this recall call 407-271-2097. | Contact the recalling firm for information | 12111 | 12,111 units | Worldwide Distribution - USA (nationwide) and the countries of Canada and Great Britain. | Terminated | September 24, 2013 | 361 | 23938 | Computer
Z-0011-2013 | 63116 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart XL M4735A Defibrillator/Monitors, used with the M3516A batteries The Heart Start XL Defibrillator/Monitor is a fully featured, defibrillator intended for use by qualified medical personnel, trained in either Advanced Cardiac Life Support or in the operation of the device, in a hospital environment. | Batteries for XL Defibrillator/Monitor | 2 | October 04, 2012 | 2012 | Philips Healthcare Inc. | HeartStart XL Defibrillator/Monitor (M4735A) used on a patient when AC power is not in use, with a battery that may fail, the device may shut down unexpectedly without providing the user with a low battery warning message or audible tone. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Philips Healthcare, sent an "URGENT-MEDICAL DEVICE RECALL/Field Safety Notice" letter dated September 6, 2012 to its customers. The letter describes the product, problem and actions to be taken by the customer and also identifies what action Philips plans to take to remedy the issue. The customers were instructed to operate the XL and AC power instead of battery power alone; use a battery from a different date of manufacture (must not have the R-2011-12 date of manufacture code) if they need to operate the XL on battery power alone; and use an alternative source for treatment or therapy (i.e., backup defibrillator/monitor) if unable to operate with AC power alone or locate a battery with a different date of manufacture than that identified. Phillips will provide a replacement battery(s) to customers as soon as new production is available.If you need any further information or support concerning this issue, contact your local Philips representative or call Quality Specialists,Recalls at 978-687-1501 or email: debra.levasseur@philips.com. | Contact the recalling firm for information | 2500 | 2500 units | Worldwide distribution: USA (Nationwide) and countries of: AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF KUWAIT LEBANON, MACAU, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VIET NAM. | Open |  |  | 2500 | Computer
Z-0009-2013 | 63064 | negative pressure wound therapy powered suction pump | OMP | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Wound Pro Apex Negative Pressure Wound Therapy Powered Suction Pump; Accuro Medical Products LLC, 2854 Hitchcock Ave, Downers Grove, IL 60515; part number 84635400182, sold inside the Apex Negative Wound Therapy system, part number 84635400181Promotes wound healing through drainage and removal of infectious material and other fluids from the wound site using continuous and/or intermittent negative pressure. Types of wounds indicated include diabetic/neuropathic ulcers, pressure ulcers, chronic wounds, acute wounds and dehisced wounds. | Negative Pressure Wound Therapy Powered Suction Pump | 2 | October 04, 2012 | 2012 | Accuro Medical Product LLC | Accuro has received reports of some Apex Pumps either failing to operate or malfunctioning, at times without notifying the user of such failure or malfunction on the unit's display screen. | DESIGN: Device Design | The firm, Accuro, sent an "Urgent: Medical Device Recall" letter dated August 30, 2012 to the customers via e-mail on August 31, 2012. The letter describes the product, problem and actions to be taken. The customers were instructed to remove all devices from use immediately; to immediately complete and return the enclosed Accuro response form via email to; (customer.service@accuromedical.com) or fax to: (1-800-866-0545) . Note: the form still needs to be returned even if you no longer have or use the device. Upon receipt, Customer Service will call the account with instructions to arrange for all devices to be returned to Accuro. General questions were directed to Customer Service at 1-800-669-4757. |  | 919 | 919 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Open |  |  | 919 | Computer
Z-0048-2013 | 63045 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza radiological image processing system. | Siemens syngo.plaza | 2 | October 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using syngo.plaza. In syngo.plaza within the online "Short Time Storage" (STS), some studies are marked with an archived icon, however, the entry in the Archive Database referring to the physical file may be missing. | DESIGN: Software Design | The firm initiated this correction on June 6, 2012 by issuing update instructions SY129/12/S to consignees. A script will be distributed to all potential affected customer systems which will prevent data loss by setting the delete protection on all affected studies. |  | 11 | 11 | Nationwide Distribution including the states of MO, FL, MA, NY, WI, GA, CA and WA. | Terminated | April 13, 2013 | 184 | 11 | Computer
Z-0049-2013 | 63079 | pump, blood, cardiopulmonary bypass, non-roller type | KFM | Nonroller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery, 100/115V (Includes charger and connecting cable)The Sarns" Centrifugal System is indicated for use in cardiopulmonary bypass procedures only. | Sarns" Centrifugal System | 2 | October 11, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2004 of variousmalfunctions of the battery backup system for the Sarns" Centrifugal System.The reports can be broadly categorized as follows:" Failure to charge" Failure to hold charge" Failure to operateIn the most severe case, a battery malfunction could result in an unexpected loss of arterial flow. In | DESIGN: Component Design/Selection | TERUMO sent an Urgent Medical Device Recall letter dated June 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. On 09/25/2012 an URGENT MEDICAL DEVICE RECALL notification and addendum was sent to all consignees. TCS issued this notification as a Safety Advisory to alert all users of the Sarns centrifugal system of the potential for malfunction and remind them to review the appropriate sections of the Operators Manual for the Sarns Delphin Battery. Providing an addendum to the Operators Manual for the Sarns Delphin Battery with additional instruction on the use and charging of a discharged battery.Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; place the addendum in the Operator's Manual; and confirm receipt of the communication by faxing or emailing the attached Customer Response Form.For questions customers were instructed to call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 1076 | 1076 | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, UT, VA, WA, WI, and WV. Puerto Rico and internationally to:ALEXANDRIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CENTRO AMERICA, CHILE, COLOMBIA, Dominican Republic, Ecuador, FRANCE, GERMANY, GREECE, Guatamala, HOLLAND, HONG KONG, India, ITALY, Japan, KL SEGUNDO, KOREA, LEICESTERSHIRE, MALAYSA, Malaysia, Mexico, Mexico, Mexico, NETHERLANDS, Panama, Peru, Philippines, RALEIGH, Republic of Georgia, SANTIAGO CHILE, Saudi Arabia, SEOUL KOREA, SHANGHAI, CHINA, Singapore, South Africa, SPAIN, SWITZERLAND, TAIWAN, Thailand, UNITED ARAB EMIRATES (UAE), Uruguay, USA, and Venezuela, | Terminated | January 15, 2014 | 461 | 1365 | Computer
Z-0051-2013 | 62059 | pump, blood, cardiopulmonary bypass, non-roller type | KFM | Nonroller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal System Control module, 100/115VThe Sarns Centrifugal System is indicated for use in cardiopulmonary bypass procedures only. | Sarns Centrifugal System | 2 | October 11, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received one report of loss ofpower to the control module and subsequently the drive motor for the Sarnscentrifugal system during cardiopulmonary bypass. | DESIGN: Component Design/Selection | TERUMO sent an Urgent Medical Device Recall dated April 13, 2012, to all affected consignees. An Addendum was sent on September 25, 2012. All US Consignees will receive a Safety Advisory and an update to the Operator's Manual via Federal Express to inform them of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of back-up system availability. TERUMO will include a customer response form and follow up with all affected users to obtain a response to the recall.For questions customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. | Contact the recalling firm for information | 847 | 847 units | Worldwide Distribution - USA (nationwide) and the countries of Bangladesh, Seoul Korea, Chile, Brazil, Venesuela, Thailand, Mexico, India, Taiwan, Canada, Dominican Republic, Guatemala, Ecuador, Germany, Republic of Georgia, Hong Kong, Japan, United Arab Emirates (UAE), Japan, Belgium, Phillipines, Singapore, and South Africa. | Terminated | November 13, 2013 | 398 | 1186 | Computer
Z-0058-2013 | 62703 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Belmont Rapid Infuser, a Fluid Management System,Model FMS2000. | Belmont Rapid Infuser | 2 | October 12, 2012 | 2012 | Belmont Instrument Corporation | Leak from a saline bag caused saline to drip along the back of the Rapid infuser close to the Power Entry Module causing an electrical flash at the power entry area. | DESIGN: Device Design | Belmont Instrument initiated recall on 9/25/12. A "Field Correction" notice was provided that described the product issue and recommended actions. US-based sales representatives visited all facilities to provide the modified power cord/shield. For units sold outside the United States, the modified power cords were sent to distributors who would ensure the new power cords were provided to all facilities within their accounts. |  | 2913 | 2,913 units | Worldwide Distribution including Nationwide (USA) and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, England, Estonia, France, Germany, Hong Kong, India, Israel, Italy, Jordan, Mongolia, Netherlands, New Zealand, Norway, Pakistan, Peru, Puerto Rico, S. Korea, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, UK, United Arab Emirates, and Vietnam. | Terminated | February 28, 2014 | 504 | 2913 | Computer
Z-0056-2013 | 63155 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 1250 and AutoMate 2550, Part Numbers ODL25125 and ODL25255.The AutoMate 1200/2500 System Series is a semi-open, pre- and post-analytical sample processing and sorting system. The base system automates the sample sorting, decapping, and archiving processes. Handling and sorting of samples includes automatic detection of the tube type and cap type (color). Optional features include an Aliquot Module for creation of labeled secondary tubes (including the detection of the presence of adequate volume for the requested aliquots) and a Recapper Module to re-seal previously decapped tubes prior to archiving. | AutoMate 1250 and AutoMate 2550 | 2 | October 12, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of secondary tubes being labeled with two different barcode labels on the AutoMate 2500 Family (previously referred to as OLA2500 Systems) units which can cause misidentification of a sample when software versions lower than version 4.1 is used. | MISBRANDING: Labeling False and Misleading | Beckman Coulter initiated an Urgent Product Correction letter dated April 11, 2012, with attached PCA Response Form to all customers who purchased the affected product. . The letter identified the product, the problem, and the action taken by the customer. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers with questions were instructed to contact their local Beckman Coulter representative. Customers with questions should call 714-993-5321.For questions regarding this recall call 714-961-4941. | Contact the recalling firm for information | 468 | 468 units total (51 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:ArgentinaAustraliaAustriaBelarusBelgiumChinaCroatiaCzech RepublicDenmarkFinlandFranceGermanyHungaryIndiaIrelandIsraelItalyJapanKorea, Republic ofKuwaitMacaoMalaysiaNetherlandsNew ZealandNorwayPanamaRussian FederationSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited Kingdom | Terminated | February 11, 2013 | 122 | 468 | Computer
Z-0073-2013 | 62516 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory :intended for use by professionals working in a clinical laboratory. | Sunquest Laboratory | 3 | October 17, 2012 | 2012 | Sunquest Information Systems, Inc. | Online calculations may not be performed as expected or may yield results that are different than expected when the calculation module is called from the Online Entry (OEM) application for a subtest whose subcolumn number on the instrument online string is divisible by 10. | DESIGN: Software Design | Urgent - Product Safety Notice: PSN-11-06,dated, 15 November 2011, was sent to customers. Costumers should follow instructions in the notification letter. To request the software correction, create a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb.Return receipt acknowledging PSN-11-06 is required .Client Support can be reached at 1-877-239-6337 for U.S. and Canadian sites and at +44 (0) 161 335 0562 for International sites. |  | 127 | 127 sites | Nationwide distribution | Terminated | October 19, 2012 | 2 | 127 | Computer
Z-0069-2013 | 63056 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; List Number 16026-01Device is a general purpose infusion pump designed to deliver fluids, solutions, medications, agents, nutritionals, electrolytes, blood and blood products for parenteral, enteral, intravenous, intra-arterial, subcutaneous, epidural | Symbiq One Channel Infuser | 1 | October 19, 2012 | 2012 | Hospira Inc. | The Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation scree | DESIGN: Device Design | Hospira sent an Urgent Device Correction letter dated August 29, 2012, to all affected customers. Customers have been informed of reports that the Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation screen prior to starting the infusion. The users were instructed to complete and fax the enclosed reply card to 1-888-943-5180 or email it to Hospira194@stericycle.com indicating the number of infusers at the facility and if they notified their users throughout the facility. Wholesalers were asked to notify their customers. If they have further distributed the affected product, they should notify their accounts and ask them to contact Stericycle at 1-877-272-6141 to receive a reply form. For questions regarding this recall call 224-212-2000. |  | 19614 | 19,614 units | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada | Open |  |  | 33681 | Computer
Z-0095-2013 | 61603 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1: Seno Advantage is a medical image review station that allows easy selection, processing, filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. When interpreted by a trained physician, mammographic images displayed on the high-resolution monitors may be used as an element for diagnosis. Furthermore filmed images from all modalities may also be used as an element for diagnosis. | Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1 | 2 | October 19, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the magnification images acquired on non-GE digital mammography systems and your Seno Advantage Workstation that may impact patient safety. The measurement values provided by the Seno Advantage Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems. If not noticed by the caregiv | DESIGN: Software Design | GE Healthcare sent a Urgent Medical Device Correction letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will install new software to correct this issue with the release of Field Modification Instruction (FMI) 12129 in the near future. This correction will not affect the performance or any other functionality of your current equipment. A GE Healthcare service representative will contact you to arrange for this modification. This activity will be performed at no cost to you.If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171, Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 188 | 188 | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, DC, FL, GA, KY,LA, MA, MO, NV, NJ, NY, NC, OH, OK, PA, SD, TN, TX, VA and WI., and the countries of UNITED KINGDOM, TUKEY, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL, NEW ZEALAND, NETHERLANDS, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, HUNGARY, GREECE, GERMANY, FRANCE, CHINA, CANADA, BRAZIL, AUSTRIA and ALGERIA. | Terminated | March 26, 2013 | 158 | 188 | Computer
Z-0121-2013 | 63424 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM/CT670; GE Discovery NM 630Product Usage: a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy. The GE Discovery NM/CT 670 system, combining Nuclear Medicine (NM) and Computed Tomography (CT) systems, is intended to produce: NM System: General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and imaging features designed to enhance image quality. | GE Discovery NM/CT670; GE Discovery NM 630 | 2 | October 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with thetable pallet of your Discovery NM/CT 670 and/or Discovery NM630 system which may impactoperation safety.At scan setup, during gantry rotation towards scanning position, a potential contact between the table-pallet-edge and the detector was identified. Such occurrence may damage the pallet front-edge and | DESIGN: Software Design | GE Healthcare, LLC sent an " Urgent Medical Device Correction" letter dated October 12, 2012 to all affected customers. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Mangers, and Mangers of Radiology / Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions or concerns contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 210 | 210 | US Nationwide Distribution in the states of: AL, CA, FL, GA, IL, IN, KS, LA, MD, MA, MN, MS, MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, VA, WI and WV.AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL,ITALY, JAPAN, REPUBLIC OF KOREA, NETHERLANDS,NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TAIWAN,THAILAND, TURKEY, UNITED KINGDOM. | Terminated | November 14, 2013 | 387 | 210 | Computer
Z-0118-2013 | 63405 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam 4.1.Radiation treatment planning software designed to analyze and plan radiation treatments in their dimension for the purpose of treating patients with cancer. | Oncentra External Beam | 2 | October 23, 2012 | 2012 | Elekta, Inc. | During the planning process the beam weights can be changed by the customers. When the beam weights are changed and renumbered and this happens during the dose calculation it is possible to save and reopen the plan without any error messages. | DESIGN: Software Design | Nucletron sent a Customer Information Bulletin 555.00248 to all affected customers. The bulletin identified the product, problem, and actions to be taken by the customers. The bulletin contained information on how an incorrect dose association can be prevented by not renumbering beams during recalculation. A Read Confirmation Form was included for customers to complete and return. Contact your local Nucleton Service Representative for questions regarding this notice. | Contact the recalling firm for information | 52 | 52 units | Nationwide Distribution-including the states of AL, CA, DC, KS, LA, MO, NB, NY, NC, ND, OH, OR, PA, RI, SC, TX, and WA. | Open |  |  | 52 | Computer
Z-0125-2013 | 61612 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.2 workstations with software version 22_01 and 22_02 A medical image review station that allows easy selection, processing, filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. | Seno Advantage 2.2 | 2 | October 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety. This issue may result in a potential miscalculation of size. | DESIGN: Software Design | Consignees were sent on 5/24/12 a GE Healthcare "Urgent Medical Device Correction" letter dated May 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers & Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. Customers should verify if their workstation is affected. GE Healthcare will correct all affected workstations by providing a software upgrade and Operator Manual update. A GE Healthcare service representative will contact you to arrange this correction. This activity will be performed at no cost to you. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 12 | 12 | Worldwide Distribution -- USA (nationwide) including states of: CA, CT, NY, TX, and MS and countries of TAIWAN, ITALY, CANADA, and BAHRAIN | Terminated | December 11, 2012 | 48 | 12 | Computer
Z-0123-2013 | 63143 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCelÂŽ DxC 600/800 withSoftware v5.0.11, Part Number B12218.The UniCelÂŽ DxC SynchronÂŽ Clinical Systems are fully automated, computercontrolled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | UniCel DxC 600/800 with | 2 | October 24, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignmentprocedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform thisalignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to theShuttle to Bar | DESIGN: Software Design | Beckman Coulter sent an Urgent Field Safety Notice dated August 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their Beckman Coulter Service Representative would contac them to reinstall the software. Customers were asked to complete and return the enclosed Response Form. For questions customers were instructed to call their Customers Support Center on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/SupportRequest.For questions regarding this recall call 714-993-5321. |  | 7 | 7 units total (both US and International) | Worldwide Distribution - USA including TX and OH. Interntionally to China and Turkey | Terminated | February 14, 2014 | 478 | 7 | Computer
Z-0128-2013 | 62315 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 4 pump, 220V/240V safety monitor includedThe Sarns Modular Perfusion System 8000 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Perfusion System 8000 | 2 | October 24, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the power distribution and battery back-up systems for the Sarns Modular Perfusion System 8000. Very few of the reports occurred during cardiopulmonary bypass; the majority of the malfunctions occurred during set-up or were reported during routine preventative maintenance. The malfunction | DESIGN: Component Design/Selection | TERUMO sent an URGENT MEDICAL DEVICE letter dated September 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed review the Medical Device Recall notice, assure that all users are aware of the notice and place the addendum in the Operator's Manual before page 1.13. Customers were also instructed to confirm receipt of the Urgent Medical Device letter by faxing, the attached Customer Response Form to 1-800-292-6551. For questions customers were instructed to call 1-800-521-2818. |  | 466 | 466 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of ARGENTINA, AUSTRALIA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, CALLO PERU, CANADA, CAROLINA PR, CHILE, China, COLOMBIA, Costa Rica, DEKALB, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, Germany, GUAYQUIL ECUADOR, Hong Kong, India, Indonesia, ITALY, Japan, , JARAMA MADRID, Lebanon, LEICESTERSHIRE, MADRID SPAIN, MADRID SPAIN, Malaysia, MEDLEY, Mexico, D.F., MIAMI, MUMBAI 26 INDIA, MUMBAI, New Zealand, Nicaragua, Pakistan, Panama, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, SELANGOR, MALAYSA, SEOUL, KOREA, Singapore, SOUTH AFRICA, South Korea, Sri Lanka, SWITZERLAND, TAIPEI, TAIWAN, Thailand, TUJUNGA, UNITED ARAB EMIRATES (UAE), URUGUAY, Venezuela, Vietnam, VIGNATE ITALY, WOODDALE, and Yemen. | Open |  |  | 2256 | Computer
Z-0122-2013 | 63081 | intraoperative orthopedic joint assessment aid | ONN | Stereotaxic instrument. | Neurology | Neurology | 510(k) | "****ORTHO SENSOR**12-123-A Knee Trial for Trithion***1560 Saw grass Corporate Pkwy***4Th floor***Sunrise, FL 33323***www.orthosensor.com***"Rx Only*** Sterile/EO***Made exclusively for Stryker Orthopaedics***.Is the graphic user interface software associated with the graphic user interface, which is loaded onto the linkstations. | ORTHO SENSOR | 2 | October 24, 2012 | 2012 | Orthosensor | Ortho sensor in Sunrise, FL is recalling the Graphics User Interface Software 1.8 (which was updated to version 1.82, validated on 02/13/12). Recall was initiated due to customer feedback which included the line of site between the knee ba lancer and the Link station was obstructed by OR metal stands or trays or the knee balancer device was placed atop of or near the metal stands or trays, interru | DESIGN: Software Design | Orthosensor did not provide a written customer notification. As the recall was reported to the FDA well after all nine of the products were corrected (the recall stemmed from an initial inspection of the firm), then no customer notification was requested. The removal and replacement of the previous version of the Graphic User Interface Software is being handled internally by members of Operations and the Sales force. The items the firm is recalling is software, and the software cannot be returned; the old version (version 1.8) would be erased. The new version 1.82 would be installed and downloaded into the equipment at the site during servicing. This does not involve the customer of which none are consignees. Effectiveness checks associated with the update of the software is being handled internally by the Operations and the Sales force.Further questions please call (813) 645-2855. | Contact the recalling firm for information | 9 | 9 | Nationwide Distribution including the states of CA, FL, MI and NY. | Terminated | December 12, 2012 | 49 | 9 | Computer
Z-0124-2013 | 63139 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartSine Samaritan(R) 300/300P PAD devices with serial numbers 0400000501 to 0700032917; 08A00035000 to 10A0070753; 10C00200000 to 10C00210106Product Usage: Samaritan 300/300P PAD is indicated for use on victims of sudden cardiac arrest who are exhibiting the following signs:UnconsciousnessNot breathingWithout circulation | HeartSine Samaritan(R) 300/300P PAD devices | 1 | October 27, 2012 | 2012 | Heartsine Technologies, Limited | Certain Samaritan 300/300P PAD devices may experience one or both of the following conditions that could affect the ability to deliver therapy to a patient in an SCA event if needed: Issue 1 (On/Off Issue): the device may turn itself on and off without input from the user. When this occurs, the normal sequence of audible instruction prompts that occurs when the device is turned on will be emitt | DESIGN: Device Design | HeartSine sent an Urgent: Medical Device Recall dated September 2012 to all affected customers. The letter identifed the affected product, problem and actions to be taken. A response form was included with the letter to be completed by customers and returned to the recalling firm's representative. In addition, HeartSine is providing a software upgrade to bring all users up to a more recent version of the software. For questions call 1-877-877-0147 or your service representative. |  | 81982 | 81,982 in total | Worldwide Distribution and US (nationwide) in the following states: AK, AL, AP, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, and WI. | Open |  |  | 81982 | Computer
Z-0166-2013 | 63381 | control, pump speed, cardiopulmonary bypass | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | SCP Pump Control Panel, Product Code: 60-02-15, Serial Numbers: 60S09121, 60S09123-60S09129, 60S09136-60S09139.The SCP is a cardiopulmonary bypass speed control device indicated for use exclusively with the COBE Revolution pump head for speed-controlled pumping through cardiopulmonary bypass circuit for typical durations of six hours or less. | Stockert Centrifugal Pump (SCP) Control Panel (S3) | 2 | October 29, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of the Stockert Centrifugal Pump (SCP) Pump Control Panel . An out of specification circuit board component within the Panel may lead to a malfunction consisting of a blank or flashing display. The flow and pressure display information may also be incorrect after the malfunction. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Sorin Group USA, Inc., sent an "URGENT FIELD SAFETY NOTICE" dated September 20, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. Customers were instructed to discontinue use of and segregate affected product for return and replacement to the firm; to complete and return the attached Customer Response Form via fax to: 303-467-6502 or email to: yvonne.feyerherm@sorin.com; and if they have transferred the affected products to a third party, to pass this information to the Sorin Group.Further questions may be directed to Sorin Group Customer Service at 1-800-650-2623. | Contact the recalling firm for information | 12 | 12 units (3 units USA and 9 units OUS) | Worldwide distribution: USA (nationwide) including states of : MI and NY; and countries of: Belgium, Canada, Japan, India, Russia, United Arab Emirates. | Terminated | January 28, 2013 | 91 | 20 | Computer
Z-0169-2013 | 63335 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQMOSAIQ is an image-enabled electronic medical record system (EMR) used for oncology workflow management. It lets users supply electronic patient charts and assemble care plans, order diagnostic tests, and prescribe medications. It also lets users import, view, annotate, manipulate, enhance, manage, and archive images. | Oncology Information System | 2 | October 30, 2012 | 2012 | Elekta, Inc. | A patient was seriously mistreated after a user made and saved an inadvertent change to the definition of a treatment filed in MOSAIQ. | DESIGN: Software Design | The firm, MOSAIQ, sent an "IMPORTANT SAFETY NOTICE- LINMSQ0003" was sent to consignees/customers on September 18, 2012. The notice described the product, problem and actions to be taken. The customers were instructed to check their Department Setup configuration settings, and re-enable the Field Edit Delta Display feature if they have disabled it. Note: sites can take this action without an upgrade/patch. The customers were also instructed to complete and return the CONFIRMATION OF RECEIPT form via fax to: 702-992-5002 in US or +44 (0)1293 654401 in OUS, Attn: Elekta OIS Support; scan and email to: support@impac.com or europe.support@impac.com; Attn: Elekta OIS Support and/or mail to: C/O Elekta OIS Support, 2310 Corporate Circle, Suite 275, Henderson, NV, USA; or C/O Elekta OIS Support, Elekta Limited, Linac House, Fleming Way, Crawley, West Sussex, RH10 9RR, UK.(Note: Product upgrade/patch is not mandatory or even necessary to reduce the related risk.)For questions, concerns, and requests for upgrade, email: support@impac.com (North America and Rest of World) or europe.support@impac.com (Europe) or call Vice President of RA/QA RNA at 770-670-2548. |  | 300 | 300 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | June 06, 2013 | 219 | 300 | Computer
Z-0186-2013 | 63104 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software for Vitrea, Vitrea fX, and Vitrea Enterprise Suite (VES): VPMC-09293C Vitrea . Model Number Vitrea 5.2, 6.0, 6.1, and 6.2VPMC-09346B Vitrea fX Model Number Vitrea fX 3.1, 5.2, 6.0, 6.1, and 6.2VPMC-11731A Vitrea Enterprise Suite Model Number Vitrea Enterprise Suite 1.3, 6.0, 6.1, and 6.2.The software is a medical diagnostic system that allows the processing, review, analysis, communication, and media interchange of multi-dimensional digital images acquired from a variety of imaging devices. It provides the ability to review digital images from original DICOM data for multiple modalities, including CT, MR, DR, CR, XA, US, NM, PET, etc. The software is not meant for primary image interpretation in mammography. The Vitrea VScore option is intended for cardiac scoring from whole body CT derived measurements, including non-invasive detection and quantification of atherosclerotic plaque. Two image processing options, EKG Gate and Auto Gate, allow the operator to select images with reduced motion artifacts when processing data for Coronary Artery Calcification Scoring. | Software for Vitrea, Vitrea fX, and Vitrea Enterprise Suite (VES): | 2 | November 01, 2012 | 2012 | Vital Images, Inc. | Vital Images has found two potential errors in calcium score values in restored snapshots of studies generated through use of Vitrea or Vitrea Enterprise Suite. This issue affects Version 3.1 of Vitrea fX; Version 5.2 of Vitrea; Versions 6.0, 6.1 and 6.2 of Vitrea and Vitrea fX; and Versions 1.3, 6.0, 6.1 and 6.2 of Vitrea Enterprise Suite; and all updates to those versions. | DESIGN: Software Design (Manufacturing Process) | Vital Images sent an "URGENT DEVICE CORRECTION" letter dated August 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter gave advise on how to resolve the two issues and advised consignees to immediately inform all potential users. Contact Customer Support at 1-800-208-3005 for questions regarding the notice. |  | 2088 | 2088 (1133 USA) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BULGARIA, CANADA CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, IVORY COAST, JAPAN, JORDAN KAZAKHSTAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MALI, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TURKEY, UKRAINE, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, AND YEMEN | Open |  |  | 2088 | Computer
Z-0187-2013 | 63471 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, Version 11. Varian Medical Systems, Palo Alto, CA. The ARIA Radiation Oncology product is a treatment plan and image management application. | ARIA Radiation Oncology Version 11 | 2 | November 01, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the ARIA Oncology Information System Version 11 where Monitor Units of one fraction may be incorrect, when partial treatment is being continued. | DESIGN: Software Design | "Urgent Medical Device Correction/ Urgent Field Safety Notice" letters were sent on September 24, 2012 by Certified Mail with return receipt to all affected users. The problem was described and recommended user actions were provided. The notification letters were to also be distributed to Varian Sales, Marketing, and Service organizations. US customers may contact the Varian help desk at 1-(888) 827-4265. |  | 205 | 205 | Worldwide Distribution - Nationwide(USA) and including the countries of Australia, Denmark and the United Kingdom. | Terminated | December 05, 2013 | 399 | 205 | Computer
Z-0206-2013 | 63430 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Pinnacle3 VCC P/N 4598 001 41341 REV A, Philips Medical Systems (Cleveland), Inc., Fitchburg, WI.Pinnacle3 Radiation Therapy Planning System is a computer software package intended to provide support for radiation therapy treatment planning for the treatment of benign or malignant disease processes. Pinnacle3 Radiation Therapy Planning System assists the clinician in formulating a treatment plan that maximizes the dose to the treatment volume while minimizing the dose to the surrounding normal tissues. The system is capable of operating in both the forward planning and inverse planning modes. The device is indicated for use in patients deemed to be acceptable candidates for radiation treatment in the judgment of the clinician responsible for patient care. Plans generated using this system are used in the determination of the course of a patient's radiation treatment. They are to be evaluated, modified and implemented by qualified medical personnel. | Pinnacle3 Radiation Therapy Planning System | 2 | November 05, 2012 | 2012 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems Cleveland), Inc. is recalling Pinnacle3 Virtual Client Connection software version 3.1 due to an improper release and subsequent distribution. | PRODUCTION CONTROLS: Process Control | The firm, Philips Healthcare, sent a "Customer Information Letter" dated 2012 Oct 05. The letter listed the affected products, problem description, how to identify affected products, advise on actions by customer / user, actions planned by Philips and further information and support. The letter stated that no action is needed on the part of the customer. The field service engineer will be scheduling a site visit to complete the replacement. Philips Healthcare will be delivering the correct VCC kit.For customers in North America if you need any further information or support concerning this issue, please contact our Customer Care Solutions Center at 1-800-722-9377. Select option 5 for "All Imaging Systems". Enter your site ID # (If you do not have a site ID#, simply pause for a moment.). Select option 5 for "Nuclear Medicine" and finally select option 3 for "Pinnacle" support. |  | 11 | 11 | Nationwide distribution: USA including states of: ME, MI, NY, and WI. | Terminated | February 21, 2013 | 108 | 11 | Computer
Z-0195-2013 | 63368 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX CO-OX analyzer, model #393-841 (all software versions prior to 1.35).The ABL80 FLEX CO-OX is a portable, automated analyzer that measures pH, blood gases, electrolytes, glucose and oximetry in whole blood. The ABL80 FLEX CO-OX system is intended for use by trained technologists, nurses, physicians and therapists. It is intended for use in a laboratory environment, near patient or point of care setting | ABL80 FLEX COOX analyzer | 2 | November 05, 2012 | 2012 | Sendx Medical Inc | The recall was initiated because SenDx Medical Inc. became aware of a software issue in the ABL80 FLEX CO-OX analyzer related to the reporting of the hemoglobin fractional parameters. | DESIGN: Software Design | SenDx Medical sent a recall letter to all their foreign customers on January 6, 2010, to informed them of the ABL80 FLEX CO-OX recall via the Radiometer Field Action Notification (FAN) system. The Field Action Notes informed the customers of the problem identified and the action to be taken. Customers were instructed to submit confirmation fax#1 to confirm that the customer letter has been distributed to all ABL80 CO-OX customers by January 31, 2010. Customers were instructed to submit confirmation fax #2 to confirm that all ABL80 CO-OX analyzers have been upgraded to version 1.35 by 6/30/10.In the U.S., the customer was provided with a software upgrade CD along with a letter advising them to the issues associated with the recall. Customers with questions were instructed to contact TechLine at (800) 736-0600 opt 2. |  | 767 | 767 units | Worldwide Distribution-- US Distribution including the state of Ohio and the countries of Switzerland, Germany, France, Poland, Austria, Italy, Japan, Australia, Denmark, China, Sweden, Norway, Spain, United Kingdom, Netherlands, and Czech Republic. | Terminated | November 05, 2012 | 0 | 767 | Computer
Z-0196-2013 | 63506 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The touchscreen display (model 94267) designed for use with the XPREZZON bedside monitor (model 91393) is width-452 mm (17.2 inches) and height- 368 mm (14.5 inches). The depth without the pedestal --- 67.7 mm (2.67 inches). The display is an integrated LCD 19-Inch. The Spacelabs Healthcare Xprezzon Bedside Monitor passively displays data generated by Spacelabs parameter modules. | Touchscreen Display Model 94267 | 2 | November 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Potential failure of Spacelabs Healthcare Touchscreen Display, Model 94267 where the alarm lights failed to illuminate upon installation. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Spacelabs Healthcare sent a URGENT - MEDICAL DEVICE CORRECTION letter dated October 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to advise their staff of this potential failure and advise them to ensure the alarm lights are functional. The firm will contact the consignee at the earliest possible date to schedule a convenient time for the hospital to upgrade the affected product at no cost.Customers can call Technical Support at 1-800-522-7025 and select 2 for any questions about this recall. |  | 169 | 169 units distributed (98 in US and 71 Internationally) | Worldwide Distribution--US (nationwide) including states of Alaska, California, Florida, Illinois, Michigan, South Carolina, and Texas., and the countries of Canada, Chile, France, Libya, Panama, Qatar, Saudi Arabia, Switzerland and United Kingdom | Terminated | March 25, 2013 | 140 | 169 | Computer
Z-0251-2013 | 63380 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Application Interfacing Outbound Phlebotomy Point of Care Interface, v3.0.0_07222004 and later. | Sunquest Application Interfacing Outbound Phlebotomy Point of Care Interface | 2 | November 06, 2012 | 2012 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest Information System has confirmed reports of an Outbound Phlebotomy Point of Care (POC) interface is combining patient orders inappropriately. | DESIGN: Software Design | Sunquest sent an "URGENT-PRODUCT SAFETY NOTICE: PSN-09-06" letter dated November 19, 2009 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were informed that Sunquest Application Interfacing would be contacting their facility for scheduling to correct the issue. |  | 9 | 9 units | Nationwide Distribution. | Terminated | November 06, 2012 | 0 | 9 | Computer
Z-0250-2013 | 63369 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory, version 7.0not available | Sunquest Laboratory | 2 | November 06, 2012 | 2012 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory version 7.0 because when saving an Accession to a non-current Episode/Event, a User-Defined field can be incorrectly replaced by the current Episode/Event User-Defined data. |  | Sunquest Information Systems, Inc. sent a recall letter dated November 3, 2011, to all affected customers. Customers were informed of the problem identified and the actions to be taken. Customers were informed that a change request CR11-00174907 has been created to address this issue. This will be available as part of Sunquest Laboratory version 7.0.1. To request the software correction, customers were instructed to create a Case via the Client Support webpage at www.sunquestinfo.com/supportweb. Then choose a new case, and finally select work queue and select software request. Customers with questions were instructed to call (877) 239-6337 for US and Canadian sites and at +44 (0) 161 335 0562 for international sites.For questions regarding this recall call 520-570-2252. |  | 14 | 14 sites | Worldwide Distribution - USA (nationwide) and Internationally to Ireland | Terminated | November 06, 2012 | 0 | 14 | Computer
Z-0252-2013 | 63270 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | STA Satellite Automated Multi-Parameter AnalyzerThis system is a fully automatic clinical instrument indicated and intended for the performance of tests on human plasma, the results of which aid in the diagnosis of coagulation abnormalities or in monitoring anticoagulant therapy. | STA Satellite Automated MultiParameter Analyzer | 2 | November 06, 2012 | 2012 | Diagnostica Stago, Inc. | Following a single customer complaint that incorrect INR values were occasionally printed for PT tests, even though the primary units (seconds) were correct, the firm investigated and confirmed the potential for printing incorrect secondary units when using an external printer and the customer printout mode on the STA SatelliteÂż. | DESIGN: Software Design | Diagnostica Stago sent an Urgent Field Safety Notice dated June 12, 2012, to all affected customers via email return receipt requested on June 15, 2012. Customers were instructed to follow the instructions in the notification and complete the Customer Verification form and return to the Hotline by fax to 973-644-9348. For additional information, customers were instructed to contact the Hotline at 800-725-0607.For questions regarding this recall call 973-631-1200, ext 2044. |  | 237 | 232 units US; 5 units Canada | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 19, 2013 | 255 | 237 | Computer
Z-0208-2013 | 52796 | chamber, hyperbaric | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | Sechrist Monoplace Hyperbaric Chamber with Electronic Control Module Model 3300E/ERThe Hyperbaric Chamber is intended to provide 100% oxygen to the patient at greater than ambient pressures. | Sechrist Monoplace Hyperbaric Chamber with Electronic Control Module | 2 | November 06, 2012 | 2012 | Sechrist Industries Inc | Sechrist conducted a field upgrade on the hyperbaulic chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of the chamber. | DESIGN: Software Design | Sechrist Industries contacted consignees by phone to notify them of a software upgarde and to schedule appointments to perform the required field upgrades. An approximate summary of the phone notification is provided below."I am contacting you to notify you of a field upgrade we are performing on ourhyperbaric chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of your chamber. The upgrade includes the installation of a new motherboard with revised software, a reset switch, and the exchange of other control panel components. The process will take approximately 2.5 hours and will be performed at no charge to you."For questions regarding this recall call 714-579-8400. |  | 73 | 73 units | Worldwide Distribution - USA (nationwide) and the countries of Bulgaria, and Canada | Terminated | November 06, 2012 | 0 | 73 | Computer
Z-0270-2013 | 63083 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable hospital bed, Prime Bed TL500, Distributed by: Primus Medical LLC, 8401 Southern Blvd, Boardman, OH 44512.Usage: Hospital Bed. | Prime Bed TL500 Bed Frame | 2 | November 08, 2012 | 2012 | Primus Medical LLC | The "Retaining washers & Clevis Pins" provided with the AC -powered adjustable hospital beds may fail under certain circumstances. The pins at the top of the leg were disengaging from the channel. This may cause the bed frame to come off tracks or collapse. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Primus Medical LLC, sent an "Important Notification Relating to Primus Medical Product TL500" letter dated March 29, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to replace the "Retaining washers & Clevis Pins" with the "Clevis Pins & Bow Tie Cotter Pins" in the TL500-001 Service Pack provided with the letter in order to avoid any potential failure and/or bed collapse and to return the" Retaining washers & Clevis Pins" after the replacement procedure is performed. If you have any questions, or need assistance in performing the replacement, call 330-965-9681. |  | 342 | 342 Beds | Nationwide distribution: USA including states of: AZ, IL, KY, MD, MI, OH, OR, PA, TX and VA. | Open |  |  | 342 | Computer
Z-0271-2013 | 63383 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Single Roller Pump 150, Item Number: 10-80-00The Stockert S5/Sorin C5 System is intended to be used during cardiopulmonary bypass for procedures lasting six hours or less. | Stockert S5/Sorin C5 System | 2 | November 08, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of theSorin Group S5 Perfusion System due to reports of a "Fault in motor controller" alarm message displayed on the system panel. This could lead to unintended pump stoppage which may result in permanent impairment or life-threatening injury. An out of specification electrical component was determi | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Sorin Group, sent an "URGENT FIELD SAFETY NOTICE" dated September 19, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. The customers were instructed to follow the provided instructions for the continued safe use of the Perfusion systems until they have been serviced or replaced by Sorin Group. The Sorin Group Service Team will contact the customer to schedule servicing and replacement of the affected product. The customers were also instructed to complete and return the Customer Response Form via fax to 303-467-6502 or by email to yvonne.feyerherm@sorin.com; assure this notice is distributed to all personnel who need to be aware of this notice; and if they have transferred the affected products to a third party, pass this information with them as well as with Sorin Group Customer Service at 1-800-650-2623. For questions regarding this notice, contact Sorin Group USA Customer Service at 1-800-650-2623. |  | 35 | 35 units | Worldwide distribution: USA (nationwide) including states of: AZ, GA, MI, MS, NC, NY, PA, SC, and TX; and countries of: Austria, Australia, Belgium, Canada, People's Republic of China, Columbia, Denmark, Finland, Germany, India, Iran-Islamic Republic of, Ireland, Japan, Lebanon, Morocco, Pakistan, Poland, Russian Federation, Saudia Arabia, Spain, South Africa, Thailand, Turkey and United Kingdom. | Terminated | January 18, 2013 | 71 | 47 | Computer
Z-0282-2013 | 62733 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 4 pump, 220V/240V safety monitor includedThe Sarns" Modular Perfusion System 8000 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Sams Modular Perfusion System 8000 | 2 | November 13, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the arterial and cardioplegia monitors for the Sarns ModularPerfusion System 8000. The malfunctions can be broadly categorized as follows monitor power malfunctions, inability to measure pressure and inaccurate and/or false pressurealarms.If the malfunction of the arterial or cardio | DESIGN: Device Design | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" notice/Safety Advisory Alert dated September 24, 2012 and an addendum to the Operator's Manual to alert all affected Customers. The notice described the product, problem and actions to be taken. The customers were instructed to review this medical device recall notice; assure that all users are aware to this notice; place the addendum in the Operator's Manual just before page 4.1 and complete and return the attached Customer Response Form via Fax to: 1-734-741-6149 or email to:www.termuo-cvs.com.If you have any questions or concerns contact Terumo CVS Customer Service at 1-800-521-1818; Monday - Friday, 8 AM - 6 PM ET. |  | 466 | 466 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY; and countries of: ARGENTINA, AUSTRALIA, BANGLADESH, BRAZIL, CALI COLOMBIA, CALLO PERU, GERMANY, GUAYQUIL ECUADOR, HONG KONG, INDIA, INDONESIA, JAPAN, LEBANON, MADRID SPAIN, MEXICO, NEW ZEALAND, SANTIAGO, SANTIAGO CHILE, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SRI LANKA, TAIWAN, THAILAND, and UNITED ARAB EMIRATES (UAE). | Terminated | July 11, 2013 | 240 | 2256 | Computer
Z-0296-2013 | 63458 | tourniquet, pneumatic | KCY | Pneumatic tourniquet. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ATS 3000 Automatic Tourniquet Systemw/HOSES AND LOP SENSORIntended to be used by qualified medical professionals to temporarily occlude blood flow in a patient's extremities during surgical procedures on those extremities. | ATS 3000 Automatic Tourniquet System | 2 | November 15, 2012 | 2012 | Zimmer, Inc. | Zimmer is conducting a voluntary removal of some Zimmer A.T.S. 3000 tourniquets after receiving complaints of both out-of-box failures and failures during surgical procedures on some of the affected units. During failure the unit will alarm, lock into a non-operational mode, and the display screen will go blank. Based on investigation, this problem appears to be due to a counterfeit chip on the |  | Zimmer sent an Urgent Device Removal Notice letter dated September 5, 2012, to all affected customers. Zimmer Distributors and Sales Leadership were notified via E-mail of the pending recallthe product under recall beginnig September 11, 2012. The notification listed the reason for the recall, risks involved, and requested return of the unit for repair To: Zimmer Surgical, Attn: QAlRA Dept. - Recall, 200 West Ohio Avenue, Dover, Ohio 44622 USA The notice asked to ensure all users be made aware of the recall, remove the affected ATS 3000 tourniquest sytem for repair and return the fax back form/acknowledgement certification to 1-888-429-7380 or email it to zimmer3611@stericycle.com. Questions should be directed to 1-888ÂŽ943ÂŽ5167 between 8 a.m. and 5 p.m. EDT.For questions regarding this recall call 574-372-4807. |  | 204 | 204 | Worldwide Distribution - USA (nationwide) and Internationally to Belgium, Czech Republic, Egypt, France, Germany, Morocco, Netherlands, Romania, Saudi Arabia, Spain, Switzerland, Turkey, and UAE. | Open |  |  | 204 | Computer
Z-0294-2013 | 63660 | pump, infusion, enteral | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump identified with the following list numbers: INFKIT2 and INFORNG. The device is used with disposable PVC administration tubing sets to provide a means for the volumetric delivery of fluids used in enteral applications.The EnteraLite Infinity enteral feeding pump with disposable PVC administration tubing sets and user accessories provide a means for the volumetric delivery of fluids used in enteral applications. | Infinity Pump | 2 | November 15, 2012 | 2012 | MOOG Medical Devices Group | MOOG Medical Devices Group is conducting a voluntary for a limited set of serial numbers of the EnteraLite Infinity and Infinity Orange Small Volume Enteral Feeding Pumps because certain diodes on the PCB assembly may result in insufficient voltage to a portion of the pump resulting in an error 13 message and shutting the pump down entirely. | DESIGN: Component Design/Selection | MOOG Medical Devices Group sent a Urgent Medical Device Recall letter dated November 8, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to contact Moog Medical Devices Group to discussthe process for product replacement at 1-800-970-2337, prompt #7.See Attachment A to determine if you have a serial number that is related to a suspect pump. Determine the number of affected pumps. Notify all of your affected customers of the recall and the protocol for getting a replacement pump. Provide serial numbers, number of product(s) to be returned and a shipping address for replacement product. Execute the pump replacement process discussed in the conversation with Moog Customer Advocacy. Package and return product via shipping label. |  | 8736 | 8,736 | Worldwide Distribution - USA (nationwide) including the countries of Canada and Lithuania. | Terminated | January 30, 2014 | 441 | 8736 | Computer
Z-0200-2013 | 63597 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Brivo XR 285Mobile general purpose radiographic imaging of the human head and body. | Mobile X Ray Series | 2 | November 16, 2012 | 2012 | GE Healthcare, LLC | It was discovered GE Healthcare has become aware of a firmware issue on the mobile radiographic product, Brivo XR285amx, Optima XR200amx, Optima XR22Oamx and Optima XR200amx with Digital Upgrade, related to Technique Accuracy as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-2-54. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will remedy the defect or bring the product into compliance with each applicable Federal and IEC standard. A Service Representative will update the firmware on the system to address the issue.Field Modification Instructions (FMI) 10870, 10872, and 10873 are describing the rework plan the GEHC Engineers will make to the imaging systems compliant with the applicable performance standard.Further questions please call (262) 513-4122. |  | 364 | 364 total units installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 384 | 364 | Computer
Z-0299-2013 | 63467 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | bioMerieux BacT/ALERT Control Module, catalog number 210147, Clinical instrumentThe product is most commonly used to isolate and recover specimens drawn from adult and older pediatric patients that are acutely ill with signs and symptoms suggestive of blood stream infection, which can be due to a variety of sources including lung, abdomen, urinary tract, endovascular, and skin. | bioMerieux BacT/ALERT Control Module | 2 | November 16, 2012 | 2012 | Biomerieux Inc | The product may mislabel patient information on culture specimen bottles due to a software miscommunication. | DESIGN: Software Design | BIOMERIEUX sent an Urgent Product Correction Notice to all affected customers starting on October 16, 2012, The letter identified the affected product, the problem, and the action to be taken by hte customer. The firm provided modified instructions for use to clarify how to properly load bottles and to provide appropriate cautions and warnings concerning bottle loading. Customers were instructed to ensure that the letter and guidance on how to properly scan and load bottles is distributed, and training is provided to all appropriate personnel within their organization. Check for lost bottles in their system; fill out and return the Acknowledgement Form in Attachment A by fax to confirm receipt of the notice. Customers were instructed to contact their local bioMerieux Customers Service Representative with any questions.For questions regarding this recall call 919-620-2682. |  | 70 | 70 modules | Worldwide Distribution - USA including AR, AZ, CA, FL, KS, LA, MD, MN, MT, NC, NV, NY, OR, PA, RI, SC, TX, VA and WA. Internationally to Argentina, Armenia, Austria, Bangladesh, Bosnia and Herzegowina, Brazil, Bulgaria, Burkina FASO, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DIVOIRE, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA, MALTA, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UNITED STATES, URUGUAY, VENEZUELA, VIETNAM, PALESTINE, AND SERBIA. | Terminated | October 28, 2013 | 346 | 6400 | Computer
Z-0385-2013 | 63607 | orthopedic stereotaxic instrument | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | The RIO (TGS 2.), Model # MAKO TGS 2.0 (Part No. 204000). The RIO is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The RIO is indicated for use in surgical knee and hip procedures, in which the use of stereotactic surgery may be appropriate, and where reference to rigid anatomical bony structures can be identified relative to a CT based model of the anatomy. These procedures include: Unicondylar knee replacement (UKA) and/or Patellofemoral knee replacement (PKA) and Total Hip Arthroplasty (THA). | The RIO (TGS 2.), | 2 | November 19, 2012 | 2012 | Mako Surgical Corporation | MAKO Surgical Group recalled their RIO System software, version 2.4 and is implementing software version 2.5 to address a software functional issue with the existing version of the system's software. Loss of tactile feedback constraining the cutting burr has been reported. | DESIGN: Software Design | Mako Surgical Corporation sent an Important Medical Device Correction letter dated August 31, 2012, to all affected customers. The letter included specifics regarding the software issue and how the correction will take place. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them to establish the appropriate time for the upgrade .Customers were asked to return by fax an Acknowledgement Form to 954-927-0446, ATTN: Quality Assurance. For questions customers were instructed to call 954-628-0447.For questions regarding this recall call 954-927-2044, ext 605. |  | 137 | 137 devices | Worldwide Distribution - USA including IL, OR, IA, AZ, CA, CO, FL, MS, OK, PA, WV, NC, TX, NY, GA, NJ, OH, WA, VA, MA, MD, WI, MI, NV, MO, AL, UT, RI, AR, LA, ID, IN, and TN and Internationally to Japan, Scotland, Italy, Singapore, Turkey, Hong Kong, and Korea. | Terminated | May 22, 2013 | 184 | 137 | Computer
Z-0403-2013 | 63504 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | qube Compact Monitor (Patient Monitor) 91390. The monitor is 26.2 cm x 31.5 cm x 13.2 cm and weighs 4.1 kg (10.3 inches x 12.4 inches x 5.2 inches - 9 lbs). It has a 26.2 cm (12.1 inches) LCD display with 1024 x 768 resolution. The Spacelabs Healthcare Qube Compact Monitor (91390), functioning as either bedside or central monitors; passively displays data generated by 2 Spacelabs Healthcare parameter modules, Flexports interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. Key monitored parameters available on the Model 91390 when employing the Spacelabs Command Module, consist of ECU, respiration, invasive and noninvasive blood pressure, Sp02, temperature and cardiac output. Additional parameters and interfaces to other systems are also available depending on the parameter modules employed. The Qube is intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other physiologic monitors via Flexport interfaces. These devices determine a) when an alarm condition is violated; b) the alarm priority (i.e. high, medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. The Qube may also function as a generic display or computer terminal. As a generic display or terminal, the patient monitors allow network-based applications to open windows and display information on other networked monitors. The Qube is also designed to communicate with a variety of external devices such as displays, network devices, serial devices, user input devices, audio systems, and local/remote recorders. The Qube is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. | Qube Compact Monitor Model 91390 | 2 | November 20, 2012 | 2012 | Spacelabs Healthcare, Llc | Spacelabs Healthcare has learned through severalreports that the input circuit may fail and the monitorwill go on battery mode. The monitor will functionnormally until the batteries become depleted and thenthe monitor will shut down. In addition the monitor mayfail to turn on due to an internal short. | PRODUCTION CONTROLS: Process Control | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated October 19, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact Spacelabs Healthcare at 1-800522-7025 option 2 for questions regarding this notice. |  | 742 | 742 total devices distributed; 530 with the U.S. and 212 Internationally. | Worldwide Distribution.On 4/16/13 the firm advised FDA that they added serial number 1390-000055 and 1390-000913. The two additional serial numbers were added because they were provided as replacements to a French customer whose devices had been returned to support the firm's investigation. The customer had already received the Urgent Medical Device correction, customer letter and required Spacelabs to provide these devices with the understanding that they would require the field corrective action | Terminated | May 22, 2013 | 183 | 742 | Computer
Z-0435-2013 | 63543 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System Platform (APS)The TerumoÂŽ Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | 100/120V AC, Advanced Perfusion System Platform (APS) | 2 | November 27, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular System (TCVS) has received reports of a situation where users experienced a total loss of functionality for some System 1 units. The reports indicate that the units went blank and shut down with no sign of power and battery backup did not initiate. The result is all pumps stop, with no safety system functionality, and the battery would not be activated. The user would be limi | DESIGN: Device Design | Terumo Cardiovascular Systems sent a Urgent Medical Device Recall Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Terumo CVS is alerting all users of Terumo System 1 of the reports of malfunction. Customers were instructed to: 1. Review this Medical Device Safety Advisory.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/email address indicated on the form.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service 1-800-521-2818, Recall Fax 1-734-741-6149Customer Service Hours: Monday thru Friday, 8 AM - 6 PM ET. | Contact the recalling firm for information | 1647 | 1647 units total | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | August 30, 2013 | 276 | 1647 | Computer
Z-0448-2013 | 63212 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima XR 220amx 30kw Mobile X-ray SystemThe GE Automatic Mobile X-Ray (AMX) Series: Brivo XR2S5amx,Optima XR200amx, Optima XR22Oamx are intended to take exposures utilizing film or computed radiography (CR), however the Optima XR22Oamx utilizes the GE Wireless Detector, which is intended to replace radiographic film screen systems in all general purpose diagnostic procedures, for digital radiography (DR). | GE Healthcare | 2 | November 29, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR22Oamx and Optima XR200amx with Digital Upgrade. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will without charge remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication through a GE Healthcare field engineer site visit. Field Modification Instructions (FMI) 10868 describes the rework plan the GEHC Engineers will make to the imaging system compliant with the applicable performance standard.This CAP appears to adequately address the problem and is hereby approved.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 217 | 217 total Installed in the US | USA (nationwide) | Terminated | January 10, 2013 | 42 | 217 | Computer
Z-0451-2013 | 63649 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly-Clark Patient Warming System - Model 1000 Control UnitThe product is a thermal regulating system, indicated for monitoring and controlling patient temperature. | KimberlyClark Patient Warming System | 2 | November 29, 2012 | 2012 | Kimberly-Clark Corporation | Error 85 alarm on the Model 1000 Control Unit was occasionally triggered inappropriately. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Kimberly Clark sent an Urgent Field Corrective Action letter dated September 16, 2010, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were asked to confirm whether their facility as the affected product by faxing the completed Field Corrective Action Response Form to 920-380-6682. Kimberly-Clark will arrange to perform the upgrade at the customer's facility. Customers with questions should call the Field Corrective Action Coordinator at 770-587-7400.For questions regarding this recall call 770-587-8393. |  | 400 | 400 units | Nationwide Distribution including AZ, AK, CA, CO, DE, DC, FL, GA. ID, IL, IN, IA, KS, KY, LA, MD, ME MA, MS, MO, MT, NE, NJ, NM, NV, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WV, and WI | Terminated | December 20, 2012 | 21 | 400 | Computer
Z-0447-2013 | 61835 | cap,cooling (infants) | MXM |  |  | Neurology |  | The product has the Catalogue/Part number 60010.Olympic Cool-Cap System is a selective head cooling system for treatment of hypoxic-ischemic encephalopathy (HIE) in infants. A cooling water cap is placed on the infant's head and the body is warmed using radiant warmers. The Cool-Cap provides 72 hours of monitored cooling. The device is labeled in parts: "***OLYMPIC COOL-CAP Cooling Unit***Natus Medical Incorporated 5900 First Avenue South ***Made in USA***Control Unit Model No. 60010***Software in this product is***Technical Service***" | Natus Medical Olympic CoolCap System | 1 | November 30, 2012 | 2012 | Natus Medical Incorporated | The Olympic Cool-Cap's Control Module has experienced a frozen screen during use. When this occurs, the on-screen information remains on display, but the system is no longer providing cooling treatment to the infants. The visible clock displayed at the upper right hand corner of the screen stops advancing. | DESIGN: Device Design | The firm, Natus Medical Incorporated, sent its customers an "URGENT: MEDICAL DEVICE" letter dated May 16, 2012 with an attached Field Safety Notice. The letter describes the product, problem and actions taken. Customers can call Natus Technical Service at 888-496-2887 for any questions about this recall. |  | 324 | 324 units | Worldwide distribution: USA (nationwide) including states of: AK, AR, AZ, CA, CO, DE, FL, GA, HI, ID, IL, IN, KY, MA, MI, MN, MS, ND, NE, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, WA, Puerto Rico, and countries of: Argentina, Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Ecuador, Egypt, Germany, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Kuwait, Latvia, Malaysia, Mexico, Morocco, Panama, Puerto Rico, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open |  |  | 324 | Computer
Z-0465-2013 | 63521 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza Radiological Image Processing SystemUsage: Radiological Image Processing System | Siemens syngo.plaza Radiological Image Processing System | 2 | December 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of an unintended behavior when using syngo.plaza and issued an Update Instruction. The Update Instruction SY050/12/S provides a software update and addresses the following issues: - Patient Rename messages for the same patient are being continually processed for several days. This fills up the 'pcvdbsrv' log and potentially affects performance of the system. - Data could | DESIGN: Software Design | The firm, Siemens, sent "Update Instruction SY050/12/S" on June 20, 2012 to its consignees/customers. The Update Instructions described the product, problem and actions to be taken. It also provides a software update.If you have any questions call 610-219-4834 or email: anastasia.mason@siemens.com. |  | 15 | 15 | Nationwide distribution: USA including states of: CA, FL, GA, MA, MN, MO, NJ, NY, OH, and WI. | Terminated | January 14, 2014 | 407 | 15 | Computer
Z-0467-2013 | 63646 | assembly, knee/shank/ankle/foot, external | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | BiOM Rechargeable Li Ion 22.2 V Battery and BiOM Lithium Battery Charger sold as an accessory to the BiOM lower leg system:Product Number: 2001249 BiOM BatteryProduct Number: 2001240 BiOM Charger | BiOMBiOM Rechargeable Li Ion 22.2 V Battery and BiOM Lithium | 2 | December 04, 2012 | 2012 | iwalk inc | Battery may over-heat during charging and smoke |  | iWalk notified customers directly by sales representatives by telephone on October 26, 2012, and instructedthem to stop use of the battery and charger and return them to iWalk. Accounts provided with Return Material Authorizations (RMA).For questions regarding this recall call 781-761-1560. |  | 30 | 21 Batteries; 9 Chargers | Nationwide Distribution including OH, NC, and NY. | Open |  |  | 30 | Computer
Z-0475-2013 | 63619 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software Version R2.0.2Product Usage:This system is used for making X-ray exposures for diagnostics | Philips Digital Diagnost R2.0.2 | 2 | December 05, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost R2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "Right patient side" marker, although this can appear on the left patient side. | DESIGN: Software Design (Manufacturing Process) | Philips Healthcare initiated a recall on October 25, 2012 via Urgent - Field Safety Notice for the Digital Diagnost X-ray System. The letter identified the affected product, problem and actions to be taken. The letter informs customers that Philips will supply a software update kit to version 2.1.4. For questions please contact your local Philips representative 1-800-722-9377. |  | 623 | 623 units | Worldwide Distribution - USA Nationwide and the countries of Argentian. Australia, Austria, Belgium, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Eastonia, Finland, France, Germany, India, Indonesia, Iraq, Italy, Japan, Latvia, Lithuania, Lybia, Malta, Mexico, Netherlands, New Zealnd,Norway, Oman, Polnad, Protugual, Quatar, Romania, Russia, Saudi Arabia,Singapore, Slovakia, Soth Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Open |  |  | 623 | Computer
Z-0480-2013 | 63495 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilette Mira, intended use as Mobile x-ray system. | Siemens Mobilette Mira | 2 | December 06, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a risk of lost images during examination. | DESIGN: Software Design | The firm initiated their recall/correction of this product on September 9, 2012 by sending a Field Safety Notification letter with Update Instruction XP054/12/S to all consignees. The letter described the steps to be taken to avoid the potential risk and included a response form to be returned to the recalling firm. |  | 21 | 21 | Nationwide distribution: USA including states of: GA, ID, IL, IA, NE, NC, OH, SD, TX, and WI. | Terminated | January 15, 2014 | 405 | 21 | Computer
Z-0484-2013 | 63312 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Data Station AMPLILINK software is an instrument user interface/ data management software used with COBASÂŽ AmpliPrep instrument, COBASÂŽ TaqManÂŽ analyzer, COBASÂŽ TaqManÂŽ 48 analyzer, COBASÂŽ AMPLICORÂŽ analyzer, and Cobas p 630 instrument for sample preparation, amplification and detection of PCR (Polymerase Chain Reaction) molecular diagnostic analytes. | AMPLILINK | 2 | December 07, 2012 | 2012 | Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. has become aware of a device-associated diagnostic software issue. | DESIGN: Software Design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification letter dated July 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Molecular Systems, Inc. has disabled this software on affected Roche platforms and discontinued the distribution of the affected software. Contact Roche Diagnostic Technical Support, 24 hours a day, seven days a week at phone number : 1-800-428-2336 if you have questions regarding this notification. For questions regarding this recall call 908-253-7569. |  | 884 | 884 units | Nationwide Distribution | Open |  |  | 9250 | Computer
Z-0481-2013 | 63518 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps with Plug-and-Play Modules that provide wired Ethernet and wireless local area networking; Hospira, Inc., Lake Forest, IL 60045; pump list number 11971, module list numbers 12101, 12102 and 12393; pump list number 11973, module list number 12380; pump list number 12391, module list numbers 12097 and 12680Product Usage: The Plum A+ Infusion System is a dual-line volumetric infusion pump that provides precise delivery of multiple i.v. therapies across the general spectrum of clinical care, including ICU/CCU, OR, PACU, ER, Med/Surg, Pediatrics, Outpatient clinics and Home care | Plum A Single Channel Infusion Pumps | 2 | December 07, 2012 | 2012 | Hospira Inc. | The volume control knob on some Plum A+ single channel infusers (located on the back of the infuser) may not function as described in the System Operating Manual - the direction for Loud and Quiet may be reversed. The manual indicates a clockwise rotation of the knob decreases the volume; however, on some of the devices the alarm volume increases when the knob is turned clockwise. | DESIGN: Device Design | The firm, Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters dated October 26, 2012 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to the firm via fax at 1-877-650-8667. The letter states that the System Operating Manual will be updated to clarify the instructions for setting the volume, and that the updated manuals will be sent as soon as they are available. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for additional information and technical assistance. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Australia, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Lavenia, South Africa, Spain, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates and Vietnam. | Open |  |  |  | Computer
Z-0495-2013 | 63709 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center/Client and Philips IntelliVue Information Center iX The intended use of the Information Center Software is to display physiologic waves, parameters, and trends, format data for strip chart recordings and printed reports, and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. The Information Center Software provides for the retrospective review of alarms, physiologic waves and parameters from its database. An additional intended use of the Information Center Software is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips IntelliVue Information Center/Client and Philips IntelliVue Information Center iX | 2 | December 10, 2012 | 2012 | Philips Healthcare Inc. | Internal audio cable may have diminished retention force due to damaged contacts, which could lead to intermittent or loss of audible alarms. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips sent an URGENT - Medical Device Correction letter dated November 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is conducting a voluntary correction to replace the audio cable on affected devices. Customer were instructed to follow the "Action to be taken by CustomerUser" section of the instructions. Customers with wquestions were instructed to contact their Phililps representataive or contact the Customer Care Solutions Center at (800) 722-9377.For questions regarding this recall call 978-687-1501. |  | 3165 | 3165 | Worldwide Distribution - USA (nationwide) and Canada. | Open |  |  | 3165 | Computer
Z-0513-2013 | 63614 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6/7 Patient MonitorManufactured in China for : Mindray DS USA, Inc.800 MacArthur Blvd.Mahwah, NJ 07430 USA.Intended to be used for monitoring, displaying, reviewing, storing and alarming of multiple physiological parameters including ECG (3-lead or 5-lead or 12-lead selectable). arrhythmia detection, ST Segment analysis, heart rate (HR) , respiration rate (RESP), temperature (TEMP), pulse oxygen saturation (SpO2), pulse rate (PR), non-invasive blood pressure (NIBP), invasive blood pressure (IBP), carbon dioxide (CO2), anesthetic gas (AG),impedance cardiograph (ICG), bispectral index (BIS), cardiac output (C.O.) and respiration mechanics (RM). | Mindray DS USA Inc. | 2 | December 11, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is an issue with the DPM 6 and DPM 7 monitors where the following features may be unavailable : Full Disclosure, Drug, Hemodynamic Calculations, Renal Calculations, Oxygenation Calculations, and Ventilation Calculations. | DESIGN: Device Design | Mindray sent a "PRODUCT CORRECTIVE ACTION" letter dated August 11, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their Mindray Service Representative at 1-800-288-2121 for assistance with this issue. | Contact the recalling firm for information | 124 | 124 units | Nationwide Distribution-USA (nationwide) including the states of WA, IL, FL, MN, SC, NV, MI, TN, PA, LA, NE, TX, and IA. | Terminated | August 02, 2013 | 234 | 124 | Computer
Z-0524-2013 | 63655 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module with v2.0 software used with Solar 8000M/l, Transport Pro or Carescape B650/B850 host monitor. The Patient Data Module (PDM) is intended to provide physiologic parameter data on adult, pediatric and neonatal patients during bedside and transport patient care episodes. | Carescape Patient Data Module v2.0 | 2 | December 14, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of two issues with the Patient Data Module using Version 2.0 software that may impact patient care. | DESIGN: Software Design | Consignees were sent on 2/17/2012 three GE Healthcare "Important Product Information" letters dated 10/17/ 2011. The letters were addressed to Healthcare Administrator/Risk Manager, Director of Biomedical Engineering, and Chief of Nursing. The letters described the Issue, Instructions, Affected Product Details, Product Correction and Contact Information. |  | 10250 | 10,250 | Worldwide distribution, including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, LA, KS, KY, LA, ME,MD, MA, MI, MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN,TX, UT,VA, WA, WT, WY and the countries of the UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAK IA, SINGAPORE, SAUDI ARABIA, QATAR, POLAND, PERU, PANAMA, NORWAY, NEW ZEALAND, NETHERLAND, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, ITALY, INDONESIA, INDIA, IRELAND, GERMANY FRANCE, FINLAND, EGYPT, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHILE CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | December 03, 2013 | 354 | 10250 | Computer
Z-0494-2013 | 63650 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000/4000/5000.The Dash 3000/4000/5000 patient monitor is intended for use under the direct supervision of a licensed healthcare practitioner. The intended use of the system is to monitor physiologic parameter data on adult, pediatric and neonatal patients. The Dash 3000/4000/5000 patient monitor is designed as a bedside, portable, and intra-hospital transport monitor that can operate in all professional medical facilities including but not limited to: emergency department, operating room, post anesthesia recovery, critical care, surgical intensive care, respiratory intensive care, coronary care, medical intensive care, pediatric intensive care, or neonatal intensive care areas located in hospitals, outpatient clinics, free-standing surgical centers, and other alternate care facilities. Physiologic data includes, but is not limited to: electrocardiogram, invasive blood pressure, noninvasive blood pressure (NBP), heart rate, temperature, cardiac output, respiration, pulse oximetry, carbon dioxide, bi-spectral index, impedance cardiography, oxygen, and anesthetic agents as summarized in the operator's manual. The Dash 3000/4000/5000 patient monitor is also intended to provide physiologic data over the UNITY NETWORKTM indirectly to clinical information systems (via our Enterprise Gateway) and allow the user to access hospital data at the point-of-care. The information can be displayed, trended, stored, and printed. The Dash 3000/4000/5000 patient monitor was developed to interface with nonproprietary third party peripheral devices that support serial data outputs. | GE Healthcare, Dash 3000/4000/5000. | 2 | December 16, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to an unexpected loss of CustomAutomatic NBP Measurement cycling associated with the Dash 3000/4000/5000 patient monitors. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | GE Healthcare sent a Important Product Information Letter dated June 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If your facility is using the Custom Auto NBP Measurement feature in conjunction with Automatic View On Alarm and/or View Other Beds options, GE Healthcare recommends that you stop using the Custom Automatic NBP Measurement feature and use the Auto NBP function instead.All Dash3000/4000/5000 patient monitors with software version 6.0 and later with the networking option enabled may be affected. You can determine the software version of the Dash monitors by selecting MORE MENUS |  | 119260 | 119,260 devices | Worldwide Distribution -- USA (nationwide) including the countries of AUSTRALIA, ANGOLA,ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM and URUGUAY | Open |  |  | 119260 | Computer
Z-0531-2013 | 63713 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand ARIA Radiation Oncology, ARIA Oncology Information System Radiation Oncology, Import Export Application, Model Number: HIT, Affected Versions: ARIA version 11.0 below 11.0.55, Build 11.0.28 and 11.0.34, Reference/FSCA Identifier: CP-09798;Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA | Varian brand ARIA Radiation Oncology | 2 | December 17, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian Medical Systems has identified an anomaly with the ARIA Oncology Information System Radiation Oncology [ARIA RO] DICOM import/export functionality where a 360 degree arc field may convert to static field during DICOM export or import. | DESIGN: Software Design | Varian sent an Urgent Medical Device Correction Urgent Field Safety Notice dated November 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that the plan was not converted unintentionally during import/export. Customers would be contacted by a Customer Service Representative when the correction was available to schedule installation on their system. Customers with questions were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager.US and Canada 1-888-827-4265.For questions regarding this recall call 650-424-5731. |  | 264 | 264 units | Worldwide Distribution | Open |  |  | 264 | Computer
Z-0528-2013 | 63652 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V12 Size: 320 mm X 320 mm X 450 mm N.W.: 3 kg G.W.: 6 kg Qty:1 Manufactured in China Mindray V21 Size: 515 mm X 335 mm X 685 mm N.W.: 8 kg G.W.: 11 kg Qty:1 Manufactured in China One product, V Series monitor, available in two sizes: The V12 has a 12 inch screen, the V21 has a 21 inch screen. The V Series Monitor is a Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. It also monitors of the following human physiological parameters: 1) ECG waveform derived from 3, 5, 6 and 12 lead measurements, 2) Heart Rate, 3) Pulse Oximetry (Sp O2), 4) ST Segment Analysis, 5) Arrhythmia Detection, 6) Non Invasive Blood Pressure (NIBP), 7) Invasive Blood Pressure (IBP), 8) Cardiac Output (CO), 9) Respiratory Gasses, 10) Respiration Rate, 11) Temperature, It is intended for intra hospital use under the direct supervision of a licensed healthcare practitioner. | Mindray | 2 | December 17, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the V Series Monitor where the monitors touch screen may stop responding to touch. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Mindray issued an "Urgent: V Series Monitoring System Corrective Action" letter dated 9/24/2012 via certified mail with return receipt verification to all hospitals.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 to arrange for relplacement of the affected product. For questions regarding this recall call 201-995-8407. | Contact the recalling firm for information | 206 | 206 units (180 units domestic, 26 units foreign) | Worldwide Distribution - USA (nationwide) and internationally to Great Britian, Colombia, and Australia. | Terminated | August 15, 2013 | 241 | 206 | Computer
Z-0560-2013 | 61593 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Advantage Workstation Volume Share 2 (versionAW4.4_04) with Volume Viewer 3 and 3.1 (version7.0 to version7.5 and version 8.1) and only when or more of these applications is installed:" VesselIQ Xpress or AVA Xpress" CardIQ Xpress Pro or Plus" CardEP" CardIQ Fusion PET or SPECTAdvantage Workstations Volume Share (v.AW4.3) and AW4.2 (version4.2-05 and above) provided with Volume Viewer 2(version6.0 to version6.11) and only when one or both of these applications is installed:" CardIQ Xpress Pro or Plus" CardEPAdvantage Workstations AW Volume Share 2 (AW version 4.4).ONLY IF ONE OF FOLLOWING APPLICATIONS IS INSTALLED:ÂŽ VesselIQ Xpress or AVA XpressÂŽ CardIQ Xpress Pro or PlusÂŽ CardEPÂŽ CardIQ Fusion PET or SPECTprovided with Volume Viewer 3 and 3.1 from versions 7.0 to 7.5 and 8.1.To verify the version installed on your system: Select the Admin menu from the Patient list page. Then select Display Configuration . You will find the version of your system in Installed Application(s) . This version will appear as vxtl_7.x or vxtl_8.x. Intended use: K993792: Smart Vessel Analysis (Smart VA) is a software post-processing option for the Advantage Workstation (AW) platform, which can be used in the analysis of 3D angiography data. It provides a number of display, measurement and batch filming/archive features and will aid physicians in studying user-selected vessels for stenosis analysis, pre/post stent planning and directional vessel tortuosity visualization. K031261: Card EP is a post processing software option for the Advantage Workstation (AW) Platform. This product can be used for the analysis of CT angiographic images for the assessment of the heart to include the atria, pulmonary veins, and coronary sinus. It provides quantitative analysis tools which include a number of display, measurement and model export capabilities. This product can be used to aid trained physicians in the visualization and assessment of cardiac anatomy. K041267: CardIQ Analysis III is a CT image analysis software package, which allows the visualization of 2D and 3D medical image data of the heart derived from DICOM 3.0 compliant CT scans for the purpose of cardiovascular disease assessment. It provides functionally for 2D/3D rendering, assessment of calcified and non-calcified plaque to determine the densities of the plaque within a coronary artery, ventricular function of the heart, and measurement tools to detect coronary artery stenosis. This product can be used to aid a trained physician to process, render, review, archive, print and visualizing cardiac anatomy and coronary vessels. CardIQ Analysis II will run on the AW workstation, scanner operator console and PACS system. K061370: CardIQ Fusion is intended to provide an optimized non-invasive application to analyze vascular anatomy and pathology, aid in the assessment of functional data e.g. PET perfusion, and aid in tailoring treatment plans based on the fused anatomical and functional information. Anatomical data could be from a set of Computed Tomography (CT) Angiographic images while functional data could be from PET, SPECT, or processed CT data. CardIQ Fusion is a software post-processing package for the Advantage Workstation (AW) platform, PET/CT and CT scanners and PACS reading stations. It is an additional tool for the analysis of 3D CT angiographic cardiac images/data providing a number of display, measurements and batch filming/archive features to study user-selected vessels. Also included is the capability to visualize reformatted CT/PET/SPECT perfusion and viability data. Finally, it provides different ways of visualizing CT anatomy fused with PET/SPECT functional information. With CardIQ Fusion, clinicians have the opportunity to overlay functional information over the (CT) anatomy of a patient's heart, and thus, they can potentially tailor their decision for that particular patient. CardIQ Fusion provides the visualization of the vessels in severa | Advantage Workstation | 2 | December 19, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has became aware of two issues with the Advantage Workstation.1) A possible mismatch between the label of the tracked vessel and the underlying image associated with the cardio- vascular applications of the Advantage workstation that may impact patient safety.2) A refresh problem of the stenosis / aneurysmmeasurement tools associated with the cardio-vascular applications of |  | GE Healthcare sent two Urgent Medical Device Correction letters dated May 23, 2012, to all affected customers. Each letter addressed a separate issue. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. Customers were instructed to use the following precautions until corrected software was installed:1. Please avoid the use of leading or trailing spaces in Vessel Names2. Be sure to press "Show Tracking" after depositing new bifurcation points3. When eating a multi-point vessel tracking, remove and replace the entire branch, not only the final end-point.4. When analyzing multi-phase cardiac images, do not delete vessel branches from a phase other than the one in which it was created.Customers were instructed that the new application version that addresses this issue will soon be installed on all systems. Customers with questions or concerns were instructed to call the following:United States 1-800-437-1171.Japan 0120-055-919For other countries they were instructed to contact their GE Healthcare Service Representative.For questions regarding this recalla call 262-513-4122. | Contact the recalling firm for information | 548 | 548 (118 USA; 430 OUS) | Worldwide Distribution - USA including AL, CA, CO, FL, GA, IL, IN, IA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX , UT, VA, WI and Internationally to the following countries:VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SOUTH AFRICA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIA, ROMANIA, REUNION, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, NORWAY, NIGERIA, NEW CALEDONIA, NEPAL, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA ALGERIA. | Terminated | January 17, 2013 | 29 | 548 | Computer
Z-0565-2013 | 63714 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators are anesthesia gas machine.In most cases, the customer ordered both a BleaseFocus anesthesia workstation and a 700/900 ventilator. In some cases, the customer may order a 700/900 stand-alone ventilator would could be used with a BleaseFocus anesthesia workstation at a later date. | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators | 2 | December 19, 2012 | 2012 | Del Mar Reynolds Medical, Ltd. | The BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators were shipped with manuals that describe a secondary O2 Supply Failure alarm which has not been incorporated into the system. The manuals were prematurely updated for a feature that is not yet available. | PRODUCTION CONTROLS: Error in Labeling | Spacelabs has sent the Urgent -Medical Device Correction-Spacelabs Healthcare BleaseFocus Anesthesia Workstation Blease 700/900 Series Ventilators letter, dated 16 November 2012, to US customers.On November 26, 2012, a customer letter (translated as necessary) wasemailed to all international subsidiaries and distributors of record.Customers are advised that the primary alarm will sound only in the event of an O2 supply failure and they will be provided with the corrected user and service manuals that do not reference a secondary O2 supply failure alarm.Customers with questions can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. | Contact the recalling firm for information | 3 | 3 in US and 3121 Outside US | Worldwide Distribution--USA (nationwide) including states of North Carolina and Oklahoma and the countries of Algeria, Argentina, Bahrain, Bangladesh, Bolivia, Chile, China, Columbia, Dominican Republic, Ecuador, Egypt, Finland, Ghana, Great Britian, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Jamaica, Jordan, Lebanon, Libya, Malaysia, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Paraguay, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sudan, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates , Uruguay, Venezuela, and Vietnam. | Terminated | July 11, 2013 | 204 | 3 | Computer
Z-0479-2013 | 63394 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System, Extended Brilliance Workspace (EBW), Model #728260 Philips Healthcare, Highland Heights, OH The Extended Brilliance Workspace (EBW) is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Computed Tomography XRay System, Extended Brilliance Workspace ( EBW) | 2 | December 21, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issued a Field Change Order to update Extended Brilliance Workspace (EBW) systems to software version 4.5.5. The update will provide a number of improvements to address nonhazardous reliability issues which will improve the overall quality and functionality of the EBW. | DESIGN: Software Design (Manufacturing Process) | Philips sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Philips is issuing a system update to all of the Extended Brilliance Workspace (EBW) systems currently running software version 4.5.4. The update will provide a number of improvements to the overall quality and functionality of the EBW systems.If you need further information or support concerning this issue, please contact your local Philips representative. In the United States and Canada, please call the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer; in all other countries please dial your local Philips Healthcare office. |  | 265 | 265 units | Worldwide Distribution--US (nationwide) including the states of AR, OH, OR, PA, TX VA & WI., and the countries of Algeria, Australia, Brazil, China, Ecuador, Finland, France, Germany, Hong Kong, India, Italy, Lebanon, Libyan Arab Jamahiriya, Russia, South Korea & United Kingdom | Terminated | December 24, 2013 | 368 | 265 | Computer
Z-0609-2013 | 63803 | [6 empty] | Diamigo i-Phone App The Diamigo app was intended for use as an educational tool in the management of diabetes. | Sanofi Aventis Diamigo iPhone App | 2 | December 22, 2012 | 2012 | sanofi-aventis US, Inc. | Diamigo, a software application, was released through the i-Phone Global Store which inadvertently allowed global access to the application as opposed to the original intended restriction of access to Brazil. |  | Apple sent a message ("push notification") on September 21, 2012 to its customers. The message describes the product, problem and actions to be taken. The customers were instructed not to use the app to calculate insulin does and should delete if from their device(s). If you have additional questions about the app or this message, contact Sanofi's offices in your country. If you are in the U.S. contact Sanofi US Medical Information at 1-800-633-1610, for Brazil please contact 0800 703 0080. |  | 1600 | 1600 downloads (140 US; 1460 Foreign) | Worldwide distribution. | Open |  |  | 1600 | Computer
Z-0614-2013 | 63814 | clinical sample concentrator | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Instrument (CAP) with AMPLILINK (AL) Software v.3.3.5Automated sample processing instrument used with CAP/CA and CAP/CTM.Roche Molecular Systems, Inc.1080 US Highway 202 SouthBranchburg, NJ 08876USAMade in Switzerland | Roche COBAS AmpliPrep Instrument (CAP) with AMPLILINK (AL) | 2 | December 23, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain COBAS AmpliPrep instruments using AMPILINK software v.3.3.4 had the sample clot detection flagging feature disabled during quality control release resulting in samples that may not be flagged as having clots. | DESIGN: Device Design | Roche sent an Urgent Medical Device Correction notification/Fax Back forms dated 10/11/12 to potentially affected customers. The letter identified the affected product, the issue, clinical significance, workaround, and required actions, along with contact information. Customers are to use the workaround until their Roche FSE has performed a Master Initialization on their affected instruments. Also, the letter should be filed for future reference. Questions should be directed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 7 | 7 units | Distributed in the states of GA, CA, NC, and MN. | Terminated | September 11, 2013 | 262 | 7 | Computer
Z-0606-2013 | 63123 | ventilator, emergency, powered (resuscitator) | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Giraffe and Panda T-Piece Resuscitation System labeled ***T-Piece Neonatal Patient Circuit Kit Disposable with Size 1 Mask***M1091335 QTY: 10 T-Piece Disposable Circuit, QTY: 10, M1091316 T-Piece Disposable Circuit Kit with Mask Size O; M1091365 QTY: 10 T-Piece Disposable Circuit Kit with Mask, Size 1, Lot 1981713 *** Ohmeda Medical, A Division of Datex-Ohmeda, Inc. A General Electric Company Laurel, MD 20723 USA MADE IN USA***Provides the basic equipment required for pulmonary resuscitation of infants. | Giraffe and Panda Warmer TPiece Resuscitation System | 1 | December 26, 2012 | 2012 | GE Healthcare, LLC | Disposable T-Piece circuits used with resuscitation systems are unable to achieve maximum Peak Inspiratory Pressure (PIP) during the pre-use check out procedure. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter, dated June 26, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: Inspect existing stock of T-piece circuits to look for the Red inspection stamp, the T-pieces with the blue stamp are acceptable for clinical use. Separate affected circuits (with Red inspection stamp) from unaffected stock (with Blue inspection stamp). Destroy affected circuits that have the Red inspection stamp. Fill out the customer acknowledgement form and fax or forward to the GE contact address on the form. Further, it is recommended that healthcare professionals continue to emphasize the recommended pre-use checkout practices, and do not use a T-piece patient circuit that does not reach the maximum PIP level of 45 +/- 5 cmH20. Clinicians should always use the built-in Airway Pressure Manometer to verify the PIP value. For questions or concerns regarding this notification, customers should call the following phone number: 1-800-345-2700. |  | 209804 | 209,804 circuits | Worldwide Distribution - USA, including DC and the states of AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, IL, KY, KS, LA, MA, MD, ME, MI, MN, MO,MS, MT, NC, NE, NJ, NM, NY, NV, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Albania, Australia, Austria, Bahrain, Bangladesh, Belgium, Canada, China, Czech Rep., Egypt, Finland, France, Germany, Great Britain, Greece, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, S. Korea, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and Arab Emirates, | Open |  |  | 209804 | Computer
Z-1042-2012 | 60866 | stimulator, caloric-water | ETP | Air or water caloric stimulator. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Leica M822 Surgical Microscope Device Controller Software (MDC version 3.6, pack F, affecting the control unit component, Article #10448446, in combination with a Leica M822 Optics Carrier-XY-Focus-Tilt-Unit, article #10448587; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, SwitzerlandThe intended use of the microscope is for improving the visibility of objects through magnification and illumination. It can be applied for observation and documentation and for human and veterinary medical treatment. | Leica M822 Surgical Microscope | 2 | February 16, 2012 | 2012 | Leica Microsystems, Inc. | Leica Microsystems received complaints stating that the zoom function of the Leica M822 surgical microscope locked up during use. | DESIGN: Device Design | The firm, Leica Microsystems, Inc., sent a "Medical Device Correction" notice dated December 21, 2011 to all affected customers by e-mail. The product, problem, and actions to be taken by the customers were addressed. Customers were instructed- DO NOT USE use the affected product until the software could be upgraded and complete and return the Acknowledgement Form via fax to: 1-847-607-3147 or Scan and send to elizabeth.culotta@leica-microsystems.com. Leica Microsystems representatives will contact the customers to make necessary arrangements to complete the software upgrade.If you have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 29 | 29 units (10 units-USA) | Worldwide distribution: USA (nationwide) including states of: California, Colorado, Illinois, New York, Oregon, South Carolina; and countries of: Australia, Japan, Korea, Pakistan, Romania, South Africa, Switzerland and Turkey. | Terminated | February 16, 2012 | 0 | 29 | Computer
Z-1099-2012 | 61180 | camera, ophthalmic, ac-powered | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Retcam3 & Retcam Shuttle with version 6.0 software.Clarity Medical Systems, Inc.5775 W. Las Positas Blvd, Suite 200Pleasanton, CA 94588General ophthalmic imaging including retinal, corneal and external imaging. Photodocumentation of pediatric ocular diseases including retinopathy or prematurity (ROP). Screening of Type-2 re-threshold retinopathy of prematurity or treatment requiring ROP. | Retcam3 & Retcam Shuttle with version | 2 | February 24, 2012 | 2012 | Clarity Medical Systems Inc | Still images extracted from video recordings may be mislabeled as to right and left eye. | DESIGN: Software Design | Clarity Medical Systems, Inc. sent a Safety Alert letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following :1. Do not use the video mode until the software has been updated.2. Install the software update using the enclosed CD. The instructions for installation are also enclosed. This software update will correct the anomaly.3. If you have extracted still images from video recordings, review all potentiallyaffected patient files to determine if any right eye images are mislabeled as left eye images and/or left eye images are mislabeled as right eye images. If any images are mislabeled, make the appropriate corrections in the annotation section of the affected file. Consider whether there is any diagnostic impact or need to re-irnage the patient.We apologize for the inconvenience. If you have questions or concerns, please contact Clarity at 800-215-6005. |  | 33 | 33 RetCam Shuttles and 20 RetCam 3 - total 53. | Worldwide Distribution -- USA (nationwide) including the states of TN, MO, TX, FL, NJ and Puerto Rico and countries of Italy, Portugal, Czech Republic, Thailand, Chile, Oman, Hong Kong, Russia, Philippines, Sweden, Finland, Ecuador, Australia and Canada. | Terminated | June 07, 2012 | 104 | 33 | Computer
Z-1015-2012 | 60995 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | B. Braun Infusomat Space Pump - US Version Infusion pump. Model number 8713050U. | B. Braun Infusomat Space Pump US Version | 1 | February 29, 2012 | 2012 | B. Braun Medical, Inc. | If the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed, the clip catch on the inside of the pump door may break. If the clip catch is broken and the door opened, free flow protection is ensured. If IV set is then removed from pump without closing the roller clamp, the clip may not re-engage the tubing, leaving an open fluid path with free flow pote | DESIGN: Device Design | B Braun sent a letter on August 16, 2011, to their customers, that detailed the proposed G03 implementation in response to reported customer complaints. The software upgrade with the door jam was agreed upon and the eight hospitals were upgraded in November 2011. |  | 5769 | 5769 | Product was distributed to hospitals in CA and MO. | Open |  |  | 5769 | Computer
Z-1277-2012 | 61175 | calibrator, multi-analyte mixture | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products --- ALT Slides --- [REF] 165 5281 --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA. The recalled device is the calibration disk contained within the VITROS Calibrator Kit 3. The is can be identified with the label "VITROS Chemist y Systems --- Universal Calibration Diskette --- [REF] 871 6607 --- Ortho-Clinical Diagnostics"For in vitro diagnostic use only. VITROS Chemistry Products Calibrator Kit 3 is used to calibrate VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System for the quantitative measurement of AcP, ALT, ALKP, AMYL, AST, CK, GGT, LDH, and LIPA. | VITROS | 2 | March 22, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received complaints of lower than expected results ( | DESIGN: Software Design | Ortho Clinical Diagnostic sent a Important Product Correction letter dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers to: Install Calibration Diskette/Assay Data Diskette DRV 5667 on your VITROS System. Following a successful recalibration, you may resume using VITROSÂŽ ALT Slides, GEN 26 remaining in your inventory. To please complete and return the Confirmation of Receipt form upon receipt of this notification. Please return this form no later than February 17, 2012. Post this notification by each VITROSÂŽ System in your facility that utilizes the VITROSÂŽ ALT Slides, or with the VITROSÂŽ user documentation. Forward this notification if you have distributed this product outside of your facility.We apologize for the inconvenience this may have caused your facility. If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 4306 | 4306 | Worldwide Distribution - USA (nationwide) | Open |  |  | 4306 | Computer
Z-1375-2012 | 61118 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Assay Protocol File (APF) Software versions for UniCel DxC 600i Systems.Part Number: A54378, UniCel DxC 600i APF, Part Number: A54379, UniCel DxC 600i AAF.The Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentration found in human body fluids. | Access Immunoassay Systems | 2 | April 06, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed a customer report that a timing conflict can occur while aspirating samples for the Access Total BhCG assay. | DESIGN: Software Design | Beckman Coulter sent a "PRODUCT CORRECTIVE ACTION" letter dated September 24, 2008 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were instructed to complete and return the enclosed response form. Customers with questions regarding this notification were instructed to contact Technical Support at (800) 854-3633 or their local Beckman Coulter representative. |  | 2948 | 2948 units total (2762 units in the US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | June 26, 2012 | 81 | 2948 | Computer
Z-1448-2012 | 60865 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Software, User Manuals, and Handi-Guides for acute and chronic peritoneal dialysis. Software versions earlier than 2.8.The Fresenius Liberty Cycler is intended for acute or chronic peritoneal dialysis. | Fresenius Liberty Cycler | 2 | April 24, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Increased risk of intraperitoneal volume (IIPV), referred to as Overfill | DESIGN: Software Design | Fresenius Medical Care sent a Urgent Product Information letter dated March 30, 2012, and April 3, 2012, to all affected customers. The letter was sent via Certified Mail, Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.Recall to occur in 2 phases:Phase I: Fresenius Medical Care notified all clinics with patients using Liberty Cyclers with software versions earlier than 2.8 that these patients are receiving updated User Manuals and Handi-Guides. These new materials will include updated descriptions of the safety warnings included with the materials provided with the release of software version 2.8. Clinics will also be informed that their patients will be receiving replacement cyclers with updated software within 12 months. These notifications via telephone and certified letter (return receipt requested). Following clinic notification, Fresenius Medical Care will contact patients operating Liberty Cyclers with software versions earlier than 2.8 by phone, informing them that they will be receiving updated User Manuals and Handi-Guides. In addition, these patients will be informed that they will be receiving replacement cyclers with updated software at some point in the future and that Fresenius Medical Care anticipates that the full timeline of cycler replacements will occur over several months. A certified letter return receipt requested will be sent to all patients after the initial phone contact containing the updated User Manuals and Handi-Guides. Phase II: Liberty Cyclers with updated software will be issued to customers using Liberty Cyclers with software versions earlier than 2.8 as inventory is available.Please complete and return the enclosed Reply Form, indicating receipt and understanding of this communicationFurther questions please call (1-800) 227-2572. |  | 10146 | 10,146 units | Worldwide Distribution -- USA (nationwide) | Open |  |  | 10146 | Computer
Z-1473-2012 | 61535 | processor, tissue, automated | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | ASP 6025 Vacuum Tissue Processor; an automated modular tissue processor; Leica Biosystems Nussloch GmbH, Heidelberger Str. 17-19, D-69226 Nussloch, Germany; The intended use of this device is as an automated system used to process tissue specimens for examination through fixation, dehydration, and infiltration. | ASP 6025 Vacuum Tissue Processor | 2 | April 27, 2012 | 2012 | Leica Microsystems, Inc. | There is a potential malfunction with the device in autorotation mode, which can lead to incorrect reagent handling of the device. The end result may lead to damage or to loss of tissue specimens. | DESIGN: Software Design | Leica Microsystems sent a Urgent Medical device Recall letter dated March 30, 2012, to all affected customers in the U.S. on the same date. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed of the potential malfunction of the device in the autorotation mode, which could lead to damage or loss of tissue specimens. The letter provided instructions on what the customer should do with the recalled product in the interim before the software fix can be installed. They were instructed to follow the steps below: 1. Remove and dispose all reagents including Paraffin and clean all bottles and Paraffin bathes.2. Conduct a "Smart Clean" (IFU, chapter 6.1.4).3. In Submenu "System Setup", here at "Program options", respectively "Station sequence", customer shall switch from "Concentration" to "Auto by RMS".4. Prestored Xylene protocols shall be used instead of autorotation protocols (IFU, chapter 5.6.2 "Time - optimized Xylene infiltration programs"). Here you find the five validated programs which need to be copied to Favorite programs (IFU, chapter 5.5.3).5. Use the Reagents, bottle set up from "Time optimized Xylene programs" (IFU, chapter 5.6.2., page 134) and implement them into submenu "Reagents" (IFU, chapter 5.2. following pages).6. In submenu "The RMS system adjusting warning thresholds rules" (IFU, chapter 5.2.2) the validated reagent thresholds need to be implemented into Submenu "Reagents" (IFU, chapter 5.2).7. Current Reagent management values are visible within Submenu "Reagent Status" (IFU, chapter 5.2.3)The customers were also informed that their Leica Microsystems representative will them to make the necessary arrangements to complete a software upgrade at no charge. The customers were also requested to sign and date the enclosed acknowledgement form and fax it back to Leica at 847-236-3747 or e-mail it to LMGRA@leica-microsystems.com within 10 working da. |  | 67 | 27 units - U.S., 40 units - ROW | Worldwide Distribution -- USA (nationwide)including states of: California, Florida, Georgia, Kentucky, Massachusetts, Minnesota, Mississippi, Nebraska, New York, North Carolina, the District of Columbia; and countries of: Australia, Belgium, France, Germany, China/Hong Kong, Italy, Japan, the Netherlands, Singapore, Switzerland, Turkey and United Kingdom. | Terminated | October 10, 2012 | 166 | 67 | Computer
Z-1867-2012 | 61768 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS(R) Behring Coagulation System, Device catalog No. OVIO03Product Usage:Multipurpose system for In Vitro coagulation studies | BCS(R) Behring Coagulation System | 2 | June 26, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed that if the Cleaner SCS vial becomes empty after processing a test for Antithrombin assay using the INNOVANCE(R) Antithrombin kit, and the test immediately following is an APTT based clotting assay several results may be shortened. | DESIGN: Device Design | An Urgent Field Safety Notice dated December 2010 was issued to all BCS(R) System owners. Firm conducted a voluntary field corrective action to notify BCS(R) System owners and to install a software "patch" to resolve the issue. Until the software patch was installed on the BCS(R) System, customers were instructed in the Safety Notice that "the potential issue of carryover can be avoided if the Cleaner SCS vial is kept filled and not allowrd to run empty.". |  | 186 | 186 | Worldwide Distribution - US (nationwide) including the states of: AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IO, KS, KY, LA, MD, MA, MI, MS, MN, MT, NE, NH, NJ, NY, NC, ND, OH, OK, OR, PA, TN, TX, UT, VA, WA, WV, and WI and the countries of: Australia, France, Canary Islands, Spain, Belgium, Germany, Switzerland, Netherlands, Italy, Guadeloupe, Ceuta, Austria, Slovenia, Croatia, Saudi Arabia, Turkey, Greece, and San Marino. | Open |  |  | 186 | Computer
Z-2053-2012 | 62114 | microscope, surgical, general & plastic surgery | FSO | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Leica M525 F50 Surgical Microscope; Article No. 10338422; Leica Microsystems (Schweiz) Ag, 201 Max Schmidheiny Strasse, Heerbrugg, Switzerland Product Usage: The intended use of this device is for improving visibility of objects through magnification and illumination. It can be applied for observation and documentation and for human and veterinary medical treatment. | Leica M525 F50 surgical microscope | 2 | July 20, 2012 | 2012 | Leica Microsystems, Inc. | Leica Microsystems evaluated data showing that a malfunction resulting in an illumination reduction can be observed when Leica M525 F50 surgical microscope is used in combination with a remote control connected to a remote video adapter or motorized dual video adapter. | DESIGN: Software Design | Leica Microsystems sent an Urgent Medical Device letter dated June 4, 2012 to the affected customers. The letter described the product, problem and actions to be taken. The customers were informed that the Leica surgical microscope M525 F50 required a software upgrade. Customers were instructed to NOT USE a remote video adapter or motorized video adapter until the required software upgrade has been installed by their Leica Microsystems representative. The customers were requested to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the serial numbers of the units at their facilities and to fax it back to Leica Microsystems at (847) 236-3747. |  | 43 | 39 Units - USA , 4 Units - Canada | Worldwide Distribution - US (nationwide) in the states of: California, Colorado, Florida, Georgia, Illinois, Indiana, Massachusetts, Michigan, Minnesota, New York, Oklahoma, Oregon, Pennsylvania, Texas, Utah, West Virginia, including Puerto Rico, and the countriy of Canada. | Terminated | October 29, 2012 | 101 | 43 | Computer
Z-2216-2012 | 62344 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART & Sunquest Laboratory SMART Select | Sunquest Laboratory SMART & Sunquest Laboratory SMART Select | 3 | August 16, 2012 | 2012 | Sunquest Information Systems, Inc. | In Online Entry (OEx) there are two scenarios where a Test result from one container may file to a Test on another container. In order for this to occur:1) Site parameter SIT 1,2,1,26 (3,217) [Restrict resulting of tests to owner container ( | DESIGN: Software Design | Sunquest sent an Urgent Product Safety Notice dated April 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. For the 103 sites that still have this issue Sunquest distribute a new maintenance release. Customers were instructed to use LabAccess Results Workstation (LARS) instead of Online Entry when entering results for instrument method codes that are defined to use replicate tests that are typically ordered on the same accession. Customers were also instructed to share this notification with appropriate laboratory personnel. Customers were questions should call 520-570-2000. For questions regarding this recall call 520-570-2252. | Contact the recalling firm for information | 103 | 103 sites | Nationwide Distribution | Terminated | September 04, 2012 | 19 | 103 | Computer
Z-2270-2012 | 62708 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory Versions 6.4 and later using Microbiology Result Entry | Sunquest Laboratory Versions 6.4 | 2 | August 24, 2012 | 2012 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory versions 6.4 and later while using Microbiology Result Entry (MRE), the specimen ribbon containing the Accession Number (AN) and the Battery Name may be missing or contain the wrong information. | DESIGN: Software Design | Sunquest sent a recall letter dated December 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your reply to this email message serves as verification you have received, read, understood and/or followed instructions outlined in the attached Product Safety Notification. Your reply assists Sunquest Information Systems to satisfy our obligation to Regulatory Agencies.Please note in your email reply if your site will not be requesting the software correction. If your site plans to request the software correction, please follow the directions noted on the attached PSN.Further questions please call (520) 570-2347. | Contact the recalling firm for information | 127 | 127 sites | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | August 27, 2012 | 3 | 127 | Computer
Z-2272-2012 | 62714 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory LabAccess Results Workstation (LARS) | Sunquest Laboratory LabAccess Results Workstation | 2 | August 24, 2012 | 2012 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory LabAccess Results Workstation (LARS) may file results to the wrong patient by incorrectly combining results. | DESIGN: Software Design | An "Urgent - Product safety notice " was issued on 11/4/10 to all customers who purchased the Sunquest Laboratory LabAccess Results Workstation (LARS). The safety notice informed the customers of the problems identified and the actions to be taken. US customers may contact 1-877-239-6337. | Contact the recalling firm for information | 151 | 151 sites | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Denmark, Ireland, the United Arab Emirates and the United Kingdom. | Terminated | August 24, 2012 | 0 | 151 | Computer
Z-2304-2012 | 60514 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX (with FLEX software) and ABL80 FLEX CO-OX (with CO-OX software) - pH Blood Gas and Electrolytes Analyzer System.Intended Use The ABL80 FLEX (with FLEX software) and ABL80 FLEX CO-OX (with CO-OX software) are portable, automated analyzers that measure pH, blood gases, electrolytes, glucose and oximetry (ABL80 FLEX CO-OX only) in whole blood. The ABL80 FLEX (with FLEX software) and ABL80 FLEX CO-OX (with CO-OX software) analyzer systems are intended for use by trained technologists, nurses, physicians and therapists. They are intended for use in a laboratory environment, near patient or point-of-care setting. | pH Blood Gas and Electrolytes Analyzer System | 2 | August 30, 2012 | 2012 | Sendx Medical Inc | The firm recalled because the glucose measurements from a patient sample that the customer felt was too low and did not reflect the patient's clinical condition. | DESIGN: Software Design | Radiometer sent a Field Safety Notice letter dated June 06, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. Also included is a Note to Users to be placed at the front of customers existing Operator's Manual. Customers were instructed to provide confirmation to Radiometer that all customers have received the letter and the Note to Users. For questions contact your Radiometer representative. |  | 3599 | ABL80 FLEX is 1680, ABL80 CO-OX is 1759 | Worldwide Distribution - US Nationwide including the state of OH. | Terminated | September 04, 2012 | 5 | 3599 | Computer
Z-2072-2012 | 62621 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS X-ray Medical Diagnostic , X-ray Equipment SystemDigital Radiography used in general radiographic examination. | EOS System | 2 | September 28, 2012 | 2012 | Eos Imaging Inc | It was discovered of several generator failures of the EOS System due to failure of its anode controller board that drives the rotation of the X-ray tube anode. |  | EOS Imaging notified users by letter on 6/19/12 and that an on site correction would be performed by an EOS Imaging Field Service engineer.EOS indicated in its notification that although this specific defect has not been observed in EOS systems installed in the U.S, they will replace the anode driver boards with new boards equipped with high current capacitors. Similarly, the generator firmware will be upgraded with the software patch that improves the cooling fans' controls. This will be replaced at no charge to the customers, during regularly scheduled preventative maintenance of the user's EOS system. This should be completed by the end of calendar year 2012. Your proposed CAP appears to adequately address the issues you raised in your notification, and is hereby approved.Your proposed customer notification letter as revised, appears to adequately provide instructions with respect to the use of the product pending the correction of the defect and it appears to provide a clear evaluation in nontechnical terms of the potential hazards related to the defect. Your proposed notification letter is hereby approved for distribution to all affected parties. Further questions please call 678-564-5400. |  | 12 | 12 EOS Systems installed in the US | Nationwide distribution | Terminated | September 16, 2013 | 353 | 12 | Computer
Z-0473-2013 | 63577 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8000 Modular Analyzer Series, Model/Catalog/Part Number: 05641489001. The Cobas 8000 is a fully automated system for clinical chemistry analysis. | Cobas 8000 Modular Analyzer Series | 2 | December 04, 2012 | 2012 | Roche Diagnostics Operations, Inc. | On cobas c 701 and cobas c 702 modules of the Cobas 8000 series analyzer, a mismatch between the reagent probe positioning and the aspirated reagent volumes under certain conditions can lead to insufficient reagent pipetting. This may affect patient sample and controls recovery for the following tests: Alkaline Phosphatase (ALP2), Alanine Aminotransferase (ALT), Bicarbonate (CO2-L), Creatinine (C | DESIGN: Software Design | On 11/01/2012 an "Urgent Medical Device Correction" notification was sent via UPS Ground (receipt required) to consignees. The device issue was described and recommended actions were provided. A Roche Field Application Specialist will install the new e-barcode during a scheduled visit. Customers also will have an option to install the new e-barcodes themselves, following the instructions provided in the notification. A Software update will be released when available. Customers may contact Technical Support at 1-800-428-2336 for more information. |  | 78 | 78 analyzers | Nationwide Distribution, including the states of AL, AZ, CA, FL, GA, IL, IN, KS, KY, MA, MD, MI, MN, NC, NJ, NY, OH, VA, WA, and WI. | Terminated | January 08, 2014 | 400 | 78 | Computer
Z-0599-2012 | 60367 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH500 Series Analyzer,Part number: 178832, 178833, 178834.LH 500 Series System Instructions for Use, PN 624602 Rev BA. The COULTER LH 500 Analyzer is a quantitative, automated hematology analyzer and leukocyte differential cell counter For In Vitro Diagnostic Use in clinical laboratories. The LH500 Analyzer also provides a semi-automated reticulocyte analysis. | Coulter LH500 Series Analyzer | 2 | January 11, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman has identified a problem with Integrated Circuit chips used in different boards throughout the analyzer. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to contact Beckman Coulter Service at (800) 526-7694 if they were experiencing failures. Customers not experiencing failures, but their system may have the suspect chips, were told they would be contacted by Beckman Coulter Service to correct their problem. Customers were also instructed to complete and return an enclosed Response Form. | Contact the recalling firm for information | 117 | 117 units total | Worldwide Distribution. | Terminated | June 01, 2012 | 142 | 117 | Computer
Z-0876-2012 | 60904 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015.The Alaris PC unit is the central programming, monitoring and power supply component for the Alaris System. | Alaris PC unit | 2 | January 26, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion has identified a potential risk with a PLD component on the Alaris PC unit model 8015 power supply board. The PLD component used in production between June 21, 2011 and August 15, 2011 has a higher standby current resulting in the battery depleting prematurely. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carefusion sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated September 13, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A Customer Reply Card was enclosed for customers to complete and return to the firm. Customers were instructed to call Carefusion Support Center at (888) 562-6018, 7 am-5 pm (Pacific) for any recall related questions. | No consumer action necessary | 1039 | 1039 units | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CT, FL, IA, IL, MI, MO, NJ, NY, OH, SC, SD, and TX. and the country of Australia. | Terminated | January 31, 2012 | 5 | 1039 | Computer
Z-1041-2012 | 59827 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V24-00400-29 cable, an accessory to the Newport HT50 ventilator.The V24-00400-29 cable is an accessory to the Newport HT50 ventilator intended to transfer the alarm signal from the ventilator to a nurse remote station. | V240040029 cable | 1 | February 21, 2012 | 2012 | Flight Medical Ltd. | The recall was initiated because Flight Medical Innovations Ltd. has identified a problem may occur when contact between an exposed unused wire and the grounding wire occurs. If this contact occurs the ventilator may reset and switch off. | DESIGN: Device Design | Newport Medical Instruments, Inc., the U.S. distributor for Flight Medical Ltd., sent a letter dated August 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use of the affected product. An Acknowledgement Form was attached to the letter for customers to complete and return via fax at (714) 427-0489 or e-mail to Regulatory@NewportMedical.com. Replacement product will be provided once the form is received. Contact Newport Medical's Technical Service Group at 800-451-3111 ext. 500 for questions regarding this notice.Updated: The firm posted a press release on 02/12/12 Isreal time, notifying customers of update regarding its voluntary recall of remote alarm cable, part number V24-00400-29, used to connect a Newport HT50 mechanical ventilator to a nurse station. The press release notified: The recalled cable may negatively affect Newport HT50 ventilators manufactured prior to 2007. Affected units may shutdown unexpectedly. The FDA has classified this action as a Class I recall.On August 23, 2011, Flight Medical sent an urgent Recall Notice to affected customers through its distributor, Newport Medical Instruments, Inc. (Newport Medical), to stop using the alarm cable immediately and to contact Newport Medical in order to receive an upgraded version of the cable. In the notification letter, the customers were notified of the affected cables and that an upgraded version of the cable would be provided to them free of charge. Customers, including Newport Medical are requested to dispose of the affected cables properly.Instructions to customers: Customer inquiries related to this action should be addressed to Newports Technical Service Group at (800) 4513111, extension 500. | Contact the recalling firm for information | 363 | 363 total units, all in the US | Nationwide Distribution. | Terminated | May 21, 2012 | 90 | 363 | Computer
Z-0487-2012 | 56540 | full-montage standard electroencephalograph | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software in use with:Harmonie-E Long Term Monitoring System (Models HSYS-REC-WD, HSYS-REC-WT; HSYS-REC-LT2; HSYS-REC-EAMP)Harmonie-S Sleep System (Model HSYS-REC-DUO)VITA/ICU Neurological Monitoring System (Model HSYS-REC-DUO-PPC)Harmonie{-S}{-E} Readers (Models HSYS-RDR-D, HSYS-RDR-T).Acquire, display, store, and archive electroencephalographic signals from the brain using a full montage array (i.e., 16 or more electrodes) and user-specified locations. | Harmonie Software | 2 | January 11, 2012 | 2012 | Stellate Systems | The following implied or stated indications in Stellate labeling with Harmoniae software have not yet been cleared by the Food and Drug Administration (FDA) for sale in the United States of America. | PRODUCTION CONTROLS: Error in Labeling | Stellate Systems sent a "DEVICE CORRECTION" letter dated April 6, 2009 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The firm provided information on how to verify if customers had the affected product. Software updates will be available to all customers in possession of the affected product. Customers were instructed to call the Customer Support line at 1-888-742-1306 with any questions or concerns they may have regarding the notification. |  | 1041 | 1,041 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Denmark, Greece, Hong Kong, Hungary, Italy, Japan, Mexico, Norway, Rep. DE Panama, Spain, Switzerland, Taiwan ROC, and Turkey. | Terminated | December 27, 2011 | -15 | 1041 | Not_Computer
Z-0523-2012 | 60423 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology |  | MAMMOMAT InspirationFull field digital mammographic x-ray system | Siemens MAMMOMAT Inspiration | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Lowering the swivel arm with "object table" while an operator's or patient's body parts are underneath the "object table" may cause the "object table" to collide with and injure the operator or patient. Therefore attention should be paid when adjusting "object table" height to avoid a possible risk of injury to the operator or the patient.The movement of the swivel arm may also cause an injur | DESIGN: Device Design | Siemens Medical Solutions USA, Inc sent a Safety Advisory Notice dated January 8, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The notification included an Addendum to the User Manual.We appreciate your understanding and cooperation with this safety advisory and ask you to immediately instruct your operating personnel accordingly. Please ensure that this safety advisory is placed in the System's Operator Manual. Your operating personnel should maintain awareness over an appropriate defined period. Further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of FL, MD, MI, NE, NY, NC, ND, OH, and SC. | Open |  |  | 12 | Not_Computer
Z-0389-2012 | 60344 | [6 empty] | Magnus Hybrid operating table columns 1180.01A1 and 1180.01B1 | Magnus Hybrid operating table columns | 2 | January 11, 2012 | 2012 | Maquet Inc. | A component of the Magnus hybrid OR system fails to respond to a position command. | DESIGN: Device Design | Maquet sent a PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION letter dated September 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. To avoid the potential of a malfunction occurring prior to their scheduled Service procedure, customers were instructed to ensure that the MAGNUS table-top is positioned at the "0" pivot position and that the MAGNUS column pivot brake is secured. If it is necessary to pivot the MAGNUS table-top prior to procedure start, ensure that the column pivot brake is completely released and then completely secured when returned to "0" pivot position. Customers were instructed to ensure that all user of the affected product and other necessary persons are informed of the product correction notice. For technical support or questions call 1-888-627-8383. | Contact the recalling firm for information | 45 | 45 UNITS | Worldwide Distribution - USA including VA, LA, MO, NC, MD, and AZ and the countries of AE, AU, CA, CH, CN, DE, DK, ES, IE, IL, SP, PL, RU, and TH. | Terminated | December 08, 2011 | -34 | 45 | Not_Computer
Z-0637-2012 | 60756 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | ThermoFisher Scientific / Microgenics brand MAS CardioImmune XL, Liquid Assayed Cardiac Marker Control, in vitro diagnostic. Catalog Numbers & Description:CAI-XLL, Level L; CAI-XL1, Level 1; CAI-XL2, Level 2; CAI-XL3, Level 3;CAI-XL4, Multi-Pack; CAI-XLS, Sample Pack; | ThermoFisher/Microgenics | 2 | January 11, 2012 | 2012 | Microgenics Corporation | Low measurements of Troponin I in the MAS CardioImmune XL Control which were outside the published control ranges. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 12/14/2011, ThermoFisher Scientific issued letters via express mail to all its direct consignees, informing them of the affected products and providing instructions on the recall. |  | 25113 | 25,113 Kits | Nationwide and Worldwide, including Germany, France, Italy and Spain. | Terminated | October 26, 2012 | 289 | 25113 | Not_Computer
Z-0506-2012 | 60581 | thermometer, electronic, clinical | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Walgreens Quick Read Digital Thermometer; an electronic digital thermometer; Model WIC551226 | Walgreens Quick Read Digital Thermometer | 2 | January 11, 2012 | 2012 | Walgreen Company | The thermometers have been reported to have inaccurate readings | DESIGN: Labeling Design | On November 23, 2011, Walgreens issued a recall notice electronically to their stores, instructing them to pull and quarantine the Thermometers and placing a Stop Scan on the products in the computer cash register system on the same date. Walgreens also removed the thermometer from their internet sales site. | Discard product or return to place of purchase for a full refund or exchange | 92000 | estimated 92,000 units | Nationwide through Walgreens retail stores and their internet site. | Open |  |  | 92000 | Not_Computer
Z-0603-2012 | 60390 | radioimmunoassay, thyroid-stimulating hormone | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The GSP 2021-0010 (Genetic Screening Processor) instrument is a fully automated, high through put batch analyzer for time- resolved and prompt fluorescence analysis of samples in microtitration plates. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. The chemistries running on the GSP instrument are used to test blood specimens dried on filter paper as an aid in screening newborns for congenital disorders. | Genetic Screening Processor | 2 | January 11, 2012 | 2012 | PerkinElmer LAS, Inc | Malfunction of the GSP disk remover unit may cause a punched disc to remain in the well during the measurement of the plate. This may cause a false negative screening result being reported for 17-OHP, resulting in failure to identify cases of congenital adrenal hyperplasia (CAH). | DESIGN: Device Design | Recall (correction) letters were sent to all customers on October 24, 2011. An improved disk remover design is under development to eliminate the reported problem. Until the final field correction is available, customers should follow the precautions provided. The needed protocol update will be performed by the local service engineer. If customers observe an increase in the occurrence of the problem, the laboratory should contact their local PerkinElmer representative. |  | 52 | 52 | Worldwide Distribution - USA, including the states of: Missouri, Florida, Texas, Oregon, Alabama, Louisiana, New Jersey, Georgia, Virginia, Kentucky and the countries of: Australia, Belgium, Brazil, Egypt, France, Germany, Italy, Netherlands, Norway, Qatar, Saudi MOH, Switzerland, and the UK. | Terminated | July 20, 2012 | 191 | 52 | Not_Computer
Z-0382-2012 | 60437 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System: Lung Optimized Treatment option.The CyberKnife treatment is indicated for the treatment planning and image guided stereotactic radiosurgery and precision radiotherapy of lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | CyberKnife Robotic Radiosurgery System: Lung Optimized Treatment option. | 2 | January 11, 2012 | 2012 | Accuray Incorporated | User Facility reported an anomaly where inhale and exhale CT pairs used for treatment planning did not represent the same magnitude of respiration that was displayed during treatment delivery. | DESIGN: Labeling Design | Accuray sent a Urgent Medical Device Correction letter dated November 7, 2011, via Federal Express to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The information provided in this letter will be incorporated in user manuals shipped with the CyberKnife system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at 1. (877) 668-8667 ( USA) 1. (408) 716-4700 ( non USA ). |  | 30 | 30 | Worldwide Distribution -- USA ( nationwide ) including the states of FL, TN, IL, NJ, CO, PA, MD, CA, MS, MN, SC, WA, OK, GA, AZ, and TX., and the countries of France and United Kingdom. | Terminated | February 27, 2012 | 47 | 30 | Not_Computer
Z-0585-2012 | 60132 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery XR650, Digital Radiographic Systems. Revolution XR/d Digital Radiographic Imaging System with Image Pasting and Autopositioning is indicated for use in generating radiographic images of human anatomy. This device is not intended for mammographic applications. | Discovery XR650 | 2 | January 11, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the patient barrier rotating arm that is a part of the Definium 8000 and Discovery XR650 image pasting systems that may impact patient and user safety. The patient barrier rotating arms that are used to position and stabilize the patient during image pasting procedures may fail to retain residual friction at the arm pivot. Thi | DESIGN: Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 17, 2011 to its customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions; discontinue use of the patient barrier and remove from service (away from patients) if either one the rotating arm locking levers is damaged (does not hold the rotating arm in place when engaged); take extra precautions to hold the patients barrier rotating arm during release of the locking lever and repositioning of the arm and store the rotating arms in the lowest pivot position (down) when not in use. The letter also stated that the patient barrier may continue to be used if the rotating arm locking lever securely holds the rotating arm in place when it is engaged.Additionally, GE Healthcare will modify the rotating arms to restore the residual friction so that the arm will retain position (rather than swing freely) while the locking levers are disengaged. Also, GE will update the product labeling to improve awareness and identify the hazard relative to proper storage.If you have any questions, call Center in US: 800-437-1171; Japan: 0120-055-919 and for other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 229 | 229 | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID,IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, NORWAY, NEW ZEALAND, NETHERLANDS, MARTINIQUE, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, HONG KONG, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE, FINLAND, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BOSNIA AND Herzegovina, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | October 16, 2012 | 279 | 860 | Not_Computer
Z-0717-2012 | 60767 | circuit, breathing (w connector, adaptor, y piece) | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Model 750 24' MRI Reusable/Disposable CircuitAccessory for use with Model 750 Portable Ventilator | Model 750 24' MRI Reusable/Disposable Circuit | 2 | January 12, 2012 | 2012 | Impact Instrumentation, Inc. | Model 750 24' MRI Reusable/Disposable Circuit does not have a 510(k). |  | Impact sent Urgent Device Product Recall letters/return response forms dated 10/19/11 to their consignees. The letter asked customers to return the affected product listed to the address provided within 10 days of receipt of the notification letter. In order to receive proper credit for the returned product, customers should follow the instructions precisely, which includes completing and returning the Recall Tracking Form along with the affected product. If customers have no product to return, they should contact Alan Giordano at agiordano@impactii.com, to inform him of the status of the circuits. |  | 564 | 564 units | Nationwide Distribution -- VA, CA, IL, NC, and TX. | Terminated | March 11, 2013 | 424 | 564 | Not_Computer
Z-0741-2012 | 60889 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer, Hydrogen Peroxide Gas Plasma Sterilizer, Chemical Sterilizer. STERRAD 200; Product Usage: The STERRAD Sterilization System is a low-temperature general-purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | STERRAD 200 Sterilizer | 2 | January 13, 2012 | 2012 | Advanced Sterilization Products | ASP determined that an internal filter in the vacuum pump on the STERRAD 200 System is subject to wear over time. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ASP will send an Urgent Medical Device Correction letter to all affected customers on January 23rd, 2012. The letter describes the product, problem and the actions to be taken. The letter informs customers that ASP has determined that the vacuum pump in the STERRAD 200 System may cause the emission of a slightly increased level of hydrogen peroxide (H2O2) vapor into the surrounding environment if not regularly serviced or replaced.The letter instructed customers to read the "Issue Description" and "Recommendations" sections in the letter, pass on the notice with any staff that work with or around the STERRAD 200, and maintain awareness of this communication until ASP has serviced the system. Also, a list of FAQ was provided. The letter states that ASP will replace, free of charge, the vacuum pumps on all STERRAD 200 Systems that have exceeded the pump manufacturer's recommended service interval. ASP has contracted with Stericycle to manage this field correction. Telephone numbers are provided in the letter: for questions related to this field correction, please contact Stericycle at 1-877-225-9750 and to report any complaints or suspected problems with the System, please contact ASP Professional Services directly at 1-888-783-7723, option # 2. | Contact the recalling firm for information | 395 | 395 units | Worldwide Distribution - USA (nationwide) | Terminated | September 23, 2013 | 619 | 395 | Not_Computer
Z-2098-2012 | 60853 | toothbrush, powered | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Arm and Hammer Spinbrush Pro Clean Sonic Recharge Made in China.Manufactured for Church and Dwight Co., Inc. Princeton, NJ 08543 USA.For dental health; plaque removal. | Arm and Hammer | 2 | January 25, 2012 | 2012 | Church & Dwight Inc | Church and Dwight plans to take actions to best convey to the public the potential small parts choking hazard with using a Spinbrush beyond the recommended 3-month brush head life. | DESIGN: Device Design | Church & Dwight Co., Inc. proposed a consumer education media campaign starting on December 22, 2011 with the following statement:"Safety Notice: Please remember to replace your brush head after 3 months of use, or if the brush is damaged or the parts become loose. Extended usage, loose parts or excessive wear could lead to brush head breakage, generation of small parts and possible choking hazard. Inspect the brush for loose parts before use."The campaign will be conducted in the following media and anticipate that the combined consumer reach of all of these vehicles to be approximately 144 million people an average of 4 times.- Television campaign will be run during the entire Q1 of 2012 (January -March). -Print ads that will appear in the December 30th issues of People, Star, US Weekly and Entertainment Weekly. -Church & Dwight Website will contain the Safety Alert notice starting December 22nd. (www.spinbrush.com)Consumer Relations IVR (Interactive Voice Response) to consumers calling the Spinbrush toll-free numbers). (800-353-3384 and 800-5610752).4-30-121) A print advertisement (Attachment A) containing the Sonic Safety Notice that will appear in each magazine (Star, People, US Weekly, and Entertainment Weekly) as noted:Five total print insertions are planned (on sale dates listed):I. 4/30 -Star2.5/14 -Us Weekly3.5/28 -People4.6/11 -Entertainment Weekly5.6/25 -Star2) A disk containing the 15 second commercial containing the safety notice.3) A document (Attachment B) of the National TV slots showing when the commercials will run in the quarter, and another with the shows during which they will air. TV ads will begin running the week of April 16th. The DRTV schedule is determined weekly and we are informed after the fact when the ads where shown. The DRTV stations are included at the bottom of the National TV listing." In total, combined print and TV are estimated to reach 163MM impressions (measurement target: women. |  | 0.347 | 0.347 million units | Worldwide Distribution. | Open |  |  | 38 | Not_Computer