Z-0914-2012 | 60893 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen(TM) Laboratory Automation Systems configured with Hettich Rotanta 46 RSC Robotic Centrifuges, Catalog # enGen -- Common/Usual Name: enGen(TM) Track System, Hettich Centrifuge -- Ortho-Clinical Diagnostics, Inc. Rochester, NY --- For in vitro diagnostic use only.The VITROS 5,1 FS Chemistry System with enGen(TM) Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric immunoassay methods). | enGen(TM) Laboratory Automation Systems | 2 | January 26, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received complaints of centrifuge buckets disconnecting from their supports during centrifugation when using the Hettich Rotanta 46 RSC Robotic Centrifuges, Models 4815, 4816 and 4817, that may be configured with enGen(TM) Laboratory Automation Systems. | DESIGN: Component Design/Selection | Ortho-Clinical Diagnostics sent an Urgent Product Correction Notification letter dated April 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to post the enclosed attachment near their centrifuge as a reference. Clean and carefully inspect the suspension groove area of the centrifuge buckets for cracks as part of their weekly maintenance procedure. Careful to check for micro-cracks that may not be easily detectable.If any cracks/micro-cracks, corrosion or damage are present, they must immediately discontinue using the centrifuge and contact OCD Customer Technical Services at 1-800-421-3311 to obtain replacements.OCD recommends that they do not exceed the manufacturer's relative centrifuge force (RCF) and run time limits for the blood collection devices used by their laboratory.Complete and return the attached Confirmation of Receipt Form.For questions regarding this recall call the Customer Service Center at 1-800-421-3311. . |  | 78 | 78 units worldwide (18 Domestic and 68 Foreign) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Chile, England, France, Germany, Italy, Mexico, Singapore and Spain. | Open |  |  | 78 | Not_Computer
Z-0918-2012 | 60707 | low energy direct current thermal ablation system | OAB | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | NanoKnife Irreversible Electroporation System. System consists of multiple components: Generator, footswitch, disposable electrode. System is manufactured by AngioDynamics, Inc. --- Catalog #20300101Intended use: the generator system with six outputs is indicated for surgical ablation of soft tissue. | NanoKnife | 2 | January 27, 2012 | 2012 | Angiodynamics Worldwide Headquarters | Angiodynamics was marketing the NanoKnife System with the ablation zone estimator feature in the US without FDA clearance. |  | AngioDynamics sent an Urgent - Medical Device Correction letter dated December 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the User's Manual that was supplied with the generator when it was purchased. Instructions include discarding their current manual and troubleshooting guide. A replacement manual will be provided with the Field Safety Notice; the troubleshooting guide will not be replaced. Customers were instructed to complete the Reply Form and fax to AngioDynamics at 518-798-1360. If customers further distributed the affected product, they should inform those accounts of the field correction. Angiodynamics will also disable the AZE feature during a schedule service visit.For questions regarding this recall call 1-800-772-6446 or e-mail customerservice@angiodynamics.com. |  | 40 | 40 units | Nationwide Distribution | Open |  |  | 40 | Not_Computer
Z-1044-2012 | 61019 | plasma, coagulation control | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | STA - System Control N + PProduct Usage:The STAÂŽ -System Control N+P kit provides a normal plasma and an abnormal plasma intended for use as two control levels for the following assays performed on analyzers of the STAÂŽbrand name suitable to these reagents: -the Reagent 1 (STAÂŽ-System Control N)is for the following tests: prothrombin time (PT), activated partial thromboplastin time (APTT), fibrinogen (Clauss method), thrombin time (TT), ReptilaseÂŽÂŽ time, factors II, V, VII, VIII, IX, X, XI and XII, antithrombin (AT), protein C, protein S, plasminogen and antiplasmin. -the Reagent 2 (STAÂŽÂŽ -System Control P) is for the tests as described above except thrombin time (TT) and ReptilaseÂŽÂŽ time. Do not use the STAÂŽÂŽ -System Control N+P kit on STA SatelliteÂŽ for the control of the APTT performed with the STAÂŽÂŽ-rTT kit. | STA System Control NP | 2 | February 16, 2012 | 2012 | Diagnostica Stago, Inc. | Factor VIII quality control results are being recovered below the assigned ranges for STA System Control N+P. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stago sent a recall letters via email dated August 23, 2011 to all affected customers . The letter identified the product, description of problem, and corrective actions to be taken. Customers were asked to immediately examine their inventory and follow the corrective actions provided. Customers were instructed to complete the response form attached to the letter. Because of incorrect manufacturer's instructions included in the initial letter another letter was emailed by the firm on August 30, 2011. For questions or technical assistance for updating the parameters please call the Diagnostica Stago 24 hour hotline at 1-800-725-0607. |  | 1259 | 1259 units | Worldwide Distribution and USA (Nationwide) and including Canada | Terminated | September 04, 2013 | 566 | 1259 | Not_Computer
Z-1046-2012 | 61102 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity CT (Computed Tomography) System, 510 (k) #K033326, Model #728326, Serial Numbers: 300003, 300004, 300005, and 300010. Product Usage: The Ingenuity CT is a Computed Tomography X-Ray System intended to produce cross-aÂŽsectional images of the body by computer reconstruction of X-Ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | InGenuity CT System | 2 | February 16, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Design defect. Philips learned of this device issue following their receipt of a customer complaint reporting that a 'gap' is created in the side covers when the gantry of the device is tilted to +30 degrees (forward tilt). A hazard to the operator/user presents itself if the operator puts his or her fingers or hand inside the gap and the gantry tilt is returned in position while the fingers or | DESIGN: Component Design/Selection | The firm, Philips, sent an 'URGENT - Medical Device Correction Ingenuity CT***' letter dated December 08, 2011. The letter describes the affected product, the nature of the problem along witth the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. The customers (operator/users) were instructed not to insert hands in the gap which will be located between the gantry and stationary support column located on either side of the CT System (illustrations are provided). The letter notifies the customers that a Philips representative will be visiting each customer site and installing a hardware upgrade through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, option 5; enter site 10 or follow prompts). | Contact the recalling firm for information | 8 | Four (4) device system units were distributed. Two (2) in U.S. and two (2) overseas | Worldwide distribution: USA (nationwide) including states of: IN and OR; and countries of: Australia and Germany. | Terminated | March 01, 2013 | 379 | 4 | Not_Computer
Z-1060-2012 | 60855 | prosthesis, wrist, hemi-, ulnar | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Ascension Orthopedics First Choice Partial Ulnar Head System ("DRUJ System"). Head sizes 14.5 mm, 16.0 mm, 17.5 mm, and 19.0 mm each with stem sizes 4.5 mm, 5.5 mm, or 6.5 mm.The First Choice DRUJ System is intended for replacement of the distal radio-ulnar joint. The Partial Head implant is indicated for replacement of the distal ulnar head for rheumatoid, degenerative, or post-traumatic arthritis | First Choice Partial Ulnar Head System ("DRUJ System"). | 2 | February 17, 2012 | 2012 | Ascension Orthopedics, Inc | A recent review conducted by Ascension Orthopedics revealed that certain modifications were made to the First Choice System and should have been preceded by a submission of a filing with the FDA. | DESIGN: Device Design | The Firm, Ascension Orthopedics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 3, 2011 via FedEx, email or fax to all customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately cease any further use of any affected product; remove from their stockrooms until it is shipped back to Ascension Orthopedics; promptly return any affected product to: Ascension Orthopedics, 8700 Cameron Road, Austin, TX 78754, Attn: First Choice DRUJ Recall Administrator, and complete and return the reply form even if you do not have affected product in stock, to Regulatory Affairs Dept. via fax to 512-836-6933.If you have any questions, please do not hesitate to contact customer service at 1-512-836-5001. | No consumer action necessary | 1303 | 1303 units released to customers | Worldwide distribution. | Terminated | January 14, 2013 | 332 | 1303 | Not_Computer
Z-1067-2012 | 60751 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE SIGNA EXCITE MR SURGICAL OPTION (MR450w Surgical Suite).The MR Surgical Option when integrated with the GE SignaÂŽ 3.0T or 1.5T magnetic resonance scanner (previously cleared K041476) can be used to produce head and whole body magnetic resonance images that are high resolution, high signal-to-noise ratio, with short scan times. The SignaÂŽ Excite MR system with Surgical Option is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal and oblique images, spectroscopic images, and/or spectra, dynamic images of the internal structures and organs of the entire body, including, but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. The images produced by the Signa" Excite system with Surgical Option reflect the spatial distribution of protons (hydrogen nuclei) exhibiting magnetic resonance. These images and/or spectra when interpreted by a trained physician yield information that may assist in diagnosis. The GE Signa" Excite system with Surgical Option can provide MR images at various stages of a surgical procedure, and provides the capability to transfer a patient to the MR system table directly from a surgical table, or patient transport system. Images can be obtained using the MR scanner body coil, or local surface coils, and the surface coils used with the integrated system may/can accommodate sterile draping for surgical procedures. | GE SIGNA EXCITE MR SURGICAL OPTION (MR450w Surgical Suite). | 2 | February 21, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the failure of the transfer board release lever associated with the OR Compatible patient transfer table of your Optima MR450w MR Surgical Suite. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare, sent a Urgent Medical Device Correction letter dated December 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed discontinue use of OR Compatible table until issue is corrected. GE Healthcare will correct all affected systems by providing a replacement table cradle incorporating design changes to the cradle release lever mechanism at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification. For further questions please call (262) 513-4122. | Contact the recalling firm for information | 2 | 2 | Worldwide Distribution -- USA including Wisconsin and the country of Austria. | Terminated | May 29, 2012 | 98 | 2 | Not_Computer
Z-1065-2012 | 60126 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | REVOLUTION XR/D DIGITAL RADIOGRAPHIC IMAGING SYSTEM WITH TOMOSYNTHESISThe Revolution XR/d with Tomosynthesis is intended for use in generating Tomographic images of human anatomy. It is not intended for mammographic use. | REVOLUTION XR/D DIGITAL RADIOGRAPHIC IMAGING SYSTEM WITH TOMOSYNTHESIS | 2 | February 21, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the hand grip that is clamped to the table top of your Revolution XRD X-ray systems that may impact patient safety. The hand grip accessory that is clamped to the table top may not be properly adjusted and can dislodge during use. Should this happen during patient loading or unloading, patient injury could occur. | DESIGN: Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to remove the hand grip accessory from use until the product is corrected and to take extra precautions when loading or unloading patients from the table while the hand grips are not in use. A GE Healthcare service representative would contact the consignee to arrange for the correction. For questions regarding this recall call 262-513-4122. | Contact the recalling firm for information | 772 | 772 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WY, WI and the countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SLOVENIA, SAUDI ARABIA, REPUBLIC OF KOREA, NORWAY, NEW ZEALAND, NETHERLANDS, MALTA, KUWAIT, JORDAN, JAPAN, ITALY, ISRAEL, HONG KONG, GERMANY, FRANCE, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, and AUSTRALIA. | Terminated | August 02, 2012 | 163 | 772 | Not_Computer
Z-1090-2012 | 61074 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian 52-leaf Standard MLC multileaf collimator;Reference/FSCA Identifier: CP-04792; Model Number(s): All models, includingH52, H54, H56, HHM; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA.The Multileaf Collimator is intended to assist the Radiation Oncologist in the delivery of radiation to well defined target volumes while sparing surrounding normal tissue and critical organs from excess radiation. | 52leaf Standard MLC multileaf collimator | 2 | February 23, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | When using the Standard and Millennium 52-leaf MLC multi-leaf collimators for beam shaping, under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceed the expected transmission dose. | DESIGN: Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides recommended actions when planning and treating dynamic fields on the 52 leaf MLCs. Contact the Help Desk at 1-888-827-4265 for questions regarding this recall. |  |  |  | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Canada, China, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Singapore, Bahrain, Egypt, Hong Kong, India, Indonesia, Israel, Malaysia, Morocco, Nepal, South Africa, South Korea, Taiwan, Turkey, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Colombia, Mexico, Panama, Venezuela, Saudi Arabia, and Australia. | Terminated | July 26, 2012 | 154 |  | Not_Computer
Z-1187-2012 | 61131 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra InterFix CT Adapter (CTA)Intended for cranial fixation during CT scanning andtreatment with the TomoTherapy HiArt (Torno) SystemThe Integra InterFix CT Adapter (CTA) kits are intended for cranial fixation during CT and and treatment with the Tomo Therapy Hi-Art system. Provides adapters for attaching: Integra head rings to the following: 1.Flat CT scanner tabletops ranging form 47cm to 54 cm 2. TomoTherapy Hi-Art system tabletop | Integra InterFix CT Adapter (CTA) | 2 | March 08, 2012 | 2012 | Integra LifeSciences Corp | Material used in the superior region of the InterFix CT Adapter (CTA) has a different density than that of the Treatment Table Couch Adapter (TTA). Resulting dosimetry can be affected if posterior fields are used. Integra estimates that a target dose decrease of 1.2% would result from a typical 360 degree TomoTherapy arc. | DESIGN: Device Design | Integra sent a Product Technical Bulletin dated January 23, 2012, to all affected consignees via FedEx, email or telephone conference calls. The revised Operator Manual was provided with a Technical Bulletin to all consignees that have been shipped this system. Consignees were asked to distribute the information in the bulletin to members of the Radiotherapy Planning Team within their institution. Customers were also asked to sign and return the enclosed Product Technical Bulletin Acknowledgement Form. For questions regarding this recall call 888-772-7378. | Contact the recalling firm for information | 33 | 33 units | Worldwide Distribution - USA including DC, IN, FL, HI, LA, NY, MD, MI, OH,OR, WI, WY,Foreign: Great Britain, Italy, and France | Terminated | October 01, 2012 | 207 | 33 | Not_Computer
Z-1199-2012 | 61179 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Nucletron Adapter and Transfer Guide Tube, part number AL 13005000, a component used with the VariSource" IX HDR Afterloader.Product Usage: The adapter is part of the VariSource Afterloader System, a computer controlled remote HDR Afterloader used to place a high activity radioactive source within a needle(s) or applicator(s) which have previously been placed for a specific clinical purpose in a patient. | Regulation Number 892.5700 | 2 | March 12, 2012 | 2012 | Varian Medical Systems, Inc. | Components of medical device used in brachytherapy may be defective and cause a shifted dose distribution to patients. | DESIGN: Device Design | Varian Medical Systems sent an Urgent Medical Device Correction - Urgent Field Safety Notice letter dated January 5, 2012 to all affected consignees. The letter identified the affected products, description of problem and recommended user actions to be taken. Consignees were requested to cease use of the affected products and contact Varian BrachyTherapy Helpdesk immediately to arrange return of their affected parts. Consignees were requested to return the enclosed proof of notification form. Consignees were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. For future reference, this document will be posted to the Varian Customer Support website: http://www.MyVarian.com. For further clarification contact your local Varian District or Regional Manager. |  | 122 | 122 total units | Worldwide Distribution - USA (nationwide) Including the states of CA, FL, IN, KY, MI, MN, MO, OH, TX, VA and the countries of Canada, Japan and Sweden. | Terminated | November 21, 2012 | 254 | 122 | Not_Computer
Z-1322-2012 | 61170 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BBL (tm) Lowenstein-Jensen Medium Deeps, Catalog #221256, packaged in boxes of ten/600 mL tubes. Product Usage: Lowenstein-Jensen Medium is used for the isolation and cultivation of mycobacteria. The medium tubed as deeps is used for the semi-quantitative catalase test as an aid to the classification of mycobacteria. | BD BBL (tm) LowensteinJensen Medium Deeps | 2 | March 27, 2012 | 2012 | Becton Dickinson & Co. | Media fails to perform as intended with quality control organism Mycobacterium kansasii. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated January 2012 to Ex-US BD sites via email and its customers via UPS ground shipment on 1/30/12. The letter identified the product, the problem and the action to be taken. The customers were instructed to discontinue use of the recalled lot numbers and discard any remaining packages. Customers were also instructed to complete and return the attached form whether or not they had any inventory remaining via fax to: BD Regulatory Compliance at 410-316-4285. The letter also informed the customers that BD has decided to discontinue this product.If further assistance is needed regarding credit, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 3900 | 3900 tubes | Worldwide Distribution - USA (nationwide) including states of: AL, CA, CT, IA, IL, KY, LA, MD, MI, MN, NC, NJ, NY, OH, PA, TN, VA, WA, and WV; and countries of : Belgium and Singapore. | Terminated | April 02, 2012 | 6 | 3900 | Not_Computer
Z-1321-2012 | 59573 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Base, 220/240V, Catalog 801764, 6 " Roller Pump Occlusion: Reference AA-2011-015 C. indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures. | Terumo Advance Perfusion System 1 | 2 | March 27, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of 27 instances in which the 6 inch Roller Pump jammed during cardioplegia delivery. There has been one report of the pump jamming during a case that concluded with an adverse patient outcome, although there was no indication that the pump jam contributed to the adverse patient outcome. | DESIGN: Component Design/Selection | "URGENT MEDICAL DEVICE RECALL" letters were sent on 3/16/12 to all consignees detailing the correction as well as also adding an Addendum to the Operator's manual for the Terumo System 1 to incorporate additional information and warnings for use when operating a 6" roller pump with dual-sized tubing. For questions call 1-800-521-2818. | Contact the recalling firm for information | 648 | 648 units | Worldwide Distribution--USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, Puerto Rico and the countries of INDIA, CANADA, CHILE, HONG KONG, EGYPT, PAKISTAN, KUWAIT, SAUDI ARABIA, JORDAN, UNITED ARAB EMIRATES, SOUTH AFRICA, SOUTH KOREA, PHILIPPINES, JAPAN, PHILIPPINES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, RUSSIA, TURKEY, BELGIUM, GERMANY, COLUMBIA, COSTA RICA, ARGENTINA, HONDURAS, COLUMBIA, DOMINICAN REPUBLIC, GUATEMALA, MEXICO, VIETNAM, and THAILAND. | Terminated | May 02, 2013 | 401 | 1642 | Not_Computer
Z-1391-2012 | 61347 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink(TM) Informatics Systems Software Version 5.0 and Software Version 5.0 Service Packs 1 through 4 used with the Dimension Vista(R) System.The EasyLink(TM) Informatics System is classified as a data calculator and processing module for use with In-vitro diagnostics devices. The EasyLink(TM) is included with the Dimension Vista(R) System, and may be used with other Siemens analyzers, as a communications and connectivity workstation for integration with laboratory information system (LIS) networks. Product Quantity Distributed (Int) 787 | EasyLink(TM) Informatics Systems | 2 | April 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | QC results may be unexpectedly associated with a Pending Control. When this situation occurs, Pending QC Controls maybe be created or reused. | DESIGN: Component Design/Selection | Siemens Healthcare Diagnostic Inc, sent a Urgent Field Safety Notice dated December 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were asked to follow these instructions. The first option is Extending the "Minimum Inactive Days" to 999 days. The second option is to proactively monitor the system for unexpected Pending QC Controls.Long Term Action Plan: A corrective and preventive action (CAPA) investigation has been issued. Currently firm is actively investigating.Further questions please call (302) 631-0516. |  | 787 | 787 | Worldwide Distribution -- USA (nationwide) and the countries of CANADA, FRANCE, GERMANY, SPAIN, PORTUGAL, ITALY, JAPAN, NORWAY, AUSTRALIA, AUSTRIA, NEW ZEALAND, BELGIUM, SWITZERLAND, NETHERLANDS, MALAYASIA, SLOVENIA, SOUTH KOREA, DENMARK, SAUDI ARABIA, SLOVAKIA, PUERTO RICO, INDONESIA and TAIWAN. | Open |  |  | 787 | Not_Computer
Z-1399-2012 | 61553 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Automatic Mobile X- Ray (AMX) series: Brivo XR285 15KW, Brivo XR285 30KW, Optima XR200 15KW, Optima XR200 30KW, Optima XR220 15KW, and Optima XR220 30KW.General Electric, Waukesha, WI Indicated for taking radiographic exposures of the skull, spinal column, chest, abdomen, extremities, and other body parts with the patient sitting, standing, or lying in the prone or supine position. | GE Automatic Mobile XRay Systems | 2 | April 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a discrepancy on the x-ray source assembly of the mobile radiographic product, related to the light field to x-ray field edge alignment. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will issue an Important Electronic Product Radiation Warning letter to customers. The letter will describe the defect and related hazards, along with the safety instructions, affected product, and product correction GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with an approved plan, the details of which will be included in a subsequent communication to customers or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the external tube port on the source assembly to address the issue.Questions or concerns can be directed to the Call Center at 800-437-1171 or your local GE Healthcare Service Representative. |  | 130 | 130 installed in US | USA (nationwide) | Terminated | January 16, 2014 | 632 | 130 | Not_Computer
Z-1462-2012 | 61514 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura X-ray System.The Allura X-ray System is an angiographic X-ray system. | Allura Xray system | 2 | April 25, 2012 | 2012 | Philips Healthcare Inc. | The monitor Ceiling suspension may fail, which may cause the monitor to drop. | PRODUCTION CONTROLS: Process Control | Philips sent an "URGENT-FIELD SAFETY NOTICE" dated March 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides information for continued safe and proper use of the equipment. Contact your local Philips representative at 866-767-2822 for further support regarding this issue. Recall expanded 5-2013 to additional 460 units distributed worldwide. | Contact the recalling firm for information | 145 | 145 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Belgium, Brazil, Canada, Czech Rep, Chili, China, Columbia, Denmark, Egypt, France, Germany, India, Indonesia, Italy, Japan, Malaysia, Mexico, Morocco, Netherlands, Oman, Russia, Saudi Arabia, Singapore, South Korea, Spain, Thailand, United Kingdom, and Vietnam. | Open |  |  | 145 | Not_Computer
Z-1617-2012 | 61694 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and TrueBeam STx V1.0, 1.5, 1.6.95 and below. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy anywhere in the body where radiation treatment is indicated. Varian Medical Systems, Palo Alto, CA. | TrueBeam and TrueBeam STx | 2 | May 16, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian received a report involving a Gantry collision and is sending a notification to remind users of the collision protection tools available with TrueBeam and TrueBeam STx, including actions the operator should be aware of. | DESIGN: Labeling Design | An "Urgent Medical Device Correction / Field Safety Letter" was distributed to all affected users on April 9, 2012, with a description of the problem and user corrective action steps. The Letter was also distributed to the Varian Sales, Marketing, and Service organizations. Customers in the US and Canada may contact (888) 827-4265 for questions.*** An additional letter was sent on 11/5/12 to additional consignees identified by Varian as being affected the the recall but not previously notified. | Contact the recalling firm for information | 207 | 207 devices 12/7/12 - amended to 216 devices. | Worldwide Distribution -- USA, and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Netherlands, New Zealand, Norway, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and the United Kingdom. | Open |  |  | 207 | Not_Computer
Z-1675-2012 | 61760 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Essenza, Nuclear Magnetic Resonance Imaging SystemProduct Usage:Nuclear magnetic resonance imaging | Siemens MAGNETOM Essenza, Nuclear Magnetic Resonance Imaging System | 2 | May 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is an error in the compatibility data sheet which is part of the System Owner Manual | PRODUCTION CONTROLS: Error in Labeling | Siemens issued a Customer Safety Advisory Notice letter to affected customers on March 30, 2012. The letter identified the affected product, problem, potential risk and what steps can the user take to avoid the potntial risk of this issue. Siemens is releasing corrected pages of the Operator Manual in order to deliver clear guidance and recommendation for the users. Customers are instructed to ensure that the safety advisory and updated pages are replaced in the System Owner Manual. |  | 97 | 97 | US Nationwide Distribution - including the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, MS, MO, NE, NJ, NY, NC, OH, OR, PA, TN, TX, WA, WI, WY and Puerto Rico. | Open |  |  | 97 | Not_Computer
Z-1821-2012 | 61877 | table, radiographic, tilting | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy DiagnostProduct Usage: Stationary Fluoroscopic system | Easy Diagnost | 2 | June 18, 2012 | 2012 | Philips Healthcare Inc. | If a lead ruler is not positioned in the same plane as the anatomical image detail to be measured, measured object distances (using the ruler) can be wrong by a corresponding factor. If, additionally, the measurement tool is not used to check the ruler, or is used in the right way (ie: calibration), such calibration error is not detected. As a result of this, a surgical intervention may be execut | DESIGN: Device Design | Philips sent an Urgent - Field Safety Notice letter dated May 11, 2012 to their affected customers. The letter identified the affected product, what the issues are, under what circumstances they may occur and the actions to be taken to avoid or minimize the occurrence of the issue Philips plans to supply an Addendum to the Instruction for Use. A Philips Service Engineer will contact the affected customers as soon as the Field Action Kit is ready to be implemented. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 6 | 6 | 6 US customers | Open |  |  | 6 | Not_Computer
Z-1829-2012 | 62052 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Multi Diagnost Eleva II with swivel cableProduct Codes;708032, 708036Multi Diagnost Eleva with FlatDetector with swivel cableProduct Codes: 708034, 708037This Philips MultiDiagnost Eleva system is intended to be used as a multifunctional/ universal system. General R/F, performed along with more specialized interventional applications. This includes the following general areas: Digestive system Skeletal system Urinary system Reproductive system Respiratory system Circulatory system Various . Arthrograms . Myelograms . Facet joint injections . Discography SialographyFluoroscopy, Radiography and Angiography | Philips Multi Diagnost Eleva II with swivel cable | 2 | June 20, 2012 | 2012 | Philips Healthcare Inc. | MultiDiagnost Eleva X--Ray system may have damage to System Cables from repeated movement | DESIGN: Device Design | Philips Healthcare Imaging issued Urgent Field Safety Notice dated May 21, 2012, to accounts informing an inspection will be performed of the unit. The Cable guiding plate will be replaced. If the cables show any damage the complete cable set wll also be replaced. No action is be taken by the customer. For further information customers should call 1-800-722-9377.For questions regarding this recall call 978-687-1501. | Contact the recalling firm for information | 433 | 433 units | Worldwide Distribution - USA (nationwide)and the countries of:NationwideForeign: AustraliaAustriaBelgiumBharainCanadaChinaColumbiaCzech RepublicDenmarkDutch AntillesEgyptEstoniaFinlandFranceGermanyHong KongIrelandIsraelItalyLatviaLebanonLuxembourgMalaysiaMexicoNetherlandsNew ZealandNorwayPhilippinesQatarRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandSyriaTaiwanThailandTurkeyUnited Kingdom | Open |  |  | 433 | Not_Computer
Z-1927-2012 | 62283 | nebulizer (direct patient interface) | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLifeÂŽ MiniSpacerÂŽ MDI Dispenser Adapter. Dual-Spray MiniSpacerÂŽ Metered Dose Inhaler (MDI) Dispenser/Adapter with 15 mm connections. The 15 mm Dual Spray MiniSpacerÂŽ is a Metered Dose Inhaler (MDI) Dispenser. It is a non-sterile, single-use item, intended to dispense aerosolized medication into a breathing circuit or ancillary circuits. MINISPACER adapters are indicated for patients on a breathing circuit, for whom aerosol medication has been prescribed, in short and long term critical care environments. The 15 mm device is installed and operated by licensed healthcare practitioners in clinical settings. The MiniSpacer is not a metering valve, and does not measure the amount of medication dispensed. The device may also be used in conjunction with a resuscitator bag or in-line with cool or heated humidification, oxygen enriched T-pieces (blow-by) systems connected to tracheal/endotracheal tube. Environments of use include short and long term critical care settings. The 15 mm Dual Spray MiniSpacers design provides standard dimension connectors (15 mm outer diameter x 15 mm inner diameter) for connection to the ventilator circuit. The device design incorporates a dual spray nozzle that allows a path for the medication to travel through the stem and directs it through the dual orifices into the airstream. Principles of Operation A healthcare practitioner installs the 15 mm MiniSpacer MDI adapter in the breathing circuit, according to facility procedures or as ordered by a physician or other licensed health care practitioner. To dispense the prescribed medication, the healthcare practitioner inserts the MDI canister stem into the devices actuator port and manually depresses the MDI canister. The MDI canisters metering valve releases the combination of medication and propellant through the valve stem and out of the devices dual spray nozzle. The plume is simultaneously ejected through upstream and downstream spray orifices along the longitudinal axis of the device. The practitioner administers the medication as prescribed, caps the MiniSpacer device, and monitors the patient response. The devices are packaged in cardboard corrugated cartons that are rated to meet burst testing up to 200 pounds per square inch. Each carton contains 30 devices; each device is packaged in a polybag with the instructions for use. These devices are sold non-sterile. | MDI Dispenser /Adapter | 2 | July 02, 2012 | 2012 | Thayer Medical Corporation | Internal data identified a molding defect in the 15 mm O.D. / I.D. MiniSpacer meter dose inhaler (MDI) adapter that results in an incomplete fill of the aerosol dispensing pathway. | DESIGN: Device Design | Thayer Medical sent an Urgent Medical Device Recall letter dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Customers were also instructed to destroy and/or dispose of the product in a proper manner and certify to Thayer Medical that they have done so by complelting the attached Certificate of Destruction. The Certificate of Destruction should be sent to Thayer Medical via fax or email. If customers prefer to return the product to Thayer Medical they should contact Customer Service at info@thayermedical.com or fax 1-520-790-5854 for instructions.For questions regarding this recall call 520-790-5393, ext 2210. |  | 19340 | 19340 | Worldwide Distribution - USA including CA, KY and Europe. | Terminated | January 15, 2013 | 197 | 19340 | Not_Computer
Z-1962-2012 | 61231 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1Electronic O2 Blender/Analyzer.The electronic gas blender provides control and monitoring of the gas flow rate and oxygen content of the gas input to the oxygenator in the perfusion circuit. It also provides monitoring of the carbon dioxide if 95-5 (%oxygen/%carbon dioxide) is used as the blending gas. | Terumo Advanced Perfusion System 1 | 2 | July 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received two complaints related to the Gas System for the Terumo Advanced Perfusion System 1. Both were reported instances of users experiencing difficulty adjusting the System 1 Gas System local flow control knob. | MISBRANDING: Labeling False and Misleading | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL-SAFETY ADVISORY" letter dated June 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. | Contact the recalling firm for information | 1095 | 1095 Units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | February 19, 2013 | 225 | 2735 | Not_Computer
Z-2056-2012 | 61242 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 220/240VThe TerumoÂŽ Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | The Terumo Advanced Perfusion System | 2 | July 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received eight complaints since September 2008 of drive motor failure of the Sarns Centrifugal System of the Terumo Advanced Perfusion System 1. A failure of the drive motor for the Terumo System 1 centrifugal system triggers error messages and alarms; therefore it is easily recognizable by the user.The potential hazards are dependent on the application | PRODUCTION CONTROLS: Equipment Maintenance | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, and an update to the Operator's Manual via Federal Express to inform consignees of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of backup system availability. Consignees were instructed to review the Medical Device Correction notice, assure that all users are aware of the notice. Place the Addendum at the beginning of the Centrifugal System Section of the Operator's Manual, Confirm receipt of the communication by faxing or emailing a completed Customer Response Form as indicated on the form. For any questions customers should call 1-800-521-2818. | Contact the recalling firm for information | 640 | 640 devices | Worldwide Distribution - USA (nationwide) and the countries of Nationwide and worldwide: ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 06, 2013 | 229 | 1632 | Not_Computer
Z-2112-2012 | 62526 | system, x-ray, film marking, radiographic | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | Beekley Altus Floral Skin Markers -Order Number: 754Product Usage:designed for Digital Mammography-Disposable skin markers for marking areas of concern or pain | Beekley Altus | 2 | July 31, 2012 | 2012 | Beekley Corporation | Digital Skin Mammography Marker separating from the backing of the spot material |  | Beekley Medical issued an Urgent Medical Device Recall letter dated July 9, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were advised to contact Beekley Medical's Customer Service Team at 1-800-233-5539 to return affected lots for replacement. Customers were instructed to complete and return the attached reply form. |  | 6420 | 6420 units | US Nationwide Distribution | Terminated | July 09, 2013 | 343 | 342910 | Not_Computer
Z-2314-2012 | 62949 | coil, magnetic resonance, specialty | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | TxRx Knee 15 Coil 3T.For use in conjunction with magnetic resonance scanners to produce diagnostic images of the knee that can be interpreted by a trained physician. | TxRx Knee 15 Coil 3T | 2 | September 06, 2012 | 2012 | Quality Electrodynamics LLC | On July 27, 2012, it was reported to the firm that the Siemens 3T 15ch Knee Coil was not functional because the coil identification code for the device was incorrect. | PRODUCTION CONTROLS: Error in Labeling | Since all of the devices were located at various Siemens distribution locations, all communications were internal to Siemens concerning bringing the coils from one warehouse to another. | Contact the recalling firm for information | 17 | 17 units | Worldwide distribution: USA (nationwide): Tennessee only and country of: Germany | Terminated | February 19, 2013 | 166 | 17 | Not_Computer
Z-2397-2012 | 62494 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS 7-color Setup beadsBD Biosciences, San Jose, CA 95131.BD FACS 7 color Setup Beads are for in vitro diagnostic use on BD FACSCanto or BD FACSCanto II systems. The beads are run with BD FACSCanto clinical software. This system provides application specific setup, including automated compensation and cytometer quality control (QC) with minimal user intervention. | BD FACS 7color Setup Beads | 2 | September 17, 2012 | 2012 | BD Biosciences, Systems & Reagents | Incorrect version of the CD was included with BD FACS 7 Color Setup Beads. | TRAINING: Employee Error | BD Biosciences sent a Important Product Recall Information letter beginning July 3, 2012, to all affected consignees via Fed Express. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed that the CD enclosed in their recent purchase of BD FACS 7-Color Setup Beads, lot number 33412, contain CD version 2.0 instead of the newer version 3.0. The letter goes on to explain the the potential risk and include a table of listed range values to determine if a potential false pass or false fail for the listed detectors has occurred. Customers are asked to check the CD with the kit to ensure that it is version 3.0 (the kit lot number ID is on the CD label). If it is not the correct version, discard version 2.0 and do not use it with this bead kit. Call BAD customers service for a replacement.If the 2.0 kit was already uses, letters request that clinicians inspect the FITC and APC-CY7 stains in the plots of the laboratory reports for the clinical applications in BD FACSCanto clinical software for compromised resolution between negative and dim populations.The verification form/effectiveness check should be completed and returned as requested. Questions or concerns may be directed to BD Customer Support at 1-855-236-2772. |  | 96 | 96 kits | Worldwide Distribution -- USA (nationwide) including the countries of Australia, Belgium, Canada, Chile, Korea, Mexico, Singapore and Taiwan. | Terminated | March 25, 2013 | 189 | 96 | Not_Computer
Z-2440-2012 | 62606 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BP7561 OPTI CCA-TS B-Lac Cassette (25 per box)Cassettes are consumables used in the OPTI CCA-TS Analyzer intended to be used for in vitro measurements of pH, PCOX2 PO2, lactate (lactic acid), total hemoglobin (tHb), and oxygen saturation (SO2), in heparinized whole blood samples (either arterial or venous) on the OPTI CCA-TS system, in either a clinical setting or point of care locations. Cassettes contain one-time use sensors that are used with the OPTI CCA-TS analyzer to perform in-vitro measurements of blood gases and metabolites. OPTI CCA-TS B-Lac Cassettes are supplied in boxes of 25 and may only used with the OPTI CCA-TS Analyzer. | BP7561 OPTI CCATS BLac Cassette | 2 | September 21, 2012 | 2012 | OPTI Medical Systems, Inc | The values reported for PCO2 by BP7561 B-Lac cassettes are biased at high concentrations when used at altitudes above 2800 feet (853 meters). | DESIGN: Process Design | OPTI Medical sent a Urgent Field Correction Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to apply correlation factors to correct the PCO2 measurement reported by the device. They were instructed to determine the elevation (altitude above sea level) for their location and to use the tables provided in the recall notification to determine the correct correlation factor. A website was provided if the consignee did not know their elevation. Steps were provided to enter the correlation factor into their OPTI CCA TS analyzer.Please contact OPTI Medical Technical Support at 1-770-510-4444 or 1-800-490-6784. | No consumer action necessary | 895 | 895 boxes (25 cassettes per box) | Worldwide Distribution -- US (nationwide) including the states of AZ, ID, TX and UT., and the countries of Argentina, Bolivia, China, Ecuador, France, Germany Greece, India, Indonesia, Iran, Ireland, Japan, Kuwait, Lithuania, Luxembourg, Malaysia, Mexico, Morocco, Nigeria, Poland, Romania, Russia, Saudi Arabia, Singapore, South Africa, Sudan Switzerland, Turkey, United kingdom, Venezuela, and Viet Nam. | Terminated | April 11, 2013 | 202 | 895 | Not_Computer
Z-2434-2012 | 63015 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Covidien Duet TRS 45 4.8 mm Universal Straight Single Use Loading Unit.Product Code: DUET4548.Single Use Loading Unit (SULU) tissue reinforcement stapler. The Endo GIA Universal Staplers with Duet TRS reloads have applications in abdominal, gynecologic and pediatric surgery for resection, transection and creation of anastomoses. They may be used for transection and resection of liver substance, hepatic vasculature and biliary structures and for transection and resection of pancreas. | Covidien Duet TRS Loading Unit | 2 | September 21, 2012 | 2012 | Covidien LP | Duet TRS may have the potential to injure adjacent anatomical structures which may result in life threatening post-operative complications. | DESIGN: Device Design | Covidien issued an "URGENT MEDICAL DEVICE RECALL" letter dated August 21, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm sent a revised customer recall notification letter via FedEx overnight delivery to all customers who purchased Duet TRS" devices from August 2009- August 2012. This letter went to all customers whether they had been sent the August 21st letter or not. Users were advised that Covidien was recalling all lots of the Duet TRS Universal Straight and Articulating Single Use Loading Units (SULU). On August 23rd 2012, Covidien notified, via e-mail, perspective surgeons that were planning to participate in the "Duet TRS Patient Registry" of the recall and that as a result the Duet TRS registry program was discontinued. Customers were asked to complete and return a Recalled Product Form via fax to (800) 895-6140. Contact Covidien Customer Service at (800) 722-8772, option 1 for questions regarding this recall. |  | 4867 | 4,867 units | Worldwide Distribution. | Open |  |  | 1078939 | Not_Computer
Z-0111-2013 | 63254 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | EVICEL Fibrin Sealant (Human) Rx Only. Manufacturer:Omrix Biopharmaceuticals Ltd. MDA blood bank, Sheba Hospital, Ramat-Gan, POB 888, Kiryat Ono 55000 Israel U.S. License No. 1603. Distributed by:Ethicon, Inc. P.O. Box 151, Somerville, NJ 08876-0151 USA.The Pressure Regulator System is positioned on the floor or on an IV pole and accepts medical gas (compressed air, nitrogen, CO2) from the central hospital system or from medical gas cylinders. | EVICEL | 2 | October 22, 2012 | 2012 | Ethicon, Inc. | Ethicon and Omrix Biopharmaceuticals have received reports of air or gas embolism occuring in association with the use of spray devices employing a pressure regulator to administer EVICEL; these events appear to be related to the use of the spray device at higher-than-recommended pressures and/or in close proximity to the tissue surface. The field correction initiated is to provide the recommend | TRAINING: Employee Error | Ethicon sent a notification letter "Urgent Product Information EVICEL Fibrin Sealant (Human) All Lots and Product Codes" dated June 27, 2012, to all affected customers. Notifications were sent via UPS on July 2, 2012. The notification informed customers of the identified risks and included instructions for use. The notification included a business reply card consignees were requested to fill out and return.Customer were notified by a notification titled "Urgent Product Information Omrix Pressure Regulator All Product Codes" that was sent by Ethicon on August 1, 2012, as a follow up to the letter dated June 27, 2012. The notification informed consignees of the identified risks and included instructions for use. Customers were notified that a company representative would visit them to confirm receipt of the notification and to ensure the enclosed label has been affixed correctly on all Omrix Pressure Regulators.For further questions please call 1-877-384-4266. |  | 4800 | 4800 pressure regulators distributed worldwide; 1500 distributed within the US | Worldwide Distribution -- USA (nationwide) and the countries of Canada, European Union, Singapore, Israel, Mexico and Brazil. | Open |  |  | 4800 | Not_Computer
Z-0290-2013 | 63314 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Mindray A3 Anesthesia Delivery System Manufactured in China for Mindray DS USA Inc. 800 MacArthur Boulevard, Mahwah, NJ 07430 This is a device used to administer to a patient continuously or intermittently a general inhalation anesthetic and to maintain a patient's ventilation. | A3/A5 Anesthesia Delivery System | 1 | November 16, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified the potential for a system leak resulting an issue related to the carbon dioxide absorber canister gasket on A3/A5 Anesthesia Delivery Systems. A potential system leak resulting from the canister gasket will likely present itself during the Automatic Circuit Leak and Compliance Test performed at startup and the Manual Leak Test recommended to be performed before each case. | DESIGN: Device Design | The firm, Mindray DS USA Inc., sent an "URGENT: A3/A5 ANESTHESIA DELIVERY SYSTEM CORRECTIVE ACTION" letter dated August 13, 2012 to their customers via Certified Mail return receipt requested. The letter described the product, problem and actions to be taken. The customers were instructed to route this letter to all potential uses of the product within their facility; increase fresh gas flow, should the leak occur during use, to compensate for any pressure loss resulting from the leak; alternatively, the absorber canister may be unlocked and locked again to reseat the gasket and verify that the canister and gasket are fully seated, a leak test is performed after replacement of the Pre-Pak or loose-fill absorbent and also be performed before each patient use until the canister gaskets are replaced on systems in their facility. A Mindray Representative will replace the canister gasket to correct the issue.If you have any questions or need to arrange for a replacement of the absorber canister gasket on the A3/A5 System(s) in your facility, contact your Mindray Service Representative, at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm).A press release was issued by the firm on 11/14/2012. | Contact the recalling firm for information | 69 | 69 units (61 US - 8 Foreign) | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Chile, Colombia, and Venezuela. | Terminated | June 04, 2013 | 200 | 415 | Not_Computer
Z-0443-2013 | 63235 | flowmeter, blood, cardiovascular | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Devices includes a Multi Port Inflator or without Multi Port Inflator (MPI). All the devices are vascular blood flow machines. Devices are distributed with pressure cuffs of varying sizes as accessories. Affected models include *2100-SX, 3000 series, 2017, 2016, 2015, 2014, 1059, 1059-A, *1059-C, *1058-C, 4000 series, 1100 series, 1060 series, 1080 series, 234 and 246. (* = current production model)All the devices are used in the study and diagnosis of vascular disease and blockages in the upper and lower extremities. | Cardiovascular blood flowmeter | 2 | December 10, 2012 | 2012 | Parks Medical Electronics, Inc | Parks Medical Electronics, Inc. was made aware of two incidents where an air supply tube from their PARKS PVR unit was inadvertently connected to an IV in the patient's arm instead of into the blood pressure cuff. |  | The firm, Parks Medical Electronics, Inc., sent an "IMPORTANT SAFETY ISSUE SERIOUS HEALTH CONSEQUENCE" letter, dated September 10 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers are advised to order LUER UPGRADE KIT WITH MPI if the Parks device includes a Multi Port Inflator (MPI), and/or order the LUER UPGRADE KIT WITHOUT MPI the Parks device does not have a Multi Port Inflator. While waiting for the kit or its installation, consignees/customers should use extreme caution when hooking up the air supply tubes to the cuff's Luer fittings. For ordering your kit, please call 800-547-6427 option 0.For responses, questions or concerns, customer should contact Parks Medical Electronics at 888-356-9522 (Phone), 503-591-9753 (Fax) or send email to support@parksmed.com. |  | 349 | 349 units (Domestic: 270 units - Foreign: 79 units) | Worldwide distribution: USA (nationwide) and countries of: Australia, Brazil, Canada, Chile, Egypt, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Korea, Mexico, Peru, Republica Argentina, Spain, Switzerland, and Venezuela. | Open |  |  | 349 | Not_Computer
Z-0508-2013 | 63682 | system, test, carbohydrate antigen (ca19-9), for monitoring and management of pancreatic cancer | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens Healthcare ADVIA Centaur CA 19-9250 Test 1049124450 Test 10491379Intended Use The ADVIA Centaur CA 19-9 Assay is an in vitro immunoassay for the quantitative measurement of the CA 19-9 tumor-associated antigen, in human serum, using the ADVIA Centaur and ADVIA Centaur XP systems. This assay is indicated for the serial measurement of CA 19-9 to aid in the management of patients diagnosed with cancers of the exocrine pancreas. | Siemens Healthcare ADVIA Centaur CA 199 | 2 | December 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Operating range of the ADVIA Centaur systems is 18C to 30ÂżC (64.4ÂżF to 86.0ÂżF), a change in room temperature may increase or decrease assay results, depending on the specific assay and ADVIA Centaur system used. | DESIGN: Device Design | Siemens Healthcare issued An Urgent Medical Device Correction Notice (for US customers) and an Urgent Field Safety Notice (for customers outside the US) dated November 5, 2012, to all affected customers. These notices inform customers who used the ADVIA Centaur BNP, BR, CA 19-9, Digitoxin, Folate, and Theophylline 2 assays to reduce their instrument operating temperaturerange to 20 -25 degrees C and for the DHEAS assay to reduce theinstrument operating temperature range to 18 -26 degrees C when theseassays are being used. These notices also direct customers to review thecontents of the correction notice with their Laboratory Director or MedicalDirector. A confirmation fax-back form was included to ensure that allcustomers worldwide have been notified and have complied with themandatory action.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local technical support representative.For questions regarding this recall call 508-668-5000. |  | 13068 | 250 test=525 and 50 test=12243 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries: ArgentinaAustraliaAustriaAzerbaijanBelgiumBosnia Herzeg.BrazilCanadaChileCOLOMBIACroatiaCzech RepublicDenmarkEgyptEstoniaFinlandFranceGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIranIrelandItalyJapanKazakhstanLatviaLithuaniaMalaysiaMexicoNetherlandsNew CaledoniaNorwayOmanP.R. ChinaPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaReunionRomaniaRussian Fed.Saudi ArabiaSerbiaSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTurkeyU.A.E.United KingdomUruguayVietnam | Open |  |  | 12768 | Not_Computer
Z-0383-2012 | 34843 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | Product is 40mm Flexible Tube. ***NON-STERILE PRODUCT***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN."W.LORENZ SURGICAL***WALTER LORENZ SURGICAL, INC. 1520 TRADEPORT DRIVE , REPLY TO POST OFFICE BOX 18009, JACKSONVILLE, FLORIDA 32229-8009 USA, 1-800-874-7711. Part # 80-3111-01 Lot 052190 QTY 76 Pieces 10/18/2005.Sold for use with:Part # 80-3111-01 Lot 052190 Qty 14. Part # SP-1651 Lot # 069330 Qty 10.Part # SP-1639 Lot # 046230 Qty 4. Part # SP-1598-13 Lot # 041960 Qty 10. Part # SP-1329 Lot # 039700 Qty 10. Part # SP-1331 Lot # 039720 Qty 12. Part # SP-1330 Lot # 039710 Qty 11. Part # SP-1633 Lot # 056630 Qty 5. Product Usage: Distraction devices are for use in bone stabilization and elongation (lengthening) when correction of oral, cranial, mandibular and maxiollofacial deficiencies or post traumatic defects require gradual bone distraction. Each implantable distraction device has a titanium alloy drive screw mechanism and LactoSorb resorbable correction plates, which are implanted by fixation with LactoSorb bone screws. | LORENZ RESORBABLE DISTRACTOR SYSTEM | 2 | January 11, 2012 | 2012 | Walter Lorenz Surgical, Inc. | Lorenz Engineer discovered flex tube could break if bent too much. | DESIGN: Device Design | W. Lorenz Surgical notified consignees of recall by an Urgent Advisory Notice letter. Consignees were requested to acknowledge receipt of recall and provide product disposition. Initial letter was sent on March 2, 2006, March 15, 2006, follow-up letters sent on April 4, 2006 and May 15, 2006. For additional information call (904) 741-4400 or fax (904) 741-4500. |  | 76 | 76 pieces | Worldwide Distribution - USA (nationwide) including the states of: GA and TX and the country of Japan. | Terminated | December 12, 2011 | -30 | 182 | Not_Computer
Z-0468-2012 | 51146 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer Hydrogen Peroxide Gas Plasma Sterilizer Product Code 10133. The STERRAD Sterilization System is a low-temperature general-purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | STERRAD NX Sterilizer | 2 | January 11, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has discovered a defect in the barcode labeling of the STERRAD NX System Cassettes 10133, Lot #08I032. This is the only lot number affected by this action. The barcodes for Lot #08I032 were printed improperly and are not capable of being read by the STERRAD unit. | PRODUCTION CONTROLS: Error in Labeling | Advance Sterilization Products sent an "URGENT: PRODUCT RECALL" letter dated December 10, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine inventory and return all sealed cassettes, using an enclosed prepaid UPS return label to: Stericycle, 2670 Executive Drive Suite A, Indianapolis, IN 46241. Additionally, a Business Reply Card and packing slip were included with the letter for customers to complete and return. Questions or further assistance was directed to ASP Customer Care Center at 888-783-7723, option 2. | Contact the recalling firm for information | 776 | 776 cases | Nationwide Distribution. | Terminated | December 16, 2011 | -26 | 776 | Not_Computer
Z-0556-2012 | 57863 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Appraise A1C Test Pak Kit (Alere), Product K7300, each kit packed in a bag or envelope and contains 1 Triad Medium Alcohol Prep Pad, Cat. No. 10-3001, NDC 50730-3001-0, and blood specimen collection supplies, forms, and instructions.Blood specimen collection | Appraise A1C Test Pak Kit (Alere | 2 | January 11, 2012 | 2012 | Heritage Labs International LLC | Kits contain recalled Triad alcohol pads | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Heritage Labs sent a Device Notification and Safety Alert letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Heritage Labs sent replacement alcohol prep pads from a different supplier not affected by the recall to user of Heritage Labs kits. Heritage Labs recommended the replacements be on "as needed" basis. Upon opening kit, users should immediately dispose of sealed TRIAD brand alcohol pads. Users are to use the replacement wipes instead of the wipes contained in the kits. For any questions call (888)764-4120. |  | 1883 | 1,883 kits | Worldwide Distribution - USA (nationwide) and Germany | Terminated | December 26, 2012 | 350 | 683376 | Not_Computer
Z-0393-2012 | 60386 | tampon, menstrual, unscented | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Kotex Natural Balance Security tampons, regular absorbency, unscented. packaged in 18 count or 36 count retail containers. Consumer Services, Kimberly-Clark Corp. Dept. KNBTR-18, P.O. Box 2020, Neenah, WI 54957-2020 USA.To absorb menstrual or other vaginal discharge. | Kotex tampons | 2 | January 11, 2012 | 2012 | Kimberly-Clark Corporation | A limited number of Kotex Natural Balance Security Unscented Tampons, regular absorbency, are being recalled after a raw material contaminated with a bacterium, Enterobacter sakazakii, was found when testing the plunger portion of the applicator. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Kimberly-Clark sent an Urgent Product Recall letter, dated 11/09/2011, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine, and discontinue use immediately. Customers were asked to complete the response form and fax to +1 (920) 380-6467. A Kimberly-Clark representative would contact them regarding return of the product, replenishement and/or account credit. For questions regarding this recall call +1 (920) 216-8152. |  | 1400 | 1400 Cases (216 tampons per case) | Nationwide Distribution | Terminated | December 13, 2012 | 337 | 1400 | Not_Computer
Z-0508-2012 | 60325 | holder, needle; orthopedic | HXK | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Offset Wire Fixation Bolt, catalog number 8180-50-006. The bolts are used with the DePuy Ace-Fischer External Fixation System and are packaged in a flexible polyethylene bag. | DePuy Offset Wire Fixation Bolt used in AceFischer External Fixation System | 2 | January 11, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy has received three complaints concerning oversized bolt threads in the Offset Wire Fixation Bolts used in the Ace- Fischer External Fixation System. This could lead to the inability to properly thread the nut onto the bolt. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | On 11/9/2011 DePuy sent "Medical Device Recall Notice" letters to Distributors and to direct account hospitals. Recall notices included instructions to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN via distributors and sales representatives. The product is going to be discontinued and no reconditioning or reworking for the product is planned. For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - pm EST.). |  | 890 | 890 | Nationwide Distribution | Terminated | March 07, 2013 | 421 | 890 | Not_Computer
Z-0350-2012 | 60434 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA AQUAFORM Comfort Science enfilcon A Daily Wear SPHERE Soft Contact Lenses (disposable), Rx only, STERILE --- CooperVision: Hamble, SO31 4RF, UK -- Scottsville, NY 14546 USA -- Juana Diaz, PR 00795 USA --- Device Listing # D033736. Packaged: 6 Soft Contact Lenses per unit carton or trial product 3 pack.AVAIRA (enfilcon A) SPHERE soft contact lenses are indicated for the correction of ametropia (myopia or hyperopia) in aphakic and non-aphakic persons with non-diseased eyes. The lenses may be worn by persons who exhibit astigmatism of 2.00 diopters or less that does not interfere with visual acuity. | AVAIRA (enfilcon A) Daily Wear Sphere Contact Lens | 1 | January 11, 2012 | 2012 | Coopervision Inc. | Presence of silicone oil residue on Avaira Sphere Soft Contact Lenses. | PRODUCTION CONTROLS: Process Control | CooperVision Inc. issued a press release on November 15, 2011. An "URGENT: MEDICAL DEVICE RECALL- AVAIRA SPHERE CONTACT LENSES" letter dated November 16, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory, stop further distribution of the affected product, contact their patients and to advise their patients to check the firm's website to confirm they have the affected product. Additionally, a response form was included with the letter for customers to complete and return via fax at 866-484-1048. Contact Coopervision at 800-341-2020 for questions regarding this recall. | Contact the recalling firm for information | 4877897 | Domestic: approx. 3,155,781 lenses; Foreign: estimated 1,722,116 lenses | Worldwide Distribution--USA (nationwide) including Puerto Rico. | Open |  |  | 4877897 | Not_Computer
Z-0454-2012 | 60359 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Dental cement under the brand name Cement-It and Natural Elegance (private label of Cement-It).The intended use of this device is to affix dental devices such as crowns or bridges onto the tooth. Cement-It and Natural Elegance (private label of Cement-It) are indicated as a luting cement bonding system for porcelain crowns, inlays, and onlays and other dental restorations, including ceramics, metal alloys and composite restoratives. | Dental cement under the brand name CementIt and Natural Elegance | 2 | January 11, 2012 | 2012 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | The firm initiated a recall because the product sets faster than specified in the "Directions for Use" and in release specifications. | OTHER/UNDETERMINED: Pending | The firm, Pentron Clinical, sent an " URGENT: MEDICAL DEVICE RECALL" letters dated April 4, 2011 and April 20, 2011 via USPS 1st class mail to its customers. The letter describes the product, problem and actions to be taken. The letter instructed customers to review the table with affected product lots and determine if the affected products was in their inventory; return the affected product for a replacement at no charge; contact Pentron Clinical Customer Service at (800) 551-0283; 1-800-243-3969 ext. 677 or email: aestridge@pentron.com directly to handle the arrangements of a quick return and replacement; and complete and fax back the enclosed Recall Return Form at 1-877-677-8844. The letter also request the customer to identify and recover the affected product lots that may have been shipped to thier customers.If you have any questions, contact Pentron Clinical Customer Service at (800) 551-0283. |  | 157 | 157 | Worldwide distribution: USA (nationwide) including states of: PA, IL, NV, WA, OH, TN, WI, NY, NJ, CT, IN FL, MI, OK, NE, HI, VA, CA, and TX; and countries of : Canada, Egypt, Isreal, Netherlands, Sweden and UK. | Terminated | January 22, 2014 | 742 | 157 | Not_Computer
Z-0417-2012 | 60447 | prosthesis, knee, femorotibial, constrained, cemented, metal/polymer | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS INTERLOK BOWED IM STEM W/SCREW 14 MM X 150 MM TI 6AL 4V ALLOY.Hip and knee joint replacement prostheses intended for primary and revision joint arthroplasty for use in cemented applications. | OSS Cemented IM Stem | 2 | January 11, 2012 | 2012 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that Part 150367, (OSSCemented 1M Stem 13x150) was incorrectly labeled as part 150368 (OSS Cemented 1M Stem14x150). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Biomet, Inc. sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 1, 2011 to the one distributor with the affected product. The letter described the product, problem, and actions to be taken by the customer. The letter instructs the customer to immediately locate and remove the affected product from inventory. Additionally, the customer was asked to complete and return a Response Form via fax to 574-372-1683. Call 574-372-3983 for questions related to this notice. |  | 3 | 3 units & 1 pending release | Distributed only in The Netherlands. | Terminated | September 10, 2012 | 243 | 4 | Not_Computer
Z-0394-2012 | 60440 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | OR table TITAN, Surgical Table.Surgical bed with Orthopedic Accessories. | Titan | 2 | January 11, 2012 | 2012 | Trumpf Medical Systems, Inc. | Product is being recalled due to unintended movement on the Titan Tables. |  | Trumpf Medical Systems sent an initial notification to customers dated January 24, 2011. A subsequent "URGENT SAFETY NOTICE-UPDATE" letter dated May 26, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. A Corrective Action Plan includes a software update to be installed by a Trumpf Technical Service Representative. Contact Technical Service at 888-474-9359 for questions regarding this notice. |  | 94 | 94 beds total for all types. | Worldwide Distribution. | Terminated | August 13, 2012 | 215 | 98 | Not_Computer
Z-0596-2012 | 60504 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, pre-filled flush solution, for IV Flush only, sterile fluid path; 10 mL Single-Use Syringe with Male Luer Lock, 100 syringes per pack, 4 packs per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list no. 1978-20lock flush solution | 0.9 Sodium Chloride Flush Syringe | 2 | January 11, 2012 | 2012 | Hospira Inc. | There is particulate matter in the fluid pathway, which has been identified as the same material as the rubber tip of the syringe plunger. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Hospira sent a "URGENT DEVICE RECALL" letter dated November 18, 2011, to all affected customers via UPS. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to check their inventory and immediately quarantine any affected product complete and return Reply Form via fax to 1-866-912-2512. Return affected product to Stericycle using the label provided with the letter and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability.Hospira expanded the recall to include thirteen additional affected lots via letter dated January 31, 2012. The recall instructions remain the same as in the November 18,2011 letter. |  | 5817600 | 5,817,600 syringes | Nationwide distribution including states of: Alabama, Alaska, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Mississippi, Missouri, Nebraska, New Hampshire, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, West Virginia and Wisconsin. | Open |  |  | 5817600 | Not_Computer
Z-0447-2012 | 60421 | drills, burrs, trephines & accessories (compound, powered) | HBF | Powered compound cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Box labeled in part, "REF: 72-1000***IQ SYSTEM IQ INTELLIGENT DRIVER 16.51CM (6.5")(L) RABALON, STAINLESS STEEL, GOLD PLATED***NON-STERILE PRODUCT*** MADE IN THE PHILIPPINES*** BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***"The Biomet Microfixation IQ Intelligent Driver is a handheld cordless device designed for three functions: drill pilot holes, high-speed insertion of 1.5x4mm titanium screws, and high-speed insertion of 1.5x4mm and 1.5x5mm Lactosorb fixation screws provided by Biomet Microfixation. The IQ Intelligent Driver is indicated to aid in the implantation of screws (Drilling holes and inserting screws). | iQ Intelligent Device | 2 | January 11, 2012 | 2012 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc., Jacksonville, FL, recalled their IQ Intelligent Device, 72-1000 due to failure of the front bit sensor to detect the correct bit that is installed. When a screw is inserted with the device in drill mode, it is possible for the screw or bone to strip. | DESIGN: Component Design/Selection | Biomet Microfixation called, emailed, or sent an Urgent Medical Device Recall letter dated July 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the remaining device(s) at their facility and return immediately to PIVOT for the required software upgrade. Customers were also instructed to return the attached reconciliation form to:PIVOT International10916 Strang Line RoadLenexa, KS 66215Attn: Julie LawrenceFor any questions regarding this recall call 1-800-874-7711 or 904-741-4400, ext 9468. | No consumer action necessary | 161 | 161 devices were distributed. | Worldwide Distribution - USA AZ, CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, LA, NC, NH, NY, OK, OR, PA, TN, TX, and WI and the countries of Argentina, Brazil, France, Italy, Japan, South Africa, and Thailand. | Terminated | March 09, 2012 | 58 | 161 | Not_Computer
Z-0446-2012 | 59231 | methyl methacrylate for cranioplasty | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Osteopore PCL Scaffold Bone Filler: Osteopore Osteoplug-C, Osteoplug, and Osteomesh. | Osteopore | 2 | January 11, 2012 | 2012 | OSTEOPORE INTERNATIONAL PTE LTD | Product sterility may be compromised. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Osteopore sent an Urgent Voluntary Product Recall notice to their consignees by letter dated April 27, 2011. | Contact the recalling firm for information | 330 | 330 | Nationwide distribution - one distributor in TX. | Terminated | June 06, 2012 | 147 | 330 | Not_Computer
Z-0461-2012 | 60542 | chromogenic in situ hybridisation, nucleic acid amplification, her2/neu gene, breast cancer | NYQ |  |  | Immunology |  | SPOT-Light" HER2 CISHÂŽ Kit, catalog #84-0150, 20 tests/kit, labeled in part ***Invitrogen Corporation 7335 Executive Way Frederick MD 21704 www.invitrogen.com***Product Usage: The SPOT-Light HER2 CISH Kit is intended to quantitatively determine HER2 gene amplification in formalin-fixed, paraffin-embedded (FFPE) breast carcinoma tissue sections and is indicated as an aid in the assessment of patients for whom HerceptinÂŽ (trastuzumab) treatment is being considered. | SPOTLight" HER2 CISH Kit | 2 | January 11, 2012 | 2012 | Life Technologies Corporation | In vitro diagnostic reagent may be contaminated with a fungal contaminant. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Life Technologies sent an URGENT DEVICE RECALL letter dated November 14, 2011 by mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return remaining inventory. Customers were advised to complete the Customer Response Sheet included for recovering inventory data and fax back to (+1) 716-774-6727 Technical Support or e-mail to techsupport@lifetech.com. Customers who had distributed the affected kits outside of their facility were instructed to immediately notify their customers of the recall. For questions call Technical Support, at (+1) 800-955-6288 opt. 2 then opt. 3. |  | 29 | 29 kits | Worldwide Distribution - USA (nationwide) in the states of CA and NV and the countries of: Canada, Venezuela, and the UK. | Terminated | February 10, 2012 | 30 | 29 | Not_Computer
Z-0445-2012 | 60565 | lens,intraocular,accommodative | NAA | Intraocular lens. | Ophthalmic | Ophthalmic |  | CrystalensÂŽ Accommodating Posterior Chamber Intraocular Lens, AT-50AO.CrystalensÂŽ Accommodating Posterior Chamber Intraocular LensPMA No.: P030002.The CrystalensÂŽ is intended for primary implantation in the capsular bag of the eye for the visual correction of aphakia secondary to the removal of a cataractous lens in adult patients with and without presbyopia. | Crystalens Accommodating Posterior Chamber Intraocular Lens | 2 | January 11, 2012 | 2012 | Bausch and Lomb, Incorporated | Bausch + Lomb voluntary initiated the recall due to a physician complaint of haze on the Crystalens AOwhich was visualized using an angled beam from a slit lamp during the post-operativeexamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Bausch + Lomb sent an "URGENT- MEDICAL DEVICE RECALL" letter dated November 14, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to remove the affected products from inventory and return the lenses to the firm. Additionally, a Recall Acknowledgement Form was attached to the letter for customers to complete and return with the affected products. Contact Bausch + Lomb at 949-521-7895 for questions regarding this recall.. | No consumer action necessary | 62 | 62 units | Nationwide Distribution | Terminated | May 23, 2012 | 133 | 62 | Not_Computer
Z-0489-2012 | 60502 | system, test, carbohydrate antigen (ca19-9), for monitoring and management of pancreatic cancer | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA19-9 AssayCat. Nos.123521 (03481938) - 250 Test kit w/o Calibrators123519 (04612750) - 50 Test kit w/o Calibrators10491244 - 250 Test kit w/Calibrators10491379 - 50 Test kit w/CalibratorsThe ADVIA Centaur CA 19-9 Assay is anÂŽ in vitro immunoassay for the quantitative measurement of the CA 19-9 tumor-associated antigen, in human serum, using the ADVIA Centaur and ADVIA Centaur XP systems. This assay is indicated for the serial measurement of CA 19-9 to aid in the management of patients diagnosed with cancers of the exocrine pancreas. The test is useful as an aid in monitoring of disease status in those patients having confirmed pancreatic cancer who have levels of serum CA 19-9 at some point in their disease process exceeding the median concentration determined for the apparently healthy cohort. CA 19-9 values must be interpreted in conjunction with all other clinical and laboratory data before a medical decision is determined. | ADVIA Centaur CA199 Assay | 2 | January 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Positive bias of 14% and higher for some individual points for CA 19-9 results generated on the ADVIA Centaur CP system when they are compared to CA 19-9 assay results generated on either the ADVIA Centaur or ADVIA Centaur XP systems | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens Healthcare Diagnostics issued a Urgent Device Notification email dated November 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers conduct a look back if they are using the ADVIA Centaur CA 19-9 interchangeably with the ADVIA Centaur or ADVIA Centaur XP systems and the ADVIA Centaur CP system and to review the contents of the recall notice with their Laboratory Director. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory actionFor further questions please call (508) 359.3825. |  | 43053 | 50 Test Kits =21,160; 250 Test Kits = 21,593 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Malyasia, Korea, Hong Kong, Singapore, Myanmar, Mexico, Egypt, Taiwan, South Africa, Brazil ,India, Japan, Argentina, Chile, Peru, Germany, Malaysia, Saudi Arabia, and Canada. | Terminated | April 12, 2013 | 457 | 43053 | Not_Computer
Z-0457-2012 | 60571 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | St. Jude Medical Riata (8Fr), Models: 1560, 1561, 1562, 1570, 1571, 1572, 1580, 1581, 1582, 1590, 1591, 1592Product Usage: These Silicone transvenous high voltage (HV) leads connect an implantable cardioverter-defibrillator (lCD or CRT-D) to cardiac tissue to monitor and regulate a patients heart rate by providing pacing therapy and/or tachyarrhythmia therapy. | Riata | 1 | January 11, 2012 | 2012 | St Jude Medical CRMD | St. Jude Medical has confirmed the failures associated with all cause insulation failure on the St. Jude Medical Riata and Riata ST Silicone Endocardial Defibrillation Leads with specific emphasis on externalized conductors. |  | St. Jude Medical sent a Medical Device Advisory Important Product Information Update letter dated November 28, 2011, Physician Communication dated December 15, 2010, Product Performance Report Extracts, and a Physician Device Advisory Notification to physicians by FedEx priority overnight service for receipt on November 29, 2011. The letters provides the customers with an explanation of the problem, the product, and the actions to be taken. The issuance of the recall letter is in conjunction with St. Jude's recent release November 2011 Product Performance Report (available online at http://sjmprofessional.com). The letter also explains about the Lead Peformance Summary, Root Cause, Clinical Implications, Rate of Occurrence from Complaints and Returns, New Peer Reviewed Publications, Data on Optim Insulated Leads (Durata and Riata ST Optim), and Recommendations and Mitigations. Physicians with any questions or concerns can contact their local St. Jude Medical representative or the Technical Services Department at (800) 722-3774. | Contact the recalling firm for information | 128000 | 128,000 units total | USA (Nationwide) and worldwide to: Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hungary, Iceland, Italy, Romania, Russia, Netherlands, Norway, Poland, Portugal, Russia, Serbia, Slovakia, Spain, Sweden, Switzerland, The Netherlands, UK, Australia, China, Hong Kong, Indonesia, Japan, Korea, Malaysia, Philippines, Singapore, Taiwan, Thailand, Vietnam, New Zealand, Argentina, Brazil, Colombia, Costa Rica, Ecuador, Mexico, Puerto Rico, Uruguay, Egypt, Greece, India, Iran, Israel, Joran, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, South Africa, Turkey, Malta, Cyprus, Iraq, Egypt, Lybia, Yemen, and U.A.E. | Terminated | May 09, 2012 | 119 | 128000 | Not_Computer
Z-0590-2012 | 59833 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DURAMAX 15.5 F x 36 cm Chronic Hemodialysis Catheter Set (With Cuff 31 cm from Tip), DuraMax Stacked Tip 36 cm Str. Basic Kit, Catalog No./REF 10302806, STERILE --- AngioDynamics, Inc. 603 Queensbury, NY 12804 USAThe AngioDynamics DuraMaxÂŽ Chronic Hemodialysis Catheter is indicated for use in attaining long-term vascular access for Hemodialysis and apheresis. | DuraMax Chronic Hemodialysis Catheter | 2 | January 11, 2012 | 2012 | Angiodynamics, Inc. | AngioDynamics received customer complaints for breakage of the tunneler sleeve during product use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, AngioDynamics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 28, 2011 and response form via Certified Mail, Return Receipt Requested, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to identify and segregate the recalled lots; complete and return the enclosed DuraMax HemoDialysis Catheter Recall Reply Form via fax to: Attn: DuraMax HemoDialysis Catheter Recall Coordinator at 518-798-1360 (even if they don't have any product); and do not use the affected product, return and ship the recalled product to ANGIODYNAMICS (replacement product will be shipped upon receipt and confirmation of the returned product). If the customers are in immediate need for replacement product, they should fax the completed Recall Reply Form, along with a copy of the product label, to Customer Service at (518) 798-1360 or call us at (800) 772-6446, x1358 or x1363.If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call ANGIODYNAMICS customer service at 1-800-772-6446 or email to: customerservice@angiodynamics.com. | Contact the recalling firm for information | 39 | 39 units | Nationwide distribution. | Open |  |  | 1124 | Not_Computer
Z-0690-2012 | 60593 | detector and alarm, arrhythmia | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Module is the Spacelabs Multi-parameter Module. It is intended for use with the Patient Care Management System (PCMS) to acquire, monitor, and process various clinical parameters from an adult or neonate/infant populations in any type of clinical environment other than home use.The Spacelabs Multi-parameter Module is intended for use with the Patient Care Management System (PCMS) to acquire, monitor, and process various clinical parameters from an adult or neonate/infant populations in any type of clinical environment other than home use. Physiological parameters that may be monitored include ECG with arrhythmia detection, respiration, invasive and noninvasive blood pressure, temperature, oxygen saturation (Sp02) and cardiac output. Acquired data may then be communicated to an information network for display, recording, editing and analysis. | Ultraview SL Command Module | 2 | January 11, 2012 | 2012 | Spacelabs Healthcare, Llc | NIBP pump inside the Ultraview SL Command Module, Model 91496, became loose. The pump is suspended by the air hoses and electrical connections, but there can be enough movement todamage internal components and cause the module to fail. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | SpaceLabs Healthcare sent a Urgent Field Safety Notice dated November 28, 2011, return receipt requested to all US customers. The letter identified the product the problem and the action needed to be taken by customer.On December 6, 2011, SpaceLabs will send a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record. Records of notification will be maintained to ensure complete notification.Customers were instructed If you hear an unusual rattling noise, tag the module, set it aside and contact your Spacelabs field service engineer. Your field service engineer will be contacting you early in 2012 to schedule a mutually agreeable time to upgrade all of your affected modules. Upgrades will be performed on-site and at no cost to you.For any questions about this corrective action program, please call Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support or, outside the United States, 1-425-657-7200 x5089. | Contact the recalling firm for information | 6722 | 6722 units worldwide (3598 units in US and 3124 units outside US) | Worldwide Distribution -- USA (nationwide) | Terminated | September 19, 2012 | 252 | 6722 | Not_Computer
Z-0519-2012 | 60518 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Smith & Nephew, SPIDER2 Limb Positioner, including accessories Battery, foot pedal and chargerCatalog Number: 72203299.Table, Operating-room, Pneumatic. | Smith & Nephew SPIDER2 Limb Positioner | 2 | January 11, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | SPIDER2 may unlock (loss of traction) resulting inpotential movement of the patient when the device is within close proximity to an RF source. Lossof traction can potentially result in the unintended repositioning of the extremity and potential for tissue or cartilage damage | DESIGN: Device Design | Smith & Nephew Advanced Surgical Devices Division sent an "URGENT- PRODUCT RECALL NOTIFICATION-URGENT" letter dated July 26, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return a form via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this recall. | Contact the recalling firm for information | 52 | 52 units | Worldwide Distribution-USA (nationwide) and the countries of Canada and France. | Open |  |  | 52 | Not_Computer
Z-0459-2012 | 60525 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR linear accelerators with G31 mother board.ARTISTE and ONCOR are manufactured by Siemens Medical Solutions, Roentgenstrasse 19-21, Kemnath, Germany.ARTISTE, ONCOR and PRIMUS family of linear accelerators deliver X-ray photon and electron radiation for therapeutic treatment of cancer. | Siemens ARTISTE and ONCOR linear accelerator | 2 | January 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Reports of gantry rotation without receiving a movement command caused by an overload within the drive signal power line. | DESIGN: Component Design/Selection | Siemens Medical Solutions USA, Inc. sent a Customer Information letter dated November 9, 2011, by federal express. The letter identified the product the problem and the action needed to be taken by the customer.The HW Update included information regarding the potential problem and provide instructions in the Safety update.Further questions call (925) 602-8157. |  | 13 | 13 active devices | Worldwide Distribution-- USA ( nationwide ) including the states of : MO, NC, OH, and OK and the countries of Turkey, Japan, Germany and Portugal | Terminated | April 15, 2013 | 460 | 13 | Not_Computer
Z-0595-2012 | 60628 | system, photopheresis, extracorporeal | LNR |  |  | Gastroenterology/Urology |  | Therakos Cellex Photopheresis SystemCellexUSA & Cellex.The photopheresis system is indicated for use in the ultraviolet-A (UVA) irradiation in the presence of the photoactive drug 8-methoxypsoralen (8-MOP) of extracorporeally circulating leukocyte-enriched blood in the palliative treatment of the skin manifestation of cutaneous T-cell lymphoma (CTCL) in person who have not been responsive to other forms of treatment. The Cellex procedural kits are designed to interface with the Cellex Photopheresis System to perform cell separation and photo activation in a single closed and sterile circuit. This is accomplished by use of an integrated centrifuge bowl and associated tubing for separation of whole blood into red blood cells, plasma, and buffy coat (leukocyte-enriched blood). All blood components are returned to the patient. The procedural kits are fully assembled as received by the customer and the customer is only required to place the components in the Cellex System to enable the system for use. | Therakos Cellex Photopheresis System | 2 | January 11, 2012 | 2012 | Therakos, Inc. | There have been three confirmed reports of post reinfusion hemolysis or hemoglobinuria in patients treated on the Therakos Cellex Photopheresis System. | OTHER/UNDETERMINED: Pending | Therakos sent an "URGENT PRODUCT CORRECTION NOTIFICATION" dated September 1, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Additionally, a Confirmation of Receipt was attached to the letter for customers to complete and return. Contact Therakos Technical Support at 1-877-865-6850 for questions regarding this notice. | Contact the recalling firm for information | 114 | 49 USA, 65 Foreign | Worldwide Distribution. | Open |  |  | 114 | Not_Computer
Z-0606-2012 | 60626 | prosthesis, knee patellofemorotibial, partial, semi-constrained, cemented, polymer/metal/polymer | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Eius Unicompartmental Knee SystemStryker Howmedica Osteonics Corp.Howmedica International S. de R.L., Raheen Business Park, Limerick, Ireland.Intended for use in Unicompartmental Knee Arthroplasty (UKA). Surgeon users have identified the following potential advantages of UKA procedures versus Total Knee Arthroplasty procedures generally: conservation of the opposite compartment, preservation of ligaments and patellofemoral articulation, reduced loss of bone stock and lower implant cost. The EIUS UKA is a Resurfacing device that requires the removal of a thin condylar cross-section intended to replace the diseased cartilage only. This allows the surgeon to conserve good quality subchondral bone. | Eius Unicompartmental Knee System | 2 | January 11, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Two National Joint Registries suggest that the EIUS Unicompartmental System is associated with a higher revision rate then the unicompartmental devices. | OTHER/UNDETERMINED: Pending | Stryker Orthopaedics sent "URGENT PRODUCT RECALL" letters dated August 17, 18 and 25, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgement Form via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions concerning this recall. |  | 26648 | 26,648 units | Nationwide Distribution. | Open |  |  | 26648 | Not_Computer
Z-0455-2012 | 60551 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Artiste Maverick Tint, Part Number: N36TB, Lot/Serial Numbers: 3416180, Product Description and Intended Use: The intended use of this device is for internal color characterization and includes 3 opaques for blocking out deep staining or metal substrates.Product Usage: The intended use of this device is for internal color characterization and includes 3 opaques for blocking out deep staining or metal substrates. | Artiste Maverick Tint | 2 | January 11, 2012 | 2012 | Kerr Corporation | The resin material separates when the product is extruded from the syringe. Although there are no safety issues related with this situation, please refrain from using the affected product. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pentron Clinical sent an URGENT: Medical Device Recall letter dated April 20, 2011 to all affected customers. The letter described the product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all product(s) subject to recall for return and exchange at no charge. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were also asked to complete and return the enclosed Recall Return Form. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangement of a quick return and replacement. |  | 100 | 100 Units | Worldwide Distribution - USA (nationwide) and the countries of: Armenia, Russia and Ukraine. | Terminated | January 09, 2012 | -2 | 100 | Not_Computer
Z-0509-2012 | 60587 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Benephit XT Infusion System, Targeted Renal Therapy Infusion System, Catalog/REF 70035, STERILE -- AngioDynamics Inc. 603 Queensbury, NY 12804 USA --- COMMON/USUAL NAME - Continuous flush catheter; CLASSIFICATION NAME: Continuous flush catheter; Device Listing # D069932Benephit Infusion Systems are intended to facilitate targeted renal therapy, or TRT, the delivery of physician-specific agents to the kidneys via the renal arteries. The Benephit Infusion Systems are indicated for use in patients undergoing medical, interventional, or surgical procedures, where the procedure carries an elevated risk of iatrogenic kidney injury for the patient. The Benephit Infusion Systems are also indicated to facilitate TRT in patients who have demonstrated symptoms of acute kidney injury, and whom arterial catheterization for TRT is feasible. | Benephit | 2 | January 11, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Benephit XT Infusion System may contain hairline cracks in the sheath hub. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Angiodynamics Worldwide Headquarters sent a Urgent Medical Device Recall Letter and Response Forms dated December 2, 2011, to all affected customers via Certified Mail on December 5, 2011. The recall is being extended to the medical user level. Customers will be instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. A Return Authorization number (RA #) will be assigned to each customer in the recall notification letter and a UPS account will be provided for return freight. A reply form was provided with the recall notification letter to be completed by each consignee. ACTIONS TO BE TAKEN BY CUSTOMER / USER: (1.) Identify and segregate the recalled lots that are in your possession. (2.) Complete the enclosed Benephit XT Infusion System Recall Reply Form and fax it to the attention of the Benephit XT Infusion System Recall Coordinator at 518-798-1360. The form lists the catalog number, lot numbers, and quantity our records indicate your facility has received. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number on the outside of the shipping box and include a copy of the Benephit Infusion System Recall Reply Form with your shipment. (4.) If you have distributed the affected product/lot, please inform your customers of this recall action immediately, and have them return the recalled units to you. --- Certified Mail will verify delivery to each domestic consignee. Non-respondents will be contacted by telephone or via e-mail. --- Corrective Action: Returned product will be quarantined and final disposition has not been determined at this time. A corrective action has been initiated by AngioDynam. |  | 55 | 55 units domestically | Nationwide Distribution including the states of New York, Pennsylvania, Virginia, Kentucky, Iowa, Illinois, Missouri, Louisiana, Texas, and Utah. | Open |  |  | 55 | Not_Computer
Z-0308-2012 | 55294 | filter, bacterial, breathing-circuit | CAH | Breathing circuit bacterial filter. | Anesthesiology | Anesthesiology | 510(k) | ViroMax viral/bacterial filter. The filter reduces the risk of infection due to cross-contamination between patientsFor use on all patient populations in conjunction with other respiratory devices containing standard 15mm and/or 22mm fittings (such as breathing circuits) to filter respiratory gases where infection from airborne bacteria and viruses is a concern. Disposable, for single patient use only. | A M Systems ViroMax viral/bacterial filter | 2 | January 11, 2012 | 2012 | A M Systems Inc | Some ViroMax viral/bacterial filters are leaking at the point where the two halves of the filter are ultrasonically bonded together. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A - M Systems sent a Urgent Notification of Product Recall/Correction dated March 31, 2010, to all affected customers. The letter identificed the product the problem and the action needed to be taken by the customer.Customers were instructed to cease use and/or distribution of the affected product and to immediately notify any of their sub-accounts with the same information. Customers are asked to provide A - M Systems with the amount distributed and amount remaining in inventory and to then destroy or discard any remaining product in their possession.Please call 360.683.8300, if you have any questions or concerns. |  | 37800 | 37800 | Worldwide Distribution -- USA ( nationwide) including the states of AZ, CA, FL, KY, MA, MI, NC, OH, OR, and WI., and the countries Canada, Mexico, and Taiwan. | Terminated | January 05, 2012 | -6 | 37800 | Not_Computer
Z-0582-2012 | 48745 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Artificial Ligament Fixation Device 15MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 25MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 20MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 30MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 35MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device #7 PE Ziploop (TM) ToggleLoc (TM).Toggleloc anchors are indicated for use for fixation of tendons & ligaments during orthopedic reconstruction procedures such as Anterior Cruciate Ligament (ACL) reconstruction. | Artificial Ligament Fixation Device 15mm, 20mm, 25mm, 30mm, & 35mm TI NS Toggleloc; & 7 PE Ziploop | 2 | January 11, 2012 | 2012 | Biomet Sports Medicine | Six sizes of ToggleLoc, Artificial Ligament Fixation Device, have been manufactured with the suture pull hole drilled in the wrong location on the device.The ToggleLoc implant is deployed with a suture that pulls the device through the tunnel. The hole being drilled in the wrong location will orient the device to pass through the tunnel with the angled end of the implant first. This may caus | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomet Sports Medicine sent an "URGENT MEDICAL DEVICE REMOVAL" notice dated May 15, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Customers should contact the firm at 800-535-8692 for questions related to this notice. |  | 203 | 203 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, ID, IL, IN, KS, TX, and WA, and the countries of Australia, Chile, France, Poland, The Netherlands and Turkey. | Terminated | December 27, 2011 | -15 | 203 | Not_Computer
Z-0597-2012 | 60456 | sterilizer, steam | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | Midmark M11 Ultra Steam Sterilizer, Distributed by: Midmark Corporation, 60 Vista Drive, Versailles, Ohio 45380.Sterilize reusable equipment that is heat and moisture stable. | Midmark M11 Ultra Steam Sterilizer | 2 | January 11, 2012 | 2012 | Midmark Corportation | During and FDA inspection on 8/23/2011, it was discovered that the firm received numerous reports of the steam sterilizer malfunctioning. | TRAINING: Employee Error | Midmark Corp. will begin sending customer communications out the week of January 9, 2012 to all affected customers. The letter will describe the product, problem, and actions to be taken by the customers. Additionally, a warning label and an illustration of where to place the label will be included with the letter. Customers may contact the firm at 1-800-643-6275. | Contact the recalling firm for information | 30255 | 30,255 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and the countries of AUSTRALIA, BELGIUM, BRAZIL, CHILE, CHINA, EGYPT, FRANCE, GUAM, HONG KONG, HUNGARY, ICELAND, INDIA, KOREA, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, POLAND, RUSSIA, SAIPAN MP, SAUDI ARABIA, SEOUL, SINGAPORE, TAIWAN, THAILAND, TURKEY, U.A.E., UNITED KINGDOM and WEST INDIES. | Open |  |  | 30255 | Not_Computer
Z-0483-2012 | 60659 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Unique Single Energy Linear Accelerator, C-Series Clinac or Trilogy, versions 7.x and 8.x, Reference/FSCA Identifier: CP-0661; Model Numbers: H14, H18, H27, H29; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA Varian Unique Single Energy Linear Accelerator: The UNIQUE is a Single Energy Linear Accelerator intended to be used for conventional radiotherapy and includes modifications to the previously cleared Varian Trilogy. The UNIQUE provides additional features, safety improvements, and usability improvements. The UNIQUE is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Varian Unique Single Energy Linear Accelerator | 2 | January 11, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An event has been reported to Varian which entails excessive connector resistance, which caused the actual jaw positions to differ from the intended jaw positions without warning the operator. | DESIGN: Device Design | Varian Medical Systems sent an Urgent Medical Device Correction - Urgent Field Safety Notice dated November 28, 2011 ***Expanded to include product #2 on December 16, 2011***, to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to establish a comprehensive quality assurance program before beginning treatments for the day, quality assurance tests must be performed. Additionally, any discrepancies in imaging field sizes vs planned field size should be brought to the attention of the medical physicist for investigation before treatment continues. Customers were asked to complete the attached Proof of Notification or Receipt Verification Card once they have read the document and return it to Varian Medical Systems. For question regarding this recall:USA and Canada 1-888-827-4265Europe +41 41 749 8844. |  | 6158 | 6158 units total | Worldwide Distribution | Open |  |  | 6158 | Not_Computer
Z-0689-2012 | 60673 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | MAQUET QUADROX iD PEDIATRIC -OXYGENATOR:BEQ-HMOD 3000; Made In GermanyThe diffusion membrane oxygenator Quadrox iD pediatric is intended for use in an extracorporeal perfusion system. The oxygenator is designed for a blood flow rate of 0.2 -2.8 I/min and is intended for the treatment of pediatric patients. Within the specified flow rate range, the device oxygenates the blood, removes carbon dioxide and regulates the blood temperature.The utilization period for this device is restricted to six hours. | Quadrox iD Pediatric Diffusion Membrane Oxygenator with Bioline Coating | 2 | January 11, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | There are reports of leakage in the Leuer connector of the blood outflow connector of the Quadrox-iD pediatric oygenator. | DESIGN: Device Design | The firm, Maquet Getinge Group, sent an "PRODUCT RECALL URGENT-MEDICAL DEVICE REMOVAL ACTION" letter dated September 22, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed tocheck their remaining inventory; remove product immediately to prevent use; remove the boxes or individually pouched oxygenators in a secure area of their facility and notify their MAQUET representative (a representative will physically remove the recalled product and replenish the removed units); and complete and return the enclosed Product Removal Disposition Form via fax to: (973)-807-9210 or email: whitney.torning@maquet.com.Should you have any questions or concerns, please do not hesitate to contact MAQUET Customer Service at 800-777-4222. |  | 168 | 168 units | Nationwide distribution: AL, CA, DE, DC, FL, GA, HI, IL, IN, LA, MI, MO, MS, NC, NE, NJ, NM, NY, OH, OK, PA, and TX. | Open |  |  | 168 | Not_Computer
Z-0604-2012 | 60479 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 101 Hematocrit/Oxygen Saturation Monitoring System. The CDI 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where continuous monitoring of the blood's oxygen saturation, hematocrit and hemoglobin levels is needed. | CDI 101 Hematocrit/Oxygen Saturation Monitoring System | 2 | January 11, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo is recalling the CDI 101 Hematocrit/Oxygen Saturation Monitoring System because they have not yet submitted a 510(k) to FDA to market the device. |  | 12/7/11 & 12/8/11 All consignees received a "Transition and Notification Guide" dated 12/6/11 and a "URGENT MEDICAL DEVICE REMOVAL" letter dated 12/4/11, which will inform them of the removal of the CD l 101 and provide the user with various options. Terumo will remove CD l 101 units from the US market over the next 18 months. |  | 366 | 366 | Nationwide Distribution - Including the states of: NV, MO, SC, WA, MI, FL, GA, NY, CT, IL, VA, NY, TN, IN, HI, TX, AL, OH, MD, AZ, CA, KS, MD, PA, and NM. | Terminated | August 30, 2013 | 597 | 366 | Not_Computer
Z-0714-2012 | 60679 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | VITROS Chemistry Products Calibrator Kit 16, REF 680 1700, 5x1 mL --- Ortho-Clinical Diagnostics, Rochester, NY USA -- Ortho-Clinical Diagnostics High Wycomber, UK. VITROS Chemistry Products Calibrator Kit 16 is used in conjunction with VITROS Chemistry Products FS Calibrator 1 to calibrate the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the quantitative measurement of rheumatoid factor (RF). | VITROS Chemistry Products Calibrator Kit 16 | 2 | January 11, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling Lot 1630 of VITROSÂż Calibrator Kit 16 because it will not likely meet its specified 12 months of stability. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A letter was sent to customers via overnight delivery on 11/10/2011. The letter informed customers they should discontinue using and discard all remaining VITROS Calibrator Kit 16, Lot 1630; Complete and return the attached Confirmation of Receipt form upon receipt of the notification. Once the fax reply is received, the firm will begin the replacement order process; forward the notification to all departments and areas that use the VITROS Calibrator Kit 16, Lot 1630; and post the notification by each VITROS Chemistry or Integrated System in your facility that uses the VITROS Calibrator Kit 16, Lot 1630 or with VITROS user documentation. |  | 768 | 768 kits | Class II Recall: Worldwide Distribution - USA (Nationwide) and the countries of Brazil, Canada, Chile, China, Colombia, England, France, Mexico, Panama, Singapore, Spain, Venezuela. | Open |  |  | 768 | Not_Computer
Z-0420-2012 | 60096 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Single Use PERPOS PLS System, Catalog #9045-01Catalog#9045-01 for Single Use PERPOS PLS System- a sterile, medical device kit consisting of the components: PERPOS Access Needle, PERPOS K-Wires, PERPOS 2-in-1 Drill, PERPOS Tap, PERPOS Device Driver, 4.5mm BONE-LOK PLS Implant, and PERPOS Pull Pin Remover. The kit is used by medical professions to stablize the spine as an aid to fusion through immobilization of the facet joints. | Single Use PERPOS PLS System | 2 | January 11, 2012 | 2012 | Interventional Spine Inc | The recall was intiated because Interventional Spine has confirmed that the PERPOS 2-in-1 Drill (6070) manufactured by a contract manufacturer for Interventional Spine, Inc., may break during use. |  | The recall communication was initiated on February 22, 2008 with Interventional Spine forwarding an Urgent Recall Notification to all their customers (via email). The notification informed the customers of the problem identified and instructions for the recall.If you have any questions, contact the Vice President, Quality, Regulatory and Clinical Affairs at (949) 525-1493 or email: jmetcalf@i-spineinc.com. | Contact the recalling firm for information | 65 | 65 | Worldwide distribution: USA (nationwide) including states of: CA, FL, MO, NY, OH, PA, TN, WA, and WI; and country of: Europe. | Terminated | February 29, 2012 | 49 | 69 | Not_Computer
Z-0453-2012 | 60481 | system, nucleic acid amplification test, dna, methicillin resistant staphylococcus aureus, direct specimen | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid brand Xpert MRSA/SA-BC US-IVD, an In-vitro diagnostic test kit, Catalog Numbers: GXMRSA/SA-BC-10, Lots Affected: 2803 and 3001; | Cepheid Xpert MRSA/S Blood Culture | 3 | January 11, 2012 | 2012 | Cepheid | Cepheid has received reports from some of its customers regarding higher than desired ( | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 11/17/11, Cepheid issued letters via mail to all its direct consignees, informing them of the affected products and providing instructions on the recall. |  | 1289 | 1289 | Product was released for distribution in US. | Terminated | January 09, 2012 | -2 | 1289 | Not_Computer
Z-0390-2012 | 60407 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Maxcem Elite(TM) self-etch/self-adhesive resin cement, Refill Kit, net wt. 10 g, REF/Part # 33872, CLEAR --- Refill Kit Contents: 2 Syringes (5 g ea), 10 Automix tips (regular) 4:1, 10 Automix tips (wide) 4:1, 10 Intra-oral tipsDental resin cement. Self-etch, self-adhesive resin cement. High bond strength without adhesive needed. Easy to clean up and can be used with all anterior and posterior ceramics, PFMs, and CAD/CAM materials. Color stable and aesthetic. Works with all substrates from opaque to translucent. | Maxcem Elite | 2 | January 11, 2012 | 2012 | Henry Schein, Inc. | The specific lots of Maxcem Elite Refill Kit appear not to be authentic. |  | The firm, Henry Schein, Inc., sent an "URGENT: Product Recall (Dental Practitioners/User Level) letter and response form dated November 10, 2011, via first class mail to its customers. The letter instructed the customers to examine their inventory to verify if they have any of the specified lots on hand. If so, remove it from their shelves. The product may be returned for credit within 30 days, to the following address: Henry Schein, Inc., 41 Weaver Road, Denver, PA 17517. Additionally, the letter instructed the customers to complete, sign, date, and return the enclosed Henry Schein Response Form via fax to: (631) 843-5557 or mail to: Regulatory Affairs Department (E-355) 135 Duryea Road, Melville, NY 11747, Attn: Regulatory-RR. The letter stated "Only returns of the above noted recalled item/lot numbers purchased from Henry Schein, Inc. will be credited to your account. In order to expedite your credit, please include a copy of your response form and invoice, if available, with your return".If you have any questions regarding this matter, please contact Recall Coordinator at 631-390-8000 ext. 8476. | Contact the recalling firm for information | 35 | 18,855 units total: Domestic - 18,820 units; Foreign - 35 units | Worldwide distribution: USA (nationwide) and countries of: Bermuda, Martinique FWI, Nassau Bahamas, St Maarten NA, Australia, and South Africa. | Open |  |  | 18855 | Not_Computer
Z-0629-2012 | 60702 | controller, infusion, intravascular, electronic | LDR |  |  | General Hospital | 510(k) | STAT 2 I.V. Controller, Catalog/REF No. S2, ASSEMBLED IN MEXICO, CONMED CORPORATION, 525 FRENCH ROAD, UTICA, NY 13502 --- This is a disposable device with a sterile, nonpyrogenic fluid path. For controlling the infusion of Intravenous fluids into the human body. | STAT 2 I.V. Controller | 2 | January 11, 2012 | 2012 | ConMed Corporation | ConMed received complaints of some units exhibiting inaccurate or inconsistent flow rates. It was determined that certain units were assembled with an incorrect component which could result in the devices exhibiting a different flow rate and might reduce the ability of the user to regulate the device at low flow rates. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | ConMed Corporation sent Urgent Medical Device Recall Letters and Response Forms dated April 14, 2011, via UPS Priority Mail pouches to all domestic consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The international consignees were sent Urgent Medical Device Recall Letters and Response Forms dated April 22, 2011, via FedEx International Priority. Users were instructed to immediately stop use of the device and return it to ConMed Corporation. --- RECALL EXPANSION: Consignees were also instructed to contact all of those organization within their facility and any other facilities that they may have supplied or given these affected products to. ConMed Corp. sent out Amended Urgent Medical Device Recall Letters and Response Forms dated May 10, 2011, to the 4 consignees via UPS PriorIty Mail for the expansion of the recall for Catalog No. S2-12N, Lot 1006234.For any questions call 315-624-3237 or email ivcontroller@conmed.com. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Canada, Panama and India | Terminated | January 15, 2013 | 370 |  | Not_Computer
Z-0492-2012 | 60412 | catheter, urological | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | McKESSON Medi-Pak Foley Catheter Tray without Catheter For Use With 30 cc Foley Catheter Sterile.Single Use.Disposable.Latex-Free 20 Trays Per Case Reorder No. 37-234 Marketed by McKesson Corporation McKesson Medical-Surgical, Richmond, VA 23228 Made in the U.S.A. Foley Catheter Tray without Catheter | McKesson MediPak | 2 | January 11, 2012 | 2012 | Nurse Assist, Inc | Nurse Assist, Inc. received a recall notice from the Triad Group. This recall has been initiated due to a recall of the PVP Solution Packs manufactured by the Triad Group that are contained in certain procedure trays. | DESIGN: Component Design/Selection | Nurse Assist sent a recall notice dated September 2, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately and quarantine the affected product. Customers were also instructed to identify their customers and notify them at once of the product recall. Customers were instructed to complete and return the enclosed response form as soon as possible. For those distributors that choose to destroy the affected product they must use the attached Notice of Destruction to report the specifics of the destroyed product. and contact Nurse Assist Customer Service for a RGA # prior to destruction. For any questions call Nurse Assist Customer Service at 1-800-649-6800. |  | 4816 | 4,816 cases | Worldwide Distribution - USA (nationwide) and the countries of Canada and Saudi Arabia. | Terminated | October 26, 2012 | 289 | 55118 | Not_Computer
Z-0676-2012 | 60406 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Drill Bit for 5.0mm Recon Screws with Large Quick Coupling. This orthopedic manual surgical instrument is a non powered hand-held device intended for medical purposes to manipulate tissue, or for use with other devices in orthopedic surgery. | Synthes drill bit for 5.0mm Recon Screws/Large QC | 2 | January 11, 2012 | 2012 | Synthes USA (HQ), Inc. | Synthes is initiating a recall of the Drill Bit for 5.0mm Recon Screws with large Quick Coupling due to the potential for the drill stop to slip on the drill bit. | DESIGN: Component Design/Selection | On November 2, 2011 the firm sent "Notice: Medical Device Recall" letters to all consignees, with return receipt request. The letter requested consignees cease use of the device and remove the product from inventory. For questions, call 610-719-5450 or contact your sales consultant. | Discard product or return to place of purchase for a full refund or exchange | 577 | 577 | Product was distributed nationwide to AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, ME, MD, MI, MN, MT, LA, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TX, UT, VA, WA, WV, WI, and WY. GOVERNMENT account is Audie Murphy VA Hospital, Chief Supply Svc Whse #1, 7400 Merton Minter Blvd, San Antonio, TX 78229. | Open |  |  | 577 | Not_Computer
Z-0296-2012 | 60099 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, 200 and 202 VentilatorsThe device is intended to be used in home, institution/hospital, and portable applications such as wheelchairs and gurneys, and may be used for both invasive and non-invasive ventilation. | Respironics Trilogy 100, 200 and 202 Ventilators | 1 | January 11, 2012 | 2012 | Respironics, Inc. | Trilogy 100, 200 and 202 units are being recalled due to low impeller press-force used during the manufacturing process. If the press force were low enough, theoretically the impeller could move up the shaft and come into contact with the inside of the blower housing/impeller enclosure. This may impact therapy. | DESIGN: Process Design | Philips Healthcare, began notifying customers via telephone on October 11, 2011. The telephone notification discussed the product, problem and actions to be taken. The customers were instructed to not remove a ventilator from a patient without providing a replacement unit; provide a fax number or email address Philips can use to send the Business Reply Form to complete and return via fax to: 800-733-9962 or email: report.request@philips.com; locate their affected devices (if the device is with a patient, replace it with a non-affected unit-Philips will provide a loaner overnight if needed)-(if the device is in their inventory, quarantine it and return it using the RA number provided in the email/fax that Philips sends them); and If loaner units are needed, they are to call 1-877-387-3311 to arrange return of your affected device and to obtain a loaner unit.If you have any questions or concerns, contact Philips Respironics at 877-387-3311. |  | 127 | 127 | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, ID, IL, KS, MI, MO, NE, NV, NJ, OH, PA, TN, and TX; and countries of: Australia, Austria, Canada, China, Denmark, Germany, Great Britain, India, Italy, Japan, Korea (Republic of [South] Korea), Netherlands, Philippines, Saudi Arabia, Spain (Espana), and Turkey. | Terminated | November 06, 2012 | 300 | 127 | Not_Computer
Z-0712-2012 | 60894 | latex patient examination glove | LYY | Patient examination glove. | General Hospital | General Hospital | 510(k) | Phantom-brand Black Latex Powder Free Examination Gloves, 100 gloves by weight, size small and Medium, Product Code PHM912, Lot No. 002-E1LF302-1, Manufacture date JUL 2011, Size SmallProduct Code PHM915, Lot No. 002-E1LF305-1, Manufacture date JUL 2011, Size MediumExamination Disposable Gloves | Phantom black latex powderfree gloves | 2 | January 11, 2012 | 2012 | Adenna Inc | The recall was initiated because Adenna Inc. distributed latex gloves that have been FDA refused. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Adenna Inc. sent a recall letter dated December 22, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to examine their inventory and quarantine the products that are subject to the recall. Customers were instructed to notify their customers if they have further distributed the products. Customers were instructed to cease using the recalled products, and Adenna will replace the products. Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with any questions can contact their account manager at (888) 3ADENNA (1-888-323-3662) for futher assistance. |  | 427 | 427 cartons | Nationwide Distribution including the states of: FL, CA, NY, NV, NC, MD, TX, OR, AZ, OH, MN, and IL. | Terminated | May 29, 2012 | 139 | 427 | Not_Computer
Z-0682-2012 | 60703 | controller, infusion, intravascular, electronic | LDR |  |  | General Hospital | 510(k) | STAT 2 PumpetteÂŽ I.V. Compensating Controller, Catalog/REF No. PS2, ASSEMBLED IN MEXICO, CONMED CORPORATION, 525 FRENCH ROAD, UTICA, NY 13502 --- This is a disposable device with a sterile, nonpyrogenic fluid path. Product Usage: The STAT 2 PumpetteÂŽ Extension sets are for use in conjunction with commonly available I.V. administration sets for controlling the infusion of I.V. fluids into the human body.The STAT 2 PumpetteÂŽ Primary Administration Sets are for use in attachment to an I.V. bag and patient catheter inserted into a vein to control the infusion of intravenous fluids into the human body.This is a disposable device with a sterile, nonpyrogenic fluid path. | STAT 2 Pumpette | 2 | January 11, 2012 | 2012 | ConMed Corporation | ConMed received 16 complaints for STAT 2 PumpetteÂż devices for inaccurate flow rates and for leakage from the body of the compensating controller. The investigation of the complaints determined that the STAT 2 PumpetteÂż Compensating Controller on several complaint samples had been assembled incorrectly. | DESIGN: Device Design | ConMed sent an "Urgent Medical Device Recall" letters and Reply Forms dated April 15, 2011 to the domestic consignees via UPS Priority Mail pouches and to the foreign consignees via FedEx International Priority. The letter identified the product, description of problem, and actions to be taken. Consignees were instructed to review their inventory and immediately stop use of the affected devices and return them to ConMed Corporation. Consignees who had distributed the affected products outside of their facility were instructed to immediately notify their customers of the recall and provide them with a copy of this notice. Consignees were asked to complete and return the enclosed Reply Form. For questions call 315-624-3237 or email ivcontroller@conmed.com. |  | 1146741 | 1,146,741 sets in total | Worldwide Distribution - USA (nationwide) - including the countries of: Australia, Barbados, Belgium, Bermuda, Canada, Egypt, Israel, India, Italy, Korea, New Zealand, Philippines and Saudi Arabia. | Terminated | November 20, 2012 | 314 | 1146741 | Not_Computer
Z-0650-2012 | 60225 | suture, absorbable, synthetic, polyglycolic acid | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MANI Needle and Suture Pack, 7-0 PGA Absorbable, Model 2091, 30 cm suture and needle, packaged in individual aluminum pack, within tyvek packaging, 12 packages per box, and labeled in part ***Mani, Inc 8-3 Kiyohara Industrial Park Utsunomiya, Tochigi, Japan***For use in short term soft tissue approximation, including use in ophthalmic surgery, but not for use in cardiovascular and neurological tissue. | MANI Needle and Suture Pack PGA Absorbable | 2 | January 11, 2012 | 2012 | Mani, Inc. - Kiyohara Facility | Surgical sutures may be weak and may not function as intended during surgical procedures. |  | The manufacturer notified consignees by phone, via fax with a letter, and e-mail on 07/20/11. The letter asked customers to check their inventory for any of the affected product, and if they had any product remaining to place the amount in the list "Returned." If there was no product left, then the amount used should be placed in the list "Used." Customers needed to complete and sign the document, being sure to send it and any remaining product to the address provided. Credit would be issued once goods were received. If customers have questions, they should contact the firm. |  | 497 | 497 boxes | Nationwide Distribution -- PA and FL. | Terminated | April 03, 2013 | 448 | 497 | Not_Computer
Z-0518-2012 | 60381 | absorber, carbon-dioxide | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | The IS Can Disposable CO2 absorber x 6 --- Carbon Dioxide absorbent for use on GE rebreathing systems --- Manufactured in the United Kingdom for: Intersurgical Incorporated --- 417 Electronics Parkway, Liverpool, NY 13088.Intended use: GE Healthcare anesthetic machines: S5 Aespire, S5 Avance, Aisys. The Intersurgical IS Can is for use with air, oxygen, nitrous oxide, halothane, enflurane, isoflurane, desflurane and sevoflurane. | IS Can Disposable CO2 absorber | 2 | January 11, 2012 | 2012 | Intersurgical Inc | Intersurgical Incorporated has issued a voluntary recall of the IS Cans Carbon Dioxide Absorber due to a possibility that loose granules may enter the inspiratory path of the breathing system. | DESIGN: Device Design | Intersurgical Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 8, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected product found. A Recall Response Form was attached to the letter for customers to complete and return via fax at 315-451-3696, Attn: Customer Service. Contact Customer Service at 1-800-828-9633 if you have questions concerning this recall. |  | 110 | 110 units | Worldwide Distribution- USA (nationwide) including the states of Florida, North Carolina, New York, Ohio, and the country of Canada. | Terminated | January 15, 2013 | 370 | 110 | Not_Computer
Z-0422-2012 | 60516 | warmer, thermal, infusion fluid | LGZ |  |  | General Hospital | 510(k) | HOTLlNEÂŽ Disposable Administration Sets, Product Codes L-70Product Usage: Single patient use Sets designed for use by trained medical professionals, in conjunction with HOTLlNEÂŽ Routine Flow Blood and Fluid Warmers (HL-90 and HL-290), for the warming and delivery of blood and I.V. solutions, at normothermic flow rates of 50 - 5,000 ml/hr. | Hotline Administration Sets | 2 | January 11, 2012 | 2012 | Smiths Medical ASD, Inc. | Disconnected/ loose Luer connectors at the patient end of the triple lumen tubing of the HOTLlNE Disposable Administration Sets and Level 1 D/DI-60HL Normothermic IV Administration Sets |  | Smiths Medical notified consignees by letter dated November 28, 2011 titled: Urgent Field Safety Notice. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to inspect inventory and segregate unused affected product. Distributors were instructed to immediately stop distributing and quarantine all inventory. The letters included a Confirmation Form to be complete and return, by fax to 781-610-9859 or by email to hotlineconnectorrecall@smiths-medical.com. For questions consignees were instructed to contact Smiths Medical at 1781-763-9330. |  | 449307 | 449,307 sets | Worldwide Distribution  (USA) (nationwide) - Argentina, Austria, Australia, Belgium, Bahrain, Bermuda, Canada, Chile, China, Columbia, Czech Republic, Germany, Denmark, Estonia, Egypt, Finland, France, Great Britain, Guernsey, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel India, Iceland, Italy, Japan, South Korea, Kazakhstan, Lebanon, Luxembourg, Mexico, Malaysia, New Caledonia, Netherlands, Norway, New Zealand, Panama, Poland, Portugal, Qatar, Singapore, Sweden, Slovak Republic, South Africa, Spain, Sri Lanka, Switzerland, Taiwan, and Vatican City State. | Open |  |  | 805506 | Not_Computer
Z-0168-2012 | 60236 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | SafeSheath(R) CSG(R) Coronary Sinus guide Braided Core Hemostatic Introducer, SafeSheath(R) Worley Telescopic Braided Series, and Situs LDS2 Lead Delivery System - Catalog numbers: CSG/WORLEY/BCor-1-09; CSG/WORLEY/BCor-2-09; CSG/WROLEY/L/BCor-1-09; CSG/Worley/R-1-09; CSGWORBC10M; CSGWORBC29M; CSGWORLBC19M; LVI/66-5.5-CR; LVI/66-5.5-RE; LVI/75-5-62-07-HO; LVI/75-5-62-07-HS; LVI/75-5-62-07-MP; LVI/75-5-62-07-RE; LVI/75-5-66-5.5-RE; LVI/Tele/B-60-07-HO; LVI/Tele/B-60-07-HS; LVI/Tele/B-60-07-MP; LVI/Tele/B-60-07-RE; RL065; RL066; RL067; and RL068This device is intended to be used for the introduction of various types of pacing or defibrillator leads and catheters. | SafeSheath(R) CSG(R) Coronary Sinus guide Braided Core Hemostatic Introducer | 2 | January 11, 2012 | 2012 | Thomas Medical Products Inc | Radiopaque (RO)/soft tip of the braided core component in the CSG kits may detach from the braided core tube. | DESIGN: Component Design/Selection | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 7, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately segregate all affected product in a manner that ensures it will not be used; return any product with reference to the return authorization (RA) number or contact customer service representative to facilitate return of product; if any affected product has been forwarded to another facility, contact that facility to arrange return, and complete and return the attached product reconciliation form.If you have any questions, you may contact the Customer Service number at 1-866-446-3003 Monday through Friday from 8:30am to 5:00pm EDT. |  | 33134 | 33,134 | Nationwide distribution: CA only. | Terminated | August 29, 2012 | 231 | 33134 | Not_Computer
Z-0373-2012 | 60214 | light, surgical, fiberoptic | FST | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Luxtec UltraLite Pro Headlight with 9ft bifurcated cableModel: AX2000BIFProduct Usage: The fiber optic headlight system Is designed to deliver illumination from a high intensity light source for surgical site illumination. | Integra Surgical | 2 | January 11, 2012 | 2012 | Integra LifeSciences Corp. | Headlight assembly (module) may become loose and separate from the headband assembly | DESIGN: Device Design | Integra Life Sciences notified consignees by URGENT DEVICE RECALL letter dated October 10, 2011, via FedEx, email or telephone conference calls. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to identify and return any affected product. Customers were requested to complete and return the attached Field Corrective Action Acknowledgement Form to Integra Surgical indicating receipt and review of this notification. For assistance and additional information contact your local and Integra Sales Representative or call Integra Surgical Customer Service at 1-800-431-1123. | Discard product or return to place of purchase for a full refund or exchange | 2181 | 2181 units | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of: Africa, Algeria, Australia, Austria, Brazil, Canada, Check Republic, China, Chile, Costa Rica, Denmark, Dominican Republic, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece Hungary, Guatemala, India, Ireland, Israel, Italy, Iraq, Japan, Jordan, Kenya, Kuwait, Lebanon, Lithuania, Macedonia, Mexico, Netherlands, New Zealand! Norway, Oman, Pakistan, Peru Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Qatar, United Kingdom, Yemeni United Arab Emirates, Uruguay, and Venezuela. | Terminated | October 01, 2012 | 264 | 7216 | Not_Computer
Z-0365-2012 | 60117 | clinical sample concentrator | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 | Abbott m2000sp Instrument | 2 | January 11, 2012 | 2012 | Abbott Molecular | Abbott Molecular has received reports in which the liquid waste sensor and components in the m2000sp lower instrument cabinet overheated and/or caught fire. | DESIGN: Component Design/Selection | Abbott Molecular sent a Field Correction Recall letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until further notice, and to complete the enclosed reply sheet and fax it to (847) 775-6728 at their earliest convenience to acknowledge receipt and understanding of the letter.For questions regarding this recall call 224-361-7133. | Discard product or return to place of purchase for a full refund or exchange | 432 | 432 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Australia, Austria, Azerbaijan, Belgium, Brazil, Burundi, Cameroon, China, Croatia, Estonia, Ethiopia, Finland, France, Gaza & Jericho, Germany, India, Ireland, Israel, Italy, Kazakhstan, Kenya, Libya, Malawi, Malaysia, Mali, Mozambique, Netherlands, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Vietnam and Zambia | Open |  |  | 432 | Not_Computer
Z-0418-2012 | 59739 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FOUNDATION HIP SYSTEM Series 460 Hip Stem STERILE R CEMENTED Sz 16.5mmDJO Surgical 9800 Metric Blvd., Austin, TX 78758The Foundation Non-Porous Stem is intended for treatment of patients who are candidates for total or hemi-hip arthroplasty because the natural femoral head and neck have been affected by osteoarthritis, inflammatory arthritis, traumatic arthritis, rheumatoid arthritis, avascular necrosis, or femoral neck fracture, and revision arthroplasty where bone loss is minimal. These devices are intended to aid the surgeon in relieving the patient of hip pain and restoring hip motion and are intended to be used with bone cement. | Foundation NonPorous Hip System | 2 | January 11, 2012 | 2012 | Encore Medical, Lp | The sterility has the potential to be compromised. | PRODUCTION CONTROLS: Packaging | Encore Medical Lp sent a Urgent Field Safety Notice dated August 17, 2011. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately quarantine these devices if you have not already done so. Contact Customer Service to obtain replacement devices as well as an RMA. Return all affected devices by October 31, 2011, using the RMA Customer Service provided in step #3.If you have any further questions you may call (512 ) 834-6330. For International Customer Service please call (512) 834-6275. |  | 169 | 169 units | Worldwide Distribution -- USA ( nationwide ) including the states of WI, NY, MO, TX, FL, CA, UT, AZ, Puerto Rico, and the country Japan. | Terminated | May 11, 2012 | 121 | 169 | Not_Computer
Z-0452-2012 | 59562 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC and Synchron LX Clinical Systems 100 ÂŽL Sample Syringe, 100UL-PKGD, Part Number 474171Product Usage: The UniCel DxC and LX Synchron Clinical Systems are fully automated, computer-controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | UniCel DxC and Synchron LX Clinical Systems | 2 | January 11, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Modular Chemistry (MC) and Cartridge Chemistry (CC) 100 uL sample syringes may experience premature wear as indicated by brown discoloration on the plunger and/or inside the glass syringe barrel. Excessive wear of the 100 uL sample syringe barrel and plunger could lead to leaking and may affect results. | DESIGN: Component Design/Selection | Beckman Coulter sent an URGENT: Product Corrective Action letter to all affected customers. The letter provided customers with an explanation of the problem identified and actions to be taken. The letter instructs customers to share this information with their laboratory staff and retain a copy as part of their laboratory Quality System documentation. In addition, the letter states that if the affected product was further distributed to provide a copy of the letter to them. Customers were also asked to complete and return the enclosed Response Form. For any questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 3278 | 3278 total (3083 in US, 195 in Canada) | Worldwide Distribution - USA (nationwide) and the country of Canada. | Open |  |  | 3278 | Not_Computer
Z-0706-2012 | 60362 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman Cranial Access Kit with XylocaineProduct Code: 82-6612 Product Usage:The Codman Cranial Access and ICP Kits are convenience kits which provide devices and the drug (XylocaineÂŽ) to allow physicians to access the cranium for various Neurosurgical procedures. | Codman | 2 | January 11, 2012 | 2012 | Codman & Shurtleff, Inc. | Incorrect version of Instructions for Use (IFU) for the Xylocaine product that is supplied with the Cranial kits. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Codman issued an Urgent Medical Device Correction Notice dated October 25, 2011 to all affected customers. The notice identified the product, problem, and actions to be taken. Customers were instructed to notify all appropriate personnel of the correction notice. The notice instructs customers to fax the attached acknowledgement form to 508-977-6403. For questions call your Customer Service or contact your local Codman representative. | Discard product or return to place of purchase for a full refund or exchange | 861 | 861 units | Worldwide Distribution - USA (nationwide) and the countries of: Dubai, Mexico, Panama, Spain and Uruguay. | Open |  |  | 18249 | Not_Computer
Z-0391-2012 | 60371 | stocking, medical support (for general medical purposes) | FQL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | Truform Therapeutic Classic Medical Style Compression Stockings 20-30 mmHG, Truform Therapeutic Classic Medical Style Compression Stockings 30-40 mmHG, A Div of SAI, Cincinnati, OH 45209Product Usage: To help relieve tired, heavy legs, and moderate to pronounced varicose veins and swelling. | Truform Therapeutic Classic Medical Style Compression Stockings | 2 | January 11, 2012 | 2012 | Surgical Appliance Industries | During an audit with the FDA, it was discovered that a new compression testing device used for testing hosiery had been implemented and that the operators were not properly trained. Additionally, the acceptance criteria for the new device were not changed from the old testing device to reflect additional data (compression profile percentages). This resulted in the device history files containing s | TRAINING: Employee Error | The firm, Surgical Appliance Industries sent an URGENT: MEDICAL DEVICE RECALL letters dated November 7, 2011 to all affected customers. The letter described the products, problem and actions to be taken. Customers were instructed to immediately examine their inventory and quarantine all product(s) subject to recall for return and exchange. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Recall Response Form. For any questions contact Quality Systems Manager, at 513-271-4594 ext. 282. | Contact the recalling firm for information | 351896 | 351896 pairs | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Canada, Colombia, Hong Kong, Peru, Singapore, Spain, and Venezuela. | Terminated | March 05, 2013 | 419 | 363715 | Not_Computer
Z-0463-2012 | 60375 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Smith & Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE ResectorCatalog Number: 7205962. | Smith& Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE Resector | 2 | January 11, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | Manufacturing issue: The distal tip of the outer sheath (return electrode) may detach and become loose in the joint during use. | PRODUCTION CONTROLS: Process Control | Smith & Nephew Inc. Endoscopy Division notified facilities on 11/7/11 via an "Urgent - Product Recall 1st Notification - Urgent" letter and/or phone. The letter identified the affected product and the reason for the recall. It also discussed the potential risks, customer actions, and instructions for the return of the product. Customers were to check their inventory and locate any of the affected product. They were to also complete the requested contact information and return the notification letter. In addition, customers were to contact Smith & Nephew Endoscopy Division Returns Group at endo.andreturns@smith-nephew.com for information on how to return the affected product and receive replacement product. Questions, contact Hoangthi Le directly by phone at 508-337-3731. | Discard product or return to place of purchase for a full refund or exchange | 1743 | 1743 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Dubai, Finland, France, Germany, Norway, South Africa, Switzerland, United Kingdom, and Venezuela. | Open |  |  | 1743 | Not_Computer
Z-0681-2012 | 60552 | vinyl patient examination glove | LYZ | Patient examination glove. | General Hospital | General Hospital | 510(k) | Zee Clear Vinyl Medical Examination Gloves, Powder Free, X-Large, Non-Sterile; 100 gloves per box, 10 boxes per case; Distributed by Zee Medical, Inc., Irvine, CA 92606, Made in China; Zee Part No.: 3063Single Use, medical exam glove worn on the hand to provide a barrier against potentially infectious materials and other contaminants | Zee Clear Vinyl Medical Examination Gloves, XLarge | 2 | January 11, 2012 | 2012 | Cypress Medical Products LLC | The gloves fail to meet the performance requirements of 21 CFR 800.20 for leakage defects. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Cypress Medical telephoned the accounts and then sent Urgent Voluntary Medical Device Recall letters dated November 14, 2011, to all affected customers. The letter informed them that the lot of gloves did not meet certain release specifications and that some of the gloves may not provide an effective barrier to infectious material. Customers were requested to cease distribution of the affected product, quarantine the gloves, and contact Cypress Medical's customer service department at 800-334-3646 for a returned goods authorization, return instruction and credit. The accounts were also requested to notify their customers of the recall, and to complete and return the enclosed reply form.For questions regarding this recall call 815-385-0100. |  | 160 | 160 cases | Nationwide Distribution including California, Indiana, New York and Texas | Terminated | October 08, 2012 | 271 | 680 | Not_Computer
Z-0655-2012 | 60435 | support, patient position | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allen C-Flex Head Positioning System non-invasive medical device used to support a patient's head during spine surgeryCatalog Numbers: A-70700 and A-70701-2 | Allen CFlex Head Positioner | 2 | January 11, 2012 | 2012 | Allen Medical Systems, Inc. | C-Flex Polar Head Positioner may have an adjustment screw loosen and affect load capacity | DESIGN: Process Design | Allen Medical Systems sent a Urgent Medical Device Recall Notification dated November 18, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Allen Medical will provide packaging materials and instructions for returning your device for repair at not cost to you. Each unit will be repaired and tested. Devices which pass all testing will be returned to the field. For further questions please call 888-521-4277 ext. 4280. |  | 83 | 83 units | Worldwide Distribution -- USA (nationwide ) and the countries of Africa, Australia, Bulgaria, China, Germany, India, Istanbul, France, Greece, New Zealand, Saudi Arabia,Singapore,Spain, Taiwan and Japan. | Terminated | November 14, 2012 | 308 | 83 | Not_Computer
Z-0386-2012 | 59812 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | ActiveCare +SFTMedical Compression SystemsOR Aqiva, ILIntended for use in: Preventing deep Vein Thrombosis; Enhancing blood circulation; Diminishing post-operative pain and swelling; Reducing wound healing time; treatment and assistance in healing of stasis dermatitis, venous stasis ulcers and diabetic leg ulcers; Treatment of chronic venous insufficiency and; Reducing edema | ActiveCare SFT | 2 | January 11, 2012 | 2012 | Medical Compression Systems | This voluntary field action is being undertaken because MCS has received a limited number of reports involving the use of a non-MCS AC/DC adaptor with the ActiveCare DVT and ActiveCare+S.F.T. devices. The firm found that an ActiveCare+S.F.T device might cease to work if connected to a non-MCS AC/DC adaptor with a higher nominal voltage. As specified in each product's User Manual the ActiveCare+S. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medical Compression Systems (DBN) Ltd, sent out "VOLUNTARY FIELD ACTION" letters dated September 15, 2011 to all affected customers. The letter included description of the device and problem. Customers were instructed to check their inventory and make a list of serial numbers of all applicable ActiveCare+S.F.T. and ActiveCare DVT devices in their inventory, include the list on the attached excel file (Field Action Verification Form) and email to maya@mcsmed.com; even if they do not have any such devices. The letter also states that within 15 days from the email, they will receive instructions for the next steps. In addition, the letter stated -YOU MAY CONTINUE TO USE THE ActiveCare+S.F.T.ÂŽ DEVICES WHILE AWAITINGTHE INSTALLATION OF THE OVP.In case you have questions or you need any assistance, you can contact the firm at (317) 485-7887 or e-mail to: maya@mcsmed.com or mahlet@mcsmed.com. | Contact the recalling firm for information | 551 | 551 units | Nationwide distribution. | Terminated | July 11, 2012 | 182 | 2768 | Not_Computer
Z-0646-2012 | 60644 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Vista(R) V-LYTE(R) Diluent (K835)Product Usage: V-LYTE(R) Diluent is for the automatic dilution of samples processed in the V-LYTE(R) Integrated Multisensor on the Dimension Vista(R) system | Dimension(R) Vista(R) VLYTE(R) | 2 | January 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed that specific lots of V-LYTE(R) Fluids may contribute to reduced V-LYTE(R) Integrated Multisensor on-board use life as demonstrated by increases in: Urine sodium QC results within 48 hours after installation of a new IMT sensor. The results can be elevated approximately 10-20% which may cause QC recovery being out of acceptable ranges. Patient Urine sodium results may also b | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Siemens sent an Urgent Field Safety Notice letter dated October, 2011 to all consignees. The letter identified the product, description of problem, and actions to be taken. Consignees were instructed to replace the affected lots of V-LYTE(R) Fluids with alternate lots if available and recalibrate. If no alternate lots were available they were recommended to replace the V-LYTE(R) Integrated Multisensor every 24 hours or when indicated by QC until the impacted lots were no longer in use. For replacement product consignees were instructed to contact the Siemens Customer Service at 888-588-3916 or your local Siemens Service Center. Consignees were advised to complete and return the enclosed Customer Response Sheet via fax to the Technical Solutions Center at 302-631-8467. For questions or concerns contact Siemens Technical Solutions Center at 800-441-9250 for further assistance. |  | 5282 | 5,282 | Worldwide Distribution - USA (nationwide) iin the states of : AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MS, MO, MT, NE, NH, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, VT, UT, VA, WA, WV, WI, WY including Puerto Rico and Washington, DC and the countries of: Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Italy, Netherlands, Norway, Portugal, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, Canada, Australia, Japan, New Zealand, Malaysia, and South Korea. | Terminated | December 24, 2013 | 713 | 19524 | Not_Computer
Z-0677-2012 | 60619 | hook, surgical, general & plastic surgery | GDG | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Weck, DuraHook ÂŽ (6mm) and ÂŽ (12mm), Neuro Elastic Retractors, 6 per Package, 10 Packages per Sales Units, 10 Sales Units per Shipper, Rx Only, Teleflex Medical, Research Triangle Park, NC 277809.For use as skin retraction devices during general surgery. DuraHooks are also cleared for use in procedures requiring retraction of dura mater. | DuraHook Neuro Elastic Retractors | 2 | January 11, 2012 | 2012 | Teleflex Medical | Teleflex Medical received complaints that the elastic bands are breaking, cracking and /or deteriorating prior to their expiration date. If the elastic bands break during use there is a chance for operative exposure, loss of the band in the operative field as well as inadvertent needle stick. |  | The firm, Teleflex Medical, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated December 6, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately discontinue distribution and quarantine any products with the listed catalog and lot numbers; communicate this recall to any of their customers who have received the product; have their customers return the affected product together with a completed Recall Acknowledgement Form to Teleflex Medical, if an alternative approach is needed, contact Teleflex Medical Customer Service for more information at 1-866-246-6990; complete and return the Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customers Service even if they and their customers have no affected stock.If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  |  | to be obtained | Nationwide distribution. | Open |  |  |  | Not_Computer
Z-0513-2012 | 60643 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Healthcare Diagnostics, Dimension Vista(R) Cardiac Troponin I Calibrator (KC678) lot 1DD085, in vitro diagnostic | Dimension Vista(R) Cardiac Troponin I Calibrator | 2 | January 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed positive bias on QC and Patients. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens healthcare Diagnostics initiated their recall of lot 1DD085 on November 17, 2011 by sending an Urgent Field Safety Notice letter to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the CTNI CAL lot 1DD085. |  | 2281 | 2281 cartons | Worldwide distribution, including USA, Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, New Zealand, Portugal, Puerto Rico, Slovenia, South Korea, Spain and Switzerland. | Open |  |  | 2281 | Not_Computer
Z-0482-2012 | 60608 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Y-Type Blood/Solution Set with 40 and 150 Micron Dual Screen Filter, Length 95" (2.4 m), Injection Site, Male Luer Lock Adapter; an Rx sterile, nonpyrogenic fluid pathway for the removal of microaggregates from whole blood or red blood cells and transfusion of platelets; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 4C6838Intended use: for the removal of microaggregates from whole blood or red blood cells and transfusion of platelets | Baxter Interlink System YType Blood/Solution Set | 2 | January 11, 2012 | 2012 | Baxter Healthcare Corp. | During production of the IV sets, solvent was inadvertently applied to the locking collar of the male luer connector, which may have damaged the component, preventing proper engagement of the locking collar to secure a connection. | TRAINING: Employee Error | The firm, Baxter, sent an "Urgent Product Recall" letter dated December 6, 2011, via first class mail to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to inspect all locations where their facility's inventory is located for products with the affected lot numbers; return the affected product; request a credit or replacement by calling Baxter Healthcare Center for Service at 1-888-229-0001 between 7:00am to 6:00pm Central time, Monday through Friday; complete and return the attached Customer Reply Form via fax to: 1-847-270-5457, or scan and email to fca@baxter.com; and if they are a dealer, wholesaler, or distributor/reseller that distributed any affected product to other facilities, please notify their customers of this recall.If you have any questions regarding the communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 1440 | 1,440 units | Worldwide distribution: USA (nationwide) including states of: Alaska, California, Colorado, Florida, Georgia, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Maine, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, Rhode Island, South Carolina, Tennessee, Virginia, Washington, Wisconsin and West Virginia; and country of: New Zealand. | Terminated | November 01, 2012 | 295 | 48480 | Not_Computer
Z-0465-2012 | 60568 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | CD HORIZON (R) Spinal System, MULTI-AXIAL SCREW FOR 5.5mm ROD W/REDUCTION HEAD, Size 6.5mm X 40mm, Mat'l TITANIUM ALLOY, REF X1010484, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132implantation | MULTIAXIAL SCREW | 2 | January 11, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | One screw may be longer then size etched on the screw. | PRODUCTION CONTROLS: Error in Labeling | Medtronic Sofamor Danek USA, Inc sent a Voluntary Field Correction letter dated November 30, 2011, via FedEx to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The firm will visit each consignee to perform a visual check of inventory.Please utilize the attached questionnare to account for the affected product and to record any patient injuries that may have occurred specifically related to the use of these items. If you need assistance, or if you have any related questions or concerns please call (800) 876-3133 extension 3173. |  | 30 | 30 units | Nationwide Distribution including the states of AL, PA, IN, CO, MI and IA. | Terminated | April 05, 2012 | 85 | 282 | Not_Computer
Z-0643-2012 | 60633 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Neulook Deluxe is packaged in a pink soft plastic box with white, black, and rainbow printing that reads,NeuLook COLOR CONTACT LENSES Give yourself a whole new look***LENS CASE INCLUDED HOT RED PWR. 0.00 LOT. BD53***EXP. 2014/08***Contents : Two sterile color contact lenses(62% Polymacon, 38% Water) Immersed in saline and one lens case. Caution : To ensure that your contacts stay comfortable and your Vision stay clear, ask your eye-care professional for product instruction. Made in Korea www.NeuLooklens.com Distributed by EYEZONE***1023 FDA 510K CLEARED***STERILE***B.C 8.6mm DIA. 14.0mm***8 806048 000468.The color contact lens colors are: Black out, vein, night, white cat, red cat devil, blue angelic, white out, and hot red.prescription contact lens | Neulook | 2 | January 11, 2012 | 2012 | Eyezone | This recall has been initiated because Pishon Trading dba Eyezone have distributed soft color lens while on FDA hold and were later found to be misbranded. | MISBRANDING: Labeling False and Misleading | Pishon Trading sent an Urgent Medical Device Recall letter dated November 9, 2011, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to examine their inventory and quarantine the recalled product, and they were instructed to return the recalled products to Pishon Trading. Customers were instructed to complete and return the enclosed response form as soon as possible. If customers had any questions, they can call (562) 906-1100.On December 16, 2011, Pishon Trading expanded their recall by sending a new recall letter to all their customers who purchased the Neulook Deluxe and Neulook Fancy contact lens to include all lots. |  | 24628 | 24,628 total pairs | Nationwide Distribution including LA, MS, NV, CA, TX, TN, AR, HI, IL, GA, and FL.. | Open |  |  | 24628 | Not_Computer
Z-0648-2012 | 60775 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Exodus(TM) 10 F (3.4 mm)/ 35 cm Standard Loop Biliary Drainage Catheter with Hydrophilic Coating, REF/Catalog No. 10013, UPN/Product No. H965100131, STERILE --- Package/Box Units: Box of 5 eaches --- Made in USA: 2301 Centennial Blvd., Jeffersonville, IN 47130 -- Manufactured for: Navilyst Medical, Inc. 26 Forest Street, Marlborough, MA 01752 USA --- The Exodus Drainage Catheter consists of a radiopaque polyurethane catheter. The distal end of the catheter is configured as a J-Tip or a locking pigtail and is coated with a hydrophilic coating. Catheter markings are provided in centimeters on the shaft to indicate the distance from the distal tip. The Exodus Biliary Drainage Catheter includes a radiopaque (RO) marker band located approximate 5 mm from the most proximal drainage hole to assist with catheter placement.Biliary Drainage Catheters are intended for percutaneous drainage of the biliary tree. | Exodus(TM) Standard Loop Biliary Drainage Catheter with Hydrophilic Coating | 2 | January 11, 2012 | 2012 | Navilyst Medical, Inc | Navilyst Medical, Inc. determined that the Exodus Standard Loop Biliary Drainage Catheters exhibited a rate of reported failures (catheter tip fragmentation) higher than anticipated. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Navilyst Medical, sent an "Urgent Medical Device Recall -Immediate Action Required" letter dated February 4, 2011, via Federal Express to the domestic consignees/customers and one foreign consignee/customer. Consignees/customers were instructed to: (1.) Immediately discontinue use and remove the recalled product from inventory; (2.) Segregate this product in a secure location for return to Navilyst Medical, Inc.; (3.) Immediately forward a copy of the recall notification to ALL sites to which they have distributed affected device (4.) promptly complete and return the Reply Verification Tracking Form (even if you do not have any product to return) via fax to: QA Manager, Navilyst Medical, Inc., 1-800-782-1357; (5.) call Customer Service at 1-800-833-9973 between 8:30.am to 7:00pm EST to obtain RMA number; (6.) package and return the recalled product; and (7.) Physicians were instructed to continue to monitor their patients and evaluate per institutional protocol for biliary catheters. If you have any questions about this recall action, your local Sales Representative is available to assist you or contact QA Manager, Navilyst Medical, Inc., directly at 518-742-4571. | Contact the recalling firm for information |  |  | Worldwide distribution: USA (nationwide) including states of: AK, AZ, CA, CO, FL, IL, KS, MI, MO, NC, NJ, NY, OH, OK, PA, PR and country of: South Africa. | Terminated | January 10, 2012 | -1 |  | Not_Computer
Z-0460-2012 | 60051 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | SRS FAST SET PUTTY 3 CC BONE VOID FILLER-STERILE, Part Number SRS-005-FSP, Lot number N001101.Bone void filler. | SRS FAST SET PUTTY 3 CC BONE VOID FILLERSTERILE | 2 | January 11, 2012 | 2012 | Synthes USA | A recent shipment carrying Norian SRS was lost by the carrier. Since the product was lost before final product release testing, Synthes is unable to confirm that Part Number SRS-005-FSP- Lot Number N001101 meets final product release requirements. The product contains all final labeling and cannot be visually distinguished from the product that has met final product release requirements. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | All Synthes accounts with a history of purchasing the product in the past two years will be notified by letter of the incident. All firm sales consultants will also receive written notification of the incident. | Contact the recalling firm for information | 131 | 131 units | Synthes USA sent a "NOTICE: MISSING SHIPMENT-NORIAN SRS" letter dated September 29, 2011 to all accounts with a history of purchasing the product in the past two years. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory to verify they have not received the affected product. Call 1-800-620-7025 ext. 5450 for questions concerning this notice. | Open |  |  | 131 | Not_Computer
Z-0641-2012 | 60462 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | LigaSure V 5 mm Dolphin Tip Open 20 cm Sealer/Divider, LS1520, Sterile EO, Manufactured for Valleylab a division of Tyco Healthcare Group LP, Boulder, CO 80301, Manufactured by Covidien Medical Products (Shanghai) Manufacturing LLC, Building No. 10, 789 Puxing Rd., Shanghai 201114, P.R. China.Hand activated sealer/dividers that can be used to seal vessels up to and including 7 mm. lymphatics, and tissue bundles. | LigaSure V 5 mm Dolphin Tip Open 20 cm Sealer/Divider | 2 | January 11, 2012 | 2012 | Covidien, PLC | Retaining pins in handle assembly may disengage and fall into surgical field or compromise handle unlatching. |  | Covidien sent an Urgent Medical Device Recall letter dated November 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify and quarantine all affected inventory, Complete the attached Recalled Product Return Form in its entirety. Fax the completed form to (203) 822-6009 and include a copy of the completed form with the products being returned. If customers did not have any units from the affected lots in their inventory, simply return the recall form, via fax or email to:SDFeedback@Covidien.com, indicating you have zero (0) units. If the affected product was purchased directly from Covidien, return product as follows: Customers should contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722-8772, option 1, to obtain a Return Goods Authorization (RGA) prior to returning the affected units. Covidien will arrange for shipment pickup and will cover all freight costs associated with the return. Ship affected products with the RGA # provided by Customer Service to:Covidien, Surgical SolutionsBuilding 4/Receiving5920 Longbow DriveBoulder, CO 80301If the affected product was purchased from a distributor, customers should contact their distributor to obtain instructions for return of their affected product. Customers will receive a credit for all products returned. For questions call 303-530-6227. |  | 20190 | 20,190 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Guatamala, Hungary, Iceland, Iran, Ireland, Israel, Italy, Japan, Jordan, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Netherlands, Norway, Oman, Panama, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Tunisia, Turkey, UAE, UK, and Vietnam. | Terminated | March 22, 2013 | 436 | 120810 | Not_Computer
Z-0421-2012 | 60454 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | MAXLOCK EXTREME EXTREMITY PLATING SYSTEM WITH VARIABLE ANGLE TECHNOLOGY. Product Usage: Variable Angle technology in MTP, TMT, MFT, MXL, DFX, and EDL plates are indicated for fractures, fusions and osteotomies for small bones in the hand, wrist, foot and ankle in both pediatric and adult patients. Variable Angle Technology in MXL and CLA plates are indicated for the fractures, fusions, and osteotomies of the clavicle and small bones in the hand, wrist, foot and ankle. | MAXLOCK EXTREME EXTREMITY PLATING SYSTEM | 2 | January 11, 2012 | 2012 | Orthohelix Surgical Designs Inc | The firm discovered a testing protocol error, in both dynamic and static cantilever comparative testing, during the initial comparative testing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Orthohelix Surgical Designs Inc. sent an URGENT RECALL NOTIFICATION letter dated August 22, 2011 to all affected customers. The letter identified the affected products, the problem, and the actions to be taken. Customers were instructed to immediately inventory, remove and return the recalled product to the firm via FedEx or any shipping method with tracking capability. Customers were also requested to notify any affected user facilities of the recall. For any questions call 330-247-1444 or 330-247-1445. |  | 2722 | 2,722 units | Nationwide Distribution (USA) - including the states of: CA, CO, CT, FL, GA, IA, IL, KY, LA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, WA, WI and WY. | Terminated | February 19, 2012 | 39 | 2722 | Not_Computer
Z-0485-2012 | 60528 | alignment laser product | RER |  |  |  |  | Raymond Laser alignment guide - Model #1096597/002 Red Laser.Intended to assist operators in performing material handling activities such as storing and retrieving pallets and material in and on racking projecting a line onto a pallet or other material storage medium for assisting the operator in determining the position of the powered industrial truck forks with respect to load. | Raymond Corporation ForkTip Laser Guide | 2 | January 11, 2012 | 2012 | The Raymond Corporation | Raymond shipped a laser alignment guide as an option on its powered industrial trucks from February 2, 2010 through April 21, 2011, but was not aware of its obligation under 21 C.F.R. 1010.2 and 1010.3 to certify and label. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Raymond will be initiating a Field Campaign to ensure that all customers with Alignment Guides have the confirmatory laser caution decals and manufacturer label, an updated installation and service manual, the operator instructions, and informational brochure. This campaign will be performed at no cost to Raymond's dealers and customers, and will include documenting the laser serial number for each truck configured with an Alignment Guide for product traceability. | Contact the recalling firm for information | 493 | 493 total | Nationwide Distribution | Open |  |  | 493 | Not_Computer
Z-0521-2012 | 60519 | enzyme immunoassay, amphetamine | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Panel Plus TCA 25 Test Kit, PN (Model Number) 92000An immunoassay used for the qualitative determination of the presence of the major metabolites of drugs of abuse and Tricyclic Antidepressants in urine. | Triage Drugs of Abuse Panel Plus TCA | 2 | January 11, 2012 | 2012 | Alere San Diego | The recall was initiated because Alere San Diego has confirmed false positive THC results obtained on Triage Drugs of Abuse Plus TCA 25 devices from lots 235991 and 238928. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A recall communication was initiated on July 1, 2011 with Alere San Diego forwarding a Product Notification and a Verification Form (via fax, e-mail or direct mail) to all their customers who purchased the Triage Drugs of Abuse Plus TCA 25 Test Kit, PN 9200, Lots 235991 and 238928. The letter provides the customers with an explanation of the problem identified and an action to be taken. The firm reminds customers that clinical consideration and professional judgment must be applied to any test result. Customers were instructed to retain the Product Notification letter for their records as documentation that they have received the information and to ensure that all users have a copy. Alere will replace the lots at the customer's request. As the customers complete the fax back form, they were instructed to indicate whether they wish to replace the affected product. Customers were instructed to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice. If customers have any further questions or need to report any concerns they have about the Triage Drugs of Abuse Plus TCA performance, they were instructed to contact Technical Services at (877) 308-8287 or by e-mail at techservice@alere.com. | Contact the recalling firm for information | 997 | 997 total kits | Worldwide Distribution -- US and to Canada. | Terminated | April 04, 2012 | 84 | 997 | Not_Computer
Z-0517-2012 | 60583 | stylet, tracheal tube | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Rusch Flexi-slip Endotracheal Tube Stylet with Soft Distal Tip, 6FR, Catalog numbers: 502501, Teleflex Medical Sdn Bhd, P.O. Box, Industrial Estate, 24600 Kamunting, Perak, Malaysia.Used temporarily to make rigid a flexible tracheal tube to aid its insertion into a patient. | Rusch | 2 | January 11, 2012 | 2012 | Teleflex Medical | Complaints of difficulty removing stylet from the endotracheal tube and possible PVC sheath separation were received. |  | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated December 2, 2011 to its Consignees/Customers. The letter described the product, problem and actions to be take. The customers were instructed to immediately discontinue use and quarantine any products with the listed catalog and lot numbers; return the product; complete and return the enclosed Immediate Attention Requested Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service (a customer service representative will contact the customer with a returns good authorization number and provide instructions for the return of product to Teleflex medical), even if they have no affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. | Contact the recalling firm for information | 73281 | 73,281 units | Worldwide distribution: USA (nationwide) and countries of: Canada and Europe. | Open |  |  | 73281 | Not_Computer
Z-0511-2012 | 60476 | automated urinalysis system | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Urine Reagent Strips (URS) 10 ParametersURS-l0-100W510(K) number: K970250Devise Listing number: D072377Each device consists of 100 strips of Urine Reagent Strips 10 parameterspacked with a desiccant pouch inside a gray bottle with Teco DiagnosticsLabel on the bottle and the bottle placed inside a Teco Diagnostics Box.For the semi-quantitative and qualitative detection of Glucose, Bilirubin, Ketone, Specific Gravity, Blood, pH, Protein, Urobilinogen, Nitrite, Leukocytes in urine. | Urine Reagent Strips (URS) 10 Parameters | 2 | January 11, 2012 | 2012 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134, Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Teco Diagnostics sent an "URGENT: MEDICAL DEVICE RECALL" letter dated October 14, 2011 to affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed customers to examine their inventory for the affected product and to quarantine any product found. A Recall Return Response Form was included with the letter for customers to complete and return. Contact the firm's Technical Support Department at (800) 222-9880 for questions regarding this recall. |  | 550 | 550 kits | Distributed in CA and Poland. | Terminated | April 09, 2013 | 454 | 550 | Not_Computer
Z-0697-2012 | 60144 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Series A Standard Patella size 28mm, one peg - 184702, Biomet Orthopedics, Warsaw, IN 46582 For use with Vanguard, Maxim, Ascent & ACG Knee system. Knee joint patellofemorotibial, cemented, polymer/metal/polymer, semi-constrained. Knee joint replacement prosthesis, intended for patella replacement of the knee joint in conjunction with a femoral and tibial component. | Biomet Series A Standard Patella size 28mm, one peg 184702 | 2 | January 11, 2012 | 2012 | Biomet, Inc. | Correcting language on product labeling which incorrectly states patella buttons are intended for use with the AGC knee system. | MISBRANDING: Labeling False and Misleading | Biomet sent a "FIELD CORRECTION NOTICE" dated October 3, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to locate the affected product and to follow the work instructions included in the letter. Additionally, a Confirmation of Completion Response Form was enclosed for customers to complete and return via fax to 574-372-1683. Questions related to this notice are directed to 574-371-3755. |  |  |  | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Argentina, Australia, Canada, Chile, Costa Rica, Europe, India, Japan, Korea, and Singapore. | Terminated | April 08, 2013 | 453 |  | Not_Computer
Z-0415-2012 | 60263 | filler, bone void, osteoinduction (w/o human growth factor) | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM Flex; Unit Size 5x5 cm; 1eaContainer: Foil Pouch with Tyvek pouchFracture repair, bridge bone gaps or fragments, use with strut grafts | Grafton DBM Flex | 2 | January 11, 2012 | 2012 | Osteotech Inc | Aseptic conditions during the production of these products may have been compromised. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Osteotech sent a recall letter dated August 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all units of the affected product. Customers were asked to contact Osteotech Customer Service at 732-578-6699 within two business days of receiving the letter to acknowledge receipt of the letter and to arrange for a return and exchange or credit for the product.For questions regarding this recall call 732-542-2800. |  | 15 | 15 units | Worldwide Distribution | Terminated | December 08, 2011 | -34 | 33 | Not_Computer
Z-0510-2012 | 60148 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | AVEA Ventilator (Catalog Code 17310, 17311, 17312).The AVEA is intended to provide continuous respiratory support in an institutional health care environment (e.g. hospitals). It may be used on neonatal through adult patients. It should only be operated by properly trained clinical personnel, under the direction of a physician. Prescription Use: Federal law restricts the sale of this device except by or on order of a physician. | AVEA Ventilator | 1 | January 11, 2012 | 2012 | Carefusion 211 Inc | The recall was initiated because Carefusion has identified a potential risk associated with certain AVEA ventilators and affected replacement parts manufactured between March 1, 2009 and June 30, 2011. Carefusion is voluntarily initiating a field correction of the affected devices to preclude the possibility of this risk. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CareFusion sent an "URGENT PRODUCT RECALL" letter dated September 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers were instructed to contact the firm by telephone to coordinate implementation of the corrective action at their site. Customers were instructed to return an enclosed response card to expedite the correction process and acknowledge receipt of the recall notification. Carefusion does not require that the customers return their devices. Customers will be contacted by a member of the CareFusion Technical Support Department to arrange for onsite remediation of the affected devices. In the interim, if any affected AVEA ventilator units in their facility exhibits a sustained Extended High Ppeak alarm followed by the opening of the Safety Valve, customers were instructed to remove the ventilator from service, provide alternate ventilation and contact Carefusion Technical Support (800) 213-2466 to report the issue. All ventilator dependent patients should be constantly monitored by qualified personnel to ensure that if a malfunction were to occur, alternate ventilation can be provided. |  | 6797 | 6797 total, 1891 in US, 4906 worldwide | Worldwide Distribution-USA (nationwide) and the countries of Cyprus, Estonia, Brazil, Kuwait, United Arab Emirates, Slovenia, Lithuania, Latvia, Belgium, Bermuda, Ecuador, Guatemala, Italy, Canada, Australia, Netherlands, Germany, United Kingdom, Spain, Czech (Rep), Nicaragua, China, Bangladesh, Morocco, Bolivia, Vietnam, Slovakia, Paraguay, Argentina, Turkey, Chile, Singapore, Japan, Colombia, Egypt, Hong Kong, Honduras, India, Croatia, Venezuela, Saudi Arabia, Greece, Hungary, Mexico, Austria, Ireland,Indonesia, Dominican Republic, Portugal, South Africa, Panama, Philippines, Libyan Jamahiriya, Jordan, Malaysia, France, Thailand, Peru, Costa Rica, Korea, and Taiwan. | Terminated | October 29, 2013 | 657 | 6797 | Not_Computer
Z-0593-2012 | 60524 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | IV Start KitsConvenience kit for starting IVs. Triad wipes are packaged within the sterile barrier and would require compromising the barrier to replace the wipe. | IV Start Kits | 2 | January 11, 2012 | 2012 | ICU Medical, Inc. | The recall was initiated because ICU Medical Inc. has become aware of the market recall of the Triad Group's alcohol prep pads, alcohol swabs, and alcohol swabsticks manufactured by Triad in the United States and marketed under various brand names. The Triad alcohol prep pads have the potential contamination of the bacteria, Bacillus cereus. The Triad Group alcohol prep pads have been copackaged a |  | The firm, ICU Medical, Inc., sent an "IMPORTANT-Important Information About Triad Group's Alcohol Prep Pads" letter dated March 2, 2011 (via email) to their distributors. Hard copy notices were shipped (via next day express delivery with receipt confirmation) to all their customers. The Important letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were informed that the Orbit and Orbit Micro subcutaneous infusion sets as well as the IV start kits are not contaminated and may continue to be used according to the directions for use. The issue is confined to the actual Triad alcohol prep pad products. Patients and healthcare providers should not use the alcohol prep pads packaged with these devices and should instead use an alternate alcohol prep pad that is not involved with the Triad Group recall. ICU Medical has halted all shipments of product that contain the Triad alcohol prep products to its distribution network. An alternate wipe has been qualified and is now being shipped in new orders. Customers were instructed to forward the Important letter to their customers and on to any end users of the devices. For Orbit and Orbit Micro devices, patients are instructed to immediately discontinue using the Triad alcohol prep pads included in the Orbit packaging and dispose of those pads in the trash. Patients should use an alternative prep pad that is not subject to the Triad recall. For IV Start Kits, distributors and users are instructed to place the enclosed label on the device prior to use. Additional labels can be obtained by contacting ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Standard time: (800) 866-9025 and select option 8, or email the following address: customerservice@icumed.com. Customers may also contact regulatory@icumed.com for MSWord template file for Avery 5160 labels and print additional labels. Customers were instructed to complete. |  | 581700 | 581,700 | Worldwide distribution: USA (Nationwide) and countries of: Austria, Belgium, Chile, Cyprus, CZ, France, Germany, Great Britain, Greece, Israel, Italy, Lithuania, Netherland, Poland, Russia, Saudi Arabia, Sweden, andSwitzerland. | Terminated | February 22, 2013 | 408 | 957990 | Not_Computer
Z-0639-2012 | 60582 | wire, orthodontic | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | TMA Archwires, Part Number 211-1402, Lot Number 10K291, Product Description and Intended Use: The intended use of this device is an archwire during orthodontic treatment. | TMA Archwires | 2 | January 11, 2012 | 2012 | Ormco Corporation | The TMA Lower Broad Arch packages labeled with part number 211-1402 actually contained TMA Upper Archwires, part number 2021007. | PRODUCTION CONTROLS: Packaging | Ormco Corporation sent a Urgent Medical Device Recall letter dated April 29, 2011, to the customers in the United States and Canada via USPS First Class mail. The recall letter for customers in Europe are being translated. Ormco Corporation sent the recall letter dated May 20, 2011, to the customers in Austrailia. The letter identified the product the problem and the action needed to be taken by the customer. As a result of this situation the customer is instructed to return the affected lot number of product. This will be replaced at no charge or credited to account.Please label your returned product " RECALLED PRODUCT" attention Customer Return. In addition, we ask for your assistance in completing the RETURN FORM and returning it by fax 909-962-5605. | Contact the recalling firm for information | 23 | 23 Packs | Worldwide Distribution -- USA including the states of Indiana and Minnesota; and the countries of Canada, Europe and Australia. | Terminated | March 08, 2012 | 57 | 23 | Not_Computer
Z-0607-2012 | 60660 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Imaging Systems, Model 882050;Intended to produce images depicting the anatomical distributions of single photon and positron emitting radioisotopes within the human body for interpretation by medical personnel. | SKYLight Imaging Systems | 2 | January 11, 2012 | 2012 | Philips Medical Systems | During preventative maintenance, a crack was discovered in the lower portion of the old style, non-common arm for the detector assembly, which could result in the detector dropping in an uncontrolled manner. | DESIGN: Component Design/Selection | Urgent Medical Device Correction/Field Safety Notices, dated Dec 08, 2011, were mailed to affected customer sites on December 9th and 12th, 2011 by Return Receipt certified mail. The letter identified the affected product, described the issue and when it could occur, along with the hazards involved. Customers were advised to stop using the system and contact a Philips service representative for an immediate inspection. Customers were also to follow the Instructions For Use to avoid the issue. Philips service will contact customers to arrange a time for Philips to inspect the system. The letter should be placed in the Instructions for Use. The Field Change Order will be released on December 16, 2011. If further information or support is needed, customers should contact Customer Care Solutions Center at 1-800-722-9377. |  | 129 | 129 | Worldwide Distribution -- USA, Canada, Italy, Spain, France, United Kingdom, Sweden, Switzerland, Germany, Japan, Denmark, and the Netherlands. | Terminated | August 13, 2012 | 215 | 129 | Not_Computer
Z-0628-2012 | 60049 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OPUS Minimagnum Implant Part Number OM-2500 is a bone anchorIntended for rotator cuff repair | OPUS Minimagnum Implant Part Number OM2500 | 2 | January 11, 2012 | 2012 | Innovatech Medical Resources L.P. | Product is being removed from the field due to an over-label issue. | EXPIRATION DATING: Incorrect or No Expiration Date | Innovatech Medical Resources L.P. decided to recall all products resterilized and overlabeled at their facility. A firm representative would visit each customer and request that they remove the overlabeled product from their shelves and secure it. Customers could call the MSI office to get a RMA (Return Materials Authorization) for a refund on this product.For questions regarding this recall call 817-714-9066. |  | 110730 | Lot Number (QTY): 110730 (7 units), 110874 (2 units), 110400 (7 units), 119256 (1 unit) and 117489 (1 unit) | Nationwide Distribution including Texas | Terminated | September 27, 2012 | 260 | 106214186 | Not_Computer
Z-0467-2012 | 60590 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | MORPHEUS CT PICC with STYLET, 5F Dual Lumen 65 cm Insertion Kit, Catalog No./REF 1200923, STERILE, AngioDynamics Queensbury, NY 12804 USA --- COMMON/USUAL NAME: Percutaneous, implanted, long-term intravascular catheter; CLASSIFICATION NAME: Percutaneous, implanted, long-term intravascular catheter; Device Listing # D095274.PICC Catheters are utilized for Short or Long Term peripheral access to the Central Venous system for intravenous therapy and power injections of contrast media. | MORPHEUS PICC | 3 | January 11, 2012 | 2012 | Angiodynamics Worldwide Headquarters | AngioDynamics distributed the affected PICC catheters that were beyond their expiration date. | EXPIRATION DATING: Incorrect or No Expiration Date | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated December 2, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. A Recall Reply Form was attached to the letter for customers to complete and return to the firm via fax at 518-798-1360. Contact your local sales representative or AngioDynamics Customer Service at 800-772-6446 for questions regarding this notice. |  | 16 | 16 units domestically | Nationwide Distribution-including the states of Florida, Georgia, Montana, and Tennessee. | Terminated | September 27, 2012 | 260 | 16 | Not_Computer
Z-0434-2012 | 60138 | prosthesis, knee, femorotibial, unicompartmental, semi-constrained, metal/polymer, mobile bearing | NRA |  |  | Orthopedic | PMA | "***BUECHEL-PAPPAS Conforming Bearing***UHMWPE***SIZE***CAT**STERILE***ENDOTEC***Endotec, Inc. 300 Sunport Lane, Suite 500 Orlando, FL 32809 USA***"SIZE 1 x 7.5mm CAT 04-33-0031; SIZE 2 x 7.5mm CAT 04-33-0032; SIZE 3 x 7.5mm CAT 04-33-0033; SIZE 4 x 7.5mm CAT 04-33-0034; SIZE 5 x 7.5mm CAT 04-33-0035; SIZE 6 x 7.5mm CAT 04-33-0036. Tricompartmental Knee Replacement | Beuchal Pappas Mobile Bearing Knee System | 2 | January 11, 2012 | 2012 | Endotec, Inc. | On 09/21/2011 Endotec, Orlando, FL initiated a recall for the Beuchel-Pappas Mobile Bearing Knee System (BP Knee). The BP Mobile Bearing Knee System consists of unapproved mobile (rotating) bearing tibial and patellar components. | PREMARKET APPROVAL: No Marketing Application | Endotec sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 21, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine and quarantine product subject to the recall. Additionally, a Response Form was included with the letter for customers to complete and return to the firm. Call 862-703-6730 for questions regarding this recall. | Contact the recalling firm for information | 243 | 243 | USA (nationwide) including the states of AZ, CA, CT, FL, NE, NJ, OH, and OK. | Terminated | September 27, 2013 | 625 | 2400 | Not_Computer
Z-0408-2012 | 60467 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | VenaCure EVLTÂŽ NeverTouch - FRS .018 " System PROCEDURE KIT Featuring Gold-Tipped Fiber, 45 cm, 45cm NeverTouch FRS .018 Procedure Kit, Catalog No./REF 11403012, STERILE -- AngioDynamics, Inc. 603 Queensbury, NY USA 12804. Product Usage: The AngioDynamics, Inc. VenaCure Procedure Kit is indicated for endovascular coagulation of the Great Saphenous Vein (GSV) in patients with superficial vein reflux, for the treatment of varicose veins and varicosities associated with superficial reflux of the Great Saphenous Vein (GSV), and for the treatment of incompetence and reflux of superficial veins of the lower extremity. | VenaCure EVLT | 2 | January 11, 2012 | 2012 | Angiodynamics Worldwide Headquarters | During an in-process inspection the hemostasis sidearm assemblies were identified to have the potential for flash to exist on the housing at the base of the threads. This flash has the potential to break off and relocate within the fluid path. | PRODUCTION CONTROLS: Process Control | AngioDynamics sent an URGENT: Medical Device Recall letters and response form (dated November 10, 2011) to all affected customers. The letter identified the affected products, the problem, and the actions to be taken by customer/user. The customers were instructed to do the following: 1) Identify, segregate and return the recalled products to ANGIODYNAMICS and 2) Complete the enclosed Recall Reply Form even if they do not have any products remaining in their possession, noting zero quantity to be returned and fax to 518-798-1360. The letter states that a returned authorization number (RA#) will be assigned to each customer in the recall notification letter, an UPS account will be provided for return freight and a replacement product will be shipped upon receipt and confirmation of the returned product. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. | Contact the recalling firm for information | 531 | 531 units | Worldwide Distribution - USA (nationwide) and the countries of: Canada, Great Britain, France, India, New Zealand, Netherlands, Saudi Arabia, Turkey, and Singapore | Open |  |  | 12205 | Not_Computer
Z-0448-2012 | 60387 | dressing, wound, drug | FRO |  |  | General & Plastic Surgery | 510(k) | Microcyn Solution with Preservatives (Rx), 16.9 fl oz (500 mL) and 33.5 fl oz. (990 mL)Ingredients: Electrolyzed Water (99.97%) Sodium Chloride (NaCl) (230 mg/L) and preservatives: Sodium Hypochlorite (Na0Cl) (40 mg/L), Hypochlorous acid (H0CL) (30 mg/L)-Non Foaming-Does not contain antibiotics- Does not facilitate resistance- No known drug/treatment interactions or contraindications-laboratory proven in solution inactivation of bacteria, viruses, fungi and spores.Performance testing: Microcyn Solution with Preservatives has been evaluated by the USP | Microcyn Solution | 2 | January 11, 2012 | 2012 | Oculus Innovative Sciences Inc | Unapproved drug claims and those not approved under 510K clearance. | PRODUCTION CONTROLS: Process Control | Oculus Innovative Sciences sent an"URGENT MEDICAL DEVICE RECALL" letter dated November 15, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to identify and remove all affected product from inventory. A Business Reply Card was attached to the letter for customers to complete and return via fax to 1-415-462-5181. Call Oculus Innovative Sciences Customer Service at 1-800-759-9305 for questions regarding this recall. |  | 4890 | 1363 x 500 mL; 2037 x 990 mL. | Nationwide Distribution. | Terminated | December 07, 2012 | 331 | 19918 | Not_Computer
Z-0605-2012 | 60658 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | ETCO2 Divided Sampling Cannula with Simultaneous Oxygen Delivery, Model number 4707 7-7-24.Salter Labs Nasal Cannula (Adult), Salter Stylewith modified nasal prongs.w/7' O2 line and 7' CO2 line with male luer lock connector.Salter Labs100 W. Sycamore Road,Arvin, CA 93203 | ETCO2 Divided Sampling Cannula with Simultaneous Oxygen Delivery | 2 | January 11, 2012 | 2012 | SALTER LABS | Mis-labeling: report of cases labeled as 4707 (cannulas with male luer lock connector) actually contained 4707F (with female luer lock connectors). | PRODUCTION CONTROLS: Process Control | The firm notified the three affected customers directly by phone, followed by a recall form for each direct customer indicating the product, part number, lot number and shipment date and customer order specific for each firm. Customers were to complete and return the form to the fax number listed or call a Customer Service Representative at 1-800-628-4690. | Contact the recalling firm for information | 240 | 240 cases | Nationwide Distribution -- TX and VA. | Terminated | January 09, 2012 | -2 | 240 | Not_Computer
Z-0735-2012 | 60700 | dc-defibrillator, low-energy, (including paddles) | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | R2 Adult Zoll Non-Radiolucent Multifunction Electrodes, Catalog/REF 3111-1721, NON-STERILE, ASSEMBLED IN MEXICO, CONMED CORPORATION, 525 FRENCH ROAD, UTICA, NY 13505.This device, when connected to a monitor/defibrillator has the ability to monitor the electro-cardio function (ECG) of the patient. The device, when electrically activated by the monitor/defibrillator can pace the heart rate of the patient and may also transmit the electrical current to the patient for cardio-conversion. This is a non-sterile, disposable device. | R2 Adult Zoll NonRadiolucent Multifunction Electrodes | 2 | January 12, 2012 | 2012 | ConMed Corporation | The product lot was assembled with an R2 Connector, rather than a Zoll Connector. | DESIGN: Device Design | ConMed Corporation sent an "URGENT: DEVICE RECALL" letter dated March 9, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to review their inventory for the affected product and to stop use of the devices immediately. A Business Reply Form was attached for customers to complete and return to the firm via fax to 315-624-3225. Additionally, instructions were provided on how to return the affected devices. Contact the ConMed Recall Coordinator at 315-624-3237 for questions regarding this notice. | Contact the recalling firm for information | 190 | 190 sets (2 pads per set in individual pouch) | Nationwide Distribution-including the states of AZ, CA, IN, and MI. | Terminated | January 12, 2012 | 0 | 190 | Not_Computer
Z-0724-2012 | 60804 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AngiodynamicsÂŽ DuraMaxÂŽ 15.5F x 55cm Chronic Hemodialysis Catheter Set (with Cuff 50cm from Tip) --- Catalog No. [REF] 10302809 --- MANUFACTURED IN USA AngioDynamics, Inc., 603 Queensbury Avenue, Queensbury, NY USA 12804Product Usage: indicated for use in attaining long-term vascular access for Hemodialysis and Apheresis. | DuraMax | 2 | January 12, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The securement cuff on the hemodialysis catheter has the potential to detach from the shaft. | PRODUCTION CONTROLS: Process Control | AngioDynamics sent an Urgent - Medical Device Recall letter dated December 30, 2011, via certified mail to all affected customers. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to examine inventory, segregate the recalled lots and return remaining inventory to the firm. The letter states that replacement product will be shipped upon receipt and confirmation of the returned product. Customers were advised to complete the Recall Reply Form included and fax to DuraMax Hemodialysis Catheter Recall Coordinator at 518-798-1360. For questions contact your local representative or call the Customer Service Manger at ANGIODYNAMICS Customer Service at 1-800-772-6446. | Contact the recalling firm for information | 191 | 191 units | Worldwide Distribution - USA (Nationwide) | Open |  |  | 4271 | Not_Computer
Z-0734-2012 | 60810 | lens, contact, (disposable) | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Outer Carton labeled in part:" 1-DAY ACUVUE MOIST BRAND CONTACT LENSES WITH LACREON FOR ASTIGMATISM***Contact Lenses for Daily Wear***D -1.50***-1.25 100***2016/07***BC 8.5***DIA 14.5***Made in USA***LOT B00BVTN0VS***"Individual Lenses labeled in part: "***1-DAY ACUVUE MOIST FOR ASTIGMATISM WITH LACREON***BC 8.5***DIA 14.5***AXIS -1.75 -0.75***080***2016/06***LOT B00BSBH***"It is intended for daily wear for the correction of refractive ametropia (myopia and hyperopia) in aphakic or non-aphakic persons with not diseased eyes who may have 1.00D or less of astigmatism. | 1Day Acuvue Moist contact lenses for Astigmatism | 3 | January 12, 2012 | 2012 | Johnson & Johnson Vision Care, Inc. | Johnson & Johnson Vision Care, Inc. recalled their 1-DAY ACUVUE MOIST Brand Contact Lenses for ASTIGMATISM due to a lot being mixed. The secondary carton lot number and power are not the same as on the primary package inside the carton. | TRAINING: Employee Error | The firm, Johnson & Johnson Vision Care, sent an "IMPORTANT PRODUCT NOTIFICATION" letter dated 1 December 2011 to all consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to check their practice/store inventory for product affected by this discrepancy in labeling; if you dispensed this product to a patient; contact the patient and replace lenses to all intended visual correction and performance with the correct lens power; complete and return ACKNOWLEDGEMENT OF RECEIPT Form via fax to: 904-443-3442 or email: VPIWEB@visus.jn.com and return any cartons labeled with the noted lot number to Customer Relations. Johnson & Johnson will replace them with the correct product.For any questions regarding this labeling situation, contact Customer Relations at 1-800-843-2020. | Contact the recalling firm for information | 152 | 152 cartons | Worldwide distribution: USA (nationwide) including states of: NY and FL and countries of: Finland, France, Germany, Italy, Netherlands, Norway, Denmark, Switzerland, Sweden, and the United Kingdom. | Terminated | June 18, 2012 | 158 | 152 | Not_Computer
Z-0742-2012 | 58889 | airbrush | KOJ | Airbrush. | Dental | Dental | 510(k) | New Label: EtchMaster ***Cordless Handpiece***P/N 797 www.GromanDental.com 1800-906-9298 HAZARD WARNING: Using CO2 cartridges from other suppliers may cause an operational hazard and will void the warrantee.Old Label: EtchMaster***Cordless Handpiece***Pats. Pend. Groman, Inc. Adapater P/N 747 WWW.DisposableAirAbrasion.com 1-800-906-9298.The Cordless Etcher air brush is powered by CO2 mini-cartridges to deliver pressurized gas to the EtchMaster pre-filled disposable micro air abrasive tips. The Cordless Etcher handpiece is quickly re-loaded with a CO2 mini-cartridge. Each CO2 mini-cartridge provides sufficient power to run 3 EtchMaster tips. The EtchMaster tips are used to abrade tooth surfaces and prosthetic devices prior to bonding procedures. | Cordless Etcher Air Brush | 3 | January 13, 2012 | 2012 | Groman Inc. | Groman Inc. of Margate, FL recalled their product Cordless Etcher Air Brush. When non-conforming cartridges are inserted into the Cordless Etcher they can induce excessive stresses on the device causing cracks in the plastic housing of the device. These crack can lead to a hazardous situation due to the high pressure of the gas contained in the CO2 cartridges. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Groman Inc. sent a "HAZARD NOTICE: IMPORTANT NOTICE ABOUT CORDLESS ETCHER HAZARD DUE TO IMPROPER USE" letter dated March 1, 2011 to all affected customers. The letter describes the product, problem, and advises the customers of the proper use of the device. Customers with questions and concerns, should contact the President of Groman Dental at 800-906-9298 ext. 709. | Contact the recalling firm for information | 395 | 395 | Worldwide Distribution-USA (nationwide) and the country of Rotterdam, Netherlands. | Terminated | January 17, 2012 | 4 | 395 | Not_Computer
Z-0756-2012 | 60798 | drills, burrs, trephines & accessories (manual) | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | 31144 Combo Kit including 1 Cranial Kit w/Cam Chuck, 1 Micro Ventricular Bolt ICP, 1 CSF Drainage System, contains sterile and non-sterile product. Integra NeuroSciences, Integra LifeSciences 3498 West 2400 South, West Valley City, Utah 84119.Product Usage: The Xylocaine component of the kit is used as an injectable local anesthetic during cranial access operations. | Integra | 2 | January 13, 2012 | 2012 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Current package insert was revised to include additional warning information related to the potential hazards of an unapproved off-label use of Xylocaine. | MISBRANDING: Labeling False and Misleading | Integra sent an Urgent: Medical Device Correction Notice letter dated December 21, 2011 to all affected customers. The letter identified the product, problem and actions to be taken. Integra is requesting that customers further distribute the communications along with the package insert to all appropriate health care practitioners at their facility. The letter instructs customers to complete and return the attached acknowledgement form by fax to 609-750-7999. For questions or concerns call 609-750-2814. | Contact the recalling firm for information | 62 | 62 units. | (USA) Nationwide Distribution | Terminated | May 10, 2012 | 118 | 30391 | Not_Computer
Z-0740-2012 | 60722 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | U-Joint Hex Head Driver 3.5mm, Part Number 228-0016, for the Escalade Acetabular Cup System. Driver is included in Part Number 270-9004-Escalade Acetabular Prep Kit and Part Number 228-9001-Triplus Acetabular Instrument Kit.The Acetabular Cup System is indicated for use in total hip arthroplasty procedures where the means of cup fixation is uncemented, biological fixation. The Hex Head Drive is used to place the screw that affixes the acetabular cup to the hip acetabulum. | UJoint Hex Head Driver | 2 | January 13, 2012 | 2012 | Ortho Development Corporation | Tip of hex-head screwdriver may break off during use in hip replacement procedure. | DESIGN: Device Design | Ortho Development Corporation sent an "URGENT!! DEVICE RECALL" letter dated December 19, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check to see if they have the affected device. Customers who have the device are to return the device to the firm, immediately. The firm will provide suitable replacements. Contact the Regulatory Affairs Manager for questions regarding this notice. |  | 37 | 37 units | Nationwide Distribution-including the states of AZ, CA, FL, MI, OR, TX, and UT. | Terminated | August 27, 2012 | 227 | 37 | Not_Computer
Z-0746-2012 | 60736 | film, radiographic | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | DVB KODAK DRYVIEW --- 35x43 cm --- 125 SHT --- REF 8723132 --- Made in U.S.A. by Carestream Health, Inc., 150 Verona St, Rochester, New York 14608.General purpose radiography. | Kodak DryView Laser Imaging Film | 3 | January 13, 2012 | 2012 | Carestream Health, Inc. | Kodak DryView DVB Film is being recalled because it does not meet the specification for maximum density. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carestream Health Inc. sent an "URGENT PRODUCT REMOVAL" letter dated December 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to identify, segregate, and return any affected product to the firm. A Product Information Form was attached to the letter for customers to complete and return via fax to 541-831-7101. Contact Technical Support at 1-800-328-2910 for questions regarding this notice. | No consumer action necessary | 24 | 24 cartridges | Worldwide Distribution-USA (nationwide) including the states of Alabama, Florida, Georgia, and Tennessee and the country of Australia. | Open |  |  | 24 | Not_Computer
Z-0747-2012 | 60548 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Olympus Electrosurgical Unit.The Olympus Electrosurgical Unit (ESG-100) is used in conjunction with cutting and coagulation electrosurgical devices intended for use in electrosurgery procedures during flexible gastrointestinal endoscopy. | Olympus Electrosurgical Unit (ESG100) | 2 | January 13, 2012 | 2012 | Olympus America Inc. | Investigations of customer complaints revealed a counterfeit capacitor was used in the manufacture of some lots of the ESG-100. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Olympus America Inc. sent an URGENT: MEDICAL DEVICE CORRECTION ACTION" letter dated November 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the affected product. A questionnaire was attached to the letter for customers to complete and return via fax at 484-896-7128. Contact the firm at 484-896-5000 for questions regarding this notice. | Contact the recalling firm for information | 245 | 245 | Nationwide Distribution. | Terminated | March 01, 2013 | 413 | 245 | Not_Computer
Z-0739-2012 | 60591 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | VenaCureEVLT Perforator Vein Ablation Kit, CATALOG #: EVLT/PVAK, STERILE -- Manufactured for: AngioDynamics Inc. 603 Queensbury, NY 12804 USA, AngioDynamics UK Limited, Building 2000, Beach Drive, IQ Cambridge, Waterbeach, Cambridge, CB259TE, United Kingdom --- COMMON/USUAL NAME: Laser Surgical Instrument; CLASSIFICATION NAME: Powered Laser Surgical Instrument; DEVICE LISTING # D053485. --- Product Usage: The 810nm Diomed Laser and EVLT Procedure Kit are intended for use in the treatment of superficial vein reflux of the greater saphenous vein associated with varicosities. The Diomed D15 plus and D30 plus and EVLT Kits are indicated for treatment of incompetence and reflux of superficial veins in the lower extremity. | VenaCureEVLT | 2 | January 13, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The product was labeled with a specific indication for the treatment of perforator veins and this specific indication had not received FDA clearance prior to the product being marketed in the US. |  | Angiodynamics sent an Urgent Medical Device Recall Letters and Response Forms dated December 2, 2011 to all affected consignees via Certified Mail on December 6, 2011. The recall was extended to the medical user level. The letter identified the product, problem, and actions to be taken. Customers were instructed not to use the product and to return recalled product to AngioDynamics, Inc. in Queensbury, NY. The letter states if the affected product/lot have been distributed, please inform your customers of this recall action immediately and have them return the recalled units. For questions or concerns call the Customer Service Manager at ANGIODYNAMICS Customer Service at 1-800-722-6446. |  | 7640 | 7640 units domestically | (USA) Nationwide Distribution | Open |  |  | 7640 | Not_Computer
Z-0738-2012 | 60857 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Linet Remote Handset for model Multicare, Eleganza 3 and Eleganza Smart, hospital bed. Product Usage:Positional hospital bed for acute care. The bed is designed for the care and treatment of inpatients at hospitals and is designed for standard and intensive care. | Linet | 2 | January 13, 2012 | 2012 | Linet Americas Inc | The premature failure of the Remote Handset buttons may result in unintended movement of the bed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Linet sent a notification letter dated December 9, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken.The letter states that firm Linet has decided to do a voluntary field correction to replace the recalled product. The replacement will be completed by the territory Representative (Linet Employee) who will replace the affected product for each customer personally. The replacement takes approximately 30 seconds and does not require the patient to be moved. For questions or concerns call 704-248-5650. |  | 29252 | 29,252 units | Worldwide Distribution - USA (nationwide) including the states of: AL, FL, KY, LA, MA, NC, NY, OH and the country of: Canada | Terminated | July 20, 2012 | 189 | 29252 | Not_Computer
Z-0743-2012 | 60803 | electrocardiograph | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Affected device is Model 98400-SL6-AHA, CardioExpress Electrocardiograph. The ECG cable part number is 01.57.471017. The electrocardiograph is only intended to be used in hospitals or healthcare facilities by doctors and trained healthcare professionals. The intended use of electrocardiograph is to acquire ECG signals from adult and pediatric patients through body surface with ECG electrodes. The electrocardiograph is only intended to be used in hospitals or healthcare facilities by doctors and trained healthcare professionals. The cardiogram recorded by the electrocardiograph can help users to analyze and diagnose heart disease. However, the interpreted ECG with measurements and interpretive statements is only intended to be used on adult patients and is offered to clinicians on an advisory basis only. | Spacelabs Healthcare CardioExpress SL6 ECG cable | 2 | January 13, 2012 | 2012 | Del Mar Reynolds Medical, Ltd. | CardioExpress SL6 ECG cable provided with the electrocardiograph is mislabeled. The left arm (LA) is mislabeled as right leg (RL). The left leg (LL) is mislabeled as left arm (LA). The right leg (RL) is mislabeled as left leg (LL). | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated September 24, 2010 to one affected customer in Massachusetts. The letter described the product, problem, and actions to be taken by the customers. Customers were instructed to review previous diagnoses made from this electrocardiograph to mitigate any inappropriate recordings or treatment made, and dispose of the mislabeled ECG cable to prevent its use.Spacelabs replaced the affected ECG cable with the new ECG cable.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 15 | 15 units | Nationwide distribution. | Terminated | January 13, 2012 | 0 | 15 | Not_Computer
Z-0760-2012 | 60737 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Apex Oxygenator, Sterile EO, Distributed in U.S. by: Sorin Group USA, Inc, 14401 W. 65th Way, Arvada, CO 80004, manufactured by: Sorin Group Italia, Via Statale 12 Nord, 86, 41037 Mirandola (MO) Italy.Intended for use in small-adult or adult surgical procedures requiring extracorporeal gas exchange support and blood temperature control for periods up to 6 hours. | Apex | 2 | January 14, 2012 | 2012 | Sorin Group USA, Inc. | Customer complaints of a leak between the blood and water compartments of the heat exchanger used in the oxygenators due to prolonged contact of the heat exchangers with sodium chloride causing micro-holes. | DESIGN: Labeling Design | Urgent Field Safety Notices dated December 2, 2011 were sent to customers. The letter identified the affected products and described the problem. The letter also provided additional directions to the Instructions for Use in order to minimize risk and clarify priming procedures. Customers were to communicate the new warning to all personnel in the use of the affected products and complete the Customer Response Form. The form should be returned via fax or e-mail. Questions should be directed towards Sorin Group Customer Service at 1-800-650-2623. |  | 19808 | 19,808 (9,550 stand alone/ 10,258 in packs) | Worldwide Distribution | Terminated | June 13, 2012 | 151 | 479778 | Not_Computer
Z-0772-2012 | 60618 | alkaline picrate, colorimetry, creatinine | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Creatinine Flex(R) Reagent Cartridges.The CREA method is an in vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma, and urine on the Dimension Vista System. | Dimension(R) Creatinine Flex(R) Reagent Cartridges | 2 | January 15, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed the potential for falsely elevated creatinine results with CREA FLex(R) reagent cartridge lots GA2196, GA2203, DC2221, DB2221, GA2229 and FA2237 when processing tests from a well that has been punctured. |  | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated November 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter recommends that customers discontinue and discard any remaining inventory of the affected product. The firm also recommends customers to test QC at least every 8 hours with lots manufactured after lot FA2237. Additionally, a Field Correction Effectiveness Check form was attached to the letter for customers to complete and return via fax to 302-631-8467. Contact the Siemens Customer Service Center at 888-588-3916 for questions regarding this notice. |  |  |  | Worldwide Distribution-USA (nationwide) including the District of Columbia and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and countries of Albania, Algeria, Angola, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belgium, Benin, Bosnia Herzegovina, Botswana, Brazil, Canada, Canary Islands, Ceuta/Mililla, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Fiji, France, French Guinea, Germany, Great Britian, Greece, Guadeloupe, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxemburg, Macedonia, Mexico, Montenegro, Morocco, Nepal, Nether Antilles, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uruguay, Venezuela, Vietnam, Yemen and Zimbabwe. | Open |  |  |  | Not_Computer
Z-0773-2012 | 60503 | calibrators, drug specific | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Tobramycin Calibrators, REF 7F93-01;Microgenics Corporation, 46360 Fremont Blvd Fremont, CA 94538;Distributed by Abbott Diagnostics, 100 Abbott Park Road, Abbott Park, Illinois 60064..Tobramycin is a commonly used antibiotic to treat bacterial infections, in particular gram negative infections. | Tobramycin Calibrators | 2 | January 17, 2012 | 2012 | Microgenics Corp | Reports that third party control material for 4 lots of Tobramycin Calibrators that are recovering low and outside of the published range. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ThermoFisher Scientific sent a product removal letter dated November 11, 2011. Letters were sent to the sole distributor of the product. The letter identified the product the problem and the action needed to be taken by the customer.If you are using or have inventory of the affected lot numbers immediately discontinue use and destroy any remaining affected lot numbers per your local waste ordinances.If you have any further questions regarding this information please call 1-877-4Abbott. Customers outside the US should contact their local customer service. |  | 1366 | 1366 kits | Worldwide Distribution -- Nationwide Distribution including the country Europe. | Terminated | May 09, 2012 | 113 | 1366 | Not_Computer
Z-0795-2012 | 60697 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Mevatron MD-2;Siemens HealthcareConcord, CA 94520Product Usage: PRIMUS and MEVATRON Linear Accelerators are intended to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | MEVATRON | 2 | January 17, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Bolts holding the trim weight may break causing the plates of the counterweight to become loose and fall into the cover. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent a Customer Information Letter on December 1, 2011 to all affected customers. The letter idenfied what is the issue, when does the issue occur and what Siemens will do to address the issue. The letter recommend that the Customer Information letter be included in the Linear Accelerator System Owner manual chapter Safety Advisory Letters. |  | 327 | 327 Active Devices - all models | Worldwide Distribution - USA (nationwide) | Terminated | December 02, 2013 | 685 | 327 | Not_Computer
Z-0774-2012 | 60670 | vehicle, motorized 3-wheeled | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Leo 3-Wheel Scooter and Leo Scooter, Distributed By:Invacare Corporation, 1200 Taylor Street, Elyria, OH 44035Intended to provide transportation for a disabled or elderly person. | LEO Scooter | 2 | January 17, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the products because the scooter backrest fails unexpectedly due to seat flexing, and this has the potential to result in a consumer falling causing death or serious injury. | DESIGN: Device Design | On 12/7/2011 the firm sent Urgent: Recall Information Letters to their customers. The letter informed customers of the affected product and the related problem. The letter also states that Invacare has developed replacement seat kits intended to replace the original seats. The replacement kit will be provided free of charge and can be ordered by completing and returning the attached response card. After receiving replacement kits, customers are asked to contact their customers that have potentially affected devices and replace the seats immediately. Questions should be directed to Customer Service at 1-800-333-6900. |  | 391 | 391 scooters | Worldwide Distribution -- USA, including the states of AK, CA, FL, GA, HI, IA, ID, IL, KS, LA, MD, MI, MN, MO, MT, NC, NY, OH, OK, OR, PA, SC, TX, UT, VA, WA, & WI, and the countries of AUSTRIA, BELGIUM, CZECH REPUBLIC, ESTONIA, FRANCE, GERMANY, IRELAND, ITALY, JERSEY (CHANNEL ISLANDS), LITHUANIA, PORTUGAL, SWEDEN, SWITZERLAND, & UNITED KINGDOM. | Terminated | July 02, 2013 | 532 | 391 | Not_Computer
Z-0775-2012 | 60883 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADDÂŽ High Volume Administration Set, REF 21-7081V-01, Sterile EO, Rx only, Made in Mexico, Smith Medical ASD, Inc., 1265 Grey Fox Road, St Paul, Minnesota 55112. Designed for use with CADD pumps to allow fluid delivery from an IV bag. EXPANDED: Catalog No./ REF 21-7381-01 | CADD High Volume Administration Set | 2 | January 17, 2012 | 2012 | Smiths Medical ASD, Inc. | Smith Medical has become aware of a small number of customer reports regarding occlusions/high pressure alarms with the use of one lot of CADD High Volume Administration Sets. EXPANDED: Smiths Medical became aware that an additional Product Reorder No. 21-738-01 and Lot No. 21X221 are also being recalled for receiving reports regarding occlusion/high pressure alarms with the use of these set | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Consignees were sent a Smiths Medical "Urgent Field Safety Notice" letter dated January 04, 2011. The letter was addressed to Risk/ Safety Managers, Clinicians, Nursing, Pharmacy, Pain Management and Anesthesia Professionals, Distributors and other users of these devices. The letter described the details on affected devices, Advise on Action to be Taken by the User and distributor and Transmission of this Urgent Field Safety Notice. Requested consignees to inspect their inventory and remove affected product and to complete and return the attached Confirmation Form. EXPANDED: The additional devices were distributed 12/14/2011 thru 1/27/2012. An Urgent Medical Device letter, dated 3/16/2012, was sent to customers beginning 3/16/2012. The letter identified the additional affected product, described what the issue is and asked that customers return all unused sets to Smiths Medical for credit or replacement. Also, the customer was asked to complete and return an Urgent Medical Device Recall Confirmation Form. If the customer is a distributor, they were asked to pass along the information to their customers. |  | 21510 | 5,385 EXPANDED 16,125 devices | Nationwide Distribution -- AZ, CA, FL, GA, IL, KS, LA, ME, MA, NC, OR, PA,SC, TX, VA, WA, and WI. EXPANDED to include AR, CO, IA, IN, MD, MI, MN, MO, NJ, OH, OK, TN, WV and Canada. | Terminated | November 06, 2012 | 294 | 16125 | Not_Computer
Z-0812-2012 | 60720 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Single Stage Venous Cannula, Carmeda Bio-Active Surface Venous Cannula, Sterilized Using Ethylene Oxide, Rx Only, Nonpyrogenic. Model # 'sCB66112, CB66114, CB66116, CB66118, CB66120, CB66122, CB66124, CB66128, CB66130, CB66132, CB66134, CB66136, CB66140, CB66236, CB66240, CB67312, CB67314, CB67316, CB67318, CB67320, CB67512, CB67514, CB67516, CB67518, CB67520, CB67522, CB67524, CB67528, CB67530, CB67532, CB67534, CB67536, CB67540, CB67636, CB67640, CB68112, CB68114, CB68116, CB68118, CB68120, CB68122, CB68124, CB68126, CB68128, CB68130, CB68132, CB68134, CB68136, CB68138, CB68140, CB69312, CB69316, CB69320,CB69322, CB69324, CB69328, CB69331, CB69428.These cannulae are intended for collection of venous blood from the right side of the heart via the superior and inferior vena cava during cardiopulmonary bypass surgery up to six hours or less. | Medtronic, Single Stage Venous Cannula | 2 | January 18, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | Instructions for Use for the DLP Single Stage Venous Cannulae are in the process of being updated. The current precaution statement referencing ventricular assist procedures will be identified as a contraindication in the updated Instructions for Use. DO NOT USE THIS PRODUCT FOR EXTENDED TERMS SUCH AS VENTRICULAR ASSIST PROCEDURES. Medtronic has received two reports of patient deaths due to p | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Medtronic sent an "URGENT MEDICAL DEVICE INFORMATION" letter dated December 2011 to all affected customers. The letter identifies the product, problem, and the action to be taken by the customers. Contract Medtronic CardioVascular LifeLine Technical Support at 1-877-526-7890 for questions related to this issue. |  | 14306 | 14,306 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the Virgin Islands and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Mauritius, Mexico, Montenegro, Morocco, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, and Yemen. | Open |  |  | 1232914 | Not_Computer
Z-0804-2012 | 60744 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Brand Name: Custom Straight AccoladeÂŽÂŽ Common Name: Handle Classification Name: Orthopedics Manual Surgical InstrumentBrand Name: Specialty Straight CommandÂŽÂŽ Common Name: Handle Classification Name: Orthopedics Manual Surgical InstrumentBrand Name: Custom Straight AccoladeÂŽÂŽCommon Name: HandleClassification Name: Orthopedics Manual Surgical InstrumentBrand Name: Custom AccoladeÂŽÂŽ Broach HandleCommon Name: HandleClassification Name: Orthopedics Manual Surgical InstrumentBrand Name: Custom AccoladeÂŽÂŽ Broach HandleCommon Name: HandleClassification Name: Orthopedics Manual Surgical InstrumentBrand Name: Specialty Locking AccoladeÂŽÂŽ Straight Broach Handle Common Name: Handle Classification Name: Orthopedics Manual Surgical InstrumentBrand Name: Specialty Woodpecker Adapted AccoladeÂŽÂŽ BroachHandleCommon Name: HandleClassification Name: Orthopedics Manual Surgical InstrumentStryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430The Handles are instruments that attach to the femoral rasps or broaches to prepare the Femoral Canal in total hip replacement procedures. The rasps have a geometry similar to the final size and shape of the implant. The straight design allows the rasp handle to be used in standard or minimally invasive approaches because it minimizes the contact with adjacent soft tissue. The Handle is impacted with a slotted mallet on the top and on the bottom portions of the striking plate, to achieve rasp insertion and extraction respectively. The handles share the same geometry in the region that is of interest in this matter. | Custom Straight Accolade Common Name: Handle Classification Name: Orthopedics Manual Surgical Instr | 2 | January 18, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become ware that there is the potential for the above referenced instruments to fracture in two pieces along the width of the body under the striking plate. | PRODUCTION CONTROLS: Process Control | Stryker Orthopaedics sent a Urgent Product Recall update dated August 18, 2011, via FedEx to all affected customers. The letter identified the product(s) the problem and the action needed to be taken by the customer.Our records indicate that you have received the above referenced product(s). It is Stryker'sÂŽ responsibility as the manufacturer to ensure that customers who may have received these affected products also receive this important communication. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form within 5 days to 201-831-6069.Please also contact the hospitals in your territory that have the affected product to arrange return of the product as replacement product becomes available.If you have any further questions please call (201) 831-5158. |  | 106 | 106 UNITS | USA (nationwide distribution) | Open |  |  | 106 | Not_Computer
Z-0806-2012 | 60613 | pump, air, non-manual, for endoscope | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus OFP Flushing Pump irrigation tubingThe Olympus OFP Flushing Pump is a peristaltic flushing pump intended for use with Olympus gastrointestinal videoscopes having an auxiliary water inlet as part of an Olympus endoscopic imaging system to facilitate washing of gastric and colonic mucosa to remove blood, feces and other organic matter, enabling improved visualization, diagnosis and therapy during endoscopic procedures. | Olympus OFP Flushing Pump irrigation tubing | 3 | January 18, 2012 | 2012 | Olympus America Inc. | Water may leak from the irrigation tubing used on the Olympus Flushing Pump | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm,Olympus, sent an "URGENT-DEVICE CORRECTION " letter dated December 12, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were informed that any current irrigation tubing in their inventory can be used with the OFP pump; however, they must inspect the tubing installed on the OFP pump for signs of water leakage; if leakage is found, discontinue use of the tubing and replace with a new irrigation tube. In addition, complete and return the enclosed questionnaire via fax to (484) 896-7128. If you do not wish to use your current inventory of irrigation tubing and would like a replacement, call Olympus Customer Care Center at 1-800-848-9024.Olympus is discontinuing the current version of irrigation tubing and has redesigned the irrigation tubing to prevent water leakage. For any additional information, contact the V.P., Regulatory/Clinical Affairs & Quality Assurance at (484) 896-5688 or email at laura.storms-tyler@olympus.com. | Contact the recalling firm for information | 10 | 64,667 packages of 10 irrigation tubes | Nationwide distribution: USA including states of: AL, AK, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, AND WY. | Terminated | September 25, 2012 | 251 | 64667 | Not_Computer
Z-0821-2012 | 60729 | cardiovascular procedure kit | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC" Convenience Kit --- Legal Manufacturer Navilyst Medical, Inc, 26 Forest St, Marlborough, MA 01752 USA --- Made in USA: 10 Glen Falls Technical Park, Glen Falls, NY 12801 USA --- each unit label contains the customer name under the name of the product, and each catalog number is specific to the customer.Product usage: For use in fluid management and/or invasive pressure monitoring systems | NAMIC Convienience Kit | 2 | January 19, 2012 | 2012 | Navilyst Medical, Inc | Navilyst Medical has received reports of damaged manifolds contained in certain NAMIC Convenience Kits which may result in air injection during certain procedures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Navilyst Medical sent an Urgent Medical Device Recall- Immediate Action Required letter dated December 19, 2011 and a Reply Verification Tracking Form to all affected consignees via Federal Express or email. The letter identified the product, description of problem, and actions to be taken Customers were instructed to 1) IMMEDIATELY DISCONTINUE USE; 2) remove the recalled product from your inventory (whether in Labs, Central Supply, Shipping and Receiving or ANY other location); 3) segregate this product in a secure location for return to Navilyst Medical, Inc; 4) immediately forward a copy of this recall notification to all sites to which you have distributed affected product; 5) complete and return the Reply Verification Tracking Form; and 6) package and return the recalled product per the instruction provided. | No consumer action necessary | 4547 | 4547 | Worldwide Distribution - USA (nationwide) including the states of Connecticut, Florida, Illinois, Indiana, Kentucky, Maryland, Michigan, New Jersey, New York, Ohio, South Carolina, Wisconsin, Virginia and the country of Japan. | Open |  |  | 4547 | Not_Computer
Z-0661-2012 | 60449 | tube, tracheostomy (w/wo connector) | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Bivona Uncuffed Neonatal Flex Tend Plus Tracheostomy Tube Sizes 2.5 - 4.0 (With a Straight Neck Flange)Product Usage: This tube is intended to provide direct airway access for a tracheostomized patient for up to 29 days. It may be reprocessed up to 5 times for single patient reuse. | Bivona Uncuffed Neonatal Flex Tend Plus Tracheostomy Tube | 1 | January 19, 2012 | 2012 | Smiths Medical ASD, Inc. | Difficulty disconnecting accessories from the Bivona Uncuffed Neonatal, Pediatric and Flextend Tracheostomy Tubes TR3 connector, which may result in a delay in initiating therapy or patient discomfort. If excessive force is used to detach the accessory, this may result in decannulation of the tube which could result in serious injury or death; especially if no replacement tracheostomy tube is imm | DESIGN: Device Design | Smith Medical sent an Urgent Field Safety Notice dated November 28, 2011, via U.S. Mail to all affected consignees. The notice identified the product, description of problem, and actions to be taken. This notice included a Customer Information Bulletin and Urgent Field Safety Notice Confirmation Form. Customers were instructed to 1) identify all affected unused product in inventory and segregate it to a quarantine location; 2) complete and return the attached Confirmation Form by fax to 219-989-7259 or email to bivona.tr3@smiths-medical.com; and 3) distributors were instructed to stop distributing and provide Smith Medical with a list of their customers, including contact details, so that Smith Medical can notify the customers of this notice. For questions contact Smiths Medical Customer Service at 800-258-5361, Option 1. | Contact the recalling firm for information | 174523 | US distribution 153,033 devices ; Foreign 21,490 devices | Worldwide Distribution - USA Nationwide: AZ, CA, CO, CT,DE, FL, GA, HI, IL, IN, KS, KY, LA, SC, TX, OR, NJ, MO, UT, MI, NY, NM, MA, MN, MS, NC, ND, NE, NH, NV, OH, OK, OR, RI, SD, TN, VA, VT, WA, WI, WV including Washinton DC, Australia, Austria, Barbados, Belgium, Brazil, Canada, Chile, Cyprus, Czech Republic, Denmark, Egypt, Finland, France, Germany, Gibraltar, Great Britain (UK), Greece, Hong Kong, Ireland, Israel, Italy, Kuwait, Libya, Malaysia, Netherlands, New Zeland, Norway, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirates, United States, Vatican City State | Open |  |  | 174523 | Not_Computer
Z-0816-2012 | 60466 | stretcher, wheeled | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Hyperbaric Chamber Gurney.The Sechrist Hyperbaric Chamber Gurney is an accessory device to the Sechrist Monoplace Hyperbaric Chambers. It is intended solely for the loading and unloading of patients into Sechrist Monoplace Hyperbaric Chambers. It is not intended for general hospital patient transport and is sold only to be used in conjunction with Sechrist Monoplace Hyperbaric Chambers. | Sechrist Hyperbaric Chamber Gurney. | 2 | January 19, 2012 | 2012 | Sechrist Industries Inc | Stryker Medical has received complaints from customers that indicate a bushing in the braking system may fracture, which could result in the brake/steer pedal seizing. If the pedal seizes, the braking system may become inoperable or steering around corners may come more difficult. While no injuries have been associated with these reported events nor has there been reported failure events resulting | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sechrist Industries sent a Urgent Medical Device Correction letter dated September 9, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1. Locate the gurneys with the serial numbers listed in this advisory.2. If any of the affected gurneys cannot be located, determine if they have been sold, transferred, or donated to another facility and notify us regarding their new location.3.Verify the braking system is working properly. If the brake system is inoperable, or if the brake and /or steer require a greater than normal force to engage, remove the stretcher from service and contact Stryker Medical at 800-Stryker, option 3. Reference RA-2009-135. Do not put the stretcher back into service until is repaired by the Stryker field representative.4. Contact Stryker Medical to schedule a service call for replacement of the bushing and mating parts of the stretcher.5.Complete retrofit of the Gurney.6.Verify correct function of the brake assembly following retrofit.7.Complete and send the Correction Verification Form contained with this notice to Sechrist.For further questions please call (714) 579-8400.. |  | 198 | 198 | Worldwide Distribution -- USA (nationwide) and the countries of Peru, Venezuela, Brazil, Egypt, Mexico, Austrailia, Denmark and Taiwan | Terminated | January 19, 2012 | 0 | 198 | Not_Computer
Z-0817-2012 | 60869 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira Blood Set, with 200 Micron Filter, 80 Inch, Non-DEHP; an Rx sterile medical device used to administer blood products; 48 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 14203-28.Intended use: for the administration of blood products. | Hospira Blood Set, with 200 Micron Filter | 2 | January 19, 2012 | 2012 | Hospira Inc. | Silicone applied to the piercing pin of the blood set is reducing retention force of the piercing pin while in use. This could cause the piercing pin to slip out of the administration container. | DESIGN: Component Design/Selection | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated January 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory and immediately quarantine any affected product. Additionally, a Reply Form was attached for customers to complete and return via fax to 1-866-912-8410 or e-mail to: hospira2573@stericycle.com. Please contact Hospira Customer Care at 1-877-946-7747 for questions regarding this recall. |  | 960 | 960 units | Distributed in the states of North Carolina and Washington. | Open |  |  | 960 | Not_Computer
Z-0815-2012 | 60897 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Handi-Fil Disposable Syringe Fill Tube P/N 302050, Plastic & Tyvec pouch, Mallinckrodt, Inc., Libel-Flarsheim, Cincinnati, OH Made in Mexico.The Handi-Fil Disposable Syringe Fill Tube is a sterile plastic tube, approximately 10 inches in length, with a "J" shaped curve the projects upward approximately two inches. The Handi-Fil Disposable Syringe Fill Tube is used to transfer radiographic contrast media from a sterile vial to an injection syringe. | HandiFil Disposable Syringe Fill Tube | 2 | January 19, 2012 | 2012 | Mallinckrodt Inc | The firm received complaints from their customers stating that they found open seals on two of their sterilized products. | PRODUCTION CONTROLS: Packaging Process Control | Mallinckrodt Inc. sent an "URGENT DEVICE RECALL" letter dated December 20, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the affected product and to complete and return an attached Business Response form via fax to (314) 654-8206 or e-mail to imaging.recalls@covidien.com. Contact the Product Monitoring Department at 800-778-7898 for questions regarding this notice. | Contact the recalling firm for information | 63900 | 63,900 units | Worldwide Distribution--USA (nationwide) including the states of AZ, CA, GA, IA, IL, KS, KY, LA, MA, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, TN, WA, WI and WV and countries of Australia, Belgium, Canada, Japan and Singapore. | Terminated | February 19, 2013 | 397 | 63900 | Not_Computer
Z-0831-2012 | 58983 | light, surgical, fiberoptic | FST | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Surgical Headlight labeled in part *Gulf Medical Fiberoptics*OLDSMAR, FLORIDA U.S.A.* Model Numbers: GMF-H001, GMF-H100, GMF-H101, GMF-H102, GMF-H103, GMF-H104, GMF-H105, and GMF-H106. The Surgical Headlight is a passive luminaire. It is illuminated by a source of light by way of the fiber optics and provides supplemental lighting for medical and surgical procedures. | Surgical Headlight | 2 | January 20, 2012 | 2012 | Gulf Medical Fiberoptics, Inc. | Gulf Medical Fiberoptics, Inc. recalled their Surgical Headlight because the cleaning instructions for the exterior surfaces of the headlight , headgear and fiber cable with 70% isopropyl alcohol had not been validated. | PRODUCTION CONTROLS: Process Control | On May 16, 2011, Gulf Medical Fiberoptics, Inc. issued to all customers an updated letter to their May 13, 2011 "Field Correction of Instructions For Use" informing them to discontinue the use of 70% isopropyl to clean the surfaces of the device. Enclosed with the letter was an updated Surgical Headlight User Guide with instructions to discard/shred/recycle the prior guide. The firm asked customers to return a form acknowledging the receipt and action from the letter. Customers with questions should contact Customer Service department manager Jacqueline Armand at 813-855-6618 ext. 427. | No consumer action necessary | 381 | 381 units | Worldwide Distribution - USA, including the states of FL, WI, WA, AL, CA, MO, MI, OH, VA, CO, TX, TN, NC and the countries of Turkey, Brazil, Belgium, South Korea, Panama, Mexico, Ireland, Taiwan (ROC), Italy, Argentina, India, Spain, Germany, and Venezuela. | Terminated | August 28, 2012 | 221 | 381 | Not_Computer
Z-0825-2012 | 60734 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F250NRe Hemodialyzer, , intended for single use in acute and chronic hemodialysisCatalog No. 0500325EThe Fresenius Optiflux F250NRe Hemodialyzers are intended for patients with acute or chronic renal failure when conservative therapy is judged to be inadequate. | Fresenius Optiflux F250NRe Hemodialyzer | 2 | January 20, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Optiflux F250NRe Hemodialyzers have increased risk of an internal blood leak | PRODUCTION CONTROLS: Process Control | Fresenius Medical Care North America contacted customers by telephone on December 27, 2011, and sent a Customer Notification letter dated December 27, 2011, via certified mail return receipts and fax - back requested to all affected customers. Customers were instructed to examine your stock of dialyzers, and isolate and discontinue use of the before mentioned lots. If you have identified dialyzers from the recalled lots, please contact Customer Service for a Return Goods Authorization. We will pick up the product you currently have in stock and supply replacement product. All products for return will be picked up during your next scheduled Fresenius delivery. Promtly fill out the fax-back reply form attached to the notification. Please do not return used dialyzers.For further questions, please call 1 (800) 662-1237. |  | 599436 | 599,436 units | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | July 08, 2013 | 535 | 599436 | Not_Computer
Z-0820-2012 | 60788 | esophagoscope (flexible or rigid) | EOX | Esophagoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | EE- Series Video Naso Pharyngo Laryngo Scope, VNL-1570STK.Intended to provide optical visualization (via a video monitor) of the Upper Airway. | Pentax Video Naso Pharyngo Laryngo Scope | 2 | January 20, 2012 | 2012 | Pentax Medical Company | The wrong silicone sealant was used and excess sealant was neglected to be removed during repair of the bronchoscope. | TRAINING: Employee Error | Pentax sent an "Urgent: Device Recall - Stop Use Immediately" letter/return response form to their customers on 9/30/2011. The letter identified the affected product and asked customers to immediately discontinue the use of their endoscope(s). Customers were to contact PENTAX at 1-800-736-8293 to schedule the return of the affected product(s). The reason for the recall was also explained, and the letter stated that there will be no charge for shipping, evaluation, or repair of the products. Customers can request loaner instruments, if necessary. |  | 1 | one unit | Nationwide Distribution. | Terminated | November 27, 2012 | 312 | 19 | Not_Computer
Z-0829-2012 | 60772 | unit, electrosurgical, endoscopic (with or without accessories) | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus Endoscopic Electrosurgical Unit working elements, model numbers WA22366A, and WA22367A.Working element for cutting, coagulation, and vaporization in conductive irrigation fluid or conductive distension medium in the Saline mode. | Olympus Endoscopic Electrosurgical Unit working elements | 2 | January 20, 2012 | 2012 | Olympus America Inc. | Less than optimal thickness in the plastic housing holding the electrode in the Olympus Working Elements | DESIGN: Component Design/Selection | Olympus America Inc. sent an "URGENT: MEDICAL DEVICE WITHDRAWAL" letter dated December 27, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to inspect their inventory for the affected product and discontinue use of any affected product discovered. Customers should contact the Gyrus ACMI Customer Service Department at 1-888-524-7266 for credit or replacement of the product. Additionally, an Olympus Working Elements Withdrawal form was enclosed for customers to complete and fax to 484-896-7128. Contact the firm at 484-896-5688 for questions regarding this notice. |  |  |  | Nationwide Distribution-including Washington, D.C. and the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MO, MT, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, and WI. | Terminated | December 21, 2013 | 701 |  | Not_Computer
Z-0824-2012 | 60539 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HemoCue Glucose 201 Microcuvettes, Article Number:110705*,110706,110715**,110716**,110717**,110718**,110719**" No affected lots of this article have been delivered to customers. Stillin stock at distributor HemoCue Inc."" Article not available in USQuantitative determination of glucose in whole blood, using a specially designed analyzer, the HemoCue Glucose 201 Analyzer. The quantitative determination of the instant blood glucose concentration in circulation supplements the clinical evidence in the diagnosis and treatment of patients with diabetes as well as monitoring of neonatal blood glucose levels. HemoCue Glucose 201 Microcuvettes are for In Vitro Diagnostic use only and are only to be used with the HemoCue Glucose 201 Analyzer. For professional use only. | HemoCue Glucose 201 Microcuvettes, | 3 | January 20, 2012 | 2012 | Hemo Cue, Inc. | A recall was initiated because Hemocue confirmed that the package insert for Hemocue Glucose 201 Microcuvettes had errors in two sections. | PRODUCTION CONTROLS: Packaging Process Control | HemoCue sent a Medical Device Correction Notification letter dated June 14, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to passed on the notice to all that need to be aware with their organization or to any organization where the affected lot numbers have been distributed. Customers were instructed to sign the Correction Notification Verification form below and return to HemoCue Inc. Customers received the revised package insert with the notification letter. Customers with any questions regarding the Correction Notification could please call (562) 668-5665. |  | 13620 | 13,620 packages | Worldwide Distribution--USA (nationwide) and the countries of Czech Republic, Denmark, Slovenia, Norway, Sweden, Switzerland, United Kingdom, Finland, Germany, Netherlands, Poland, Romania, and Tanzania. | Terminated | March 02, 2012 | 42 | 13620 | Not_Computer
Z-0826-2012 | 60807 | subsystem, proportioning | FKR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NXStage PureFlow SL SAK-303 - 2 dialysate sacks (SAK) with concentrate for use with PureFlow SL, RxThe NxStage PureFLow SL is an optional accessory to the NxStage System One that prepares dialysate for use during hemodialysis, as prescribed by the physician. | NXStage PureFlow SL SAK303 | 3 | January 20, 2012 | 2012 | NxStage Medical, Inc. | Report of high conductivity failure alarms with the PureFlow SL due to excess concentrate by weight in the product (SAK) dialysate Preparation System | TRAINING: Employee Error | On December 13, 2011, NxStage sent replacement product for overnight delivery and initiated phone contact with all consignees of the product . In addition, NxStage sent an "IMPORTANT MEDICAL DEVICE RECALL" notice dated December 14, 2011 to its consignees/customers. The notice describes the product, problem and actions to be taken. The customers were instructed to use SAKs from a different lot or premixed dialysate bags (RFPs), until their replacement SAKs arrived; set aside all SAKs from listed lot number; record their inventory; dispose of the SAKs following their normal practices, and complete and return the DEVICE PRODUCT RECALL REPLY FORM even if they do not have any of the affected product on hand via postage-prepaid envelop provided. If you have any questions about this notice, contact NxStage Customer Support at 866-Nxstage ( 697-8243). | Contact the recalling firm for information | 1192 | 1,192 devices | Nationwide distribution. | Terminated | April 17, 2012 | 88 | 1192 | Not_Computer
Z-0823-2012 | 60910 | mandibular condyle prosthesis | MPL | Mandibular condyle prosthesis. | Dental | Dental |  | The 45mm Left Standard Offset Mandible is one component in the Total Temporomandibular Joint (TMJ) Replacement System used to functionally reconstruct a diseased and/or damaged temporomandibular joint.The component is labeled in part: "REF: 24-6646***TMJ SYSTEM LEFT STANDARD OFFSET MANDIBULAR COMPONENT***QTY 00001***45MM / 7 HOLE CO-CR-MO, TI-6AL-4V ALLOY***BIOMET MICROFIXATION 1520 Tradeport Drive Jacksonville, FL 32218 USA www.biometmicrofixation.com***Product Recall: The Total Temporomandibular Joint (TMJ) Replacement System is implanted in the jaw to functionally reconstruct a diseased and/or damaged temporomandibular joint. The Total TMJ Replacement System is a two-component system comprised of mandibular condyle and glenoid fossa components. Both components are available in multiple sizes as right and left side specific designs and are attached to bone by screws. Mandibular (Condylar) Prosthesis- The prosthesis is designed to replace the articular surface of the mandibular condyle. The mandibular prosthesis is offered in 3 different sizes: 45mm, 50mm, and 55mm, designated left and right. The mandibular prosthesis is offered in 3 styles:Standard, Offset, and Narrow, to fit a diverse range of mandibular sizes and shapes. The mandibular prosthesis is made of Cobalt Chromium Alloy. The under-surface of the prosthesis is coated with titanium plasma spray for increased bony integration to the mandibular prosthesis. | 45mm Left Standard Offset Mandible | 2 | January 20, 2012 | 2012 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc., Jacksonville, FL recalled Part # 24-6546, 45mm Left Standard Mandible and Part # 24-6646, 45mm Left Standard Offset Mandible due to a lot of 24-6246 being etched & labeled as 24-6646 and the reverse, a lot of 24-6646 being etched and labeled as 24-6546. | TRAINING: Employee Error | Biomet Microfixation Inc sent an URGENT-Medical Device Recall notice to all consignees. The letter informs the customers to discontinue use of the items and return them to Biomet Microfixation.They also ask all customers to complete the inventory reconciliation sheet and return it by fax to the Global Complaint Coordinator. The letter also states that if the product was further distributed, the letter must be forwarded on to the firm's customers. For questions call 1-800-874-7711 or 904-741-4400, ext 468. | No consumer action necessary | 2 | 2 pieces | Worldwide Distribution - USA (nationwide) in the state of Florida and in the countries Germany and Guatemala. | Terminated | February 06, 2012 | 17 | 2 | Not_Computer
Z-0830-2012 | 60874 | injector, contrast medium, automatic | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Illumena 150mL Linden Luer Syringe with Handi-Fil Straw (p/n 900103) & Illumena 150mL Regular Luer Syringe with Handi-Fil (p/n 900101), Made in Mexico, Liebel-Flarsheim Company, Cincinnati, OHThe Liebel Flarsheim Illumena 150mL Syringes with Handi-Fil Straw are designed for use with the Illumena Injector System. The Angiomat Illumena is designed to inject a radiopaque contrast medium into the vascular system for Angiographic or CT procedures. | Illumena 150mL Linden Luer Syringe (p/n 900103) & Illumena 150mL Regular Luer Syringe (p/n 900101) | 2 | January 20, 2012 | 2012 | Mallinckrodt Inc | Syringes packaged and labeled as the Illumena 150mL Linden Luer Syringe with Handi-Fil Straw, product 900103, contained the Illumena 150mL Syringe with Handi-Fil Straw, product 900101. | PRODUCTION CONTROLS: Packaging Process Control | The firm, Covidien, sent an "URGENT DEVICE RECALL" letter dated October 11, 2011 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to check their inventory, if any of the affected lots are found, discontinue use of the lots immediately; complete and return the BUSINESS RESPONSE FORM even if they do not have any of the affected lots via fax to (314) 654-8206 or email to imaging.recalls@covidien.com by October 28, 2011; return product and notify any consignees that this product may have been distributed to.If you have any questions about this recall, contact our Product Monitoring department at 800-778-7898. |  | 900101 | 187,400 units (900103 syringe: 129,200; 900101 syringe: 58,200) | USA (nationwide) including states of: AZ, CA, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI & WV; and countries of: Australia, Belgium, Brazil, Canada, Singapore, & Uruguay. | Terminated | July 19, 2013 | 546 | 187400 | Not_Computer
Z-0835-2012 | 59181 | screw, fixation, bone, non-spinal, metallic | NDJ | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Package of cross-drive bone screws labeled in part: "***REF #95-1504***"1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, 5/ PK***1.5 X4MM TITANIUM***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN***BIOMET MICROFIXATION***1520 TRADEPORT DRIVE***JACKSONVILLE, FL 32218 USA***www.biometmicrofixation.com***""***REF: 91-1504***"1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW***1.5 X4MM TITANIUM***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN.***BIOMET MICROFIXATION***1520 TRADEPORT DRIVE***JACKSONVILLE, FL 32218 USA***www.biometmicrofixation.com***"Products subject to recall:REF #95-1504***"1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, 5/ PK, Lot 622120. REF: 91-1504***"1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, 1/PK, Lot 622120. Internal fixation screw intended to aid in the alignment and stabilization of bone in the oral cranio-maxillofacial skeletal system. | Cross Drive Screw | 3 | January 23, 2012 | 2012 | Biomet Microfixation, Inc. | On 11/08/2010 Biomet Microfixation, Jacksonville, FL initiated a recall of the Biomet 1 count (REF 91-1504, Lot 622120) and 5 count (REF 95-1504, Lot 622120) packages of 1.5MM System High Torque (HT) Cross-Drive Screw 5/Pk and the Biomet 1 count (REF 99-7204, Lot 202450) and 5 count packages (REF 01/7204, Lot 202450) of the 1.5MM System Center-Drive Screw, 5/Pk because of a possibility that thes | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Biomet Microfixation, Jacksonville, FL, sent an Urgent Medical Device Safety Alert letter dated November 8, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to fill out the attached Inventory Reconciliation sheet and FAX to 904-741-9425. If consignees further distributed this product, they must notify the recipient of this action by providing the recipient a copy of the notice. For questions call 800-874-7711 or 904-741-4400, ext 468. | Contact the recalling firm for information | 295 | 295 | Worldwide Distribution - USA including CA, GA, OH, TX, and VA and the countries of S Africa and Japan | Terminated | January 31, 2012 | 8 | 779 | Not_Computer
Z-0838-2012 | 60691 | drug mixture control materials | DIF | Clinical toxicology control material. | Toxicology | Toxicology | 510(k) | Multi-Analyte Urine DAU Control (15 mL);Level 2, For in-vitro diagnostic use only.Lin-Zhi International, Inc.670 Almanor Ave.Sunnyvale, CA 94085Multi-Analyte Urine Drug of Drug Abuse (DAU) Controls are for use as controls in the quantitative and semi-quantitative calibration of Amphetamines Enzyme, Opiate Enzyme, Phencyclidine Enzyme, Cocaine Metabolite Enzyme, Benzodiazepine Enzyme, Barbiturate Enzyme, Methadone Enzyme and Propoxyphene Enzyme on a number of automated clinical analyzers. | MultiAnalyte Urine DAU Control (15 mL); | 2 | January 23, 2012 | 2012 | Lin-Zhi International Inc | Customer complaint investigation found that the Multi-Analyte Urine DAU Level 2 control was incorrectly labeled and actually contained Multi-Analyte DAU Low Level Calibrators. | PRODUCTION CONTROLS: Error in Labeling | Lin Zhi International, Inc. sent a Recall notification letter dated August 5, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to determine whether they wanted credit or a replacement product. The incorrectly labeled product should be disposed of properly.For questions regarding this recall call 408-732-3856. |  | 107 | 107 items kitted | Nationwide Distribution including MA, FL, NC, NH, WA, and IN. | Terminated | January 23, 2012 | 0 | 107 | Not_Computer
Z-0837-2012 | 58551 | antibody igm,if, cytomegalovirus virus | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | SeraQuest***CMV IgM TEST SET For the detection of CMV IgM Antibodies 96 Test Set Cat. No. 01-150 Store 2-8 C For In Vitro Diagnostic Use***Contents: 1 96 wells bag with holder, 1 Negative Control*, 1 Positive Control*, 1 Substrate, 1 Stop Solution, 1 Conjugate*, 1 Calibrator 1*, 1 Calibrator 2*, 1 Diluent for Specimen*, 1 Wash Concentrate*, *Contains Sodium Azide Lot No. E1700 Doral, FL 33122 USA Exp. Date: May 11.BIORAD***CMV IgM EIA For the detection of CMV IgM Antibodies Xn-HARMFUL Cat. No. 25178 96 Test Set Store 2-8 C ID: grey For In Vitro Diagnostic Use*** Contents: I CMV IgM Microplate (96 wells), 1 Negative Control* (0.6 mL), 1 Positive Control (0.6 mL), 1 Substrate (12 mL), 1 Stop Reagent (12 mL), 2 Conjugate* (12 mL), 1 Calibrator 1* (0.6 mL), 1 Calibrator 2* (0.6 mL) , 1 Diluent* (25 mL), 1 Wash Concentrate* (30 mL), *0.1% Sodium Azide Lot No. E1700 Distributed by: Bio-Rad Laboratories-Redmond, WA 98052. | SeraQuest CMV IgM | 3 | January 23, 2012 | 2012 | Quest International, Inc. | Quest International recalled their CMV IgM Test kit because the device failed the test run due to lack of activity in the coated wells. | PRODUCTION CONTROLS: Storage | On September 28, 2010, Quest International, Inc. issued a Memo to its Customers who received the CMV IgM Lot E1700. The memo identified the affected product and stated the problem with the device. The memo asked customers to let the firm know if they had any remaining inventory from the affected lot and the product would be replaced immediately. Questions should be directed to 305-592-6991. | No consumer action necessary | 315 | 315 | Nationwide Distribution -- Washington and California. | Terminated | May 31, 2012 | 129 | 315 | Not_Computer
Z-0841-2012 | 50776 | device, neurovascular embolization | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Axium Detachable Coil System Model Number: QC-2-2-3D.The AXIUM DETACHABLE COIL SYSTEM consists of a platinum embolization coil attached to a composite implant delivery pusher with a radiopaque positioning marker and a hand-held I.D. (instant detacher) which when activated detached the coil from the delivery pusher tip. Some of the AX IUM Detachable coils are enlaced with PGLA microfilaments or Nylon microfilaments. I.D. (instant detached) is sold separately. | The AXIUM DETACHABLE COIL SYSTEM | 3 | January 24, 2012 | 2012 | Micro Therapeutics Inc, dba ev3 Neurovascular | ev3 Neurovascular has identified an incorrect expiration date on the label of one lot of ev3 Axium coils Catalog QC-2-2-3D. This lot 5871057 shows an expiration date of 0211-04. The expiration date should read 2011-04. There are no functional issues with the coils in question as the appropriate expiration date is 2011. | TRAINING: Employee Error | ev3 Neurovascular sent a "NOTIFICATION OF RECALL" letter dated November 6, 2008 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to stop use and segregate the affected product from inventory. A reply form was attached to the letter for customers to complete and fax to Customer Service at (763)398-7001. Customer Service will issue an RGA number for the return of the affected product. Contact the Quality Systems Manager at (949) 680-1305 for questions regarding this notice. |  | 16 | 16 | Worldwide Distribution-USA (nationwide) including the states of CA, CT, NE, NV, and WI. and the countries of Argentina, Brazil, Colombia, and Egypt. | Terminated | January 26, 2012 | 2 | 16 | Not_Computer
Z-0846-2012 | 60750 | enzyme immunoassay, barbiturate | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Barbiturate Enzyme Immunoassay; R1Antibody/Substrate Reagent (5000 tests and 500 tests);Contains a mixture of monoclonal and polyclonal anti-barbiturate antibody, glucose 6-phosphate (G6P) and nicotinamide adenine dinucleotide (NAD), stabilizers, and with sodium azide as preservative.and Barbiturate Enzyme Immunoassay R2;Enzyme Conjugate Reagent (5000 tests and 500 test packaging);Contains glucose-6-phosphate dehydrogenase (G6PDH) labelled with barbiturate in buffer with sodium azide as preservative.For in-vitro diagnostic use only.Lin-Zhi International Inc.670 Almanor Ave.,Sunnyvale, CA 94085Product Usage: Drug testing | Barbiturate Enzyme Immunoassay | 2 | January 24, 2012 | 2012 | Lin-Zhi International Inc | Catalog #0140 and 0141 Barbiturate EIA small and large kits were not correctly made to specification such that the two antibodies used in production did not have curves that correctly aligned. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Lin-Zhi sent a letter via e-mail dated September 8, 2011 to customers with information regarding the recall. The e-mail identified the product description, lot number, description of nonconformity and corrective actions to be taken. Each customer and distributor was then contacted by regulatory affairs personnel by phone on September 9, 2011 or September 12, 2011. Customers and distributors were contacted to determined whether they wanted a replacement of the affected product. Customers that wished for replacements of the affected product was asked to correctly dispose of the product with the lot numbers listed. For questions or concern call 408-732-3856. | Contact the recalling firm for information | 84 | 72 small kits, 12 large kits. | Worldwide Distribution - USA (nationwide) including the states of: AL, CA, FL, IN, MA, NC, NH, TX and the countries of: Germany and United Kingdom. | Terminated | January 25, 2012 | 1 | 84 | Not_Computer
Z-0843-2012 | 60682 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The SternaLock Blu Implant Tray includes plates, screws, screw measuring device, and blades. The SternaLock Blu Implant Tray would have been found within the following outer containers:Silver Instrument Tray, with a White rack, and there is no Power Driver warning. Labeled in part: "Manufactured for Biomet Microfixation***Biomet Microfixation 1520 Tradeport Drive Jacksonville, FL 32218 USA www.biometmicrofixation.com***#73-1301***STERNALOCK IMPLANT TRAY***256X 125.5 MM (10.080 x 4.940 ") (LXW)***Polypropylene, Aluminum, Stainless Steel***"A 3" tall container, with blue rubber handles, and a Power Driver warning on the container lid. Labeled in part: "Manufactured for Biomet Microfixation***Biomet Microfixation 1520 Tradeport Drive Jacksonville, FL 32218 USA www.biometmicrofixation.com***#73-1302***STERNALOCK BLU IMPLANT TRAY SEMIMODULAR***256X 125.5 MM (10.080 x 4.940 ") (LXW)Polypropylene, Aluminum, Stainless Steel***"A 5" tall container, with a wide metal latch to secure the lid, and there is no Power Driver warning. Labeled in part: "Manufactured for Biomet Microfixation***Biomet Microfixation 1520 Tradeport Drive Jacksonville, FL 32218 USA www.biometmicrofixation.com***#73-1300***STERNALOCK BLU IMPLANT SEMI MODULAR TRAY***228.6X 128.8MM (9.00 x 5.07 ") (LXW)Polypropylene, Aluminum, Radel***".The Biomet Microfixation SternaLock Blu System is intended for use in the stabilization and fixation of fractures of the anterior chest wall including sternal fixation following sternotomy and sternal reconstructive surgical procedures to aid in the alignment and stabilization of bone. There SternaLock instrument tray has been designed specifically to store the implants for this system. | SternaLock Blu Implant Tray | 2 | January 24, 2012 | 2012 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc., Jacksonville, FL recalled their SternaLock Blu Implant Tray on 11/23/11, model numbers 73-1300 through 73-1302 due to the possibility of an inaccurate reading on the Screw Measuring Area. If used incorrectly, the screw measuring area inside the SternaLock Blu Implant Tray, could lead to the improper selection of a screw 2mm longer than needed. | DESIGN: Device Design | Biomet Microfixation, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides a "Recommended Technique" to be followed to correct the problem until replacement trays are available. Contact the Quality Improvement Manager at 1-800-874-7711 for questions regarding this notice. | No consumer action necessary | 109 | 109 trays | Worldwide Distribution-- USA (nationwide) including the states of AL, AZ, CA, CO, DE, FL, GA, HI, IL, IN, KY, MD, MO, MS, NC, NE, NJ, NY, OK, OR, PA, SC, TN, TX, WA, and WI. and countries of Argentina, Australia, Canada, Denmark, Finland, Germany, South Africa, and Spain. | Terminated | April 03, 2012 | 70 | 109 | Not_Computer
Z-0842-2012 | 60475 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems CAPIOX (r) SX18 Hollow Fiber Oxygenator LOT: HF22 ÂŽ. 2010-04 The X-coated Capiox SX18 Hollow Fiber Oxygenators with Detachable Hardshell Reservoir are used to exchange gases between extracorporeal blood and a gaseous environment during cardiopulmonary bypass procedures. The integral heat exchanger is used to warm or cool the blood or perfusion fluid and the detachable hardshell reservoir is used to store blood during extracorporeal circulation from both the venous line and the cardiotomy line. The reservoir contains a venous section that is comprised of a filter and defoamer to facilitate air bubble removal. | CAPIOX SX18 Hollow Fiber Oxygenator w/Detachble Hardshell Reservoir | 3 | January 24, 2012 | 2012 | Terumo Cardiovascular Corporation | Complaints were received of expired X-coated Capiox SX18 Hollow Fiber oxygenators. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | All customers were contacted by telephone between 6/28/10 and 7/2/10. The manufacturer mailed "Urgent Medical Device Removal " letters to customers on 01/16/12 with receipt confirmation requested. Customers were informed that expired product had been shipped. Use of expired product was not recommended and product return was requested for refund. Contact Terumo CVS Customer Service 1-800-521-2818 for questions. | No consumer action necessary | 16 | 16 units | Nationwide Distribution including states of: WI, MO,and KY. | Terminated | February 29, 2012 | 36 | 16 | Not_Computer
Z-0856-2012 | 60701 | electrode, electrodcardiograph, multi-function | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | R2 Adult R2 Non-Radiolucent Multifunction Electrodes, Catalog/REF No. 3200-1715, NON-STERILE, MADE IN USA, CONMED CORPORATION, 525 FRENCH ROAD, UTICA, NY 13507This device, when connected to a monitor/defibrillator has the ability to monitor the electro-cardio function (ECG) of the patient. The device, when electrically activated by the monitor/defibrillator can pace the heart rate of the patient and may also transmit the electrical current to the patient for cardio-conversion. This is a non-sterile, disposable device. | R2 Adult R2 NonRadiolucent Multifunction Electrodes | 2 | January 25, 2012 | 2012 | ConMed Corporation | The specific lot was assembled without the rivet cover. | TRAINING: Employee Error | ConMed Corporation sent a Urgent Medical Device Recall Letter dated March 25, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately stop use of the device and return it to ConMed Corporation. Please review your inventory for any of the devices listed on Attachment I. If you do not have any devices to return, please follow the instructions in the letter and complete Attachment II, indicating you have no devices and fax it to 315-624-3225. Please complete the attached Business reply form. We apologize for any inconvenience this will cause you or your staff.For further questions please call (315) 797-8375. | Contact the recalling firm for information | 850 | 850 sets (2 pads per set in individual pouch) | Worldwide Distribution -- USA (nationwide) and the countries of Canada, Belgium, Germany and Hong Kong. | Terminated | January 26, 2012 | 1 | 850 | Not_Computer
Z-0851-2012 | 60858 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-Knee II Durasul All-Poly Patella, Size 3, 7mmThe Zimmer Natural-Knee II Durasul Patella product is a semi-constrained total knee prosthesis consisting of four anatomically designed components: the femoral, tibial, articular surface, and patellar prosthesis. The all-polyethylene patella is made from Durasul Highly Crosslinked Ultra High Molecular Weight Polyethylene. The N-K II Durasul polyethylene patella components are indicated for patient conditions of non-inflammatory degenerative joint disease (NIDJD), e.g., avascular necrosis, osteoarthritis, and inflammatory degenerative joint disease (IJD), e.g., rheumatoid arthritis; correctable valgus-varus deformity and moderate flexion contracture; those patients with failed previous surgery where pain, deformity, or dysfunction persists; and revision of previously failed knee arthroplasty. | NaturalKnee II Durasul AllPoly Patella, Size 3, 7mm | 2 | January 25, 2012 | 2012 | Zimmer Inc. | The firm has 16 reports of patella peg shear out of 109,386 distributed units (0.0137%). The failures have been reported where the pegs on the back side of the patella have sheared at the peg/patella interface. Analysis shows that the highly cross-linked polyethylene, Durasul, patella was shown to have a lower fracture fatigue strength than the Sulene counterpart and that the cement technique used | DESIGN: Component Design/Selection | Zimmer sent an Urgent Device Recall letter dated December 20, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop distribution of the affected product and return all lots in inventory. Customers should provide a list of all hospitals and surgeons that have used the affected product by completing the User Facility and Health Care Professional Form and fax a copy to (574) 372-4265. If customers distributed the product further they should provide a copy of the recall letter to those customers.For questions regarding this recall call 1-877-946-2761. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Germany, France, Austria, Sweden, Switzerland, South Africa, and the UK | Open |  |  |  | Not_Computer
Z-0860-2012 | 60789 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Aristos FX Plus solid state x-ray imager.Image-intensified fluoroscopic x-ray system. | Siemens AXIOM Aristos FX Plus | 2 | January 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Fixation bolts for the mounting of the X-ray tube holder became mixed with an incorrect bolt type. | PRODUCTION CONTROLS: Process Control | Siemens Healthcare sent a "SAFETY ADVISORY NOTICE" dated December 12, 2011 to the affected customer. The letter identifies the product, problem, and actions to be taken by the customer. The firm prepared a modification, which will be provided to the affected customer under Update Instruction XP055/11/S. The firm will resolve this potential issue by preventively exchanging all bolts for the fixation of the X-ray tube holder of the concerned AXIOM Aristos FX Plus unit. |  | 1 | 1 | Product distributed in one hospital in California. | Terminated | January 14, 2014 | 720 | 1 | Not_Computer
Z-0847-2012 | 60757 | calibrators, drug mixture | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Multi-Analyte Urine DAU Calibrator (15 mL);Multi-Analyte,Intermediate Calibrator;for in-vitro diagnostic use only.Contains 300 ng/mL benzoylecgonine, 750 ng/mL d-methamphetamine, 4000 ng/m: morphine, 50 phencyclidine, 500 ng/m: oxazepam, 500 ng/mL secobarbital, 700ng/mL propoxyphene and 600 ng.mL methadone in urine buffer with dosium azide as preservitive.Drug abuse testingLin-Zhi International, Inc.'670 Almanor Ave.Sunnyvale, CA 94085. | MultiAnalyte Urine DAU Intermediate Calibrator | 2 | January 25, 2012 | 2012 | Lin-Zhi International Inc | Catalog # 0804 Multi-Analyte Urine DAU Intermediate Calibrators were incorrectly labeled and mistakenly filled with Catalog # 0805 DAU High Calibrators. | PRODUCTION CONTROLS: Labeling Mix-Ups | Lin -Zhi International sent a Recall letter dated October 11, 2011, and a e-mail to all affected customers. This was followed up by a personal phone call from Lin-Zhi personnel. The letter identified the product the problem and the action needed to be taken by the customer. The customer will be contacted to determine whether or not they were affected by this non-conformity. If they have been affected, they will be offered replacement Multi-Analyte Urine Intermediate Calibrators which are currently estimated to have a four week wait time before they become available.For further questions please call (408) 732-3856. | Contact the recalling firm for information | 49 | 49 items | Worldwide Distribution -- USA (nationwide) including the states of NC,MA, NH and TX. and the country of Germany, | Terminated | January 25, 2012 | 0 | 49 | Not_Computer
Z-0868-2012 | 60743 | calibrators, drug mixture | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Multi-Analyte Urine DAU Calibrator (15 mL HDPE bottle inside a cardboard box)For in vitro diagnostic use only.Lin-Zhi International, Inc..670 Almanor Ave.Sunnyvale, CA 94085intended use: Drug abuse detection. | MultiAnalyte Urine DAU Cutoff Calibrator | 2 | January 25, 2012 | 2012 | Lin-Zhi International Inc | Multi Analyte Urine DAU Cutoff Calibrator Gives a decreased absorbance reading due to a loss of the BZO (oxazepam) analyte from the solution and may cause false positive readings. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Lin-Zhi International, Inc., sent a notice by email dated August 23, 2011 to all customers. The notice described the product, problem and actions to be taken. Lin-Zhi will contact the customers to determine whether they want replacement BZO calibrators and controls immediately or if they are willing to wait a minimum of four weeks for a new lot of the Multi-Analyte Urine Cutoff Calibrators. The customers were instructed to contact Lin-Zhi with their replacement request.The notice also stated that the Lin-Zhi has plans to create more stringent QC standards and further R&D; regarding stabilization of this analyte in solution in its calibrators of controls containing BZO.If you have any questions or concerns regarding how this situation is being handled or the outcome, contact Lin-Zhi at (408)732-3856; email: customerservice@lin-zih.com or webpage: www.lin-zhi.com. | Contact the recalling firm for information | 45 | 45 items | Worldwide distribution: USA (nationwide) including states of: AL, FL, GA, IN, NC, NH, and TX; and country of: Germany. | Terminated | January 25, 2012 | 0 | 45 | Not_Computer
Z-0863-2012 | 60180 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Gas cylinders used in Gunnell manual wheelchairs with dynamic recline feature, including: Rehab Kidster( Kidster). Product Usage: The purpose of these wheelchairs is to provide mobility to individuals limited to a seating position. They may also be used as attendant propelled transport devices in a health care environment such as a hospital, nursing home or extended care facility. | Rehab Kidster ( Kidster) | 2 | January 25, 2012 | 2012 | Gunnell, Inc. | Gunnell has received complaints that the gas cylinders used in Gunnell manual wheelchairs would not hold a firm position, or as described by users, the back support or canes were spongy in the Rehab Tough and Tilt (TNT), Bariatric Rehab Tough and Tilt (BTNT), Rehab Recline and Mobility (RAM), Bariatric Rehab Recline and Mobility (BRAM), and Rehab Kidster. | DESIGN: Component Design/Selection | Gunnell, Inc. sent a Medical Device Correction letters dated November 1, 2011 to all consignees. The letter identified the product, problem, and actions to be taken. The letter states that Gunnell Inc will provide free replacement cylinders and reimburse for each correction verified. The notice instructs customers to complete and return the Response Form. Customers were instructed to notify their customers by including a copy of the correction notice. Question and concerns should be addressed to andy.evancho@gunnell-inc.com. or call 1-800-551-0055, attention Andy Evancho. |  | 23 | 23 (22 listed) | Worldwide Distribution - USA (nationwide) nad the countries of: Australia and Canada | Terminated | August 29, 2013 | 582 | 70 | Not_Computer
Z-0862-2012 | 60586 | mixer, cement, for clinical use | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker PCD Precision System 13G Needle with Spineplex, sterile, single use, disposable , 4 units per box. units per box. Stryker Instruments, Kalamazoo, MI Part Number: 0505-585-000, Lot:1122301213G Match Ground Four facet Tip Introduction Needle - PIN 0306-130-01113G Bevel Tip Stylet - PIN 0306-131-012The intended use of the Stryker PCD System is for mixing bone cement and delivering it percutaneously. Bone cement is mixed in the chamber then transferred directly into an extension nozzle which has been attached to the introduction needle in the fracture site by a luer lock fitting. Cement is then dispensed directly into the fracture site under fluoroscopy | Stryker PCD Precision System 13G Needle with Spineplex | 2 | January 25, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Product was packaged with a four inch needle in stead of the correct five inch needle. There is a potential that the needle may not be able to reach the fracture site, which may require additional medical intervention. In addition, if the stylet needs to be changed and the additional stylet that comes in the package is used, this additional stylet is one inch longer than the needle. which could l | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Stryker sent an Urgent Recall Notification letter dated December 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. If product was further distributed, customers should forward the recall notification letter and the attached Business Reply Form (BRF) to all affected locations. Immediately check inventory and quarantine any affected product found. Mark product as "RECALLED PRODUCT". Indicate on the BRF the quantity of affected PCD kits they are returning and fill in the form completely. Fax the Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762. Upon receipt of the BRF, Stryker will contact them to coordinate the return of all affected product that they have on hand. For questions regarding this recall call 269-389-4354. . |  | 11 | 11 kits of the 192 manufactured contain the wrong needle size. | Nationwide Distribution including FL, PA, NY, OH, IL, KY, MT, WA, CA, MO, HI, IN, SC, MS, WV, KS, TX, and WY | Terminated | August 14, 2012 | 202 | 11 | Not_Computer
Z-0861-2012 | 60342 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 5100-03-707 -- MOST Option STEM DISTRACTORMATL: SS NONSTERILE QTY-1Designed to facilitate intraoperative decisions with a variety of options for femoral and tibial components, implant connections and fixation of the implants. | 510003707 Most Option Stem Distractor | 2 | January 25, 2012 | 2012 | Zimmer Inc. | Fractured tips: Any use of force in a certain direction to the flat side of the instrument may result in fracturing of the mating geometry, creating the potential for metal debris at the surgical site. | MISBRANDING: Labeling False and Misleading | Urgent Device Correction Letters, dated October 28, 2011, which included revised Surgical Technique to be used in the Zimmer Most Option Wedge Stem Distractor were sent November 15, 2011 to all Distributors requiring Correction letters be sent to all Hospital Risk Managers and Surgeons receiving devices. Directions were provided to destroy all previous versions of the surgical technique, and review the attached revised surgical technique, with warnings of proper direction of force needed for safe separation of mated components. A FAX -Back Certificate of Acknowledgement is to be sent to (574) 372-4265. Call Zimmer at 1-877-946-2761 with questions or concerns. | Contact the recalling firm for information | 70 | 70 | Worldwide Distribution -- US, including states of TX, IN, GA, WI, MN, NJ, MO, IL, VA, MI, PA, NC, OR, FL, KS, NY, AZ, CO, and UT and country of Switzerland. | Terminated | September 24, 2013 | 608 | 70 | Not_Computer
Z-0873-2012 | 55696 | device, neurovascular embolization | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Axium Detachable Coil System QC-2-1-Helix The ev3 Axium Detachable Coil System is indicated for the endovascular embolization of intracranial aneurysms and other neurovascular abnormalities, such as arteriovenous malformations and arteriovenous fistulae. | Axium Detachable Coil System QC21Helix | 3 | January 26, 2012 | 2012 | Micro Therapeutics Inc, dba ev3 Neurovascular | ev3 has initiated a voluntary removal of two lots of Axium Detachable Coil System for product mislabeling as the lots were swapped in production. | TRAINING: Employee Error | Micro Therapeutics Inc, dba ev3 Neurovascular performed product recovery according to the following outline.a) Quality Assurance Department (QAD) was used to provide details of product distribution for the affected lots.b) Based on the Health Hazard Assessment, the risk was considered undesirable. c) The affected product was part of a limited market release for Physician Preference Testing. As the product was within control of ev3 employees, no public notification was performed.d) Quality Systems contacted the sales representatives identifying the part number, lot number and number of units sent, and a request to segregate and return the affected product was initiated. e) Product returned to ev3 Neurovascular was held in Restricted Stock pending complete device accountability and disposition.f) The ev3 Neurovascular Quality Assurance department scrapped the affected product.g) The ev3 Neurovascular Regulatory Affairs department was responsible for assuring documentation and appropriate notification to regulatory bodies as required. |  | 13 | 13 units | Worldwide Distribution - USA (nationwide) and the countries of the UK, and France. | Terminated | January 31, 2012 | 5 | 27 | Not_Computer
Z-0875-2012 | 58935 | negative pressure wound therapy powered suction pump | OMP | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ***REF 66800694***Qty 10***smith&nephew;***RENASYS Port Foam Dressing Kit Port***STERILE***Rx Only.***Made in Mexico for/Fabrique au Mexique pour: Smith & Nephew, Inc., St. Petersburg, FL 33716***Smith & Nephew Medical Limited Hull HU3 2BN England***Lot 123456.Product Usage: The product is a suction device used as an accessory to the foam dressing kits, which is used in conjunction with Smith &#38; Nephew Negative Pressure Wound Therapy (NPWT) systems to deliver negative pressure to a wound. | Renasys Port | 2 | January 26, 2012 | 2012 | Smith & Nephew Inc. | On 12/14/2010, Smith & Nephew, Inc. Wound Management Division in St. Petersburg, FL recalled their RENASYS PORT product code number 66800694 due to small holes in the primary pouch of some of the packaged units of which constituted a breach of the sterile barrier. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smith & Nephew, Inc., Wound Management sent an URGENT DEVICE RECALL letter dated December 14, 2010, to all affected customers. The letter identified the product, problem, and actions to be taken. Customers were instructed to immediately identify and isolate the recalled lots in their inventory, complete and fax the response form, and return any affected products to APL Logistics, Inc. in Suwanee, GA. For questions in general, please contact the Customer Care Center at 1-800-876-1261. |  | 71329 | 71329 units | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, GA, MA, MI, MO, MS,.NC, NY,OR, TN, and WI. and the countries of: Australia, Canada, China, England, Germany, Ireland and Singapore. | Terminated | April 30, 2012 | 95 | 71329 | Not_Computer
Z-0913-2012 | 60812 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Visum LED Surgical Lights, Ceiling Mounted Surgical Light. The Visum LED surgical light is intended to illuminate the operative site during surgical procedures with high intensity light. | Visum LED Surgical Lights | 2 | January 26, 2012 | 2012 | Stryker Communications Corp | Stryker has identified a potential structural issue with the Visum LED surgical light(s). During a complaint investigation, Stryker became aware that there was a possibility that incorrect assembly could result in the light head beginning to separate from the suspension arm. In the event the light separates and falls from the arm, it could lead to serious injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm sent "Urgent: Medical Device Correction" letters dated 11/07/2011 to their domestic and international customers. The letter informed customers of the remote risk of the light detaching from the arm and recommended continued use of the product. A representative will contact them about scheduling an appointment to inspect the lights. Contact Technical Services at 1-866-841-5663 for questions. |  | 1143 | 1143 Devices | Worldwide Distribution - USA, and the countries of Australia, Canada, Hong Kong, Brazil, Sweden, India, Japan, Korea, Ecuador, Venezuela, Mexico, Denmark, Great Britain, and New Zealand. | Terminated | August 17, 2012 | 204 | 1143 | Not_Computer
Z-0803-2012 | 60270 | [6 empty] | Salter Labs Bubble humidifier, with 6 PSI safety valve, 350cc, Latex Free; Single Patient use; used with supplemental oxygen therapy. Model Numbers 7600-0-50, E7600-0-50, NP7600-0-50 and special order patient kits that contain 7600 humidifiers. Patient kits have brand names of Salter Labs, AdvaCare; Allcare Medical; Memorial Home Services; Therapy Support Inc; Mendo-Lake Home Respiratory Services; MedAssurance Inc.; Abundant Home Care.Product Usage: Humidifier - single patient use. used with supplemental oxygen therapy. | Salter Labs Bubble humidifier | 1 | January 26, 2012 | 2012 | SALTER LABS | A manufacturing defect was discovered internally which indicates that in a small number of humidifier lids, the pop-off may not occur at the appropriate pressure and can result in a leak. | PRODUCTION CONTROLS: Process Control | Salter Labs contacted each customer directly on October 21, 2011 to inform them of the recall. Upon speaking to the direct customers, Salter Labs also provided a Recall Customer Reply Form for each direct customer to complete and return by fax to 661-854-6816 or 1-800-628-4690. Customers were instructed to immediately examine their inventory of the recalled product, remove and quarantine all affected products, indicate if product would be returned and asked if product was further distributed. If so, the second page of the response form was to be completed and returned. Customers were informed that a representatives from Salter Labs can assist, if needed, in returning all affected products to Salter and completing the attached forms. Salter's direct customers are instructed to contact their customers, should the product have undergone further distribution. | Contact the recalling firm for information | 163 | 6,060 units ( 163 cases) | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, IL, MO, NJ, OH, OK, TX and the country of: Italy | Terminated | February 15, 2012 | 20 | 6060 | Not_Computer
Z-0871-2012 | 60823 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components: a) Silastic Tubing, 0.062 ID x 0.032 Wall x 50 Ft, component C-602-285, vendor part number SFM3-2850; b) Silastic Tubing, 0.040 ID x 0.023 Wall x 50 Ft, component C-SFM3-2050, vendor part number SFM3-2050; c) Pathology Container with Lid, 8 oz., component 3701Z, vendor part number 3701; d) Silicone Tubing, 0.187 ID x 0.313 OD x 50 Ft., component C-31010SG, vendor part number 31010; | Presource Custom Sterile Surgical and Procedure Kits | 2 | January 26, 2012 | 2012 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturer(s) does not have all necessary FDA registrations or filings needed for this product. | PREMARKET APPROVAL: No Marketing Application | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated December 12, 2011 to the affected hospital consignees via UPS next day air. The letter informed them that their Presource kits listed in the attached spreadsheet contain various components also listed on the attachment that are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels, instructing the user not to use the affected component upon opening the kit, to be affixed to the front side of each affected Presource kit and shipping carton in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact Presource Sales Operations at 800-766-0706. | Contact the recalling firm for information | 7213 | 7,213 each | Nationwide Distribution -- California, Georgia, Illinois, Massachusetts, Michigan, New York, North Carolina, Ohio, Pennsylvania, Tennessee, Texas and Wisconsin. | Open |  |  | 7213 | Not_Computer
Z-0884-2012 | 60909 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | AngioDynamics MorpheusÂŽ Smart PICC CT 55cm 6F Dual Lumen with REVERSE TAPER --- Max CT Flow Rate: 5mL/SEC --- Injector Pressure Limit: 300 PSI --- Catalog No. [REF] 12102617 --- Made in the USA --- [Manufacturing symbol] AngioDynamics, 603 Queensbury Avenue, Queensbury, NY USA 12804indicated for short or long term peripheral access to the central venous system for intravenous therapy and power injections of contrast media. | AngioDynamics Morpheus Smart PICC CT | 2 | January 26, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The red luer on the Morpheus CT PICC (Peripherally Inserted Central Catheter) has the potential to crack during use of the product. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | AngioDynamics, Inc. sent an Urgent - Medical Device Recall letter dated January 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and segragate the recalled affected product.Complete the enclosed Morpheus Ct PICC Recall Reply Form and fax it to the attention of the Morpheus CT PICC Recall Coordinator at 518-798-1360. If customers do not have any affected product remaining in their possession they should fill out the attached form noting zero quantity to be returned and fax the form to 518-798-1360. If customers distributed the affected product, they should inform their customers of this recall action immediately and have them return the recalled units to them. For any questions regarding this recall call 1-800-772-6446 or e-mail customerservice@angiodynamics.com. | Contact the recalling firm for information | 85 | 85 units | Nationwide | Open |  |  | 10392 | Not_Computer
Z-0917-2012 | 60799 | calibrator, multi-analyte mixture | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROSÂŽ Chemistry Products DT Calibrator Kit, REF/Catalog No. 195 7927, IVD --- Common/Usual Name: VITROS DT Calibrator Kit -- Ortho-Clinical Diagnostics, Inc. Rochester, NY -- Ortho-Clinical Diagnostics High Wycombe, UK VITROS Chemistry Products DT Calibrator Kit is an in vitro diagnostic specially formulated for use as calibrators for ALB, ALKP, ALT, AMYL, AST, TBIL, NBIL, BUN/UREA, Ca, CHOL, CK, Cl-, CO2, CREA, CRSC, Fe, GGT, GLU, HDLC, K+, LAC, LDH, LIPA, Mg, Na+, NH3, PHOS, TP, TRIG, urCR, and URIC on VITROS DT Chemistry Systems. | VITROS Chemistry Products DT Calibrator Kit | 2 | January 27, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) was notified by customers of negatively biased quality control results obtained while using VITROSÂż Chemistry Products NBIL DT Slides, GEN 72. OCD performed an internal investigation and confirmed that negatively biased NBIL results could occur for quality control, proficiency and patient samples. As a result, VITROSÂż Chemistry Products DT Calibrator Kit Lots 37 & | DESIGN: Component Design/Selection | Ortho Clinical Diagnostics (Johnson & Johnson Company) sent an Important Product Correction Notification letter dated October 13, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to recalibrate and instructing them to recalibrate VITROSÂŽ Chemistry Products NBIL DT Slides, GEN 72 with updated DT calibrator values for the NBIL assay provided on CDM 0164. J&J; Foreign affiliate consignees were notified by e-mail on October 13, 2011, informing them of the issue and of the availability of the updated CDM (0164).Consignees were instructed to complete and return the attached Confirmation of Receipt form upon receipt of the notification even if they do not use VITROSÂŽ NBIL Slides in their facility.If the affected product was distributed outside the consignees facility they were instructed to forward the notification letter to affected customers.For any questions regarding this recall call 1-800-421-3311. | No consumer action necessary | 12750 | 12,750 units (Lot 37: domestic - 1403, foreign - 3488; Lot 38: domestic - 50, foreign - 30) | Worldwide Distribution - USA (nationwide) Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, and Venezuela | Open |  |  | 12750 | Not_Computer
Z-0915-2012 | 60685 | dental cement w/out zinc-oxide eugenol as an ulcer covering for pain relief | MZW | Dental cement. | Dental | Dental | 510(k) | Brand name: Maxcem Elite, Product: a dental cement,Part Number: 34418. The intended use of this device is for cementation of all indirect restorations including ceramic, resin and metal-based inlays, onlays, crowns, bridges, posts and veneers. Additional Indications include core-buildup material, pit and fissure sealant and cementation of crown restorations to implants | Maxcem Elite, a dental cement | 2 | January 27, 2012 | 2012 | Kerr Corporation | The recall was initiated because Kerr Corporation has confirmed that the use of the affected lot of Maxcem Elite may not provide users with a minimum of 2 minutes of workÂżtime thereby causing the product to prematurely set by 15 seconds. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Kerr Corporation sent a Urgent Medical Device Recall letter dated March 14, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. It has recently come to our attention that use of the affected lot of product may not provide users with a minimum of 2 minutes of workÂŽtime thereby causing the product to prematurely set by 15 seconds. If this were to happen prior toplacing restorations on the tooth or core build-up the restorations may not adhere properly. If you have any of the affected product listed above please contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. Please fax back the enclosed Recall Return Form in order to confirm your receipt of this recall notification regardless of whether you have any product in your inventory. |  | 67 | Four hundred and sixty-three (463) | Worldwide Distribution -- USA (nationwide) and the countries of Canada, Australia and New Zealand. | Terminated | January 31, 2012 | 4 | 463 | Not_Computer
Z-0921-2012 | 60811 | urinary homocystine (nonquantitative) test system | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homogeneous Enzymic Homocysteine Reagent Two Part Reagent Method. Catalog Number E02-057-02.A homocysteine test system is an in vitro device intended to measure total homocysteine quantitatively in serum or plasma. Homocysteine measurements are used in the diagnosis and treatment of hyperhomocysteinemia. | Homogeneous Enzymic Homocysteine Reagent Two Part Reagent Method | 2 | January 27, 2012 | 2012 | Catch Incorporated | Homogeneous Enzymic Homocysteine Reagent Two Part Reagent Method will not calibrate due to Reagent R1 had prematurely decayed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Catch Inc. sent an "URGENT-PRODUCT RECALL" letter dated 27 June 2011 to one sole customer. The letter described the product, problem, and actions to be taken by the customer. A reply form was attached to the letter for the customer to complete and return via fax at 425-402-8954. Please contact the firm at 425-402-8960 for questions regarding this recall. | No consumer action necessary | 47 | 47 kits | Product distributed in PA. | Terminated | July 11, 2012 | 166 | 47 | Not_Computer
Z-0925-2012 | 60845 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components:a) Jobri Circumcision Strap, Small, 2 x 8 inch, component part number 086-0001;b) Jobri Circumcision Strap, Large, 2 x 10 inch, component part number 086-0002;The restraint is used to secure an infant during a circumcision procedure.Product Usage:The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then sterilized and are ready to be used by the customer for the specific procedure. The restraint is used to secure an infant during a circumcision procedure. | Presource Custom Sterile Surgical and Procedure Kits | 2 | January 27, 2012 | 2012 | Cardinal Health | The Presource kits were assembled with a Jobri Infant Restraint Strap which does not have the appropriate 510(k) pre-market clearance. |  | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated December 23, 2011 to the affected hospital consignees via UPS next day air. The letter identified the product, problem, and actions to be taken. Enclosed with the letter were recall labels, instructing the customer not to use the affected component upon opening the kit, to affix recall label to the front side of each affected Presource kit and shipping carton currently in their inventory. The customers were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit/replacement the consignees were instructed to contact the folowing: Hospitals - Presource Sales Operations at 800-766-0706; Federal Government facilities - Government Customer Service at 800-444-1166; and Distributors - Distributor Management Group at 800-635-6021. For any questions, require additional labels or desire special assistance relating to this recall contact Cardinal Health Quality Systems, 800-292-9332. |  | 15369 | 15,369 kits in total | Nationwide Distribution (USA) including the states of: Alabama, Arizona, California, Connecticut, Florida, Georgia, Hawaii, Illinois, Indiana, Kentucky, Maryland, Missouri, North Carolina, Nebraska, New Jersey, New Mexico, New York, Oregon, Pennsylvania, Tennessee, Virginia, Washington and West Virginia. | Open |  |  | 15369 | Not_Computer
Z-0920-2012 | 60844 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The Artificial Bur HypoPack is a Temporary Fascial Prothesis with Hypobaric Wound Dressing. For Single Patient Use only.Contents: 1. Large Bur Patch (Sterile: EO), 20cm x 40cm Loop Sheet, 20cm x 40cm Hook Sheet.2. 60cm x 85 cm Steri Drape3. Drain with Connector4. 4 x Gauze RollsSterile: Contents are sterile for maximum two years unless enclosed package has been opened or damaged.Used for a fascia expander prosthesis to temporarily close the abdominal cavity when regular abdominal closure is not possible; to prevent perfusion impairment of abdominal organs from increases intra-abdominal pressure; and for advanced intra-abdominal pathology requiring more than one abdominal operation to control the disease process. | Bur Hypo Pack | 2 | January 27, 2012 | 2012 | Novomedicus Llc | On 06/22/2011 Novomedicus, FL initiated a recall of the Artificial Bur Hypopack Lot 1010.1, 1102.1, 1103.1 and 5643.2. Cytotoxicity testing revealed a biocompatibility issue. The test results did not represent a health risk. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Novomedicus sent a Medical Device Correction letter dated June 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy the inventory, or return the inventory to Novomedicus. If customers decided to return the remaining inventory, Novomedicus would provide TWO replacement kits for each one returned at no additional cost. Customers were asked to complete the attached Return Response form and return to NovoMedicus using the self-addressed and stamped envelope included. For questions regarding this recall call 941-375-4108. |  | 37 | 37 kits | Wordwide Distribution - USA including IL, AZ, PA, CT, MO, WI, NY, TX, OR, ME, FL, VA, and MS and the countries of Mexico and Qatar. | Terminated | August 28, 2012 | 214 | 47 | Not_Computer
Z-0922-2012 | 60964 | mesh, surgical | FTM | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MEDPOR TITAN Cranial Temporal with Template - Right, Catalog Number 81038. MEDPOR Biomaterial with embedded Titanium Mesh Implants are intended for non-weight bearing applications of craniofacial reconstruction/cosmetic surgery and repair of craniofacial trauma. | MEDPOR TITAN | 2 | January 27, 2012 | 2012 | Howmedica Osteonics Corp dba Stryker Craniomaxillofacial | The template that is packaged with this implant was not packaged according to specification. The template was packaged in a single pouch instead of a double pouch. | PRODUCTION CONTROLS: Packaging Process Control | "Urgent Product Recall" Notification letters and Product Accountability Forms have been sent on January 13, 2012 domestically via FedEx and Internationally via DHL. Customers were informed of the product problem and provided risk mitigation information. Return of the Product Accountability/Acknowledgement form was requested. Questions were directed to 770-254-4423. | Contact the recalling firm for information | 10 | 10 units | Worldwide Distribution - USA, the state of Michigan, and the countries of Australia and Canada. | Terminated | June 21, 2012 | 146 | 10 | Not_Computer
Z-0929-2012 | 60806 | nasopharyngoscope (flexible or rigid) | EOB | Nasopharyngoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Product is a lithium ion battery sold as both a stand-alone accessory, and as part of endoscopy system sets. Stand alone labeling: Vision Sciences Rechargeable Lithium Ion Batteries, Replacement (2) [REF] 07-3052; [MFG Symbol] Vision-Sciences, Inc., 40 Ramland Road South, Orangeburg, NY 10962 USA.Accessory to endoscopy kits used to charge/recharge lithium ion batteries. | Vision Sciences Battery | 2 | January 29, 2012 | 2012 | Vision-Sciences, Inc. | Vision-Sciences is recalling the Vision-Sciences Battery Charger due to reports relating to the potential for over-heating and possible burning of the lithium-ion batteries being charged. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Vision Sciences sent an "URGENT: PRODUCT RECALL NOTIFICATION" letter dated December 27, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to remove the affected product from inventory. A Tracking/Verification Form was provided for customers to complete and return to the firm. Contact the firm at 800-874-9975 for questions regarding this recall. |  | 3628 | 3628 batteries | Worldwide Distribution | Open |  |  | 5246 | Not_Computer
Z-3126-2011 | 58719 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Part Number: 81600Access 2 Immunoassay Systems Part Number: 81600NAccess Immunoassay Systems LXi 725 Part Number: 386200Subsequent Product Codes: JGSThe Access/DxI Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access Immunoassay Systems | 2 | January 30, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because certain Access immunoassays can produce inaccurate results when performed within the published UniCel DxI Immunoassay Systems* room temperature operating specification of 18 degrees C to 32 degrees C (64.4 degrees F to 89.6 degrees F). An increase in room temperature causes assay results for some assays to decrease, while results for other assays increase with an i | OTHER/UNDETERMINED: Pending | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter includes information on the new room temperature operating range for all UniCel DxI Immunoassay Systems and Access Immunoassay Systems during system operation. Customers were instructed to review their historical room temperature records to determine if their Access Immunoassay System and UniCel DxI Immunoassay System were operating outside the revised temperature range. An attached Response Form was included for customers to complete and return.Please contact Beckman Coulter Customer Support Center for questions regarding this notice at 1-800-854-3633. |  | 11266 | 11,266 units | Worldwide Distribution--USA (nationwide) and Puerto Rico and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Lebanon, Liechtenstein, Malaysia, Mayette, Mexico, Morocco, Nambia, Netherlands, New Zealand, Norway, Philippines, Poland,, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, United States, Venezuela and Vietnam. | Terminated | December 21, 2012 | 326 | 14648 | Not_Computer
Z-0930-2012 | 60320 | impactor | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | The SIGMAÂŽ HP Fixed Bearing Tibial Tray Impactor is part of a set of tools designed specifically for the installation of the PFC SIGMAÂŽ Knee. SIGMA HP Fixed Bearing Tibial Tray Impactors Packaged in poly bag with optional protection from sharp or fragile points, distributed by DePuy International Leeds, UK. Instruments are supplied non-sterile and do not have specified expiration date or shelf life.Usage: The Tibial Tray (with protective cover) and instrument are assembled outside of the patient in theatre; the tray is introduced and impacted into the cement containing tibia using the impactor. Once the tray is fully seated in the cement containing tibia, the instrument is intended to be removed. | SIGMA HP Fixed Bearing Tibial Tray Impactor | 2 | January 30, 2012 | 2012 | DePuy Orthopaedics, Inc. | All lots of DePuy, SIGMA HP Fixed Bearing Tibial Tray Impactors are being recalled because the impactor can become intentionally locked on the Tibial Tray implant during surgery. | DESIGN: Component Design/Selection | The firm, DePuy, sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" email dated December 13, 2011 to its customers. The email notice described the product, problem and actions to be taken. The customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN; follow the instructions provided and complete and return the Fax Back Response card via fax to: 574-372-7567 or email: kseppa@its.jnj.com by December 23, 2011.For questions about recall information provided, please contact Manager of Customer Quality at 574-372-7333 (M-F; 8 am - 5 pm EST.) For product-related questions from Sales force, please contact Group Product Director at 574-372-7538 (M-F; 8 am - 5 pm EST.) or your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact DePuy's Scientific Information Office at 1-888-554-2482 (M-F; 8 am - 5 pm EST.). | Contact the recalling firm for information | 1794 | 1794 | Worldwide distribution: USA (nationwide) and countries of: Canada, Ireland and South Africa. | Terminated | July 12, 2013 | 529 | 1794 | Not_Computer
Z-0931-2012 | 60850 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DurationÂŽ packaging distributed byStrykerÂŽ Orthopaedics, Limerick, IrelandStrykerÂŽ IrelandCarrigtwohill Industrial EstateCarrigtwohill County Cork, IrelandProduct Usage:Orthopaedic implants utilized for primary and revision hip and knee (total and unilateral); total shoulder and elbow replacement; replacement; and oncological surgical procedures related to tumor resection. All devices are manufactured from UHMWPE and are processed through post-sterilization (gamma irradiation) heat treatment for stabilization (i.e. DurationÂŽÂŽ process). Launched in 1996, DurationÂŽÂŽ Stabilized Polyethylene is produced by a patented process utilizing a combination of nitrogen packaging to reduce shelf oxidation, gamma irradiation (3Mrads) to sterilize and crosslink the polyethylene, and a post-irradiation stabilization step during which the components undergo a heat treatment process to promote further crosslin king and to enhance product wear resistance. In order to maintain the nitrogen environment over the shelf-life of the product, the packaging consists of double, nested PET blisters sealed with foil (non-gas permeable) Iidstock. | Stryker | 2 | January 30, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | As part of StrykerÂż's quality process, a packaging discrepancy was identified in specific lots of StrykerÂż's DurationÂż packaging. | PRODUCTION CONTROLS: Packaging | Stryker sent an Urgent Product Recall notification letter and Product Recall Acknowledgment Form dated September 6, 2011 to all affected consignees via FedEx with return receipt. The letter identifed the product, problems, potential hazard, and instructions on how to handled the affected product. Consignees are instructed to Return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics. Consignees are advised to completed and fax back the attached Product Recall Acknowledgment Form within 5 days to (201) 831-6069. For questions call (201) 972-2100. |  | 39483 | 39,483 units | Nationwide Distribution (USA) | Open |  |  | 39483 | Not_Computer
Z-0933-2012 | 60968 | cleaner, air, medical recirculating | FRF | Medical recirculating air cleaner. | General Hospital | General Hospital | 510(k) | Product is a 15x16x27" (WxDxH) air purifier device with one of two different brands next to the control panel: HealthWay or CleanStation. HealthWay brand is sold under model number Deluxe 20600-3 Air Purification System. The CleanStation brand is sold under model number CS20000.The devices are intended to remove particles from the air for medical purposes | HealthWay or CleanStation | 2 | January 31, 2012 | 2012 | Healthway Home Products, Inc. | The Company is recalling certain serial numbers of HealthWay Model 20600-3 and CleanStation Model CS20000 air purifiers because of a potential risk of electrical arcing that could cause overheating or fire within the units. | DESIGN: Device Design | Healthway Home Products, Inc. sent letters to distributors and end-users on January 9, 2012. The letter identified the product, the problem, and the action to be taken by the distributor. Distributors were instructed to immediately cease the distribution or sale of any of the HealthWay 20600-3, and CleanStation CS20000 units with the affected serial numbers. Distributors were also instructed to isolate any of these devices currently in their inventory in a secure area until they receive further instructions from the recalling firm. Distributors were asked to compile a list of end-users who purchased the product; this information is being provided to a third-party engaged to contact end-users on behalf of the recalling firm. --- End users are instructed to immediately stop using the product and go to www.airpurifierrecall.com and follow the prompts to enter the model number and serial number of the unit. If the unit is subject to the recall, the user will be asked to provide their name, address, and contact information. Users will receive a follow-up call from HealthWay to discuss the recall process. Customers that have purchased units with the serial numbers set forth below will have 3 options: (1) a retrofit of the unit, adding an arc-suppression circuit and a more robust high energy connector, at no cost; (2) a retrofit of the unit, adding an arc-suppression circuit and a more robust high energy connector, including a basic cleaning of the unit and new filter set for $99.99 (a 60% discount); or (3) a replacement, brand-new unit, including a remote control and 5-year limited warranty for $199.99 (an 85% discount). --- Customers without access to the internet or who to prefer to receive the information by mail may call 855.230.0401. --- On 01/19/2012, the firm began conducting field corrections by disabling the high energy zone on the HealthWay 10600-2 models. End users will be sent letters informing them to immediately stop using the device and contact the f. | No consumer action necessary | 9107 | 9107 units | Worldwide Distribution - USA (nationwide) and Canada | Open |  |  | 9107 | Not_Computer
Z-0932-2012 | 59920 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Solanas Posterior Stabilization System Part Number: 63920.Solanas Posterior Stabilization System facilitates the surgical correction of spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. The Solanas Posterior Stabilization System is a spinal fixation system intended to improve stability of the cervical and thoracolumbar area of the spine (C1-T3). | Solanas Posterior Stabilization System | 2 | January 31, 2012 | 2012 | Alphatec Spine, Inc. | These lots were not manufactured correctly in that the instrument does not have the durability required to repeatedly perform over its full service life and, therefore, may experience breakage in the tips. | DESIGN: Device Design | Alphatec Spine contacted all affected customers via telephone. The communication included discussion of the product, problem, and actions to be taken by the customers. Customers were instructed to remove the affected product from inventory and to return the product to the firm. Arrangements were made for replacement product to be sent to the customers. Contact the firm at 1-800-922-1356 for questions regarding this recall. | No consumer action necessary | 27 | 27 units | Nationwide Distribution-including the states of CA, CT, FL, GA, NC, NE, NJ, TX, and UT. | Terminated | January 31, 2012 | 0 | 27 | Not_Computer
Z-0832-2012 | 60615 | bassinet, hospital | NZG | Pediatric hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Bassinet Models: BSNT/01; BSNT/02, BSNT/03, BSNT/04The Nemschoff Perinatal Bassinet products are wheeled carts with a top tray intended for an infant tray with mattress. The bassinets are intended for medical purposes that consist of a bed designated for the use of a pediatric patient. | Bassinet | 1 | February 01, 2012 | 2012 | Nemschoff Chairs, Inc. | To correct issues with casters, door hinges, drawer slides, and the bassinet surround. Also, to add labeling to the product regarding specific instructions for cleaning the bassinet tub and drawer weight limits. | DESIGN: Device Design | The firm, Nemschoff, sent an "Important Medical Device Correction" letter and email dated October 14, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check inventory, quarantine and hold product until they receive instructions for servicing the units; once corrections are complete on the units at their locations, document all corrections on the Correction Acknowledgment Letter and return to Nemschoff, Inc. as instructed. The customers were also instructed to complete and return the enclosed Receipt Acknowledgment Form by October 21, 2011.Nemschoff will provide instructions and assistance with locating and repairing the units.If you have any questions, call Nemschoff, Inc at (800) 203-8916 and ask for our Quality Representative or your Customer Service Representative. | Contact the recalling firm for information | 1214 | 1214 units. | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, IL, IN, KS, LA,MA, MD, MI, MO, MS, NC, ND, NJ, NY, OH, OK, SC, TX, VA, WA, WI, and WY; and country of: Canada. | Terminated | November 27, 2012 | 300 | 1214 | Not_Computer
Z-0934-2012 | 60911 | system, imaging, optical coherence tomography (oct) | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General & Plastic Surgery | 510(k) | Imalux Niris OCT Imaging System Probe, Model Number 1300, Serial Number F464019, IMALUX Corporation, Cleveland, OHThe Niris Imaging System probe is used with the Niris Imaging System. The Niris Imaging System employs Optical Coherence Tomography (OCT), which can be used to construct high spatial resolution. The probe, is attached to the Imaging Console, and is used to direct light to and from the patient tissue. A small electromechanical scanning mechanism in the probe moves the optical beam laterally across the tissue surface while simultaneously acquiring in-depth backscattering profile at each lateral position. The probe is not labeled or sold as sterile. | Imalux Niris OCT Imaging System Probe | 3 | February 01, 2012 | 2012 | Imalux Corporation | During an FDA inspection, the firm was notified that the wrong tubing was used during the manufacture of the probe for the unit. | PRODUCTION CONTROLS: Process Control | Imalux Corp contacted their one customer via telephone on December 14, 2011. They informed the customer of the problem and requested the probe returned. The probe involved was returned and received by Imalux on December 22, 2011, and will be stored offsite in quarantine.Any further questions please call (216) 502-0755 ext. 301. |  | 1 | One Unit | Nationwide Distribution only to Texas. | Terminated | January 09, 2014 | 708 | 1 | Not_Computer
Z-0938-2012 | 60900 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | KIMBERLY CLARK MIC GASTRO-ENTERIC FEEDING TUBE, 22F.Gastro-Enteric Feeding Tube. | KIMBERLY CLARK MIC GASTROENTERIC FEEDING TUBE | 2 | February 02, 2012 | 2012 | Kimberly-Clark Corporation | The product is misbranded, i.e. the Gastric (center) port may be mislabeled as the Jejunal port, and the Jejunal (side) port mislabeled as the Gastric port. | MISBRANDING: Labeling False and Misleading | Kimberly-Clark Health Care sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter and e-mail to all affected customers. The communications described the product, problem, and actions to be taken by the customers. Customers were instructed to locate and quarantine any affected product. A Product Recall Response Sheet was enclosed for customers to complete and return to the firm via fax at 920-380-6682. Contact the firm at 770-587-7400 for questions regarding this recall. | No consumer action necessary | 50 | 50 devices | Worldwide Distribution-USA (nationwide) including the states of CA, CO, ID, IL, MD, MO, NC, NJ, NY, PA, SC, UT, VA, WA, WI, and WV and the country of Canada. | Terminated | July 02, 2012 | 151 | 50 | Not_Computer
Z-0939-2012 | 60016 | oxygen administration kit | OGL | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Broselow/Hinkle Pediatric Emergency SystemVital Signs Inc.20 Campus RoadTotowa, NJ 07512 USAThe system supplies you with all the information and equipment needed for a young patients' emergency resuscitation requirements, all contained in a color coded, easy-to-use kit | Vital Signs | 2 | February 02, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | Povidone Iodine Pads included in the Vital Signs Broselow Pediatric Emergency Kit may be contaminated with Elizabethkinga meningoseptica. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | GE Healthcare sent a recall letter/return response form dated October 17, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to not use the affected product, isolate and discard all affected product.Fill out the attached Confirmation form and fax back to 800-535-7923 per the instructions to obtain replacement Busse IV Start Kits. Also, consignees were instructed to return the attached response form even if no recalled product was in their inventory.If they forwarded any affected product to any other healthcare institutions, they should forward a copy of the letter to those institutions.For any questions regarding this recall call 1-800-932-0760. |  | 6675 | 6675 kits | Nationwide and Foreign | Terminated | January 15, 2013 | 348 | 6675 | Not_Computer
Z-0940-2012 | 61030 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Aera/Skyra, model numbers 10432914 and 10432915.Nuclear magnetic resonance imaging system. | Siemens MAGNETOM Aera/Skyra | 2 | February 02, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring firm has discovered that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary. The potential existed for a cable connection to become overheated, leading to the emission of smoke and possible material damage. Firm preventively checked the gradient cabl | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens Medical Solutions sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated November 24, 2010 to all affected customers. The letter described the product, problem, and actions to be taken. Contact the firm at 610-219-4834 for questions regarding this recall. |  | 128 | 128 | Nationwide Distribution-including the states of AK, AZ, CA, CT, DE, FL, GA, IL, IN, KS, MD, MA, MI, MN, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, and WI. | Terminated | January 31, 2014 | 729 | 128 | Not_Computer
Z-0937-2012 | 60969 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF 16 PET/CT System & GEMINI TF 64 PET/CT System, Philips Medical Systems, Cleveland, OHThe device is a diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, head, heart and other organ imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode, dynamic, and gated acquisitions. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a "fused" (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: -The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. -The detection, localization, and staging of tumors and diagnosing cancer patients. -Treatment planning and interventional radiology procedures. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left-ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. | Philips Medical Systems GEMINI TF 16 PET/CT System &38; GEMINI TF 64 PET/CT System | 2 | February 02, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS cabinet that contained a melted battery which caused damage to other batteries lower in the cabinet, leading to battery acid leaking on the floor. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Philips Medical Systems (Cleveland) Inc communitcated the problem to the facilities via telephone. No customer letter was sent since the correction was already performed by the supplier. | Contact the recalling firm for information | 2 | 2 | Nationwide Distribution including Maryland and Colorado | Terminated | June 18, 2013 | 502 | 2 | Not_Computer
Z-0947-2012 | 58937 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***CAT 7205305***DYONICS***Arthroscopic Blade***3.5mm FULL RADIUS, 13cm long***Fits Power/EP-1 Motor Drives***Beige***SAMPLE LATEX-FREE***Made in the USA***Manufacturer: ASCENT HEALTHCARE SOLUTIONS***Lakeland, FL*** Lot number: 199468***Expiration Date: 12/2011.***CAT 7205306***DYONICS***Arthroscopic Blade***4.5mm FULL RADIUS, 13cm long***Fits Power/EP-1 Motor Drives***Yellow***SAMPLE LATEX-FREE***Made in the USA***Manufacturer: ASCENT HEALTHCARE SOLUTIONS***Lakeland, FL*** Lot number: 185659***Expiration Date: 08/2011.***CAT 7205307***DYONICS***Arthroscopic Blade***5.5mm FULL RADIUS, 13cm long***Fits Power/EP-1 Motor Drives***Orange***SAMPLE LATEX-FREE***Made in the USA***Manufacturer: ASCENT HEALTHCARE SOLUTIONS***Lakeland, FL*** Lot number: 197755***Expiration Date: 12/2011.Reprocessed arthroscopic shavers are indicated for use in orthopedic surgical procedures of the joints, jaw or sinuses where cutting and removal of soft and hard tissue or bone is needed in patients requiring orthopedic surgery. | DYONICSArthroscopic Blade | 2 | February 03, 2012 | 2012 | Ascent Healthcare Solutions | On 01/26/2009, Ascent Healthcare Solutions, Lakeland, FL initiated a recall of various models of Dyonics, Stryker, and Linvatech Reprocessed Arthroscopic Shaver Blades. The sterility of the devices may be compromised due to a potential breach in the seal of the Tyvek cover to the plastic tray. | PRODUCTION CONTROLS: Packaging Process Control | Ascent Healthcare Solutions sent an "URGENT MEDICAL DEVICE VOLUNTARY RECALL" letter dated January 26, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to examine their inventory, remove and quarantine all affected devices. A Recall Effectiveness Check Form was attached for customers to complete and return via fax to 1-866-325-4760. Contact the firm at (888) 888-3433 ext. 5519 for questions regarding this recall. | No consumer action necessary | 330 | 330 devices | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, and WY, and the country of Canada. | Terminated | March 09, 2012 | 35 | 3702 | Not_Computer
Z-0965-2012 | 53127 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Total Knee Junior Model Number 1100Product Usage:The Total KneeÂŽ Junior is an external polycentric prosthetic knee with a geometric locking system. | Total Knee | 2 | February 06, 2012 | 2012 | Ossur H / F | Some units of the Total Knee Junior produced between November 25th, 2008 and through April, 2009 contain balancing units that may fail under extreme use over a long period. The fault is noncritical and no incidents or injuries resulting from this situation have been reported to Ossur at this stage. Over long term use (estimated to be in the four to six year range) the prosthetic device may fail a | DESIGN: Component Design/Selection | Ossur sent a Safety Notice dated August 10, 2009 by mailed to all affected consignees. Ossur informed consignees of the above mentioned public reason for recall and asked consignees to take the following action: Identify the users of each of the Total Knee Junior model devices listed above and check inventory and isolate any recalled devices on hand but have not fitted on users. The letter states that Sue Fitzgerald from customer service will be contacting consignees in the next few days to coordinate replacement of the affected devices. For question call 1-(800) 219-2150, Ext: 4109. In addition, Ossur website: www.ossur.com we will be posting updates as appropriate. | Contact the recalling firm for information | 141 | 141 units | Worldwide Distribution - USA (nationwide) in the state of: AL, AR, AZ, CA, CO, DE, FL, GA, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, SC, TX, and WV and the countries of: Australia, Canada, China, France, Germany, Hong Kong, Israel, Japan, Malaysia, Netherlands, New Zealand, Portugal, Puerto Rico, Saudi Arabia, South Africa, Spain, Sweden, Taiwan, United Kingdom and United States. | Terminated | February 06, 2012 | 0 | 141 | Not_Computer
Z-0964-2012 | 60391 | enzymatic esterase--oxidase, cholesterol | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Analyzer, Model Number: 10-004, 14-874, 14-875, in vitro diagnostic. | Alere Cholestech LDX System | 2 | February 06, 2012 | 2012 | Alere San Diego | Alere has identified that a bias on results for certain analytes tested on the Alere Cholestech LDX System may exist due to humidity variability. | DESIGN: Device Design | Alere sent an Updated Urgent Recall Notification Letter in March 2012 to all their customers who are in possession of one or more Cholestech LDX Systems. The updated letter provides the customers with important information about the problem identified and the actions to be taken. Customers are instructed to operate their device in a controlled environment that maintains humidity within 40-60% relative humidity if the device has not been upgraded with ROM pack v3.40. If humidity cannot be monitored, the device is no longer CLIA waived. Customers are instructed to complete the customer verification form. A previous notice from Alere sent in November 2010 instructs users to contact Alere to receive the ROM pack upgrade. Customers may email LDXROM@alere.com to request an electronic form to complete and return via email. Customers can expect their ROM packs approximately 10 to 14 business days after Alere receives their request. If customers have transferred ownership to another party, they can email Alere at LDXROM@alere.com with the updated contact information. If customers are not currently using their Cholestech LDX System, it is recommended that the ROM Pack upgrade be installed prior to restarting use of any Cholestech LDX analyzer. Customers with questions regarding this communication can call Alere at (877) 441-7440, option 1 or email at LDXROM@alere.com. |  | 22000 | 22,000 meters | Worldwide distribution including USA | Terminated | August 22, 2013 | 563 | 22000 | Not_Computer
Z-0963-2012 | 60852 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Sterile surgical kits: DeRoyal SELECTIVE ARTERIOGRAM PACK, REF 89-7284.02 Lot Number: 26166359; DeRoyal UNIVERSAL PACK, REF 89-7237.01 Lot Number: 26013627; DeRoyal NEURO BASIC PACK, REF 89-7229.01 Lot number: 26013635; DeRoyal C-SECTION TRAY, REF 89-5256.06 Lot Number: 26185242; DeRoyal C V MINOR TRAY, REF 89-5277.02 Lot Number: 25793211; DeRoyal COMBINED EXTREMITY PACK, REF 89-6773.02 Lot Numbers: 25744022, 26010670, 26194641; DeRoyal DISKECTOMY TRAY, REF 89-6196.03 Lot Number: 26097373; DeRoyal MAJOR DRAPE PACK, REF 89-7255.01 Lot Number: 26011736; DeRoyal EXTREMITY PACK PGYBK, REF 89-5275.05 Lot Number: 26063270; DeRoyal SCOPE SHOULDER TRAY, REF 89-5653.03 Lot Number: 26227422; DeRoyal KNEE SCOPE PACK, REF 89-5654.03 Lot Number: 26020510; DeRoyal BILATERAL DRAPE, REF 89-5379.02 Lot Number: 26042735, 26282421; DeRoyal LAPAROSCOPIC CHOLE PACK, REF 89-5259.05 Lot Number: 25817201; DeRoyal MAJOR HIP PACK, REF 89-4718.04 Lot Number: 25840313; DeRoyal SPINE PACK, REF 89-4150.06 Lot Number: 25933077; DeRoyal ARTHROSCOPY TRAY, REF 89-5317.02 Lot Number: 26078463; DeRoyal LARGE VASCULAR TRAY, Lot Number: 26078447; DeRoyal SMALL VASCULAR TRAY, REF 89-3694.13 Lot Number: 261449399; DeRoyal ASC VASCULAR PACK PGYBK, REF 89-6284.05 Lot Number: 25698986; DeRoyal ASC BREAST PACK, REF 89-6285.06, Lot Number: 26078578; DeRoyal ASC FOOT/ANKLE PACK, REF 89-6014.07 Lot Numbers: 26096856, 26147916; DeRoyal ACUTE BURN EXCISION TRAY, REF 89-3330.05 Lot Number: 26078367; DeRoyal RECONSTRUCTION TRAY, REF 89-3331.04 Lot Numbers: 25817383, 26019720; DeRoyal FOOT PACK, REF 89-6467.05 Lot Number: 25880438; DeRoyal TRAUMA PACK, REF 89-6776.04 Lot Number: 26010688; DeRoyal KNEE ARTHROSCOPY PACK, REF 89-5257.08 Lot Numbers: 26125434, 26149102; DeRoyal CHILDRENs ORTHOPEDIC TRAY, REF 89-3479.08 Lot Numbers: 25873676, 25905092 DeRoyal KNEE SCOPE, REF 89-2910.09 Lot Number: 25695611; DeRoyal SPINE PACK PGYBK, REF 89-6720.01 Lot Number: 26022048; DeRoyal OPEN SHOULDER PROCEDURE PACK, REF 89-6629.01 Lot Number: 25839187; DeRoyal ORTHO TOTAL BASIC TRAY, REF 89-2650.15 Lot Numbers: 25893423, 26010207, 26077111; DeRoyal ORTHO TOTAL HIP/KNEE TRAY, REF 89-2644.17 A & B Lot Number: 25894012; DeRoyal MINOR/PLASTIC PACK, REF 89-6994.02 Lot Number: 25853544; DeRoyal OPEN HEART PACK, REF 89-6996.02 Lot Number: 25853608; DeRoyal GEN SURG UNIV PACK, REF 89-6991.02 Lot Number: 25696032; DeRoyal PERI VASCULAR PACK, REF 89-6998.03 Lot Number: 25853587; DeRoyal VASCULAR PACK, REF 50-4157.29 Lot Number: 25732507; DeRoyal OPEN HEART B, REF 50-13010.19 Lot Number: 26009978; DeRoyal GENERAL ORTHO PACK, REF 89-6883.02 Lot Number: 25695911; DeRoyal HIP COUNTABLES TRAY, REF 89-5382.03 Lot Number: 26020464; DeRoyal BASIC ORTHO TRAY, REF 89-3389.11 Lot Number: 26010291; DeRoyal TOTAL KNEE PACK, REF 89-7114.02 Lot Number: 26012780; DeRoyal ACF TRAY, REF 89-7660.01 Lot Number: 26022656; DeRoyal HAND PACK, REF 89-7549.02 Lot Numbers: 26042330, 26147713; DeRoyal ORTHO PACK, REF 89-7546.02 Lot Numbers: 26025242, 26025717; DeRoyal THORACOTOMY PACK, REF 89-7536.02 Lot Number: 26022161; DeRoyal LAPAROSCOPIC GYN PACK, REF 89-7500.02 Lot Number: 26096063; DeRoyal ELG/AAA PACK, REF 89-7495.02 Lot Number: 26256371; DeRoyal PEDI OPEN HEART PACK, REF 89-7486.02 Lot Number: 26100624; DeRoyal TOTAL KNEE PACK, REF 89-7484.01 Lot Number: 25875196; DeRoyal TOTAL KNEE PACK, REF 89-7466.01 Lot Number: 26013301; DeRoyal CUSTOM TOTAL HIP PACK, REF 89-7465.03 Lot Number: 26096039, 26166244; DeRoyal PERIPHERAL VASCULAR PACK, REF 89-7451.02 Lot Number: 26166252; DeRoyal NEURO PACK, REF 89-7396.01 Lot Number: 26061573; DeRoyal LABOR & DELIVER PACK, REF 89-7448.03 Lot Number: 26166236; DeRoyal SHOULDER ARTHROSCOPY PACK, REF 89-7446.02 Lot Number: 26096549; DeRoyal GENERAL MINOR PACK, REF 89-7365.01 Lot Numbers: 25838838, 26095765; DeRoyal CUSTOM GENERAL MAJOR PACK, REF 89-7357.01 Lot Nu | DeRoyal Surgical Kits and trays | 2 | February 06, 2012 | 2012 | DeRoyal Industries Inc | Custom surgical kits contain recalled Kimberly Clark surgical drapes which were found to contain metal shavings in the folds of the drapes. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DeRoyal Industries Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL OF KIMBERLY-CLARK THREE-QUARTER DRAPE 79147" letter dated June 14, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify and quarantine all affected product. Further instructions by the firm asks customers to complete and return a Recall Response Spreadsheet via fax to 865-362-3744. Customers should contact their Deroyal Sales Representative for questions or assistance regarding this recall. | Contact the recalling firm for information | 3881 | 3881 kits | Nationwide Distribution. | Terminated | June 27, 2013 | 507 | 3881 | Not_Computer
Z-0966-2012 | 61018 | monitor, fetal ph | LLT |  |  | Obstetrics/Gynecology | Contact ODE | Bravo pH Monitoring capsule, UPC 7290101361695 (FGS# 0312) and UPC 7290101361688 (FGS# 0313).Product Usage:The Bravo pH Monitoring System is intended to be used for gastroesophageal pH measurement and monitoring of gastric reflux in adults and children form 4 years of age. The Bravo pH Capsule can be attached following either endoscopy or manometry. | Bravo pH Monitoring capsule | 2 | February 06, 2012 | 2012 | Given Imaging Ltd. | Failure of the Bravo capsule to attach to the esophagus or alternatively, failure of the capsule to detach from the placement device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Given Imaging sent a Urgent Product Recall letter to all affected consignees. The letter identified the product, description of problem, actions to be taken and advised customers to read this important information packet in its entirety. Customers were advised to complete the Customer Response Sheet included and return per the instructions provided. Also attached to the letter was a Customer Bulletin: Update to Vacuum Settings. For questions call 1-800-448-3644 option 3. | No consumer action necessary | 21885 | 21,885 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Belgium, Brazil, Denmark, Canada, Cyprus, Finland, Germany, Hong Kong, India, Indonesia, Israel, Italy, Japan, Mexico, Netherlands, Norway, New Zealand, Panama, Portugal, Singapore, Sri Lanka, South Africa, Spain Thailand, Turkey, United Kingdom and Venezuela. | Terminated | September 14, 2012 | 221 | 21885 | Not_Computer
Z-0967-2012 | 60987 | stabilizer,heart | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | ACROBAT Stabilizers-ACROBAT V Vacuum Off-Pump System (OM-91 OOS), ACROBAT SUV Vacuum Off-Pump System (OM-9000S), ACROBAT Mechanical Off-Pump System (OM-6000S)Product Usage:The OM-9100S Acrobat V Vacuum Off-Pump System is intended for use during performance of cardiac surgical procedures through a sternotomy incision. The stabilizer isolates and provides local immobilization of a vessel on the beating heart. | XPOSE 4 Access Device (XP4000) XPOSE 3 Access Device (XP3000) positioners. | 2 | February 06, 2012 | 2012 | Maquet Cardiovascular, LLC | There may be a small perforation in the sterile pouch of select lots of the ACROBAT V, ACROBAT SUV, and ACROBAT Mechanical Stabilizers and the XPOSE 3 and XPOSE 4 Positioners. | PRODUCTION CONTROLS: Packaging Process Control | Maquet sent an Urgent Device Removal Immediate Action Required letter dated November 16, 2011 and a Field Action Response forms via Federal Express to to affected customers. The letter identifies the problem, product. risk factors and actions to be taken. Customers were instructed to examine their stocks immediately to determine if they have any of the recalled product. If so, discontinue dispensing (distributing) the lot and complete the Field Action Response form attached to complete and fax the attached Field Action Response form. In addition, it provides instructions for how to return the recalled product to MAQUET. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. If you have any additional questions, contact your local MAQUET Cardiovascul'ar sales representative or Customer Service at 1-888-880-2874 (6AM -5PM PST). |  | 9436 | 9436 unit total (US AND WORLDWIDE) | Worldwide Distribution - USA Nationwide | Terminated | January 10, 2014 | 704 | 9436 | Not_Computer
Z-0943-2012 | 60422 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Foot Rest Kit (pN 6252700-003) for Models 6251 and 6252 Stryker (EMS Stair -PRO), OE 2065939, 1739631, 2066066, 6066803 and 20666941 with optional , Stryker Medical, Portage, MI 49002. Intended to provide the patient greater comfort and a sense of security. | Stryker EMS Stair PRO | 2 | February 06, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | The Foot Rest Kit may have the wrong pivot spacers making the joint more susceptible to relaxation and prone to fatigue. There is a moderate risk of injury to the patient should the handles become detached during use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Stryker Medical spoke with all customers who received these shipments prior to 10/31/2011 requesting information on kits affected and recalled foot rest units installed. December 5, 2011 Urgent Medical Device Recall Letters, with instructions on how to to either install a new Foot Rest Kit ( enclosed ) or dispose of any kits that were not installed were sent to both customers who could have installed recalled foot rest units. Customers needing help or assistance with installation are asked to call Stryker Technical Support a 1-800-327-0770. Customers were to complete and return a postage paid postcard confirming receipt. Customers with questions or concerns are asked to contact Kristin Dudek at 269-389-6927 M-F m 8 a.m. - 5 p.m. (EST). | Contact the recalling firm for information | 3 | 3 units | Nationwide Distribution -- MI & MA. | Terminated | August 28, 2013 | 569 | 3 | Not_Computer
Z-0971-2012 | 60794 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROSÂŽ Chemistry Products Calibrator Kit 24, Catalog No. 680 1708. Common/Usual Name: VITROS Calibrator Kit 24. For in vitro diagnostics use only. VITROS Chemistry Products Calibrator Kit 24 is used to calibrate the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the quantitative measurement of albumin in urine. Ortho-Clinical Diagnostics Inc. | VITROS Chemistry Products Calibrator Kit 24 | 2 | February 07, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received numerous complaints of lower than expected results for control fluids and patient samples upon initial calibration or following a recalibration using VITROSÂż Chemistry Products Calibrator Kit 24 and VITROSÂż Chemistry Products mALB Reagent. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | All consignees were sent "Important Product Correction Notification" communications (dated 10/27/11) beginning on 10/27/11 informing them of the issue, and instructing them to discontinue use of their remaining inventory of VITROSÂŽ Chemistry Products Calibrator Kit 24, VITROSÂŽ Chemistry Products mALB reagent, and VITROSÂŽ Chemistry Products mALB Performance Verifiers. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. |  | 4436 | 4436 units worldwide (domestic - 2201; foreign - 2235) | Worldwide Distribution - USA, including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open |  |  | 4436 | Not_Computer
Z-0979-2012 | 60577 | catheter, percutaneous, cardiac ablation, for treatment of atrial fibrillation | OAE |  |  | Cardiovascular |  | EZ Steer ThermoCool Nav Part Number D-1292-01-S Diagnostic/Ablation Catheter instructions for use.The Biosense EZ STEER~ THERMOCOOL~ NAV Diagnostic/Ablation Deflectable Tip Catheter is a multi-electrode luminal catheter with a deflectable tip designed to facilitate electrophysiological mapping of the heart and to transmit radiofrequency (RF) current to the catheter tip electrode for ablation purposes. The catheter shaft measures 7.5 F with 8 F ring electrodes. For ablation, the catheter is used in conjunction with an RF generator and a dispersive pad (indifferent electrode) for the treatment of: - Type I atrial flutter in patients age 18 or older. - Recurrent drug/device refractory sustained monomorphic ventricular tachycardia (VT) due to prior myocardial infarction (Mil in adults. - Drug refractory recurrent symptomatic paroxysmal atrial fibrillation, when used with compatible three-dimensional electroanatomic mapping systems. The EZ STEER~ THERMOCOOL~ NAV Diagnostic/Ablation Deflectable Tip Catheter provides location information when used with CARTO~ EP Navigation Systems. | EZ Steer ThermoCool Nav | 2 | February 07, 2012 | 2012 | Biosense Webster, Inc. | Biosense Webster, Inc. has recently become aware of the omission of the following two Warning and Precaution statements on the Instructions For Use (IFU) for the above catheter:1. If phrenic nerve location is a concern, precautionary measures are recommended to evaluate the proximity of the nerve to the ablation electrode, such as pacing maneuvers.2. To minimize risk of atrio-esophageal fist | DESIGN: Labeling Design | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated Novembe 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to continue use of the affected product and that no further action is required on their part. Consignees were advised that the updated IFU can be found at www.jnjgatewayifu.com. Consignees were instructed to contact their local Biosense Webster representative if they have any questions or require additional information.For questions regarding this recall call 909-839-8604. |  | 24713 | 24,713 units in total | Nationwide Distribution | Terminated | February 08, 2012 | 1 | 24713 | Not_Computer
Z-0975-2012 | 60919 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU Clinical Chemistry Analyzers: Model Number: AU600 Clinical Chemistry Analyzer,AU640 Clinical Chemistry Analyzer,AU640(e) Clinical Chemistry Analyzer,AU680 Clinical Chemistry Analyzer.The AU Clinical Chemistry Analyzers are fully automated photometric analyzers intended for clinical laboratory use. Applications include colorimetric, turbidimetric, latex agglutination, and homogeneous enzyme immunoassay. | AU Clinical Chemistry Analyzers | 2 | February 07, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the AU Clinical Chemistry Analyzers may experience cuvette overflow. |  | Beckman Coulter sent an "UPDATE TO PCA -14563: URGENT PRODUCT CORRECTION" letter dated April 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter is a communication which describes an alternative approach to detecting cuvette overflow. Additionally, a response form was enclosed for customers to complete and return to the firm. Contact the Customer Support Center at (800) 854-3633 or on the Beckman Coulter website at www.beckmancoulter.com/customersupport/support/default.asp for questions regarding this notification. |  | 8225 | 8225 units | Worldwide Distribution-including USA (nationwide). | Open |  |  | 8225 | Not_Computer
Z-0978-2012 | 60011 | blade, saw, surgical, cardiovascular | DWH | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Terumo, Sarns Sternal Saw I System, Catalog 6090. The Sarns Sternal Saw I is indicated for use in splitting the sternum during thoracic surgery. | Sarns Sternal Saw I System | 2 | February 07, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received one report of minute metal shavings leaking from the Sarns Sternal Saw I System at the point where the saw blade is inserted into the saw handpiece. | PRODUCTION CONTROLS: Equipment Maintenance | On 11/16/2011, 11/23/11, 02/3/12, 03/22/12, and 4/24/12 the firm sent a letter entitled "URGENT MEDICAL DEVICE RECALL" to it's domestic direct accounts that included the name of the device, catalog and lot numbers, dates of distribution, and instructions to discontinue use and return the device to the firm for destruction, with a Fax-back response form to be submitted. Medwatch reporting information was also included in the letter. Questions and comments are to be addressed by Customer Service @ 1-800-521-2818, M-F, 8-6 EST. | Contact the recalling firm for information | 10186 | approx. 10186 | Worldwide Distribution - USA, and the countries of Colombia, Argentina, Lebanon, Belgium, Chile, India, Sri Lanka, Jordan, Egypt, Mexico, Brazil, Taiwan, Honduras, Canada, Germany, Singapore, Republic of Georgia, Hong Kong, Japan, UAE, Philippines, Thailand and South Africa. | Terminated | December 06, 2012 | 303 | 10186 | Not_Computer
Z-0972-2012 | 60662 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens CLINITEST hCG -Pregnancy Test (qualitative), Radio immunoassay,Human Chorionic Gonadotropin (hCG) test system. Package of 25. Product Code 1760/ Catalog# 06484105).For in vitro diagnostic use as a qualitative method in the rapid detection of human chorionic gonadotropin (hCG) in urine specimens. | Siemens CLINITEST hCG | 2 | February 07, 2012 | 2012 | Siemens Healthcare Diagnostics | A potential for false positive or increase in borderline results with patients and commercially available negative control material. |  | Siemens issued Overnight Mailings of Urgent: Recall Notices on December 15, 2011 for delivery beginning December 16, 2011 to Distributors and end users. Customer bulletin has been issued instructing customers to discontinue use of product and discard impacted inventory. Customers are to contact their distributor regarding credit for unused product. For technical questions contact Siemens Technical Solutions Center at 877-229-3711. |  | 48578 | 48,578 Kits (25 tests / kit) | Worldwide Distribution -- US, Austria, Germany, Great Britain, Denmark, Norway, Sweden, Poland, Netherlands, Luxembourg, Belgium, Finland, France, Greece, Ireland, Italy, Portugal, Qatar, Saudi Arabia, Slovakia, Slovenia, and Switzerland. | Open |  |  | 48578 | Not_Computer
Z-0984-2012 | 60800 | enzyme immunoassay, valproic acid | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | VITROSÂŽ Valproic Acid (VALP) Reagent, REF/Catalog No. 6801710 --- Ortho- Clinical Diagnostics Rochester, NY USA -- Ortho- Clinical Diagnostics High Wycombe, UK --- An in vitro diagnostic. For in vitro diagnostic use only. Used to quantitatively measure valproic acid (VALP) concentration in human serum and plasma. | VITROS Valproic Acid (VALP) Reagent | 2 | February 08, 2012 | 2012 | Ortho-Clinical Diagnostics | Results using VITROSÂż Chemistry Products VALP Reagent were negatively biased compared to other manufacturers. OCD confirmed that patient results were also affected by this issue. | PRODUCTION CONTROLS: Manufacturing Material Removal | All domestic end-user consignees were initially sent "Important Product Correction Notification" recall communications (dated 5/10/11) by Federal Express overnight mail on 5/10/11. The letters informed users of the issue, and advised them to discontinue using their VITROS VALP reagent until adjusted calibrator values are provided for GEN 14 of VITROS VALP Reagent. The adjusted calibrator values for VITROS VALP Reagent were sent to customers via Assay Data Disk (ADD) Data Release Version (DRV) 5632. Customers were to refer to the Technical Support Communication for details. Users were to confirm receipt of the notice by completing and returning the Fax Reply Form - Confirmation; confirm that all customers or distributors were informed of the issue and instructed to discontinue use of affected product; and lastly confirm that all customers or distributors were sent adjusted calibrator values. |  | 5800 | Approximately 5800 units worldwide (domestic - 4549, foreign - 1287) | Worldwide Distribution -- US, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open |  |  | 5800 | Not_Computer
Z-0148-2012 | 59982 | monitor, patient position, light-beam | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Signa, Discovery and Optima. Signa HDe 1.5T; Signa Infinity TwinSpeed MR System, 1.5T Infinity TwinSpeed, Signa Lx; Signa Infinity MR System with EXCITE Technology, Excite II, Signa Excite 1.5T Expert; Signat 3T Infinity with Excite; Signa 3.0T with EXCITE; Signa Excite 1.5T and Signa Excite 3T, Signa Excite HD, Signa HD 1.5T and Signa HD 3.0T; Signa HDx; Signa HDxt; Signa Vibrant; Signa HDi; Signa Horizon Cx aka MR/I; Discovery MR450; Discovery MR750; Optima MR450w; Signa Cvi, and Signa Nvi.The Optima MR45Ow is a whole body magnetic resonance scanner designed to support high resolution and high signal-to-noise ratio images in short exam times. it is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique anatomical images, spectroscopic data, parametric maps, or dynamic images of the structures or functions of the entire body. The indication for use includes, but is not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. Depending on the region of interest being imaged, contrast agents may be used. The images produced by the Optima MR450w reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and spectra, when interpreted by a trained physician yield information that may assist in diagnosis. | GE MR Systems | 2 | February 08, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of missing labels on the alignment lasers. | DESIGN: Labeling Design | GE Healthcare will execute a field action in which GE Healthcare Field Engineers will perform an inspection of all suspect MR systems. Any systems lacking any of the 3 laser warning labels will be remediated by applying the missing labels). We expect this field action will be completed by June 30, 2012.1. Notification of all distributors and purchasers be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. The FDA district office coordinator noted below is to be included in the notifications.2. Corrections be made at no cost to the purchasers. | Contact the recalling firm for information | 4614 | 4614 | Nationwide Distribution | Terminated | June 29, 2012 | 142 | 4614 | Not_Computer
Z-0986-2012 | 60938 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra OS Osteoconductive Scaffold Putty (2.5 cc and 5 cc). Integra OS Putty, combined with bone marrow aspirate, is intended for use as a bone void filler to fill voids or gaps of the skeletal system. Following placement in the bony void or gap (defect), Integra OS Putty is resorbed and replaced with bone during the healing process. | Integra | 2 | February 08, 2012 | 2012 | Integra LifeSciences Corp | Integra OS Osteoconductive Putty is labeled with incorrect storage conditions. | PRODUCTION CONTROLS: Error in Labeling | Integra issued an "Urgent: Product Recall Notification" to customers and sales specialists on 10/7/2011. They were advised of the problem and provided instructions on how to handle the affected product. |  | 45 | 45 units | Nationwide Distribution - Including the states of CA, OK, MD and VA. | Terminated | May 21, 2013 | 468 | 45 | Not_Computer
Z-0987-2012 | 59092 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ***REF C9944A***"FULL RADIUS RESECTOR"***2.9mm***Qty 6***STERLING***STERILE***Rx Only.*** Made in the USA***Manufacturer: CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***ConMed Linvatec Biomaterials, Ltd.Hermiankatu 6-8L 33720 Tampere, FINLAND.***Lot number # 95987.Soft Tissue resection and removal. | ConMed Linvatec 2.9MM Full Radius Resector | 2 | February 10, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL recalled catalog number, C9944A, lot # 95987, 2.9MM Full Radius Resector Blade in November 2009 because there was a possibility that the device was assembled with an incorrect outer tube. | TRAINING: Employee Error | A Linvatec Corporation dba ConMed Linvatec sales representative contacted the affected customers by telephone on November 12, 2009. The customers were also contacted by the Recall Coordinator. The product, problem, and action to be taken by the customers were discussed. Customers with the product in their inventory were instructed to return the product to the firm. Contact the firm at 727-399-5205 for questions regarding this recall. | Contact the recalling firm for information | 30 | 30 Units | Product Distributed in the states of Georgia and New York. | Terminated | February 13, 2012 | 3 | 30 | Not_Computer
Z-0988-2012 | 60902 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Gamma camera systems: AXIS SPECT Systems & IRIX SPECT Systems, Model #: AXIS- 882130, IRIX- 882140. Philips Healthcare.Used to do nuclear medicine diagnostic imaging of various body organs. | AXIS SPECT Systems & IRIX SPECT Systems | 2 | February 10, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was informed that the detector buckets may experience weakness over time. As a result, this could cause the detector bucket to separate from the mounting plate of the gantry and eventually result in the detector falling. | DESIGN: Process Design | On 11/15/2011 Philips sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected product, the problem, the hazard involved, and the actions to be taken by users and the Philips. Customers are to place the letter in their User Documentation. Philips service will arrange a time to inspect customers' systems and conduct the appropriate field safety correction. Customers are to contact their local Philips representative if they need further support or information. In the US and Canada, customers can contact Philips Healthcare Customer Care Solutions Center at 1-800-722-9377. |  | 795 | 795 units | Worldwide Distribution -- USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PRO, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, & WY and the territory of Puerto Rico; and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Egypt, Finland, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Japan, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, & United Kingdom. | Terminated | February 22, 2013 | 378 | 795 | Not_Computer
Z-0990-2012 | 60693 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Flow Tec Nano-Hybrid Flowable Affected ProductSyringe Part No.Value Pack Part. No.ShadeBEN3784-514BEN3786-754A1BEN3784-532BEN3786-763A2BEN3784-541BEN3784-738A3BEN3784-550BEN3784-747A3.5BEN3784-569BEN3784-756A4BEN3784-578BEN3784-765B1BEN3784-596BEN3784-774B2BEN3784-603BEN3784-783C1BEN3784-612BEN3784-792C2BEN3784-621BEN3784-809C3BEN3784-649BEN3784-818D2BEN3784-722BEN3784-818Universal OpaqueBEN3784-685BEN3784-827A0BEN3784-701BEN3784-836Incisal.The intended use of this device is as a light cured, low viscosity, hybrid composite material which is compatible with a variety of bonding agents and luting materials during the course of dental restorations. Artiste Flowable and Flow Tec (private label of Artiste Flowable) are indicated for restorations of carious lesions, porcelain repairs, small core build-ups, fill-in of surface enamel irregularities, etc. | Artiste Flowable and Flow Tec (private label of Artiste Flowable) | 2 | February 11, 2012 | 2012 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed that the Flow Tec Nano-Hybrid Flowable Composite material has been found to be difficult to extrude. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pentron Clinical sent an initial "URGENT: MEDICAL DEVICE RECALL" letter dated April 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and fax back an enclosed recall return form in order to confirm their receipt of the recall notification. A second "URGENT: MEDICAL DEVICE RECALL" letter dated November 10, 2011 was sent to all affected customers because Pentron Clinical expanded the recall. On December 8, 2011 another recall communication was forwarded to all affected customers to change the wording on the letter to make it more clear to the customers about the recall expansion. Contact Pentron Clinical Customer Service at 1-800-551-0283 for questions regarding this recall. |  | 77289 | 77,289 Units total | Worldwide Distribution: USA (nationwide) and the countries of Canada, Egypt, France, Germany, Italy, and Spain. | Terminated | October 05, 2012 | 237 | 77289 | Not_Computer
Z-1008-2012 | 60929 | electrode, ion based, enzymatic, creatinine | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | S1827 Crea Calibration solution 1, 944-135, Lot DV-02 Radiometer Medical, Bronshoj DenmarkProduct Usage:The S1827 Crea Calibration Solution 1 (944-135) Lot DV-02 is a solution used by the ABL 8x7 series creatine analyzers to establish calibration points for measured parameters. | Crea Calibration solution | 2 | February 13, 2012 | 2012 | Radiometer America Inc | When the user activates the dosicap (containing the crea additive) and mounts it onto the bottle, the threads will not seal correctly causing approximately 10% of the solution to leak out when the bottle is placed horizontally for dissolving the additive. This will cause the concentration of creatinin and creatin in the calibration solution to be wrong, and results in a bias on the Crea results. | DESIGN: Device Design | Radiometer sent a Medical Device Recall letter dated December 28, 2011 to their customers. The letter provided details on the affected devices, description of the problem and actions to be taken. The firm instructed their customers to immediately examine their stock and return any remaining product to Raiometer America Inc. for disposal. The letter states that Radiometer will replace all bottles of the affected lots, at no charge. Customers were asked to complete the attached Recall Response Fax Form and fax to 800-736-0601. For questions regarding the Recall contact Radiometer America Technical support at 1-800-736-0600 opt. 4. |  | 7 | 7 boxes/6/175mL bottles | Nationwide Distribution (USA) including the states of: NY, PA and TN. | Terminated | January 17, 2014 | 704 | 7 | Not_Computer
Z-1000-2012 | 59936 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Artiste Flowable Value Pack Part Number N28VA, Syringe Lot Number 165741The intended use of the device is as a light cured, low viscosity, hybrid composite material which is compatible with a variety of bonding agents and luting materials during the course of dental restorations. Artiste Flowable and Flow Tec (private label of Artist Flowable) are indicated for restorations of carious lesions, porcelain repairs, small core build-ups, fill-in of surface enamel irregularities, etc. | Artiste Flowable Value Pack | 2 | February 13, 2012 | 2012 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed the Flow Tec Nano Hybrid Flowable Composite material has been found to be difficult to extrude. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sybron Dental sent a Recall Letter dated November 15, 2010 to consignees informing them of the above mentioned public reason for recall and asking them to determine if they have any affected products in their inventory and requesting their cooperation in returning any affected product which they still have in their stock. Product was to be returned and replaced at no charge.They also requested that consignees complete and fax back an enclosed return from , whether or not they had any product to return.Consignees were told to call Pentron Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement.Authorized Petron Clinical distributors were asked to identify and recover the affected product lots that may have been shipped to their customers. |  | 771 | 771 units in total | Worldwide Distribution - USA (nationwide) and the countries of Cyprus, Germany, Italy, Poland, Canada, Ecuador, and Kuwait | Terminated | February 13, 2012 | 0 | 771 | Not_Computer
Z-1011-2012 | 60835 | [6 empty] | Zoll pedi.padz Reduced Energy Multi-Function Electrodes (MFE)Part number: 8900-0401Used in connection with Zoll M Series and E series Defibrillators specifically for pediatric patients. | Zoll pedi.padz Reduced Energy MultiFunction Electrodes (MFE) | 2 | February 13, 2012 | 2012 | Bio-Detek, Inc. | Lack of an adequate silicone seal around the periphery of the attenuator enclosure may allow fluid to enter. | PRODUCTION CONTROLS: Process Control | Zoll Medical issued Urgent Device Correction notices dated 12/30/11 via UPS/Registered Mail. The letter described the issue and identified the affected product. ZOLL recommended that customers alert their staff of the issue, that the electrodes should be kept away from a fluid environment, to use caution in handling the electrodes if they are subjected to a fluid environment and maintain safe defibrillation practice, and seek a replacement if available. Also customers should contact Zoll 1-800-348-9011 for product replacement. The firm plans to replace the affected lots free of charge. |  | 3080 | 3,080 units | Worldwide Distribution -- US and Canada. | Terminated | December 11, 2012 | 302 | 3080 | Not_Computer
Z-1003-2012 | 61038 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Custom surgical kits:DeRoyal(R) TRAUMA CRANIOTOMY PACK, REF 89-7639.01, STERILE EO, Manufactured by: DeRoyal Industries, Inc., Powell, TN 37849DeRoyal(R) TRAUMA CRANIOTOMY PACK PGYBK, REF 89-7639.01, STERILE EO, Manufactured by: DeRoyal Industries, Inc., Powell, TN 37849DeRoyal(R) TRAUMA CRANIOTOMY PACK PGYBK, REF 89-7639.02, STERILE EO, Manufactured by: DeRoyal Industries, Inc., Powell, TN 37849surgical kits | DeRoyal(R) TRAUMA CRANIOTOMY PACK | 2 | February 13, 2012 | 2012 | DeRoyal Processing Center | The firm distributed surgical kits which contained DermaHook Neuro Elastic Retractors which were subsequently recalled by Teleflex. |  | DeRoyal Industries sent a Customer letter dated January 6, 2012, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer. An email was also sent on January 10, 2012, to the affected customer. The customer was requested to follow these instructions. 1) Identify any affected inventory using the product and lot numbers on Attachment 1.2) Attach the provided warning label to each affected kit and tray.3) Forward this recall to any of your end users that may have affected product.4) Complete and return Attachment 1 by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter, even if you no longer have affected inventory.If you have questions or need assistance with the recall, including additional labels, please call (865) 362-1037 . We apologize for any inconvenience this may caused you. |  | 214 | 214 units | (USA) including Indiana | Terminated | November 05, 2012 | 266 | 214 | Not_Computer
Z-1012-2012 | 61044 | bone grafting material, synthetic | LYC | Bone grafting material. | Dental | Dental | 510(k) | Model #03001-02: Periophil Biphasic, Synthetic Biphasic Dental Bone Graft Granules, 1.0 cc. Model # 03003-01 Periobase Biphasic: Periobase, Synthetic Biphasic Dental Bone Biphasic Granules, 0.5 cc. Model #03002-01: Periobone Biphasic, Synthetic HA/TCP Dental Bone Graft Biphasic Granules, 0.5 cc. Periophil Biphasic (synthetic grafting material) is intended for use as a bone grafting material to fill, augment, or reconstruct periodontal or oral/maxillofacial defects. Cytophil, Inc., 2485 Corporate Circle, Suite 2, East Roy, WI 53120. | Periophil Biphasic Periophil Biphasic, Periobase Biphasic, and Periobone Biphasic | 2 | February 13, 2012 | 2012 | Cytophil, Inc | Cytophil Corporation is recalling Periophil Biphasic Periophil Biphasic, Periobase Biphasic, and Periobone Biphasic because devices were inadvertently shipped after receiving an inadequate dose of radiation and this product may not be sterile. | PRODUCTION CONTROLS: Process Control | Consignees were sent an "Urgent Medical Device Recall" letter dated January 11, 2012 by Cytophil. The letter described the problem and the product involved in the recall. The firm advised consignees to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, they were requested to identify their customers and notify them at once of this product recall. For questions, call 262-642-2765. | Contact the recalling firm for information | 547 | 547 | Nationwide Distribution - Including the states of WI, GA, IL, SC, MO, and TX. | Terminated | May 01, 2012 | 78 | 547 | Not_Computer
Z-1010-2012 | 61004 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Surgical lighting systems iLED and TruLight in use with the ALC+ function. The TRUMPF surgical lights are intended to locally illuminate an operating or examination area of the patients body with a high intensity light. | TruLight ALC, Trulight 5520 or 5320 lighting system. | 2 | February 13, 2012 | 2012 | Trumpf Medical Systems, Inc. | The ALC+ function may not be properly locked during use resulting in the possibility of the handle falling off into the surgical field causing injury to the patient. | DESIGN: Device Design | The firm, TRUMPF Medical Systems, Inc., sent an "Urgent Safety Information" notice/package dated January 9, 2012 to its customers. (Note: Some were delivered by TRUMPF representatives in person and some via mail with Urgent Safety identification on the envelope). The notice describes the product, problem and actions to be taken. TRUMPF will replace all sterilizable handles of this type in the field with new handles, in addition, TRUMPF will send new handles and verify disposal of the old version at the customers location. Distribution of new handles will start January 2012. The customers were instructed to ensure that all users and other relevent persons are informed of the Urgent Safety Information; store the information until the measure has been completed and complete and return the Confirmation of User Receipt form via mail to TRUMPF Medical Systems, Inc. 1046 LeGrand Blvd., Charleston, SC 29063 or email: Lindsey.ronnenberg@us.trumpf-med.com.If you have any questions, call 843-822-6939. |  | 8733 | 8,733 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, CT, DE, FL, GA, IL, IN, MA, MD, MO, NC, NM, NJ, NY, OH, PA, TX, VA, VT, WA, and WI; and countries of: Australia; Austria, Belguim, China, Cyprus, Denmark, Estonia, Finland, France, Germany, Great Britain, Hong Kong, Indonesia, India, Ireland, Italy, Japan, Kazakhstan, Kenya, Korea, Latvia, Luxembourg, Mexico, Netherlands, Norway, Pakistan, Peru, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Seychelles, Singapore, South Africa (Zuid Afirka), Spain, Sweden, Switzerland (Confederation of Helvetia), Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 09, 2012 | 239 | 8733 | Not_Computer
Z-1009-2012 | 60988 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD 30G1/2" Precision Glide" NeedleNeedles are used for aspiration and injection of medications | Becton Dickinson 30G1/2" Precision Glide" Needle | 2 | February 13, 2012 | 2012 | Becton Dickinson & Company | Becton Dickinson has identified clogged needles. | DESIGN: Device Design | Becton Dickinson sent Urgent Product Recall letters dated October 18, 2011 and October 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to deterine if they have any of the affected product. If they have affected product customers should return the product to Beckton Dickinson. For any questions regarding this recall call 201-847-4344. |  | 480000 | 480,000 units | Nationwide Distribution including IL and NJ | Terminated | December 11, 2012 | 302 | 480000 | Not_Computer
Z-0997-2012 | 60901 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DUET TRS 60 4.8MM STRAIGHT SULUProduct Code: DUET6048Product Usage:The ENDO GIA UNIVERSAL Staplers with Duet TRS SULUs have applications in abdominal, gynecologic, pediatric and thoracic surgery for resection, transection and creation of anastomosis. They may be used for transection and resection of liver substance, hepatic vasculature and biliary structures. | Covidien | 2 | February 13, 2012 | 2012 | Covidien LP | Duet TRS Straight and Articulating Single Use (SULU)Loading Staplers used in thoracic surgery may result in serious injury or death |  | Covidien issued an Urgent Medical Device Recall letter dated January 12, 2012 to all affected customers. The letter identified the affected products, problem, and actions to be taken. Customers were instructed to immediately discontinue use of the device in thoracic surgery, Immediately advise all surgeons/surgery personnel of this recall and return products intended for thoracic use to Covidien by contacting Customer Service at SDFeedback@Covidien.com or 1-800-722-8772, option 1, to obtain a Return Goods Authorization prior to returning the affected units. Customers were advised to complete the Recall product return form in its entirety and return it with the returned units. For questions contact your Covidien representative or Quality Assurance at 203-492-5232. |  | 9687 | 9,687 units | Worldwide Distribution - USA (nationwide) and the countries of: ÂŽeskÂŽ republika, ASSAGAY, Australia, Austria, Belgium, Bosnia, Herzegovina, Bulgaria Cairo,Egypt, Canada, Canary Islands, China, Croatia,Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France,Germany, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Ireland,Israel, Italy, Japan, Jordan, Kingdom of Bahrain, Korea, Republic of Kuwait,Latvia, Lebanon, Libya, Luxembourg, Nederland, Nelspruit, New Redruth,New Zealand, Norway, Nouvelle Caledonie, Philippines, Poland, Portugal, Qatar, Reunion, Russia, Romania, Rustenburg, Saudi Arabia, Serbia,Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Turkey, Taiwan, Thailand, Tunisia, Uae, Uae - Dubai, Umhlanga Rocks,United Kingdom, Vietnam and Wales. | Open |  |  | 666880 | Not_Computer
Z-1014-2012 | 50701 | elisa, antibody, west nile virus | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgM Capture ELISAThe Panbio West Nile Virus IgM Capture ELISA is for the qualitative presumptive detection of IgM antibodies to West Nile virus in serum as an aid in the clinical laboratory diagnosis of West Nile virus infection in patients with clinical symptoms consistent with encephalitis/meningitis. Positive results must be confirmed by plaque reduction neutralization test (PRNT), or by using the current Centers for Disease Control and Prevention (CDC) guidelines for diagnosis of this disease. Assay performance characteristics have not been established for testing cord blood, neonate, prenatal screening, general population screening without symptoms of meningoencephalitis or automated instruments. The user is responsible for establishing these assay performance characteristics. | West Nile Virus IgM Capture ELISA | 2 | February 14, 2012 | 2012 | Inverness Medical | It has come to our attention that users of these product lots may experience an elevation in the rate of false positivity when compared with previous product lots. Internal quality assurance testing has shown that the products are performing within the release specifications, although there has been a slight shift in the cut-off which may lead to an increase in the number of positive results obta | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Inverness Medical sent an Urgent In Vitro Diagnostic Recall letter dated October 6, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use, destroy any unused or partially used test kits of the affected lot(s), and complete the table and certification of compliance. Customers were asked to fax back the completed form to inverness Medical at 1-609-527-8054. Replacement stock would be forwarded to them at no cost when the new lot becomes available.For questions regarding this recall call 609-627-8038. |  | 289 | 289 Kits | Worldwide Distribution including ND, PA, TX, MN; and the country of Canada | Terminated | February 15, 2012 | 1 | 289 | Not_Computer
Z-1013-2012 | 59094 | stretcher, wheeled | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | CI healthcare***Model: CI8092***S/N: 5-60-85-044 MFD: JUN, 2010 MADE IN CHINA***cihealthcare.comThis stainless steel stretcher is used for transporting patients. | CI Healthcare Stainless Steel Patient Transport Stretcher | 2 | February 14, 2012 | 2012 | Cardiopulmonary Instrumentation, Inc | Cardiopulmonary Instrumentation, Inc. dba CI Healthcare, Miami, FL, is recalling their Stainless Steel Patient Transport Stretcher, model # CI-8092, Lot 002, because of a malfunction that occurs when weight is applied to the backrest. This malfunction causes the backrest support bar, when secured in the first hook, to fall to a flat position. | DESIGN: Device Design | CI Healthcare sent an "URGENT RECALL NOTIFICATION" letter dated February 17, 2011, to its customer. The letter describes the product, problem and actions to be taken. The customer were instructed to do the following: 1) Examine their inventory and isolate the product(s). 2) If they have distributed any of the products to other services or facilities, please forward this information as appropriate. 3) Discontinue use of the product until a correction has been properly executed and 4) Complete and return the "Recall Response Form" as soon as possible. Should you have any questions, please call Robert Nunez at (305) 592-8196. | Contact the recalling firm for information | 72 | 72 Units | Nationwide Distribution - to one distributor in Puerto Rico | Terminated | April 19, 2012 | 65 | 72 | Not_Computer
Z-1032-2012 | 60996 | catheter, percutaneous, cardiac ablation, for treatment of atrial flutter | OAD |  |  | Cardiovascular |  | The 1500T9 Cardiac Ablation Generator, RO#lBl-89000 is part of the Therapy Cardiac Ablation System, along with the Cool Point irrigation Pump, the Therapy 1300 Series Steerable Ablation Catheters, the Therapy Cool Path Ablation Catheters and accessories. The generator is a microprocessor-controlled device that produces a continuous unmodulated radiofrequency (RF) output delivered from the generator between its tip electrode and a dispersive pad for the treatment of atrial flutter. | 1500T9 Cardiac Ablation Generator | 2 | February 14, 2012 | 2012 | St Jude Medical Inc | Two St Jude Medical 1500T9 RF Generator units were distributed to customers while still under FDA review as part of a PMA supplement to the approved PMA applications. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | The notification letter was dated January 5th, 2012 and titled "Product Recall RF Generator 1500T9-CP". The letter informed customers that a St Jude Medical 1500T9-CP RF Generator unit had been distributed to them while currently under FDA review as part of a PMA supplement application. The letter requested that the use of the impacted generator be discontinued. Customers were requested to return the field action form and return the unit identified in the letter. The contact number provided in the letter is: 651-756-2000. |  | 2 | 2 units | Nationwide Distribution, including the states of Kansas and Maryland. | Open |  |  | 2 | Not_Computer
Z-1016-2012 | 60947 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | DYNASTY(R) BIOFOAM(R) SHELL, SIZE: 46mm GROUP B, REF DSBFGB46, Rx ONLY, STERILE R, Wright Medical Technology, Inc., hip prosthesis acetabular component.The instrument is used as a hip acetabular shell. | DYNASTY(R) BIOFOAM(R) SHELL | 2 | February 14, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that particulate debris from the packaging spacer could be unintentionally implanted into a patient and could potentially lead to complications that might necessitate a revision surgery. | PRODUCTION CONTROLS: Packaging | Wright Medical sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. A Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact Customer Service at 1-800-238-7117 for questions regarding this recall. |  | 160 | 160 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Denmark, and Italy. | Terminated | March 13, 2013 | 393 | 5342 | Not_Computer
Z-1040-2012 | 61098 | stent, iliac | NIO |  |  | Cardiovascular |  | S.M.A.R.T. CONTROL NITINOL STENT SYSTEM ILIAC: Self-Expanding Vascular Stents delivered by catheter for the indication of improving luminal diameter in patients with symptomatic atherosclerotic disease of the iliac arteries. catalog numbers are CxxyyyMB, CxxyyySB, CxxyyyML, and CxxyyySL, where xx is the stent diameter in mm and yyy is the stent length in mm; but not including catalogs where yyy is 120 or 150 . The Removal only includes Catalog Numbers in the above sequence with S.M.A.R.T.ÂŽ CONTROLÂŽ in the logo/description. All unexpired lots from 15112447 through 15525516. | S.M.A.R.T. CONTROL NITINOL STENT SYSTEM ILIAC | 2 | February 15, 2012 | 2012 | Cordis Corporation | Cordis Corporation initiated a recall of S.M.A.R.T. CONTROL Nitinol Iliac Stent System due to a potential sterility breach in the pouches. | PRODUCTION CONTROLS: Packaging | All affected consignees were notified by an "Urgent Medical Device Recall (Removal) Letter" dated January 31, 2012 and mailed by overnight delivery on February 1st and 2nd, 2012. Consignees were instructed to identify and quarantine affected product; review, complete, sign, and return an Acknowledgement Form; return the affected product per instructions. Questions should be directed to Customer Service at (800) 551-7683. |  | 132000 | 132,000 units | Worldwide Distribution - USA, and the countries of Japan, Canada, Mexico, South Korea, and Taiwan. | Open |  |  | 150000 | Not_Computer
Z-1033-2012 | 61129 | expiratory resistance valve, intranasal, for obstructive sleep apnea | OHP | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Provent -SR - Sleep Apnea Therapy-SR (standard resistance).Prescription only, Single Use;Packed 30 units per boxManufactured in the USA by Ventus Medical Inc.1301 Shoreway Road, Suite 425;Belmont, CA 94002Used in the treatment of obstructive sleep apnea. | ProventSR or Provent50 | 3 | February 15, 2012 | 2012 | Ventus Medical, Inc. | A discrepancy was found during inspection of a returned unit that the lot number and SKU on cartoon did not match. | PRODUCTION CONTROLS: Labeling Mix-Ups | Ventus Medical sent a Urgent Medical Device Correction Letter dated February 3, 2012, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.There is no health hazard associated with the affected devices or with the continued use of the affected devices.Ventus Medical wanted Customers to be aware of the problem so they can respond to patients if inquiries are received. A sample letter to patients was also provided, should physicians choose to provide it to patients. Questions may be directed to Ventus Medical Territory Business Manager at 650-632-4199. |  | 288 | 288 units | Nationwide Distribution including the states of OH, TX, IL, OR, CA, FL and NY. | Terminated | February 15, 2012 | 0 | 288 | Not_Computer
Z-1037-2012 | 61117 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Orthofix rotating cutter used in spinal fusion procedures with the FORZA Spacer System. Catalog Number 89-0452 (8mm Rotating Cutter), 89-0454 (10mm Rotating Cutter), and 89-0030 (FORZA Discectomy Case). Manual surgical instrument comprised of stainless steel and titanium. | Rotating Cutter | 2 | February 15, 2012 | 2012 | Orthofix, Inc. | The 8mm and 10mm Rotating Cutters may fracture during use, specifically during disc space preparation. If an 8mm or 10mm Rotating Cutter does Fracture during use, it may result in a surgical delay estimated to be 30 minutes or greater due to preparing the disc space using other instruments or methods and/or retrieving the instrument fragment. | DESIGN: Device Design | Orthofix mailed recall letters to the distributors on 02/06/2012. Orthofix does not sell or consign these instruments to hospitals. | No consumer action necessary | 52 | 52 | Nationwide and Puerto Rico | Terminated | June 11, 2012 | 117 | 52 | Not_Computer
Z-1038-2012 | 60942 | electrode, pacemaker, temporary | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Temporary Pacing Wire, BIPOLAR, Models BM, V, VE and VF, Rx Only, STERILE/EO, Single use only -- Mce (MEDICAL CONCEPTS EUROPE) --- Manufacturer ECM B.V. Oost-Om 54, P.O. Box 53, NL 5420 AB GEMERT --- Shipping or Unit Package: 20 individual units packaged per box with 6 boxes per case. --- Device Classification Name: Electrode, pacemaker, temporary. ---The device is supplied sterile and intended for single use. Used to pace the heart during or after open heart surgery. The electrodes are placed on the heart; the other end of the pacing wire is inserted through the skin of the patient with a thoracic needle. The thoracic needle can snap-off at the end where the pacing wire is attached, after snap-off the remaining pins are connected to an external pacemaker. For temporary atrial and ventrical pacing and sensing for contemplated implant duration of 7 days or less. | Temporary Pacing Wire | 2 | February 15, 2012 | 2012 | European Custom Manufacturing B.V. | The crimping operation during product production was not performed properly. As such, the pacing wire may separate from the connection pin, which causes non-functionality. | DESIGN: Process Design | ECM sent an "ECMbv RECALL OF MCE TEMPORARY PACING WIRES" letter dated May 2, 2011 to all affected customers. Customers were also informed of a potential problem via e-mail and/or phone between April 18 and 21, 2011. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to place the affected product on hold and to return the product to the firm. For questions on this matter, contact ECM or MCE Customer Service at 1-877-886-1985. | Contact the recalling firm for information | 2120 | 2,120 individual units (106 boxes) | Nationwide Distribution-including the states of AL, AR, CA, FL, GA, KY, NC, NV, NY, TX, VA, and WI. | Terminated | February 15, 2012 | 0 | 2120 | Not_Computer
Z-1048-2012 | 60599 | shunt, central nervous system and components | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | "***REF NL82059-01***INTEGRA NEUROSCIENCES***Touhy Needle***Non Sterile for Single Use Only***14 g, 3.5"***Plainsboro, NJ***"Needle for lumbar peritoneal shunt | INTEGRA NEUROSCIENCESTouhy NeedleNon Sterile for Single Use Only | 2 | February 16, 2012 | 2012 | Integra Neurosciences PR, Inc. | The current packaging configuration for the non-sterile shunt accessories could lead to the potential for confusion regarding product sterility. | PRODUCTION CONTROLS: Labeling Mix-Ups | Integra sent an Urgent Product Recall Notification letter dated November 28, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out and sign the attached Recall Acknowledgement Form noting the part numbers, lot numbers and quantity of the affected product in their inventory. Customers were asked to contact Integra LifeSciences Corporation International Customers Service at 609-936-5400 to arrange for Returns. If customers do not have any of the affected product in their inventory they are required to complete and sign the form and indicate they have NO affected product in their inventory. Customers should fax or email their Recall Acknowledgement Form to:609-750-4259 or LAPAC_CS@Integralife.comFor questions call 609-936-2495. |  | 320 | 320 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Canada, Italy, Australia | Terminated | October 29, 2013 | 621 | 1303 | Not_Computer
Z-1043-2012 | 60982 | test, time, partial thromboplastin | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent Product Usage:Liquid purified soy and rabbit brain phosphatides with plasma activator. An activated partial thromboplastin reagent with increased sensitivity to lupus-like inhibitors for use in the determination of the activated partial thromboplastin time (APTT) and related coagulation procedures. | Dade(R) Actin(R) FSL Activated PTT Reagent | 2 | February 16, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The listed lots of Dade(R) Actin(R) FSL Reagent show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice letter dated December 2011 to all customers of the affected lots. The letter identified the product, description of problem, and actions to be taken by the user. DadeÂŽ ActinÂŽ FSL customers were instructed to exercise one of two options: 1) If customers would like to continue working with their Dade(R) Actin FSL reagent lot (if one of the impacted lots) they must verify on a regular basis (monthly) that their therapeutic range is still valid, e.g. by using a Heparin containing control. In the case that it is not valid, they are to adjust their existing therapeutic range or establish a new one. 2) If customers decide not to continue working with the Dade(R) Actin(R) reagent lot (if one of the impacted lots) they are to contact their Siemens representative at 1-800-242-3233, option 5, for a replacement lot. Once they have established a new normal, therapeutic, and quality control ranges, they are to discard the remainder of the old lot. Customers were advised to complete the response form included to confirm receipt of Safety Notice and fax back to (302) 631-8467. If you have any questions or concerns, please contact the Technical Solutions Center at 1-800-242-3233, option 1. |  | 24672 | 24,672 Kits | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, Oman, Philippines, Poland, Portugal, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam. | Open |  |  | 24672 | Not_Computer
Z-1045-2012 | 60990 | kit, screening, staphylococcus aureus | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | SLIDEX Staph kit (250 Tests).The SLIDEX Staph-Kit is a rapid latex and red blood cell agglutination test for the identification of Staphylococcus aureus strains from culture media. | SLIDEX Staph kit | 2 | February 16, 2012 | 2012 | bioMerieux, Inc. | Reports of false negative results in organism identification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomerieux sent an "URGENT PRODUCT REMOVAL NOTICE" dated December 21, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer. The letter instructs the customer to stop using and to destroy the affected products. The firm recommends steps to follow when interpreting results. An acknowledgement and Product Replacement Form was enclosed for the customer to complete and return to the firm. Contact Biomerieux Customer Service representative at (800) 682-2666 for information regarding this notice. | Contact the recalling firm for information | 100 | 100 kits | Distributed to one customer in California. | Terminated | September 17, 2013 | 579 | 100 | Not_Computer
Z-0833-2012 | 57842 | cabinet x-ray, industrial | RCE |  |  |  |  | GE eXplore Locus Micro CT cabinet x-ray system model RS9-80.CT Scanner for laboratory mice or other research. | GE eXplore Locus Micro CT. | 2 | February 17, 2012 | 2012 | GE Healthcare, LLC | Product fails to comply with the federal performance standard for cabinet x-ray systems (Title 21 CFR 1020.40(c)(1)(i)). Specifically, the x-ray emission limit of 0.5 milliroentgen in one hour could be exceed. Highest emission measured was 3 mR/hr. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will notify purchasers of the problem, the affected product, instructions for users, and specify the maximum emission rate measured. GE Healthcare will bring defective systems into compliance free of charge. This will be done via Field Modification Instructions IFMI) 11101 which is scheduled for deployment on March 1, 2011 with an estimated completion date of July 1, 2011. | Contact the recalling firm for information | 21 | 21 | Worldwide Distribution | Terminated | September 06, 2012 | 202 | 21 | Not_Computer
Z-1059-2012 | 60972 | spinal vertebral body replacement device | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ENDURAMESH CORPECTOMY SPACER SYSTEM 013MM, 017MM, 023MM by varying heights. Dist. By: Lucero Medical, PO Box 67, Richfield, Ohio 44826Product Usage:Device Description: The Enduramesh is a vertebral body replacement system which provides structural support of the vertebral bodies following an anterior or far lateral corpectomy, aiding in spinal fusion. The device may be used with a cleared supplemental fixation system based on surgeon discretion and preference. The Enduramesh vertebral body replacement system consists of a single cylindrically shaped titanium mesh cage. The hollow core of the cage allows for packing bone graft and the circular holes throughout the device promote bone fusions. Various diameters and heights are available to accommodate variability among patients. Indications for Use: The Enduramesh is a vertebral body replacement system intended for use in the thoracic and lumbar spine (T1-S1) to replace a collapsed, damaged or unstable vertebral body due to tumor or trauma/fracture. The device is intended for use with either allograft or autoograft. The device may be used with a cleared supplemental fixation system based on surgeon discretion and preference. | ENDURAMESH CORPECTOMY SPACER SYSTEM | 2 | February 17, 2012 | 2012 | Lucero Medical LLC | During an FDA inspection the firm was informed that the device marketed is not the device cleared for the 510K associated with the device. The product is sold unlabeled. | DESIGN: Device Design | Lucero Medical LLC made a personal visit to facilities in regards to an Urgent Medical Device Recall as of January 12, 2012. The firm instructed the customer that they were removing and quarantining all product from the facility. If you have any further questions, please contact Lucero Medical, PO Box 69, Richfield, Ohio 44826, Director of Quality Assurance/Natalie Williams/ (440) 670-9583. |  | 190 | 190 units | USA Nationwide Distribution - in the state of OH. | Open |  |  | 190 | Not_Computer
Z-1055-2012 | 60513 | reagents,specific,analyte | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | ImmunoCAP Allergen Ri206, Cockroach American, Periplaneta americanaIt is an analyte specific reagent and no intended use, however, it is typically used in vitro quantitive assay for the measurement of cockroach allergen IgE in human serum or plasma. | ImmunoCAP Allergen Ri206 | 3 | February 17, 2012 | 2012 | Phadia US Inc | The firm found this lot of ImmunoCAP Allergen Ri206, American Cockroach does not contain sufficient levels of some cockroach allergens e.g. tropomyosin. Approximately 2% of the positive samples have a risk to fall below the Limit of Quantization (o .1 kUA/L). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | ThermoFisher sent an URGENT RECALL NOTIFICATION letter dated December 22, 2011, to all affected consignees. The letter identifed the product, the problem, and the action to be taken by the conignee. Consignees were recommended to cease testing on the affected lot of product and, if so desired, substitute with ImmunoCAP Allergen i6, German Cockroach, until ImmunoCAP Allergen Ri206 was in stock. For any questions regarding this notification, contact Technical Support at 800.346.4364, option 2. |  | 10 | 780; all sent to US, with 10 assays per sleeve | Nationwide to include: MN, VA, NC, MO, SC, TX, MA, MI, FL, IL, PA, MD, UT, CO, GA, CA, | Terminated | February 04, 2013 | 353 | 780 | Not_Computer
Z-1052-2012 | 60974 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica Bond Ready-to-Use Primary Antibody CD2 (11F11); 7 mL vial; an in vitro diagnostic primary antibody for use as part of an antibody panel for the characterization of T cell disorders; Catalog No. PA0271; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. For the qualitative identification by light microscopy of human CD2 molecule in formalin-fixed, paraffin-embedded tissue by immunohistochemical staining using the automated Bond system. | Leica Bond ReadytoUse Primary Antibody CD2 (11F11) | 2 | February 17, 2012 | 2012 | Leica Microsystems, Inc. | The primary antibodies to CD2 (clone 11F11) and CD7 (clone LP15) have been adulterated. | PRODUCTION CONTROLS: Labeling Mix-Ups | Leica Microsystems, Inc. sent a Medical Device Field Notification letter dated January 4, 2012, via first class mail on January 10, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the product and to return the product to the firm. Customers who have received the affected lot numbers of the listed products were requested to return any unused/part used reagents to the Buffalo Grove, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the account's Leica Representative. For further questions please call (847) 405-0123. |  | 121 | 114 - 7 mL vials | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | April 18, 2012 | 61 | 339 | Not_Computer
Z-1062-2012 | 61052 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa DX-D Imaging Package - (The detector power cables described in the Report of Correction and Removal are connected to portable detector panels used in the DX-D systems)Product Usage:DX-D Imaging Package ha the same intended use as the predicative devices: namely to provide diagnostic quality images to aid the physician with diagnosis. The device uses Agfas fla panel detectors with amorphous silicon scintillators and NX workstations with MUSICA imaeg processing to create radiographic images of the skeleton (including skull, spinal column and extremities), chest, abdomen and other body parts. The product is not indicated for use in mammography. | DXD Imaging Package | 2 | February 19, 2012 | 2012 | AGFA Corp. | The firm received a complaint that when using the DX-D Mobile DR X-Ray unit they noticed the detector cable was hot. | DESIGN: Device Design | AGFA Corp. sent an "URGENT FIELD SAFETY NOTICE" letter dated January 16, 2012 to affected consignees. The letter described the affected product, concerns, issues and actions to be taken. Customers have been requested to place the Varian warning label included in a highly visible and suitable position on the cable of the portable detector. Customers were instructed to complete and return the feedback form by fax to 864-421-1664, as soon as possible. For questions about this matter contact your local AGFA HealthCare Technical Support at 877-777-2432. | Contact the recalling firm for information | 42 | 42 units | Nationwide Distribution (USA) including the states of: FL, IL, IN, KS, MI, MN, MS, NE,NY, OH, SC and WV. | Terminated | May 10, 2012 | 81 | 42 | Not_Computer
Z-1061-2012 | 60863 | endoscope, neurological | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | ASSISTO Arm System Table-mounted self retaining instrument holder system consisting of stainless steel tubular, articulated arms. Product Usage:ASSISTO Arm Systems with a holding power of up to 50N are designed as universal support systems for ventilator tubing and breathing tubes. When equipped with the appropriate accessories these arm systems may also be used to hold endoscopes. They are usually proved with non-insulated clamp sockets. All ASSISTO Arm Systems with a holding power of up to 100N are designed as universal support and holding systems for endoscope holders and detractors or other accessories. These systems are attached to the operating table with an electrically insulated clamp socket and are designed to hold the accessory in the desired position for a short period of time (which may exceed 60 minutes. | ASSISTO | 2 | February 19, 2012 | 2012 | Geomed Medizin-Technik Gmbh & Co. | The product may fracture during surgical procedures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Geomed Medizin-Technik Product sent an Urgent Voluntary Medical Device Recall notification on December 9, 2012 and December 16, 2012 to affected customers. The letter identified the affected products, reason for recall and actions to be taken. Customers were instructed to (1) fill out the enclosed return form and fax back immediately at +49 (0)7461-935593, (2) return the affected products to the address provided, including the completed return form, (3) use the UPS account number 03710R provided to return the products free of charge and (4) immediately forward a copy of this recall notification to third parties to which you have distributed affected products. For questions and support, please contact Mr. Hanno Haug, VP -AMB at +49 (0)7461 9355-30 or by email: hhaug@geomed.de. |  | 1205 | 1,205 units | Worldwide Distribution - USA (nationwide) including the countries of: Australia, Belgium, China, Denmark, Finland, France, Germany, Great Britain, India, Italy, Japan, Lebanon, Libya, Lithuania, Luxemburg, Netherlands, Russia, Sweden, Spain and Turkey. | Open |  |  | 1205 | Not_Computer
Z-1070-2012 | 60828 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Children's Medical Ventures 1 ml Amber Oral Syringe. The product is a Liquid Medication Dispenser. | Children's Medical Ventures 1 ml Amber Oral Syringe | 3 | February 21, 2012 | 2012 | Benlan, Inc. | The tip caps may come loose from the syringe, potentially resulting in leakage of medication from the syringe. | PRODUCTION CONTROLS: Process Control | The firm initiated their recall of the specified lots of product in November, 2011 by sending a "Customer/Distributor Medical Device Correction" letter to their consignee. The letter requested they examine their inventory, quarantine any affected product and complete a verification and response form to be returned to the firm. Questions were directed to (905) 829-5004 ext 233. |  | 143250 | 143,250 | Nationwide Distribution, USA - including only the state of Pennsylvania | Terminated | March 19, 2013 | 392 | 143250 | Not_Computer
Z-1071-2012 | 60971 | prosthesis, elbow, constrained, cemented | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Discovery Elbow System Humeral Condyle Set- Hexalobular, REF 114700, Lot 797550 Alloy, Sterile, Biomet Orthopedics Warsaw IN 46581.Intended for primary and revision joint arthroplasty for use in cemented applications. | Discovery Elbow System Humeral Condyle Set Hexalobular | 2 | February 21, 2012 | 2012 | Biomet, Inc. | A Discovery Condyle Kit w/ Hexalobula was opened during an elbow joint replacement surgery, and was found to contain two male condyles. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | December 14, 2011 Urgent Medical Device Recall Notices were sent to Distributors and hospital OR managers describing the product, problem, risks associated with the recall, and actions to be taken. The notice included Instructions to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, prior to returning product by priority mail to the attention of Angi Dickson, Biomet , 56 East Bell Drive Warsaw, IN, USA 46580 with a copy of the response form. Receipt of the notice should also be confirmed by calling 800-348-9500, extension 3983 or 3009. Questions related to this notice should be directed to 574-371-3983 or 574-371-3009 Monday through Friday 8 AM to 5 PM. | Contact the recalling firm for information | 10 | 10 units | International Only Distribution -- Netherlands & UK. | Terminated | December 12, 2012 | 295 | 10 | Not_Computer
Z-1069-2012 | 61124 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | AngioDynamics SPECIAL ORDER Morpheus Smart PICC CT --- 4F Single Lumen 65cm --- Max CT Flow Rate: 4mL/sec --- Kit Contains: --- 1 each Nitinol Guidewire, 0.46mm (0.018in.) x 130cm --- Catalog No. [REF] 52110902 --- [Mfg symbol] AngioDynamics, Inc., 603 Queensbury, NY USA 12804.Indicated for short or long term peripheral access to the central venous system for intravenous therapy and power injections of contrast media. | Morpheus Smart PICC | 2 | February 21, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Morpheus Smart PICC label states the kit contains a Nitinol guidewire; however the kit contains a Platinum guidewire. | PRODUCTION CONTROLS: Labeling Mix-Ups | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated February 2, 2012 to the single customer. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs the customer to identify and segregate the affected product. A Morpheus Smart PICC Recall Reply Form was enclosed for the customer to complete and return via fax to 518-798-1360. Contact AngioDynamics Customer Service at 1-800-772-6446 for questions regarding this notice. | Contact the recalling firm for information | 25 | 25 units | Distributed to one customer in Wisconsin. | Open |  |  | 25 | Not_Computer
Z-1080-2012 | 61036 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CORFLO-Ultra Jejunal Tube with CORLOCK-Twoomey Y-Adaptor for use with the CORFLO-Max PEG; enteral feeding tube,10 Fr, 36" (91 cm), 3 Gram weighted bolus; individually packaged, 5 tubes per case; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #30-7361; For delivery of enteral nutritionProduct Usage: For delivery of enteral nutrition | CORFLOUltra Jejunal Tube with CORLOCKTwoomey YAdaptor | 2 | February 22, 2012 | 2012 | Corpak Med Systems | The wrong size Twoomey adapter was packaged with the Jejunal tube. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Corpak MedSystems telephoned customers on January 26, 2012 and sent a follow-up Recall Notification letter dated January 27, 2012 via e-mail reiterating the telephone conversations. Customers were informed of the urgent product recall, the problem and the actions to be taken. Customers were instructed to inspect inventory, segregate any affected product and contact Corpak at 1-800-323-6305 to arrange for the return and replacement of the units found. Customers were advised to complete and return the enclosed acknowledgment form, indicating the number of affected units found in their inventory. Distributors were requested to notify their customers. For questions contact Corpak MedSytems at 1-800-403-3400. |  | 145 | 145 units | Worldwide Distribution - USA (nationwide) including the states of Georgia, Indiana, Kentucky, Maryland, Michigan, Minnesota, Tennessee, Utah and the countries of to Sweden and New Zealand. | Terminated | March 21, 2012 | 28 | 145 | Not_Computer
Z-1081-2012 | 60748 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | LH Lithium Heparin tubes, Greiner bio-one, North America, 4238 Capital Drive, Monroe, NC 28110.Venous blood collection tubes. | LH Lithium Heparin tubes | 2 | February 22, 2012 | 2012 | Greiner Bio-One North America, Inc. | The lithium heparin tubes may have a contamination with EDTA potassium. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Greiner Bio-One sent a "RECALL OF LITHIUM HEPARIN TUBES WITH EDTA CONTAMINATION" letter dated December 13, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete an attached Product Disposition Site Confirmation Form and to fax the form to 800-726-0052. This form will confirm that customers have discarded/destroyed all affected product. The firm will replace the product after receipt of the form. Contact the firm at 800-515-8112 for questions regarding this recall. | Contact the recalling firm for information | 2082000 | 2,082,000 pieces | Nationwide Distribution. | Terminated | June 21, 2012 | 120 | 2082000 | Not_Computer
Z-1077-2012 | 60962 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic |  | Precision Aspheric Lens (PAL) Intraocular LensModel: EC-3 PAL,Serial Number: 1169381106, 1169391106, 1169421106The device is indicated for primary implantation in the capsular bag of the eye for the visual correction of aphakia in adult patients in whom a cataractous lens has been removed. | EC3 Precision Aspheric Lens Intraocular Lens | 3 | February 22, 2012 | 2012 | Aaren Scientific, Incorporated | The recall was initiated because Aaren Scientific has confirmed that a specific group of lenses were shipped without a complete expiry date showing on the label. | EXPIRATION DATING: Incorrect or No Expiration Date | Aaren Scientific sent a recall communication letter dated August 2, 2011, August 3, 2011, and August 30, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Aaren Scientific is requesting that you return the lenses and we will rework the product and return them back to you. Please let me know and I will issue an RGA number and inform you of our FedEx account number to reference.In the meantime, we are conducting an internal investigation regarding this issue. Once this has been closed out then we will notify you of our findings.For further questions please call (909) 937-1033 ext. 108. |  | 3 | 3 units | Worldwide Distribution - USA including CA and TN, and the countries of India and Lebanon. | Terminated | August 10, 2012 | 170 | 156 | Not_Computer
Z-1075-2012 | 60959 | antigens, all groups, streptococcus spp. | GTY | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Strep A Twist Rapid Test Cassette (K023766), IST-502T Twist Cassettes.Brand Name (Catalog Number): ACON Rapid Device Strep A Twist (IST-502T); CONSULT Diagnostics Strep A Twist, PSS (5005); Medi-Lab Performance Strep A Test-Twist, McKesson (32-502T); MooreBrand Strep A Twist Cassette, Moore (82793); and SP Strep A Rapid Twist Cassette Test, Cardinal Health (B1077-29).Type of Packaging (ie. box, plastic bag): Individually pouched devices (number of tests per kit varies from 5 to 50), the kits also include: disposable extraction test tubes, dropper tips, sterile swabs, 1 vial Strep A reagent A (10 mL): 2M Sodium Nitrite, 1 vial Strep A reagent B (10 mL): 0.4M Acetic Acid, 1 vial Strep A positive control (1 mL): Nonviable Strep A; 0.09% NaN3, 1 vial Strep A negative control (1 mL): Nonviable Strep C; 0.09% NaN3 in a labeled kit box.The Strep A Rapid Test Cassette and Twist devices are qualitative, lateral flow immunoassays for the detection of Strep A carbohydrate antigen in a throat swab, providing results within 5 minutes. The tests utilize antibodies specific for whole cell Lancefield Group A Streptococcus to selectively detect Strep A antigen in a throat swab specimen. The Cassette device uses a drop-wise application method while the Twist device has an onboard well to extract the throat swab specimen. The Strep A Rapid Test Devices are rapid chromatographic immunoassays for the qualitative detection of Strep A antigen from throat swab specimens to aid in the diagnosis of Group A Streptococcal infection. This test is indicated for professional and point of care use only. | Strep A Twist Rapid Test Cassette | 2 | February 22, 2012 | 2012 | Alere San Diego | Alere San Diego is expanding the scope of the voluntary recall of the Strep A Rapid Test Cassette and Twist devices from the single lot that was recalled in January 12, 2011 to all lots of product that were on the market within the expiry period. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 1, 2011 to its customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to retain the recall letter for their records as documentation that they have received the recall information. Customers were instructed to follow the directions and return the attached Verification Form. Customers were instructed to use the table in the Verification Form to document all inventory at any of their inventory locations. Customers may use multiple forms to document all sites, products, and lots. Alere will issue credit against PO#ASD811STA for any unused portion of the product based on completion and return of the attached Verification Form. Customers with any technical questions about the information contained in the notification were instructed to contact Alere by phone (877) 866-5309, fax (877) 866-9304 or email: ca2services@Alere.com. |  | 2797600 | 2,797,600 devices (228 lots) | Worldwide distribution: USA (nationwide) and countries of : Austria, Canada, Cyprus, Denmark, Finland, Germany, Greece, Hungary, Iceland, Italy, Poland, Portugal, Romania, Slovenia, Spain, Switzerland, and United Kingdom. | Terminated | June 28, 2012 | 127 | 4657030 | Not_Computer
Z-1084-2012 | 60967 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | UltraVision Detection System;Anti-Polyvalent, Alk-Phos/Fast Red (Ready to Use)Manufactured by: Lab Vision Corporation46360 Fremont Blvd, Fremont, CA 94538Product Usage: These products detect specific antibodies bound to an antigen in tissue sections. The specific antibody is located by a biotin-conjugated secondary antibody. This step is followed by the addition of streptavidin-enzyme conjugate that binds to the biotin preset on the secondary antibody. The specific antibody, secondary antibody, and streptavidin-enzyme complex is then visualized with an appropriate substrate/chromogen. | UltraVision Detection System, AntiPolyvalent, AlkPhos/Fast Red (Ready to Use) | 2 | February 22, 2012 | 2012 | Lab Vision Corporation | Firm was notified by vendor that a component in kits was not meeting internal specifications. Declination in product performance confirmed by Thermo Fisher. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Lab Vision Corporation of Thermo Fisher Scientific sent a Product Recall Notification letters dated January 11, 2012 to all affected customers. The letter identified the affected products, reason for recall and necessary actions to be taken. Customers were instructed to examine inventory, discontinue use, destroy any remaining inventory, follow normal procedure to order an alternate lot of the affected products, retain a copy of this letter for their laboratory records, if the affected products were further distributed provide a copy of this letter to them and complete the attached Customer Recall Response Form and return within 5 days. For questions regarding this information, U.S. customers should call Technical Support at 510-771-1595. Customers outside the U.S. should contact their local area Customer Service. |  | 31 | 31 kits in total | Worldwide Distribution - USA (nationwide) FL, IL, MA, TX and countries of Denmark, Turkey, Italy, Poland, Canada, Japan, Germany, Egypt and Greece. | Terminated | July 20, 2012 | 149 | 31 | Not_Computer
Z-1094-2012 | 60824 | labware, assisted reproduction | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | ORIGIO MIDATLANTIC DEVICES CRYOPETTE (r) Rx only Single use only CRY-PETTE-CL LOT 123456 2013-10-10 Exp. date *** MADE IN THE USA STERILE R ***Origio, Inc. 3400 Hunter's Way, Charlottesville, VA 22911 www.origio.comThe Cryopette (r) is a cryopreservative storage device that is intended for use in vitrification procedures to contain and maintain human blastocyst stage embryos. | Cryopette | 2 | February 23, 2012 | 2012 | ORIGIO HUMAGEN PIPETS INC. | Cryopettes were rupturing during the warming process following vitrification. | DESIGN: Device Design | ORIGIO sent a Field Safety Corrective Action letter dated November 21, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ORIGIO feels it is important to eliminate this risk completely. As soon as we are confident that this danger of rupture has been eliminated, we will make Cryopettes available again. Please use the attached UPS call tag to return any unused Cryopettes to ORIGIO Inc. Upon receipt, we will issue a credit to your account.We sincerely apologize for the problems this issue has caused in your laboratories. For further questions please call (434) 979-4000. |  | 73850 | 73850 Cryopettes | Worldwide Distribution - U.S. included states of AL, AR, DE, PR, NY, MA, NH, VT, IL, CT, NJ, PA, VA, WV, NV, NC, SC, GA, FL, OH, MD, MS, IN, MI, PR, WI, MN, SD, MT, MO, TX, CO, UT, AZ, CA, HI, OR, AK, WA and in countries of Turkey, Saudi Arabia, Kuwait, Jordon, So. Korea, Taiwan, Brazil, So. Africa, Oman, Kazakhstan, Honduras, Slovenia, Serbia, Poland, Czech Republic, Romania, Croatia, Ukraine, Netherlands, Bulgaria, Hungary, Russia, Lebanon, Singapore, Malaysia, Vietnam, China, Thailand, Nigeria, Japan, Spain, Greece, Canada, Denmark, France, United Kingdom, Italy, Germany, and Australia, | Terminated | February 04, 2013 | 347 | 73850 | Not_Computer
Z-1097-2012 | 61206 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | BD Oral Dispensing Syringe Oral Syringe Tip Caps various sizes ( 1 mL, 3 mL, 5 mL. 10 mL BD, Franklin Lakes, NJ 07417 USA Made in USA www.bd.comRef Numbers 305207, 305218, 305210, 305220, 305208, 305218, 305219Product Usage:Used for the oral administration of medications | BD Oral Dispensing Syringe Oral Syringe Tip Caps | 3 | February 23, 2012 | 2012 | Becton Dickinson & Company | There are complaint reports for loose tip caps in various types of oral syringes | DESIGN: Device Design | Becton Dickinson (BD) sent an Urgent: Product Recall letters dated October 5, 2011 to all affected consignees. The letter identifies the affected product, problem and actions to be taken. Customers are instructed to follow the instructions on how to handle the affected product: 1) Check their inventory to determine if they have any Oral Syringe Tip Caps from the impacted lot numbers. 2) If they have any of the impacted product return the Oral Syringe Tip Caps to BD for replacements as directed on the attached recall packing list. 3) If they use the Oral Syringe Tip Caps and require replacement, follow the instructions on the attached form. 4) If they have no inventory of the affected lots, indicate this on the website. |  | 22827500 | 22,827,500 units | (USA) Nationwide Distribution | Terminated | February 01, 2013 | 344 | 22827500 | Not_Computer
Z-1104-2012 | 61048 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits containing Ethicon Ethilon Black Nylon Nonabsorbable Surgical Sutures in new Tyvek packaging; surgical trays; Cardinal Health, McGaw Park, IL 60085; The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then ethylene oxide sterilized and are ready to be used by the customer for the specific procedure. The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then ethylene oxide sterilized and are ready to be used by the customer for the specific procedure. | Presource Custom Sterile Surgical and Procedure Kits | 2 | February 24, 2012 | 2012 | Cardinal Health | The Presource kits were assembled with Ethicon Sutures in new Tyvek packaging which could no longer be placed in kits that will undergo ethylene oxide sterilization. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 26, 2012 to the affected consignees via UPS next day air. The letter identified the product, problem, and actions to be taken. Enclosed with the letter were recall labels, instructing the customer not to use the affected sutures and to remove and destroy the sutures upon opening the kit, to affix recall label to the front side of each affected Presource kit and shipping carton currently in their inventory. The customers were also requested to complete and fax back to Cardinal at 847-689-9101 or 614-652-9648 regardless of whether or not you have inventory. For component credit/replacement the consignees were instructed to contact the following: Presource Sales Operations at 800-766-0706; Government Customer Service at 800-444-1166; and Distributor Management Group at 800-635-6021. For any questions, require additional labels or desire special assistance relating to this recall contact Cardinal Health Quality Systems, 800-292-9332. |  | 18300 | 18,300 kits | Worldwide distribution: USA (Nationwide) including the states of: Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, West Virginia, Wisconsin and Washington; and country of Canada. | Open |  |  | 18300 | Not_Computer
Z-1101-2012 | 60981 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Oral-Eze (R) Oral Fluid Collection System, manufactured by Capitol Vial, Inc., Auburn, AL 36832Product Usage:IVD for drugs of abuse | OralEze (R) Oral Fluid Collection System | 2 | February 24, 2012 | 2012 | Capitol Vial Inc | Low buffer fill volumes on Oral-Eze, an oral fluid collection system. | DESIGN: Device Design | Thermo-Fisher Scientific/Capital Vial sent a Recall Notification letter dated October 7, 2011 to the sole consignee. The letter identified the affected product, problem, result impact and necessary actions to be taken. The consignee was instructed to discontinue use of the affected product and follow the necessary actions provided. If you have any questions regarding this information, call Michael Hect, Manufacturing Manager through Customer Service at (800) 772-8871. | Contact the recalling firm for information | 135000 | 135,000 units | (USA) Nationwide Distribution - in the state of KS | Terminated | August 24, 2012 | 182 | 135000 | Not_Computer
Z-1103-2012 | 61082 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Medical Zippie TS wheelchairTo empower physically challenged person by providing a means or mobility. | Zippie TS wheelchair | 2 | February 24, 2012 | 2012 | Sunrise Medical (US) LLC | One Zippie TS Wheelchair was made per customer order was shipped with transit brackets for use a seat in a motor vehicle, however the chair was not configured with a back tested seat for motor vehicle use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Sunrise Medical contacted customer by telephone on January 17, 2012 to arrange for removal of transit brackets. The Sunrise Sales rep then visited the customers on January 18, 2012 and removed the bracket from the chair making it appropriate for intended use.If you need any additional information or have any questions, call 559-348-2572 or email: laurie.roberts@sunmed.com. | Contact the recalling firm for information | 1 | 1 chair | Nationwide distribution: OK only. | Terminated | February 27, 2012 | 3 | 1 | Not_Computer
Z-1100-2012 | 61235 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, part number: N97F. The intended use is for cementation of porcelain, resin, and metal-based inlays, onlays, crowns, bridges, and posts. | Breeze SelfAdhesive Resin Cement | 2 | February 24, 2012 | 2012 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the Breeze Self-Adhesive Resin Cement exhibits faster than expected gel and set times. If the work time is exceeded when using this material, it could possibly result in inadequate bond strength. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An "Urgent Medical Device Recall" letter was sent on 12/17/2010 via USPS 1st class mail to all their customers. The letter provides the customers with an explanation of the problem identified and the action to be taken. Customers were instructed to complete and fax back the enclosed recall return form. Customers with questions were instructed to call (800) 551-0283. | Contact the recalling firm for information | 622 | 622 units total (12 units in US) | Worldwide Distribution - Nationwide (USA) and the countries of Vietnam, Jordan,Turkey, and Italy. | Terminated | February 24, 2012 | 0 | 622 | Not_Computer
Z-1106-2012 | 61049 | chair, with casters | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Premier Care Recliner Model 5570. Brochure: "***Premier Care Recliner 5570/5574***Winco***ISO 9001 2008 Certified***Key Features***Urethane armrests***Multiple side-table options***3-position lock-bar*** Fully upholstered***OPTIONAL Heat or Heat & Massage***OPTIONAL LiquiCell***1.800.237.3377***5516 SW 1st Lane***Ocala, Florida 34474***www.wincomfg.com***Proudly Made in the USA.***"Press on type label: "***Manufactured By Winco Ocala, FL 34474***Model No. 557***Serial No. 557-...***"Press on type label: "***Model No. 557***Serial No.557-...***".Long-term multi-positional recliner. | Premier Care Recliner Model 5570 | 2 | February 26, 2012 | 2012 | Winco Mfg., LLC | On 10/17/2011 Winco Mfg. LLC, Ocala, FL initiated a recall of The Premier LifeCare Recliner and Premier Care Recliner Models 5400 and 5570 distributed between January 4, 2011 and October 12, 2011. The Position Lock Pin(s) could potentially fail if exposed to excessive force when the chair is in the Locked Recline position. This failure will result in an inoperable Position Lock Bar feature. | DESIGN: Device Design | Winco Mfg. sent a "POSITION LOCK PIN RECALL NOTIFICATION" letter dated October 24 , 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an enclosed response form and provide product disposition. Winco will ship retro-fit parts to install on the affected product. Instructions will explain step by step, how to implement the fix. Contact Winco at 800-237-3377 for questions regarding this recall. |  | 79 | 79 chairs | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, KS, MA, MI, MO, MS, NC, NM, NY, OK, PA, SC, SD, TN, TX, VA, WI, and the country of Canada. | Terminated | May 04, 2012 | 68 | 133 | Not_Computer
Z-1110-2012 | 60790 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM/CT 670The GE Discovery NM/CT 670 system is a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy. The GE Discovery NM/CT 670 system, combining Nuclear Medicine (NM) and Computed Tomography (CT) systems, is intended to produce: NM System: General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and imaging features designed to enhance image quality. | GE Discovery NM/CT 670 | 2 | February 27, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with the collimator cart and lead cover of your Discovery NM/CT 670 and/or Discovery NM 630 system that may impact operator safety. 1.The cart locking pin release and cart disengagement could result in the collimator falling down and could lead to operator or service personnel injury.2.Another issue is that the | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated January 13, 2012 to its Consignees/Customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to continue to use product, ensure that the collimator exchange monitor the collimator cart are properly locked in place; pay attention to any unexpected cart motion during the exchange procedure; cease the process immediately if the cart has moved from the docking spot, dock cart in place again, and complete the exchange; and if a 3rd party service engineer is servicing rather than the firm, GE, advise according to the instructions in letter.A GE Healthcare Service representative will perform the required hardware update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 79 | 79 | Worldwide distribution: USA (nationwide) including states of: CA, FL, GA, IL, NJ, NY, NC, TX, and WI; and countries of: AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, FRANCE, GERMANY, INDIA, ISRAEL, ITALY, KOREA, JAPAN, NETHERLANDS, NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, UNITED KINGDOM. | Terminated | December 06, 2012 | 283 | 89 | Not_Computer
Z-1109-2012 | 61006 | culture media, antimicrobial susceptibility test | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL[TM] Lysine Iron Agar, catalog #211363,500 g bottle 2014-03-31. Lysine Iron Agar was developed for use in the differentiation of enteric organisms based on their ability to decarboxylate or dominate lysine and to form hydrogen sulfide. Becton, Dickinson and Company, Sparks, MD 21152 USA 38800 Le Pont de Claix, France.Lysine Iron Agar was developed for the use for differentiation of enteric organisms, in particular Salmonella species due to its ability to decarboxylate or deaminate lysine and form hydrogen sulfide. | BBL TM LYSINE IRON AGAR | 2 | February 27, 2012 | 2012 | Becton Dickinson & Co. | Lysine Iron Agar may not produce the correct biochemical reactions with Quality Control organisms over the shelf life of the product.. | DESIGN: Process Design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated January 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining packages. A form was attached for customers to complete and return via fax to: Attention: Regulatory Compliance at 410-316-4258. Contact the BD Technical Services Department at 1-800-638-8663 for questions regarding this recall. |  | 33 | 33 bottles | Worldwide Distribution-USA (nationwide) including the states of AL, AR, CA, CO, CT, DC, FL, GA, IL, IN, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TX, VA, and WI, and the countries of Belgium, Canada, Chile, China, Columbia, Costa Rica, Mexico, El Salvador, Ecuador, Japan, Manila, Singapore, Peru, Taiwan, and Venezuela. | Terminated | March 30, 2012 | 32 | 297 | Not_Computer
Z-1117-2012 | 60831 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROSÂŽ Chemistry Products PHBR Slides, REF 822 1384 --- COMMON/USUAL NAME: VITROS PHBR (Phenobarbital) Slides --- Ortho-Clinical Diagnostics, Inc. Rochester, NY --- Ortho-Clinical Diagnostics High Wycombe, UK --- This is an in vitro diagnostic. For in vitro diagnostic use only. VITROSÂŽ Chemistry Products PHBR Slides quantitatively measure Phenobarbital (PHBR) concentration in serum and plasma using VITROS 250/350/950 and 5,1 FS Chemistry Systems and the VITROS 5600 Integrated System. | VITROS Chemistry Products PHBR Slides | 2 | February 28, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in control fluid values when using various lots of VITROSÂż Chemistry Products PHBR Slides (Phenobarbital) from Coating 0053. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ortho Clinical Diagnostics sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated June 23, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using and discard all remaining lots of the affected product. A Confirmation of Receipt Form was attached so customers may complete and return to the firm. The firm recommends that customers consult with their Laboratory Medical Director and requesting physician to resolve any concerns they may have regarding previously reported results. Contact Customer Technical Services at 1-800-421-3311 for questions regarding this notification. | Contact the recalling firm for information | 2977 | Domestic - 2345 sale units, Foreign - 632 sales units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open |  |  | 2977 | Not_Computer
Z-1112-2012 | 61123 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics BENEPHIT CV Infusion System Targeted Renal Therapy Infusion System --- Catalog No. [REF] 60035 --- Manufactured in USA --- [Mfg symbol] AngioDyanamics, 603 Queensbury Avenue, Queensbury, NY USA 12804.Intended to facilitate targeted renal therapy, or TRT, the delivery of physician-specific agents to the kidneys via the renal arteries. | BENEPHIT CV Infusion System | 2 | February 28, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Benephit CV Infusion System contains a Sheath Y Connector that has the potential to develop cracks. | DESIGN: Component Design/Selection | Field Safety Notices (FSNs), entitled Urgent - Medical Device Recall, were sent to consignees on 02/06/2012 via USPS certified mail, return receipt requested. The letters identified the affected devices, the issue, hazard involved, and actions to be taken by the customer and AngioDynamics. The FSN directs consignees to identify and segregate the recalled lots, returning any affected product to AngioDynamics, Inc., Queensbury, NY 12804. The Recall Reply Form should be completed and faxed to 518-798-1360. --- If customers have distributed the affected product/lot, they are instructed to inform their customers of this recall action immediately, and then return the recalled units. --- If customers have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please contact the local sales representative or call Julie Blair, Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. Customers may also e-mail questions to customerservice@angiodynamics.com. |  | 11 | 11 units | Nationwide Distribution -- Florida, Iowa, Maryland, New York, and Texas. | Open |  |  | 19 | Not_Computer
Z-1115-2012 | 61239 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Brand Name: Fusio Liquid Dentin, part number N21SB, N21SC.Product Usage:The intended use of this device is as a self-etch, self-adhesive flowable resin composite indicated for light-cured dental restorations. Applications include self-etch base liner, restorations of carious lesions, Class I, III and V cavity preparations and pitt and fissure sealant. | Fusio Liquid Dentin | 2 | February 28, 2012 | 2012 | Kerr Corporation | A recall was initiated because Pentron Clinical has confirmed that the Fusio Liquid Dentin is not as flowable as ideally expected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pentron Clinical sent a Medical Device Recall letter dated September 8, 2010 (via USPS 1st class mail) to all affected customers. The letter identified the affected product, the reason for the recall and the actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete and fax back the enclosed recall return form. Customers were advised to contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 382 | 382 units total (207 units in US) | Worldwide Distribution - USA (nationwide) in the countries of Austria, Bulgaria, France, Great Britain, Netherlands, Japan, Lebanon, Malaysia and South Africa. | Terminated | February 28, 2012 | 0 | 382 | Not_Computer
Z-1116-2012 | 61031 | fluorometer, for clinical use | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Genetic Screening Processor (GSP) Product code: 2021-0010. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. | Perkin Elmer | 2 | February 28, 2012 | 2012 | PerkinElmer Life and Analytical Sciences, Wallac, OY | GSP instruments with defective washer modules may cause a false negative screening result being reported for T4, TSH and IRT. For 17-OHP, a false positive result may be generated if the failure occurs for a normal sample. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Perkin Elmer issued an "Urgent Medical Device Recall" letter dated 1/25/12 to all customers and informed them about the possible issue and requested permission for service personnel to check their instrument. The customers were instructed that the instrument could not be used and the washer modulerequired replacement. Customers were told to expect to be contacted by the firm's local representative for further details. |  | 10 | 10 units | Worldwide Distribution - USA, including the states of Florida, Ohio, Oregon, Texas and the countries of Brazil, Belgium, France, Germany, and the UK. | Terminated | March 06, 2013 | 372 | 10 | Not_Computer
Z-1126-2012 | 61184 | needle, fistula | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems (a Nxstage Company) Hemodialysis Fistula Needle Set with MasterGuard Ant-Stick Needle Protector 15 GA x 1"Ref. Number: D9-2007MG.Medisystems AVF with MasterGuard anti-stick needle protector is indicated for use as a vascular access device in hemodialysis or other extracorporeal procedures. It is intended for use in conjunction with a blood tubing set equipped with a standard medical male luer-locking connector. | Hemodialysis System and Accessories | 2 | February 29, 2012 | 2012 | Medisystems Corp | Mislabeled outer carton: Outside shipping cartons is stamped 15Ga x 1", product contained within properly labeled as 17Ga x 1" needle | TRAINING: Employee Error | Medisystems Corporation issued notice of recall to distributors on January 31, 2012. Distributors to notify affected end users by letter. Each letter was accompanied by a response form and instructions on how to return the response form. End users were given the option of returning involved product for credit or replacement. Contact Medisystems Quality Assurance Manager at (978) 332-5973 for information. |  | 20750 | 20,750 units (83 cases) | USA (nationwide) including the states of AZ, CA, CO, OR, NV, NY, TN, and WA. | Terminated | September 12, 2012 | 196 | 20750 | Not_Computer
Z-1129-2012 | 59285 | research laser | REK |  |  |  |  | DirectFRAP (Direct Flourescence Recovery After Photobleaching) Imaging System is an extension unit for Axio Observer that is used for microscopic analyses of dynamic processes within living cells. The DirectFRAP module is a component used in conjunction with Laser Scanning Microscopes, Non-Medical Devices for Research Use. The DirectFRAP Imaging System is a Class 3B laser system. The DirectFRAP Imaging System is a laser manipulation slider which allows coupling of laser light simultaneously with the conventional fluorescence illumination into the reflected-light beam path of the microscope system. Depending on the application, the laser light is used for the targeted bleaching of fluorophores for FRAP experiments or for the conversion of fluorophores within the cell. | DirectFRAP (Direct Flourescence Recovery After Photobleaching | 2 | February 29, 2012 | 2012 | Zeiss, Carl Inc | The manufacturer confirmed that the opto-semiconductor sensors can be triggered by exposure to incandescent lighting. The safety interlock is in place to prevent unintended exposure to Class 3B laser radiation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carl Zeiss MicroImaging LLC, through trained personnel, first visited each user site, provided additional user instructions, re-labeled the components to provide adequate instruction and warning and further placed a label over the access to the locking screw which secures the component to the microscope. Carl Zeiss Inc. issued "IMPORTANT LASER SAFETY NOTICE: URGENT: FIELD CORRECTIVE ACTION" notices to each user by express courier in late April 2011, Carl Zeiss personnel will confirm that the relabeling has been performed and updated instructions for use have been provided. For questions, please contact the company.The firm will implement 2 actions. Service Representatives will visit each user site, provide additional user instructions, relabel the components to provide adequate instruction and warning, and add a label over the access to the locking screw which secures the component to the microscope. Secondly, when new replacement sensors are available, Carl Zeiss field service technicians will perform the component replacements. Modified componentsare anticipated to be available within 6 months and replacement completed by April 2012. |  | 11 | 11 units | Worldwide Distribution -- USA (nationwide) | Terminated | December 11, 2012 | 286 | 111 | Not_Computer
Z-1122-2012 | 61187 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Mallinckrodt Optistar Elite 60 mL Syringe with Lateral Flow Needle, Liebel-Flarsheim Company, Cincinnati, OH 45237. Made in Mexico. Devices packaged in Plastic and Tyvec pouch, with 50 pouches per case.Sterile, The purpose of this product is the injection of contrast media by means of power injector system during a diagnostic procedure. | 60 mL Syringe w/Lateral Flow Needle | 2 | February 29, 2012 | 2012 | Mallinckrodt Inc | The machine used to package the product was missing a tool, which could prevent the packaging from sealing properly, causing the sterility of the product to be compromised. | PRODUCTION CONTROLS: Packaging Process Control | Covidien sent an Urgent Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product immediately. Customers were to use the attached business response form to report the quantity of each lot that is currently in their stock. If customers did not have any of these lots in stock, they should indicate by checking the appropriate box on the business report form. When completed, customers should email the completed form to covidien2600@stericycle.com or fax to 888-257-7910. For any questions customers should call 866-761-9571. |  |  |  | Worldwide Distribution - USA including Distributed worldwide: USA including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, Puerto Rico and the countries of Canada, Australia, Belgium, Brazil, Chile, Ecuador, Guatemala, Mexico, Panama, Singapore, and Uruguay. | Terminated | December 24, 2013 | 664 |  | Not_Computer
Z-1133-2012 | 61185 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | BardPort M.R.I. Implanted Port with Attachable 6 Fr. ChronoFlex Open-Ended Single-Lumen Venous Catheter and Peel-Apart Introducer Kit. Subcutaneous implanted intravascular infusion port and catheter is used to provide repeated access to the vascular system for the delivery of medications, intravenous fluids, parenteral nutrition solutions, and blood products. They are also indicated for the withdrawal of blood samples. | BardPort M.R.I. Implanted Port | 3 | March 01, 2012 | 2012 | Bard Access Systems | The end flap label on the kit incorrectly describes the product as "Bard Port" MRI DUAL Implanted Port instead of a "Bard Port MRI Implanted Port" as correctly indicated on the main label of the kit. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | A "Labeling Discrepancy Notification" letter was sent to US consignees via FED EX on February 16, 2012 and on February 20, 2012 for the foreign consignee explaining the labeling discrepancy and providing a visual example of the incorrect label. The customers were given the option of continuing to use the product with the understanding the end flap label was incorrect or they could return the affected products for replacement (Call 1-800-290-1689 to arrange a replacement). | Contact the recalling firm for information | 545 | 545 units (Lot RESF0856 - 150 units; Lot RESH0067 - 295 units; Lot RESH1150 - 100 units) | Worldwide Distribution -- USA, including the states of CA, CT, FL, KS, IL, MI, MN, OH, OK, SD, TN, TX, VA, WA and the country of Belgium. | Terminated | March 30, 2012 | 29 | 545 | Not_Computer
Z-1132-2012 | 60888 | bed, air fluidized | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Clinitron Rite Hite Bed headboard and foot side rails.The Clinitron Rite Hite Air Fluidized Therapy Unit is a patient management system that combines air fluidized therapy and low airloss therapy technologies together on a low-height frame. The lower body section provides air fluidized therapy, and the upper body section provides low airloss therapy. | Clinitron Rite Hite Bed headboard and foot side rails. | 3 | March 01, 2012 | 2012 | Hill-Rom, Inc. | The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed headboards failed to meet the safety standard force requirements. The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed outer rail, which is extruded Aluminum tubing, is intended to be one piece joined together by a weld at the lower center. The parts receiv | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hill-Rom, Inc. sent a Modification Notice Urgent Medical Device Correction Notification dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The modification notice instructed all customers to find and replace all headboards and food side rails that do not meet design intent. All technicians will inspect all affected product in the service centers and all at home and rite hite beds for an "x" or "L" stamp. The technicians will inspect the beds that are still in use during scheduled maintenance or before each rental. All will be completed before 01-JUL-12.For support call 1-800-445-3720. |  | 173 | Foot Siderail = 70; Headboards = 103 | Worldwide Distribution -- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV, and WY. and the country of Canada. | Terminated | December 05, 2012 | 279 | 173 | Not_Computer
Z-1130-2012 | 60980 | enzyme immunoassay, barbiturate | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Product Brand Names: TriageÂŽ Drugs of Abuse Panel 8 Test Kit (K924476), TriageÂŽ Drugs of AbusePanel Plus TCA (K973784), & TriageÂŽ Drugs of Abuse Panel Plus PPX (K014247)Model Number: 90008, 92002, & 92000Type of Packaging (i.e. box, plastic bag): For all affected device types: 25 individually pouched devicesin a labeled kit box.The TriageÂŽ Drugs of Abuse Panel plus Tricyclic Antidepressants utilizes a patented immunochemistry,ASCENDTM MULTIMMUNOASSAYÂŽ (AMIATM). Each AMIATM assay is a competitive binding immunoassayin which a chemically labeled drug (drug conjugate) competes with drug which may be present in the urine forantibody binding sites. After a brief incubation, the reaction mixture is transferred to the membrane in theDetection Area. Free drug conjugate that is displaced from antibody binding sites by drug in the urine, binds to azone of monoclonal antibody that is immobilized on the membrane. The membrane is washed to remove theunbound conjugate and clear the background. Test results are visually read.The TriageÂŽ Drugs of Abuse Panel plus Tricyclic Antidepressants is an immunoassay used for the qualitative determination of the presence of the major metabolites of drugs of abuse, (Phencyclidine, Benzodiazepines, Cocaine Metabolite, Amphetamines, THC, Opiates, Barbiturates), and Tricyclic Antidepressants in urine. | Drugs of Abuse Screen | 2 | March 01, 2012 | 2012 | Alere San Diego | Alere San Diego received a rise in customer complaints regarding false positive Barbiturates (BAR) results on the Triage Drugs of Abuse Panel Plus TCA. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Alere sent a "PRODUCT NOTIFICATION" letter dated January 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return to the firm. Contact Technical Services at 1-877-308-8287. | Contact the recalling firm for information | 693054 | 693054 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, EU, Guatemala, Japan, Lebanon, and Saudi Arabia. | Terminated | July 05, 2013 | 491 | 693054 | Not_Computer
Z-1095-2012 | 61070 | bronchoscope accessory | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Inspira AIR Balloon Dilation System,size 18 x 40;Manufactured by Acclarent, Inc.1525-G O'Brien Drive, Menlo Park, CA 94025Product Usage: Dilation of airway tree. | Airway balloon catheter | 1 | March 01, 2012 | 2012 | Acclarent, Inc. | Acclarent received reports of difficulty deflating the balloon during the procedure, which could potentially result in airway obstruction. | DESIGN: Component Design/Selection | Acclarent sent an "Urgent Voluntary Product Recall" letter dated January 31, 2012 by US mail, return receipt requested to all affected customers. Additionally, customers known to have scheduled procedures using the affected product were notified by phone to advise them of the recall. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use and return any remaining in stock per the instructions provided. The letter provides additional information and photographs illustrating how to identify the recalled product. Customers were advised to complete the enclosed Business Reply Card regardless of whether they have or do not have product subject to this recall in their possession and mail to Stericycle. Customers were instructed to distribute this information to all staff within their department who uses the affected product. For further questions call 1-866-781-1173 or contact your Acclarent Sales Representative. |  | 633 | 633 units | Nationwide Distribution (USA) only sixe 18 x 40Worldwide distribution other sizes.Distirbution to France, Canada, Saudi Arabia | Terminated | August 29, 2013 | 546 | 633 | Not_Computer
Z-1137-2012 | 61073 | stimulator, muscle, powered | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | PerfectorProduct Usage: Muscle stimulation | Perfector | 2 | March 01, 2012 | 2012 | Ion Genius, Inc. | Wires do not meet performance standards under 21 CFR 898 and have unprotected electrode configurations. Additionally, devices were sold/distributed without 510(k) or PMA | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ion Genius sent a Medical Device Product Recall letters to all affected customers on January 6, 2012. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to check their inventory for the affected products and return immediately to Ion Genius for replacements. Customers returning affected products should call the firm so that a shipping label can be provided.Customers were instructed to be completed the Response Form and returned to Ion Genius by fax at 808-395-0787. For questions and concerns contact Xanya Sofra-Weiss at 808-222-2664 or email x@arasysperfector.com. | Contact the recalling firm for information | 72 | 72 cables in total, amounts of other devices not reported | Worldwide Distribution - (USA) Nationwide Distribution | Open |  |  | 72 | Not_Computer
Z-1136-2012 | 61221 | bracket, metal, orthodontic | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Brand name: Orthos CM Bracket, Orthodontic Metal Bracket,Part Numbers: 454-1210, 454-0210.The intended use of this device is a fixed attachment on a tooth which holds an archwire during orthodontic treatment. | Orthos CM Bracket, Orthodontic Metal Bracket, | 3 | March 01, 2012 | 2012 | Ormco Corporation | The recall was initiated because Ormco has confirmed that manufacturing error resulted in the Orthos CM Brackets being incorrectly manufactured with a torque of -9Âż. The packaging was labeled with +9Âż torque; however the brackets contained inside the packaging are actually Orthos CM Brackets with a torque of -9Âż. | PRODUCTION CONTROLS: Process Control | Ormco Corporation sent an Urgent Field Safety Notice dated February 13, 2009, to all affected customers. The recall communication for consignees in the United States, Canada, Colombia, HongKong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa and Croatia was sent on February 13, 2009 via USPS First Class Mail. The recall communication for consignees in Japan was sent on February 17, 2009. European consignees were sent recall notification on February 23, 2009. The Australian consignees and New Zealand consignees were the sentrecall communication on February 25, 2009. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 to receive an RMA number. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Avenue, Glendora, CA 91740.Customers were instructed to label their return product "RECALLED PRODUCT Attention: Customer Returns". Customers were instructed to complete the enclosed Return Form and return it by fax to (909) 962-5605.For questions regarding this recall call 909-962-5600. |  | 304 | 304 units | Worldwide Distribution - USA (nationwide) and worldwide to Canada, Colombia, Hong Kong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa, Croatia, Japan, France, Bulgaria, Germany, Estonia, United Kingdom, Hungary, Ireland, Italy, Netherlands, Australia, and New Zealand. | Terminated | March 02, 2012 | 1 | 304 | Not_Computer
Z-1143-2012 | 61100 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS AVL Poly Tibial Bushing SetOSS AVL Tibial Bushings are designed to hold the anti-lux yoke in place for OSS procedures. | OSS AVL Poly Tibial Bushing Set | 2 | March 02, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after complaints were received where OOS AVL Bushing Sets were opened during OSS knee replacement procedures, and were found to contain two tabbed bushings. | DESIGN: Device Design | BIOMET Orthopedics sent an URGENT MEDICAL DEVICE recall notice dated February 7, 2012, to all affected consignees. The recall notice instructed consignees to immediate locate and remove the devices from circulation and to fax back the response form to 574-372-1683. If product was further distributed, consignees must notify hospital personnel of this action via the enclosed notice. Consignees were charged with the location and return of the affected product. For questions regarding this recall call 574-372-3983 or 574-371-3009. |  | 27 | 27 units distributed US and foreign | Worldwide Distribution - USA including IN. KS, NJ, CA, and OHInternationally: The Netherlands | Terminated | February 06, 2013 | 341 | 27 | Not_Computer
Z-1147-2012 | 61125 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Titanium SURFIX ALPHA SCREWS, a component of the Newdeal Hallu Lock Plate System. Surfix Alpha Variable angle locking system exists in different diameters and lengths. The product can be sold as a replenishment part individually packaged sterile, or as one (1) unit which is kitted within an instrumentation tray at which point, the product is non-sterile until the tray in which it is kitted is sterilized a the designated facility. The product is indicated for use in fixation of fractures, osteotomies or arthrodesis of the first metatarso-phalangeal joint. | Titanium Surfix Alpha Screws | 2 | March 02, 2012 | 2012 | Integra LifeSciences Corp. | Manufacturing error. The part number which corresponds to the length dimension of : 14 mm. was incorrectly etched on to screws with the actual measurement of: 12 mm. | PRODUCTION CONTROLS: Labeling Mix-Ups | Integra contacted the Sales Representatives and All USERS on December 21, 2011, with a 'SUBJECT: MEDICAL DEVICE RECALL' letter. The letter was sent via FedEx and/or E-mail. The letter describes the product, problem, and the actions to be taken by the customers and the firm. The customers are instructed to immediately examine their inventories for the presence of the recalled product, and complete and return the attached, MEDICAL DEVICE RECALL ACKNOWLEDGMENT AND RETURN FORM via E-mail or FAX #1-609-0750-7999. Any questions concerning the recall document are to be addressed to the customers' regional managers or to the E-mail address and/or telephone number provided on the recall notification. | Contact the recalling firm for information | 105 | 105 units | Nationwide Distribution, USA - including the states of OH, FL, WI, CT, OK, TN, OR, ID, CA, PA, MD, VA, FL, MS, NJ, LA, IA, GA, AL, MA, NY, MN, TN, IL, AR, NE, AZ, MI, WA, OR, NC, KY, MO, UT, NV, IN, SC, CO and the countries of the United Kingdom, France, Switzerland, and Austria. | Terminated | February 19, 2013 | 354 | 105 | Not_Computer
Z-1146-2012 | 59959 | material, impression | ELW | Impression material. | Dental | Dental | 510(k) | Correct VPS Vinyl Polysilocane Impression Material, Product Number Q01B. The intended use of this device is as a dental impression material. | Correct VPS | 3 | March 02, 2012 | 2012 | Sybron Dental Specialties | Pentron Clinical is voluntarily recalling the following lot of mislabeled Correct VPS Medium Body Tubes. It was discovered that the individual tubes inside the Correct VPS Medium Body Tube boxes are missing the lot/expiration date sticker. | PRODUCTION CONTROLS: Packaging Process Control | Pentron Clinical sent "Urgent Medical Device Recall" letters dated November 17, 2010 to consignees explaining the above mentioned public reason for recalls and asking them to determine if they have any of the affected products in their in stock. Product that is returned will be replaced at no charge. They also asked that they complete and fax back an enclosed recall return form.Questions can be directed to Customer Service at (800) 551-0283. | Contact the recalling firm for information | 176 | 176 units | Worldwide Distribution, USA including the states of FL, NY, WA, CA, IL, MA, NC, NJ, AZ, FL, KY, CT, PA, IN, NJ, PA, TN, NV, TX and the countries of Canada and Pakistan. | Terminated | March 14, 2012 | 12 | 176 | Not_Computer
Z-1142-2012 | 61265 | agent, tooth bonding, resin | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Brand name: Bond-1 Primer/Adhesive, a resin tooth bondingagent.Item # Item Description N01I Bond-1 Prim/Adhesive Kit N01IAA Bond-1 Primer/Adhesive 4mlN01IAB Bond-1 Primer/Adhesive 6mlN04 Lute-it w/ Bond-1 kitN04B Lute-t w/Bond-1 Intro KitN31B Bond-1 C&B; Prim/AdhesiveN33 Cement-it Universal C&B; Resin.The intended use of this device is adhesion to dentin of various polymeric filling materials (composites) used with other conditioners or combination of conditioners for bonding of composite to metal including amalgam, gold, semi-precious and non-precious alloys, porcelain and glass and luting of same to Dentin and Enamel. | Bond1 Primer/Adhesive | 3 | March 02, 2012 | 2012 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that various lots of Bond-1 Primer/Adhesive have been found to have gelled and therefore becomes unusable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pentron Clinical sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 7, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to complete and return an enclosed Recall Return Form via fax at 1-877-677-8844. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 30276 | 30,276 units total | Worldwide Distribution-USA (nationwide) and the countries of Germany, Canada, Ecuador, Egypt, Hong Kong, India, Israel, Japan, Kenya, Lebanon, Morocco, Saudi Arabia, Singapore, El Salvador, Thailand, Turkey, Taiwan, UAE, Vietnam, Armenia, Austria, Belgium, Bulgaria, Byelorussia, Cyprus, Germany, Spain, Russia, Finland, England, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Poland, Romania, Sweden, and Serbia. | Terminated | March 02, 2012 | 0 | 30276 | Not_Computer
Z-1144-2012 | 61041 | ankle arthroplasty implantation system | OYK | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TibiaxysÂŽ Ankle Arthrodesis Plate Compression Guide, Model #159635, part # 159 635, Batch #: EAHZ, EAJ1, ELGV, ENQE & ENQH.Product Usage:The Compression guide is a re-usable instrument used in the implantation of the Tibiaxys product. The Tibiaxys system is intended for fixation of bone fractures or for bone reconstruction (arthrodesis, osteotomies, and fractures of ankle joint, distal tibia and fibula). | Tibiaxys Ankle Arthrodesis Plate Compression Guide | 2 | March 02, 2012 | 2012 | Integra LifeSciences Corp. | During 6/30/2010 to 8/17/2011 the firm received five (5) incident complaints of breakage of the threaded part of the compression guide which potentially occurred when: 1) a compression is applied between talus and tibia. (Step 3 of the product description) or when 2) the surgeon screws the compression guide into the plate which most likely was attributed to improper surgical techniques used durin | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra sent a "Field Safety Notice" letter dated November 17, 2011 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review the Surgical Techinique Manual and follow the recommended guiidelines to avoid breakage. ensure that this information is distributed to their surgical teams, and fill out the enclosed "Field Safety Notice Acknowledgement and Return Form" and return to bonnie.kuzbyt@integralife.com or fax to 1-609-750-7999. For questions call 609-936-6932 or email at sean.luland@integralife.com. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 168 | 168 units | Worldwide Distribution - USA (nationwide) including the states of; AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KY, MA, MD, MI, MO, MN, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA & WI. Product was also shipped to the following countries: Austria, Belgium, Canada, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Saudi Arabia, South Africa, Spain, Sweden, Switzerland & United Kingdom. | Terminated | February 19, 2013 | 354 | 168 | Not_Computer
Z-1145-2012 | 59740 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 42-422561 SIGNATURE QTY, 1 TKA GDE/MDL (Prosthesis, knee) SET 04-05 DGHUMPHR-NOV52LKA-J_K-00432022 SLIDEX DISTAL-MP/PREMIER TIBIAL CUSTOM-MADE DEVICESignature Personalized Patient Care System is intended to be used as a surgical instrument to assist in the positioning of total and partial knee replacement components intra-operatively and in the guiding and marking of bone before cutting and to guide cutting, provided that anatomic landmarks necessary for alignment and positioning of the implant are identifiable on patient imaging scans. | Biomet Signature " TKA GDE/MDL SET 0405 | 2 | March 02, 2012 | 2012 | Biomet, Inc. | Biomet has initiated this action following investigation that identified two (2) specific lots of part number 42-422561 were incorrectly labeled. One (1) unit (lot 025640) was received back from distributor's office as a product complaint. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, BIOMET Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated January 23, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:-Immediately locate and remove the above identified Signature " Guide and Model Set.-Carefully follow the instructions on the enclosed FAX Back Response Form .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.- Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action, via the enclosed recall letter. This letter MUST be given to hospital personnel responsible for receiving recall notices. Confirm receipt of this notice by calling, 800-348-9500, extension 3983. Questions related to this notice should be directed to 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 2 | 2 units | Nationwide distribution:Virginia and Ohio only. | Terminated | September 06, 2012 | 188 | 2 | Not_Computer
Z-1148-2012 | 60657 | retainer, screw expansion, orthodontic | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AdvanSync is an orthodontic device intended for the orthodontic treatment of Class II malocclusions. The AdvanSync part numbers and lot numbers affected by this recall are as follows:Part Number: 600-4000 (Kits-include 10 Right 600-4000 and 10 Left Assemblies). Lot Numbers: 070990588, 070976869, 060971515, 060971514, 04094278DPart Number: 600-4040 (5-Pack, Left Assembly). Lot Numbers: 080901333, 080902356, 070981416, 060971299, 060976084Part Number: 600-4041 (5-Pack, Right Assembly). Lot Numbers: 080902357, 080901335, 070981417, 060971304, 060976079 | AdvanSync, an orthodontic expansion screw | 2 | March 05, 2012 | 2012 | Ormco Corporation | Manufacturing Error: There is a potential for a braze joint failure in the upper telescopic rod assembly. Although an incident could occur in which the eyelet loses connection from the rod, there is no risk of swallowing or aspirating the appliance because it will still be attached to the patient's crown. | PRODUCTION CONTROLS: Process Control | Ormco sent a Urgent Medical Device Recall letter dated October 9, 2009,( via USPS 1st class mail) to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are instructed to return the affected lot numbers of the 5-Pack Left and Right Telescopic Rod Assemblies (600-4040 and 600-4041) and the Left and Right Telescopic Rod Assemblies contained within the AdvanSync Kit (600-4000). The affected parts are identified by the items contained inside the wells titled, "Left Rods" and "Right Rods." The remaining items within the kit are not affected by this recall notification; therefore, there is no need to return the entire kit.Customers were instructed that any product which they have in stock will be replaced at no charge or credited to their account. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 (ext. 1) directly to handle the arrangements of a quick return and replacement. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Ave., Glendora, CA 91740. Customers were instructed to label their returned product "RECALLED PRODUCT Attention: Customer Returns". In addition, customers were asked to complete the enclosed Return Form and return it by fax at (909) 962-5605. |  | 3635 | 3,635 Units (237 kits, 315-pack left assemblies and 305-pack right assemblies) | Worldwide- USA (nationwide) and the countries of Europe, Australia, Canada and New Zealand. | Terminated | March 05, 2012 | 0 | 3635 | Not_Computer
Z-1160-2012 | 60285 | accessories, photographic, for endoscope (exclude light sources) | FEM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Mobile Workstations (Endoscopy Carts), Models WM-DP1, WM-NP1 and WM-WP1The mobile workstations are intended for use in medical facilities under the direction of a trained physician, and has been designed to be used with a range of Olympus equipment to facilitate GI endoscopy, endoscopic ultrasound, respiratory and surgical endoscopic | Olympus Mobile Workstations | 2 | March 05, 2012 | 2012 | Olympus America Inc. | Replacement wheels may break off from the base of the cart | PRODUCTION CONTROLS: Equipment Maintenance | OLYMPUS sent an Urgent - Device Correction letter dated October 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. OLYMPUS required customers indicate on the attached questionnaire the contact information for their facility so that they can schedule a convenient time for an Olympus Field Service Representative to visit. Service personnel would inspect, and replace, as necessary any user installed casters on the WM-P1 cart. For questions regarding this recall call (483) 896-5688. |  | 5 | 5 | Nationwide Distribution including AL, CA, NY and WV. | Terminated | January 14, 2014 | 680 | 5 | Not_Computer
Z-1176-2012 | 61055 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque CD7 (MRQ-12);Mouse Monoclonal Antibody,Antibody is bottle in HDPE bottle with screw cap lid with container label. Then put in cardboard Box with outer box label.Cell Marque, Rocklin, CA 95677,intended for qualified laboratories to qualitatively identify by light microscopy, the presence of associated antigens in section of formalin fixed, paraffin embedded tissue section using manual or automated IHC test methods. Use of this antibody is indicated as an aid in diagnostic determination of human tumors within the context of the patient's clinical history and other diagnostic tests evaluated by a qualified pathologist or physician. | Cell Marque CD7 (MRQ12); | 2 | March 06, 2012 | 2012 | Cell Marque Corporation | Letter received from supplier that material received in October 2009 was labeled as CD7, but was actually CD 2. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Cell Marque sent a Field Notification letter dated January 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the attached form indicating they have received the notification. Customers were also asked to destroy any existing inventory of the affected product. If the product was shipped to their customers, customers should confirm that they have contacted those customers. Customers were asked to complete, sign and date and return the attached Acknowledgement Form by faxing it to 1-916-746-8989 or by scanning it and emailing it to bdevos@cellmarque.com.For questions regarding this recall call 1-800-665-7284 or 1-916-746-8900. . |  | 110 | 110 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Greece, Turkey, Italy, The Netherlands, Spain, Portugal, France, Israel, S. Korea, Pakistan, Morocco, Russia and Brazil. | Terminated | October 19, 2012 | 227 | 110 | Not_Computer
Z-1174-2012 | 60937 | calibrators, drug mixture | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Synchron Phenobarbital (PHE) Reagent, P/N 469785, Classification Product Code: KLT.Subsequent Product Codes: DIP, KLS, LCD, LCQ, LCR, LFN.PHE reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems Drug Calibrator 1, is intended for quantitative determination of Phenobarbital concentration in human serum or plasma. | Synchron Phenobarbital (PHE) Reagent | 2 | March 06, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed reports of variability in results with the lots of Phenobarbital (PHE) reagent and calibrator. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 26, 2009 to all customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Customers were asked to complete and return a response form. Contact Beckman Coulter Customer Support Center (Hotline) at (800) 854-3633 for questions regarding this recall. |  | 31559 | 31,559 units total (30,617 units in the US) | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | December 05, 2013 | 639 | 31559 | Not_Computer
Z-1165-2012 | 61088 | fastener, fixation, biodegradable, soft tissue | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Depuy Mitek Lupine Loop Plus Anchor w/OrthocordCatalog Number: 210708The Lupine Anchor System is indicated for use in soft tissue bone fixation. | Depuy Mitek Lupine Loop Plus Anchor w/Orthocord | 2 | March 06, 2012 | 2012 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Lupine BR Anchor w/ Orthocord may be assembled with an incorrect suture configuration and result in a loose repair | PRODUCTION CONTROLS: Process Control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated January 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check all inventories to determine if they have any affected product. Customers were asked to refer to the instructions to report their inventory status, return and obtain replacement or credit for the product. For questions regarding this recall call 1-877-379-4871. |  | 32408 | 32408 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, AE, Argentina, Australia, Austria,Belgium,Brazil, Bahrain, Chile, China, Columbia,Cyprus, Check Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Finland, Great Britain, Guadeloupe, Hong Kong, Hungary,Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Luxembourg,Latvia, Martinique, Montenegro, Mexico, Maylasia, Netherlands,Norway, New Zealand, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Tunisia, Turkey,Uruguay, Venezuela, and South Africa | Open |  |  | 190853 | Not_Computer
Z-1172-2012 | 61224 | file, pulp canal, endodontic | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | K3 Engine File G-Pack,Product: Endodontic Pulp Canal File, Part Number: 830-9925.For use in root canal preparation. | K3 Engine File GPack | 3 | March 06, 2012 | 2012 | Ormco Corporation | A recall was initiated because SybronEndo has confirmed that the K3 Engine File G-Pack has label packaging and file mis-marking error. | PRODUCTION CONTROLS: Process Control | A recall communication was initiated on 07/31/2008 with SybronEndo forwarding an Urgent Medical Device Recall letter (via USPS 1st class mail) to all their customers who purchased the K3 Engine File G-Pack (Part No. 830-9925). The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact SybronEndo Customer Care at (800) 346-3636 directly to handle the arrangements of a quick return and replacement. Customers were instructed to complete the Return Form and return any affected product in their inventory. | Contact the recalling firm for information | 199 | 199 units | Worldwide Distribution -- US, including the territory of St. Thomas, and countries of Australia, Canada, New Zealand, China, Singapore, India, and Chile. | Terminated | March 08, 2012 | 2 | 199 | Not_Computer
Z-1177-2012 | 61262 | unit, operative dental | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Pelton & Crane Spirit 3000 dental operative unit with the TLC SAFEdrive option, 1727 Fuehauf Dr. Charlotte, NC 28273Product Usage:The intended use of the dental operative unit is for supplying power and to serve as a base for other dental devices, such as a dental hand piece, air/water syringe, oral cavity evacuator, a suction operative unit, and other dental devices. The TLC SAFEdrive option is an optional feature that monitors the operation of the high speed hand piece and gives an audible and visual signal to let the user know if the hand piece is potentially damaged or heating up. | Pelton & Crane | 2 | March 06, 2012 | 2012 | Pelton & Crane Company | Modification of the Pelton & Crane Spirit 3000 unit disabled the TLC SAFEdrive feature rendering it ineffective in detecting whether the handpiece is potentially damaged or heating up. | TRAINING: Employee Error | Pelton & Crane sent an Urgent Medical Device Recall letter dated February 20, 2012 to all affected customers. The letter identifies the affected product, problem and actions to be taken. The letter states that a Pelton & Crane representative will be contacting the customers to schedule a service appointment with the end user of the device so that the affected TLC SAFEdrive circuit board can be replaced and returned to Pelton & Crane. Customers were instructed to complete the attached Recall Acknowledgement/Return Form and fax it to Pelton & Crane, attention Frank Ray, at (704) 587-7250. The letter also requested customers to mail the original Recall Acknowledgement/Return Form in the enclosed self addressed stamped envelope. | Contact the recalling firm for information | 2 | 2 units | (USA) Nationwide Distribution | Terminated | March 28, 2012 | 22 | 2 | Not_Computer
Z-1179-2012 | 61114 | hearing aid, bone conduction | LXB | Hearing Aid. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | SoundBite Hearing System; Microphone Tube Size 0. For use in patients with moderately severe, severe or profound sensorineural hearing loss in one ear and normal hearing in the other or patients with conductive hearing loss where the pure tone average bond-conduction hearing threshold is better than or equal to 25 dB HL. Sonitus Medical, San Mateo, CA. | SoundBite Microphone Tube | 3 | March 07, 2012 | 2012 | Sonitus Medical Inc. | Sonitas received reports of stress cracking on the SoundBite Microphone Tube. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Letters were sent to all direct SoundBite customers on February 9, 2012 via certified mail. Customers were asked to remove the old microphone tubes and replace with the new enclosed microphone tubes. Return mail instructions were provided for the old microphone tubes. Customer questions were directed to 650-251-4185. | Contact the recalling firm for information | 676 | 676 microphones, all sizes. | Class III Recall - Nationwide Distribution, including the states of CA, AZ, PA, MI, UT, NC, NY, WI, TX, FL and DC. | Terminated | March 23, 2012 | 16 | 676 | Not_Computer
Z-1194-2012 | 61182 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON Rheumatoid Factor (RF) Reagent, Part Number: 475136A rheumatoid factor immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the rheumatoid factor (antibodies to immunoglobulins) in serum, other body fluids, and tissues. | SYNCHRON Rheumatoid Factor (RF) Reagent | 2 | March 08, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the all in date lots of Synchron Rheumatoid Factor (RF) Reagent might generate false positive results or calibration failures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice/Product Correction letter dated June 20, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that Beckman Coulter is developing a new RF assay. In the meantime due to current issues related to shelf life and low end sensitivity, Beckman has decided to take the product off the market effective immediately.Beckman will keep the customers updated on their progress and advise when the new assay becomes available.Customers were instructed to complete and return the enclosed Response Form with in 10 days. Customers with questions regarding this notice can contact their local Beckman Sales representative or contact Customer Support Center at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the US and Canada. |  | 22081 | 22,081 units total (12,462 units in US) | Worldwide Distribution - USA (nationwide) and worldwide to foreign countries:Australia, Belgium, China, Costa Rica, Czech Republic, Finland, France,French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy,Japan, Kuwait, Lebanon, Lithuania, Macao, Mexico, Netherlands, NewZealand, Poland, Russian Federation, Serbia, Slovakia, South Africa, Spain,Switzerland, Taiwan, Turkey, United Arab Emirates, and the United Kingdom. | Open |  |  | 22081 | Not_Computer
Z-1193-2012 | 61092 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | PointCare NOW CD4NOW Gold Reagent Pack, or Gold Pack,Part Number :30034In Vitro Diagnostic : Assay reagent intended for use on the PointCare NOW In Vitro Diagnostics Instrument to provide CD4 positive lymphocytes cell count | CD4NOW Gold Reagent Pack | 2 | March 08, 2012 | 2012 | Pointcare Technologies, Inc. | Stability not assured at labeled temperature 2-42ÂżC | PRODUCTION CONTROLS: Storage | The firm, PointCare, issued a "Customer Service Advisory (CSA-00016)" dated November 2, 2011 to its customers to inform them of the new storage requirements of 2-30ÂŽC as opposed to the previously claimed 2-42ÂŽC. The letter describes the product, problem and action to be taken. The customers were instructed to correct the labeling on any Gold Reagent Packs and Gold Reagent and ensure that the Gold Reagent Packs and Gold Reagent they currently have in stock is stored within the 2-30ÂŽC storage range. Future lots will be labeled with the updated storage temperatures.A member of our customer support team will provide additional information relative to this advisory.Should you have any questions or require further assistance, you may email us at customersupport@pointcare.net or call our offices directly at 1-508-281-6925. | Contact the recalling firm for information | 463 | 463 packs | International only: including countries of : Abuja, Haiti, Indonesia, Natal, Nigeria, Senegal, South Africa, Uganda, and United Kingdom. | Terminated | April 26, 2012 | 49 | 463 | Not_Computer
Z-1192-2012 | 61279 | endoscope, neurological | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Visionsense Stereoscopic Vision System, Brand name VSII, Model/Part Numbers 172-0030 and 172-0040 --- Classification Name: Neurological Endoscope --- NOTE: Model Numbers 172-0032, 172-0044, 172-0070, 172-0072, 172-0080 were also included in the firm's recall action but these never entered US commerce. These were distributed directly from Israel to the foreign consignees. --- The VSII is intended for viewing internal surgical sites during general surgical procedures and for use in visualization of ventricles and structures within the brain during neurological surgical procedures. These devices do not carry an expiration date. However, they are automatically taken out of service and replaced after 40 uses. | VSII | 2 | March 08, 2012 | 2012 | Visionsense, Ltd. | A complaint was received and a unit returned that exhibited the metal and glass shaft window becoming dislodged during scope cleaning. Additionally, during in-process visual inspection of the metal and glass shaft window sub-assembly from a specific batch that was received from the supplier it was noted that the window at the tip of the scope was not seated properly. It was determined that the wi | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Visionsense Ltd. sent URGENT DEVICE RECALL and URGENT DEVICE RECALL/FIELD SAFETY NOTICE letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer and included a return "fax" form in the recall communication package. The letters were sent via UPS express service. All direct consignees were contacted, and follow up phone calls conducted if needed. The recall required the return of all affected units for inspection and/or repair. Visionsense provided a return authorization number and packing instructions to the consignees. |  | 21 | Domestic - 17 units; Foreign - 4 units | Worldwide Distribution - USA including Arizona, Georgia, Pennsylvania, California, and the country of Turkey | Terminated | March 08, 2012 | 0 | 21 | Not_Computer
Z-1190-2012 | 61173 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Right Long Tactile Probe, Part #962011S, Medtronic Navigation, Louisville, CO 80027. Trackable handheld instruments used in spine surgeries for preparation of the pedicle. These instruments are used in conjunction with the Medtronic Navigation StealthStation platform and various software applications. | Right Long Tactile Probe | 2 | March 08, 2012 | 2012 | Medtronic Navigation, Inc. | When subjected to bending, twisting, and malleting, the probe tips break off when used in the sacrum for which they are not indicated for use. These probes are approved for use in the lumbar vertebra. | DESIGN: Device Design | An Urgent - Product Safety Notice customer letter was sent via FED EX on 2/6/12 and 2/6/12 to direct consignees in the US identifying the affected product and the issue with it. The letter states that Medtronic Navigation is removing the product from clinical use and will provide a sales credit to customers. In the letters to Medtronic field personnel, it states that local sales and service personnel will be identified to remove the affected product from customer sites and return them for scrapping. Customers are to contact a sales representative if they have any questions regarding the letter. | Contact the recalling firm for information | 5 | 5 units | Worldwide Distribution -- US and Republic of South Korea. | Terminated | March 26, 2012 | 18 | 11 | Not_Computer
Z-1195-2012 | 61145 | electrophoretic separation, alkaline phosphatase isoenzymes | CIN | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCESS OstasePart Number: 37300Used for the quantitative measurement of bone alkaline phosphatase (BAP). | ACCESS Ostase | 2 | March 09, 2012 | 2012 | Beckman Coulter Inc. | The Access Ostase reagent lot 021281 shows instability resulting in falsely high patient results. | DESIGN: Process Design | A recall communication was initiated on 6/15/2011 with Beckman forwarding an "Urgent Product Correction" (PCA) letter to all their customers who purchased the Access Immunoassay Systems Ostase Reagent Kits. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the Access Ostase reagent lot identified on the PCA letter and discard all remaining inventory. As stated in the Instructions for Use for Access Ostase, results should be interpreted in light of the total clinical presentation of the patient.All results generated with reagent lot number 021281 where patient results are higher than expected should be evaluated. The affected lot has been removed from Beckman's inventory, and alternate lots will be provided for all current and future replenishment orders or requests.Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions can contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call(800) 854-3633 in the US and Canada. | Contact the recalling firm for information | 318 | 318 units total (44 units in US) | Worldwide Distribution -- US, including states of FL, IN, VA, CA, and KY and countries of China, Japan, and Panama. | Terminated | March 09, 2012 | 0 | 318 | Not_Computer
Z-1198-2012 | 60892 | dispenser, cement | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Flat Proximal Pressurizer.Bone cement pressurizers are intended to help seal the delivery path between the delivery nozzle and bone in order to pressurize the cement during delivery. | Flat Proximal Pressurizer. | 2 | March 12, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The firm identified two complaints have been reported indicating that the product inside the package did not match with the label information. The product was labeled as Flat Proximal Pressurizer P/N 0206-556-000 while the product inside the package was Femoral Pressurizer P/N 0206-566-000. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Stryker Instruments sent an "Urgent Medical Device Recall notification" dated January 13, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to : Immediately check all stock areas or operating room storage and quarantine any affected product found. Mark product as RECALLED PRODUCT . Please indicate on the BRF the quantity of affected Pressurizers you are returning and fill in the form completely. Sign the BRF (even if you do not have any affected product).Note: Your signature on the BRF indicates that you received and understand this notification and are returning all affected product from your facility.Fax the Business Reply Form to Stryker Instruments Regulatory Department,866-521-2762.Upon receipt of the BRF, Stryker will contact you to coordinate the return of allaffected product that you have on hand. Send back all affected product using the pre-paid shipper provided to you by Stryker.Upon receipt of the product, a replacement will be issued to your account. |  | 10 | 10 boxes (60 units) | Worldwide Distribution -- Nationwide, including the states of AL and OK and the country of Canada. | Terminated | July 24, 2013 | 499 | 10 | Not_Computer
Z-1201-2012 | 61160 | wheelchair, powered | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, Model #'s: TDXSP-CG, TDXSP-CG-GT, TDXSP-MCG, TDXSP-MCG-GT, Distributed By:Invacare Corporation, Elyria, OH 44035The Invacare TDX power wheelchair is a battery powered, motor driven device with the intended function and use of providing mobility to those persons limited to a sitting position that have the capability of operating a power wheelchair. It is a rigid or non-folding type power wheelchair, with mid wheel drive capability. | TDX SP Power Wheelchair | 2 | March 12, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because use of this actuator could result in reduced cross sectional current carrying area due to wear/chafing of the wires. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Invacare, sent an "URGENT: RECALL INFORMATION" letter dated February 10, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. Invacare has developed retrofit kits intended to address this issue. The consignees were instructed to contact the customers that have potentially affected product and perform the necessary updates. The consignees were also instructed to complete and return the Provider Response Form to order the retrofit kit.If you have any questions, please contact Customer Service at 1-800-333-6900. |  | 102 | 102 wheelechairs | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CT, FL, GA, IA, IL, IN, KY, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, RI, SD, TX, VA, WA, and WI; and countries of: Australia, Canada and New Zeland. | Terminated | March 01, 2013 | 354 | 102 | Not_Computer
Z-1202-2012 | 61156 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | da Vinci S. Surgical System, da Vinci Si and da Vinci Si-e Surgical Systems.Intuitive Surgical, Inc.Sunnyvale, CA 94086.The Intuitive Surgical Endoscopic instrument Control System is intended to assist in the accurate control of Intuitive Surgical Endoscopic Instruments. | da Vinci S. Surgical System, da Vinci Si and da Vinci Sie Surgical Systems | 2 | March 13, 2012 | 2012 | Intuitive Surgical, Inc. | The holding brake, a component of the da Vinci S Surgical System, may allow passive uncontrolled motion due to gravity during specific power-off conditions. | DESIGN: Device Design | Intuitive Surgical sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 3, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provided new instructions for use. Additionally, a Device Correction Response Form was enclosed for customers to complete and return via fax to 408-523-1390. Contact Intuitive Surgical Customer Service at 1-800-876-1310 for questions regarding this notice. |  | 1966 | 1966 systems are in commerce at hospital locations. | Worldwide Distribution. | Terminated | May 09, 2012 | 57 | 1966 | Not_Computer
Z-1204-2012 | 61296 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The Reprocessed Mitek Arthroscopic Wand is considered a soft tissue ablators. Soft tissue ablators are radiofrequency (RF) surgical tools designed for removal and dissection of tissue in arthroscopic surgeries. A generator or controller serves as the power unit. Soft tissue ablators are available in a wide range of sizes and angles; some models include thermal or suction functions.The outer box of the Reprocessed Mitek Arthroscopic Wand is labeled in part: "***ASCENT HEALTHCARE SOLUTIONS Lakeland, FL***Contents sterile by EO unless package is opened or damaged.***Reprocessed Device for single use. Reprocessed By Ascent Healthcare Solutions.***Cat #: 225301 MITEK 3.5mm SIDE EFFECT ELECTRODE***Lot #: 674290 Expiration Date: 04/2012***"Individual pouch labeled in part: "***Cat #: 225301***MITEK 3.5mm SIDE EFFECT ELECTRODE****IFU at www.ascenths.com***Contents: One Expiration Date 04/2012***Lot # 674201 ID# 0***Reprocessed by Ascent Healthcare Solutions Not affiliated with the original manufacturer.***ASCENT HEALTHCARE SOLUTIONS Lakeland, FL***" | Reprocessed Mitek Arthroscopic Wand | 2 | March 13, 2012 | 2012 | Stryker Sustainability Solutions, Inc. | On 06/09/2009, Stryker Sustain ability Solutions, formerly Ascent Healthcare Solutions, recalled lots 674201 & 674290 of their Reprocessed Mitek Arthroscopic Wand due to the lots not being tested in accordance with the functional test requirement during the time frame of 04/14/09 to 05/25/09. | PRODUCTION CONTROLS: Process Control | Ascent Healthcare Solutions notified their customers by presenting a letter dated June 1, 2009 from an Ascent Representative. The representative was to recover any products that may have been affected. Questions were directed to 1-888-888-3433. | Contact the recalling firm for information | 48 | 48 | Nationwide Distribution (USA), including the states of: TX, NY, MN, CA, and CT. | Terminated | May 07, 2012 | 55 | 48 | Not_Computer
Z-1203-2012 | 61275 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Coupler System IIMizuho Orthopedic Systems, Inc.Union City, CA 94587The 5873 Coupler II is designed for use specifically with the Mizuho OSI Modular Base and the Spinal System Table Tops (AKA Jackson Spinal Table Top). The Coupler II mounts onto the Spinal Surgery Top by using the Spinal Top Adaptor Brackets. The 5873 Coupler II is used to allow positioning of a patient on the table utilizing either a skull clamp or horseshoe headrest for cranial stabilization or support during surgical procedures. The Coupler II can be accommodate either a radiolucent or aluminum skull clamp manufactured by Integra Life Sciences, referred to as a MayfieldÂŽ clamp or the radiolucent or aluminum skull clamp referred to as the DOROÂŽ which is manufactured by pro med instruments Inc. | Coupler System II | 2 | March 13, 2012 | 2012 | Mizuho Orthopedic Systems Inc | Mizuho determined that current labeling of The Coupler II system required revisions after investigation of a complaint. | DESIGN: Labeling Design | Mizuhosi OSI sent an Urgent Medical Device Recall Communication letter dated February 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Mizuho OSI provided their customers of the affected product an updated Owner's manual. Customers were instructed to ensure they remove their existing Owner's Manual from use and repalce with the current Owner's Manual. If customers were unsure of the device's use, they should refer to the included Owner's Manual and ensure all users are properly trained, able to assess the device's status prior to and after each use. If customers were unsure of the status of their device, or of its use, they should call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.Customers were asked to respond that the recall has been received, read, understood and communicated to those affected by this activity. A response should be sent via e-mail to couplerII@mizuhois.com.For any questions regarding this recall call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.was sent out on February 27, 2012 to all affected sites. This was after the original was sent to CDRH for review on January 31, 2012, but approval of revisions was not received until 2/27/12. |  | 1077 | 1077 devices | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, France, Germany, Hong Kong, Israel, Japan, South Korea, Lebanon, Malaysia, Netherlands, New Zealand, South Arabia, Singapore, Germany, South Africa, Spain, Sweden, UA. E., United Kingdom, and Australia. | Terminated | September 25, 2012 | 196 | 1077 | Not_Computer
Z-1207-2012 | 61103 | skin prep tray (kit) | OJU | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Wet PVP-I Small Procedure Preoperative Skin Prep Tray containing 1 - 59 mL bottle Povidone Iodine Cleansing Solution (Scrub) USP 7.5% and 1 - 59 mL bottle Povidone Iodine Topical Solution (Paint) USP 10% with Powder-Free Esteem Gloves; 20 trays per case; a sterile, latex-free, single use only tray, solution not exposed to sterilant; U.S. components packaged in Mexico except as indicated: Gloves and Cotton-Tipped Applicators made in China, Sponge Stick made in Mexico; CareFusion, McGaw Park, IL 60085 USA; tray catalog number 4420The Povidone Iodine patient prep trays are used where chlorhexidine gluconate (CHG) products aren't appropriate or indicated for patient preoperative prep (traditional two-step scrub and paint using cleansing solution and topical solution) on intact skin and mucous membranes, and for washing lacerations and abrasions. | CareFusion Wet PVPI Small Procedure Preoperative Skin Prep Tray | 2 | March 14, 2012 | 2012 | Carefusion 2200 Inc | The povidone-iodine (PVP-I) solution bottles in the trays may have loose caps, allowing leakage and also elevated levels of ethylene oxide (EO) sterilant and/or by-products within the PVP-I solution, which may represent a patient safety hazard. |  | CareFusion sent e-mails and follow-up Urgent Recall Notices dated February 7, 2012 to all of the affected customers on the same date, requesting that the customers follow the steps outlined below as mitigation until product can be replaced.1. Please immediately destroy/discard or return any affected product you may have in-stock at your facility. Do NOT further distribute these products.2. Forward this RECALL NOTICE to any customers to whom these products were sold notifying them of the potential risk (reference attachment for affected production lot numbers). End users must be advised to discard or return any CareFusion Skin Prep Tray with the noted product/lot number; distributors should manage this return process per their internal procedures. End users MUST NOT use these products. 3. Complete and return the attached Acknowledgement Form indicating that they have received this RECALL NOTICE and that they will communicate the Recall Notice to all customers who have purchased CareFusion Preoperative Skin Prep Tray catalog numbers 4420, 4468, and 4469A. 4. Contact CareFusion Customer Service at (800) 523-0502 ext. 53576 for details on return and/or replacement of stock. |  | 38440 | 38,440 trays | Nationwide Distribution including Arizona, California, Colorado, Connecticut, Florida, Illinois, Missouri, Nebraska, Nevada, New York, Ohio, Puerto Rico, Rhode Island, South Carolina, South Dakota, Texas, Utah and Virginia. | Open |  |  | 1644150 | Not_Computer
Z-1205-2012 | 61270 | cement, dental | EMA | Dental cement. | Dental | Dental | 510(k) | Brand Name: TempSpan CMT Temporary Cement, a temporary cement.Part Number: N69CA.The intended use of this device is to affix a temporary dental devices such as crowns or bridges to a tooth. | TempSpan CMT Temporary Cement | 2 | March 14, 2012 | 2012 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the TempSpan Temporary Cement may have black plastic material on the surface of the paste, and it may set faster than specified. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Pentron Clinical sent an Urgent Medical Device Recall letter dated April 20, 2011, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax back the enclosed recall form to 1-877-677-8844 in order to confirm receipt of the customer notification. Customers were instructed to contact Pentron Clinical Customer Service at (800) 551-0283, option 1, directly to handle the arrangements of a quick return and replacement.For questions regarding this recall call 714-516-7400. |  | 627 | 627 units total (601 units in the US) | Worldwide Distribution - USA (nationwide) and the countries of Canada, Philippines, Italy, Netherlands, and the UK. | Terminated | March 14, 2012 | 0 | 627 | Not_Computer
Z-1213-2012 | 61090 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System.Intended to be used for palliative treatment of biliary strictures produced by malignant neoplasms. | Stent Delivery Catheter | 2 | March 15, 2012 | 2012 | Idev Technologies, Inc. | The firm is taking this action because the device was one of eight that may have been mislabeled with the incorrect size (length). | PRODUCTION CONTROLS: Labeling Mix-Ups | IDEV Technologies sent an "URGENT PRODUCT RECALL" letter dated February 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to sign the letter and return a copy to the firm via fax to 281-525-2001 to indicate that they have received the notification. Contact the firm at 281-525-2052 for questions regarding the notice. |  | 7 | Seven devices | Nationwide Distribution-USA (nationwide) including the states of AL, IL, NJ, NV, and PA. | Terminated | September 09, 2013 | 543 | 7 | Not_Computer
Z-1215-2012 | 61093 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Integra CUSA EXcel Sterile Tips Rx Only Single Use OnlyProduct Usage: For usage in surgical procedures where fragmentation, emulsification and aspiration of soft and hard (e.g. bone) tissue is desirable. | Integra | 2 | March 15, 2012 | 2012 | Integra LifeSciences Corp | A small number of non-conforming CUSA Excel 23 KHz Standard Tip and flue (item Number C4601S) were released for distribution that did not meet Integra's in-process requirements for release. It was discovered during in-process inspection that the max power reading exceeded the specification of 80 watts, causing the stroke range of the tip to exceed it required range by up a value up to 0.002 inches | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Integra sent an Urgent Product Recall Notification letter dated October 14, 2011 to all affected customers via US Mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, discontinue use, quarantine and return any available affected lot numbers. Customers were instructed to distribute this communication to any customers to whom they have shipped the affected product. Customers were advise to complete the enclosed Recall Acknowledgement and Return Form and return it as soon as possible to Integra Customer Service. For questions contact Integra Customer Service at 609-936-2659. |  | 24 | 24 | Worldwide Distribution | Terminated | October 19, 2012 | 218 | 24 | Not_Computer
Z-1211-2012 | 61348 | table, examination, medical, powered | LGX | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General Hospital | 510(K) Exempt | MPI Rapid Response HUT Table, Model #RRHUT 1017. The firm name on the label is Medical Positioning Inc., Kansas City, MO.Designed for multiple uses including a conventional examination table, unexplained syncope, neurocardiogenic syncope, chronic fatigue syndrome, and most noninvasive procedures. | MPI Rapid Response HUT Table, Model RRHUT 1017 | 2 | March 15, 2012 | 2012 | Medical Positioning Inc | The warning label regarding electrical shock was not placed on the device warning to have the power switch in the "OFF" position prior to connecting the unit to electric power. | PRODUCTION CONTROLS: Error in Labeling | Medical Positioning Inc. sent a "RAPID RESPONSE HUT TABLE LABELING CORRECTION NOTIFICATION" letter dated March 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included an Owner's Manual insert page which clearly states the warning and how to avoid the electric shock hazard and a warning tag to attach to the product's power cord near the plug end using the provided wire tie. Contact the firm at 816-595-2356 for questions regarding this notice. | Contact the recalling firm for information | 5 | 5 tables | Worldwide Distribution-USA (nationwide) including the states of IN, KS, WI, NY, IL, NC, OH, and MS and the country of Thailand. | Terminated | July 11, 2013 | 483 | 9 | Not_Computer
Z-1220-2012 | 61007 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | DuoDiagnost, model 707015, Universal RF systemProduct Usage: is a mult-functional diagnostic X-Ray system intended for Fluorosopic, Radiographic, Angiographic, and Interventional applications. | DuoDiagnost | 2 | March 16, 2012 | 2012 | Philips Healthcare Inc. | The device is currently not compliant with FDA's amended design criteria for this system, the product is not commercially available in the US. However, it was inadvertently shipped to Puerto Rico due to a logistical error. |  | Philips sent an Urgent - Field Safety Notice dated December 26, 2011 to the two affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter states that customers may continue to use the DuoDiagnost system until the new systems are replaced by Philips. A Philips representative will contact customers to arrange for the replacement of current system. For further information or support contact your local Philips representative. |  | 2 | 2 units | USA Nationwide Distributed in Puerto Rico | Terminated | November 01, 2013 | 595 | 2 | Not_Computer
Z-1221-2012 | 59992 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Disposable Blades size Mac 4 (part number 50674, 50674/F, 50674/BMI) and Mac 3 (part numbers 50675, 50675/F, 50675/BMI) Product Usage: Disposable laryngoscope blades. | Crystal Metal Disposable Blades | 2 | March 18, 2012 | 2012 | Penlon, Ltd. | Disposable laryngoscope blades incorrectly labeled with wrong size. | PRODUCTION CONTROLS: Error in Labeling | InterMed Penlon sent and Urgent Field Safety Notice letter dated September 21, 2011 to affected customers. The letter identified the affected product, problem and advice on actions to be taken by user. Customers were instructed to complete the Reply Form and send to the address provided. | Contact the recalling firm for information | 24940 | 24,940 cartons | Worldwide Distribution - USA (nationwide) in the countries of Belgium, France, Greece, Ireland, Italy, Lithuania, Netherlands, Norway, Portugal, Switzerland, Canada, China, Hong Kong and Malaysia. | Terminated | June 19, 2012 | 93 | 24940 | Not_Computer
Z-0910-04 | 35768 | [6 empty] | Medtronic MiniMed Paradigm Quick-set Plus Infusion Sets. Models MMT-359S6; MMT-359S9; MMT359L6; & MMT-359L9. (differences in models is variation in length, 6 versus 9mm) | Medtronic MiniMed Paradigm Quickset Plus Infusion Sets | 1 | March 19, 2012 | 2012 | Medtronic MiniMed | Possibility of interrupted insulin flow. This posting is for a recall that occurred in March 2004, also posted on the FDA website in 2004. |  | Firm will ship all users (customers) one box at a time every three weeks for the next several months. A recall notification letter was issued to users and to healthcare professionals, and distributors. Letters were mailed from March 2-5, 2004. |  |  | Undetermined | Nationwide | Open |  |  |  | Not_Computer
Z-1226-2012 | 61183 | accessories,arthroscopic | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Endoscopic Gastrocnemius Release (EGR) System. The device consists of one (1) unit which is kitted within an instrumentation tray. The Integra EGR System includes a range of instruments for endoscopic approach to the Gastrocnemius Aponeurosis Recession. The cutting instrument features a retracting blade for cutting the gastrocnemius aponeurosis. Associated instruments include an elevator, a cannula, and a cannula obturator, which are all consumable instruments provided sterile. The EGR System is indicated for the treatment of posterior heel cord contracture (Equinus) in those patients who fail to respond to conservative management. | EGR (Endoscopic Gastrocnemius Release System | 2 | March 19, 2012 | 2012 | Integra LifeSciences Corp. | The firm received several complaints from customers reporting that the blade of the EGR system may not retract back into the device following surgery. The firm attributes this problem to user error or improper surgical technique on the part of the physician. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm sent a 'A Medical Device Correction Notice' with an attached Acknowledgement and Return Form to all consignees on 1/6/2012 via FedEx and E-mail. The Correction Notice informs customers of the firm's receipt of several complaints regarding the Endoscopic Gastrocnemius Release (EGR) System and the problem with non-retraction of the blade back into the device during surgery. Although no adverse health consequences or patient injuries have been reported to date, the notice is being issued by the firm in order to reinforce the importance of following the additional clarification to the Surgical Technique Manual. According to the firm, this action should serve to reduce the possibility in which the blade may fail to retract into the device. Questions were directed to (609) 936-6832. |  | 850 | 850 units | Nationwide Distribution, including the states of: MI, CA, TX, VT, OK, OH, NC, MN, CT, CO, TN, NY, NJ, MO, FL, WA, AR, MT, VA, IL, PA, NE, GA, AZ, WI, OR, MD, LA, IA, AL, MA, ID, KY, UT, NV, IN and SC. | Terminated | January 23, 2014 | 675 | 850 | Not_Computer
Z-1227-2012 | 61243 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Covidien Duet TRS Articulating 60- 3.5MM, Single Use Loading UnitRef: DUET6035A The ENDO GIA Universal Staplers with DUET TR S Loading Units have applications in abdominal, gynecologic and pediatric surgery for resection, transection, and creation of anastomosis. They may be used for transaction and resection of liver substance, hepatic vasculature and biliary structures and for transection and resection of pancreas | DUET TRS Loading Units | 2 | March 19, 2012 | 2012 | Covidien LP | Missing component results in the stapler not firing. | PRODUCTION CONTROLS: Process Control | The firm, Covidien, sent a "Medical Device Recall" letter dated February 16, 2012 to customers via federal express. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of Lot# N1H003LX of Duet TRS Articulating Loading Unit and segregate affected product; return inventory to Field Returns Department, 195 McDermott Road, North Haven, CT 06473 and contact customer service at SDFeedback@Covidien.com or (800) 722-8772, option #1 for RGA # prior to returning; and complete and return the Recall Product Return Form via fax to (203) 822-6009 (all affected units must be returned with completed form through the Distributor). If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance at 203-492-5232, or Customer Service at 800-722-877 option 1 customer service. | Contact the recalling firm for information | 1158 | 1158 units | Worldwide distribution: USA (nationwide) and country of: Venezuela. | Open |  |  | 1158 | Not_Computer
Z-1225-2012 | 61138 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith & Nephew BIOSURE HA Interference Screw, 7 mm x 25 mmCatalog Number: 72201772.The BIOSURE HA Interference Screw is indicated for fixation of bone-tendon-bone or soft tissue grafts during anterior/posterior cruciate ligament (ACL/PCL) reconstruction procedures. | BIOSURE HA Interference Screw | 2 | March 19, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | Mislabeled: The device was labeled as a standard thread screw (P/N 72201772) on pouch and box,however the device contained in the packaging is a reverse thread screw (P/N 72201773). | PRODUCTION CONTROLS: Packaging Process Control | Smith & Nephew Endoscopy sent an "URGENT-PRODUCT RECALL 1st NOTIFICATION" dated February 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to inspect their inventory and quarantine any affected product. For instructions on returning product, contact the Smith & Nephew Endoscopy Division Returns Group at 800-343-5717 ( 0ption 3) or email to endo.andreturns@smith-nephew.com. Additionally, customers were asked to complete the bottom portion of the letter and to fax it to 1-508-261-3620. Contact the Smith & Nephew Endoscopy Division for questions regarding this notice at 1-508-261-3731. | Contact the recalling firm for information | 194 | 194 units | Worldwide Distribution-USA (nationwide) and the states of AK, AZ, CA, CO, CT, IA , MD, MI, MN, MO, MS, NE, OH, OK, PA, SC, and VA and the countries of Australia, Canada, Germany, South Korea, Spain, Switzerland, Thailand, and United Kingdom. | Open |  |  | 194 | Not_Computer
Z-1223-2012 | 61274 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instrumentation Tibial Resection Tower Universal with alignment tower geometry, REF 32487557, LOT 2561334, Biomet UK Ltd, South Wales, UKProduct Usage: Usage: Used to guide a saw blade in resecting the proximal tibia. Intended to be used with a tibial alignment guide. | Vanguard Universal Tibial Resection Block | 2 | March 19, 2012 | 2012 | Biomet U.K., Ltd. | This recall was initiated following a report that a bag labeled as REF. 32-487557/Lot no 2561334 (Vanguard Tibial Resection head universal w/ alignment tower) actually contained a Ref. 32-487002/Lot no. 2561329 (Vanguard Distal cutting Block w/out handle attachment). Investigation found packaging of the affected product may contain the incorrect item. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | Biomet U.K. sent a Field Safety Corrective Action notice dated February 14, 2012 to all affected customers. The notice identified the affected products, problem and actions to be taken. Customers were instructed to locate, discontinue use and quarantine the recalled products pending return to Biomet. The letter recommended customers provide the notice to any other organization who may have received the affected instruments. Customers were instructed to complete and return without delay the attached "Fax-Back Response Form" acknowledging receipt of notice and indicating the quantity of products to be returned. For questions call +44 (0) 1656 761658, fax: +44 (0) 1656 645454 and e-mail uk.complaints@biomet.com. |  | 4 | 4 units | Worldwide Distribution - in the countries of France, Japan, Italy, Netherlands, UK , no devices distributed in the USA | Terminated | November 06, 2012 | 232 | 4 | Not_Computer
Z-1229-2012 | 61122 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Synchron CX Immuno-Protein Cal, part number: 442840, other product code: CKAThe Beckman Coulter SYNCHRON CX Immune Protein Calibrator Set, in conjunction with SYNCHRON Immuno Protein Reagents, is intended for use on SYNCHRON CX Systems for the calibration of IgA, IgG, IgM, and Transferrin (TRF). | SYNCHRON CX(r) systems, ImmunoProtein Calibrator Set | 2 | March 20, 2012 | 2012 | Beckman Coulter Inc. | A recall was initiated because Beckman has confirmed that the standardization of the Immuno Protein Calibrator Set has been re-established to IFCC Reference Preparation for Plasma Proteins, BCR-470 beginning with lot number M901330. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice letter dated August 24, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers with questions regarding the product correction action were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-993-8916. |  | 1923 | 1,923 units total (1,882 units in US) | Worldwide Distribution - USA (nationwide) and Canada | Terminated | April 17, 2012 | 28 | 1923 | Not_Computer
Z-1230-2012 | 61197 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 Systems, Product Code M-4800-01The intended use of the CARTO 3 System is catheter-based cardiac electrophysiological (EP) procedures. | CARTO 3 Systems | 2 | March 20, 2012 | 2012 | Biosense Webster, Inc. | Some CARTO 3 Systems may get radio frequency interference through certain interconnected cable configurations resulting in jumping or disappearing catheter icon and/or error messages during ablation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Three Urgent Customer Notification recall communications were initiated 2/9/12 with Biosense forwarding the notification to customers with affected backplane cards (electrophysiology lab managers, a notification for customers with affected back plane cards (risk managers), and a notification for customers with affected backplane cards and smarttouch module along with customer acknowledgement forms. The letters provides the customers with an explanation of the problem identified and steps to take in the event that the issue is observed. If the steps provided do not resolve the issue, customers are to contact their Biosense Webster representative. A Biosense Webster representative will contact customers to schedule an appointment for the replacement of the Backplane card.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. If customers have questions, they can contact their BWI representative. | Contact the recalling firm for information | 134 | 134 units total (32 units in US) | Worldwide Distribution -- US, Japan, China, Ireland, Germany, Italy, Slovenia, Indonesia, Poland, Spain, United Kingdom, Switzerland, Belgium, and Canada. | Terminated | September 06, 2012 | 170 | 134 | Not_Computer
Z-1271-2012 | 61029 | fluorometric method, cpk or isoenzymes | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITEÂŽ/IMMULITEÂŽ 1000 CK-MB LKMB --- Catalog number: LKMB1 (100 tests), LKMB5 (500 tests) --- Test code: CMB Color: Dark Blue.For in vitro diagnostic use with the IMMULITE and IMMULITE 1000 Analyzers --- for the quantitative measurement of creatine kinase isoenzyme MB (CK-MB) in heparinized plasma or serum, as an aid in patient management and the assessment of prognosis of myocardial infarction. | IMMULITE/IMMULITE 1000 CKMB | 2 | March 21, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed that the median for the reference range with current kit lots of the CK-MB assay differs from the guidelines published in the Instructions for Use (IFU). This difference applies to all IMMULITE platforms. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL" letter dated January 2012 to all affected customers. The letter identifies the product , problem, and actions to be taken by the customers. Siemens recommends that customers reestablish their own medians and reference ranges, in line with the recommendations made in the IFU for these products. A Field Correction Effectiveness Check form was included for customers to complete and return to the firm. Customers with questions or who need additional information should contact their Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution. | Open |  |  |  | Not_Computer
Z-1231-2012 | 61104 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 905463, 2.9 mm Micromax Anchor Long Shaft, Size 2 Blue/White Maxbraid (Lactosorb L-15 resorbable copolymer ) Sterile, EO; Biomet Sports Medicine, Warsaw, INProduct Usage: The Biomet Sports Medicine Soft Tissue Anchoring Devices are resorbable repair devices used to attach soft tissue to bone. | Micromax Anchors | 2 | March 21, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action following an internal audit which discovered the component level package insert was included in the product packaging in place of the end level product package insert (01-501072).This meant that the Indications, Contraindications, Possible Adverse Effects, and some Warnings are not present with the packaged device. | PRODUCTION CONTROLS: Packaging | Biomed Inc. sent an Urgent Medical Device Recall Notice dated January 25, 2012 to distributors, hospitals, and OR managers via US mail. The notice identified the product, description of problem, possible adverse health consequences, and actions to be taken, including directions to immediately locate devices and stop device use ,with directions for return. A copy of the correct package insert was included with directions to report any procedures performed outside of indications for use. Receipt confirmation was requested by calling 800-348-9500, extension 3983 or 3009. The notice included a list of units invoiced to the account with a FaxBack Response form to be filled out and sent to the Attention of Angie Dickson, Biomet, 56 East Bell Drive Warsaw, In 46580. Questions concerning the recall should be directed to 574-371-3983 or 574-371-3009, M-F 8-5PM(EST). |  | 787 | 787 | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Brazil, Chile, Colombia, Israel, Mexico and Netherlands. | Terminated | April 12, 2013 | 387 | 4212 | Not_Computer
Z-1239-2012 | 61214 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Ostase QC, Part Number: 37309.The Access Ostase QC is intended for monitoring system performance of the Access Ostase assay. | Access Ostase QC | 2 | March 21, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed that the Access Ostase Calibrator and QC lots do not meet expiration date claims. The Access Ostase Reagent Kit lots identified were quality control tested and released using Access Ostase Calibrator and QC lots that contain the implicated raw material. | DESIGN: Process Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 9, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard all remaining inventory. In addition, a Response Form was enclosed for customers to complete and return. Contact Beckman's Customer Support Center at (800) 854-3633 for questions regarding this notice. |  | 4961 | 4961 units total (1321 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belgium, Canada, China, Colombia, Czech Republic, Denmark, France, Georgia, Germany, Greece, Hong Kong, Hungary, India,Ireland, Italy, Japan, Jordan, Lebanon, Liechtenstein, Mexico, Morocco, Netherlands, New Zealand, Panama, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | June 26, 2012 | 97 | 4961 | Not_Computer
Z-1214-2012 | 61298 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | D-SPECT" Cardiac Scanner System -- COMMON/USUAL NAME: Emission Computed Tomography System --- SPECTRUM DYNAMICS --- Spectrum Dynamics (Israel) Ltd. 22 Bareket St. North Industrial Park, P.O. Box 3033, Caesarea, 30889 IsraelD-SPECT is an emission computed tomography system intended to detect the location and distribution of gamma ray radionuclides in the body and produce cross-sectional images through computer reconstruction of the data. The device includes display equipment, patient and equipment supports, component parts, and accessories. D-SPECT is primarily intended for cardiac applications. D-SPECT supports radionuclides within the energy range of 40 -170 keV. | DSPECT" | 2 | March 21, 2012 | 2012 | Spectrum Dynamics | The firm received a report related to the detector housing on the D-SPECT Cardiac Scanner. Specifically, it was reported that if the detector housing securing mechanism is compromised, the detector housing may pull away from its secured position on the gantry and potentially come intocontact with a patient. The reported event was discovered during set up and thus did not involve a patient or p | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spectrum Dynamics sent a "Voluntary Field Correction" letter dated July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken.To implement this fix, a BC Technical Field Service Engineer contacted users to schedule a service visit. In the meantime, Spectrum Dynamics did not believe that this issue warranted the discontinuation of use of the D-SPECT Cardiac Scanner. However, in an abundance of caution, if the detector housing appeared to be loose, or appeared to be disconnecting from the gantry, do not use the system. Please contact your local BC Technical Field Service Engineer immediately.Enclosed is a verification form. Please fax the completed form at your earliest convenience to 415.663.4225. For further information or support concerning this issue, please call 925.922.2615. |  | 44 | Domestic- 37 units, Foreign - 7 units | Worldwide Distribution -- USA (nationwide) and the countries of France and United Kingdom | Open |  |  | 44 | Not_Computer
Z-1275-2012 | 57924 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | CRYO-206V and CRYO-206F Procedure Kit, which includes: Cryo Probes, Urethral Catheter, and temperature probes.Kit is intended to be used with the CryoCare Surgical System for freezing/ablating tissue by extreme cold temperatures. | CRYO Care Surgical System | 2 | March 21, 2012 | 2012 | Healthtronics, Inc. | HealthTronics has initiated this recall of these products because there have been instances where the temperature probes (aka TempProbes) in the kit may be mislabeled. In some instances, the label near the prove does not match the label at the connector. If the labeling error went undetected, the TempProbes would not be monitoring temperature in the area believed to be under surveillance. | PRODUCTION CONTROLS: Labeling Mix-Ups | Endocare sent an Urgent: Device Recall letter dated February 27, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The letter instructed consignees to immediately cease all distribution, examine their stock and the stock of accounts that have been supplied and remove any affected devices from all inventories. If the product have been further distributed contact Endocare for instruction. If consignees have any of the affected products: 1) Contact Customer Service at 877-367-4873 to obtain a RMA number, 2) Complete the Verification Section on page two, and 3) Return this letter with the affected product to the address provided. For questions about this recall contact Endocare, Inc at 877-367-4873. | Contact the recalling firm for information | 9274 | 9274 units | Worldwide Distribution--USA (nationwide) and the countries of Barbados, Brazil, Canada, China, Italy, Jakarta, Malaysia, Madrid Spain, Taiwan, and United Kingdom. | Terminated | March 19, 2013 | 363 | 9274 | Not_Computer
Z-1269-2012 | 61294 | microscope, surgical, general & plastic surgery | FSO | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Touch Panel/Control Unit Carrier of the Swingarm M820/M844 C19, article no. 10448194MO; a component of the Leica M820 C19 Surgical Microscope, article 10448196, and the Leica M844 C19 Surgical Microscope, article 10448197; ceiling mounted surgical microscopes; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, Switzerland;; The intended use of the microscope is for improving the visibility of objects through magnification and illumination. It can be applied for observation and documentation and for human and veterinary medical treatment. | Surgical microscope | 2 | March 21, 2012 | 2012 | Leica Microsystems, Inc. | The screws, which lock the bottom plate to the spindle, pull through allowing the carrier arm with touch panel/control unit of the surgical microscope to disconnect, which potentially could result in them falling. | DESIGN: Device Design | The firm, Leica Microsystems, sent a "Medical Device Correction" letter dated February 27, 2012 to the affected customers via e-mail on the same date. The letter described the product, problem and action to be taken. The customers were informed that the Leica M820 and M844 C19 surgical microscopes require upgrading. An upgrade kit will be provide for replacement of the flange and axis of the touch panel/control unit carrier. The Customers were instructed to NOT USE the affected Leica surgical microscope until the unit has been upgraded by their Leica Microsystems representative and to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the serial numbers of the unit at their facility and fax it back to Leica at 1-847-236-3747. If they have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: California, Connecticut, Kansas, Missouri, New York, North Dakota, Ohio, Pennsylvania, Texas and Utah; and countries of: Japan and South Africa, | Terminated | May 18, 2012 | 58 | 33 | Not_Computer
Z-1246-2012 | 61314 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | ENSEALÂŽ 5 mm Diameter Tissue Sealer G2 14 cm Length Straight Jaw, Model # NSLG2S14 and ENSEALÂŽ 5 mm Diameter Tissue Sealer G2 25 cm Length Straight Jaw, Model # NSLG2S25. Ethicon Endo-Surgery, LLC Guaynabo, Puerto Rico 00969The ENSEALÂŽ G2 Tissue Sealers are indicated for bipolar coagulation and mechanical transection of tissue during laparoscopic and open procedures. The devices are bipolar electrosurgical instruments for use with an electrosurgical generator. They are intended for use during open or laparoscopic, general and gynecological surgery to cut and seal vessels, and to cut, grasp and dissect tissue during surgery. Indications for use include open and laparoscopic, general and gynecological surgical procedures (including urologic, thoracic, plastic and reconstructive, bowel resections, hysterectomies, cholecystectomies, gall bladder procedures, Nissen fundoplication, adhesiolysis, oophorectomies, etc.), or any procedure where vessel ligation (cutting and sealing), tissue grasping and dissection is performed. The devices can be used on vessels up to (and including) 7 mm and bundles as large as will fit in the jaws of the instruments. | ENSEAL 5 mm Diameter Tissue Sealer G2 14 cm Length Straight Jaw | 2 | March 21, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENSEALÂż G2 Curved and Straight Tissue Sealers due to two potential issues that may occur related to the activation button: (i) Continuous Activation: The ENSEALÂż device/system may continue to activate when the activation button is released prior to reaching the end of cycle (tone 3). (ii) No Activation | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ethicon Endo-Surgery sent an Urgent Medical Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, fill out the attached Business Reply Form (BRF) and return to and return the affected product to:Ethicon Endo-SurgeryATT: ENSEAL G2 Recall 4545 Greek RoadCincinnati, Ohio 45242Customers were instructed to choose one of the following response options:Return the Business Reply Form to their sales representativeCall 1-800-873-3636, Option 6Fax the BRF to 1-513-337-4138For any questions regarding this recall call 513-337-3419. |  | 84 | 84 pieces | Worldwide Distribution - USA including AZ, CA, CT, FL, IL, IN, KS, KY, MA, MI, MO, NC, NY, OH, PA, TN & TX and the countries of Arab Emirates, Austria, Belgium, France, Germany, Great Britain, Italy, Jordan, Kingdom of Saudi Arabia, Lebanon, Portugal, Slovenia, Sweden and Turkey. | Terminated | February 15, 2013 | 331 | 1035 | Not_Computer
Z-1263-2012 | 60753 | kit, wound dressing | MCY | Elastic bandage. | General Hospital | General & Plastic Surgery | Enforcement Discretion | Wound Management Tray, catalog# 900-540 | Wound Management Tray | 2 | March 21, 2012 | 2012 | Customed, Inc | The firm has determined that there is a possibility that packaging seals may be compromised. The issue could result in an injury to the patient due to product contamination or loss of sterility condition. | PRODUCTION CONTROLS: Packaging | Customed Inc sent an Urgent Medical Device Recall letter dated December 8, 2011 to distributors via e-mail; then by certified (return and receipt) mail. The letter identified the product, description of problem, and actions to be taken. The letter included the following instructions:1. Product should be removed from sale (Quarantine them) and cease distribution. Products affected shall NOT be used.2. Return the enclosed acknowledgement form to Customed, Inc. Recall Coordinator whether or not the consignee has affected products. 3. If products has been futher distributed, sub-recall products - It is extremely important to notify your customers of this action by including a copy of this leller along with your formal notification. Please, instruct your customer to follow your directions for credit/replacemenl if they purchased the product from you. For qucstions. contact Cuslomed, Inc. Recall Coordinator at 787-622-5151 Ext. 7510. |  | 400 | 400 units | USA Nationwide Distribution - including Florida and Puerto Rico. | Open |  |  | 184527 | Not_Computer
Z-1278-2012 | 61351 | components, wheelchair | KNN | Wheelchair component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Permolock C3, Docking system for Power wheelchair in vehicle.Intended use: locking device | Permolock C3 | 2 | March 22, 2012 | 2012 | Permobil Inc | The fasteners used between June 3, 2011 and January 31, 2012 to attach the Permolock C3 pins to wheelchairs manufactured in the United States do not contain the required nylon ring, which allows for a more secure attachment. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Permobil, sent an "URGENT DEVICE RECALL" letter dated February 27, 2012 via fax to all customers. For those customers where Permobil did not have a valid fax number, the notices were mailed via first class United States mail. The letter describes the product, problem and actions to be taken. Permobil will designate Permobil personnel to complete an update to the Permolock C3 device.If a Permobil representative has not contacted you already, please call 800-736-0925 to schedule the service. |  | 147 | 147 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, MD ,MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, VT, WV, and WY; and countries of: Australia, Canada, France, Netherlands, Sweden, and United Kingdom. | Terminated | August 13, 2013 | 509 | 147 | Not_Computer
Z-1276-2012 | 61312 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Osteonics Shoulder Glenoid SpannerHowmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430.The Solar Upper Extremity System Hemi and Total Shoulder Arthroplasty system is designed to address the most common arthritic disorders affecting the shoulder. The spherical reamer is used to prepare the glenoid fossa and provide a uniform surface for instrumentation and implant placement. The glenoid fossa is then sized using surface templates to confirm the appropriate coverage, size and curvature has been determined. To achieve this in surgery, the spherical reamer (5900-2050/-2051/-2052/-2053) must be attached to the Angled Glenoid Driver (5900-2040) by threading the reamer's male threads into the driver's female threads. The spherical reamers come in sizes 5, 7, 9 & 11, each with a captive drill head screw (5900-2054/-2055). The spanner features two kidney shaped male lobes that mate with two similar kidney shaped female slots in the screw's head and is used to tighten and loosen the spherical reamers with the angled driver. These instruments are part of the Solar Total Shoulder Instrument set and can be found in Surgical Protocol LSPUE-4 dated 10/04. Each set contains one spanner that universally fits all size reamers. | Osteonics Shoulder Glenoid Spanner | 2 | March 22, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Two product complaints reported that the Glenoid Reamer Spanner (Lot F3W7017) would not properly connect to the Glenoid Spherical Reamer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 21, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgment Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions regarding this notice. |  | 5 | five units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 19, 2012 | 272 | 5 | Not_Computer
Z-1248-2012 | 61356 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Medtronic EnTrust, Dual Chamber Implantable, Cardioverter Defibrillator with Atrial and Ventricular Therapies, models D153ATG and D154ATG. The device is indicated for use in ICD patients with atrial tachyarrhythmias, or who are at significant risk of developing atrial tachyarrhythmias. The device is intended to provide ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life-threatening ventricular arrhythmias. In addition, the device is intended to provide pacing, cardioversion, and defibrillation for treatment of patients with symptomatic, drug-refractory atrial fibrillation and/or life-threatening ventricular tachyarrhythmias. Sterile. Medtronic, Inc., Minneapolis, MN 55432 USA. | EnTrust AT ICD | 2 | March 22, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | A small percentage of EnTrust ICDs may not meet expected longevity or provide at least three months of device operation between the Elective Replacement Indicator (ERI) and End of Life (EOL) due to a more-rapid-than-expected drop in battery voltage. |  | An "Urgent Medical Device Correction" letter was sent to physicians beginning 3/6/2012. The letter described the issue and provided patient management recommendations. Additionally the letter stated that ongoing updates to the product performance report will be posted on a Medtronic website. | Contact the recalling firm for information | 68804 | 68,804 devices worldwide | Worldwide Distribution -- USA, and the countries of: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | July 27, 2012 | 127 | 68804 | Not_Computer
Z-1281-2012 | 61420 | device, vein location, liquid crystal | KZA | Liquid crystal vein locator. | General Hospital | General Hospital | 510(K) Exempt | VeinViewer (R) by Luminetx, Model Number VV1.1 GS, Part No. P00800-L, Luminetx Corporation, Memphis, TN 38104Product Usage: CTS Usage:Non- invasive electronic visual aid device designed to project an image of a superficial subcutaneous vascular structures on the surface of the skin | VeinViewer | 2 | March 22, 2012 | 2012 | Christie Medical Holding Inc | Instrument may tip over if not moved correctly. | TRAINING: Employee Error | Christie Medical Holdings, Inc. sent an Urgent Medical Device Correction by letter on March 14, 2012 to all affected customers. The letter identified the product, problem and actions to be taken. The letter includes "Correct Transport" and "Do Not Pull labels to be applied to the unit. Customers were instructed to immediately examine their inventory, quarantine all product subject to correction and apply the labels prior to use. A Service Bulletin is enclosed that should be kept and placed with the User Guide for future reference and a Response Form to be completed and return as soon as possible. For questions contact Technical Service at 901-721-0300. | Contact the recalling firm for information | 600 | 600 units | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, China, Brazil, Czech Republic, Canada, Germany, Denmark, Spain, Great Britain, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, South Korea, Kuwait, Mexico, Malaysia, Netherlands, Norway, Philippines, Russian Federation, Saudi Arabia, Singapore, Thailand, Turkey, Taiwan, and United Arab Emirates | Terminated | October 23, 2012 | 215 | 1050 | Not_Computer
Z-1279-2012 | 61205 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman ICP Skull Bolt Kit, Product Code: 82-6638. Single Use, disposable drill kit for gaining access to the cranium to allow for insertion and retention of intracranial cranial pressure sensor, included in the kit. | Codman ICP Skull Bolt Kit | 2 | March 22, 2012 | 2012 | Codman & Shurtleff, Inc. | Codman ICP Skull Bolt wrench is out of specification and cannot be used to turn the screw on the drill | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Codman issued "Urgent - Voluntary Product Recall Notification" letters on 2/9/12 via Fedx in the US and requested accounts examine inventory and return product. An Acknowledgement Form was requested to be returned. Outside the US affiliates were notified to recall to their accounts. Questions were directed to Customer Service (800) 225-0460. | No consumer action necessary | 4056 | 4056 kits | Worldwide Distribution - USA, and the countries of Canada, Austria, Brazil, Belgium, Columbia, China, Denmark, Estonia, Spain, Finland, France, Great Britain, Germany, India, Italy, Japan, Lithuania, Netherlands, New Zealand, Oman, Mexico, Portugal, Russia, Singapore, SIovenia, and Switzerland. | Open |  |  | 4056 | Not_Computer
Z-1287-2012 | 60718 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Soft-Flow Aortic Cannula; Aortic cannula, 7.0 mm (21 Fr) OD with 3/8" connector, 12" (31 em) long.Product Usage: CTS Usage:The Sarns-Aortic Cannula is indicated for use In perfusion of the ascending aorta during cardiopulmonary bypass surgery. The X-Coating is intended to reduce the adhesion of platelets to the blood contacting surfaces of the device during cardiopulmonary bypass procedures. | Sarns SoftFlow Aortic Cannula | 2 | March 23, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During in-house inspection, Terumo Cardiovascular Systems discovered one Soft-Flow Aortic Cannula with a hole in the tip, just above the suture ball. | DESIGN: Process Design | Terumo sent an URGENT MEDICAL DEVICE RECALL letter dated February 23, 2012 via Federal Express to all US customers. Terumo Foreign Affiliates received the recall notification via email on February 24th, 2012. The letters identified the affected products, reson for removal, potential hazard, affected population and actions to be taken. Customers were instructed to review the Medical Device Removal notice, assure that all users are aware of the notice, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number provided or email to tcvs.recall@terumomedical.com and call Terumo CVS Customer Service for a Returned Goods Authorization (RGA) number. For questions call Terumo CVS Customer Service at 1-800-521-2818. |  | 248 | 188,282 Cannulae and 248 Custom Tubing Packs | USA Nationwide Distribution including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of: CANADA, CHILE, India, Japan, Singapore, AUSTRALIA, Malaysia, BELGIUM, BELGIUM, BRAZIL, CHILE, BRAZIL, BRAZIL, and Philippines. | Terminated | April 03, 2013 | 376 | 10454 | Not_Computer
Z-1284-2012 | 61278 | system, endovascular graft, aortic aneurysm treatment | MIH |  |  | Cardiovascular | PMA | Talent Captivia Stent Graft System - FreeFlow configurations only.Model numbers affected have the following designation:TFXXXXXXXXCP and TBXXXXXXXXCP.Medtronic Inc.Santa Rosa, CA 95403.The TalentÂŽ Thoracic Stent Graft System is designed to treat thoracic aortic aneurysms using an endovascular approach. When placed within the target lesion, the stent graft provides an alternative conduit for blood flow within a patient s vasculature by excluding the lesion from blood flow and pressure. | Talent Thoracic Stent Graft System | 2 | March 23, 2012 | 2012 | Medtronic Vascular, Inc. | Complaints received related to the inability/difficulty to release the proximal (FreeFlow) bare stent using the tip capture mechanism. | DESIGN: Process Design | Medtronic, Inc. sent an Urgent Medical Device Safety Information letter dated February 2012, to all affected customers. The letter identified the product, the problem, and the corrective action taken by Medtronic. Medtronic updated the Instructions for Use (IFU) manual.For questions regarding this recall call 707-566-1110. |  | 1523 | 1523 units | Nationwide Distribution | Terminated | July 12, 2012 | 111 | 1523 | Not_Computer
Z-1228-2012 | 61268 | ventricular (assisst) bypass | DSQ |  |  | Cardiovascular |  | HeartMate II Left Ventricular Assist System (HM II LVAS) Sealed Outflow Graft Bend Relief.The Sealed Outflow Graft Bend Relief is provided as a component of the HM II Sealed Outflow Graft, which in turn is a component of the HM II LVAS Implant Kit with Sealed Grafts. A short configuration of the HM II LVAS Sealed Outflow Graft Bend Relief is provided as a stand-along product.Thoratec Corporation6035 Stoneridge DrivePleasanton, CA 94588.Product Usage:The HM II LVAS is intended to provide hemodynamic support in patients with end-stage, refractory left ventricular heart failure; either for temporary support, such as a bridge to Cardiac Transplanation or myocardial recovery, or as permanent destination therapy. The HeartMate II is intended for use inside or outside the hospital. The outflow of the HM II LVAS blood pump is connected to the aorta with a flexible vascular graft. The outflow graft is provided with a bend relief, a three or four in long tube of reinforced ePTFE surrounding the outflow graft proximal to the pump that is designed to prevent kinking of the outflow graft. | HeartMate II Left Ventricular Assist System | 1 | March 23, 2012 | 2012 | Thoratec Corporation | A trend of disconnected bend reliefs on the sealed outflow graft used with the HeartMate II LVAS has been identified | DESIGN: Component Design/Selection | Thoractec Corporation sent an Urgent Medical Device Correction letter on February 23, 2012 via overnight courier and e-mail to all affected consignees. The Urgent Medical Device Correction notification letters identify the affected products, description of the problem with illustrations, symptoms of problem and immediate actions to taken. All users are asked to review the attached labeling, including new caution statements and share with all surgical personnel. Consignees are instructed to completed and returned the Acknowledgement Form indicating that the information was received and understood. Additionally, an update concerning the Worldwide Medical Device Correction Notification is enclosed with each letter.Questions should be directed to Thoratec Regulatory Affairs at 925-847-8571. |  | 3852 | 3852 units in commerce | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Denmark, Germany, Greece, France, Hong Kong, Kuwait, Ireland, Israel, Italy, Lebanon, Liechtenstein, Lithuania, Luxembourg, Malaysia, Netherlands, Turkey, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand and United Kingdom. | Terminated | November 08, 2012 | 230 | 3852 | Not_Computer
Z-1282-2012 | 61305 | culture media, general nutrient broth | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | Fildes Enrichment, BBL brand, catalog #211866, in 5 ml tubes,10 tubes/package, manufactured by BD Diagnostic Systems, 250 Schilling Circle, Cockeysville, MD 21030.Fildes Enrichment may be used to enrich a variety of media for the cultivation of various microorganisms. | Fildes Enrichment | 2 | March 23, 2012 | 2012 | Becton Dickinson & Co. | In vitro diagnostic reagent is defective and may fail to enhance the growth of microbiological cultures as intended. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Becton Dickinson (BD) sent an "URGENT PRODUCT RECALL" letter dated February 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining affected product. BD will issue replacements for the discarded product. Additionally, an Acknowledgment Form was attached for customers to complete and return to the firm via fax to 410-316-4258. Contact BD Technical Services Department for questions regarding this notice. |  | 6930 | 6930 tubes | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Japan, Mexico, and Taiwan. | Terminated | June 18, 2012 | 87 | 6930 | Not_Computer
Z-1315-2012 | 61282 | [6 empty] | QRS Diagnostic Universal ECG 12 Channel ECG (IEC), Rx only, REF Z-7000-0400, Manufactured for QRS Diagnostic, 14755 27th Ave North, Plymouth, MN 55447 USA.Product Usage:The Universal ECG is an Electrocardiograph Diagnostic System designed to provide measurements and interpretative statements of a patients ECG waveform. The interpreted ECG with measurements and diagnostic statements are provided to the physician on an advisory basis only, for the physician to over read and validate or change the ECG interpretation. The ECG measurements and interpretive statements may be reviewed, edited, distributed electronically or printed. No real time monitoring or alarm capability is provided. | QRS Diagnostic Universal ECG 12 Channel ECG (IEC), | 2 | March 23, 2012 | 2012 | Qrs Diagnostic | QRS Diagnostic has initiated a recall involving Universal ECG 12 Lead IEC ECG Devices. This recall has been initiated because the chest lead wires on these devices may have been improperly labeled. Incorrect labeling of the lead wires could result in improper lead wire placement on the patient which may result in inaccurate waveforms being displayed to the user and an inaccurate interpretatio |  | QRS Diagnostic sent an Urgent Medical Device Recall letter dated February 9, 2012 via e-mail and mail to all affected consignees. The letter identified the affected product, description of the problem, and actions to be taken. If the affected product have been further distributed, consignees are instructed to notify them at once of the product recall. Consignees were instructed to contact Uni-com immediately for a replacement device. Uni-com will provide this replacement at no charge and will arrange to have the device returned to QRS Diagnostic for correction. Additionally, consignees were instructed to complete the attached Medical Device Correction Receipt Confirmation Form and return to QRS Diagnostic via fax at country code =1-763-559-2961 or email support@QRSDiagnostic.com. For questions contact Technical Support at support@QRSDiagnostic.com or by phone 800-465-8408. | Contact the recalling firm for information | 13 | 13 devices | Worldwide Distribution and the countries of: Belgium, France, Italy, Mali and Senegal. | Terminated | August 02, 2012 | 132 | 13 | Not_Computer
Z-1286-2012 | 61207 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | Product is labeled in part: "***REF: 01-7036, LOT 999999, "1.5MM" SYSTEM PLATE L-SHAPE, 3 X 3 HOLE, 100 DEGREE, QTY 00001, CE 0086, 2010-04 *** NON-STERILE PRODUCT ***, LEFT, X-LONG, 0.6 MM TITANIUM, BIOMET MICROFIXATION, 1540 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA, 01-736, LOT 999999, QTY 00001, L-SHAPE, 3 X 3 HOLE, 100 DEGREE, LEFT, X-LONG, 0.6 MM, CE 0086.***"Product Usage - CTS Usage:Bone-plate - internal fixation devices intended to aid in the alignment and stabilization of bone in the oral cranio-maxillofacial skeletal system. | Bone Plate | 2 | March 23, 2012 | 2012 | Biomet Microfixation, LLC | On 4/26/10 Biomet Microfixation, Jacksonville, FL initiated a recall of their 1.5 3/3 hole 100 degree left X-long L plate, Model #01-7036, Lot #193690. Product is labeled incorrectly. | PRODUCTION CONTROLS: Error in Labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated April 21, 2010 to all consignees affected. Each consignee was notified by E-mail, letter or telephone call the week of April 26, 2010. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to fill out the attached Inventory Reconciliation sheet and FAX it to the attention of the Global Complaint Coordinator at 904-741-9425. If the affected product have been further distribute, consignees are instructed to notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. | Contact the recalling firm for information | 23 | 23 | USA Nationwide Distribution including the states of: California and Texas. | Terminated | April 18, 2012 | 26 | 23 | Not_Computer
Z-1316-2012 | 61222 | ldl & vldl precipitation, cholesterol via esterase-oxidase, hdl | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products --- dHDL Slides --- Ortho Clinical Diagnostics, a Johnson&Johnson; company --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA --- VITROS dHDL Slides are used to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. --- Product codes: 6801895 (60 slide count); 6802469 (18 slide count) --- indicated to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. | VITROS | 2 | March 26, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) has revised the VITROS dHDL Slides Instructions for Use (IFU) and removed EDTA plasma as a recommended sample type. | DESIGN: Device Design | Ortho Clinical Diagnostics sent an Important Product Information letter dated February 2, 2012, to all affected customers via Federal Express overnight mail. The letter informed customers of the issue, notifying them of the removal of EDTA plasma as a recommended sample type and providing the revised VITROSÂŽ dHDL Slides Instructions for Use. J&J; Foreign affiliate consignees were notified by email informing them of the issue on 02 February 2012. Customers were instructed to replace the current pages in their VITROS Chemistry Products Instructions for Use Manual with the updated instructions for Use and Table of Contents.Customers were asked to return the enclosed Confirmation of Receipt form regardless of whether their laboratory uses the affected product. For any questions regarding this recall call 1-800-421-3311. | Contact the recalling firm for information | 11.364 | 120,225 of 60-slide count, 11.364 of 18-slide count | Worldwide Distribution | Open |  |  | 120235 | Not_Computer
Z-1323-2012 | 61271 | unit, heating, powered | IRQ | Powered heating unit. | Physical Medicine | Physical Medicine | 510(K) Exempt | The LactoSorb Heat/Contouring Pen is a disposable alkaline-battery powered device. Product is labeled in part: "***REF 915-3100 LOT***LACTOSORB* SYSTEM HEAT/CONTOURING PEN AND TIPS***QTY 00001***STERILE R STERILE EO***PVC, BRASS***WARNING: DO NOT USE IN THE PRESENCE OF FLAMMABLE MATERIALS. FIRE COULD RESULT.***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA***#915-3100 LOT***QTY 00001, HEAT/CONTOURING PEN AND TIPS***PVC, BRASS***".Product Usage: Two tips are provided with the body of the device to cut / contour the edges of the LactoSorb plates by heating the material. | LactoSorb Heat/Contouring Pen | 2 | March 27, 2012 | 2012 | Biomet Microfixation, LLC | On 11/15/2011 Biomet Microfixation, Jacksonville, FL initiated a recall of their Acticort Heat/Contouring Pen, Model #915-3100, Lot #861200. Product is incorrectly labeled | PRODUCTION CONTROLS: Error in Labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated November 15, 2011 via E-mail, letter or telephone call to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each consignee was Instructioned to take the following actions: 1) If you have any unopened boxes of the product in question, please place this notice with it; 2) Please fill out the attached Inventory Reconciliation sheet and FAX to the attention of the Global Complaint Coordinator; and 3) If you have further distributed this product , you must notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions please contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. | Contact the recalling firm for information | 42 | 42 | Worldwide Distribution - USA (nationwide) including the state of Texas and the country of Japan. | Open |  |  | 42 | Not_Computer
Z-1324-2012 | 61177 | media, reproductive | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sydney IVF Culture Oil, REF, K- SICO- 50, 50 mL. and K-SICO-200, 200 mL,Store at 2-8 C, Sterile, Cook Medical, William A Cook Australia LTD, Product Usage:Culture oil is intended to be used as an oil overlay for cultures of gametes, zygotes, or embryos in assisted reproductive technology (ART) and micro manipulation procedures. It is used to reduce the osmotic stress due to evaporation and to reduce pH fluctuations. | Sydney IVF Culture Oil | 2 | March 28, 2012 | 2012 | Cook, Inc. | Cook Medical is recalling Sydney IVF Culture Oil due to a number of recent customer complaints that have identified it as a possible cause of embryo degeneration at day 5/6. | PRODUCTION CONTROLS: Process Control | Cook Medical sent an Urgent Recall Notices letter dated February 15, 2012 to all affected customers. The letter identified the recall product, reason for the recall and actions to be taken. The notice included directions to pull the recalled product from stock, isolate it and forward the notice to any sub-accounts. Customers were requested to complete and return the recall acknowlegement and product tracking form by to Fax: 812-829-1801 or Email: julie.cooper@cookmedical.com. For questions call 800-457-4448 x7241. | Contact the recalling firm for information | 54 | 54 total from both lots | Worldwide Distribution - USA (nationwide) including the states of IL, KS, KY, MA, NJ. TN, TX, UT, VA, WI, WV, and the country of Barbados. | Terminated | July 03, 2012 | 97 | 54 | Not_Computer
Z-1319-2012 | 61377 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Boston Scientific Soloist Single Needle Electrode: 18 cm long, 16 gauge, Sterile, packaged with the Soloist Electrode Power Cord. Boston Scientific, Nalick, MA.The Soloist Needle Electrode is intended for use in conjunction with a Boston Scientific Corporation radiofrequency. (RF) generator for the thermal coagulation necrosis of soft tissues, including partial or complete ablation of nonresectable liver lesions. | Boston Scientific Soloist | 3 | March 28, 2012 | 2012 | Boston Scientific Corporation | Device found labeled with the incorrect expiration dates. The Soloist Single Needle Electrode has been qualified for a 3 year expiration date. Product manufactured between the dates of May 24, 2011 and November 8, 2011 has an incorrect expiration date of 20 years labeled on the pouched needle electrode and 5 years labeled on the carton. |  | Boston Scientific sent a Urgent Medical Device Recall letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed to cease use, quarantine and return unused product with directions to forward the notification to the hospital level. Customers were directed to fill out the Reply Verification Tracking Form and fax to 508ÂŽ683ÂŽ5578, for reimbursement. Returned goods are to be sent to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. For further questions please call (508) 683-4678. | Contact the recalling firm for information | 682 | 682 (1555 used) | USA (nationwide) and the countries of Canada, Columbia, Viet Nam, Germany, France, Great Britain, Italy, Jordan, Kuwait, Netherlands Pakistan, Portugal, S. Arabia, Slovakia, Spain, Sweden and Tunisia. | Terminated | November 14, 2012 | 231 | 682 | Not_Computer
Z-1334-2012 | 61329 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion CO2 sampling lines and water traps under the label:Omnistream CO2 Sampling Lines:OmniLine O2 Adult Part No.007609;OmniLine O2 Pediatric Part No 007610;Smart OmniLine Plus O2 Adult Part No.010177;Smart OmniLine Plus O2 Adult (package of 100 units) Part No.010213;Smart OmniLine O2 Pediatric Part No. 007606;;Smart OmniBloc O2 Part No. 010946Smart OmniLine Plus part No. 010172Smart OmniLine Plus (package of 100 units) Part No 010212;Smart OmniLine Guardian O2 Part No. 012531;Smart OmniLine Guardian O2 Long Part No. 012532OmniVentLine Set Part No. 012495 | Oridion | 2 | March 29, 2012 | 2012 | Oridion Medical 1987 Ltd. | CO2 sampling lines may have paper and plastic particulate present in the packaged units at levels that may exceed visual specification. | PRODUCTION CONTROLS: Packaging Process Control | Oridion notified Hospitals, Clinics, EMS with an "Important Communication" letter to clinicians dated February 12, 2012. The letter was intended to make US clinicians aware of an important issue relating to certain Oridion CO2 sampling lines, which may not fully meet quality requirements. In a few cases, paper and plastic particulate may be present in the packaged units at levels that exceed the visual specification. Recommended actions were included in the letter. Questions were directed to 1(781)-972-1252. |  | 2 | 2 million | Nationwide Distribution | Terminated | September 13, 2013 | 533 | 2 | Not_Computer
Z-1331-2012 | 61276 | preservative, cytological | LEA | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Surgipath Cyto Jar Fixative; a 120 mL clear, leakproof container prefilled with 45 mL of preservative containing polyethylene glycol for the convenient collection, fixation and preservation of all cytologic specimens, 50 jars per case; for in vitro diagnostic use only; Leica Biosystems, 5205 Rt. 12, Richmond, IL 60071; Made in USAThe Cyto Jar is for the convenient collection, fixation and preservation of all cytologic specimens. The 120 mL clear, leakproof container is prefilled with 45 mL of preservative containing polyethylene glycol. It can be used for rinsing aspiration needles or bronchial brushes and for the direct collection of sputum specimens. | Cyto Jar Fixative | 3 | March 29, 2012 | 2012 | Leica Biosystems Richmond Inc. | Product was distributed with the incorrect lot number and expiration date. The lot number and expiration date should have been listed at 073013 and instead was noted as 073014 which is beyond the maximum expiration dating of 18 months for this product. | DESIGN: Process Design | The firm, Leica Microsystems, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 17, 2012, via FedEx on February 20, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken. A second letter was sent to the customers on February 29, 2012, clarifying that the incorrect lot number and expiration date was 073014, and should have been 073013. The customers were instructed to discontinue use of the product; return the product and complete and return the enclosed Acknowledgement Form via fax to 815-678-2216 within one week of receiving the notification.If you have any questions, contact the Regulatory Affairs Manager at toll free (800) 225-8867 or email: richmond@leica-microsystems.com. |  | 11 | 11 cases | Nationwide distribution: USA including states of: Michigan, Missouri and New York. | Terminated | September 20, 2012 | 175 | 11 | Not_Computer
Z-1330-2012 | 61375 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sorin C5 (perfusion system)The S5 Perfusion System is used to perform, control and monitor extracorporeal blood circulation during an operation for up to six (6) hours. | Sorin C5 (perfusion system) | 2 | March 29, 2012 | 2012 | Sorin Group Deutschland GmbH | The pump speed control knob is attached to the shaft of an encoder that is used to adjust the speed. Insufficient lubrication within the encoder increases friction over time making it more difficult to turn the pump speed control knob. Excessive friction causes premature wear of the encoder shaft which can result in a stuck shaft. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sorin Group sent an Urgent Field Safety Notice dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Before each use of the device during the set-up procedure, rotate all pump speed control knobs back and forth to check for smooth operation. If abnormal friction or any binding is detected, please do not use the system until it is repaired.If an affected product is currently in use with a patient and an increase of friction is detected please prepare for a replacement pump.Please complete and return the attached Confirmation Form by fax to 303-467-6502.Customers were instructed to ensure that the Urgent Field Safety Notice is distributed to all personnel in their organization who need to be aware of the notice. If the product was transferred to a third party, customers were asked to share the information with them as well as Sorin Group USA Customer Service at 1-800-850-2623.For questions regarding this Urgent Field Safety Notice, contact Sorin Group USA Customers Service at 1-800-650-2623.All affected customers and distributors were notified of the Field Correction by certified mail beginning on February 23, 2012. They will be contacted by Sorin Service Group to arrange an appointment to replace affected shaft encoders on site. |  | 39 | 39 | Worldwide Distribution | Terminated | November 27, 2012 | 243 | 1147 | Not_Computer
Z-1188-2012 | 61194 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Aerolase Corporation. The products involved are the FriendlyLight YAG lasers, a Model Family first identified and reported.The products are marketed under names including FriendlyLight Er:YAG, LightPod Era, FriendlyLight Nd:YAG, LightPod Neo, LightPod Neo XL, LightPod Neo XT, LightPod Forte, LightPod Allure, and LightPod Nova. | Aerolase lasers | 2 | March 29, 2012 | 2012 | Aerolase | It was discovered that customers were not supplied with a procedure and schedule for calibrating the FriendlyLight YAG lasers. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Aerolase issued a letter to customers, dated 11/30/2011, identifying the affected product and the reason for the recall. Aerolase is providing, at no charge, the Appendix to the Operating Manual for the LightPod lasers. This Appendix describes the procedure and frequency for checking whether the laser is within calibration, but it is not a procedure for customers to actually calibrate the laser themselves; any attempt to do so by any parties other than a company technician employed by Aerolase would void the warranty. Please contact us any time at 914- 345-8300 with any questions. | Contact the recalling firm for information | 176 | 176 | Worldwide Distribution | Open |  |  | 176 | Not_Computer
Z-1355-2012 | 61054 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Critical Care Bed Model 2031, 230 V , 50/60 Hz, Max Load = 225 Kg,, Stryker, Portage, MIProduct Usage:EPIC II (model 2031 ) Critical Care beds Med/Surg beds are AlC powered hospital beds, intended for clinician use, which consist of built-in motors and remote controls that can be operated by the patient and/or clinician to adjust the height and surface contour of the beds. The EPIC II model is primarily used in a Critical Care environment, but may be used in Med/Surg environments. Model 2031 represent International Voltage configurations of the 2030 model respectively. | Epic Critical Care Bed (Model 2031) | 2 | March 30, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that units manufactured between September 1, 2004 through November30, 2007, did not receive updated preventive maintenance information. (Failure to maintain the bed brakes including the brake plate kits, may result in reduced brake holding force. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical int | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Stryker Medical sent an Urgent-Medical Device Correction letter dated March 12, 2012 to all consignees. The letter identified the affected products, problem, updated preventive maintenance information and requested the following immediate actions: 1) Locate the beds listed in this notice, 2) Ensure that the brakes are holding properly, 3) If the brakes are not holding, contact Stryker (1-800-327-0770, option 4) to order the maintenance parts, 3) Continue to periodically inspect brake functionality, as noted on enclosed preventive maintenance schedule/checklist, 4) Order replacement brake components, as needed, 5) Return the enclosed post card to confirm receipt of this notification, 6) If you have loaned or sold any of the beds listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location, 6) If you have disposed of any of the beds and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers. For questions or concerns, please, contact Sandy Cliche at 269-389-6025. |  | 507 | 507 | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: DEDEDO GU, HAMILTON ON, SANTIAGO, SHATIN, NT , CAPITOL FEDERAL, ITAPEVI SP, ATHENS, OSAKA, SEOUL, MEXICO DF MX, MONTREUX, BALERNA, NORTH POINT, WARSZAWA, CANOVANAS, KOWLOON BAY, SINGAPORE, NEWBURY, and JOHANNESBURG. | Terminated | August 29, 2013 | 517 | 121982 | Not_Computer
Z-1349-2012 | 61143 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Surestream CO2 Sampling Lines-Sure VentLine Set Adult/PediatricPart Number: 010986-03Product Usage:The intended use of SureStream sampling lines is to convey a sample of a patients expired breath to a capnograph in order to measure CO2 partial pressure in intubated and non intubated patients. | Oridion | 2 | March 30, 2012 | 2012 | Oridion Medical 1987 Ltd. | Surestream CO2 Sampling Lines-Connector may cause breakage of the microswitch inside the CO2 inlet | DESIGN: Device Design | Oridion issued a Technical Note communication in March 2011 to field staff and end users globally. The Technical Note identified the affected product and provided instructions for user on the proper insertion of Surestream products. For questions contact your local distributor or Oridion at technicalsupport@oridion.com. |  | 481 | 481 boxes (12025 untis) | USA Nationwide Distribution including the states of: AZ, CA, FL, HI, LA, MA, MO, NJ, NY, OH, OR, TX and WI. | Terminated | June 04, 2012 | 66 | 841 | Not_Computer
Z-1336-2012 | 61362 | tubes, vacuum sample, with anticoagulant | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BacT/ALERT Blood Collection AdapterThe BacT/ALERT Adapter Cap is intended for use as an aid to facilitate filling of blood into the BacT/alert Blood Culture Bottle when using a butterfly blood collection set. | BacT/ALERT Blood Collection Adapter | 2 | March 30, 2012 | 2012 | bioMerieux, Inc. | The BacT/ALERT Blood Collection Adapter Cap is no longer compatible with Becton Dickinson (BK) Vacutainer Safety Lok Blood Collection Set. The incompatibility is the blood collection set needle backing out of the adapter cap that results in an exposed sheathed needle that may lead to an accidental needlestick. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | bioMerieux sent an Urgent Product Correction Notice letter dated March 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions: 1. Ensure the letter and guidance on how to maintain control of the luer connector by securing it between the thumb and forefinger are distributed to all appropriate personnel within their organization. 2. Fill out and return the Acknowledgement Form by Fax to confirm receipt of the recall notice. For additional assistance or questions customers were instructed to contact their local bioMerieux Customer Service Representative at (800) 682-2666, option 3, then option 3 again. Industry customers should call 800- 634-7656, option 3. For questions regarding this recall call 314-731-8537. |  | 274800 | 274,800 boxes (32,862,600 caps) | Worldwide Distribution - USA (nationwide), Australia, Brazil, Canada, Chile, Columbia, Costa Rica, France, China, Indonesia, Korea, Mexico, Peru, Singapore, Ecuador, Dominican Republic, Philippines, Guatemala, Curacao, Thailand, United Arab Emirates, Andorra, Austria, Burkina Faso, Congo, Switzerland, Cameroon, Germany, Algeria, Estonia, Spain, France, Great Britain, French Guiana, Croatia, Israel, India, Italy, Kazakhstan, Lithuania, Latvia, Macedonia, Malta, New Caledonia, Netherlands, French Polynesia, Poland, Portugal, Reunion, Russia, Sweden, Slovenia, Chad, Turkey, South Africa, Saudi Arabia, Congo, CoteDIvoire, Denmark, Djibouti, Finland, Greece, Guadeloupe, Libyan Arab Jamahiriya, Jordan, Morocco, Martinique, Norway, Senegal, Togo, Yemen, Bangladesh, Georgia, Ghana, Uganda, and Qatar. | Open |  |  | 274800 | Not_Computer
Z-1339-2012 | 61293 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F160NRe HemodialyzerCatalog number: 0500316EIntended for patients with acute or chronic renal failure when conservative therapy is judged to be inadequate. | Optiflux F160NRe Hemodialyzer | 2 | March 30, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of internal blood leaks affecting the hemodialyzers. | PRODUCTION CONTROLS: Process Control | Fresenius North America notified consignees on 3/6/12 by recall notification letter titled: Urgent FMCNA Optiflux F160NRe Hemodialyzer Recall, by certified mail with return receipt. The letter identified the affected product and described the reason for the recall. Customers were instructed to examine their stock immediately to determine whether they have any of the affected lot on hand. If customers have the affected lot, they have been instructed to immediately discontinue use and place all unused units in a secure area for return to Fresenius Medical Care North America. A Reply Response from is to be completed and returned. Customers have been instructed to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. Contact 1-800-323-5188. | Contact the recalling firm for information | 30576 | 30,576 units | Nationwide Distribution | Open |  |  | 30576 | Not_Computer
Z-1351-2012 | 61440 | accessories, implant, dental, endosseous | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Standard Plus Implant 04.8mm Wide Neck SLA 10mmArticle Number: 043.652S.Dental implants are intended for surgical placement in maxillary and/or mandibular arches to provide support for prosthetic restorations in edentulous or partially edentulous patients. Dental implants are for single-stage or two-stage surgery. Dental implants are intended for immediate placement and function on single-tooth and/or multiple tooth applications when good primary stability is achieved, with appropriate occlusal loading, in order to restore chewing function. Multiple tooth applications may be rigidly splinted. In the case of edentulous patients, four or more implants must be used. | Straumann Standard Plus Implant | 2 | March 30, 2012 | 2012 | Straumann USA, LLC | Mislabeled: Vial lid label on some of the dental implant vials from this lot incorrectly state 04.8 SP 10 RN instead of 04.8 SP 10 WN | PRODUCTION CONTROLS: Packaging Process Control | Straumann sent a "FIELD CORRECTIVE ACTION NOTICE" dated March 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. Additionally, a Customer Confirmation Form was enclosed for customers to complete and return. Contact the firm at 978-747-2514 for questions regarding this notice. | Contact the recalling firm for information | 346 | 346 units | Nationwide Distribution. | Terminated | May 31, 2012 | 62 | 346 | Not_Computer
Z-1360-2012 | 61234 | equipment, traction, powered | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Removal of Literature/Brochures for Antalgic-Trak. "***What is Antalgic-Trak?***Articulating Non-Surgical Spinal Decompression****Antalgic-Trak By Spinetronics***Note: "Decompression" as used herein refers to unloading due to distraction and positioning.***Antalgic-Trak is an improved form of non-surgical spinal decompression*. | AntalgicTrak Literature | 2 | April 02, 2012 | 2012 | Spinetronics, LLC | On 08/11/2011 Spinetronics, LLC initiated a recall of Antalgic-Trak Product Brochures which contained language and/or claims that were not cleared under the original 510k, and represented a major change or modification in the intended use of the device. | MISBRANDING: Labeling False and Misleading | Customer notification letters were mailed by certified mail via USPS. The medical user/customer was instructed to either return or destroy any remaining recalled product literature. Customers were asked to complete and return a response form noting the number of available brochures. | Contact the recalling firm for information | 1600 | 1600 brochures | Worldwide Distribution, including Nationwide (USA) and the country of Canada. | Terminated | April 20, 2012 | 18 | 1600 | Not_Computer
Z-1364-2012 | 61488 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM VerioUse: Nuclear resonance imaging system | Siemens MAGNETOM Verio | 2 | April 02, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring, Siemens has discovered that the gradient cable connections on a few systems did not meet their specifications after installation. The cable connections overheated, leading to the emission of smoke and possible material damage. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Siemens Medical Solutions USA, Inc., sent a "Customer Safety Advisory Notice" including Update Instruction MR028/10/S dated March 2012 to affected customers via hand delivery by Siemens Service Engineer . The notice described the product, problem and actions to be taken. The customers were instructed to create a COBR image and to make a copy of the attached "Completion Protocol" fill it out and file in their corresponding binder.Siemens Service Engineer performed an all gradient cable connection visual inspection as well as torque check on the connecting screws at the sites and the factory prior to their release to the field. Siemens also issued a current Update Instruction MR024/11/S. If you have any questions regarding this report, please contact the Quality Engineer, Product Complaints & Field Support at 610-219-4834; email: anastasia.mason@siemens.com or Director, Product Complaints & Field Support at 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 245 | 245 | Nationwide distribution to Washington, DC and the following States: AL, AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI. Government Accounts include: BROOKE ARMY MEDICAL CENTER, 3851 ROGER BROOKE DR, FORT SAM HOUSTON, TX 78234; NATIONAL INSTITUTE OF HEALTH, 10 CENTER DRIVE, BETHESDA, MD 20892; NAVAL MEDICAL CENTER, 34800 BOB WILSON DR, SAN DIEGO, CA 92134; US AIR FORCE REGIONAL HOSPITAL, 5955 ZEAMER AVE, ANCHORAGE, AK 99506; VA MEDICAL CENTER, 3350 LA JOLLA VILLAGE DR, SAN DIEGO, CA 92161; VA MEDICAL CENTER, 830 CHALKSTONE AVE, PROVIDENCE, RI 02908; VA MEDICAL CENTER, 109 BEE ST, CHARLESTON, SC 29403; VA MEDICAL CENTER, 10000 BAY PINES BLVD, BAY PINES, FL 33744; VA MEDICAL CENTER, 800 BAILIE DR, AUGUSTA, GA 30901; VA MARTINEZ 612, 150 MUIR RD, MARTINEZ, CA 94553; VA MEDICAL CENTER, 1501 SAN PEDRO SE, ALBUQUERQUE, NM 87108; VA MEDICAL CENTER, 1670 CLAIRMONT RD, DECATUR, GA 30033; VA MEDICAL CENTER - ST LOUIS, 915 N GRAND AVE, ST LOUIS, MO 63106; and VA MEDICAL CENTER HARLINGEN, 2106 TREASURE HILLS BLVD, HARLINGEN, TX 78550. "Other" is National Institute of Health. | Open |  |  | 245 | Not_Computer
Z-1357-2012 | 61333 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | The Polarus Plus Humeral Rod (part number HR-0824-S) is intended as an intramedullary fixation device for humeral fractures. It is sealed in an internal sterile blister pack within a box that is labeled in part: "Polarus Humeral Fixation System Size: 8mm X 240mm Polarus Plus Humeral Rod***Titanium***REF HR-0824-S***LOT 266220***RX only***Manufactured or distributed by: ACUMED 5885 NW Cornelius Pass Road Hillsboro, OR 97124 USA***" | The Polarus Plus Humeral Rod | 2 | April 02, 2012 | 2012 | Acumed LLC | The packaging contains the wrong product. It does not contain model HR-0824-S as the package indicates. | PRODUCTION CONTROLS: Labeling Mix-Ups | Acumed sent an Urgent Notice Device Recall letter dated March 5, 2012, to all affected consignees. A follow-up notification letter was sent to the domestic customers on March 9, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to further notify their consignees about the recall and to cease use of and return the affected product. Consignees were requested to respond using the enclosed response form. The form should be faxed to 503-520-9618. For questions regarding this recall call 503-627-9957, ext. 1302. | No consumer action necessary | 12 | 12 rods | Worldwide Distribution - USA (nationwide) and the countries of Italy, Ireland, and Great Britain. | Terminated | July 25, 2012 | 114 | 12 | Not_Computer
Z-1362-2012 | 61327 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-i ventilator systemProduct Usage: Intended for treatment and monitoring of patients in the range of neonates, infants and adults with respiratory failure or respiratory insufficiency. | Servoi ventilator system | 2 | April 03, 2012 | 2012 | Maquet Inc. | The Servo-i and Servo-s ventilator systems equipped with specific revision os of the AC/DC convertor can result in electric shock during servicing. | DESIGN: Device Design | Maquet sent a Product Recall Urgent-Medical Device Corrective Action letter dated December 22, 2011 to all affected customers. The letter identified the affected products, patient risk, product issue and actions to be taken. Customers were instructed to complete and return the attached Customer Fax Response form as soon as possible. Maquet Field Service Technicians will update all affected products at all the concerned customer sites. Customers are advised untill Maquet has updated all the affected products (all of which are still under warranty), inform all the concerned biomedical staff about the content of this Device Correction Letter. For technical support call 1-888-627-8383. |  | 3136 | 647 US; 2462 Foreign; 27 Spare Parts Foreign | USA Nationwide Distribution | Terminated | October 22, 2013 | 567 | 3186 | Not_Computer
Z-1371-2012 | 52174 | system/device, pharmacy compounding | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | The EM 2400 is a Pharmacy Compounding System consisting of three sub-systems for compounding solutions for intravascular administration.Pharmacy compounding system for mixing ingredients for intravascular administration. | ExactaMix 2400 Compounder (EM 2400) | 2 | April 05, 2012 | 2012 | Baxa Corporation | Remote possibility of an unintended ingredient delivery as a result of a damaged valve actuator on the EM 2400. | DESIGN: Device Design | The firm, Baxa, sent a "Exacta-Mix 2400 Compounder Safety Alert" dated June 3, 2008 to all customers. The alert describes the product, problem and actions to be taken. The customers were instructed to not use the main module if any of the 24 valve actuators are damaged; do not infuse any bags if damage is noticed after they have been pumped; do not try to repair or adjust a damaged or bent valve actuator; immediately alert all compounder users of this issue to ensure that when the device is used, the valve actuators are always inspected for damage;" inspect all valve acutators before setup and after tear-down to ensure they are not damaged or bent. Do not use the compounder if the valve actuators are damaged or bent."; and if a damaged or bent valve is noticed after using the EM2400, immediately hold all pumped bags and contact Baxa Technical Support. If you need additional assistance, please contact Baxa Technical Support (800-678-2292). Technical Support is available 24 hours a day, 7 days a week. | Contact the recalling firm for information | 1000 | 1000 units (approximately) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, China, New Zealand, and United Kingdom. | Terminated | April 10, 2012 | 5 | 1000 | Not_Computer
Z-1335-2012 | 61499 | laser vision | RFK |  |  |  |  | U-SPECT-II/CT and VECTor/CT cabinet x-ray products.Laboratory research using small animals. | USPECTII/CT and VECTor/CT cabinet xray products | 2 | April 05, 2012 | 2012 | Milabs Bv | These products have doors which are not equipped with at least one safety interlock which physically disconnects power to the high voltage power supply for the x-ray source. Additionally, these products do not have a certification label and the identification labels do not include the month and year of manufacture. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | A notification will be sent to all affected purchasers by April 15, 2012. The manufacturer will bring the products into compliance with the applicable performance standards within 60 days. Specifically at least one door safety interlock will be replaced per door and revised labels will be produced and applied to the system. These corrections will be made in accordance with MILabs instructions by a MILabs installation and service engineer. |  | 2 | 2 | Nationwide Distribution -- Indiana and North Carolina. | Terminated | August 29, 2012 | 146 | 2 | Not_Computer
Z-1386-2012 | 61405 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Nano Blood Glucose Meter No product labeling, beyond the labels applied to the meter, was provided with the meter as these meters were never intended to be distributed. A representative copy of product labeling that would have been included in a finished device kit is attached.The Accu-Chek Aviva Nano Blood Glucose meter is intended to be used with Accu-Chek Aviva test strips for the quantitative measurement of glucose (sugar) in fresh capillary whole blood samples drawn from the fingertips or palm. The system is intended to be used by a single person and should not be shared. The system is intended for self-testing outside the body (in vitro diagnostic use) by people with diabetes at home as an aid to monitor the effectiveness of diabetes control. The system should not be used for the diagnosis of or screening of diabetes or for neonatal use. Alternative site testing should be done only during steady-state times (when glucose is not changing rapidly). | ACCUCHEK Aviva Nano Blood Glucose Meter | 2 | April 09, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Accu-Chek Aviva Nano Blood Glucose meters that were manufactured for the purpose of validating a manufacturing process for product to be distributed outside the US were distributed in the US through eBay by an unauthorized third party. These meters were not intended for distribution, were not verified to meet final product specifications and were not distributed to customers with the necessary pr | TRAINING: Employee Error | Roche sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to immediately return their ACCU-CHEK Aviva Nano Meter to the firm in an enclosed pre-paid envelope via USPS. Contact Roche Product Removal Services at 1-888-345-5352 between 8:00 am and 5:00 pm for questions regarding this notice. |  | 82 | 82 meter distributed on Ebay | Nationwide Distribution-USA (nationwide) including the states of AZ, NY, RI, TN, NJ, IN, TX, AL, NV, CA, OR, KY, IL, AR, NM, NE, MD, MO, GA, and SD. | Terminated | November 06, 2012 | 211 | 82 | Not_Computer
Z-1390-2012 | 61393 | fluorometric, cortisol | JFT | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Cortisol AssayCat. Nos.04610138 (1 0321 026)- 250 Test kit04610049(10309078)- 50 Test kit04611509 (1 0335503)- 250 REF Test kitADVIA Centaur Systems Cortisol: For in vitro diagnostic use in the quantitative determination of cortisol in serum or urine using the ADVIA Centaur and ADVIA Centaur XP systems. | ADVIA Centaur Cortisol Assay | 2 | April 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Cortisol Assay reagent kit lots do not meet onboard stability and calibration interval claims | DESIGN: Component Design/Selection | The firm, Siemens Healthcare, sent an "Urgent Field Safety Notice" via e-mail to all affected customers on March 6, 2012 in the United States and International. The Urgent Field Safety Notice describe the product, problem, and actions to be taken. The notice instructed the customers to perform daily calibration on all Cortisol reagent kits to maintain consistently accurate quality control and patient values; to complete and return the confirmation fax-back form 'Field Correction Effectiveness Check' via fax to the Technical Solutions Center at (302) 631-7597 and forward this notification to whomever they many have distributed this product.If you have any questions, contact Quality Engineer 5- Quality Systems and Compliance at 508-660-8540 or www.siemens.com/diagnostics. |  | 40023 | Total: 40023. 50 Test Kits = 21734;250 Test Kits= 15879;250 REF Test Kits= 2410 | Worldwide distributed: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRALIA, BANGLADESH, BD, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, EGYPT, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, JAPAN, KOREA, MALAYSIA, MEXICO, NEW ZEALAND, PAKISTAN, PHILIPPINES, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, TAIWAN, THAILAND, URUGUAY, VENEZUELA, and VIETNAM. | Open |  |  | 40023 | Not_Computer
Z-1384-2012 | 61178 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SenDx Medical IncABL80FLEXBlood Gas, pH, Electrolyte, and Metabolite SystemThe ABL80 FLEX is a portable, automated analyzer that measures pH, blood gases, electrolytes, glucose, and hematocrit in whole blood. The ABL80 FLEX system is intended for use by trained technologists, nurses, physicians and therapists. It is intended for use in a laboratory environment, near patient or point-of-care setting. | SenDx Medical Inc. ABL80FLEX | 2 | April 09, 2012 | 2012 | Sendx Medical Inc | The product can cause the analyzer to report values outside the published performance specifications for pH, pCO2, cNa+, cK+, cCa2+ and cCl-. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | SenDx Medical sent Customers notification letters dated January 30th, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. We have recently become aware of a possible condition in a limited number of ABLBO FLEX analyzers configured with a specific Control PCe lerinted Circuit Board), which can cause the analyzer to report values outside the published performance specifications for pH, pCOz, cNa*, cK*, cCa2* and cCl-.Customers were instructed to immediately implement Alternative 1 or 2 until the signal data PCB is replaced. Alterative I:Stop the use of all affected analyzers until corrective actions can be performed on these analyzers.Alternative 2:Verify proper function of the affected analyzers immediately following everypatient sample measurement by measuring a manual QC sample.Customers were asked to complete the fax form acknowledging receipt of the notice and provided a contact for Technical Support at 1-800-736-0600 opt.4. |  | 863 | 819 analyzers and 44 signal data PCBs | One US consignee in state of OH. Other consignees are international such as Australia, France, Netherlands, Denmark, China, United Kingdom, Turkey, Switzerland, Japan, India, Poland, Spain, and Germany. Distribution list submitted with supporting doc.s | Terminated | July 13, 2012 | 95 | 863 | Not_Computer
Z-1383-2012 | 61483 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Luminos dRF solid state x-ray imager Solid state x-ray imager (flat panel/dig. imager) | Siemens Axiom Luminos dRF solid state xray imager | 2 | April 09, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens became aware of a potential issue with the AXIOM Luminos dRF. Firm became aware of a potential issue with torque wrenches - during a regular scheduled calibration of the wrench, it was discovered that the wrench was out of the specification. The wrench was over-torquing approximately 21% for low-end readings (50 ftlbs) and 9% for high-end readings (250 | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens initiated a recall/correction on February 16, 2012 by issuing Update Instruction XP056/11/S to the affected customers. All screws on the unit support on AXIOM Luminos dRF units will be replaced by a Siemens service engineer.For information regarding this recall call 610-219-4834. | Contact the recalling firm for information | 2 | 2 | Nationwide Distribution including California | Open |  |  | 2 | Not_Computer
Z-1392-2012 | 61395 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | BioCare Medical Vina Green Chromogen Kit; 25 mL, 100ml IVD. Kits include: Vina Green Chromogen 1. 0 mL; Vina Green Buffer, 25 mL and Vina Green Dropper Bottle. BioCare Medical, Concord, CA 94520.Vina Green is intended for both IHC and ISH applications including HPV, CMV, EBV, Kappa, TTF-1, Ki-67 and other targets/antigens such as blood and lymphatic vessels and basal and myoepithelial cells. | Vina Green(tm) Chromogen Kit | 2 | April 10, 2012 | 2012 | Biocare Medical Llc | Reports of sporadic failure of specific lots of Vina Green Chromogen Kit, either through crystal formation that obscures positive staining, color not right, or a complete fading of staining by the chromogen. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | "Medical Device Recall Notification" letters were sent on March 13, 2012 to all direct account. Letters informed customers of the reason for recall and the specific lots and catalog numbers involved, along with the reason for recall, how to identify the recalled product, and the actions to be taken. Customers were to immediately discontinue use of the affected product lot(s). They should dispose of the product according to national, state, and local laws and regulations. The enclosed Return Response Form should be completed and returned. Customers will be credited upon receipt of the Response Form. |  | 74 | 74 kits | Worldwide Distribution -- USA, including the states of WI, KY, CA, GA, IN, MA, IL, MD, NC, LA. AZ, TX, NC, NJ, MO, NY, PA, OH, DC, MI, WA, and CO, and the countries of Canada, Sweden, Germany, Italy, and Switzerland. | Terminated | September 20, 2012 | 163 | 74 | Not_Computer
Z-1387-2012 | 61249 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Elecsys Troponin I ImmunoassayFor the in-vitro quantitative determination of cardiac troponin I in human serum and plasma. Elecsys Troponin I is intended to aid in the diagnosis of myocardial infarction. | Troponin I Immunoassay | 1 | April 11, 2012 | 2012 | Roche Diagnostics Operations, Inc. | An investigation has confirmed a considerably reduced recovery of Li-heparin plasma samples for Troponin I lot 163176 and Troponin I STAT lot 163177. In the worst case, Li-Heparin sample recovery may be as low as 50% of the serum recovery. Correct serum recovery has been confirmed for the affected lots.The issue is caused by the raw material lot (Poly-L-Lysin) instability. Poly-L-Lysin is the | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Roche Diagnostics Corporation, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 12, 2012 to all customers that received the affected lots of Elecsys Troponin I or Elecsys Troponin I STAT. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the affected product; discard the affected product from their inventory according to their site's local regulations; if their facility has distributed the affected product to other sites, ensure that this letter is provided to those sites; complete and return the attached fax form via fax to 1-888-912-8457 and file this letter for future references.If you have any questions about the information contained in this letter, please contact Roche Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-428-2336. |  | 739 | 739 kits | Worldwide distribution.Argentina, Kenya, Austria, South Korea, Belgium, Kuwait, Bahrain, Morocco, Brazil, Mexico, Switzerland, Nigeria, Chile, New Zealand, Columbia, Philippines, Germany, Pakistan, Algeria, Poland, Egypt, Qatar, Spain, Tunisia, France, Turkey, United Kingdom, Taiwan, India, Ukraine, Iran, Venezuela, Italy, South Africa, Jordan, and Armenia. | Terminated | February 06, 2013 | 301 | 2027 | Not_Computer
Z-1394-2012 | 61505 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, IMAGER II Angiographic Catheter, Sterilized with ethylene oxide gas. Made in Ireland, Manufactured for Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537.The following product is being recalled: "Material Number/(OUTER CARTON LABEL OF 5pk)"/"Material Number(INNER-POUCH LABEL FOR EACH SINGLE CATHETER)"/ Material Description/ Catalog Number.M001314001M001314000IMAGER II/5/ST/65/038 BX531-400M001314011M001314010IMAGER II/5/ST/100/038 BX531-401M001314021M001314020IMAGER II/5/ST/65/035 BX 531-402M001314031M001314030IMAGER II/5/ST/100/035 BX 531-403M001314041M001314040IMAGER II/5/BARNH/100/038 BX531-404M001314051M001314050IMAGER II/5/BERN/100/038 BX531-405M001314061M001314060IMAGER II/5/BERN/100/035 BX 531-406M001314071M001314070IMAGER II/5/BERN/40/035 BX 531-407M001314081M001314080IMAGER II/5/BERN/40/038 BX531-408M001314091M001314090IMAGER II/5/BERN/65/035 BX 531-409M001314101M001314100IMAGER II/5/BERN/65/038 BX531-410M001314111M001314110IMAGER II/5/WEIN/100/038 BX531-411M001314121M001314120IMAGER II/5/WEIN/100/035 BX 531-412M001314131M001314130IMAGER II/5/H1/65/038 BX531-413M001314141M001314140IMAGER II/5/H1/100/038 BX531-414M001314151M001314150IMAGER II/5/H1/100/035 BX 531-415M001314161M001314160IMAGER II/5/H3/100/038 BX531-416M001314171M001314170IMAGER II/5/H3/100/035 BX 531-417M001314181M001314180IMAGER II/5/H1H/100/038 BX531-418M001314191M001314190IMAGER II/5/H1H/100/035 BX 531-419M001314201M001314200IMAGER II/5/H3H/100/035 BX 531-420M001314211M001314210IMAGER II/5/JB1/65/038 BX531-421M001314221M001314220IMAGER II/5/JB1/100/038 BX531-422M001314231M001314230IMAGER II/5/JB1/100/035 BX 531-423M001314241M001314240IMAGER II/5/JB2/100/038 BX531-424M001314251M001314250IMAGER II/5/JB2/100/035 BX 531-425M001314261M001314260IMAGER II/5/JB3/100/038 BX531-426M001314271M001314270IMAGER II/5/JB3/100/035 BX 531-427M001314281M001314280IMAGER II/5/MAN/100/038 BX531-428M001314291M001314290IMAGER II/5/MAN/100/035 BX 531-429M001314301M001314300IMAGER II/5/SIM1/65/038 BX531-430M001314311M001314310IMAGER II/5/SIM1/65/035 BX 531-431M001314321M001314320IMAGER II/5/SIM1/100/038 BX531-432M001314331M001314330IMAGER II/5/SIM1/100/035 BX 531-433M001314341M001314340IMAGER II/5/SIM2/100/038 BX531-434M001314351M001314350IMAGER II/5/SIM2/100/035 BX 531-435M001314361M001314360IMAGER II/5/SIM3/100/038 BX531-436M001314371M001314370IMAGER II/5/SIM3/100/035 BX 531-437M001314381M001314380IMAGER II/5/SIM4/100/038 BX531-438M001314391M001314390IMAGER II/5/HN1/100/038 BX531-439M001314401M001314400IMAGER II/5/HN1/100/035 BX 531-440M001314411M001314410IMAGER II/5/HN2/100/038 BX531-441M001314421M001314420IMAGER II/5/HN2/100/035 BX 531-442M001314431M001314430IMAGER II/5/HN3/100/038 BX531-443M001314441M001314440IMAGER II/5/HN3/100/035 BX 531-444M001314451M001314450IMAGER II/5/HN4/100/038 BX531-445M001314461M001314460IMAGER II/5/HN4/100/035 BX 531-446M001314471M001314470IMAGER II/5/HN5/100/038 BX531-447M001314481M001314480IMAGER II/5/HN5/100/035 BX 531-448M001314491M001314490IMAGER II/5/CK/100/035 BX 531-449M001314501M001314500IMAGER II/5/GEN/100/035 BX 531-450M001314511M001314510IMAGER II/5/GEN+2 SIDEHOLES/100/035 BX 531-451M001314521M001314520IMAGER II/5/CHGB/65/038 BX531-452M001314531M001314530IMAGER II/5/CHGB/65/035 BX 531-453M001314541M001314540IMAGER II/5/CHGC/65/038 BX531-454M001314551M001314550IMAGER II/5/CHGC/65/035 BX 531-455M001314561M001314560IMAGER II/5/CHG2.5/65/038 BX531-456M001314571M001314570IMAGER II/5/CHG2.5/65/035 BX 531-457M001314581M001314580IMAGER II/5/CONTRA 2/65/038 BX531-458M001314591M001314590IMAGER II/5/CONTRA 2/65/035 BX 531-459M001314611M001314610IMAGER II/5/DUCK/65/035 BX 531-4 | Boston Scientific, IMAGER II Angiographic Catheter, | 2 | April 11, 2012 | 2012 | Boston Scientific Corporation | Boston Scientific Corporation is initiating a Medical Device Recall of certain lots/batches of the Imager" II Catheter (Urology) and the Imager" II Angiographic Catheter . An investigationhas found that, during shipping, some units in the 5-pack box may experience a physical interaction between the carton and the area onthe pouch where the hub is located. This interaction was found to cause pi | DESIGN: Packaging Design/Selection | Boston Scientific sent an " Urgent Medical Device Recall" letter dated March 14. 2012. the letter was addressed to "Dear Risk/Materials Manager. The letter described the problem and the product involved in the recall. Advised consignees to cease the distribution of any remaining unused product affected. Advised distributors that the recall notification should be forwarded to their customers. Requested consignees to complete and return the Reply Verification tracking Form enclosed by fax to 508-683-5578.For questions regarding this recall 763-494-1700.Recall Expanded 4/24/12:Imager II Angiographic customers:Eight total number of customers not notified initially (3 USA and 5 APA customers were sent the same initial letter with a a revised send date of April 18, 2012, 329 total number of US Customers notified initially with an incomplete product list were sent a revised notification letter. Urology: Imager II Torque customers:1 US Customer was not notified initially. This customer was sent the same initial letter with a revised send date of April 18, 2012.Fifteen US Customers notified with incomplete product list. These customers were sent a revised notification letter. | Contact the recalling firm for information | 406450 | 406,450 (416,555 USA, 389,895 OUS) | Worldwide Distribution - USA (nationwide) and the countries of Arab Emir, Azerbaijan, Austria, Australia, Azerbaijan, Belgium, Bulgaria, Bahrain, Bermuda, Brazil, White Russia, Canada, Switzerland,Chile, China, Colombia, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Estonia, Egypt, Spain, Finland,France, Great Britain, Greece, Guatemala, Hong Kong, Honduras, Croatia, Hungary, Ireland, Israel, Italy,Jordan, South Korea, Kuwait, Lebanon, Lithuania, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway,New Zealand, Oman, Panama, Peru, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Fed., SaudiArabia, Sweden, Singapore, Slovakia, El Salvador, Syria, Thailand, Tunisia, Turkey, Trinidad, Tobago, Taiwan,Ukraine, Vietnam, South Africa, Dutch Antilles, Sri Lanka, Libya, and Yemen | Open |  |  | 417790 | Not_Computer
Z-1395-2012 | 61408 | device, angioplasty, laser, coronary | LPC |  |  | Cardiovascular |  | 12Fr SLS II Laser Sheath. The laser sheath is intended for use as an adjunct to conventional lead extraction tools in patients suitable for transvenous removal of chronically implanted pacing or defibrillator leads constructed with silicone or polyurethane outer insulation. | 12Fr SLS II Laser Sheath | 2 | April 11, 2012 | 2012 | Spectranetics Corp. | Isolated lots of the 12Fr SLS II Laser Sheath can have a split in the outer jacket. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spectranetics sent Recall Notices to all customers on March 14, 2012. The product issue was described and requested actions were included. Questions were directed to Customer Service at 1-800-231-0978 (US). | Contact the recalling firm for information | 615 | 615 | Worldwide Distribution -- USA , and the countries of Australia, Belgium, Canada, Czech Republic, France, Germany, Italy, Ireland, Japan, Netherlands, Russia, Sweden, Switzerland, United Kingdom and Venezuela. | Terminated | May 02, 2012 | 21 | 615 | Not_Computer
Z-1396-2012 | 61434 | container, empty, for collection & processing of blood & blood components | KSR | Empty container for the collection and processing of blood and blood components. | Hematology | Hematology | 510(k) | DeRoyal (R) Umbilicup, REF 72-8000, Rx Only, STERILE EO, Manufactured by DeRoyal, Powell, TN 37849Product Usage: Collection of cord blood | Umbilicup | 2 | April 12, 2012 | 2012 | DeRoyal Industries Inc | Device's needle may become dislodged from the cup during shipment or during use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DeRoyal Industries sent Recall communication by letter to all affected consignees of the Non-Sterile Bulk and Single Sterile Units on January 20, 2012 and of the Kits and Trays on Febraury 6, 2012. The letters identified the affected products, problem and actions to be followed. Customers were instrructed to complete the enclosed Notice of Destruction form and return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available. Distributors were instructed to forward this recall to any of their end users that may have affected product.If you have questions or need assistance with the recall, please call 865.362.1020 or Robby Lockett at 865.362.1034. |  | 2472 | 2472 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, Dominican Republic, South Africa, and the United Arab Emirates. | Terminated | July 11, 2013 | 455 | 9434 | Not_Computer
Z-1420-2012 | 52120 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa EVA TPN bags for administering parenteral nutrition solutions. Sold under three brand names: Exacta Mix EVA Bags, Exacta Mix Dual-Chamber EVA Halobag, and Baxa Calibration Bags. | Baxa EVA TPN bags | 2 | April 13, 2012 | 2012 | Baxa Corporation | Spike port flor EVA TPN Bags contains low levels of DEHP which could leach out. | MISBRANDING: Labeling False and Misleading | Baxa sent a Customer Notice letter dated May 8, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Current Baxa product is mislabeled as non-DEHP but this does not present a risk to patients. Baxa is committed to providing products that are non-DEHP. We are working with our supplier to replace the spike port with a non-DEHP as soon as possible. In the meantime, product labeling will be updated to reflect this finding.For further questions please call (303) 690-4204. |  | 50 | 347,494 cases/ 50 bags each | Worldwide Distribution -- USA (nationwide) | Terminated | April 13, 2012 | 0 | 347494 | Not_Computer
Z-1401-2012 | 61446 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Non-DEHP Microbore Administration Set with Non-Vented Spike Packaged AssemblyThe Curlin Ambulatory Infusion Pump System with disposable PVC Administration Tubing Sets and user accessories, provide a means for the volumetric delivery of fluids used in Parenteral, Enteral, Epidural, Subcutaneous and Intravenous applications. | Curlin Infusion Administration Set | 2 | April 13, 2012 | 2012 | MOOG Medical Devices Group | Ambulatory administration sets used with the Curlin Ambulatory Pumps including the 4000 CMS Ambulatory Pump, 6000 CMS Ambulatory Pump and the Pain Smart IOD Ambulatory Pump may leak during use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Moog Medical Device Group sent a Medical Device Recall Notification letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product from their inventory, contact the distributor or Moog Customer Service at 1-800-970-2337, prompt #7, then return product to distributor or Moog.For additional questions regarding this recall call Moog's Customer Service Department at (800) 970-2337. |  | 11628 | 11,628 | Worldwide Distribution - USA including AZ,,, CA,,, FL,,, GA,,, KS,,, IL,,, MA,,, MD,,, MI,,, OH,,, PA,,, TX,,, VA,,, WI,,, and the countries of Germany,,, Italy,,, Morocco,,, New Zealand,,, and the United Arab Emirates,,, | Terminated | July 26, 2012 | 104 | 43176 | Not_Computer
Z-1432-2012 | 53049 | enzymatic method, ammonia | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Ammonia (AMM) ReagentProduct Usage:AMM reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems Ammonia Calibrators, is intended for the quantitative determination of Ammonia concentration in human plasma. | Beckman Coulter Synchron Ammonia (AMM) Reagent | 2 | April 16, 2012 | 2012 | Beckman Coulter Inc. | The Ammonia Reagent lot M808278 may fail calibration or cause a low shift in Quality Control recovery on some SYNCHRON Systems. |  | Beckman Coulter mailed a Product Corrective Action (PCA) Letter on April 13, 2009 dated April 10, 2009 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers to discontinue use and discard appropriately the lot number(s) of the affected product and contact their Customer Service Representative at (800) 526-3821 (US), or (800)463-7828 for reagent replacement. Customers were instructed to complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain records.Customers were told to call the Customer Support Center (Hotline) at (800) 854-3633 with any questions regarding the PCA letter. |  | 3388 | 3,388 kits | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 18, 2012 | 2 | 3388 | Not_Computer
Z-1430-2012 | 60661 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Double foot switch, type KF 2 1S/1 S-MED-AP, manufactured by the company Seute Schaltgerate. GmbH & Co. KG for Ziehm Imaging GmbH Germany, and used with Ziehm Imaging GmbH Mobile C-arm as a component thereof. Two versions were made: 2 pedal version KF 2 1S/1 S-MED-AP foot switch and 4 pedal foot switch version KF 2 1S/1 S-MED-AP This type of foot switch includes a pair of the dual foot switches assembled on a single metal bracket. The following labels may appear on the foot switch:Labeled in part: "Fluoro- Store***ZIN Part no. # 58371***ZIR Part no. # 10-2985*** Left Pedal: Initiate Radiation***Right Pedal: Store Image***"Labeled in part: "Fluoro - Cine***ZIN Part no. # 58415***ZIR Part no. # 10-2987*** Left Pedal: Initiate Radiation***Right Pedal: Switch into Cine mode***"Labeled in part: "Fluoro - Mag***ZIN Part no. # 58413***ZIR Part no. # 10-2989*** Left Pedal: Initiate Radiation***Right Pedal: Switch through magnification***"Labeled in part: "Fluoro -Snap***ZIN Part no. # 58417***ZIR Part no. # 10-2991*** Left Pedal: Initiate Radiation ***Right Pedal: Initiate Snapshot***"Labeled in part: "DSA - CINE***ZIN Part no. # 58419***ZIR Part no. # 10-2993*** Left Pedal: Switch into DSA mode ***Right Pedal: Switch into Cine mode***"Labeled in part: "Cine - Store***ZIN Part no. # 58421***ZIR Part no. # 10-2995*** Left Pedal: Switch into Cine mode***Right Pedal: Store Image***"Labeled in part: "Cine - LONG:Endo ON/OFF & SHORT:Freeze*** ZIN Part no. # 58423***ZIR Part no. # 10-2997*** Left Pedal: Switch into Cine mode***Right Pedal: Long: Switch Endo On/Off Short: Freeze the endoscopic image on live screen***"Labeled in part: "Cine- Endo ON /OFF***ZIN Part no. # 58425***ZIR Part no. # 10-2999*** Left Pedal: Switch into Cine mode***Right Pedal: Switch Endo On / Off***"Labeled in part: "Fluoro - Swap***ZIN Part no. # 58427***ZIR Part no. # 10-3031*** Left Pedal: Initiate Radiation***Right Pedal: Swaps the image between the live and reference screens***"Labeled in part: "DSA / MSA/RSA - OFF***ZIN Part no. #58429***ZIR Part no. #10-3033***Left Pedal: Switch DSA/MSA/RSA On***Right Pedal: Switch DSA/MSA/RSA Off***"The KF 2 1S/1S-MED- AP foot switch is a component of the Mobile C-arm and is used by the physician to activate Fluoroscopic X-Ray for imaging by pressing the foot switch. | Double Foot Switch | 2 | April 16, 2012 | 2012 | Ziehm Imaging Inc | On 04/29/2011, Ziehm Imaging, Inc., Orlando, FL recalled the double foot switch, type KF 2 1S/1S-MED-AP, a component used with Ziehm Imaging GmbH Mobile C-arm due to the potential of unintentional continued activation of fluoroscopic X-ray as the foot switch may not become inactive when the pedal is released. | DESIGN: Component Design/Selection | Ziehm Imaging, Inc., Orlando, FL sent an Urgent Recall Notice letter dated April 11, 2011, that included a return receipt request, envelope, and a reply questionnaire. to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. They also called consignees. After contact, Ziehm Imaging, Inc. shipped a replacement foot switch with instructions for the customer/end user on how to install the replacement foot switch. If no signed confirmation was received, Ziehm Imaging, Inc. visited the consignee and replaced the foot switch. Customers were asked to fax the completed response form to 866-839-7893. | Contact the recalling firm for information | 49 | 49 units | Nationwide Distribution including Florida, South Carolina, Georgia, Alabama, Virginia, Ohio, Massachusetts, New York, Pennsylvania, Kentucky, Connecticut, Rhode Island, Indiana, Illinois, Nebraska, Minnesota, Arizona, California, Oregon, and Washington. | Terminated | April 18, 2012 | 2 | 49 | Not_Computer
Z-1429-2012 | 61455 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Folding Shower Bench;Features a nonskid surface and handholds for added security. holes allow excess water to drain away easily. Sturdy legs have non-skid rubber feet. Supports up to 200 lbs.North Coast Medical Inc.8100 Camino ArroyoGilroy, CA 95020Comfort and safety in shower | Norco Folding Shower Bench | 2 | April 16, 2012 | 2012 | North Coast Medical Inc | Due to a consumer complaint, North Coast Medical determined that due to a design flaw, there was not enough plastic support material to retain bracket bolts when subjected to a force greater than 150 lbs. | DESIGN: Device Design | North Coast Medical, Inc. sent a Recall Notification Letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and discard the affected product. Complete the enclosed Form and return the form even if they no longer have the bath bench. Return the form by fax to 408-775-1430. Customers should contact the Quality Department if they have any questions, concerns or need any assistance following the instructions.For questions regarding this recall call 800-821-9319, ext 121. |  | 370 | 370 benches | Nationwide Distribution | Terminated | July 20, 2012 | 95 | 370 | Not_Computer
Z-1428-2012 | 61352 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | PantaÂŽ Arthrodesis Nail Support device, 510 (k) #K050882. One (1) unit of product is kitted within an instrumentation tray. The product is available in six (6) sizes. All nails are color-coded for easy size identification. Bony fixation is achieved using two (2) tibial screws, two (2) calcaneal screws and one (optional) talar screw. The nail, cross-locking screws and end cap are manufactured from a titanium alloy: TI-6AI-4V ELI ISO 5832-3, ASTM F136. The PANTA Arthrodesis Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. The PantaÂŽ Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. Examples include: " Post-traumatic and degenerative arthritis involving both ankle and subtalar joints. " Rheumatoid arthritis. " Revision of failed ankle arthrodesis with subtalar involvement or with insufficient talar body. " Revision of failed total ankle arthroplasty with subtalar intrusion. " Talar deficiency conditions (requiring a tibiocalcaneal arthrodesis). " Avascular necrosis of the talus. " Neuroarthropathy or neuropathic ankle deformity. " Severe deformity as a result of talipes equinovarus, cerebral vascular accident, paralysis or other neuromuscular disease. " Severe pilon fractures with trauma to the subtalar joint. | Panta Arthrodesis Nail | 2 | April 16, 2012 | 2012 | Integra LifeSciences Corp. | The firm received multiple complaints from their user customers (surgeons) who reported encountering difficulty inserting a calcaneal or tibial screw through the holes of the PANTA nail. |  | Integra issued a Field Safety Notice with an attached Acknowledgement and return Form to all Integra Sales Representatives and Customers via Fed Ex and E-mail on May 19, 2011. The Field Safety Notice informs the customers of the firm's receipt of information concerning several reported incidents that lead to the misalignment of the calcaneal of tibial when used with the Panta Arthrodesis Nail Support Device. The notice goes stated that Integra determined that the misalignment of these screws may be attributed to improper surgical technique and to that end, the Field Safety Notice is being issued in order to reinforce the importance of following the recommended guidelines our lined in the product's Surgical Technique Manual (a copy of the referenced guidelines are provided with the Field Safety Notice (FSN). The FSN goes on to direct the reader/customer to applicable section sections of the manual to which particular attention should be directed in order to minimize the potential for misalignment of the screws. The customers are asked to distribute the information provided to the surgical teams at their various institutions. Moreover, they are also asked to sign and return the enclosed Field Safety Notice Acknowledgement and Return Form so as to document their receipt of the notification and the material provided. Should the customer(s) have questions, they are directed to contact Sean Luland, Director of Marketing-Lower Extremity Reconstruction at Integra LifeScience at: 513-533-7929, for additional information.For questions regarding this recall call 609-275-0500. | Contact the recalling firm for information | 64 | 64 units were distributed | Nationwide Distribution including MD, IL, NU, CA, GA, CO, OH, SD, AR, FL, TX, MA, MO, NC, OK, KY, NJ, UT, VA, MN, DC, PA, WA, CT, AZ, IL, MI, AL, NE, TN, NM, NV, LA, OR, DE, WI, ND, NH, MS, SC, IN and Puerto Rico. | Terminated | February 19, 2013 | 309 | 64 | Not_Computer
Z-1424-2012 | 61304 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Attune Impaction Handle Warsaw, IN 46582Packaging: Product is packed within a polyethylene bag, with protection added as needed for sharp and/or fragile points.Product Usage: The ATTUNE Impaction handle is intended to act as a removable interface for several of the instruments that may be used throughout a procedure. The Attune Impaction Handle is used to impact, insert, and extract various Attune Instruments. The handle interfaces with the Keel Punch, Tibial Tower, Evaluation Bullet, Fixed Bearing Tibial Impactor, Rotating Platform Tibial Impactor and the Femoral Impactor. | DePuy Attune Impaction Handle Warsaw | 2 | April 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Recall Notice for nine lots of the Attune" Impaction Handle due to the potential for the handle trigger to break during surgery. The Attune Impactor is provided as a part of a set of tools and is designed specifically for the installation of the Attune Knee | DESIGN: Device Design | DePuy Orthopaedics, Inc. sent an URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE letter dated March 16, 2012 to all consignees affected which extended to the Depuy Distributor level, surgeon and hospital level. The affected DePuy Distributors were notified via email on March 16, 2012. The sales representatives notify hospitals and surgeons by mail or in person with written communication. The letter identified the affected product, reason for recall notice, clinical implications and actions to be taken. Consignees were instructed to complete and return the Response cards to DePuy by fax to 574-372-7567 or email to:kseppa@its.jnj.com. When the new impaction handles are available, the DePuy sales representative will be responsible for providing the new handles, removing and returning the affected handles. For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact DePuy's Scientific Information Office at 1-888-554-2482 (M-F; 8 am - 5 pm EST.) For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - 5 pm EST.). |  | 136 | 136 impaction handles | USA Nationwide Distribution including the states of: MA, MD, IL, IA, MN, WA, CA, FL, NC, VA, and IN. | Terminated | January 09, 2014 | 633 | 136 | Not_Computer
Z-1427-2012 | 61538 | endoscope and/or accessories | KOG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hudson RCI, BiteGard Oral Bite Block, Single Patient Use, Rx Only, Teleflex Medical, Research Triangle Park, NC. The Teleflex Medical BiteGard Molar Bite Black is a disposable single patient use device that fits between the patients molars to prevent the patient from closing their mouth entirely and from obstructing the Laryngeal Mask Airway (LMA) tube, in the case where one is used. | Hudson RCI | 2 | April 16, 2012 | 2012 | Teleflex Medical | Teleflex Medical received complaints that the bite block can become separated from the handle. If separation occurs, medical intervention may be required to remove the bite block from the patient's trachea or esophagus/GI tract. | PRODUCTION CONTROLS: Process Control | Consignees received "Urgent Medical Device Recall Notification" letters on or about 04/03/2011. The letters described the issue and provided recommended actions regarding remaining product and returns. Questions were directed to Customer Service at (866) 804-9881 (US). | Contact the recalling firm for information | 333724 | 333,724 units | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Singapore, China, Denmark, Netherlands, India, Australia, Japan and the United Kingdom. | Open |  |  | 333724 | Not_Computer
Z-1425-2012 | 52748 | diazo colorimetry, bilirubin | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Total Bilirubin, 6L45-20 &amp;amp;amp; 6L45-40.Used for the quantitative analysis of total bilirubin in human serum or plasma of adults and neonates. | Clinical Chemistry Total Bilirubin | 2 | April 16, 2012 | 2012 | Abbott Laboratories Inc. | A neonatal sample that was analyzed, produced results with values that were incorrectly less than the expected limit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Abbott Laboratories plans to release a Product Correction letter to customers instructing customers to follow the steps provided in order to revise their Total Bilirubin assay parameters. The letter explained the reason for the recall, along with the patient impact. Customers were to keep the letter as part of their laboratory records and also forward a copy to any other laboratories that they may have forwarded Total Bilirubin kits. U.S. Customers should contact Customer Service at 1-877-4ABBOTT, if they have questions. | Contact the recalling firm for information | 11282 | 11,282 | Worldwide Distribution -- USA, Guatemala, Chile, Brazil, Canada, Germany, Hong Kong, Thailand, Puerto Rico, Australia, New Zealand, Trinidad & Tobago, Bermuda, Curacao, Cayman Island, and Jamaica. | Open |  |  | 11282 | Not_Computer
Z-1431-2012 | 61547 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Zuma Mobility System Models: Zuma Mobility Assist 412045-01 and Zuma Mobility Trainer 412046-01.Product Usage: Human Non-AC-Powered Patient Lift | Zuma Mobility System | 2 | April 16, 2012 | 2012 | KCI USA, Inc. | The Zuma Mobility Assist and Trainer have the potential to bind during lowering operation. | DESIGN: Device Design | KCI sent an Urgent - Voluntary Medical Device Correction notification letters dated April 6, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. KCI will contact affected customers to schedule a time to upgrade the device. For questions contact KCL's Customer Technical Service Center at 1-800-275-4524. Select option 3 followed by option 2. | No consumer action necessary | 4 | 4 units | USA Nationwide Distribution including the states of ID and FL | Terminated | June 19, 2012 | 64 | 4 | Not_Computer
Z-1426-2012 | 61342 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PFC Steinman Pin/Drill Pack Sterile (WPC 667-2012) Drill Bit and Steinman Pin Packet (1 unit)The PFC Steinman Pin/Drill Pack Sterile is part of the PFC Modular Knee Instrument system. The single pack (1 box) contains 2 pieces of 1/8 Target Drill Bit product code 513-4193-01 and 4 pieces of Steinman Pin product code 086-9117-00. The Drill Bits are 0.125 in diameter and 5 long. The Steinman Pins in this pack are 0.125 in diameter and 3 long. The pack is provided sterile to the customer and is intended for one time use. The contents of the pack provide stable fixation of the femoral or tibial cutting blocks to the bone while a cut is made. | PFC Steinman Pin/Drill Pack Sterile (WPC 6672012) | 2 | April 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | Sterility of the pins is at risk because product packaging, which is designed to include 2 pouches (inner pouch and outer pouch) around the pin mounting card only, contains a single pouch, which could compromise product sterility. | PRODUCTION CONTROLS: Packaging Process Control | The firm, DePuy Orthopedics Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated March 12, 2012 to its customers via mail or in person with written communication. The notice described the product, problem and actions to be taken.The customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-01 on the outside of the box.The customers were also instructed to complete and return the Reconciliation Form to their Depuy Sales Representative or fax to 574-372-7567. Follow-ups for the Reconciliation Form for non-responders will be conducted with the Distributors until the form is returned. For Hospitals that are now out of business or refuse to sign the response card, the Distributor will indicate this on the Reconciliation form and return.For questions about the recall information provided, please contact Manager of Customer Quality, at 574-372-7333 (M-F; 8am - 5pm EST). For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact Depuy's Scientific Information Office at 1-888-554-2482 (M-F; 8am - 5pm EST). | Contact the recalling firm for information | 40 | 40 of each lot, 120 total | Nationwide distribution: USA including states of: FL, LA, TX, OR, CA, AZ, KY, GA, MS, and PA. | Terminated | March 06, 2013 | 324 | 40 | Not_Computer
Z-1421-2012 | 61403 | tape and bandage, adhesive | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General & Plastic Surgery | 510(K) Exempt | Sterile wound closure strips, packaged as follows:DeRoyal EPISEAL, Wound Closure Strips, REF 46-202, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-203, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-204, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-206, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-212, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-213, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-214, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-215, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849DeRoyal EPISEAL, Wound Closure Strips, REF 46-216, STERILE R, Distributed by: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849; Manufacturer: DeRoyal Industries, 200 DeBusk Lane, Powell TN 37849 | Episeal wound closure strips | 3 | April 17, 2012 | 2012 | Deroyal Industries, Inc. Lafollette | Packages of Episeal wound closure strips and surgical kits containing them were recalled because the adhesive backing on Episeal wound closure strips may not release from the backing rendering the product unusable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DeRoyal Industries sent a Recall of Episeal Wound Closure Strip letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. Please identify and quarantine any affected inventory using the product and lot numbers listed above.2. Render affected product unusable, destroy and discard according to your facilitys guidelines.3. If you are a distributor, forward this recall to any of your end users that may have affected product or provide us with the list and we will contact them.4. Please complete the enclosed Notice of Destruction form, indicating the affected products you found in your inventory and destroyed. PLEASE COMPLETE AND RETURN THIS FORM EVEN IF YOU NO LONGER HAVE THE PRODUCT IN INVENTORY.Return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available.If you have questions or need assistance with the recall, please contact DeRoyal Customer Support at 1-800-251-9864.We apologize for inconvenience this may cause you. |  | 126720 | 126,720 units | Worldwide Distribution -- USA (nationwide) | Terminated | July 10, 2013 | 449 | 205661 | Not_Computer
Z-1433-2012 | 61465 | photometric method, magnesium | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista(R) MG Flex(R) Reagent CartridgeThe Dimension Vista(R) MG method is an in vitro diagnostic test for the quantitative measurement of magnesium in human serum, plasma and urine on the Dimension Vista(R) System | Dimension Vista(R) MG Flex(R) Reagent Cartridge | 2 | April 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a negative shift of approximately 0.2mg/dL [0.08 mmol/L] in patient and QC samples when using Vista Magnesium | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated February 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter explained the correction needed and instructed customers to forward the letter to anyone they may have distributed the affected product.For questions or concerns call 800-441--9250.For questions regarding this recall call 301-631-6299. |  | 9178 | 9,178 | Nationwide Distribution including Puerto Rico and within the US to Washington, DC and the following States: AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Open |  |  | 9178 | Not_Computer
Z-1437-2012 | 61622 | set, perfusion, kidney, disposable | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ViaSpan Cold Storage Solution 1000mL Bag distributed under the Barr/Duramed Pharmaceuticals label (10 x 1000mL bags).Disposable kidney perfusion set. | ViaSpan Cold Storage Solution 1000mL Bag | 2 | April 18, 2012 | 2012 | Teva Pharmaceuticals USA, Inc. | Lack of assurance of sterility | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Teva Pharmaceuticals USA sent an "URGENT MEDICAL DEVICE RECALL-RETAIL LEVEL" letter dated March 30, 2012. The letter identifies the problem, product, and actions to be taken by the customers. The letter instructs customers to check their inventory and to discontinue distribution of the affected product. A Stock Response Form was enclosed for customers to complete and return via fax to 817-868-5362. Contact Teva Customer Services at 800-545-8800 for questions regarding this notice. | Contact the recalling firm for information | 4070 | 4,070 | Nationwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CT, DE, FL, GA, IA, IL, KS, KY, LA, MA, MD, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV and Washington, D.C. | Terminated | October 25, 2012 | 190 | 4070 | Not_Computer
Z-1434-2012 | 55440 | bromcresol green dye-binding, albumin | CIX | Albumin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Albumin BCG (AlbG) List Number 7D53-21Clinical Chemistry Albumin BCG 510(k) K981758, Product Code CIXThe Albumin BCG assay is used for the quantitation of albumin in human serum or plasma | Abbott Clinical Chemistry Albumin BCG (AlbG) | 2 | April 18, 2012 | 2012 | Abbott Laboratories Inc. | Observed particulate matter in some reagent cartridges. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Abbott Laboratories sent a Product Recall notification letter dated March 10, 2010, to all affected customers. The notification letter contained an explanation, patient impact, neccassary actions, and contact information.Customers were instructed to determine if they currently have any of the affected product in their inventory. Discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures. Completed the enclosed Customer Reply Form. If customers forwarded any of the affected product they should provide a copy of the notification letter to their customers.For any questions customers should call Customer Service at 1-877-422-2688.For questions regarding this recall call 626-440-0705. | Contact the recalling firm for information | 2718 | 2718 (U.S. 652) | US nationwide and internationally to Germany, Hong Kong, Australia, New Zealand, Trinidad/Tobago, and Jamaica. No dsitribution of Canada or Mexica. | Terminated | April 18, 2012 | 0 | 2718 | Not_Computer
Z-1435-2012 | 61620 | implant, eye sphere | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | MEDPOR PLUS SST EZ 22mm Sphere Implant, Howmedica Osteonics Corp, 15 Dart Road, Newnan, GA 30265. MEDPOR PLUS Orbital Implants are intended to restore the volume of an enucleated or eviscerated eye globe. | MEDPOR PLUS | 2 | April 18, 2012 | 2012 | Howmedica Osteonics Corp dba Stryker Craniomaxillofacial | The MEDPOR PLUS SST EZ 22mm sphere implants are not within specification of 22mm +/- 8, therefore implant is not easily released from the syringe. | PRODUCTION CONTROLS: Process Control | The sole distributor was notified by "Urgent Product Recall" letter on March 20, 2012. "Urgent Medical Device Recall" letters were sent to Stryker Customers on March 26, 2012. The letters described the product issue, hazards, and actions recommended to mitigate the risk. Customer number provided: (800) 962-6558. |  | 31 | 31 units | US Distribution, the state of Michigan. | Terminated | August 10, 2012 | 114 | 31 | Not_Computer
Z-1441-2012 | 49059 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | OMG (Observational Mechanical Gateway) Connector - A Model Number SC-9305The Precision Spinal Cord Stimulator System is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs, including unilateral or bilateral pain associated with following: failed back surgery syndrome, intractable low back pain and leg pain. The Precision System includes an implantable l6-output, multi-channel stimulator (lPO - Implantable Pulse Generator) with a rechargeable battery power source. The lPG, commonly implanted in the abdomen or buttock area, can be configured to accept one or two leads. The leads are implanted in the epidural space, adjacent to the spinal column. The stimulation delivered at the distal end of the lead is intended to mask the pain signals. | The Precision Spinal Cord Stimulator System | 2 | April 19, 2012 | 2012 | Advanced Bionics Corporation | The recall was initiated because there is a possibility that the OMG Connector-A is incorrectly packaged with an OMG Connector-M Instruction manual. | PRODUCTION CONTROLS: Labeling Mix-Ups | Algotec Limited initiated a Field Safety Notice Urgent Medical Device Correction letter dated April 22, 2008, with Boston Scientific forwarding a recall letter to one customer who received the recalled product. The letter instructed the consignee to replace the previously packaged manual with the OMG Connector-A manual that was supplied with the notification. The replacement of the manuals is to be confirmed via the Verification form.Customers were instructed to complete and fax the Verification Form to their local Boston Scientific Customer Service to Fax 00 44 1442 411732 by 06 May 2008.For questions regarding this recall call 661-949-4481. | Contact the recalling firm for information | 6 | 6 Units | Worldwide Distribution including Great Britian | Terminated | April 19, 2012 | 0 | 6 | Not_Computer
Z-1440-2012 | 49410 | stimulator, nerve, transcutaneous, for pain relief | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Synaptic 3600 and 3800 with Cranial Electrical Stimulation manufactured by Synaptic USA, Aurora, CO.Transcutaneous nerve stimulation for pain relief and other medical treatment uses. | Synaptic | 2 | April 19, 2012 | 2012 | Synaptic Corp | Device was distributed without marketing approval. | PREMARKET APPROVAL: No Marketing Application | Synaptic Corp, sent a Urgent Medical Device Recall letter dated July 14th, 2008, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.The customer was instructed to return your 3600 and or 3800 devices so that the cranial stimulator can be removed and destroyed. Please call 1-800-685-7246 to schedule pick up of your device..Complete the enclosed Confirmation Form and return it ( postage paid) to Synaptic Corp. It is important that you return the attached response confirmation form even if you do not have the (CES) module in your model.Should you have any questions about this recall please call 1-800-685-7246 ext. 101. | Contact the recalling firm for information | 59 | 59 units | Worldwide Distribution -- USA (nationwide) and the country UK. | Terminated | April 19, 2012 | 0 | 59 | Not_Computer
Z-1444-2012 | 61453 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | VERTEX MAX (R) Cannulated, VERTEX MAX (R), and VERTEX (R) Reconstruction System Adjustable Drill Stop, REF 6860460, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132Product Usage:Manual orthopedic adjustable drill stop | adjustable drill stop | 2 | April 19, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Drill stops may be bypassed when the device is used with powered drilling instruments. | DESIGN: Device Design | Medtronic notified their US Sales force on March 5, 2012, of the upcoming field action. The written "Voluntary Recall" communications were disseminated March 13 and March 14 to the US consignees. The letter identified the affected products, problem and actions to be taken. Customers are instructed to 1) disseminate this information to personnel within their facility as appropriate; 2) review work area, district sales office, and sample inventory for affected product; 3) complete the attached Risk Manager Questionnaire and return any recalled product; and 4) Sign and acknowledge the attached Sales Rep Questionnaire indicating receipt of this letter. For questions or concerns, contact your Medtronic Sales Representative or Global Quality Department at 800-876-3133, extension 3197. | Contact the recalling firm for information | 247 | 247 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Barbados, Brazil, Netherlands, Hong Kong, Japan, Korea, Columbia, Chile, Canada, India, Jamaica, Switzerland, Costa Rica, Ecuador, and Mexico. | Terminated | April 30, 2013 | 376 | 4198 | Not_Computer
Z-1438-2012 | 61089 | prosthesis, knee, femorotibial, constrained, cemented, metal/polymer | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF. CP113456, MAK OSS No Head Lock Pin, NO HEAD / WITH PLUG CO-CR-MO ALLOY/ARCOM UHMWPESTERILE, R, Biomet Orthopedics Warsaw, IN 46581.The MAK OSS No Head Metal Lock Pin is to be inserted into the anterior hole of the tibial bearing yoke to secure the OSS Axle in the correct position within the OSS Femoral component assembly. | MAK OSS No Head Lock Pin | 2 | April 19, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after receiving complaints of the MAK OSS No Head Lock Pin dislodgingfrom its intended position (post-initial surgery). An investigation was initiated and the available lots ofthe Lock Pin were evaluated and measured against print specifications. | DESIGN: Component Design/Selection | Biomet sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to continue monitoring implanted patients for disassociation of the pin component and report to Biomet. The Alert included a complete listing of device invoice history and patient information (if supplied). Contact the firm for confirmation of receipt of the notice and for questions at 574-372-3983. |  | 64 | 64 units distributed | Distributed in California and Pennsylvania. | Terminated | February 07, 2013 | 294 | 64 | Not_Computer
Z-1439-2012 | 61188 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Optilock, Locking Screw 3.5 mm x 28 mm, non sterile, REF 37728, Biomet Trauma Parsippany, NJ 97054.Cortical Locking Screw for securing a Proximal Humeral Plate to the bone in the Optilock system. | Optilock, Locking Screw 3.5mm x 28mm | 2 | April 19, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after a complaint was received that item 36728, Non-Locking Screw 3.5mm x 28mm was in a package labeled as 37728, Locking Screw 3.5mm x 28mm. | PRODUCTION CONTROLS: Labeling Mix-Ups | A March 15 2012 Urgent Medical Device Recall Notice was sent to Distributors, Operating Room Managers, and Physicians with directions to locate and remove recalled items with instructions to provide the recall notice to any sub-accounts. The letter also identified the affected product and explained the reason for the recall. OR Managers and Physicians are to confirm recall notice receipt by calling 800-348-9500, extention 3983 or 3009 and make arrangements for product return. Distributors are to complete and FAX back the response tracking form to the attention of Angi Dickson, Biomet @ 574-372-1638. Questions should be directed to 574-372-3983 or 574-371-3009 M- F, 8 AM- 5 PM, (ET). | Contact the recalling firm for information | 45 | 45 | Worldwide Distribution -- USA, including NJ, TX, VA, MI IL, KT, & Puerto Rico, and country of the Netherlands. | Terminated | August 29, 2013 | 497 | 45 | Not_Computer
Z-1446-2012 | 61637 | enzyme immunoassay, methadone | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen (MTD) Panel, Model # 94400.Type of Packaging: 25 individually pouched devices in a labeled kit box.The Triage TOX Drug Screen is a fluorescence immunoassay intended to be used with the Triage Meters for the point-of-care qualitative determination of the presence of drug and/or the major metabolites above the threshold concentrations of up to 10 distinct drug classes, including assays for acetaminophen/paracetamol, amphetamines, methamphetamines, barbiturates, benzodiazepines, cocaine, methadone, opiates, phencyclidine, THC and tricyclic antidepressants in urine. The acetaminophen/paracetamol assay will yield positive results when acetaminophen/paracetamol is ingested at or above therapeutic doses. | Triage TOX Drug Screen (MTD) Panel | 2 | April 20, 2012 | 2012 | Alere San Diego dba Biosite Innovacon Hemosense Inc | The recall was initiated because Biosite Inc. received an increase in customer calls regarding unexpected positive THC results when using the Triage TOX Drug Screen. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biosite sent a "PRODUCT ADVISORY NOTICE" letter dated March 12, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter states that customers may continue to use the product. Customers were instructed to complete and fax an enclosed verification form to Technical Services at 1-858-695-7100. Customers with any questions about the information contained in the notification were instructed to contact Biosite Inc. at (888) 246-7483. |  | 1258002 | 1,258,002 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada and European Union. | Terminated | April 23, 2012 | 3 | 1258002 | Not_Computer
Z-1447-2012 | 50072 | nadh oxidation/nad reduction, ast/sgot | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT/AEROSET Activated Aspartate AminotransferaseThe Activated Asparate Aminotransferase (Activated AST) assay is used for the quantitation of asparate aminotransferse in human serum or plasma | ARCHITECT/AEROSET Activated Aspartate Aminotransferase | 2 | April 23, 2012 | 2012 | Abbott Laboratories Inc. | This letter is to inform you that some samples may generate A-AST results of less than ( | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Laboratories sent a Product Correction letter dated August 5, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Update your assay parameters for A-AST1. Follow the steps in the appropriate attachment: - Use Attachment A for ARCHITECT cSystems - Use Attachment B for AEROSET.2. All A-AST results of less than ( |  | 1646 | 1,646 | Worldwide Distribution including Germany | Terminated | April 23, 2012 | 0 | 1646 | Not_Computer
Z-1449-2012 | 60595 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ref, S331140 Selex M2A Magnum Modular Head, 40 mm head diameter plus 3 mm neck,1 Taper, Sterile, Biomet Orthopedics, Warsaw, IN Product Usage:Intended for use in non-cemented primary and revision hip joint replacement | Selex M2A Magnum Modular Head, 40 mm | 2 | April 23, 2012 | 2012 | Biomet, Inc. | Devices with a specification for a +6mm neck length were incorrectly labeled for a +3mm neck length specification. | PRODUCTION CONTROLS: Process Control | Biomet contacted consignees with device inventory by telephone on September 13, 2011 requesting all associated product be quarantined immediately. An Urgent Medical Device Recall Notice was sent via U.S. mail March 23, 2012 to consignees with directions to immediately locate, discontinue use of the affected product and return to Biomet. The letter instructed consignees to carefully follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. The notification letter to consignees contained directions to confirm receipt of notice by calling 800-348-9500, extension 3983 or 3009. Questions concerning recall are directed to 574-372-3983 or 574371-3009 M-F 8 A.M. to 5 P.M. (ET). |  | 5 | 5 devices | USA Nationwide Distribution including the states of: MI, NC., OK and WI. | Terminated | February 21, 2013 | 304 | 5 | Not_Computer
Z-1461-2012 | 61457 | chair, with casters | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | NOCTURNAL ELITE CARE CLINER 6980(STD)/6990(X-LG) products are labeled in part: "***NOCTURNAL ELITE CARE CLINER 6980/6990***Winco***ISO 9001-2008 Certified***An infinite position recliner, independent leg rest and lay-flat all in one chair!***With lay-flat positioning and an independent leg rest, Winco's Nocturnal Elite Care Cliner provides optimal comfort for use in nocturnal dialysis, oncology, acute care and other patient care areas.***HEAT MASSAGE (Optional)***KEY FEATURES***Left side swing-arm for modified entry and ease of cleaning***User-adjustable positioning***Attendant controlled Trendelenburg foot release (on left & right side)***Left side fold down table***Right side Pivot Table***Rear push handle**Lay Flat***Hide-away foot rest***Independent leg rest***LiquiCell for added comfort***1.800.237.3377***5516 SW 1st Lane***Ocala, Florida 334474***www.wincomfg.com***". The Care Cliner is a multi position recliner with a steel frame used primarily in clinics, hospitals and treatment centers for recovery, dialysis, oncology, examination, infusion. Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala, FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"Note: Model 6988 is a version of the 6980 without the footplate. We have a single, but high volume customer, that requested the chair without the footplate. We put the model in the system so as to call out a specific Bill of Materials to leave off the footplate. Since it was a unique request, there is not a piece of literature for it. | Care Cliner | 2 | April 25, 2012 | 2012 | Winco Mfg., LLC | Winco Manufacturing LLC is recalling the Care Cliner, XL Care Cliner, Drop Arm Care Cliner, Elite Care Cliner, and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Winco Mfg. sent an "URGENT DEVICE RECALL" letter dated February 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete the Caster Recall Replacement Form and return it to the firm's Customer Service Department via fax to 352-854-9544. Contact the Customer Service Department at 800 237-3377 for questions regarding this notice. |  | 96 | 96 Casters (32 Chairs) | Worldwide Distribution-USA (nationwide) and the countries of Canada and Singapore. | Terminated | July 27, 2012 | 93 | 5132 | Not_Computer
Z-1455-2012 | 61346 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Quintex Cervical Plating System - Screws. The Quintex Cervical Plating System is intended for the treatment of cervical spinal instability. The Quintex Cervical Screws are available in three different versions: semi-constrained, constrained, and dynamic. The screws all have a head diameter of 4.0mm and vary in length from 10mm to 22mm. Rescue screws are also available with a 4.5mm head diameter and lengths of 11-19mm. | Quintex Cervical Plating System Screws | 2 | April 25, 2012 | 2012 | Aesculap Implant Systems | The firm has received complaints that the locking ring that sits in the head of the screw can come out of the screw head when engaged with the screwdriver if the screwdriver is improperly angled. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | The firm initiated this recall by sending an "Important Recall Notification" letter to consignees on January 20, 2012. This letter instructed consignees not to use the product and requested that all stock be returned to the recalling firm promptly. Customers may contact 1-(866) 229-3002 for questions. |  | 891 | 891 | Nationwide Distribution, including the states of CA, CO, FL, MO, and TX. | Terminated | December 13, 2012 | 232 | 891 | Not_Computer
Z-1456-2012 | 61473 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Certain Low-Profile, Angled Abutment, Model Number ILPAC3217. Product is labeled in part: "***Biomet 3i***4555 Riverside Drive , Palm Beach Gardens, FL 33410 USA***REF ILPAC3217***LOT 2011091179***Certain Low Profile 17 Abutment***3.4mm(D) X 2mm(H)***Abutment:; Pillar; Abutment; Abutment; Pilastro; Pillar***CE dose***Use by 2016-09***Sterile using Radiation***RX Only***Do not reuse***Consult accompanying documents***Do not re-sterilize***Do not use if package is damaged***ILPAC3217***2011091179***"P-LBLNY3P***Rev. D***".Intended to provide a means to attach crown restorations or implant frameworks to implants. | Certain LowProfile, Angled Abutment | 2 | April 25, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i initiated a recall for a small percentage of the 33mm (ILPAC3217) abutment packages that may contain the 5mm (ILPAC5217) abutment. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Biomet 3i , LLC sent Urgent: Medical Device Recall notices on 1/10/2012 to each customer in via e-mail, facsimile, telephone or postal mail. The letters identified the affected product and the reason for the recall. Each customer was instructed to check their respective inventory for the affected product and return any unused product to Biomet 3i at the address provided. Biomet 3i will issue replacement product at no charge. Questions or concerns should be directed to 1-800-342-5454. | No consumer action necessary | 400 | 400 | Worldwide Distribution -- Canada, Chile, Columbia, France, Germany, Hong Kong, India, Israel, Italy, Japan, Mexico, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, UK, US, and Vietnam. | Terminated | April 30, 2012 | 5 | 400 | Not_Computer
Z-1472-2012 | 61409 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS&amp;amp;amp;amp;amp;reg; 5,1 FS Refurbished System, Product Codes: 6801375 &amp;amp;amp;amp;amp;amp; 6801890Product Usage:For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Products Slides, VITROS Chemistry Products MicroTip Reagents and VITROS Immunodiagnostic Products Reagents. | VITROS | 2 | April 26, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics is recalling VITROS Chemistry Systems due to possible internal hard drive failure. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ortho Clinical Diagnostics (OCD) sent an Important Product Correction Notification letter dated March 15, 2012 via FedEx overnight courier to all affected customers. Foreign affiliates were informed by email on 15 March 2012 of the issue and instructed to notify their consignees of the issue and actions. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to: 1) Consider performing the backup procedure more frequently to help prevent the loss of data on your system; 2) Contact OCD Customer Technical Service representatives if a hard drive failure occurs on your system; A Field Engineer will be dispatched to replace the hard drive on your system; 3) Complete and return the Confirmation of Receipt form upon receipt of this notification; and 4) Post this notification by each VITROSÂŽ System in your facility or with the user documentation. For questions, refer to the Questions and Answers section. For additional questions, contact Customer Technical Services at 1-800-421-3311. |  | 25 | 25 | Worldwide Distribution | Open |  |  | 785 | Not_Computer
Z-1467-2012 | 61516 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage Profiler SOB Panel, PN 97300. The test is used as an aid in the diagnosis of myocardial infarction (injury), an aid in the diagnosis and assessment of severity of heart failure, an aid in the risk stratification of patients with heart failure, an aid in the assessment and evaluation of patients suspected of having disseminated intravascular coagulation or thromboembolic events including pulmonary embolism and an aid in the risk stratification of patients with acute coronary syndromes. | Alere Triage Profiler SOB Panel | 2 | April 26, 2012 | 2012 | Alere San Diego | Alere San Diego, Inc. is initiating a voluntary recall of the Alere Triage CardioProfiler Panel PN 9700CP, Lot W49569 and the Alere Triage Profiler SOB Panel PN 97300, Lot W48990. Alere San Diego, Inc. has determined that these lots have an increased frequency of Troponin I results | PRODUCTION CONTROLS: Process Control | A recall communication was initiated on 3/15/12 with Alere San Diego, Inc. forwarding a "Urgent Medical Device Recall" letter to all their customers who purchased the Alere Triage CardioProfiler Panel PN 97100CP, Lot W49569 and Alere Triage Profiler SOB Panel PN 97300, Lot W48990. The letter informed the customers about the problems identified and the actions to be taken. Customers with any questions about the information contained in this notification should contact Alere San Diego, Inc. at (877) 308-8287. |  | 1316 | 1,316 kits | Worldwide Distribution, including Nationwide (US) and the countries of Brazil, China, India and Kuwait. | Terminated | December 04, 2012 | 222 | 2396 | Not_Computer
Z-1476-2012 | 60698 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Micro-Tech Informer Plus Silver Monitor, Model 81830. The firm name on the label is Stanley Security Solutions, Inc., Lincoln, NE.Product Usage:The system is a bed exit system designed to indicate, by alarm or signal, when someone attempts to leave the bed. The system consists of a control unit that is mounted on the bed and a pressure sensitive bed sensor pad that is placed across the width of the bed under the mattress or just beneath the sheets and overlays. The control unit is powered by means of a plug-in type power supply and an alarm cord is used to interface the control unit with the nurse call system. | MicroTech Informer Plus Silver Monitors | 2 | April 28, 2012 | 2012 | Stanley Security Solutions, Inc. | Connection between the call cord and the receptacle may not make a correct connection resulting in a continuous or intermittent (false) signal to the nurse call system | DESIGN: Component Design/Selection | Stanley Healthcare Solutions, Inc. contacted affected customers via phone and email on December 13, 2011. A follow-up certified letter was sent on December 14, 2011. Customers were additionally sent and "Important Recall Notice letter on December 16, 2011 after errors in the original letters were detected. The letter identified the affected product, problem, actions to be taken, and included a reply form for the customer to complete and return. The recalling firm requested their customers to stop using the monitor and contact them for return of the unit. Contact information (phone, fax, & email) was included. For question contact Micro-Tech Customer Service at 1-800-206-1045. |  | 824 | 824 units | Worldwide Distribution - USA Nationwide including the states of CT, FL, KY, MN, NC, NE, NH, NJ, NY, PA, TX and the country of Canada (5 units). | Terminated | September 24, 2013 | 514 | 824 | Not_Computer
Z-1478-2012 | 61658 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | REF 28234 Troch Nail, Long 11MM, Right 34 CMPeritrochanteric Nail System, Rx SterileBiomet Trauma Persippany, NJRod, Fixation, Intramedullary and AccessoriesProduct Usage:An intramedullary fixation rod is a device intended to be implanted that consists of a rod made of alloys such as cobalt-chromiummolybdenum and stainless steel. It is inserted into the medullary (bone marrow) canal of long bones for the fixation of fractures. | Trochanteric Nail. Long, | 2 | April 30, 2012 | 2012 | Biomet, Inc. | The Trochanteric Nail, lot 486130, wasmislabeled as a right nail orientation but the contents are actually a left nail orientation | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet sent an Urgent Medical Device Recall Notice dated April 18, 2012 to all affected consignees. The letter identified the affected product, reason for the recall including risks involved and actions to be taken. The notice provided directions to immediately locate and return the affected product. Receipt Confirmation was requested by calling 800-348-9500 Ext 3009 or 3983. FAX Back Response Form should be completed and returned to 574-372-1683. Questions should be direct to 574-371-3009 or 574-372-3983, M-F, 8 AM to 5 PM. | Contact the recalling firm for information | 30 | 30 units in total | Worldwide Distribution - Nationwide (USA) | Terminated | March 08, 2013 | 312 | 30 | Not_Computer
Z-1477-2012 | 61635 | chamber, oxygen, topical, extremity | KPJ | Topical oxygen chamber for extremities. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | EPIFLO; Transdermal continuous oxygen therapy device; single use; disposable medical device. Model Number EPIFLO-15. Mfr.By: Neogenix, LLC, dba Ogenix, 3401 Enterprise Pkwy, Suite 340, Beachwood, OH 44122.The EPIFLO System is intended to provide topical oxygen to treat 1.) skin ulcerations due to diabetes, venous stasis, post-surgical infections and gangrenous lesions 2.) Decubitus ulcers 3.) amputations/infected stumps 4.) skin grafts 5.) burns 6.) frostbite. | EPIFLO | 2 | April 30, 2012 | 2012 | Ogenix Corporation | Ogenix Initiated a voluntary recall for the EPIFLO-15 devices due to a mislabeling of the expiry date for this lot. Instead of March 2013, this lot carried May 2013 as the expiry date. | PRODUCTION CONTROLS: Labeling Mix-Ups | Ogenix sent a letter dated March 16, 2012 to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to confirm the status of the devices and ship product with the recalled lot back to Ogenix and complete the form provided with the letter. | Contact the recalling firm for information | 38 | 38 units | Worldwide Distribution-USA (nationwide) including the states of FL, MN, OH and the county of Turkey. | Terminated | February 19, 2013 | 295 | 38 | Not_Computer
Z-1480-2012 | 61653 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems Winged Infusion Set, 22Ga, 1",Latex-Free Winged Infusion SetProduct Code: MBS-2210Huber type needle set used for the infusion of medical fluids into implantable ports. | Churchill Medical Systems 22Ga Winged Infusion Set | 2 | May 01, 2012 | 2012 | Churchill Medical Systems, Inc. | Mislabeled: MBS-2210 code on case label -22Ga Winged Infusion Set, ; however, the product label states 20Ga MBS-2010 code listed on devicelabel. | MISBRANDING: Labeling False and Misleading | The firm, VYGON, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 16, 2012 to its customers via overnight, courier service. The letter described the product, problem and actions to be taken. The customers were instructed to check their stock for the products included within the scope of the recall; count and document their affected inventory; contact Vygon's Customer Service Department at 1-800-473-5414 for return merchandise authorization number; return the affected product; complete and immediately return the Acknowledgement and Inventory Return Form via fax to 215-672-6740 and send the original form with their shipment . Distributors were instructed to communicate this recall to their customers and provide them with the Acknowledgement and Inventory Return Form for completion and return.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@yygonus.com. | Contact the recalling firm for information | 552 | 552 units | Worldwide distribution: USA (nationwide) including states of: IN, KS, KY, NY,MA, MI , and TX; and country of: Canada. | Terminated | March 28, 2013 | 331 | 552 | Not_Computer
Z-1481-2012 | 61411 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Coloplast, Restorelle DirectFix, Posterior Polypropylene Mesh, Catalogue Number 501460, Rx Only, Manufacturer, Coloplast A/S 3050 DenmarkRestorelle is a ploypropylene mesh that may be used for the repair of abdominal wall hernia, including inguinal, fermoral, and incisional, and uterovaginal proplapse and other fascial deficiencies that require support material. It may be used in open or laparoscopic abdominal procedures or for repair by the vaginal route. | Restorelle DirectFix Posterior Mesh | 3 | May 02, 2012 | 2012 | Coloplast Manufacturing US, LLC | Coloplast is conducting a field correction based on the conclusion that two lots of Restorelle DirectFix Posterior mesh were mislabeled. The incorrect size was printed on the patient label. Specifically the patient label depicts the size as 16x11 cm, however, the correct size is actually 14x15 cm. It is important to note that both the product and box labels were correct. | PRODUCTION CONTROLS: Error in Labeling | Coloplast sent a notification letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed their coloplast representative is providing replacement device (s) and new patient labels. For further questions please call 1 (800) 258-3476. | Contact the recalling firm for information | 134 | 134 | USA ( nationwide) including the states of AL, AZ, AR, CA, CO, FL, GA, ID, IL, KY, MD, MS, MO, NY, PA, TN, VA,WA and WI. | Terminated | October 11, 2012 | 162 | 134 | Not_Computer
Z-1485-2012 | 61602 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | VALOR(TM) TTC FUSION, INSTRUMENT KIT, REF 4150-KIT1, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002.The VALOR" Ankle Fusion Nail System is intended to facilitate tibiotalocalcaneal arthrodesis to treat severe foot/ankle deformity, arthritis, instability, and skeletal defects after tumor resection. These include Neuro-osteoarthropathy (Charcots Foot), Avascular Necrosis of the talus, failed joint replacement, failed ankle fusion, distal tibia fracture non-unions, Osteoarthritis, Rheumatoid Arthritis, and Pseudoarthrosis. | VALOR(TM) TTC FUSION, INSTRUMENT KIT | 2 | May 02, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that the handle of the targeting guide may rotate and/or break under normal usage. | DESIGN: Device Design | Wright Medical Technology, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their stock and to return the affected product to the firm. Additionally, a Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact the firm at 901-867-4324 for questions regarding this notice. |  | 133 | 133 units | Worldwide Distribution-USA (nationwide) and the country of Australia. | Terminated | January 28, 2013 | 271 | 133 | Not_Computer
Z-1488-2012 | 61464 | laryngoscope, nasopharyngoscope | EQN | Nasopharyngoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | BOMImed Reusable Integrated Fiber Optic Laryngoscope Blades, MacIntosh and Miller styles - all sizes. Instrument is used to aid with intubation. | BOMImed Reusable Integrated Fiber Optic Laryngoscope Blades | 2 | May 03, 2012 | 2012 | Bomimed Inc | On 9/1/2010 BOMImed, Inc. Winnipeg, MB, Canada initiated a recall of their BOMImed Integrated Reusable Laryngoscope Blades (Macintosh & Miller styles - all sizes), Model #OL-32D0 to OL-32D4 & OL-32E00, OL-32E0, OL-32E1, OL-32E2, OL-32E3, OL-32E4. | DESIGN: Process Design | BOMImed sent a "Medical Device Safety Alert & Correction Action" letter by Fax and Mail on 11/25/2010 to the consignee. The product issue was described and recommended actions to mitigate risk were provided. Contact number is 1-800-667-6276 ext. 234. | Contact the recalling firm for information | 30 | 30 units | US Distribution to the state of California. | Terminated | May 08, 2012 | 5 | 30 | Not_Computer
Z-1489-2012 | 61613 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EG-3870UTK, Pentax Ultrasound Video Gastroscope.The EG-3870UTK, Ultrasound Video Gastroscope, is intended to provide optical visualization via a video monitor of, ultrasonic visualization of, and therapeutic access to the Upper Gastrointestinal Tract, including, but restricted to the organs, tissues, and sub-systems; Esophagus, Stomach, Duodenum, Small Bowel, and underlying areas. The instrument is introduced per orally when indications consistent with the need for procedure are observed in adult and pediatric patient populations. | EG3870UTK, Pentax Ultrasound Video Gastroscope | 2 | May 03, 2012 | 2012 | Pentax Medical Company | Endoscope ultrasound balloon evacuation tube may become clogged with patient body fluid if users fail to properly follow instructions for pre-procedure check, operating the suction valve during procedure and pre-cleaning procedure. | PRODUCTION CONTROLS: Equipment Maintenance | Pentax sent an "URGENT MEDICAL DEVICE RECALL" letter on March 14, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. A response form was included for customers to complete and return to the firm. Contact the firm at 800-431-5880 ext. 2064 for questions regarding this notice. | Contact the recalling firm for information | 194 | 194 units | Nationwide Distribution. | Open |  |  | 194 | Not_Computer
Z-1482-2012 | 61656 | defibrillator, automatic implantable cardioverter, with cardiac resynchronization (crt-d) | NIK |  |  | Cardiovascular |  | Cardiac Resynchronization Therapy (CRT) Leads.Trade Name: QuickSite, QuickSite XLModel Numbers: QuickSite and QuickSite XL: Models 1056T, model 1058TIntended use: Left Ventricular Cardiac Resynchronization Therapy Lead | QuickSite, QuickSite XL | 2 | May 03, 2012 | 2012 | St Jude Medical CRMD | Visual observations of externalized conductors due to abrasion of the silicone insulation in the distal portion of these QuickSite and QuickFlex leads. | DESIGN: Component Design/Selection | A recall communication was initiated on 4/3/12 with St. Jude Medical forwarding a Medical Device Advisory to all doctors who have received the St. Jude QuickSite and QuickFlex Left Ventricular CRT leads models 1056T, 1058T, 1156T, and 1158T. The letter provides the doctors with the problems identified and actions to be taken. Doctors with any questions or concerns are instructed to contact their local St. Jude Medical representative or their Technical Services Department at (800) 722-3774.A Physician Device Advisory Notice for the QuickSite and QuickFlex Left Ventricular CRT Leads was issued on 4/3/12 to physicians who have received the recalled products. The notice provides the physicians with the problem identified and actions to be taken. Physicians are instructed to return any packaged product in hospital inventory to SJM Sales Representatives. For implanted product, due to the nature of the issue, externalized conductors may be observed with no electrical or performance anomalies. Therefore, prophylactic explant is not recommended. It is recommended that physicians continue to monitor their patients' implanted system at regularly scheduled intervals. A Medical Advisory and Inventory Management Information letter was sent on 4/3/12 to hospital administrators to notify them of the problems identified and the actions to be taken. | Contact the recalling firm for information | 101000 | 101,000 units (59,000 units in US) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Bahrain, Balkan, Baltic, Belarus, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Mauritius, Montenegro, Morocco, Netherlands, New Caledonia, New Zealand, Norway, Oman, Paraguay, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Senegal, Serbia, Serbia Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, UK, United Arab Emirates, US, Venezuela, and Yemen. | Terminated | July 17, 2012 | 75 | 101000 | Not_Computer
Z-1492-2012 | 61413 | coil, magnetic resonance, specialty | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Optima MR450W with GEM.Product Usage - The GEM Option for 1 .5T MRI systems is a set of receive-only RF surface coils designed for use with 1 .5T MRI systems manufactured by GE Healthcare. The GEM Option for 1 .5T MRI systems is indicated for use for: head, neck, brachial-plexus, spine, pelvis, hips, prostate, abdominal, cardiac, lower extremities, blood vessels, and long bone imaging. The nucleus detected is hydrogen. | Optima MR450W with GEM. | 2 | May 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is conducting a recall on Optima MR450W with GEM. During scanning of a patient, if an improperly torqued gradient connection exists, it may cause the gradient lead to produce excessive heat. If a second and compounding issue of improperly routed Rear Pedestal Air Supply Hose is also present, routed in such a way to cause the hose to come in contact with the overheated gradient lead | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare is not sending a Communication Letter to customers. There is no plan to contact customers with an affected Optima MR450W with GEM system. A GE Healthcare Field Service Engineer will be dispatched to sites to inspect and correct all potentially affected systems. As an effectiveness check, GE Healthcare will track all service requests for the field action to completion. For any questions or additional information, call 262-548-2608. | Contact the recalling firm for information | 41 | 41 | Worldwide Distribution - USA (Nationwide) including the states of :CO, FL, IA, MA ,MN, NJ, NY, NC, OH, TX. and the countries of BRAZIL., BRITAIN, FINLAND, FRANCE, GREAT GERMANY, ITALY, SAUDI ARABIA, SWEDEN, SPAIN, KUWAIT, and SWIITZERLAND. | Terminated | August 09, 2012 | 97 | 41 | Not_Computer
Z-1491-2012 | 61680 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeShield Primary Plumset, CLAVE port, CLAVE Y-Site, 104 inch, 15 drops/mL, 48 sets per case; a sterile Rx device; Hospira Inc, Lake Forest, IL USA. Made in Costa Rica; product 120300412, list number 12030-12. Product Usage: For administration of I.V. fluids. | LifeShield Primary Plumset | 2 | May 04, 2012 | 2012 | Hospira Inc. | The Lifeshield Primary Plumsets may cause a cassette test failure alarm when the set is loaded onto the infusion pump. A set which results in a cassette test failure alarm may result in delay in therapy, requiring medical intervention. | PRODUCTION CONTROLS: Process Control | Hospira, sent an "URGENT DEVICE RECALL" letter dated April 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product to return for credit. A Reply Form was enclosed for customers to complete and return via fax to 1-888-386-2073 or e-mail to: Hospira8503@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for information regarding product availability. |  | 145875 | 145,875 sets | Worldwide Distribution - USA (nationwide) including the states of: Alabama, Colorado, Georgia, Illinois, Indiana, Iowa, Kentucky, Maryland, Massachusetts, Michigan, Minnesota, Missouri, North Carolina, North Dakota, Oklahoma, Pennsylvania, Wisconsin and the country of Australia. | Open |  |  | 145875 | Not_Computer
Z-1496-2012 | 61709 | methyl methacrylate for cranioplasty | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Doctors Research Group Kryptonite Bone Cement packaged in 5cc and 10 cc Kits:Product Codes:KRYP-US-Z-05 (5cc Kit)KRYP-US-Z -10 (10cc Kit)Product Usage:Kryptonite Bone Cement is a resinous material indicated for use in repairing cranial defects | Kryptonite Bone Cement | 2 | May 08, 2012 | 2012 | Doctor's Research Group, Inc. | Bone cement strength and stiffness fall below specification at body temperature | DESIGN: Component Design/Selection | Doctors Research Group (DRG) issued a Voluntary Product Recall Notification letter dated April 2012 to all affected customers via email or Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were instructed to identify, segregate and return any affected products in their possession. Customers were asked to fill out and return the enclosed verification form Distributors were requested to notify their customers. For questions contact DRG at 203-262-9335. | No consumer action necessary | 3100 | 3100 kits | USA Nationwide Distribution | Open |  |  | 3100 | Not_Computer
Z-1498-2012 | 61610 | manual antimicrobial susceptibility test systems | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Dried Overnight Gram Negative Panels;Panel type: Neg Breakpoint Combo Type 34, catalog number: B1017-404. Siemens Healthcare Diagnostics,West Sacramento, CA. | MicroScan Dried Overnight Gram Negative Panels; Panel type: Neg Breakpoint Combo Type 34 | 2 | May 08, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Three lots were confirmed to not meet performance claims due to random non-biological contamination which is only apparent (precipated) after inoculation of the panels. | PRODUCTION CONTROLS: Process Control | Customer Notification letters were sent to all customers of the three MicroScan panels by traceable mail. The letter identified the affected product and asked customers to discard any remaining inventory in addition to contacting customer service for replacement product. The letter also recommended that customers consider the need to review previous test results, conduct patient follow-up, and or repeat tests as needed. Customers were to complete and return the customer reply card to confirm that they received and understood the recall communication. | Contact the recalling firm for information | 680 | 680 boxes (20 panels per box) | Worldwide Distribution -- US, Turkey, Poland, and Great Britain. | Terminated | May 08, 2012 | 0 | 2050 | Not_Computer
Z-1500-2012 | 61646 | enzyme linked immunoabsorbent assay, mumps virus | LJY | Mumps virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | AtheNA Multi-Lyte MMV Test System- A93111GA microparticle-based immunoassay intended for the qualitative presumptive detection of IgG class antibody to the Mumps virus in human serum using the AtheNA Multi-Lyte Test System. The test system is intended to be used for determination of a previous infection with the Mumps virus. | AtheNA MultiLyte MMV Test System AtheNA MultiLyte MMRV Test System A93111G | 3 | May 08, 2012 | 2012 | Zeus Scientific, Inc. | Select lots of AtheNA Multi-Lyte MMV Test System have lost reactivity and may produce false positive results in seronegative patients who have not been previously affected with or vaccinated against Mumps. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Zeus Scientific, sent an "URGENT NOTICE; Voluntary Product Recall" letter dated March 8, 2012 their customers via Fed Ex. The letter described the product, problem and actions to be taken. The customers were instructed to remove and discard any remaining inventory of the listed product; if you are not the end-user and/or you have distributed these products to customers, notify them accordingly; and complete, sign and return the verification notice via fax to: 908-526-2058 or email a PDF to support@zeusscientific.com.If you have any questions, call 908-526-3744 or 800-286-2111. |  | 180 | 180 kits | Nationwide distribution. | Terminated | July 03, 2013 | 421 | 180 | Not_Computer
Z-1504-2012 | 53828 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | Caridian BCT COBE Spectra System Blood Warmer, Catalog number 950000000. | COBE Spectra | 2 | May 09, 2012 | 2012 | CaridianBCT, Inc. | Potential for air entry into the fluid pathway of the COBE Spectra due to a leak at the return luer connection to an elevated bloodwarmer when directions for use are not followed. | DESIGN: Software Design | Caridian BCT sent a letter to customers dated June 2011. It was a Field labeling addendum to the COBE Spectra Apheresis System Operator's Manual stating: "When connecting a blood warmer tubing set to the return line, ensure that the tubing connection is tight. Put the luer connection no higher than 20 in (50 cm) above the return access to prevent the possibility of air entering the tubing." This correction was distributed to all consignees and is included with the current version of the owner's manual. |  | 2459 | 2459 machines | Worldwide distribution including USA, Canada, Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Netherlands, UK, Hong Kong, Hungary, India, Italy, Iran, Ireland, Israel, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Egypt, Finland, France, Germany, Japan, Malaysia, Mexico, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Romania, Russia, Spain, Saudi Arabia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Ukraine, Uruguay, Vatican, Vietnam, and Greece. | Terminated | May 11, 2012 | 2 | 2459 | Not_Computer
Z-1502-2012 | 61384 | test, sickle cell | GHM | Sickle cell test. | Hematology | Hematology | 510(k) | BioTron Sickle Cell Reagent Set, Screening test for the Presence of Hemoglobin S 2 x 100 ml Sickle Cell Buffer, 2 vials sickle Cell Reagent, Size 100 tests, Store 2-8 C, BioTron Diagnostics Inc, Hemet California, USA .For the qualitative determination of positive or negative Sickle Cell Anemia present in Hemoglobin , for in vitro diagnostic use only. | Biotron | 2 | May 09, 2012 | 2012 | MichClone Associates, Inc. | Two complaints received concerning the powder vial included in the 100 test kit observed producing pressure while sealed. When the powders are opened, the pressure from within the vial releases a small amount of powder into the air.This powder. if released near the face and inhaled can cause minor irritation resulting in sneezing and nasalcongestion. Investigation found no affect on the perf | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Michclone Associates, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify affected lots, notify sub accounts, and complete and return the response form to arrange return of unused product and replacement . Forms can be returned by FAX 248(307)-9498, mail: Michclone Associates Inc. 680 Ajax Drive, Madison Heights, MI 48118 or Email: toben@greatlakesdiag.com. Questions and concerns should be addressed to Technical Support Department at 248-307-9494. | Contact the recalling firm for information | 250 | 250 | Nationwide Distribution. | Terminated | December 13, 2012 | 218 | 770 | Not_Computer
Z-1505-2012 | 61675 | activator, ultraviolet, for polymerization | EBZ | Ultraviolet activator for polymerization. | Dental | Dental | 510(k) | DISPOSA-SHIELD. The Disposable Curing Shield For Your Curing Light. The disposable Disposa-Shield are sold in different sizes from 6 mm, 8 mm, 9 mm, 11 mm and 13 mm.The devices are Intended to be place over the tip of a dental curing light wand to protect the doctor's and patient's vision from exposure to high intensity light; soften contact of curing light against the teeth, and keep light guide cleaner, therefore maintaining good levels of light intensity for proper and fast curing The label of the O'Ryan Disposa-Shield listed the following: -Protect Doctor's And Patient's Vision From Exposure To High Intensity Light. -Soften Contact Of Light Guide Against Teeth. -Keep Light Guide Cleaner, Therefore, Maintaining Good Levels Of Light Intensity For Proper And Faster Curing.There are 25 devices per clear plastic bag when distributed in bulk and two devices per clear plastic zip lock type bag when distributed with curing lights.The "Sterilize Before Use" label is placed on a clear plastic bag containing 25 pieces of Disposa-Shield units distributed with all shipment of bulk Disposa-Shield.The "Sterilize Before Use" label is NOT placed on plastic zip lock type bag when distributed with two Disposa-Shield with curing lights.The product is labeled in parts:"***DISPOSA-SHIELD***The Disposable Curing Shield For Your Curing Light***STERILIZE BEFORE USE***O'RYAN INDUSTRIES***INSTRUCTIONS FOR INSTALLING DISPOSA-SHIELD***1. Remove and dispose of old DISPOSA-SHIELD***2. Sterilize wand tip***3. Simply stretch new DISPOSA-SHIELD over tip of wand***800-426-4311 or (360) 892-0447***(360) 892-6742 Fax***info@oryanindustries.com***'.The private label is labeled in parts: "***DISPOSA-SHIELD***The Disposable Curing Shield For Your Curing Light***ALMORE INTERNATIONAL , INC***STERILIZE BEFORE USE***". | DISPOSASHIELD | 2 | May 09, 2012 | 2012 | O'Ryan Industries Inc | Disposa-Shield Disposable Curing Shield is recalled due to lacking 510K, PMA or IDE, and inadequate quality system. |  | The firm, O'Ryan Industries Inc., began to contact consignees by telephone and sent email on April 25, 2012. O'Ryan sent a recall notification letter dated May 7, 2012 to their customers through the Purchasing Agents and Order Department Personnel. The letter described the product, problem, and actions to be taken by the customers. Customers were advised to return the Disposa-Shield back to O'Ryan Industries.Customers can call O'Ryan Industries at 800-426-4311 or fax to 360-892-6742. |  | 11376 | 11,376 units | Nationwide distribution: USA including states of NC, MN, OR and WI. | Terminated | August 10, 2012 | 93 | 11376 | Not_Computer
Z-1520-2012 | 61577 | enzyme linked immunoabsorbent assay, herpes simplex virus, non-specific | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Product is labeled in part: "***is immunosimplicity***HSV 1&2 IgM***Enzyme Immunoassay Test Kit***For the Detection of HSV 1&2 IgM antibodies***Catalog No: 720-350***Exp. Date:***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: ***For In Vitro Diagnostic Use***Caution, consult accompanying documents.***Manufacturer***Diamedix***A Subsidiary of IVAX Diagnostics, Inc. MIAMI, FL 33127 USA-1-800-327-4565***Biohazard***"The Is-HSV 1&2 IgM test kits are packaged in cardboard boxes with individual components secured within a foam insert.The Diamedix Is-HSV 1&2 IgM Test Kit is intended for the qualitative detection of IgM antibodies to herpes simplex virus (HSV) type 1 and/or type 2 in human serum by indirect enzyme immunoassay. This test can aid in the diagnosis of a primary or reactived infection with HSV. The performance of this assay has not bee established for use in neonates, infants, or on cord blood, and immunocompromised patients. | HSV 1&2 IgM Enzyme Immunoassay Test Kit | 3 | May 10, 2012 | 2012 | Diamedix Corporation | On 01/16/20112 Diamedix initiated a recall of IS-HSV 1&2 IgM Test Kit, Catalog # 720-350, Lot 60701, Exp. June 30, 2012. The Positive Control for this lot is Out of Specification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Diamedix sent Recall notification letters dated January 16, 2012, to all affected customers by Certified Mail. Consignees were asked to review their inventory and then contact Diamedix to request replacement product. Customers were instructed to call 305-324-2314 with questions. For questions regarding this recall call 305-324-2560. |  | 442 | 442 kits | Nationwide Distribution and Puerto Rico. | Terminated | May 18, 2012 | 8 | 442 | Not_Computer
Z-1521-2012 | 61685 | device, nerve conduction velocity measurement | JXE | Nerve conduction velocity measurement device. | Neurology | Neurology | 510(k) | Ultra-DriveTool, 6 MM Plug Puller, REF 423868. Recommended Power Setting 70 For Single Use Only, Sterile, Biomet Orthopedics, Warsaw, IN 46581Incorrect Picture Label, Trochanter Plate (50-60 mm) Mallory/HeadThe Ultra-DriveTM System is used to aid in the removal of bone cement (PMMA) and noncemented implants in conjunction with revision arthroplasty. The Ultra-Drive" device converts standard electrical energy into mechanical energy, via an ultrasonic transducer in the Ultra-Drive" handpiece. This energy is transmitted through individually tuned tool tips in acoustic waves. causing the tool tip to vibrate at a controlled high speed. This action breaks down bone cement and cuts through cancellous bone. | UltraDriveTool, 6 MM Plug Puller | 3 | May 10, 2012 | 2012 | Biomet, Inc. | Incorrect product picture label was on the product box. | DESIGN: Process Design | Biomet Orthopedics sent an Urgent Medical Device Recall Notice dated April 18,2012, to all affected customers. The letter identified the reason for the recall including directions to immediately locate, discontinue use of the device, and return product to Biomet, 56 East Bell Drive, Warsaw, IN 46580. Consignees were asked to confirm receipt of the notice by calling 1-800-348-9500 extension 3983 or 3009. A FAX Back Response Form was to be filled out and sent within 3 days of receipt to 574-372-1683. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m.For questions regarding this recall call 574-371-1021. | Contact the recalling firm for information | 30 | 30 units | Worldwide Distribution - USA (nationwide) and Europe, Austrailia | Terminated | April 18, 2013 | 343 | 30 | Not_Computer
Z-1507-2012 | 61713 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product is labeled in part: Inner Label "***S.M.A.R.T. Transhepatic Biliary Stent***REF Cat. No.***Lot No.***Use By***Cordis S.M.A.R.T. Nitinol Stent Transhepatic Biliary System***Assembled in Mexico***Cordis Corporation 14201 North West 60th Avenue Miami Lakes, Florida 33014, USA***Cordis S.M.A.R.T. TRANSHEPATIC BILIARY STENT***Cordis a Johnson & Johnson company***Endovascular***" Catalog C06120MB Lot No's. 15530111, 15535366, 15556361, 15560541, 15562698. Catalog C06150MB Lot No's. 15530115, 15536150, 15556366, 15560543. Catalog C07120 MB Lot No's. 15530114, 15556368, 15560545, 15562701. Catalog C07150MB Lot No's. 15530113, 15556371, 15560546. Catalog C08120MB Lot No's. 15530112. Catalog C08150MB Lot No's. 15535365, 15556374, 15562702. Indicated for palliation of malignant neoplasms in the biliary tree. | Transhepatic Biliary Stent System | 2 | May 10, 2012 | 2012 | Cordis Corporation | Cordis Corporation is recalling 32 lots of Cordis S.M.A.R.T. Nitinol Stent Transhepatic Biliary System and S.M.A.R.T. Control Nitinol Stent Transhepatic Biliary System. There is a potential for a sterility breach due to holes or tears in the Tyvek packaging. | PRODUCTION CONTROLS: Packaging Process Control | Cordis Corp sent a " Urgent Medical Device Recall (Removal) letter dated April 9, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A Response Form was provided with each notification. Customers were asked to complete and return Response Form. The customers were instructed to :" Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used." Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form." Either return any affected product with reference to the RA# on the form, orcontact your local sales representative, or contact Customer Service (see Available Assistance section) to facilitate return of the affected product.Replacement product will be provided in a short time frame. Credit will beprovided for customer-owned inventory upon request." Pass on this notice to anyone in your facility that needs to be informed." If any product listed below has been forwarded to another facility, contact thatfacility to arrange return." Maintain awareness of this notice until all product listed above has beenreturned to Cordis.For questions related to the issue or product replacement, please contactyour local sales representative, or contact the special Customer Servicenumber (800) 551-7683, Monday through Friday from 8:30 AM to 8:00 PMEDT. |  | 984 | 984 uits | Worldwide Distribution-- USA (nationwide) and the country Canada. | Terminated | December 24, 2013 | 593 | 1262 | Not_Computer
Z-1509-2012 | 61630 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT STAT Troponin-I Reagent Kit; list numbers 2K41-28 (500 test kit); an in-vitro diagnostic test kit; each kit contains bottles of microparticles, conjugate and assay diluent; Abbott Laboratories, Abbott Park, IL 60064ARCHITECT STAT Troponin-I is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cardiac troponin-I in human serum and plasma on the ARCHITECT i System with STAT protocol capability. Troponin-I values are used to assist in the diagnosis of myocardial infarction (MI) and in the risk stratification of patients with acute coronary syndromes (including unstable angina and non-ST elevation) with respect to relative risk of mortality, myocardial infarction, or increased probability of ischemic events. | ARCHITECT STAT TroponinI Reagent Kit | 2 | May 10, 2012 | 2012 | Abbott Laboratories | The lot of Troponin-I is demonstrating a shift in expected results in some cases, and varies from kit to kit. A concentration change of up to +/-55% could occur when switching kits without recalibrating. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required"letter dated April 11, 2012 to all Troponin-I customers who received the affected reagent lot internationally. The customers were directed to determine if they are currently using and/or have inventory of lot 74264UN11; to discontinue use and destroy any remaining inventory of the lot; to follow their facility's policies and procedures for communicating with the heath care providers they serve; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; to retain a copy of the letter for their records and to complete and return the Customer Reply Form (sent to Abbott international affiliates via e-mail on 4/11/2012) via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. Any questions were directed to their local Customer Service representative or call the field action coordinator at 847-938-1923 or email: albert.chianello@abbott.com. | Contact the recalling firm for information | 8723 | 8,723 kits | International distribution only: including countries of: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia & Herze, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Mali, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela and Vietnam. | Open |  |  | 8723 | Not_Computer
Z-1516-2012 | 61661 | test,natriuretic peptide | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute care(TM) NT - proBNP TestPak reagent (CPBNPM) The Stratus(R) CS Acute care(TM) NT - proBNP method (pBNP) is an in vitro diagnostic test for the quantitative measurement of N-terminal pro-brain natriuretic peptide (NT0proBNP) in heparinized plasma. In individuals suspected of having congestive heart failure (CHF), measurements of NT-proBNP are used as an aid in the diagnosis and assessment of severity. The test is further indicated for the risk stratification of patient with acute coronary syndrome and heart failure. | Stratus(R) CS Acute care(TM) NT | 2 | May 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a low frequency of calibration failures due to slope errors that are not resolved with routine maintenance with the specific SCS Pro-BNP lots listed. This issue only causes the slope error failures but does not affect patient results when the calibration is within specifications. | DESIGN: Component Design/Selection | The firm, SIEMENS, sent a "Urgent Field Safety Notice" dated March, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue use of the lots once revised expiration date occurs; contact their local Siemens Customers Service Center at 888-588-3916 for replacement; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the TECHNICAL SOLUTIONS CENTER at (302) 631-8467 as soon as possible, and forward this notification to anyone to whom they may have distributed this product. The remaining inventory at the firm was destroyed.If you have any technical questions regarding this information, please contact the Siemens Technical Solutions Center at 800-405-6473. | Contact the recalling firm for information | 1899 | 1,899 cartons | Nationwide distribution: USA including states of: AK, AR, CA, CO, CT, DC, FL, HI, ID, IN, IL, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NH, NM, ND, NY, OH, OK, OR, PA, SD, TX, UT, VT, WA, WV, WI, and WY. | Terminated | March 19, 2013 | 313 | 1899 | Not_Computer
Z-1515-2012 | 61481 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental Trabecular Metal Implant, TMM4B13, 4.1mm x13MM; 3.5 mm implant TM, MTX Rx, Sterile R, , Zimmer Dental, Carlsbad, CA 92008, USAThe Trabecular Metal Dental Implant is designed to replace one or more missing teeth in the maxilla or mandible and is used during immediate loading or for loading after a conventional healing period. Immediate loading is indicated when there is good primary stability and an appropriate occlusal load (i.e. the contacting surfaces of opposing teeth, especially the biting or chewing surfaces). The 4.1mmD Trabecular Metal Implants are intended to be splinted to additional implants when used in the posterior region. | Zimmer 4.1 Trabecular Metal" Dental Implant | 2 | May 10, 2012 | 2012 | Zimmer Inc. | This Zimmer Dental voluntary device recall resulted from an investigation into the February 2012 complaint involving an apical tip of a 4.1mm D Trabecular Metal Implant which separated from the implant assembly during surgery on a patient with a dense (Type D1), thick, inferior border. The recall correction includes adding precautions to the Instructions for Use regarding the placement of a 4.1mm | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Zimmer sent an Urgent Medical Device Correction notification letter dated April 24, 2012, with an attached Technical Bulletin listing precautions present in the new Informaiton for Use sent to distributors and customers via E-mail and FED EX. Notifications identified the issue and risks found with the device listing responsibilties, precautions, and procedural precautions. Notifications contained directions to forward information to any sub accounts and clinicians. Zimmer had telephone follow-up with each distributor and customer to assure information was received and is understood. Questions should be directed to Technical Service (800)8511-7019 orCustomer Service(760) 929-1300 or Regulatory Affairs at 1-800-854-7019 .For questions regarding this recall call 574-372-4807. |  | 195 | 195 | Worldwide Distribution - USA (nationwide) and the countries of Germany, Italy, France Egypt, Spain, Russia, Bulgaria, Turkey, Israel, UAE and CANADA | Terminated | December 03, 2012 | 207 | 1209 | Not_Computer
Z-1519-2012 | 61522 | test, qualitative and quantitative factor deficiency | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Instrumentation Laboratory (IL)HemosIL von Willebrand Factor Activity Assay KitPart Number: 0020004700The assay is intended for the quantitative determination of von Willebrand Factor Activity (VWF Activity) in human citrated plasma on IL Coagulation Systems | HemosIL von Willebrand Factor Activity Assay kit | 2 | May 10, 2012 | 2012 | Instrumentation Laboratory Co. | Falsely elevated levels may be reported with the HemosIL von Willebrand Factor Activity assay. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Instrumentation Laboratories mailed Recall Letter via USPS to US and Canadian customers on April 11, 2012, through Beckman Coulter. The customer mailing and tracking of the Urgent Recall Notification and Recovery Procedure in the United States and Canada will be handled through the distributor, Beckman Coulter. Accounts are requested to check inventory for the product lots of Hemosll von Willebrand Factor Activity, PN 0020004700, listed in the table and document their destruction on the providedtracking form and return the completed form. Questions or concerns regarding this notification, please call Beckman Coulter Technical Service at (800) 526-7694.On March 29, an e-mail communication was circulated to the IL Affiliates in Europe and Japan requesting them to translate the Urgent Field Safety Notification [title of document outside US and Canada] and product destruction tracking form, mail the notification to their customer base and notify theirCompetent Authorities..Fpr questions regarding this recall call 781-861-4467. | Contact the recalling firm for information | 9320 | 9320 Kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS BELGIUM, BRAZIL, CHILE, CHINA, CZECH REPUBLIC, DENMARK FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF KUWAIT LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS NORWAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, and the UNITED ARAB EMIRATES | Terminated | August 24, 2012 | 106 | 9320 | Not_Computer
Z-1506-2012 | 61517 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Siemens Axiom Sensis XP diagnostic programmable computerDiagnostic programmable computer | Siemens Axiom Sensis XP | 2 | May 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in the connection of a Sensis EP system with Carto(R) 3 dedicated cable set (part number 14411343) with the Carto 3 system from Bionsense Webster. |  | Siemens sent a recall/correction letter dated February 22, 2012, via Update Instruction AX049/11/S of this product to all affected customers. The letter informed customers of the potential issues and provides additional information on the implemented resolution. An Update Instuction AX048/11/S will be available immediately. The update instruction provides a new version of the Carto 3 addendum to the Operator Manual and the exchange of the previous stim cables with the latest cable.For questions call 610-219-4834. |  | 3 | 3 | Nationwide Distribution including FL, NC, and SC. | Terminated | January 24, 2014 | 624 | 3 | Not_Computer
Z-1518-2012 | 61681 | gauze/sponge, internal, x-ray detectable | GDY | Nonabsorbable gauze for internal use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | 8" x 4" 16-ply Dual X-Ray Gauze, hospital fold, banded in 10's; 200 pieces per bag, 10 bags per case; Distributed By: International Medsurg Connection, Schaumburg, IL USA 1-847-619-9926; Made in China; item IMSGZXX8416-10For use to control bleeding and absorb fluid. | 8" x 4" 16ply Dual XRay Gauze | 2 | May 10, 2012 | 2012 | International Medsurg Connection, Inc. | The lots of X-Ray Gauze are contaminated with foreign material, including remnants of cotton seeds and lint. | PRODUCTION CONTROLS: Process Control | International Medsurg Connection notified their direct accounts via telephone, e-mail and/or fax on April 9, 2012, informing them of the foreign material contamination of the gauze, and requesting them to quarantine all affected lots in inventory, cease distribution of the affected product, and notify any customers at the retail level to remove and discard these products. Any questions were directed to (847) 619-9926. | Contact the recalling firm for information | 15 | 15 cases | Nationwide Distribution including Illinois, South Carolina, Texas and Virginia | Open |  |  | 291 | Not_Computer
Z-1539-2012 | 52454 | endoscope, neurological | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Mitaka Point Setter, holding system for surgical instruments.Holding surgical instruments during operating procedures. | Point Setter | 2 | May 11, 2012 | 2012 | Mitaka USA, Inc. | Use of a potentially weaker spring in the device than required to meet specifications to hold the Point Setter ARM in place during surgery. | DESIGN: Component Design/Selection | Mitaka USA sent a letter dated to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included instructions for returning the product. Contact the Customer Support Department at 800-421-0837, ext. 7577 for questions regarding this notice. | Contact the recalling firm for information | 18 | 18 units (3 internationally and 15 within the U.S.) | Worldwide Distribution-USA (nationwide) including the states of AZ, FL, CA, SC, OH, WA, TX, VA, MI, MO, GA, and NC and the country of Canada. | Terminated | May 22, 2012 | 11 | 18 | Not_Computer
Z-1528-2012 | 61644 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardiovascular Procedure Kit-manufactured to the design and specifications of the purchaser/user) for the application it is intended.Cardiovascular Custom Procedure Kit AdultProduct Code: 71060-01 | CAPIOX FX Hollow Fiber Oxygenator | 2 | May 11, 2012 | 2012 | Terumo Cardiovascular Systems Corp | Loose connections (i.e., connections between PVC circuit tubing and Oxygenator ports) may result in leaks and patient blood loss | DESIGN: Process Design | Terumo Cardiovascular Systems issued an Urgent Medical Device Recall notification on 4/12/12 to all user facilities/consignees who have received affected devices. The letter identified the affected product, stated the reason for the recall, the potential hazard involved, and the correction. Users are to review the letter, ensure that all users are made aware of the issue, discontinue use of the affected product, and return it to Terumo CVS. In addition, users are to return the response form to the manufacturer. Affected units will be replaced with new units. Contact Terumo CVS Customer Service @ 1-800-521-2818 or tcvs.recall@teruomedical.com. | Contact the recalling firm for information | 22 | 22 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, MO, and NY and the country of Canada. | Terminated | January 23, 2013 | 257 | 843 | Not_Computer
Z-1522-2012 | 61786 | sterilizer, steam | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | The Evolution steam sterilizers, manufactured by Steris Corporation, utilize a loading system to assist the operator in transporting, loading and unloading items to be sterilized. It consists of two major components: -Loading Car: Shelving unit which holds the items to be sterilized; and the -Transfer Carriage: Wheeled unit which loads/unloads the loading car and transports it from one location to another. Product shipped in crate.Product Usage:The Evolution Transfer Carriage is a wheeled unit which loads/unloads the sterilizer loading car and transports it from one location to another. | Evolution steam sterilizers | 2 | May 11, 2012 | 2012 | Steris Corporation | Through Customer feedback and field service experience, STERIS learned that operators may experience difficulty when inserting or removing equipment from the Evolution sterilizer using the transfer carriage due to the following issues with the transfer carriage. 1) Deformation of the latch pin or cracked ball bearings preventing the roller from turning.2) The rollers on the ends of the pivot a | DESIGN: Device Design | Steris Corporation sent an Urgent Voluntary Field Correction Notice dated March 23, 2012 via FedEx with tracking numbers for delivery confirmation to each affected customer. The letter identified the affected product, problem and actions to be taken. The letter stated a STERIS Field Service Representatives will contact each affected customer to schedule a service visit to upgrade the transfer carriages. For technical questions or questions regarding STERIS visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. | Contact the recalling firm for information | 523 | 523 | Worldwide Distribution - USA Nationwide including AL, AZ, CA, CO, CN, DE, FL, GA, IL, IN, KA, KY, LA, MD, MA, MI, ME, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, & WY and the countries of Brazil and Canada. | Terminated | June 11, 2013 | 396 | 523 | Not_Computer
Z-1536-2012 | 61527 | bed, therapeutic, ac-powered, adjustable home-use | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | LUMEX A Graham-Field Brand, Bed End Sert. Low/Full-electric, Plastic Panels, Item 690-7604-000 Manufactured for : GF Heath Products, Inc., Atlanta, Georgia 30360 Made in China.LUMEX A Graham-Field Brand, Bed End Sert. Low/Full-electric, Walnut Panels, Item 690-7004-000 Manufactured for : GF Heath Products, Inc., Atlanta, Georgia 30360 Made in China.Therapeutic bed intended for home use. | LUMEX A GrahamField Brand, Bed End Sert | 2 | May 11, 2012 | 2012 | GF Health Products, Inc./d.b.a.Basic American Metal Products | GF Health Products, Inc., is conducting a recall on the Lumex Patriot Full Electric Homecare "Bed Ends" manufactured 3/27/09, 5/9/09, 5/30/09 and 6/20/09. The supplier deviated from the print specifications of a machined pin to a sheared pin. This deviation has the potential to provide room for the pin to move during use. If failure were to occur, the bed could descend rapidly causing non-typic | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | GF Health Products, Inc. sent an "URGENT CORRECTION AND REMOVAL" letter dated April 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately cease use of the affected product, secure the product, and to prevent the product from further distribution to customers and end users. Contact the firm at 1-800-556-6508 for questions regarding this notice. | Contact the recalling firm for information | 3095 | 3095 pairs of bed ends | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI and the countries of EL SALVADOR, CHILE, GUATEMALA, TRINIDAD, TOBAGO, and BAHAMAS. | Terminated | July 15, 2013 | 430 | 3095 | Not_Computer
Z-1537-2012 | 61636 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | HemaSyst, Dual Cannula Tip with Adapter 16 ga. x 4" (10.2 cm). Product is labeled in part:"***Pouch label***HAEMACURE***HemaSyst***tm***Contents: One ***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604*** CE 0297***LOT 111144***2014/10***40680b/20004-02***CAUTION: U.S. Federal Law restricts this device to sale by or on the order of a licensed physician (or properly licensed practitioner).***Contents of unopened undamaged package are sterile***CE 0297***Manufactured by: ***Micromedics*** St. Paul, MN 55121-1385 USA***Distributed by:***Haemacure Corporation***Sarasota, FL 34236 USA***Toll Free 1-877-872-4583***www.haemacure.com***Box label***Contents: 5 packages of 1 each***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604***CE 0297***STERILE EO***LOT 11144***2014/10***40681B/20544***".Sterile, single-use instrument for the application of two liquids. | HemaSyst, Dual cannula Tip with Adapter 16 ga. x 4" (10.2 cm) | 2 | May 11, 2012 | 2012 | Haemacure Corporation | This product has been recalled due to the lot number being different on the box and pouch. The lot number on the box label was incorrectly printed as 11144 instead of 111144. The pouch label referenced the correct lot number, 111144. | PRODUCTION CONTROLS: Labeling Mix-Ups | Haemacure Corporation, Sarasota, FL initiated letters via e-mail on 11/16/2009 with delivery and read receipt confirmation to their two consignees. The letters identified the affected product and stated the reason for the recall. The letter also states that the firm will replace each order upon receipt of the mislabeled boxes at no charge to the customer. Customers were asked to return their unused devices. Questions should be directed to customer service at 1-877-872-4583. | Contact the recalling firm for information | 7 | 7 boxes | Nationwide Distribution -- Illinois and California. | Terminated | May 14, 2012 | 3 | 7 | Not_Computer
Z-1540-2012 | 61629 | total,prostate specific antigen(noncomplexed&complexed;) for detection of prostate cancer | MTF |  |  | Immunology |  | Dimension Vista(R) Total Prostate Specific Antigen Flex(R) reagent cartridgeProduct Usage:The TPSA method is an in vitro diagnostic test for the quantitative measurement of total prostate specific antigen (PSA) in human serum and plasma on the Dimension Vista(R) System | Dimension Vista(R) Total Prostate Specific Antigen Flex(R) reagent cartridge | 2 | May 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed a report of leaking Flex(R) reagent cartridge lot 11300BB from Dimension Vista(R) TPSA; when a leaking Flex(R) is used, patient test results may be falsely depressed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated April 14, 2012 to all Dimension Vista(R) System customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to complete the attached form and fax it to 302-631-8467 to indicate that you have received the notification. For technical questions or concerns, contact the Siemens Technical Solutions Center at 800-441-9250. |  | 1103 | 1,103 cartons | USA Nationwide Distribution including the states of: AL, AZ, CA, CO, CT, FL IL, LA, MD, MI, MN, MS, MT, NC, NH, NJ, NV, NY, OH, PA, SD, TN, TX, UT, WI ,WV including Puerto Rico. | Terminated | April 02, 2013 | 326 | 1103 | Not_Computer
Z-1609-2012 | 61468 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00-5970-031-23, Nextgen Complete Knee System, Prolong highly crosslinked polyethylene Cruciate Retaining (CR) Micro Articular Surface Striped Yellow 23 mm Height , Sterile R, Rx ,Zimmer, Inc. Warsaw, INThe NexGen Articular Surface serves as a spacer between the baseplate of the Intended Use. tibia and the metal implant that covers the end of the femur. The micro articular surfaces are made of either conventional or prolong polyethylene in thicknesses of 9, 10, 12, 14, 17, or 20mm. The articulating surface of the device is conforming to the geometry of the NexGen micro femoral components. The articular surface locks into the tibial baseplate and is contained by a peripheral rail. The articular surface is explantable without removing the tibial base plate. The NexGen Complete Knee System ,integrated, with an extensive cruciate retaining, cruciate substituting, and fully constrained component configurations, design specific, conforming surfaces, femoral and tibial augmentation. The femoral, patellar, articular surface, and tibial base plate components should be design specific to the condition and availability of the bone stock, and the condition and functionality of soft tissues. This allows the surgeon to select the most appropriate component combination to meet specific patient needs. Articulating surfaces on micro femoral and tibial components have a smaller bearing spacing than those on the standard and macro components. | Nexgen Complete Knee System | 2 | May 14, 2012 | 2012 | Zimmer Inc. | Zimmer has received 5 complaints where a NexGen CR micro articular surface was used with a standard CR femur, even though the compatibility chart and product labeling indicates that these components are NOT compatible. Investigation identified additional instances where a CR micro articular surface and standard CR femur were sold together on the same purchase order. CR micro articular surfaces are | DESIGN: Labeling Design | Zimmer sent an Urgent Correction Notice dated March 5, 2012, to all affected customers via telephone, E-mail and letter notifying users of the appropriate compatibility of micro sizes of NexGenÂŽ CR and CR-Flex Articular Surfaces and informing them of known cases of incompatible devices being implanted, including associated risks. Customers and surgeons were asked to acknowledge receipt by calling 1-888-912-7349. Questions and concerns or help in notiying accounts about this correction were to be addressed to 1-888-548-8514.For questions regarding this recall call 574-372-4807. |  | 83322 | 83,322 all sizes | Worldwide Distribution - USA (nationwide) and the countries of Poland, Taiwan, Denmark, Korea, Singapore, China, Canada, Australia, Japan, France, Italy, UK, Germany, Spain, Belgium, Portugal, Switzerland, Sweden, India, Thailand, Peru, The Netherlands, Malaysia, and Austria. | Open |  |  | 83322 | Not_Computer
Z-1543-2012 | 61699 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PATIENT SPECIFIC INSTRUMENTSLPS-FLEX PIN GUIDES (FEMUR AND TIBIA)NONSTERILE ~ MUST BE STERILIZED PRIOR TO USEPatient ID--0242-LFManufactured by Materialise, Plymouth, MI Distributed by Zimmer, Warsaw IN.Patient specific instrument is intended to be used as a surgical instrument to assist in the positioning of Total Knee Replacement components intraoperatively and in guiding the marking of bone before cutting provided that anatomic landmarks necessary for alignment and positioning of the implant are identifiable on patient image scans. | LPSFlex Pin Guides | 2 | May 14, 2012 | 2012 | Materialise USA LLC | March 23, 2012 a field representative reported patient specific -0242-L tibia guide case packaging contained a different patient specific -0864-R tibia guide. The two patient specific products had been switched in packaging. Neither products had been delivered to the end user. | PRODUCTION CONTROLS: Labeling Mix-Ups | Materialise telephoned both field representatives on March 23, 2012 and confirmed recalled product mix-up. Materialise requested both cases be returned and replacements were provided. No affected product was received by the end user. | Contact the recalling firm for information | 1 | 1 unit | Nationwide Distribution including the states of Ohio and Minnesota. | Terminated | January 23, 2013 | 254 | 1 | Not_Computer
Z-1546-2012 | 61500 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | RF IgA Calibrator, a component of the EL-RF/3 Rheumatoid Factor Test Kit; an in vitro diagnostic; TheraTest Laboratories, Lombard, IL 60148; component catalog #12833, kit catalog #303-305. The TheraTest EL-RF/3 is intended as an aid to the diagnosis of Rheumatoid Arthritis (RA). | RF IgA Calibrator | 2 | May 14, 2012 | 2012 | Thera Test Laboratories, Inc. | The RF IgA Calibrator component of the EL-RF/3 Rheumatoid Factor test system may be unstable, resulting in a drop in the RF IgA calibrator OD value which could invalidate an RF IgA test run. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | TheraTest sent "Urgent Device Recall" letters dated March 28, 2012 to customers that received the affected kit, requesting them to return or destroy the remainder of the product. TheraTest issued a replacement kit to all customers that had unused product as well as any material they consumed which resulted in a failed run. Customers were also instructed to complete the enclosed confirmation form and fax the completed form back to TheraTest. Any questions were directed to TheraTest at 800-441-0771. | Contact the recalling firm for information | 207 | 207 kits | Nationwide Distribution (US) | Open |  |  | 207 | Not_Computer
Z-1542-2012 | 61475 | exerciser, non-measuring | ION | Nonmeasuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Shoulder PulleyDistributed by North Coast Medical, Gilroy, CA.Used to improve range of shoulder movement and coordination following a stroke, shoulder injury or surgery. | Norco Shoulder Pulley | 2 | May 14, 2012 | 2012 | North Coast Medical Inc | Potential for plastic D-ring to break during use and cause injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | North Coast Medical began contacting customers of the recall on July 13, 2011 by telephone. A phone script was used to inform customers of the affected product, problem, and actions to be taken by the customers. Information was provided for customers to return affected product. Affected devices were re-worked and inspected. Contact the firm by e-mail at www.ncmedical.com for questions regarding this recall. |  | 2364 | 2364 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Switzerland, Germany, Great Britain and the Netherlands. | Terminated | September 06, 2012 | 115 | 2364 | Not_Computer
Z-1545-2012 | 61804 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Instructions For Use: "***SmartLOX***Instructions for Use Cervical Plate System***STERILIZATION***The SmartLOX Cervical Plate System components are supplied clean and not sterile. All implants and instruments should be cleaned and sterilized prior to surgery. AORN recommended practices for in hospital sterilization should be followed. Sterilization testing of components has shown the following recommendations for sterilization are effective Method: Steam***Cycle: Pre vacuum***Temperature 270 F (132 C)***Exposure Time: 3 minutes***Contact Information: Captiva Spine Inc. 967 Alternate AA #1 Jupiter, FL 33477. Tel: 877-772-5571***CAPTIVA SPINE***".Intended for anterior screw fixation of the cervical spine. These implants have been designed to provide stabilization as an adjunct to cervical fusion. | Spinal Intervertebral body Fixation Orthosis | 2 | May 14, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine has initiated a correction of the SmartLOX Instructions for Use (IFU) for SmartLOX Cervical Plate System. The pre vacuum sterilization cycle [270 (132 C) for 3 minutes] has not been validated. | MISBRANDING: Labeling False and Misleading | Captiva Spine sent a "VOLUNTARY CORRECTION NOTIFICATION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides steps for the correction and a verification form for customers to complete. Customers should contact the firm at 877-772-5571 for questions concerning this notice. | Contact the recalling firm for information | 14 | 14 units | Nationwide Distribution including the states of FL, IN, UT, GA, PA, TX, AL, NC, MI, and MN. | Terminated | November 18, 2013 | 553 | 14 | Not_Computer
Z-1612-2012 | 61418 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical Naturalyte Acid Concentrate for bicarbonate Dialysis 45x , 55 gallons (208.2 Liters)Cat. No. 13-2251-0.NaturaLyte liquid acid concentrate is formulated to be used with a three stream hemodialysis machine which is calibrated for acid and bicarbonate concentrates. | Naturalyte Acid Concentrate | 2 | May 15, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Sodium content does not meet product requirements | PRODUCTION CONTROLS: Process Control | Fresenius Medical Care North America sent an "URGENT PRODUCT RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine whether they have any affected product. A fax back form was included for customers to complete and return. Contact FMCNA Customer Service Team at 1-800-323-5188 for questions regarding this notice. |  | 154 | 154 units | Nationwide Distribution. | Terminated | January 28, 2014 | 623 | 154 | Not_Computer
Z-1614-2012 | 59861 | sealant,polymerizing | NBE |  |  | General & Plastic Surgery |  | Coseal Surgical Sealant.Coseal is an adjunctive surgical sealant. | Coseal Surgical Sealant. | 2 | May 16, 2012 | 2012 | Baxter Healthcare Services L.P. | The recall is being conducted as a precautionary measure due to out of specification results observed at the 24 months time point (out of 24-month licensed shelf life to expiration) during a Coseal stability study. The parameter that is out specification is an indicator of possible failure for the product to gel appropriately . Coseal's failure to gel does not represent risk for the patient's life |  | Baxter Healthcare sent an"URGENT PRODUCT RECALL" letter dated July 27, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop the usage/distribution of the affected product and to quarantine any units that are in inventory. Instructions for obtaining replacement product were attached to the letter. Contact the firm at 1-800-423-2090 for questions concerning this recall. | No consumer action necessary | 11160 | 11160 | Nationwide Distribution. | Terminated | May 31, 2012 | 15 | 11160 | Not_Computer
Z-1613-2012 | 61657 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Galil Medical, 17G Cryoablation Needle Holder FPRPR 4014, Rx only, Sterile EO.Product Usage:Galil Medicals 17G Cryoablation Needle Holder (Needle Holder) is intended for use only with a Galil Medical Needle Testing Device (NTD). The NTD, a component used when performing cryogenic destruction of tissue with a Galil Medical Cryoablation System, is intended for performing a Needle Integrity and Functionality Test in preparation for a cryoablation procedure, and also serves to organize and support the cryoablation needles and thermal sensors in a sterile environment, prior to use. | 17G Cryoablation Needle Holder | 2 | May 16, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling 17G Cryoablation Needle Holder, FPRPR 4014. This voluntary recall has been initiated due to the possibility of excessive bioburden on the Needle Holders. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Galil Medical sent an "Urgent Medical Device Recall" letter dated April 12, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken. The letter instructs customers to return all unused affected products, use a revised needle preparation procedure and forward and communicate this Recall Notice and instructions. Customers were instructed to complete and return the attached Response Form with returned product or fax to 1+877-510-7757. For any questions contact Galil Medical Customer Service. | Contact the recalling firm for information | 26 | 26 kits | Worldwide Distribution - USA (nationwide) including the states of CA, NY, NC, WA and the countries of CANADA, DENMARK, FRANCE, NETHERLANDS, TAIWAN and UNITED KINGDOM. | Terminated | August 16, 2012 | 92 | 26 | Not_Computer
Z-1621-2012 | 61819 | barrier, animal source, intraoral | NPL | Bone grafting material. | Dental | Dental | 510(k) | a translucent, rectangular, resorbable, collagen membrane sheet derived from bovine tissue.Colla Guide Collagen Membrane Distributed by: Riemser Inc., PO Box 12339 Research Triangle Park, NC 27709-2339 USA Kensey Nash Corporation, 735 Pennsylvania Drive, Exton, PA 19341 USA QTY 1 Sterile Rx only15mm X 20mmQuantity 1 STERILE EO Sterilized using ethylene oxide Do not store product above 30ÂŽC (86ÂŽF) Do not re-use.CollaGuide Collagen Dental Membrane is indicated for simultaneous use of guided bone regeneration (GBR)-membrane and implants, augmentation around implants placed in immediate extraction sites, augmentation around implants placed in delayed extraction sockets, localized ridge augmentation for later implantation, alveolar ridge reconstruction for prosthetic treatment, filling of bone defects after root resection, cystectomy, removal of retained teeth, guided bone regeneration in dehiscence defects and guided tissue regeneration procedures in periodontal defects. | CollaGuide Collagen Dental Membrane | 2 | May 17, 2012 | 2012 | Kensey Nash Corp | Recall has been initiated due to concerns regarding the sterility of the product. The firm initiated the recall because of use of inadequately sterilized product might result in patient infection, | TRAINING: Employee Error | Reimser, Inc. sent an "URGENT PRODUCT RECALL-COLLAGUIDE COLLAGEN DENTAL MEMBRANE" letter dated March 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to cease use, secure and quarantine affect product. Additionally, a Recall Acknowledgement Form was attached for customers to complete and return to the firm via fax to 1-800-505-4982. Contact the firm at 484-713-2197 for questions regarding this recall. | No consumer action necessary | 6387 | Total all sizes: 6387 units; p/n 20650-01 - 1487 units | Worldwide Distribution. | Terminated | April 11, 2013 | 329 | 6387 | Not_Computer
Z-1618-2012 | 61715 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cutting Edge Acetabular Spherical Reamer 44 M, Stryker Howmedica Osteonics Howmedica Osteonics Corp. 325 Corporate Drive Mahwah, NJ 07430 USA.The CuttingEdge acetabular reamers are used to ream the acetabulum and prepare the bone for the insertion of the shell during hip arthroplasty. Surgical procedures for shells used with CuttingEdge instruments indicate that reaming should begin with a Spherical Reamer that is 4mm smaller than the templated or gauged size of the planned shell. Per the procedures, reaming should increase in size by 1mm increments until final sizing is achieved. Surgical judgment is used to assess bone stock, bone quality, desired level of press fit, and proper amount of reaming. | Cutting Edge Acetabular Spherical Reamer | 2 | May 17, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | A 46mm acetabular reamer was lasermarked as a 44m reamer and distributed to the field. | DESIGN: Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 29, 2011 to all affected customers. Notificationletters and Product Accountability Forms were sent to Stryker accounts via FedEx on January 3, 2012 with return receipt.Hospital Risk Management and Chief of Orthopaedics Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Surgeon Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Notification sent via e-mail toCanada.EMEA, LatinAmerica.Asia Pacific on December 23, 2011.The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to return the affected product. A Product Recall Acknowledgment Form was attached to the letter for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. |  | 134 | 134 units | Worldwide Distribution. | Terminated | July 17, 2013 | 426 | 134 | Not_Computer
Z-1624-2012 | 61801 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Pelleve Non-Ablative Wrinkle Treatment Handpiece, Product Codes D110FSHP, D115FSHP (reusable) --- Pelleve --- Elegance through science --- Contents: One (1) Pelleve Handpiece (Reusable) --- Non-ablative treatment of mild to moderate facial wrinkles and rhytids for skin phototypes I-IV. | Pelleve | 2 | May 17, 2012 | 2012 | Ellman International, Inc. | Ellman International is recalling its Pelleve reusable handpiece because of possible cable failure and sparking. | PRODUCTION CONTROLS: Equipment Maintenance | Ellman will notify all consignees initially via UPS shipment with an Urgent: Medical Device Recall letter containing notification of the recall and instructions. The letter identifies the affected product, problem, and instructs customers to identify and return the handpieces in the enclosed prepaid mailer. Customers are to complete and return the Product Recall Verification Form and fax it to the number provided. In addition, customers are provided with a toll-free number, 1-877-522-7141, to call in the event that they have questions or require additional instruction. Customers will initially be provided with a fixed-life disposable GlideSafe handpiece and may choose to receive a replacement reusable handpiece with modified (more robust) cable, or additional GlideSafe handpieces. |  | 1088 | 1,088 units | Worldwide Distribution | Open |  |  | 1088 | Not_Computer
Z-1619-2012 | 61343 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Alignment HandleTriathlon InstrumentsCAT #: 6541-2-807Nav Tibial Alignment Handle Triathlon InstrumentsCAT #; 6541-2-808Howmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430The tibial alignment handles Cat# 6451-2-807 and 6541-2-808 are part of the Triathlon Knee Instruments set. The alignment handle is used after the tibial resection portion of a total knee arthroplasty procedure. The alignment handle helps holding and orienting the tibial template and trial, which are used to estimate the size and assess the orientation of the final implant. The implant will provide the new bearing surface of the knee, replacing the damaged joint. | Triathlon Navigation Tibial Alignment Handle | 2 | May 17, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker received a Product Experience Report (PER) that a swivel block was improperly oriented on a Triathlon Tibial Alignment Handle, catalog no. 6541-2-807. | DESIGN: Device Design | Stryker sent an Urgent Product Field Action letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product . Customers were asked to fax back the attached Poroduct Field Action Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-3523. |  | 6126 | 6,126 Units | Worldwide Distribution - USA (nationwide) and the countries of Asia Pacific, Canada, Europe, Middle East and Africa (EMEA), and Japan | Open |  |  | 6126 | Not_Computer
Z-1625-2012 | 61859 | catheter (gastric, colonic, etc.), irrigation and aspiration | KDH | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 50 Portal Vein Cannula REF 61.110.10.030 Size: Fr 10 Length: 12 Inches LL Connector LOT 120493 NON STERILE MEDDEV USA, INC. 100 Highland Park Village/Suite 200 DALLAS, TEXAS 75205For use in humans. No indication of use is on the label. However, the device can be used for infusion. | 50 Portal Vein Cannula | 2 | May 17, 2012 | 2012 | MEDDEV USA, INC. | The device was marketed without a 510(k) and is not sterile. |  | MEDDEV USA, Inc decided to recall and notified their customer by phone on 05/10/12. Customers were instructed to return the entire shipment to MEDDEV .For questions regarding this recall call 214-692-3559. | Contact the recalling firm for information | 20 | 20 units | Nationwide Distribution including Illinois | Terminated | August 15, 2013 | 455 | 20 | Not_Computer
Z-1626-2012 | 61860 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 8F Plastic Dignity Mid-Sized CT Port w/Silicone Filled Suture Holes and w/Attachable ChronoFlex Polyurethane Catheter. Catalog # MRCTI8000S. The Power Injectable, Implantable Infusion Port is an implantable single fluid reservoir port with a polyurethane (8F) catheter attachable for application by the inserting physician. | 8F Plastic Dignity MidSized CT Port | 2 | May 18, 2012 | 2012 | Medical Components, Inc dba MedComp | 50 trays of Dignity Low Profile CT Port were labeled as Dignity Mid-Sized CT Port. | TRAINING: Employee Error | On April 12, 2012, MedComp sales representatives and the foreign distributor were notified via e-mail with a "Product Alert" notification. The issue was described and requested actions were provided to recipients. |  | 50 | 50 trays | Worldwide Distribution - - USA, including the states of CA, NJ, AZ, PA, TN and the country of Spain. | Terminated | July 18, 2013 | 426 | 50 | Not_Computer
Z-1627-2012 | 61750 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The ECHELON 60 Endoscopic Linear Cutter - Straight, compact ECHELON 60 Endoscopic Linear Cutter - Straight, and the Long ECHELON 60 Endoscopic Linear Cutter - Straight are sterile, single patient use instruments that deliver staples while simultaneously dividing tissue between rows. ECHELON 60 instruments with white, blue, gold, and green reloads deliver two triple-staggered rows of staples. The instruments have a staple line that is approximately 60 mm long and a cut line that is approximately 57 mm long. A staple retaining cap on the reload protects the staple leg points duringshipping and transportation. Sterile products are individually packed in a transparent tray sealed with TyvekÂŽ. Each tray is placed in an individual carton. 3 individual cartons are placed in a corrugated shipper.The ECHELON 60 Endoscopic Linear Cutter  Straight, Compact ECHELON 60 Endoscopic Linear Cutter  Straight, and the Long ECHELON 60 Endoscopic Linear Cutter  Straight are intended for transection, resection, and/or creation of anastomoses. The instruments have application in multiple open or minimally invasive general, gynecologic, urologic, thoracic, and pediatric surgical procedures. They can be used with staple line or tissue buttressing materials. | The ECHELON 60 Endoscopic Linear Cutter | 2 | May 18, 2012 | 2012 | Ethicon Endo-Surgery Inc | There is a possibility of a component breakage, which may cause the device to stay locked on tissue after firing. If the component exhibits this issue, the device does not open either with the anvil release button or with the manual firing release lever. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated March 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have the affected product and to remove the affected product. Fill out the Business Reply Form and fax it to 1-513-337-4138 within 3 business days, even if they do not have the affected product in their inventory. Ethicon will automatically replace any affected product returned by June 29, 2012. Customers were asked to share this information with appropriate staff.If customers should need clinical or product support call 1-800-873-3636, option 6, or contact their local sales representative.For questions regarding this recall call 513-337-7000. |  | 3120 | 3120 | Worldwide Distribution - USA including AL, AR, CA, CO, FL, IA, IL, KS, LA, MD, MO, MS, NC, NM, NY, OH, SC, TX & VA; and the countries of Argentina, Ecuador, Mexico, Australia, Egypt, Singapore, Belgium, Greece, Slovenia, Brazil, India, South Africa, China, Israel, United Arab Emirates, Colombia, Japan & Venezuela. | Terminated | March 22, 2013 | 308 | 3120 | Not_Computer
Z-1629-2012 | 61444 | processor, cervical cytology slide, automated | MKQ |  |  | Pathology |  | BD Focal Point"100V, Catalog number: 490650The BD Focal Point" Slide Profiler (formerly the AutoPapÂŽÂŽ System) is an automated cervical cytology screening device intended for use in initial screening of cervical cytology slide. | BD FocalPoint" | 3 | May 21, 2012 | 2012 | Tripath Imaging, Inc. | The cable tie mounts that adhere to the inside of the FocalPoint do not adhere as intended and can therefore become detached. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | BD Diagnostics first notified consignees by letter on October 6, 2011. BD Diagnostics also sent a letter dated February 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Field Representatives were advised to inspect all cables, adhesive cable ties and connections to ensure they are secure. Inspections and repairs, if necessary, would be carried out at the next service call for each instrument, or no later than the next regularly scheduled preventative maintenance visit. For questions consignees were instructed to contact BD Diagnostics Technical Support at 1-877-822-7771 option 2 in the US, their local distributor, or their local BD office for any questions.For questions regarding this recall call 918-563-1624. |  | 71 | 71 units | Worldwide Distribution - USA including AZ, CT, IL, IN, FL, MN, NC, OK, OR, and TX and the countries of Australia, Belgium, Brazil, Canada, Japan and Thailand | Terminated | December 03, 2012 | 196 | 85 | Not_Computer
Z-1628-2012 | 61663 | tube, tracheostomy (w/wo connector) | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone Cuffed Wire Reinforced Extended Connect Pediatric Perfect Fit Neck Flange Tracheostomy Tube, Size 4.0 mm I.D. Catalog Number 353-040.Intended for use in providing direct tracheal access for airway management. | Tracheostomy tubes | 2 | May 21, 2012 | 2012 | Arcadia Medical Corporation | A defect in 3 tracheostomy tubes. The defect observed was a separation in the bond between the proximal junction of the neck flange and the cannula. | PRODUCTION CONTROLS: Process Control | Arcadia notified the 3 consignees with a MEDICAL DEVICE VOLUNTARY RECALL notification dated 4/18/12. The notification identified the affected product, along with the reason for the recall. Consignees were instructed to cease distribution/use of any unopened product that they have remaining and return to firm following the steps provided. Arcadia will replace the product at no charge to the customer. Consignees were to confirm receipt of the recall notification by emailing Trena Depel, trena@arcadiamedical.com.Phone: 925-918-0595Fax: (415) 762-5489. |  | 1103307 | 01103307 (50 units) and 01107055 (73 units) | Worldwide Distribution - US, including Mariemount, OH and countries of UK and France. | Terminated | August 13, 2012 | 84 | 123 | Not_Computer
Z-1620-2012 | 61568 | extractable antinuclear antibody, antigen and control | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Product is labeled in part: "***is immunosimplicity***ENA-6 Screen***Enzyme Immunoassay Test Kit***Catalog No: 720-230***Exp. Date: 2012-05-31***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: S0000***For In Vitro Diagnostic Use***Caution, consult accompanying documents***Manufacturer DIAMEDIX***a Subsidiary of IVAX Diagnostics, Inc.***Miami, FL 33127 USA-1-800-327-4565***Biohazard***"For the qualitative screening of human IgG antibodies to extractable nuclear antigens (ENA) in human serum by indirect enzyme immunoassay as an aid in the diagnosis of certain autoimmune disorders. This test system screens for antibodies to Sm, Sm/RNP, SSA, SSB, Scl-70 and Jo-1 in one well. Positive samples should be evaluated further using tests designed for each ENA antibody. | ENA6 Screen Enzyme Immunoassay Test Kit | 2 | May 21, 2012 | 2012 | Diamedix Corporation | On 06/08/2011 Diamedix Corporation initiated a recall of Is-ENA-6 Screen Test Kit, Catalog # 720-320 Master Lot S0000, Exp. Date May 31, 2012. The labeling on the kit box has an expiration date of 2012-05-31 while the components have an expiration date of 2012-04-30 for an incorrect expiration date. | PRODUCTION CONTROLS: Error in Labeling | Diamedix Corporation sent a Customer notification letter dated June 8, 2011, to all affected customers via Certified Mail. Customers were instructed to search their inventory and to return affected product to Diamedix.Diamedix apologizes for any inconvenience this has caused our valued customers. Please contact the Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any kits you might still have of the lot in question. Please rest assured that Diamedix values you as a customer and is in the process of taking the proper corrective and preventive actions. |  | 312 | 312 Test Kits | Worldwide Distribution -- (USA) nationwide including the states of FL, NY, NC, MS, KY, MA and Puerto Rico and the countries of Italy, Chili and Venezuela. | Terminated | November 15, 2013 | 543 | 312 | Not_Computer
Z-1647-2012 | 48179 | catheter, electrode recording, or probe, electrode recording | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | EZ STEER Catheter D-1267-04-S / 8mm Non-Navigation Bi-Directional CatheterCardiovascular catheter for use in electrophysiological diagnosis and ablation. | EZ STEER Catheter | 2 | May 22, 2012 | 2012 | Biosense Webster, Inc. | The recall was initiated because the catheter unit was shipped with a label expiration date of 9/30/2010, which is 11 months beyond the correct expiration date of 10/31/2009. |  | The firm, Biosense Webster, sent a "Medical Device-Voluntary Field Removal" letter dated December 17, 2007 to its customer. The letter identified the product, problem and an action to be taken. The customer was instructed to complete and return the Product Voluntary Field Removal Certification Form via fax to Quality Compliance at (909) 839-7207or mail using the enclosed self-addressed stamped envelope. If you have additional questions, please follow up with your local Biosense Webster representative.or contact Customer Service at (909) 839-8500. | Contact the recalling firm for information | 1 | 1 unit | Nationwide distribution: CA only. | Terminated | May 23, 2012 | 1 | 1 | Not_Computer
Z-1645-2012 | 61770 | test,prostate specific antigen,free,(noncomplexed) to distinguish prostate cancer from benign conditions | MTG |  |  | Immunology | PMA | Access Hybritech Free PSA, Part Number: 37210Free PSA is used to aid in the differentiation between prostate cancer and benign prostatic conditions when total PSA levels are between 4 and 10 ng/mL. | Access Hybritech Free PSA | 3 | May 22, 2012 | 2012 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has confirmed a change to the recommended diluents for the Access Hybritech free PSA reagent. | DESIGN: Process Design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed regarding a change to the recommended diluents for the Access Hybritech free PSA reagent. Patient safety is not affected by this change." Do not dilute elevated free PSA patient samples with Wash Buffer II." Do not dilute elevated free PSA patient samples using the UniCel DxI onboard dilution feature." Dilute elevated free PSA samples with Access Hybritech free PSA Calibrator S0 only." A review of historical patient results is not required because there is no effect on clinical interpretation.* The UniCel DxI onboard dilution feature is only available in limited geographies." An updated version of the UniCel DxI Assay Protocol File (APF) will remove the Access Hybritech free PSA (dfPSA) onboard dilution assay from the list of available assays." The Access Hybritech free PSA IFU will be updated to remove Wash Buffer II as an acceptable diluent for elevated patient samples. Correct any procedure manuals or other documents created in your laboratory that list Wash Buffer II as an acceptable diluent for the Access Hybritech free PSA assay.Please share this information with your laboratory staff. If you have forwarded any of the affected product listed above to another laboratory, please provide a copy of the letter to them.Please complete and return the enclosed response form within 10 days so we are certain you received this important communication.If you have any questions regarding this notification, please contact our Customer Support Center" Via our website, http://www.beckmancoulter.com/customersupport/support" Via phone, call 1-800-854-3633 in the United States and CanadaOutside the United States and Canada, contact your local Beckman Coulter Representative. We apologize for the inconvenie. | Contact the recalling firm for information | 90965 | 90,965 units total (24,951 units in US) | Worldwide Distribution -- USA (nationwide ) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia, Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French, Polynesia, Georgia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Libyan, Arab, Jamahiriya, Macao, Mexico, Moldova, Republic of Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, Emirates, United Kingdom, Uruguay ,Venezuela, Viet Nam, Yemen and Zimbabwe. | Terminated | March 18, 2013 | 300 | 90965 | Not_Computer
Z-1651-2012 | 61812 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MLCi/MCLi2.Radiation therapy. | Linear Accelerator | 2 | May 23, 2012 | 2012 | Elekta, Inc. | The Y2 diaphragm block was found loose. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to carry out the recommended weekly physics check of the X-ray and light field size calibration. If an error is found, customers are instructed to contact their local Elekta service engineer. |  | 386 | 386 | Nationwide Distribution including Puerto Rico. | Open |  |  | 386 | Not_Computer
Z-1648-2012 | 61811 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine SmartLOX Cervical Plate System. Products are labeled in part: "***Part Number CP40012***12mm Drill Bit***Lot Number 89427-2 R***" "***Part Number CP40014***14mm Drill Bit***Lot Number 89427-3-R***" The Captiva Spine SmartLOX Cervical Plate System is intended for the anterior screw fixation of the cervical spine. These implants have been designed to provide stabilization as an adjunct to cervical fusion. | Drill Bits for Cervical Plate System | 2 | May 23, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine is recalling the SmartLOX Cervical Plate components 12 & 14 mm Drill Bits, Single Barrel Variable Drill Guides, and Double Barrel Fixed Drill Guides. The components are part of the instrumentation kit used to complete an anterior screw fixation of the cervical spine. There is a potential risk for the drills to not properly fit through the drill guides due to an overlap in the dime | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Captiva Spine notified customers of the recall by letter on November 23, 2011, and advised that the affected products would be replaced. Customers were asked to return the drill guides and unused drills with the provided return labels.For questions customers should call 877-772-5571.For questions regarding this recall call 561-277-9480. | Contact the recalling firm for information | 212 | 212 units | Nationwide Distribution including TX, IN, UT, AZ, GA, CA, AL, MI, and MN. | Terminated | November 18, 2013 | 544 | 317 | Not_Computer
Z-1650-2012 | 61701 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Celsus prosthetic foot, Product Number 103658-00. The Celsus is intended for low impact level lower limb prosthesis users weighing up to 300 lbs. | Celsus Foot | 2 | May 23, 2012 | 2012 | College Park Industries, Inc. | An incorrect shorter part ( pyramid) was installed in the foot. A foot of this size and configuration requires a longer pyramid. Customers requiring the extra firm configuration may have a minimal risk of component failure, which could be implicated in a potential fall. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | All customers who received affected product were initially contacted via phone on April 5, 2012 and made aware of the recall, including instructions to notify customers immediately and arrange replacements. A "Recall Notice" letter dated April 24, 2012 was sent to customers describing the problem and the recommended actions. Contact reference person is Kimberly Light, 1 800-728-7950, info@college-park.com. | Contact the recalling firm for information | 11 | 11 | Worldwide Distribution - USA, including the states of CA, MA, ME, WA, TX, VA, AZ and the country of England. | Terminated | July 31, 2012 | 69 | 11 | Not_Computer
Z-1653-2012 | 61732 | media,coupling,ultrasound | MUI | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Other-Sonic Generic Ultrasound Transmission Gel 5 Liters HypoallergenicProduct Usage is a non-sterile gel used in ultrasound procedures to improve the transmission of the ultrasound signal from the transducer to the body. | OtherSonic Generic Ultrasound Transmission Gel | 1 | May 24, 2012 | 2012 | Pharmaceutical Innovations, Inc. | Other-Sonic GENERIC ULTRASOUND TRANSMISSION GEL is being recalled due to possible microbial contamination. An FDA Safety Communication was issued for this product on April 18, 2012. | TRAINING: Employee Error | Phamaceutical Innovations Inc. sent an Urgent Message Notice letter dated March 02, 2012 with a followup letter dated March 9, 2012 and an Inventory Withdrawal letter dated March 23, 2012 via email and fax to all affected customers. Customers were instructed to examinne inventory, refrain from distribution and return remaining inventory for a replacement. For questions call 973-242-2900. | Contact the recalling firm for information | 40164 | 40,164 units in distribution worldwide - unable to obtain units in distribution Nationwide | Worldwide Distribution - United States (nationwide) including the states of: FL, MA, MS, NJ, OH, TN and VA, and the countries of: Canada and Australia. | Open |  |  | 40164 | Not_Computer
Z-1654-2012 | 61621 | pump, blood, extra-luminal | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PlasmaPrO"Plasma Pump Model PP-05Product Usage:The PlasmaPro Plasma Pump is a pump to be used with Asahi Plasma Separators to perform therapeutic plasma exchange. | PlasmaPro Plasma Pump | 2 | May 24, 2012 | 2012 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc., Palm Harbor, FL, initiated a voluntarily withdrawal on November 29, 2010 of their five (5) units of the Plasma Pro Plasma Pump and later classified as a recall (removal) rather than a withdrawal, Model # PP-05, Serial Numbers: E08003041, E08003042, E08003043, E08003044 and E08003045. | PREMARKET APPROVAL: No Marketing Application | Apheresis Technologies sent a Letter by FedEx to their two consignees. The letter notified customers of the voluntary recall initiated, affected product and description of the problem. For questions call 727-784-2353 or 1-800-382-5556. | Contact the recalling firm for information | 5 | 5 units | US Nationwide Distribution including the states of Colorado and Utah. | Terminated | May 25, 2012 | 1 | 5 | Not_Computer
Z-1655-2012 | 61825 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, GuideLiner V2 Catheter, Rx only, Sterile EO, Catalog Number 5571, 6F, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. Intended to be used in conjunction with guide catheters to access discrete regions of the coronary and or peripheral vasculature and too facilitate placement and exchange of guidewired and other interventional devices. | GuideLiner V2 Catheter | 2 | May 24, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions has became aware of a potential defect with their 6F GuideLiner V2 catheters (Model 5571) with lot number 555038. Use of this product could result in the catheter shaft separating from the pushwire on the device. This may cause a procedural delay and/or injury possibly requiring medical intervention. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Vascular Solutions sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product , immediately remove the product and place the product in a secure area. Additionally, customers were asked to complete and return a Field Action Customer Inventory Form to the Customer Service Department via fax at 763-656-4251 or e-mail at customerservice@vasc.com. Customers may contact their local sales representative for questions concerning this recall. | Contact the recalling firm for information | 330 | 330 | Nationwide Distribution including the states of AL, AZ, AR, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LAM ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, WV, and WI. | Terminated | July 19, 2012 | 56 | 330 | Not_Computer
Z-1673-2012 | 61540 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | HPBLUE: HARMONICÂŽ Blue Hand Piece; HP054: HARMONICÂŽ Hand Piece; Non-Sterile products are individually packed in a poly bag and then placed into an individual carton. Mfr. By: Ethicon Endo-Surgery, LLc, 475 Calle C, Guaynbo, PR 00969.Indicated for soft tissue incisions when bleeding control and minimal thermal injury are desired. The instruments can be used as an adjunct to or substitute for electrosurgery, lasers, and steel scalpels. | HARMONIC Blue Hand Piece & HARMONIC Hand Piece | 2 | May 25, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary recall due to some lot numbers of the HARMONICÂż Hand Piece (HP054) and the HARMONICÂż Blue Hand Piece (HPBLUE), which impacts their compatibility with the HARMONICÂż Generator 300 (GEN04). The problem occurs after successful set up and initial use of HARMONICÂż Hand Pieces (HP054 and HPBLUE) with the Generator 300 (GEN04). The procedure is interrupted when | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Ethicon Endo-Surgery Inc sent an URGENT: Product Correction Notification letter dated March 2012 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Labels were enclosed that are used to identify the recalled product. Customers were instructed to fill out the Business Reply Form and fax it to Stericycle at 1-888-912-7091 within three business days for a credit.For questions regarding this recall call 513-337-3419. |  | 1896 | 1896 instruments | Worldwide Distribution-USA including: AK,AL,AR,AZ,CA,CO,CT,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NCND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI,WV, and WY.and the countries of: Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Ecuador, Egypt, Greece, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Korea (South),Kuwait, Lebanon, Malaysia, Mexico, Netherlands, Antilles, Pakistan, Panama, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Taiwan, Thailand, Turkey, United Arab Emirates Uruguay, Venezuela & Yemen. | Terminated | March 05, 2013 | 284 | 1896 | Not_Computer
Z-1674-2012 | 61532 | irrigator, powered nasal | KMA | Powered nasal irrigator. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Water Pik Sinusense battery operated device including a blue semi-transparent water reservoir, a white hand held button initiated injector device, and a soft seal nozzle affixed to the white injector. The product is sold individually in white, blue red and orange cartons with a photo of the device as well as a stream of water flowing through it. The contents of the container include the device as well as a package insert with utilization instructions.Product Usage: Sinus washing device to keep sinus passages clear of congestion. | Waterpik | 2 | May 25, 2012 | 2012 | Water Pik, Inc. | Customer complaints related to cross contamination of sinus rinse solution between the battery compartment and solution reservoir used for rinsing nasal passages. | DESIGN: Device Design | Water Pik has generated a Press Release letter dated May 2012. The letter included safety information and instructions for the consumer to contact Waterpik for a replacement model if they experience water intrusion into the battery compartment of their device. For additional information go to www.waterpiksinusense.com/exchange. | Contact the recalling firm for information | 105257 | 105257 individual units | US Nationwide Distribution. | Terminated | August 07, 2012 | 74 | 105257 | Not_Computer
Z-1677-2012 | 61338 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration HA Reduced Neck 255 mm Long StemManufactured by Stryker Howmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430.The Restoration HA cementless Hip System is designed to address the problems of revision surgery where extensive fixation both distally and proximally is required for support of the prosthesis. A comprehensive size range mates a variety of proximal cross sections and lengths with 10 distal cylindrical diameters to provide for patient matched sizing in revision circumstances. The stem design incorporates a physiologic 127 degree neck stem angle, neck length ranges and a C-Taper head to provide the surgeon with the ability to restore near anatomic head position for proper leg length and biomechanical function. Restoration HA femoral component accepts C-Taper heads, 22mm, 26mm, 28mm, and 32mm with a negative 5mm to positive 10mm offsets. | Restoration HA Reduced Neck | 2 | May 29, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | An incorrect laser mark was noted on specific lots of product, incorrectly implying a stem size of 9, when the product is actually size 8. | PRODUCTION CONTROLS: Error in Labeling | Stryker sent an "URGENT PRODUCT RECALL" letter dated October 27, 2011 and October 28, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs the customers to return the product to the firm. An Acknowledgement Form was included for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. | Contact the recalling firm for information | 31 | 31 Units | Nationwide Distribution and the country of Canada. | Terminated | August 22, 2013 | 450 | 31 | Not_Computer
Z-1681-2012 | 61821 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion, Tissue Fusion Handpiece 5 mm O.D., 16 cm Shaft, Catalog Number 60-9520-001, Sterile/EO.The Altrus Thermal Tissue Fusion System is comprised of a dedicated energy source and disposable handpiece used to ligate (seal) and divide (cut) blood vessels and tissue bundles that fit into the jaws of the handpiece. | Altrus Thermal Tissue Fusion | 2 | May 29, 2012 | 2012 | ConMed Electrosurgery | The devices were distributed prior to the firm completing performance qualifications for their package sealing and adhesive bonding. |  | ConMed Electrosurgery sent an "URGENT: DEVICE RECALL" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem,and actions to be taken by the customers. The letter instructs customers to review their inventory and to complete and return a Business Reply Form via fax to 303-699-9854. | No consumer action necessary | 18 | 18 units | Class II Recall-Nationwide Distribution-including the states of OH, GA, IL, PA, AZ, KY, FL, NY, MO, TN, TX, AL, IN, SC, MD, CT, NJ, and CA. | Terminated | June 05, 2012 | 7 | 237 | Not_Computer
Z-1678-2012 | 61814 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Altrus Energy Source REF 60-9500-120 Tissue Fusion Energy Source ConMed Corporation 525 French Rd, Utica, NY 13502, MDSS GmbH, Germany. The Altrus Thermal Tissue Fusion System is comprised of a dedicated energy source and disposable handpiece used to ligate (seal) and divide (cut) blood vessels and tissue bundles that fit into the jaws of the handpiece. | Altrus Energy Source | 2 | May 29, 2012 | 2012 | ConMed Electrosurgery | An investigation of the ConMed Electrosurgery Altrus Energy Source revealed that the device could have been calibrated incorrectly resulting in error messages or delays in surgery. |  | ConMed Electrosurgery sent an "URGENT: MEDICAL DEVICE" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to review their inventory for the affected product and to complete and return a Business Reply Form to the firm via fax to 303-699-9854. |  | 122 | 122 units | Nationwide Distribution and the country of Canada. | Terminated | June 19, 2012 | 21 | 122 | Not_Computer
Z-1676-2012 | 61756 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SOLUS Anterior Lumbar Interbody Fusion SpinalSpacer System, Solus Proximal/Distal Deployment Wrenches, 510(k) K102402.The Alphatec Solus Anterior Lumbar Interbody Fusion (AUF) System is indicated for spinal fusion procedures in skeletally mature patients with degenerative disc disease (DOD) at one or two contiguous levels from L2-S1 with up to Grade 1 spondyiolisthesis or retrolisthesis at the involved level(s). DOD is defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies. These patients should be skeletally mature and have had six months of non-operative treatment. The Alphatec Solus implant is intended to be used with autograft. The device is intended for use with supplemental fixation that is in addition to the integrated blades. The Alphatec Solus implant is intended for use with supplemental spinal fixation. Specifically, the Alphatec Solus implant is to be used with the Alphatec's ZodiacÂŽ Spinal Fixation System, Aspida" Anterior Lumbar Plating System, or the ILLlCOÂŽ MIS Posterior Fixation System. Intervertebral body fusion device. | Alphatec Souls Proximal/Distal Deployment Wrench | 2 | May 29, 2012 | 2012 | Alphatec Spine, Inc. | Alphatec Spine has received several complaints (n+4) on the Alphatec Solus Distal Deployment Wrench for twisting or bending, which has been discovered either during or after instrumentation use. A fifth complaint received in which it was reported that the spline tip of the Alphatec Solus Proximal Deployment Wrench had sheared from the instrument upon use. | DESIGN: Device Design | Alphatec Spine sent an Urgent Medical Device Recall Notification dated March 30, 2012, to all affected customers. The notification letter informed customers of the removal/recall and included instructions on what to do with the recalled products. Initial letter was delivered via FedEx, and followed by telephone call with a total of 4 attempts. Customers were instructed to return all of the affected product. Upon receipt of of the Medical Device Recall Notification, customers were instructed to contact Alphatec Spine's Customer Service Department immediately for instructions on how to return the instruments and fill out the lower portion of this letter to confirm that they have read the notification and have taken all necessary removal actions: Customers should return a signed copy of the form to:Alphatec Spine, Inc.5818 El Camino Real,Carlsbad, CA 92008ATTN: Rand JadanForms may also be returned via email to rjadan@alphatecspine.com or fax to 760-431-0289.For questions regarding this recall call 760-431-9286. |  | 18 | 18 devices | Worldwide Distribution - USA (nationwide) and Germany | Terminated | June 20, 2012 | 22 | 18 | Not_Computer
Z-1672-2012 | 61895 | tubing, fluid delivery | FPK | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Non-DEHP Microbore Administrative Set with Female Luer Adapter Packaged Assembly, product code: 340-4176, Method of Sterilization using Ethylene Oxide.The CurlinÂŽ Ambulatory Infusion Pump System with disposable PVC Administration Tubing Sets and user accessories, provide a means for the volumetric delivery of fluids used in Parenteral, Enteral, Epidural, Subcutaneous and Intravenous applications. | Curlin Infusion Administration Set | 1 | May 30, 2012 | 2012 | MOOG Medical Devices Group | Moog Medical Device Group became aware on March 23, 2011 via customer complaint, of a reversed pump segment on an Administration Set. This malfunction could reverse the flow of fluid or medication backwards from what was intended. | DESIGN: Process Design | The firm, Moog Medical Device Group, sent an "Urgent Field Safety Notice" dated May 21, 2012 via certified letter and a press release direct to consignees/customers. The notice describes the product, problem and actions to be taken. The customers were instructed to do the following: If purchased by distributor -remove suspect product from inventory; contact your distributor to arrange replacement product and request a shipping return label as per the distributor's process and ship product to their location.If purchased by Moog-remove suspect product from inventory, contact Moog Customer Service at 1-800-970-2337, prompt #7 and Moog customer service will provide a call tag for convenient product return.For additional information or clarification, please contact the Post Market Surveillance Manager at (801) 264-1001, ext. 224 or via email at: lbrewer@moog.com. |  | 4080 | 4,080 sets. | Nationwide distribution: USA including states of: AZ, CA, FL, IL, KS, MA, MI, MD, NC, ND, NY, OK, OH, and TX. | Terminated | February 22, 2013 | 268 | 544900 | Not_Computer
Z-1683-2012 | 57371 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | Terumo BCT Lakewood, CO 80215, Catalog No 61000 Spectra Optia Apheresis System, Country of Origin: US.Product Usage:For use in humans. Within the US marketplace, this device is approved for use in Therapeutic Plasma Exchange. Internationally, this device is approved for use in Mono Nucleo Cell engraftment. | Terumo BCT | 2 | May 30, 2012 | 2012 | CaridianBCT, Inc. | Defective product may result in decreased efficacy. | DESIGN: Component Design/Selection | Terumo BCT sent a Voluntary Medcal Device Safety Alert letter dated May 18, 2012 to each of their consignees via certified letter. The letter identified the affected product, problem, actions to be taken and instructions for healthcare providers until a corrective action is identified and implemented. Customers were instructed to share this safety alert with all those who need to be aware and complete the attached Acknowledgement Form and fax to +1.303.876.9277 or email to Terumo BCT. For questions contact your Terumo BCT Representative, the Terumo BCT Support Center at (U.S. Toll Free +1.877.339.4228) or +1.303.231.4357, or your local Terum BCT Customer Service Office. |  | 372 | 372 units | US Nationwide Distribution and the country of Canada. | Terminated | November 27, 2012 | 181 | 372 | Not_Computer
Z-1688-2012 | 61858 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance iCT and iCT SP | 2 | May 30, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On 3/5/2012, Philips Healthcare was notified of one incident involving a Brilliance iCT scanner that resulted in damage to the gantry and slight injury to the user when the heat exchanger detached from the cooling unit due to some missing welds. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated May 8, 2012 to all affected customers. The letter identifies the product and problem. No action is required by the customers. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Philips Healthcare will inspect the cooling units on all iCT and iCT SP systems for this issue and if the welds on a cooling unit do not pass the inspection, Philips will replace the cooling unit free of charge. | No consumer action necessary | 372 | 372 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, VT, WA and WV and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Indonesia, Iraq, ISRAEL, Italy, Japan, Jordan, Korea, Lebanon, Mexico, Netherlands, New Zealand , Norway, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Arab Emirates and Uzbekistan. | Terminated | February 26, 2013 | 272 | 372 | Not_Computer
Z-1686-2012 | 61672 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E Series SPECT System; Single-Photon Emission Computed Tomography. The Siemens Symbia E system is a high performance dual detector variable angle system with flexible patient positioning for planar imaging, whole-body scans, and SPECT studies. | Symbia E Series SPECT System | 2 | May 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | Siemens has identified a potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. There is a remote chance that during the motor replacement the radial drive gear may have been replaced incorrectly. If the gear is installed incorrectly the mechanical brake may not engage creating a patient hazard if the sys |  | On May 15, 2012, Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated April 11, 2012 to all affected customers. The letters informed customers of the potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. Customers were informed that a Siemens Service Engineer will be inspecting the radial drive gear on their camera system to ensure it is functioning correctly. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 9 | 9 units | Worldwide Distribution, US (Nationwide) and the countries of Australia, Belgium, Canada, China, France, Germany, Italy, Japan, Republic of Korea, Mexico, Netherlands, Portugal, Spain, Taiwan and the United Kingdom. | Open |  |  | 146 | Not_Computer
Z-1690-2012 | 61762 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 - Model #728246, CT Scanning System is subject to recall/correctionProduct Usage:The Brilliance 6, 16, 40 64, Big Bore, iCT, iCT SP and Ingenuity CT are Computed Tomography X-Ray Systems that are intended to produce cross-sectional images of the body by computer reconstruction of X-Ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment , patient, and equipment supports, components and accessories. | Brilliance 16 Computed Tomography XRay Scanning System | 2 | May 30, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Component defect. Following their receipt of a customer complaint, Philips Medical determined that the metal blocks which support the scissor slide rails underneath the scanning table (couch) of several of the firm's CT scanning systems were not manufactured to specification and may crack. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips issued an 'URGENT Medical Device Correction' letter dated March 27, 2012 to all affected customers. The letter identified the affected products, problem, hazard involved, how to identify affected products, and actions to be taken by customer/user. The letter informed customers that Philips intends to dispatching service personnel to each customer site in order to correct this issue. In the meantime, the letter recommend customer to contact their Philips representative if they hear unusual noises when moving the couch vertically, or if there is unintended vertical motion of the couch. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompts. |  | 10 | 10 system units were distributed in the U.S. | Worldwide Distribution - U.S. (Nationwide) including the states of: CA, FL,, HI, IA, IN, KY, LA, MA,, MI, MO, MS. NM, NV, OH, PA, PA, SC, TN, TX, WA, and Puerto Rico | Terminated | February 22, 2013 | 268 | 35 | Not_Computer
Z-1694-2012 | 61908 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Monoject Insulin Safety Syringe 3/l0mL 29 X 1/2"Product ID: 8881511144Indications for Use: Injection of U-1 00 insulin | Monoject Insulin Safety Syringe 3/l0mL 29 X 1/2" | 2 | May 31, 2012 | 2012 | Covidien LLC | Safety shield may separate from the syringe, exposing the needle, which could potentially result in a needle stick | PRODUCTION CONTROLS: Equipment Maintenance | Covidien sent an Urgent Product Recall letter dated May 21, 2012, via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using any affected product. Distributors were requested to notify their customers. Customers were instructed to complete the Monoject Insulin Safety Syringe recalled product return form even if they have no product to return. The form should be faxed to 1-800-895-6140 or e-mailed to sdfeedback@covidien.com. Customer Service questions regarding the recall please use one of the following methods.Telephone:1-800-722-8772 option I (Weekdays, 8am to 6:30pm Eastern)e-mail:sdfeedback@covidien.com. |  | 952000 | 952,000 syringes | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Belgium, and Panama. | Terminated | June 10, 2013 | 375 | 28069500 | Not_Computer
Z-1698-2012 | 61623 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Transforminal Lumbar Interbody Fusion (TLIF) PEEK Inserter used in the Genesys Spine Apache" Interbody Fusion System.The TLIF PEEK inserter is used to introduce the TLIF PEEK implant into the patients spine by threading the implant onto the tip of the inserter shaft. Once the implant is seated within the vertebrae the inserter is removed by unthreading the inserter from the implant. | Transforminal Lumbar Interbody Fusion (TLIF) PEEK Inserter | 2 | June 01, 2012 | 2012 | Genesys Orthopedic Systems, LLC | The inserters have the potential to break. | PRODUCTION CONTROLS: Reprocessing Controls | Genesys Orthopedic Systems, LLC decided to recall the affected product and called their customers on August 24, 2011. The firm removed the inserters from distribution from 8/24-31/2011. | Contact the recalling firm for information | 4 | 4 inserters | Nationwide Distribution including Florida and Texas. | Terminated | December 04, 2012 | 186 | 4 | Not_Computer
Z-1701-2012 | 61755 | radioimmunoassay, immunoreactive insulin | CFP | Immunoreactive insulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000 --- Insulin --- Catalog Number: L2KIN2 (200 test), L2KIN6 (600 tests) --- Test Code: INS, Color: Orange --- Product Usage:For in vitro diagnostic use with IMMULITE and IMMULITE 1000 Systems Analyzers --- for the quantitative measurement of insulin in serum or heparinized plasma, for the management of diabetes. | IMMULITE 2000 | 2 | June 02, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed a negative bias in the Immulite Insulin assay. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Siemens sent an Urgent Device Recall letter dated April 2012 via fax to affected U.S. customers and an Urgent Field Safety Notice to customers outside the U.S. The letter identified the affected kit lots, reason for voluntary recall, risk to health and actions to be taken by the customer/user. Customers were instructed to discard any inventory of the affected kit lots and complete and return the attached Field Correction Effectiveness Check questionnaire within thirty (30) days to request a no-charge replacement kits. Customers were requested to post this field correction by each IMMULITE system to alert users to the situation. For questions contact your Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution - US (nationwide) | Open |  |  |  | Not_Computer
Z-1707-2012 | 61823 | heparin, vascular access flush | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin Lock Flush Solution 100 Unit/mL, USP3 ml, 4 ml, 5mL, 6 ml fill in a 10 ml Flush Syringe ZRExcelsior Medical Neptune, NJ 07753(800) 487-4276Made in USANDC 63807060055.Used to maintain patency of vascular access devices designed for intermittent infusion therapy. Export product for sale in Colombia are used for locking hemodialysis catheters, for heparinization of extracorporeal systems, in hemodialysis, and for peritoneal dialysis. | Heparin Lock Flush | 2 | June 05, 2012 | 2012 | Excelsior Medical Corp | Some lots of Heparin Lock Flush, USP Syringes (3 mL fill in 10 ml syringe, 4 mL fill in 10 ml syringe, 5 mL fill in 10 ml and 6 ml in 10 ml syringe), failed Anti-Factor IIa assay for potency. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using the affected product and to complete and return an attached Customer Reply Form via fax to (866) 807-3714. Contact Excelsior Medical Customer Service at (800) 487-4276 for questions regarding this recall. A second letter which included additional lots was sent out on 5/25/2012 with the date of 5/24/2012. A return response form was also included. | Contact the recalling firm for information | 34230500 | 34,230,500 units (596 lots) | Worldwide Distribution-USA (nationwide) and the countries of Brazil, Canada, and Colombia. | Terminated | September 12, 2013 | 464 | 34230500 | Not_Computer
Z-1706-2012 | 61766 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia 3 Anti-Torque Key Manual Surgical Instrument; Catalog Number 48237026. Product Usage: :Reusable instrument used to aid in the final tightening of the spinal construct. Intended Use: The Anti-Torque Key allows the Torque Wrench to align with the tightening axis and maximizing the torque needed to lock the implant assembly during final tightening of the spinal construct. | Xia 3 AntiTorque Key | 2 | June 05, 2012 | 2012 | Stryker Spine | Stryker received four reports regarding a handle which separates at the point where it is welded to the shaft on the Xia 3 Anti-Torque Key during surgery. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Stryker Spine sent an Urgent Product Recall letter dated January 12, 2012 with an attached Customer Response forms via FEDEX to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory to identify the affected product and return using the prepaid mailing label provided. The letter asked customers to complete the Customer Response Form and fax to 201.760.8370. For questions call 201.760.8298. | Contact the recalling firm for information | 95 | 95 units | US Nationwide Distribution | Terminated | March 21, 2013 | 289 | 95 | Not_Computer
Z-1704-2012 | 61550 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Knee walker, one unit per box. Manufacturer Kenstone Metal (Kunshan) Co., LTD. The intended use is to provide additional mobility to individuals who have difficulty walking. | Knee Walker | 2 | June 05, 2012 | 2012 | Invacare Corporation | Adverse event data prompted testing of device, which revealed that the rod supporting the knee pad was not made to specifications. Upon placing a load on the knee pad, the rod may break or bend causing the walker to collapse. As a result, the user may fall. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Invacare notified all consignees and distributors with a "Urgent: Recall Information" letter dated May 16, 2012. The product issue was described and recommended actions were provided to customers. Questions were directed to 1-888-257-7912. |  | 1640 | 1640 | Worldwide Distribution--USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of China. | Terminated | March 01, 2013 | 269 | 1640 | Not_Computer
Z-1772-2012 | 61767 | pin, fixation, threaded | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.0 MM DIAMETER 50 MM THREAD LENGTH EXTERNAL FIXATION HALF PIN HA TI-6AL-4V, REF 71932962, QTY: (1), STERILE R, Smith & Nephew, Inc., Memphis, TN 38116Orthopedic surgical device. | 6.0 MM DIAMETER 50 MM THREAD LENGTH EXTERNAL FIXATION HALF PIN HA TI6AL4V | 2 | June 06, 2012 | 2012 | Smith & Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, voiding sterility. | PRODUCTION CONTROLS: Packaging | Smith & Nephew Sales Representatives and Smith & Nephew International Distributors were notified of the problem with the affected product and the recall via email on April 13, 2012. All affected direct consignees were notified via Fed Ex on April 13, 2012. The Urgent Medical Device Recall letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the product and quarantine them immediately for return. Customers were instructed to complete the Inventory Return Certification Form included with the letter. Customers should fax or email the form to 901-566-7975 or fieldactions@smith-nephew.com. The affected product should be returned to:Smith & NephewAttn: Global Field ActionsSafety Affairs1450 Brooks RoadMemphis, Tennessee 38116If the affected product was distributed further customers were instructed to provide a copy of the recall letter to these customers. |  | 4 | 4 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Malaysia, Germany, Japan, Portugal, United Arab Emirates, Spain, France, South Africa, Norway, Sweden, Denmark, Switzerland, Belgium, China, The Netherlands, Finland, Mexico, Italy, UK, Venezuela, Colombia, Thailand, Chile, South Korea, India, and Turkey | Open |  |  | 112205 | Not_Computer
Z-1773-2012 | 61954 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad Cyclesure 24 Biological Indicator, P/N: 14324The STERRAD CYCLESURE 24 Biological Indicator is intended to be used as a standard method for frequent monitoring of the STERRAD Sterilization System cycles. | Sterrad Cyclesure 24 Biological Indicator | 2 | June 06, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated by Advanced Sterilization Products because they recently determined that certain lots of raw materials that were rejected during incoming inspection acceptance test were used for the manufacture of a specific lot range of Sterrad Cyclesure 24 Biological Indicator. | PRODUCTION CONTROLS: Process Control | Advanced Sterilization Products a Johnson & Johnson company sent an Urgent Product Recall letter dated March 16, 2012, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to read the "Details on the Affected Devices/Description of the problem" section in the letter. Customers were instructed to immediately identify and set aside all product listed above in a manner that ensures product will not be used. The lot number and the part number are printed on the label placed around the ampoules. Customers were instructed to return any affected product in accordance with the "Product Return Instructions" section of the letter. Customers were instructed to complete the requested information on the enclosed postage paid business reply card and return it to Stericycle even if they do not have the affected products listed on the letter in stock. Customers with questions were instructed to contact Stericycle at (877) 492-4795.For questions regarding this recall call 949-453-6400. | Contact the recalling firm for information | 2284 | 2,284 cases (137,040 Biological Indicators) | Worldwide Distribution - USA (nationwide) and the countries of AR, AU, BE, BR, CA, EC, HK, JP, MX, and PR | Terminated | October 17, 2013 | 498 | 2284 | Not_Computer
Z-1783-2012 | 61925 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine FuseLOX Lumbar Cage. Products are labeled in part: "FuseLOX Lumbar Cage Lumbar Interbody Fusion Device***Description:***Part Number:***Lot Number:***Quantity: 1***Material: PEEK-OPTIMA LT1, Tantalum ASTM F560***Non-Sterile-Single Use Only (See package insert for sterilization parameters and labeling.)***CAUTION: Federal (United States) law restricts these devices to sale by or on the order of a physician.***Manufactured and distributed by: CAPTIVA SPINE***captivaspine.com***967 Alternate A1A #1 Jupiter, FL 33477 Tel: 877-772-5571 Fax: 866-318-3224 Email: info@captivaspine.com***"FPC - F is FuseLOX, P is Plif, C is ConvexProducts subject to recall: Convex - 7mm x 25mm Part Number: FPC2507 Lot Number 07110002. Convex - 8 mm x 25mm Part Number FPC2508 Lot Number 07110003. Convex - 9mm x 25mm Part Number: FPC2509 Lot Number 07110004.Convex - 10mm x 25mm Part Number: FPC2510 Lot Number 07110005. Convex - 11mm x 25mm Part Number: FPC2511 Lot Number 07110006. Convex - 12mm x 25mm Part Number: FPC2512 Lot Number 07110007. Convex - 13mm x 25mm Part Number FPC2513 Lot Number 07110008. Convex - 14mm x 25mm Part Number FPC2514 Lot Number 07110009. Products are packaged individually and shipped as replacement parts to be placed in the FuseLOX Lumbar Caddy which is placed in the FuseLox Lumbar Sterilization tray. The sterilization tray is distributed on consignment.The FuseLOX Lumbar Cage implant is an intervertebral body fusion device indicated for intervertebral body fusion at one level or two contiguous levels in the lumbar spine from L2 to S1 in patients with degenerative disc disease (DDD) with up to Grade 1 spondylolisthesis at the involved level(s). DDD is defined as back pain of discogenic origin with degeneration of the disc confirmed by patient history and radiographic studies. These patients should be skeletally mature and have had six months of non-operative treatment. The FuseLOX Lumbar Cage Implant is designed for use with autograft to facilitate fusion and is intended for use with supplemental fixation systems cleared for use in the lumbar spine. | Spinal Invervetebral Disc Space Orthosis | 2 | June 07, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine, Inc. initiated a recall of FuseLOX Lumbar Implants, Trials, and Inserters after FDA advised the firm that there is inadequate support for implemented design changes to the cleared device Pivotec Lumbar Interbody Fusion Device. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Finished Device Change Control | Captiva Spine sent a Voluntary Correction Notification letter dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and return affected devices. Customers were also instructed to complete and return a Tracking/Verification Form.For questions customers were instructed to call 877-772-5571.For questions regarding this recall call 561-277-9480. |  | 136 | 136 units | USA (nationwide) including FL,TX, UT, and NJ. | Terminated | December 24, 2013 | 565 | 254 | Not_Computer
Z-1788-2012 | 37654 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Revolution XR/d; GE Healthcare, 3000 North Grandview , Waukesha, WI, 53188Model numbers: 2259988-2 & 2351505.The Revolution XR/d Digital Radiographic Imaging System is indicated for use in generating radiographic images of human anatomy. It is intended to replace radiographic film/screen system in all general-purpose diagnostic procedures. This device is not intended for mammographic applications. | Revolution XR/d | 2 | June 08, 2012 | 2012 | General Electric Med Systems LLC | An incident was reported that an operator's finger was pinched between the fixed roller and the gliding tabletop of a Proteus XR/a Table while moving the tabletop with the thumb placed over its edge and fingers underneath | DESIGN: Device Design | GE Healthcare sent an "URGENT SAFETY INFORMATION" letter dated May 31, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included new operator instructions and warning labels. |  | 3796 | 3,796 units | Worldwide Distribution. | Terminated | July 31, 2012 | 53 | 3796 | Not_Computer
Z-1792-2012 | 61853 | apparatus, suction, single patient use, portable, nonpowered | GCY | Nonpowered, single patient, portable suction apparatus. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Lipofilter 3000, model ASP-CAN-2, labeled in part ***MicroAire, 1641 Edlich Drive, Charlottesville, VA 22911, USA***Product Usage:A fat collection canister intended for use during liposuction procedures. | Lipofilter 3000 | 2 | June 09, 2012 | 2012 | MicroAire Surgical Instruments, LLC | Medical device was not manufactured under Good Manufacturing Procedures and a user may be unable to sterilize the device according to procedures. |  | MicroAire Surgical Instrument sent a Device Removal Notice letter dated April 26, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were advised to recover the affected product from inventory and quarantine. Customers were instructed to contact MicroAire Customer Service Department at 800-975-8000 or 434-975-8000 to arrange for return and replacement of the affected product. For questions call 434-975-8336 or email toddm@microaire.com. |  | 15 | 15 devices | US Nationwide Distribution including the states of CA, OH and NY. | Terminated | November 21, 2012 | 165 | 15 | Not_Computer
Z-1794-2012 | 61951 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | PlasmaJet Open Surgery Handpiece (Part of the PlasmaJet Neutral Plasma Surgery System)The PlasmaJet system is a neutral plasma surgery system that is designed for cutting and coagulation in open surgery and laparoscopic surgery. | PlasmaJet Open Surgery Handpiece | 2 | June 11, 2012 | 2012 | Plasma Surgical Inc. | Potential failure in the glued seal between the stem and the tip of the handpiece leading to the possible leakage of a very small drop of distilled water coolant from the handpiece in use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Plasma Surgical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 27, 2012 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their physical inventory; locate and quarantine or remove from service all identified product; and complete and return the attached CUSTOMER RESPONSE FORM via fax to designated persons at: (404) 678-4395 or mail to: Plasma Surgical Inc., 1125 Northmeadow Parkway, Suite 100, Roswell, Georgia 30076 (even if they no longer have any of the product in their inventory). A Plasma Surgical Representative will contact the customers to arrange for the return of the products and product replacement, should they have any of the physical inventory (product).Should you have any queries concerning this matter, please do not hesitate to contact, Vice President of Sales and Marketing, imawhinney@plasmasurgical.com or call (678) 578-4390. |  | 36 | 36 units | Worldwide distribution: USA (nationwide) including states of: CA and NY; and country of: United Kingdom. | Terminated | November 01, 2012 | 143 | 36 | Not_Computer
Z-1796-2012 | 61717 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra Mozaik Putty2.5cc, 5 cc, 10 cc Product Usage:The product is a bone graft substitute composed of collagen (Type 1) and ceramic (IP3-Tricalcium phosphate) matrix in putty and strip forms. The product is intended for use as a bone void filler to fill voids or gaps of the skeletal system in the extremities, spine. and pelvis not intrinsic to the stability of the bony structure. Integra Mozaik is also indicated for use in the treatment of surgically treated osseous defects or osseous defects created from traumatic injury to the bone. Following placement in the bone void or a defect Integra Mozaik is reabsorbed and replaced with bone during the healing process. | Integra | 2 | June 12, 2012 | 2012 | Integra LifeSciences Corp. | The Instructions for Use {HIFU} {Putty Part Number 7700713001 and Strip Part Number 770151300 l for CE marked Mozaik Strip and Putty products contains an incorrect translation of the term "bone marrow aspirate". The term is incorrectly translated in Spanish and French, | TRAINING: Employee Error | Integra LifeSciences sent a Field Safety Notice letter dated October 11, 2011 to their sole foreign consignee. The letter identified the affected products, problem and actions to be taken. The consignee was asked to remove the existing Instructions for Use (IFU) and replace it with the corrected IFU. The consignee was instructed to complete and return the attached Distributor Rework and Acknowledgment Forms provided with the customer notification letter as documentation that instructions were followed and all required actions were completed. For any questions, call: +33 (0)4 37 47 5900 , fax: +33 (0)4 37 47 5999; or email: carole.robin@integralife.com. |  | 42 | 42 units | Worldwide Distribution | Terminated | February 14, 2013 | 247 | 64 | Not_Computer
Z-1801-2012 | 61945 | heparin, vascular access flush | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Vygon VY(TM)Flush Heparin Lock Flush Syringes, 100 Units per mL, 5mL in a 10mL Flush Syringe, 30 syringes per box, Rx only Prefilled syringes are used to maintain the patency of indwelling catheter lumens | Vygon VY(TM)Flush Heparin Lock Flush Syringes | 2 | June 13, 2012 | 2012 | Vygon Corporation | Manufacturer issued a voluntary recall on multiple lots of VY(TM)Flush Heparin Lock Flush Syringes therefore Vygon has subsequently initiated a recall on Vygon branded Heparin Lock Flush Syringes and kits containing the syringes. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | VYCON sent an URGENT MEDICAL DEVICE RECALL letter dated May 10, 2012, to all affected customers. The firm expanded their recall on May 25, 2012 by sending a second letter with additional lots and products. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory and quarantine any affected product. Customers should contact Vycon's Customer Service Department at 1-800-473-5414 for a Return Merchandise Authorization number. Customers were instructed to enter the number on the space provided on the Recall Acknowledgment and Inventory Return Form, sign and date form and fax to Vycon at 215-672-6740. Customers should return the affected product as well as the original signed Recall Acknowledgment and Inventory Form using the instructions from Customer Service involving product pick-up. All inventory will be credited. If product was further distributed the Urgent Medical Device Recall letter should be provided to their customers.For questions regarding this recall call 215-672-6740. |  |  |  | Nationwide Distribution including AZ, CA, DE, FL, IL, IN, IA, KS, KY, MD, MA, MI, MN, MO, NJ, NY, OH, PA, UT, VA, WA, WV, and WI. | Terminated | March 11, 2013 | 271 |  | Not_Computer
Z-1490-2012 | 61355 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- Common/Classification Name: Infusion Pump -- Class 2 medical device -- SIGMA LLC, Medina, NY 14103 Product Usage:The Spectrum Infusion Pump is intended to be used for the controlled administration of intravenous fluids. These fluids may include blood, blood products or mixtures of pharmaceutical drugs for required patient therapy. The Spectrum is used in conjunction with legally marked intravenous administration sets and medications provided by the user. The Master Drug Library is a software package that will add additional features to the Spectrum Infusion Pump. The Master Drug Library will permit electronic communications with the Spectrum pump and other external peripheral devices. The intended use of the Spectrum pump includes common drug error prevention, through the stand alone settings features of the pump. This includes drug parameter limits and associated drug name identification. With the Master Drug Library, the intended use is to reduce user errors associated with drug selection, drug dose rates, drug dose concentrations, and patient identification associated with the prescribed drug. | Spectrum | 1 | June 13, 2012 | 2012 | Sigma | Recall expansion; the infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. Sigma's evaluation of subsequent complaints involving failed bearings indicated that the loss or degradation of bearing lubricant to be a more significant contributor to bearing failure than initiall | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sigma sent an URGENT Medical Device Recall Notifications letter dated July 6, 2011 via UPS mail the consignees informing them of the expansion of the Spectrum Infusion Pumps recall. Some consignees were notified by telephone. The letter identified the affected product, problem, actions to be taken, clinical precautions and recommendations. Additional information regarding this voluntary recall is available on Sigma's web site www.sigmapumps.com/Recall. Customers were provided the following instructions: (1) Confirm that you have the affected serial number(s) ranging from 700000 to 794213 in inventory using the Customer Response Form. Refer to attached list of known affected devices in your possession. Revise this list as required. (2) Immediately remove affected pumps from neonatal use until B. Braun BC 1000 back check valves are employed. (3) Respond to this notification within five (5) working days of receipt by faxing the enclosed Customer Response Form to the number noted on the form, calling (866) 482-2893 between the hours of 8am and 8pm (EDT) or e-mailing a scanned version of form to sigma2407@stericycle.com. (4) Within three weeks of receipt of the Customer Response Form, (which verifies the device serial numbers in your facility), SIGMA will issue a customized recall report for each facility identifying the specific affected serial numbers, SIGMA internal contacts, and the plan for coordination of pump return and repair. For any adverse reactions experienced with the use of this product should be reported to SIGMA at (800)-356-3454 ext. 300 and to FDA MedWatch either by: Phone at 1-800-FDA-1088, by fax at 1-800-FDA-0178; Mail at MedWatch, FDA 5600 Fishers Lane, Rockville, MD 20857, or MedWatch website by linking to: https://www.accessdata.fda.gov/scripts/medwatch/medwatch-online.htm. |  | 29 | 63,225 pumps total (corrected number): Domestic - 63,032 pumps; Canada - 29 pumps; Okinawa - 164 pumps | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of Canada and Okinawa | Open |  |  | 63225 | Not_Computer
Z-1810-2012 | 61848 | gown, surgical | FYA | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | T4-Hytrel Zipper Toga Product Usage - The togas are intended to provide a barrier between the operating room and the members of the surgical team in order to help protect against contamination and/or exposure of infectious body fluids and harmful microorganisms. | T4 Hytrel Zipper Toga | 2 | June 14, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape, that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids. | PRODUCTION CONTROLS: Process Control | Stryker sent the May 29,2012 URGENT MEDICAL DEVICE RECALL NOTIFICATION and the June 14, 2012 REVISED URGENT MEDICAL DEVICE RECALL Notice to all consignees. The letters identified the affected products, problem and actions to be taken. The Revised letter instructed customers to find, quarantine, make product unusable and destroy all implicated product . Replacement products may not immediately available. Customers were asked to complete and sign the Business Reply Form (BFR) and fax to Stryker Instruments Regulatory Department at 866-521-2762. If the affected products have been further distributed, forward this notification letter and the attached BRF to all affected locations. Questions should be addressed to Angela Ragainis, 269-389-4354; angela.ragainis@stryker.com. |  | 846 | 16, 846 (last 3 years) 35,492 all years | Worldwide Distribution - US (nationwide) including the states of: AK, AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV and the country of SWITZERLAND | Open |  |  | 16846 | Not_Computer
Z-1811-2012 | 61737 | processor, cervical cytology slide, automated | MKQ |  |  | Pathology |  | PrepStain Syringing Pipettes 240 (Catalog #490517) | PrepStain Syringing Pipettes | 3 | June 15, 2012 | 2012 | Tripath Imaging, Inc. | The PrepStain Syringing Pipettes may not dispense the minimum volume of 7.5 mL of the sample due to leaking. This event resulted in an increase in complaints of leaking syringing pipettes | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostics sent a "BD PrepMate Syringing Pipette" letter dated March 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm encouraged customers to refer to the Operator's Manual and inspect the centrifuge tubes to verify that the correct amount of sample has been transferred. Customers experiencing this issue were instructed to discard the pipettes and contact BD Diagnostic Technical Support for a replacement. Contact the firm at 1-877-822-7771 for questions regarding this notice. | Contact the recalling firm for information | 46112784 | 46,112,784 for all products in Recall Event | Worldwide Distribution. | Terminated | December 03, 2012 | 171 | 46112784 | Not_Computer
Z-1816-2012 | 62130 | rasp, surgical, general & plastic surgery | GAC | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Accolade Offset Rasp HandleHowmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430 USA Stryker France ZAC Satolas Green Pusignan Av de Satolas Green 69881 Meyzieu Cedex FranceSingle offset rasp handle designed to mate to the Accolade femoral broaches to prepare the femoral canal in a total hip anthroplasty using the Accolade Hip System. The offset rasp handle incorporates an offset angle designed to minimize interference with surrounding soft tissue. | Stryker Howmedica Osteonics Accolade Offset Rasp Handle | 2 | June 15, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for the above noted Accolade Rasp Handle Assembly to fracture upon use. | DESIGN: Device Design | Stryker Orthopaedics sent an email notification to all consignees on December 29, 2011, which was followed by an Urgent Product Recall letter dated January 4, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the attached Product Recall Acknowledgement Form and fax back to 201-831-8089 within 5 days. Customers were also asked to contact the hospitals in their territory that have the affected product to arrange return of the product. Customers were instructed to return the affected product to the attention of Regulatory Compliance, Stryker Orthopedics, 325 Corporate Drive, Mahwah, New Jersey 07430 with the fluorescent orange PRODUCT REMEDIATION sticker attached indicating the Product Remediation # RA 2011-161. Mark the outer box with the words "Product Recall."For questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5970. | Contact the recalling firm for information | 41 | 41 units | Nationwide Distribution | Terminated | June 13, 2013 | 363 | 41 | Not_Computer
Z-1787-2012 | 61639 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aestiva/5 7900 SmartVent, anesthesia machine. Used in Datex-Ohmeda Aestiva/5 Anesthesia Systems. | Aestiva/5 7900 SmartVent | 1 | June 17, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the interlock slide mechanism missing from the 3-vaporizer manifold of your Aestiva 7900 Anesthesia Device. This Potential for two (2) vaporizers delivering agent at the same time. This could manifest as over delivery of a single agent if both vaporizers contain the same agent or as delivery of more than one agent. This m | PRODUCTION CONTROLS: Process Control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 10, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer or Technical Support at 800-345-2700 for questions regarding this notification. | Contact the recalling firm for information | 8 | 8 | Nationwide Distribution - including the states of AL, FL, GA, MO, and VA. | Terminated | July 09, 2012 | 22 | 8 | Not_Computer
Z-1819-2012 | 61868 | vibrator, therapeutic | IRO | Therapeutic vibrator. | Physical Medicine | Physical Medicine | 510(K) Exempt | US-001 Rev A brochure is a two sided tri-fold printed marketing brochure describing the approved labeling indications for use, precautions and general operational information of the Exer-Rest device, Models AT3800 and AT4700. Ever-Rest US Brochure US-001 Rev A. Brochure: "***AT acceleration THERAPEUTICS***At Non-Invasive Medical Solutions (NIMS) we pride ourselves on being at the forefront of developing drug free wellness solutions coupled with accredited peer-reviewed clinical research...***Whole Body Periodic Acceleration (WBPA) therapy has been demonstrated to naturally activate the release of beneficial mediators like Nitric Oxide (NO) throughout the vasculature.***Nitric Oxide [NO] supports the functioning of endothelial cells that line the arterial walls. NO signals vascular smooth muscle to relax, thereby increasing blood flow through arteries and veins. With age arterial production and availability of Nitric oxide declines.***NiMS 4400 Biscayne Blvd., Miami, FL 33137 P 305.575.4200 F 305357534201***info@nims-inc.com www.nims-inc.com***NIMS, Acceleration Therapeutics and Exer-Rest are registered trademarks of Non-Invasive Monitoring Systems, Inc.***US-001 Rev A***"Printed marketing brochure | EverRest US Brochure US001 Rev A. | 2 | June 18, 2012 | 2012 | Non-Invasive Monitoring Systems, Inc. | Non-Invasive Monitoring Systems Inc (NIMS) is voluntarily recalling promotional materials which contain unsubstantiated direct and indirect claims. | MISBRANDING: Labeling False and Misleading | NiMS sent an Urgent Promotional Material Recall letter dated May 29, 2012, to all affected customers. NIMS contacted the 18 facilities by phone starting on May 29, 2012 and sent the 73 end users letters dated May 29, 2012. NiMS instructed customers to immediately examine their inventory and quarantine any of the US-001 or USA-002 brochures. If customers further distributed brochures customers were instructed to notify their customers at once by including a copy of the recall notification letter. Customers were requested to destroy all Exer-Rest promotional materials. For any questions customers were instructed to call 305-575-4204.For questions regarding this recall call 305-575-4201. |  | 3665 | US-001 Rev A brochure: 3665 | Nationwide Distribution | Terminated | July 16, 2012 | 28 | 3756 | Not_Computer
Z-1822-2012 | 62120 | agent, tooth bonding, resin | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Peak SE Primer Kits (4 syringes per kit) manufactured by Ultradent Products, Inc, South Jordan, UT, Self Etch Primer. The product is contained within prefilled, labeled syringes in a transparent plastic container identified with an aqua, red, orange, blue and black label with black and blue print writing. Product code: 5135, 872.3200 KLE, liquid and powder mixture is a syringe-in-a-syringe delivery system.The product is used for most bonding needs in restorative dentistry. Bonds chemical cure luting and chemical cure composites, namely PermaFlo DC. Peak SE does not require a separate etching step for quality adhesion. Peak SE is also conducive for bonding to: Dentin and enamel, porcelain, metal and composite. | Peak SE Primer Refill | 2 | June 18, 2012 | 2012 | Ultradent Products, Inc. | Ultradent is recalling lot B6Z4K of Peak SE Primer as a result of a field complaint, dated May 10, 2012, stating that a syringe of Peak SE Primer was mislabeled as Peak LC Bond Resin. These products are used by Dental professionals for bonding needs in restorative dental procedures. | PRODUCTION CONTROLS: Labeling Mix-Ups | Ultradent Products Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discard any affected product from their 5135 kit. Customers were asked to complete and return a Recall Return Response form to the firm by fax at 801-553-4609. |  | 172 | 172 kits (4 syringes per kit) | Worldwide Distribution. | Terminated | August 15, 2012 | 58 | 172 | Not_Computer
Z-1824-2012 | 57878 | heparin, vascular access flush | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HEPARIN LOCK FLUSH SOLUTION, USP, 10 USP Units/mL, 1 mL Multiple Dose Vial, plastic vial, vial size 3 mL, NDC 63323-544-01, Product Code 504401, Rx onlyTo maintain patency of an indwelling venipuncture device designed for intermittent injection or infusion therapy or blood sampling. | HEPARIN LOCK FLUSH SOLUTION | 2 | June 19, 2012 | 2012 | App Pharmaceuticals Llc | CGMP Deviations: Incomplete documentation associated with test results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | APP sent an URGENT DRUG RECALL LETTER and Response Form dated May 20, 2010, to all affected consignees. The consignees contacted were at the retail/ medical facilities/hospital IeveIs, including intermediate wholesale IeveIs. Direct distributor consignees were instructed to notify their customers immediately of this recall action, and direct them to discontinue distributing or dispensing the affected lots, plus return the affected products and lots to APP Pharmaceuticals, LLC in Bensenville, IL. Direct consignees with affected lots on hand were instructed to immediately discontinue distributing or dispensing the affected products and return them to APP Pharmaceuticals, LLC in Bensenville, IL 60106. Questions regarding the information on how to return product can be made to the firm's Quality Assurance Department at 1-866-716-2459.For questions regarding this recall call 847-969-2700. |  | 133150 | 133,150 vials | Nationwide Distribution and Puerto Rico | Terminated | June 20, 2012 | 1 | 637651 | Not_Computer
Z-1774-2012 | 61843 | generator, oxygen, portable | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | NIDEK Medical M5C5/ MARK5 NUVO/STD, 230 V~50/60 Hz - 420 W.Oxygen Concentrator. | NIDEK Medical M5C5/ MARK5 NUVO/STD, 230 V50/60 Hz 420 W | 1 | June 19, 2012 | 2012 | Nidek Medical Products Inc | Capacitor failure may result in a fire hazard and loss of supplemental oxygen. | DESIGN: Component Design/Selection | Nidek Medical Products Inc sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 1, 2012 to affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter included a response form for customers to complete and return to the firm. Contact the the firm at 205-856-7200 ext. 242. | Contact the recalling firm for information | 7889 | 7889 units | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ENGLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, KAZAKHSTAN, KINGDOM OF SAUDI ARABIA, KUWAIT, LEBANON, LIBYA, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPPINES, PORTUGAL, REPUBLIC OF CHINA, REPUBLIC OF SOUTH AFRICA, RUSSIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH AUSTRALIA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, THE NETHERLANDS, TOGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, and VIETNAM. | Open |  |  | 46972 | Not_Computer
Z-1823-2012 | 62125 | catheter,intravascular,therapeutic,short-term less than 30 days | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow International, Inc. 7FR 3-Lumen Arrowg+ard Blue Plus(R) Pressure Injectable Central Venous Catheter - Catalog number CDC-42703-XP1AThe multiple-lumen pressure injectable catheter permits venous access to central circulation and injection of contrast media. The ARROWg+ard(R) antimicrobial surface is intended to provide protection against catheter-related infections. | Arrow International, Inc. 7FR 3Lumen Arrowgard Blue Plus(R) Pressure Injectable Central Venous Cat | 2 | June 19, 2012 | 2012 | Arrow International Inc | Arrow is recalling the product due to the possibility of voids in the catheter hub which could lead to an interlumen crossover. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Arrow International sent a Urgent Medical Device Recall Notification letter dated June 11, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot numbers listed.2. To return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, Attn: Customer Service. This will allow us to document the amount of the product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization ( RGA) Number and will provide instructions for the return of product to Arrow International.3. If you have no affected stock, please complete t he enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 1085 | 1,085 | USA (nationwide) AL, AZ, AR, CA, CO, FL, GA, IL, IN, KS, LA, MA, MI, MO, NV, NH, NC, OH, PA, TX, VA, WA, and WI. | Open |  |  | 1085 | Not_Computer
Z-1828-2012 | 61376 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Implant, Dental, Endosseous, Acid Etched, ICE, Osseotite, Internal Connection. . Product is labeled in part: "***Certain 3.25, 13***REF IOSM313 LOT*** OSSEOTITE Certain Implant***3.25 x 13mm***CE 0086***Use by yyyy-mm***Sterile by radiation*** STERILE R***Do not re-sterile***RX only Caution: Law prohibits dispensing without prescription***do not use if package is damaged***do not reuse***consult instructions for use***Sterile unless package is opened or damaged.***Product Usage:The BIOMET 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means of prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Endosseous Dental Implants | 2 | June 20, 2012 | 2012 | Biomet 3i, LLC | On November 3, 2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Osseotite Certain Implant, Model # IOSM313, Lot #2010111529. Dental implant lot may not have an internal thread. | DESIGN: Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated November 04, 2011 via e-mail, facsimile, telephone or postal mail to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each customer was instructed to check their respective inventory for the affected product and return any unused item(s) to Biomet 3i for prompt replacement. For question or concerns contact Biomet 3i at 1-800-342-5454. | No consumer action necessary | 219 | 219 units in total | Worldwide Distribution - US(nationwide) and the countries of: Brazil, Canada, Columbia, Italy, Lebanon, Spain, Switzerland, and UK. | Terminated | June 25, 2012 | 5 | 219 | Not_Computer
Z-1832-2012 | 61742 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | SARNS HIGH-FLOW AORTIC ARCH CANNULA: 5.2 MM (16 FR) 00 WITH 3/8" CONNECTOR, 7.5" (19 CM) LONG.The Sarns High Flow Aortic Arch Cannula is indicated for use in perfusion of the ascending aorta during cardiopulmonary bypass surgery. | SARNS HIGHFLOW AORTIC ARCH CANNULA | 2 | June 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of one lot of the Sarns High-Flow Aortic Arch Cannula, Terumo Cardiovascular Systems (Terumo CVS) found a foreign substance on the outer surface of some cannulae connectors.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface, the substance was likely deposited during the molding process, but the exact composition of the parti | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advises customers to discontinue use of the affected product and to return the product to the firm. A Response Form was attached for customers to complete and return. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 133793 | 133,793 (each) | Worldwide Distribution-USA (nationwide) including the states of AL , AR , AZ, CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , MD , ME , MI , MI , MN , MO , MT , NC , ND , NE , NH , NJ , NM , NV , NY , OH , OK , OR , PA , RI , SC , SD , TN , TX , UT , VA , WA , WI, and WY and the countries of Australia, Belgium, Brazil, Canada, Chile, Germany, Honduras, Hong Kong, India, Japan, Mauritius, Nicaragua, Philippines, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates (UAE), and Vietnam. | Terminated | July 11, 2013 | 386 | 133793 | Not_Computer
Z-1845-2012 | 61942 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Covidien ROTICULATOR 55-4.8 Single Use StaplerREF# 017614The RETICULA TOR TM Single Use Staplers have applications in abdominal, gynecologic and pediatric surgery for resection, transection, and creation of anastomosis | ROTICULATOR Single Use Stapler | 2 | June 21, 2012 | 2012 | Covidien LP | Potential for the sterility barrier to be compromised. | DESIGN: Packaging Design/Selection | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 16, 2012 via federal express to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (note: units from affected lots may have been incorporated into Covidien BEST PRACTICE KITS; evaluate kits for affected lots). All affected ROTICULATOR units must be returned. The customers were also instructed to complete and return the attached ROTICULATOR Single Use Stapler Recalled Product Return form via fax to: (800) 895-6140 or email to: SDFeedback@Covidien.com, even if you do not have any of the affected units, and include a copy of completed form with products being returned. Customers who received product directly from Covidien should contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722-8772, option 1, prior to returning affected products to Covidien, Field Returns Department, at 195 McDermott Road, North Haven, CT 06473; if received from distributor, return form and contact distributor directly.If you have any questions or concerns, please do not hesitate to contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 13531 | 13531 units | Worldwide distribution: USA (Nationwide) and countries including; Africa, Argentina, Australia, Belgium, Brazil, Canada, China Costa Rica, CROATIA, Czech Republic,Dominion Republic, ESPAÂŽA France, Germany, Greece Guadeloupe, Guatemala, Honduras Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Korea ,Malaysia, Mexico, Peru, Portugal, Puerto Rico, Spain, Sri Lanka, Switzerland, Taiwan , Thailand, and UK. | Terminated | February 12, 2014 | 601 | 123856 | Not_Computer
Z-1850-2012 | 61935 | snare, flexible | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Profile Single-Use Pediatric Snare 11mmUPN Outer Box M00562531Inner Pouch: M00562530 ,The single-use Profile Polypectomy Snare consists of a flexible wire cable loop which can be extended and retracted from the snare's flexible outer sheath using a three-ring handle. When passed through an endoscope and activated, the snare delivers a monopolar electrical current to cut and cauterize tissue with the loop | Boston Scientific Profile SingleUse Pediatric Snare | 2 | June 21, 2012 | 2012 | Boston Scientific Corporation | Difficulty in extending snare loop from the catheter | PRODUCTION CONTROLS: Process Control | Boston Scientific issued an Urgent Medical Device Recall Action letter via Federal Express Priority mail on May 16, 2012. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. Scan and fax to 508-683-5578.For questions regarding this recall call 508-683-8000. |  | 2760 | 2760 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Austria, Belgium, France, GB, Germany, Italy, Netherlands, Qatar, Spain, Sweden, and Switzerland | Terminated | February 28, 2013 | 252 | 6070 | Not_Computer
Z-1855-2012 | 62127 | barrier, absorable, adhesion | MCN |  |  | Obstetrics/Gynecology |  | seprafilmÂŽ Procedure Pack ADHESION BARRIERRe-order Number: 5086-02Product Usage:Seprafilm Adhesion Barrier is indicated for use in patients undergoing abdominal or pelvic laparotomy as an adjunct intended to reduce the incidence, extent and severity of postoperative adhesions between the abdominal wall and the underlying viscera such as omentum, small bowel, bladder, and stomach, and between the uterus and surrounding structures such as tubes and ovaries, large bowel, and bladder. | Genzyme | 2 | June 21, 2012 | 2012 | Genzyme Corporation | Sterility of product may be compromised due to packaging defect | PRODUCTION CONTROLS: Packaging Process Control | Genzyme issued an Urgent: Medicatical Device Product Recall Notification Letter dated June 13, 2012 to all affected customers. Stericycle Expert Solutions, will be acting on behalf of Genzyme to coordinate notifications, communications, and logistics for returns of all product subject to this recall. The letter identified the affected products, problen and actions to be taken. Customers are requested to check their inventory and return all affected products per the instructions provided. Customers are instructed to complete the enclosed Business Reply Form and return via fax transmission to 866-912-2519. For questions call Stericycle at: 877- 473-7069. | Contact the recalling firm for information | 3880 | 3880units | Worldwide Distribution - US (nationwide) and the country of S. Korea | Terminated | July 24, 2013 | 398 | 5165 | Not_Computer
Z-1860-2012 | 61712 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Go Bed II+ Model FL28C Beds are intended for medical purposes in both general patient care as well as critical care environments. Each uses built-in electric motors and controls which can be operated by the patient, or caregiver, to adjust the height and/or surface contour of the bed.The devices listed above also include movable and latchable side rails which can be lowered or raised to help prevent patient fall and subsequent injury. The Beds include an optional Scale (Models: FL20E, FL25E, FL28C, FL28EX) as well as an optional Bed Exit (1 Zone) or Chaperone | Stryker GoBed II | 2 | June 22, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | An adverse trend of scale and bed exit service reports/complaints were detected for the GoBed+, GoBed II, Rose, and MA204 model beds. Investigation found an abnormally high failure rate of loadcells, an integral component in scale and bed exit function. The scale faults are manifested as a fluctuating scale reading, a scale error causing the scale system to be disabled until serviced, or an inac | DESIGN: Component Design/Selection | The firm, Stryker Medical, sent an "URGENT- Medical Device Correction" letter dated May 17, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to observe the beds for any malfunctions according to Preventive Maintenance Checklists in maintenance manual; report scale or bed exit malfunctions to Stryker (1-800-327-0770, option 2) to arrange for field service or to order service parts, at no charge. The customers were also instructed to forward this notice and advise Stryker of any subaccounts if the product was further distributed. A postage paid return postcard is to be filled for tracking and to make arrangements for repair. If you have any urgent questions or concerns, call 269-389-6466, Monday-Friday 8a.m-5p.m. (ET). |  | 8009 | 8009 | Worldwide distribution:USA (nationwide) and countries including: Argentina, Bertec, Brasil, Brazil, Canada, Chile, China, India, Latin America, Mexico,Osteonics SA, Polska, Portugal, Puerto Rico, and Singapore. | Open |  |  | 27319 | Not_Computer
Z-1858-2012 | 61899 | gauze/sponge,nonresorbable for external use | NAB | Nonresorbable gauze/sponge for external use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Hartmann Sterilux Ref 56910000 Sterile Premium gauze sponge 4"x4" 12-ply 24 boxes/case 1200 total spongesFor use in wound cleansing, debriding of wounds, wound packing, covering and swabbing | Sterilux | 2 | June 22, 2012 | 2012 | Hartmann USA, Inc | Micro holes were detected in the laminate of the sterile barrier system of sterilized Hartmann Econolux and Hartmann Sterilux gauze products | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, HARTMANN USA, sent an "URGENT: RECALL NOTICE" dated May 14, 2012 to its Consignees/Customers. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, block further shipments of affected items and retrieve additional affected stock from their customers; to confirm initial receipt of the notice by emailing recall@harmanninfo.com; complete and return the Affected Products and Stock Levels form via fax to: HARTMANN USA customer service at (1-803-325-7606); contact customer service to obtain a Returned Goods Authorization; and take adequate measures to ensure this notice is distributed appropriately.If you have any further questions or issues related to this recall, please contact HARTMANN USA customer service or your sales representative at 1-800-243-2294 or email: recall@hartmanninfo.com. | Contact the recalling firm for information | 604711272 | Lot 102147399 - 421 Cases; Lot 102250396 - 360 Cases; 200105390 - 522 Cases; 200206394 - 390 Cases | Nationwide distribution: USA including states of: AZ, CA, FL, GA, KY, MO, MS, NC, PA, SC, TN, TX, and VA. | Open |  |  | 1693 | Not_Computer
Z-1864-2012 | 62331 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Advanta 2 Bed.The Advanta 2 Bed is intended for low to moderate acuity patients in the medical/surgical area of the hospital. | Advanta 2 Bed. | 2 | June 22, 2012 | 2012 | Hill-Rom, Inc. | During an evaluation of Advanta 2 siderail samples returned from the field it was found that the screws used to attach the siderail were being stripped from the plastic cavity of the siderail. | DESIGN: Process Design | Hill-Rom sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm. Contact Hill-Rom Technical Support at 800-445-3720 for questions regarding this notice. |  | 7843 | 7843 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, and Singapore. | Open |  |  | 7843 | Not_Computer
Z-1866-2012 | 61955 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Formula Shaver Handpiece (with buttons);Formula 180 Shaver Handpiece (with button)Rx only, Made USA.:Stryker Endoscopy5900 Optical CourtSan Jose, CA 95138Components within the Crossfire System that provided abrasion, resection, debridement and removal of bone and soft tissue through the Formula Shaver and blade, and the ablation and coagulation of soft tissue, as well as hemostasis of blood vessels through its SERFAS Energy electrosurgical probe. | Formula 180 Shaver Hand Control | 2 | June 25, 2012 | 2012 | Stryker Endoscopy | Stryker Endoscopy determined that there may not have been adequate determination of required routine checks to ensure accuracy and fitness for use of specific serial numbers of Formula and 180 Shaver Hand Control. | PRODUCTION CONTROLS: Process Control | The firm, Stryker Endoscopy, sent an "Urgent: Device Removal" letter dated May 21, 2012 to all affected customers via Federal Express. Stryker Sales Reps were notified by e-mail on the same day. Stryker international sites were notified by e-mail on May 24, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory for part and serial number; quarantine any of the affected product, call Stryker Endoscopy's Customer Service at 1-800-624-4422 for replacements and return affected product.If you have any questions regarding this letter please contact us at shavers@stryker.com or call 1-800-624-4422. |  | 34 | 34 devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, IA , MI, NV, NC, OK, TX, and WI; and countries of: China and Great Britain. | Terminated | September 20, 2012 | 87 | 34 | Not_Computer
Z-1827-2012 | 62108 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius GranuFlo (powder) Acid ConcentrateProduct Codes:OFD1201-3B, OFD1251-3B, OFD2123-3B, OFD2201-3B, OFD2220-3B, OFD2223-3B,OFD2225-3B, OFD2231-3B, OFD2251-3B, OFD2301-3B, OFD2323-3B, OFD2325-3B, OFD3201-3B, OFD3231-3B, OFD3251-3B, OFD3301-3B.For the treatment of acute and chronic renal failure during hemodialysis procedure. | Fresenius GranuFlo (powder) Acid Concentrate | 1 | June 25, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of Alkalosis with acetate containing dialysis acidconcentrates | MISBRANDING: Labeling False and Misleading | Fresenius Medical Care sent an "IMPORTANT PRESCRIBING INFORMATION" letter dated March 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form was enclosed for customers to complete and return via fax to 781-699-9635. Contact Customer Service at 1-800-323-5188 or Medical Information at 1-855-616-2309 for questions regarding this recall. |  |  |  | Nationwide Distribution and the country of Guam. | Open |  |  |  | Not_Computer
Z-1830-2012 | 61915 | catheter,intravascular,therapeutic,short-term less than 30 days | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Two-Lumen Central Venous Catheterization Set with Blue FlexTip(R) ARROWgard Blue(R) CatheterProduct Usage:The multiple-lumen catheter permits venous access to central circulation. | Arrow TwoLumen Central Venous Catheterization Set with Blue FlexTip(R) ARROWgard Blue(R) Catheter | 1 | June 26, 2012 | 2012 | Arrow International Inc | Lidstock of product does not contain chlorhexidine contraindication and contains wording "contains no medication" on the label. Product insert does reflect medicated status. | PRODUCTION CONTROLS: Labeling Mix-Ups | Arrow International sent an Urgent Field Corrective Action letter dated May 3, 2012 to their US consignee. The letter identified the affected product, problem and actions to be taken. The letter instructed the consignee to quarantine any affected inventory and relay the information to their Arrow Sales Representative. The letter stated that the Arrow Sales Representative would affix new labels to each kit within their inventory. The consignee was instructed to sign the Acknowledgement Form provided by the Sales Representative which would then be submitted to Arrow. An Urgent Field Safety Notice letter dated May 15, 2012 was sent to foreign consignees with field safety corrective action instructions to follow. If you require additional information or clarification regarding this matter contact your sales representative. |  | 634912 | 634,912 | Worldwide Distribution - United States (nationwide) include the state of Texas. and the countries of: ARMENIA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, HUNGARY, INDIA, ITALY, IRELAND, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LITHUANIA, LUXEMBOURG, MONTENEGRO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, QATAR, RÂŽUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, TURKEY, UKRAINE, , UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | March 22, 2013 | 269 | 634912 | Not_Computer
Z-1868-2012 | 62238 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | MRCTT80041 MedComp 8F Plastic Pro-Fuse(R) CT Low Profile Port with Attachable Chronoflex(R) Polyurethane Catheter.Power Injectable Implantable Infusion Port insertion kit. | MedComp 8F Plastic ProFuse(R) CT Low Profile Port with Attachable Chronoflex(R) Polyurethane Cathet | 2 | June 26, 2012 | 2012 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect port and does not match the label. The kit is labeled as an 8F Low Profile Pro-Fuse CT Port. The kit contains an 8F Midsize Dignity CT Port. | PRODUCTION CONTROLS: Packaging Process Control | MedComp sent a "PRODUCT ALERT" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm requested the return of all affected product. Contact the firm at 215-256-4201 ext. 2225 for questions regarding this notice. | Contact the recalling firm for information | 30 | 30 | Worldwide Distribution-USA (nationwide) including the states of GA and WA and the countries of Ireland and Turkey. | Terminated | May 20, 2013 | 328 | 30 | Not_Computer
Z-1869-2012 | 61889 | radioassay, vitamin b12 | CDD | Vitamin B12test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) System B12 Flex(R) Reagent Cartridge (K6442)B12 Flex Reagent Cartridge used with the Dimension Vista System for the in vitro diagnostic test for the quantitative measurement of Vitamin B12 in human serum and plasma. | Siemens Dimension Vista(R) System B12 Flex(R) Reagent Cartridge | 2 | June 27, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that the presence of Intrinsic Factor Blocking Antibody in some patients can result in a false and significant elevation in the reported concentration of Vitamin B12 on the Dimension Vista System. This issue affects all lots of B12 Flex reagent cartridges on the Dimension Vista System. | DESIGN: Device Design | Siemens Healthcare Diagnostics Inc. sent a Urgent Device Recall letter to all affected customers dated May 2012. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed customers to immediately discontinue use of the device, quarantine and dispose of all inventory, ensure that your Medical Director reviews this letter as well as the attached physician letter. Please complete the attached Field Correction form and fax to 302-631-8467. Further questions please call 302-631-6311. |  | 59735 | 59,735 | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Great Britian, Italy, Netherlands, Norway, Portugal, Slovakia, South Africa, Spain, Switzerland, United Arab Emirates, Canada, New Zealand, and South Korea. *** Center Recommended Depth is Retail*** | Open |  |  | 59735 | Not_Computer
Z-1898-2012 | 62272 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Adult Divided Cannula end tidal C)2 sampling w/male luer-lock connector.7 ft, oxygen tube, 7 ft. CO2 tubeRX only.For Single Patient Use only.Distributed by Tri-AniumAssembled in Mexico.Usage: Delivery of Anesthesia. | Adult Divided Cannula end tidal C)2 sampling w/male luerlock connector | 2 | June 27, 2012 | 2012 | SALTER LABS | A complaint was received, and later verified, that product labeled as SO-1296 with male luer lock connector actually contained product with a female luer lock connector. | PRODUCTION CONTROLS: Packaging | The firm, Salter Labs, called their customers to discuss the recall and followed up with a "PRODUCT RECALL" letter dated May 29, 2012. The letter was distributed to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any of the lots in their possession to Salter Labs; contact Salter Labs Customer Service at 1-800-235-4203 to arrange for return of the product, and complete and return the attached form/record via fax to: 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 242 | 242 cases | Nationwide distribution: USA including states of: FL, IN, KS, PA and TX. | Terminated | July 27, 2012 | 30 | 242 | Not_Computer
Z-1897-2012 | 62056 | agent, absorbable hemostatic, collagen based | LMF | Absorbable hemostatic agent and dressing. | General & Plastic Surgery | General & Plastic Surgery |  | Ethicon Surgiflo Hemostatic Matrix Kit with Thrombin Plus Flex TipDistributed by Ethicon Ethicon Inc., Somerville, NJ 08876 USAManufactured by Ferrosan Medical Devices Ferrosan Medical Devices A/SSydmarken 5, DK-2860 Soeberg, DenmarkMixed with sterile saline or thrombin solution is indicated for surgical procedures (except opthalmic) for hemostasis when control of capillary, venous and arteriolar bleeding by pressure, ligature, and other conventional procedures is ineffective or impractical. | Ethicon | 2 | June 27, 2012 | 2012 | Ethicon, Inc. | The lyophilized thrombin component of the Surgiflo Hemostatic Matrix kit with Thrombin has the potential to be non-sterile due to an undetected air leak which occurred during the manufacturing process. | PRODUCTION CONTROLS: Packaging Process Control | Ethicon Inc. sent a Urgent Voluntary Product Recall letter dated April 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you have any product subject to this voluntary recall, please remove it from your inventory immediately and return it per the following instructions.Return product to Stericycle using the enclosed pre-paid shipping label.Shipping fees wil be charged directly to Ethicon through the use of this shipping label. No reimbursements will be made for the shipping charge if this label is not used. Extra shipping labels may be obtained by calling 1-866-338-2611. Credit will be issued for impacted product that is returned.Further questions please call 1-866-338-2611, or contact your Ethicon Biosurgery Sales Representative.Thank you for your coorporation and immediate assistance.. |  | 27396 | 27,396 kits | USA (nationwide) | Open |  |  | 27396 | Not_Computer
Z-1902-2012 | 62017 | colorimetry, acetaminophen | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | MULTIGENT Acetaminophen for use with Architect/AerosetCat. No. 2K99-20.Assay is intended for the quantitative determination of acetaminophen in human serum or plasma. | MULTIGENT Acetaminophen for use with Architect/Aeroset | 2 | June 27, 2012 | 2012 | Sekisui Diagnostics Llc | Discoloration of the acetaminophen enzyme reagent (R1), generates a high calibration factor and erratic control recovery. | PRODUCTION CONTROLS: Storage | Sekisui Diagnostics Llc sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop using the affected product and to destroy any remaining inventory. A Confirmation of Notification Form was attached for customers to complete and return via fax to 902-628-6504. Contact the firm at 902-628-0984 for questions regarding this notice. | Contact the recalling firm for information | 1600 | 1600 kits | Illinois | Terminated | November 14, 2012 | 140 | 1600 | Not_Computer
Z-1880-2012 | 61731 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | INDIANA/OHIO HEARTProduct Usage:The Sarns Cardioplegia Cannula and Cardioplegia/Vent Catheters are indicated for placement in the aortic root to deliver cardioplegia solution or blood to the heart, to vent the left heart, and to aspirate air from the aorta during cardiopulmonary bypass procedures. | Sarns | 2 | June 27, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of the Sarns Antegrade Cardioplegia Cannula, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of a foreign substance on the inner surface of some cannulae tips.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface. The substance was likely deposited during the molding process, but the exact composition of the | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Terumo sent an Urgent Medical Device Recall letter dated May 9, 2012 to All US first consignees will receive a removal notification letter via Federal Express. The letter identified the affected products, reason for correction and actions to be taken. Customers were instructed to discontinue use of affected products and return for replacement. The letter advised customers to assure that all users are aware of this notice and to fill out and return the attached Response Form. For questions call Terumo CVS Customer Service at 1-800-521-2818 or by fax at 1-800-292-6551. |  | 53139 | 53,139 (each) in total | Worldwide Distribution - US (nationwide) including the states of: AL, AR , AZ , CA, CT , FL, IA , IL, IN , KS , LA , MI , NM, MO , NC, NE , NJ , NV, NY , OH, OK , PA, RI , SC , TN, TX, WA , WI, WV and the countries of: BELGIUM, BRAZIL, Byelorussian SSR, CAMBODIA , Dominican Republic, Honduras, Japan, Malaysia, Mexico, Nicaragua, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad and Tobago, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | April 24, 2013 | 301 | 53139 | Not_Computer
Z-1871-2012 | 61670 | toothbrush, ionic, battery-powered | MMD | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Colgate Adult Motion Whitening Battery Toothbrush UPC 035000691507Dist. byColgate-Palmolive Co.New York, NY 10022Made in China.For teeth cleansing. | Colgate Adult Motion Whitening Battery Toothbrush | 2 | June 27, 2012 | 2012 | Colgate Palmolive Company | Brush battery cap was forcibly blown off at the end of the brush handle. |  | Colgate-Palmolive Company sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter requests the termination of sale and segregation of the affected product for return. Contact your Colgate-Palmolive representative to coordinate return of the affected product or if you have any questions concerning this matter. | Contact the recalling firm for information |  |  | Nationwide Distribution. | Terminated | June 19, 2013 | 357 |  | Not_Computer
Z-1895-2012 | 61952 | prosthesis, shoulder, non-constrained, metal/polymer cemented | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm; the products are medical instruments intended for use when performing surgery of the shoulder. | Bigliani/FlatowThe Complete Shoulder Solution Trabecular Metal Glenoid Instr Drill w/Stop | 2 | June 27, 2012 | 2012 | Zimmer, Inc. | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm. Catalog number 47-4301-031-00 was actually packaged and labeled as catalog number 47-4307-031-00, lot 62071190. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Zimmer sent URGENT DEVICE REMOVAL NOTIFICATION emails to distributors on May 12, 2012 to remove the product from all locations in their territories. This was an initial notification to prevent further distribution and use of the product. Zimmer Inc. sent follow-up notifications dated June 7, 2012 to the distributors, surgeons and hospital user accounts describing the product issue and providing recommended actions. Customers were directed to (877) 946-2761 for assistance or questions. | Contact the recalling firm for information | 87 | 87 units | Nationwide: TX, FL, MS, MI, OH, KS, TN, KY, OK, UT, CO, CA, IL, NC, NJ, ND, and WI. | Terminated | September 12, 2013 | 442 | 87 | Not_Computer
Z-1903-2012 | 61966 | enzymatic method, creatinine | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Creatinine test strips.PTS Panels Creatinine Test Strips are intended to measure creatinine in whole blood. Creatinine measurements are used in the diagnosis and treatment of renal (kidney) diseases and in the monitoring of renal dialysis. This system is intended for the professional use for accurately and reproducibly measuring creatinine in human whole blood obtained from a finger stick sample | PTS Panels Creatinine test strips. | 2 | June 28, 2012 | 2012 | Polymer Technology Systems, Inc. | During a routine product investigation, a creatinine sample with a value less than 1.0 mg/dl was tested on a version 2.60 CardioCek PA meter and the meter provided a result of 0.08mg/dl, which is not the reportable measuring range for the CardioChek PA meter for the creatinine assay. | DESIGN: Software Design | Polymer Technology Systems, Inc sent a URGENT FIELD CORRECTION letter dated May 15, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers to discontinue use and dispose of the product. PTS Inc. will provide appropriate replacement for any product that is properly identified on the attached Customer Acknowledgement and confirmation letter,If you have any questions please contact Customer Service toll-free at 1-877-870-5610, or +1-317-870-5610 ( outside the US ). | No consumer action necessary | 1650 | 1650 vials | Worldwide Distribution -- Nationwide Distribution including the states of CA and VA., and the countries of China, Poland, Ireland and France. | Terminated | July 20, 2012 | 22 | 1650 | Not_Computer
Z-1909-2012 | 61749 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ULTIMA*LX STM,POL 10/12TAP/SZ4Packaging: The product was packaged within inner and outer PETG blister trays and sealed with coated Tyvek lids. Foam protectors were in direct contact with the stem to help immobilize the product within the package. The sterile barrier system was then packed within a carton and shrink wrapped.Intended Use: The ULTIMA LX Cemented Polished Femoral Hip Stems product are forged CoCr polished cemented stems which are used as part of a total hip arthroplasty. These stems were indicated for use with all DePuy bearings and acetabular cups. | ULTIMALX STM | 2 | June 29, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Field Safety Notice for the LUSTER Cemented Polished Femoral Hip Stems and the ULTIMA LX Cemented Polished Femoral Hip Stems that were implanted with any metal-on-metal (MoM) articulation (including a head of CoCr and an acetabular cup where the cup is of an all metal construction (monoblock or modular)).The LUSTER Cemented Polished Femoral Hip Stems and the U | DESIGN: Device Design | DePuy sent an "URGENT RECALL INFORMATION-FIELD SAFETY NOTICE" letter dated May 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 574-372-7333 for questions regarding this notice. | Contact the recalling firm for information | 431 | 431 ULTIMA LX Cemented Polished Femoral Hip Stems in 78 lots | Nationwide Distribution-including the states of CA, CO, IA, NC, OH, TX, WA, WI, and WV. | Terminated | March 05, 2013 | 249 | 2414 | Not_Computer
Z-1905-2012 | 61555 | wheelchair, powered | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, TDX SR Power Wheelchair, TDX SI Power Wheelchair. Model Number(s): TDXSI, TDXSI-2, TDXSI-2-S, TDXSI-CG, TDXSI-HD, TDXSI-HD-S, TDXSIV, TDXSIV-2, TDXSIV-2-S, TDXSIV-HD, TDXSIV-HD-S, TDXSP, TDXSPBASE, TDXSP-CG, TDXSPEURO, TDXSP-MCG, TDXSP-MCG-GT, TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-MCG, FDX, FDX-CG, FDXEURO-14, FDX-MCG. Mrf. By: Invacare Corporation, 1200 Taylor Street, Elyria, OH 44035.Intended to provide mobility to a person restricted to a sitting position. | TDX SP Power Wheelchair, TDX SR Power Wheelchair, TDX SI Power Wheelchair | 2 | June 29, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because of a potential risk of fire and serious injury. The post fuse assemblies could potentially have a short between the positive ring terminal and the terminal battery bracket which would create an alternate current path around the fuse. | TRAINING: Employee Error | Invacare sent an "URGENT: RECALL INFORMATION" letter dated June 7, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. THe letters instructs customers to examine their inventory for the affected product and quarantine any product found. AProvider Response Form was enclosed for customers to complete and return via fax to 888-345-0289. Contact the firm at 888-641-9738 for assistance with this recall. | No consumer action necessary | 680 | 665 affected power wheel chairs and 15 post-fuse assembly service parts have been identified. | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Puerto Rico and the countries of Australia, Germany, New Zealand. | Open |  |  | 680 | Not_Computer
Z-1904-2012 | 61979 | system, test, carbohydrate antigen (ca19-9), for monitoring and management of pancreatic cancer | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT CA 19-9XR Reagent Kit; ARCHITECT CA 19-9 Reagent Kit, list 291-20 (4 x 100 tests), 2K91-25 and 2K91-27 (1 x 100 tests); Product Usage:The ARCHITECT CA 19-9XR assay is a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of 1116-NS-19-9 reactive determinants in human serum or plasma on the ARCHITECT i System. The ARCHITECT CA-19-9XR assay is to be used as an aid in the management of pancreatic cancer patients in conjunction with other clinical methods. | ARCHITECT CA 199XR Reagent Kit | 2 | June 29, 2012 | 2012 | Abbott Laboratories | The six ARCHITECT CA 19-9XR reagent lots are demonstrating a shift up in patient results. The Abbott Controls however do not detect the upward shift and will test within range. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required" letter dated May 30, 2012 to all Architect CA 19-9XR customers who received the affected reagent lots. The customers were directed to immediately discontinue use and destroy any remaining inventory of the lots listed; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. For questions were call your local Customer Service representative or call the field action coordinator at 847-938-1923. | Contact the recalling firm for information | 8672 | 316 kits - USA, 8,356 kits ROW | Worldwide Distribution - US (nationwide) the states of: Alabama, Arkansas, Arizona, California, Florida, Iowa, Indiana, Louisiana, Michigan, Minnesota, New Jersey, Nevada, New York, Ohio, Oklahoma, Pennsylvania, South Dakota, Texas, Virginia, Washington, including Puerto Rico, and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Gaza & Jericho, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kyrgyzstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open |  |  | 8672 | Not_Computer
Z-1936-2012 | 61998 | quality control slides | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics 217 Osseo Avenue North, St. Cloud, MN 56303, Gram Stain Control Slides, SLO3-10, ATCC Licensed Derivative, MediMart Europe, 11 rue Emile, Zola B.P. 2332, 38933 Grenoble Cedex 2 France.They are microscope slide preparations contains specific organisms populations of known and predictable characteristics. These slide support formal quality assurance programs by serving as quality control challenges to demonstrate and document satisfactory performance of staining regents and methods used in staining procedures and to document personnel's ability to perform the procedure and to properly interpret the microscopic test results. | KWIKQC Gram Stain Slides | 2 | July 03, 2012 | 2012 | Microbiologics Inc | Microbiologics is recalling a number of lots of KWIK-QC Gram Stain Slides. This product gram negative control organism (E.coli) is not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could potentially lead to misdiagnosis and improper treatment of a patient. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Microbiologics sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers not to use or distribute the affected product and to quarantine and return to the firm any remaining product. A Customer Response Form was attached for customers to complete and return to the firm. Contact Customer Service at 320-229-7057 for questions regarding this recall. |  | 13720 | 13,720 slides (10 slides per box) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, KY, ME, MA, MI, MN, MS, NJ, NM, NY, PA, PR, SC, TX, VT, and VA. and the countries of SWITZERLAND, ITALY, ISRAEL, IRELAND, COLOMBIA, SWEDEN, INDIA, SINGAPORE, SERBIA, SAUDI ARABIA, CANADA, TAIWAN, SOUTH AFRICA, FRANCE, BRAZIL, MALTA, MACEDONIA, MALAYSIA, FINLAND, ECUADOR, VENEZUELA, REPUBLIC OF KOREA, NETHERLANDS, PERU, UNITED ARAB EMERITES, TURKEY, POLAND, and LEBANON. | Terminated | January 08, 2013 | 189 | 13720 | Not_Computer
Z-1934-2012 | 62240 | countersink | HWW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4 Cervical (S4C) Revision Instruments (FW099R and FW101R). The S4C Revision Instruments assist in the removal of an occipital cervical screw during revision surgery. | Aesculap S4 Cervical (S4C) Revision Instruments (FW099R and FW101R) | 2 | July 03, 2012 | 2012 | Aesculap Implant Systems LLC | Aesculap Implant Systems, LLC has initiated a voluntary removal of the S4 Cervical (S4C) Occiput Revision Instrument that is part of the S4C Occipital Instrument Set. The instruments may not thread into the screw properly during a revision surgery due to tolerance issues between the two instruments. | DESIGN: Component Design/Selection | The firm initiated their recall of this product by issuing an "Important Product Removal Notification" letter to all consignees dated March 7, 2012. The letter described the problem, advised them not to use the product, and instructed consignees to return the instruments to the recalling firm with the attached Distribution Inventory Sheet. Customer Service may be contacted if questions at 1-866-229-3002. |  | 15 | 15 | Nationwide distribution, including the states of CA, CO, IL, LA, NV, OH, and TX . | Terminated | August 08, 2012 | 36 | 15 | Not_Computer
Z-1933-2012 | 62137 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | PathFinder NXT Pivoting Percutaneous Rod HolderThis device is an instrument used during orthopedic surgery. The intended use of the pivoting percutaneous rod holder is to attach to the rod to facilitate rod insertion, positioning, and orientation through the extender sleeves, under the fascia. The rod is secured to the holder by tightening a captive set screw against the rod end in the tip of the holder using the set screw driver. After positioning the rod and securing at least one pedicle screw closure top against the implanted rod, the holder is removed from the rod end by loosening the captive set screw with the set screw driver. | PathFinder NXT Pivoting Percutaneous Rod Holder | 2 | July 03, 2012 | 2012 | Zimmer, Inc. | Complaints have been reported where, during surgery, the surgeon was unable to loosen orremove the inserted rod from the rod holder instrument. The complaint investigation showedthe hex mating feature of the set screw was broken and material missing. There were noreports of fragments of the set screw being left in the patient or that there was any patient oruser injury, but these events we | DESIGN: Device Design | The firm, Zimmer Spine, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the notification and ensure they are aware of the content; to immediately remove and discontinue use of Pivoting Percutaneous Rod Holder (part number 3564-1) from Pathfinder NXT instrument sets in their possession and follow the PathFinder NXT Surgical Technique; to notify their Zimmer Spine customer service representative to obtain return authorization; and return the complete Surgeon Contact Certification Form within 10 working days from receipt of notification to ronald.musselman@zimmer.com or Fax to 1-512-258-0995.If you have any questions, call Zimmer Spine Customer Service at 866-774-6368. |  | 77 | 77 units | Worldwide distribution: USA (nationwide) and countries of: Australia, France and Germany. | Terminated | February 14, 2013 | 226 | 77 | Not_Computer
Z-1941-2012 | 57350 | vehicle, motorized 3-wheeled | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 9 with Power Seat (SC609PS)Motorized wheelchair with elevating seat. | Victory 9 with Power Seat | 2 | July 03, 2012 | 2012 | Pride Mobility Products Corp | The main harness connector was manufactured without the required back shield sealing protection. As a result, there is the potential for water to seep through the nonconforming harness. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Three (3) pieces of information was distributed in the notification. They are: (1) Letter of Correction to the Provider (2) Letter of Correction for the Consumer (provider to issue to consumer) (3) Product Insert for Victory 9 Controller Interface Harness Replacement. The letter of Correction for Victory 9 with Power Sear (SC609PS), dated November 2010, identified the affected product and the reason for the correction. Customers were to record the serial number(s) and contact their provider's office to make an appointment to have the harness replaced. |  | 186 | 186 | Worldwide Distribution | Terminated | March 22, 2013 | 262 | 186 | Not_Computer
Z-1935-2012 | 62447 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Vacuette Quickshield Complete Plus, 21G x 1 (0.8 x 25 mm), Greiner bio-one.The Quick shield Complete Plus is intended to be used only with Vacuette Blood collection Tubes as a system in routine venipuncture procedures. They are for single-use only and should only be used by adequately trained healthcare personnel in accordance with instructions for use. | Vacuette Quickshield Complete Plus | 2 | July 03, 2012 | 2012 | Greiner Bio-One North America, Inc. | : The Quickshield Complete Plus may have a defect (split at the base), which can cause the green safety shield attachment of the Quickshield holder to be loose and unstable. | PRODUCTION CONTROLS: Process Control | greiner bio-one sent a notification letter dated April 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Product Disposition Site Confirmation form and fax to 800-726-0052 or email to patech@gbo.com. This form will confirm the affected product has been discarded/destroyed. The firm will replace the product after the completed form is returned to them. Customers with questions should call 800-515-8112.For questions regarding this recall call 704-261-7823. |  | 19.5 | 19.5 cases (5,850 pieces) and 244 cases (73,200 pieces)= 263.5 (73258 pieces total) | Nationwide Distribution including IL, MS, OH, PA, TN and VA. | Terminated | February 27, 2013 | 239 | 263.5 | Not_Computer
Z-1939-2012 | 61874 | injector, vertebroplasty (does not contain cement) | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Spine Side Fire Needles 11G X 6IN Side-Fire Needle, part of the Confidence Spinal Cement System Product Code 283904611.The Confidence Side Fire Needle is used to place bone cement in an intra-vertebral fashion at a desired location. the cement extrudes fro the needle through a port on the side of the needle. | Depuy Spine Side Fire Needles | 2 | July 03, 2012 | 2012 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Additional Instructions for Use and Surgical Technique for the Confidence Needle due to reports of breakage with the side fire needles | DESIGN: Labeling Design | Depuy Spine sent a "CORRECTION NOTICE" to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Notification Form was attached for customers to complete and return via fax to 508-828-3070. For questions about Confidence Needles, please contact your local DePuy Spine Sales Representative. If you need assistance in locating your local Sales Representative please contact JJHCS CIS @ 1 -800-255 -2500 . | Contact the recalling firm for information | 530 | 530 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, China, Croatia, Czech Republic, France, Germany, Greece, Hungary, Italy, Lebanon, Netherlands, New Zealand, Poland, Russian Federation, Singapore, South Africa, Spain, Switzerland, Turkey, UAE, and UK. | Open |  |  | 8896 | Not_Computer
Z-1943-2012 | 61525 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Single Shot Epidural Trays, Sterile, Integra Pain Management, West Valley City, Utah 84119, each kit contains sterile instruments, drugs and other medical materials that will support the administration of spinal anesthesia for use by a surgeon, nurse or other licensed medical professional. Model number is 3403644.The intended use of this tray is to provide the surgeon, nurse or other licensed medical professional, a kit of sterile instruments, drugs and other medical materials that will support the administration of anesthesia, be it local, spinal or general administration. | Single Shot Epidural Tray | 3 | July 05, 2012 | 2012 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is recalling one lot of Single Shot Epidural Trays as the product was inadvertently shipped missing injectable components. | PRODUCTION CONTROLS: Packaging Process Control | Integra sent an Urgent - Medical Device Recall letter dated June 7, 2012, to one domestic consignee via Federal Express with a recall acknowledgement form followed up with a telephone call. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the Recall Acknowledgement Form including quantity of affected product in their inventory. Customers were also instructed to contact Integra Pain Management Customer Service to arrange for returns of all affected product. Product reconciliation will be performed to account for all affected products returned. Returned product will be reprocessed to correct for the missing components.Customers with questions were instructed to call Integra Pain Management Customer Service at 1-800-241-2210.For questions regarding this recall call 609-936-2311. | Contact the recalling firm for information | 4 | 4 cases x 10 units per case | Nationwide Distribution including MA | Terminated | August 07, 2012 | 33 | 4 | Not_Computer
Z-1956-2012 | 61793 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | CIVCO VirtuTRAX Instrument Navigator, Reusable non-sterile general purpose driveBAY, trakSTAR, electromagnetic sensor (8mm) with VirtuTRAX Instrument navigator and (8.9 tapered to 1.3 x 147cm) telescopically-folded CIV-Flex sensor covers (5), Ref. 610-1080. The firm name on the label is CIVCO, Kalona, IA. The kit contains sensor component part number 410-240.The general purpose sensor within the kit is intended to provide physicians with tools for electromagnetic tracking of instruments with respect to image data. The sensors are designed to attach to custom reusable tracking bracket for leading transducers and other medical devices | CIVCO VirtuTRAX Instrument Navigator | 2 | July 06, 2012 | 2012 | Civco Medical Instruments Inc | The general purpose electromagnetic sensor was incorrectly calibrated causing potential variability in the projected position of the needle path with the ultrasound device. | DESIGN: Component Design/Selection | CIVCO sent an Urgent Medical Device Recall letter dated May 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return all general purpose sensors as the product may not function as intended. To return the affected product customers should contact the OEM Sales Coordinator at 319-248-6521 to schedule the return of their device. Upon receipt of the snesor, CIVCO will arrange for recalibration. Once recalibration is completed, the sensor will be returned to the customer at no charge. Customers were asked to forward this communication to all their customers who purchased the affected product. Once CIVCO has received the customer list, they will contact the end-user to request they return their general purpose sensor directly to CIVCO for recalibration. For questions customers should call 800-441-6741.For questions regarding this recall call 319-656-4447. |  | 24 | 24 kits | Worldwide Distribution - USA including WI, CA, MA, IN, WA, and NJ. Foreign distribution was made to Germany, Canada, Denmark, Hong Kong, Ukraine, Japan, Italy, Korea, and France. | Open |  |  | 384 | Not_Computer
Z-1945-2012 | 61807 | incubator, neonatal | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i, (Atom Infant Incubator Model 100) Atom Medical Corporation 3-18-15 Hongo, Bunkyo-Ku, Tokyo, Japan 113-0033 Usage: An incubator for infants. | Dual Incu i | 2 | July 06, 2012 | 2012 | Atom Medical Corporation | There is a potential for the device's canopy to move unintentionally when using the incubator mode or radiant warmer mode. | DESIGN: Device Design | Atom Medical sent an Urgent Field Safety Notice dated May 1, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify all affected product in their inventory and quarantine them until the rework was complete. Customers were also asked to identify any customers to whom they further distributed the product and to contact their customers at once and arrange a time in whcih the filed service technician could perform the rework. Customers were asked to return the Recall Response Forms to Atom Medical (kirautc@atomed.co.jp, tsugino@atomed.co.jp,or matsumoto@atomami.co.jp). All forms should be completed and returned by January 31, 2013.For questions regarding this recall call 512-327-9997. | No consumer action necessary | 377 | 377 units | Worldwide Distribution - USA including PA. Internationally to Japan, Kuwait, Saudi Arabia, Lenanon, Italy, Spain, China, East Europe, Slovakia, Poland, Czech Republic, Ukraine, Egypt, Russia, Algeria, Colombia, Germany, Malaysia, Switzerland, Australia, Europe, Portugal, Iran, Israel, Philippines, UAE, Korea, and Netherlands | Terminated | January 31, 2013 | 209 | 377 | Not_Computer
Z-1946-2012 | 61832 | computer, blood-pressure | DSK | Blood pressure computer. | Cardiovascular | Cardiovascular | 510(k) | Safedraw Venous Close Loop System, Blood Sampling Set, packaged in a tyvek bag, Model/Part Number 688695, Argon Critical Care Systems, Singapore Pte. Ltd., 198 Yishun Avenue 7, Singapore 768028.The blood sampling set is intended to be used for closed-loop blood sampling from the patient without the system being exposed to the outside environment. | Safedraw Venous Close Loop System | 2 | July 06, 2012 | 2012 | Argon Medical Devices, Inc | Some of the kits may exhibit a potential defect with the pouch seal. The pouch issue may compromise the sterile barrier. | PRODUCTION CONTROLS: Packaging | Argon Medical Devices sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Recall Response Form was attached to the letter for customers to complete and return to the firm via fax to 903-577-9393. Contact the firm at CRS@argonmedical.com for questions regarding this recall. | No consumer action necessary | 10 | 10 pieces | Distributed only in Missouri. | Terminated | January 14, 2013 | 192 | 10 | Not_Computer
Z-1931-2012 | 61927 | test,natriuretic peptide | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP PN 98000XRIntended to be used as an aid in the diagnosis and assessment of severity of congestive heart failure (also referred to as heart failure). | Alere Triage BNP PN 98000XR | 1 | July 06, 2012 | 2012 | Alere San Diego, Inc. | Certain lots of the affected products may have significantly decreased precision relative to the package insert, which could result in an increased frequency of false positive or false negative results. These false positive and false negative results are unpredictable within lots and may not be detected by Quality Control testing. | DESIGN: Process Design | Phase 1 of the recall included a recall communication being initiated on May 22, 2012 with Alere forwarding a recall letter to all their customers who purchased the Triage CardioProfiler Panel, Triage Cardiac Panel, Triage Profiler SOB Panel, Triage BNP, and Triage D-dimer. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use of the affected product, and to complete and fax the enclosed verification form within 10 days. Phase 2 of the recall included a recall communication letter being forwarded to customers on 6/11/12. The second letter included additional lots of Triage products that go beyond the lots included in an earlier recall notification letter. The letter informed the customers that the additional lots have significantly decreased precision relative to the package insert which could result in an increased frequency of false positive or false negative results. A third letter dated June 12, 2012 included 3 more lots for the Triage Cardio Profiler Panel PN 97100CP and Alere Triage Profiler SOB Panel PN 97300. The letter informed the customers that the additional lots have an increased frequency of Troponin I results |  | 30083 | 30,083 kits | Nationwide Distribution | Terminated | October 18, 2013 | 469 | 60074 | Not_Computer
Z-1947-2012 | 62116 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Encore Mobile Patient Lifter; manufactured by Arjo Ltd., St. Catherine Street, Gloucester, GL1 2SL, England, U.K.; Models KKB5022-21US and KKB5042-21US.The Encore is designed to be used as a standing and raising aid, it is useful for short transfers from one sitting position to another, and to elevate a patient for toileting, repositioning, changing of incontinence pads or wound dressings, standing practice etc. it is not intended for long periods of suspension or transportation. | Encore Mobile Patient Lifter | 2 | July 06, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received a limited number of customer reports concerning injuries related to an exposed metal support bracket when the proactive knee pad assembly is removed. | MISBRANDING: Labeling False and Misleading | ArjoHuntleigh sent an "URGENT FIELD CORRECTION RECALL" notice dated June 25, 2012, to all affected customers via UPS Ground on June 29, 2012. The notice informed the accounts of the potential for severe patient injury if the proactive knee pad support is removed from the lift without the removal of the knee pad support bracket. The accounts were reminded that the entire knee pad and support bracket assembly must be removed from the lift when required to do so for exercise. An additional warning statement with diagrams was included in the notice. The accounts were requested to disseminate the information to all users of the affected lifts, to include a copy of the notice in the Instruction for Use Manuals for their lifts, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 7522 | 2,552 units in the U.S., 4,970 units in the rest of the world | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, Bermuda, Canada, Cyprus, Czech Republic, Denmark, Dubai, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Hong Kong, Iceland, Ireland, Israel, Italy, Japan, Kenya, Korea, Luxembourg, the Netherlands, New Zealand, Northern Ireland, Norway, Poland, Portugal, Quatar, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland and Turkey | Open |  |  | 13617 | Not_Computer
Z-1967-2012 | 52194 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Rigid Suction Wand, sterile. Model numbers S033 and SPC2081.Product Usage: These suction catheters are typically used to remove fluid on the external surface of the heart during cardiac surgery and therefore there is little risk of intravascular involvement. | Rigid Suction Wand | 2 | July 09, 2012 | 2012 | Edwards Lifesciences, LLC | Edwards is recalling Suction Wand models SPC2081 and S033 because the plastic tip attached to the metal suction wand may become detached during routine use. | PRODUCTION CONTROLS: Process Control | Edwards Lifesciences sent an Urgent Field Safety Notice Product Recall letter dated June 2012 to all affected customers. The letter identified the affected product, description of the issue and actions to be taken. Customers were instructed to examine their inventory for the affected product, quarantine and return any remaining in stock and complete the enclosed acknowledgment form with fax instructions to assure that the notification was received. For question call Edwards Customer Service at (800) 424-3278 from 8:00AM-4:30PM Pacific Time. | No consumer action necessary | 1524 | 1524 units | Worldwide Distribution - US (nationwide) including the states of CT, TX, CA, AZ, MO, and MD and the countries of Europe and Japan. | Terminated | September 10, 2012 | 63 | 1524 | Not_Computer
Z-1979-2012 | 62389 | warmer, thermal, infusion fluid | LGZ |  |  | General Hospital | 510(k) | Smiths Medical HOTLINE 3 Blood and Fluid Warmer,Reorder Numbers: HL-390Product Usage:The Micro-switch is used on the HOTLINEÂŽ 3 Blood and Fluid Warmer. This product, for use in conjunction with HOTLINEÂŽ 3 L-370 Fluid Warming Set are intended for the delivery of warmed blood and intravenous fluids at normothermic temperatures. These products are intended for use by appropriately trained healthcare professionals in clinical environments | Smiths Medical HOTLINE 3 Blood and Fluid Warmer | 2 | July 09, 2012 | 2012 | Smiths Medical ASD, Inc. | HOTLINE 3 Blood and Fluid Warmer( Reorder No. HL-390 ) disposable is not correctly installed , the microswitch may become damaged and unusable | DESIGN: Device Design | Smiths Medical sent an Urgent Medical Device Correction -Field Safety Notice letter dated June 21, 2012 via U.S. Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter included an Urgent Medical Device Field Safety Notice Acknowledgement Form ("Confirmation Form") that customers were instructed to complete and return to Smiths Medical for carrying out the correction and affectivity tracking purposes. The letter stated that a Smiths Medical Representative will visit the affected sites to correct the units with a shelf upgrade kit, which includes instructions and all the parts necessary to make the correction to their affected devices. Customers were provided with a form to record the serial numbers of the devices corrected so that this can be returned to Smiths Medical. For questions regarding this information call Smiths Medical Quality Department at 1-781-763-9335.. |  | 2247 | 2247 units | US Nationwide Distribution | Terminated | February 20, 2014 | 591 | 2247 | Not_Computer
Z-1978-2012 | 62319 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Military Kit Portable Clinical Analyzer ; Abbott Point of Care, Princeton, NJ 08540Model 200. The 37 degree tests for sodium, potassium, ionized calcium, hematocrit, pH, PCO2, PO2, and the i-Stat 200 analyzer as port of the i-Stat System, are intended for use in the in vitro quantitization of arterial, venous, or capillary whole blood 510 K940918 Model 300. The i-Stat Model 300 Portable Clinical Analyzer is intended to be used by trained medical professional for use with i-Stat test cartridges and MediSense blood glucose test strips. i-State cartridges comprise a variety of clinical chemistry tests and test panels. 510K 001387 | Military Kit Portable Clinical Analyzer | 2 | July 09, 2012 | 2012 | Abbott Point Of Care Inc. | Some i-Stat Analyzers become uncomfortably hot to touch in the are of the battery. This is a recall extension. | DESIGN: Device Design | Abbott Point Care sent an Urgent Product Recall letter dated September 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm the number of affected product at their facility that were using the Old Battery Carrier. Fill out the affected Buesiness Reply Card and submit their facility information as per the instructions on the card. Upon the receipt of the New Battery Carriers, replace the Old Battery Center with the New Battery Center. Customers were instructed to immediately dispose of all Old Battery Carriers per local regulations. If customers forward the affected product to another facility, they should provide a copy of the recall letter to them and advise them to replace the carriers if required. For any questions call 800-366-8020, Option 1 or your Abbott Point of Care representative.For questions regarding this recall call 609-454-9272. |  | 215804 | 04J48-01-56 units; 215000-695 units | Worldwide Distribution - USA (nationwide) | Terminated | January 28, 2013 | 203 | 751 | Not_Computer
Z-1972-2012 | 62441 | warmer, thermal, infusion fluid | LGZ |  |  | General Hospital | 510(k) | enFlow IV Fluid Warmer (510k name: EFLOW IV FLUID WARMER, MODEL100, EFLOW DISPOSABLE CARTRIDGE, MODEL 200 AND EFLOW AC POWER SUPPLY,MODEL 120).The Enginivity enFlow TM IV Fluid Warmer is intended for warming blood, blood products and intravenous solutions prior to administration. It is intended to be used by healthcare professionals in hospital, clinical and field environments to help prevent hypothermia. | enFlow IV Fluid Warmer | 2 | July 09, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | The EnFlow Warmer Strapdoes not meet the requirements of the biocompatibility standard (ISO 10993) for products that mayhave contact with skin for less than 24 hours. | DESIGN: Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included customer instructions and a Customer Response Form to be completed and returned to the firm. Contact Customer Service at 1-800-932-0760 for questions regarding this recall. |  | 11803 | 11803 cases | Nationwide Distribution. | Terminated | August 19, 2013 | 406 | 11803 | Not_Computer
Z-1961-2012 | 62346 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Stryker brand Arthroscope and Hardware Set Tray (sterilization tray); Model/Part Number: 233-032-116; Product is manufactured and distributed by Stryker Endoscopy, San Jose, CAThe Arthroscope and Hardware Set Tray is used to hold and protect surgical devices during the sterilization process. Specifically this tray is used for arthroscopes and arthroscopic hardware devices. The tray consists of an interlocking tray and lid which are perforated to allow for the passage of sterilizing agent from outside the tray to the inside of the tray where the devices are held. | Stryker | 2 | July 09, 2012 | 2012 | Stryker Endoscopy | The gravity steam, Ethylene Oxide (EO) and STERRAD 100S sterilization methods are being removed from the "Instructions for Use"; due to the firm's evaluation that identified original validation was not done at worst case conditions. | PRODUCTION CONTROLS: Process Control | Stryker Endoscopy sent a Urgent Device Correction letter dated June 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to:1. Discard any old Instructions for Use for your Arthroscope and Hardware Set Tray.2. Download the new IFU from the URL link: http://stryker.com/arthrotray .Complete and sign the enclosed acknowledgement of receipt form and fax to(408) 754-8378 or scan in and email to trays@stryker.com.* It is important to send the self addressed confirmation form back to Stryker.Stryker is committed to delivering the highest quality products to its customers. We sincerely apologize for any inconvenience this action may cause.If you have any questions regarding this letter please contact us at 1-800-624-4422 or via email at trays@stryker.com.. | Contact the recalling firm for information | 6388 | 6388 units | Worldwide Distribution--USA (nationwide) | Terminated | October 23, 2012 | 106 | 6388 | Not_Computer
Z-1959-2012 | 62406 | catheter, peripheral, atherectomy | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | CSI Stealth 360-degree Orbital Atherectomy Device (OAD), Part Number 70056-11, Model Number PRD-SC30-125S. Sterilized with Ethylene Oxide. The Stealth 360-degree Orbital Atherectomy Device (OAD) is part of the Stealth 360-degree Orbital PAD system.The Stealth 360-degree Orbital PAD system is a percutaneous orbital atherectomy system indicated for use as therapy in patients with occlusive atherosclerotic disease in peripheral arteries and who are acceptable candidates for percutaneous transluminal atherectomy. | Stealth 360degree Obital Atherectomy Device | 2 | July 09, 2012 | 2012 | Cardiovascular Systems, Inc. | CSI has initiated a recall on specific lot numbers of Stealth 360-degree Orbital Atherectomy devices which have the potential for an insufficient crown solder bond that may result in the crown becoming detached from the driveshaft. | TRAINING: Employee Error | Phone calls to affected sites began June 15, 2012 to quarantine the product. These same sites were also notified with an Urgent Medical Device Recall letter, dated June 18, 2012, via FedEx and e-mail (where available) on June 18, 2012. The letter notified each customer to remove from use and return the recalled product. The customer was instructed to complete a Customer Acknowledgement Form and return it with the device(s) being returned in the Returned Goods Authorization kit provided. Customers with questions can contact customer service at 1-877-274-0901.For any questions regarding this recall call 651-259-2805. | Contact the recalling firm for information | 27 | 27 | Nationwide distribution including CA, FL, IL, MI, MN, PA, and TX. | Terminated | July 19, 2012 | 10 | 27 | Not_Computer
Z-1981-2012 | 62325 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Pathfinder NXT Fixed Percutaneous Rod Holder, Part numbers 3573-2 (straight handle) Stainless Steel / Silicone Rubber, Rx only, non sterile.Product Usage:PathFinder NXT is a minimally invasive pedicle screw system. The PathFinder NXT System consists of various screws and rods intended to provide temporary stabilization following surgery to fuse the spine. The intended use of the Fixed Angled Percutaneous Rod Inserter is to insert a rod percutaneously through the extender sieeves to position into the implanted screw heads for securing with closure tops forming a construct that provides temporary stabilization following MIS or mini-open surgery to fuse the spine. | Pathfinder NXT | 2 | July 10, 2012 | 2012 | Zimmer, Inc. | Zimmer is recalling Pathfinder NXT Fixed Percutaneous Rod Holder, because complaints have been received that the rod holder tip maycrack or break when the rod is manipulated during surgery. | DESIGN: Device Design | Zimmer sent an Urgent Medical Device Correction and Removal notification dated June 5, 2012 to distributors and all affected accounts of Zimmer Spine. The notice identified the affected products, problem and actions to be taken. Distributors were asked to hand deliver Recall Notices to Surgeons and return the completed Surgeon Contact Certification Form within 10 working days to ronald.musselman@zimmer.com or FAX to 1-512-258-0995. Notices to Surgeons and User facilities included revised directions/precautions for the Instructions for Use and attached Surgical Techique. Devices will be removed once newly design parts become available.For any related questions or assistance about this please contact Zimmer Spine Customer Service at 866-774-6368. | Contact the recalling firm for information | 249 | 249 ( angle handle and straight) in total | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Australia, Belgium, France, Gelmany, Italy, Spain, Switzerland and the United Kingdom. | Open |  |  | 249 | Not_Computer
Z-1982-2012 | 61923 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Definition PM Femoral Hip Stem Catalog Numbers 6265-3-099 through 6265-3-107. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430. A subsidiary of Stryker Corp. Stryker France, ZAC Satorlas Green Pusignan, Av De Satolas Green 69881 Meyzeiu, Cedex France. Made in Ireland.The Definition PM femoral hip stem is a forged, collared, straight stem with an integral premantle (PM) design. The Poly Methyl Methacrylate (PMMA) and Barium Sulphate (BaSO4) premantle is integral to the design as it aligns, centralizes and pressurizes the stem within the femoral canal. The forged stem is first grit blasted proximally and then the premantle is applied through an infection molding process. | Stryker Orthopaedics | 2 | July 10, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for cracking of the PMMA coating on their PM COCR Hip Stem product. | TRAINING: Employee Error | Stryker Orthopaedics sent an Urgent Product Recall letter dated January 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of any affected product. Customers were also asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers should call 201-972-2100.For any questions regarding this recall please call 201-831-5970. | Contact the recalling firm for information | 16913 | 16,913 units | Nationwide Distribution | Open |  |  | 16913 | Not_Computer
Z-1984-2012 | 62138 | shunt, central nervous system and components | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Pudenz Flushing Valve with Integral Connectors, catalog number NL850-1330. Integra Neurosciences, Plainsboro, NJ. Device is used for the treatment of hydrocephalic patients when shunting cerebrospinal fluid from the ventricles of the brain. | Pudenz Flushing Valve | 2 | July 10, 2012 | 2012 | Integra LifeSciences Corp | The pouch label of catalog NL850-1330 indicates an incorrect model and closing pressure of the device. The incorrect label reads: "Model: 12 mm-Medium Pressure" instead of "Model: 16 mm-Low Pressure". | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm initiated a voluntary recall by notifying the two consignees initially via email and then with an "URGENT: MEDICAL DEVICE RECALL" dated April 5th or 9th, 2012. The letter informed them of the labeling error and provided instructions regarding affected product. The International Customer Service number was provided 609-936-5400. | Contact the recalling firm for information | 39 | 39 units | International Only Distribution, including the countries of Thailand and Japan. | Terminated | August 02, 2012 | 23 | 39 | Not_Computer
Z-1989-2012 | 62397 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Axcess Cranial Screw (sometimes referred to as the Axcess Cranial Pin) is a self-tapping, self-drilling screw and is an accessory used in conjunction with the Axcess Cranial Tri-Pin Kit.The Axcess Bone (Cranial) Pin is intended to be used when rigid fixation of a tracking device to bone is required. | GEHC OEC Axcess Cranial Pin | 2 | July 11, 2012 | 2012 | GE OEC Medical Systems, Inc | GE Healthcare is aware of a few instances where the Axcess Cranial Screw has broken during insertion. There is a possibility that breaking of the screws could lead to a health hazard. | DESIGN: Device Design | GE Healthcare sent an "URGENT REMOVAL NOTICE" dated November 8, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer Service at 1-800-874-7378 for questions regarding this recall. | Contact the recalling firm for information | 15 | 15 kits | Distributed in Washington, D.C., the states of AL, CA, FL, NY, OR, and VA | Terminated | July 11, 2012 | 0 | 15 | Not_Computer
Z-1991-2012 | 62332 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Button Switchpens, w/Extendable Needle Electrode 10 mm Exposure, Single Use device, Electrosurgical./Extendable Pencils, 510(k)#K936304, Model Number: 75800. The device is packaged in a sterile Tyvek peel pouch within a cardboard box. The cartons contain 25 pouches per box.Pens/Pencils and Handles -- Olsen Medical Pens/Pencils and handles are designed to hold active electrodes Single Use device intended for coagulating and cutting tissue during surgical procedures. | Electrode Electrosurgical/Extendable Pencil | 2 | July 11, 2012 | 2012 | PSC Industries Inc | The Tyvek pouch label for this device states: 'with holster'; however, this product does not come with a holster. Firm officials met with FDA on 5/17/2012. At which time, they reported various labeling errors with their Button Switch Pen Electrode Electrosurgical /Extendable Pencils. | PRODUCTION CONTROLS: Labeling Mix-Ups | Olsen Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated June 5, 2012. The letter identifies the product, problem, and actions to be taken by the customers. The letter asks that the customer immediately examine their inventory for the presence of the suspect product and quarantine the product, if found. Customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall. The customers are also asked to call a toll-free telephone number in order to received a Return Material Authorization (RMA) number and so identify if they would like credit or replacement inventory upon returning the product. The customer is also asked to complete and return the enclosed 'Recall Response Form' and to fax the completed form to the fax number on the form or to send the completed from back to the recalling firm via parcel post. |  | 725 | 725 units (125=Lot #023392 & 600 units Lot #022438) | Nationwide Distribution-including DC and the states of CA, GA, IN, KY, LA, MT, TX, UT, and WI. | Terminated | January 17, 2014 | 555 | 775 | Not_Computer
Z-1994-2012 | 61958 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Bayonet Forceps 7.75" (196.9 mm) 0.5 mm Tips Non-Stick Electrosurgical/Bipolar Forceps/Reusable, Model #90-7004 (k) #K982705. The surgical device is package non-sterile into small plastic tubes.Olsen Medical Biopolar Reusable Bayonet Forceps and cords are intended for use in coagulating tissue during surgical procedures. | Bipolar electrosurgical bayonet forceps 7.75" (196.0mm) 0.5 mm tips | 2 | July 12, 2012 | 2012 | PSC Industries Inc | Firm officials reported to CIN-DO that their Bipolar Bayonet Forceps 7.75" are misbranded in that the manufacturer symbol was replaced by the expiry symbol on the packaging. |  | The firm, Olsen Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 5, 2012 to their customers via Certified mail along with representative labeling to illustrate the specific labeling issue associated with the various types and/or sizes of Bayonet Forceps surgical instruments that the customers received. The letter described the product, problem and actions to be taken. The letter instructed the customers to immediately examine their inventory for the presence of the suspect product and quarantine the product, if found; to discontinue use immediately and promptly call (800) 251-3000 for a 'Return Material Authorization (RMA) number; to complete and return the enclosed 'Recall Response Form' and fax the completed form to the fax number on the form or to send the completed form back to the recalling firm via parcel post, and for customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall.If you have any questions call (615) 964-5515 M-F 8:00 am - 5:00 pm CST. | Contact the recalling firm for information | 7 | 7 units | Worldwide distribution: USA (nationwide) including states of: CA, CT, FL, GA, MD, MI, NC, OR, SC, and TX and countries of: Canada and Philippines. | Open |  |  | 267 | Not_Computer
Z-2012-2012 | 62281 | st2 assay | OYG | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Presage ST2 Assay, catalog numbers BC-1065 and BC-1065E.The Critical Diagnostics Presage ST2 Assay kit is an in vitro diagnostic device that quantitatively measures ST2 in serum or plasma by enzyme-linked immunosorbent assay (ELISA) in a microtiter plate format. The Presage ST2 Assay is indicated to be used in conjunction with clinical evaluation as an aid in assessing the prognosis of patients diagnosed with chronic heart failure. | Presage ST2 Assay | 3 | July 13, 2012 | 2012 | Critical Diagnostics | The recall was initiated because Critical Diagnostics has confirmed that it is possible to obtain inaccurate test results for specimens when using the Presage ST2 Assay kits. | DESIGN: Component Design/Selection | Critical Diagnostics sent a Medical Device Recall letter dated March 5, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.Customers were instructed to please use Federal Express (FedEx) to return the kits, and charge to account: 305599069 Replacement kits will be sent to you.Customers with questions about this notice were instructed to contact Critical Diagnostics' customer support at (877) 700-1250 ext 3. |  | 33 | 33 units | USA including the state of California and the countries of Austria, Spain, United Kingdom, and Italy. | Terminated | July 16, 2012 | 3 | 33 | Not_Computer
Z-2008-2012 | 62457 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Custom surgical kits and trays:GEO-MED TOTAL KNEE PACK PGYBK, REF 89-6664.04GEO-MED TOTAL HIP PACK, REF 89-6664.03surgical kits | DeRoyal Surgical Kits and Trays | 2 | July 13, 2012 | 2012 | DeRoyal Industries Inc | Surgical kits and trays contain Stryker Hytrel Togas which were recalled because the clear tape, that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DeRoyal initiated the recall on 6/16/2012 by notifying customers by overnight delivery of letters describing the problem, instructing the customer how to handle inventory on hand, and requesting the customer to destroy the product, fill out the provided notice of destruction form, and return the form to DeRoyal.For questions customers were instructed to call 865-362-1020 or 865-362-1034.For questions regarding this recall call 865-362-2334. |  | 303 | 303 units | Nationwide Distribution including SD and NE. | Terminated | November 09, 2012 | 119 | 303 | Not_Computer
Z-2011-2012 | 62024 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Bipolar Forceps model number: S02-XXX, S03-XXX, P02-XXX, and P03-XXX. "Instructions For Use" (QF7.5.1A.R01 thru QF7.5.1A.R06 Care and Handling of Bipolar Forceps. Bipolar Forceps designed to grasp, manipulate and coagulate selected tissue for use in general surgical procedures powered by a high frequency generator. | Bipolar Forceps | 2 | July 13, 2012 | 2012 | Stingray Surgical Products, LLC | On 05/23/2012, Stingray Surgical Products, LLC., Boca Raton, FL initiated a recall initiated due to "Instructions For Use" of Bipolar Forceps not being cleared by FDA and incorrect "Instructions For Use" . The sterilization instructions have been altered without the clearance of the FDA. | PRODUCTION CONTROLS: Packaging | Stingray Surgical Product sent an URGENT Instructions for Use Recall letter dated Mary 23, 2012, to all distributors. The letter identified the product, the problem, and the action to be taken by the client. Distributers were requested to convey the recall to their users along with fowarding copies of the firm's recall notification letter. IFU's will be destroyed or discarded by either user or distributor since returns will not be necessary. Distributors were asked to complete the form at the bottom of the letter and return via e-mail or fax. For questions clients were instructed to call 561-210-7582.For questions regarding this recall call 561-210-7582. | Contact the recalling firm for information | 6326 | 6326 | Worldwide Distribution - USA (nationwide) | Open |  |  | 6326 | Not_Computer
Z-2014-2012 | 61463 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | EtacÂŽ Ono Walker, EtacÂŽ Avant Walker, and EtacÂŽ Salsa Walker.These products are four-wheeled walkers intended as a walking aid for indoor and outdoor use for persons weighing up to 125 KG. | Etac Ono, Avant, and Salsa Walkers | 2 | July 14, 2012 | 2012 | Etac Supply Center Ab | The recall has been initiated because there is the potential that users may fall and become injured, with the potential for serious injury, including bone fracture. One death occurred in a user fall incident, but it is unclear whether the fall was caused by the loss of a wheel. Among the walkers distributed worldwide, EtacÂż has received a total of 49 reports from Europe and Japan of wheels falling | DESIGN: Device Design | Etac sent an "URGENT DEVICE RECALL " letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue selling and distributing the affected product. A Reply Postcard was enclosed for customers to complete and return. Contact the firm at 847-277-7705 for questions regarding this recall. | Contact the recalling firm for information | 226695 | 204,323 units of the Etac Ono Walker, 15,601 units of the Etac Avant Walker, and 6,771 units of the ETAC Salsa Walker | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, China, Colombia, Germany, Denmark, Estonia, Spain, Finland, France, Great Brittain, Greenland, Greece, Hungaria, Ireland, Israel, Iceland, Italy, Japan, South Korea, Kuwait, Lithuania, Luxembourg, Latvia, Malaysia, Netherlands, Norway, Poland, Portugal, Rumania, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, and Thailand. | Terminated | June 18, 2013 | 339 | 226695 | Not_Computer
Z-2016-2012 | 62051 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Arthrodesis Nailing System. Manufacturer: Stryker Trauma GmbH, Professor-Kuntscher-Str. 1-5, 24232 Schonkirchen Germany. Distributed in the USA by Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430. The T2 Arthrodesis Nail is intended for the treatment of Knee Arthrodesis. The design of the T2 Arthrodesis Nail features a unique curvature which incorporates both femur ante curvature and knee valgus bend. Indications: Aseptic failed total knee arthro-plasty with or without bone loss; Failed external fixation, non unions and malunions; Periarticular fractures where repair is not possible; Pathologic fractures and tumor resections; Pseudoarthrosis and correct osteotomy; Ipsilateral femur fractures; and Open and closed femoral fractures. | Stryker Osteosynthesis | 2 | July 16, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that there exists a potential for breach of the sterile barrier of packaging associated with certain lots of T2 Knee Arthrodesis Nails. | PRODUCTION CONTROLS: Packaging | Stryker Orthopaedics sent Recall letters/Customer Response forms dated 1/26/2012 and 1/30/2012 on 1/27/2012 and 1/31/2012 via FEDEX to their branches, hospitals and surgeon accounts respectively. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to contact their Stryker representative to arrange for return of any affected product. Customers were asked to fax back the attached Product Recall Acknowledgment Form as soon as possible to 201-831-6069..For any questions customers should call 201-972-2100.For questions regarding this recall 201-831-5970. | Contact the recalling firm for information | 532 | 532 units. | Nationwide. | Terminated | July 17, 2013 | 366 | 532 | Not_Computer
Z-2017-2012 | 62451 | gown, surgical | FYA | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MicroCool Breathable Impervious Surgical Gown, (Large and X-Large), Sterile, Made in Honduras.The Kimberly Clark Micro COOL Surgical Gowns are sterile single use surgical apparel intended to be worn by heathcare professionals to help protect the patient and the healthcare worker from the transfer of microorganisms, body fluids and particulate matter. | MicroCool Breathable Impervious Surgical Gown | 2 | July 16, 2012 | 2012 | Kimberly-Clark Corporation | Complaints were received of gowns pilling or balling, linting and abrades, blood strikethrough and fibers falling off the gowns. |  | Kimberly Clark sent an Urgent Medical Device Recall letter dated September 17, 2009. to all affected consignees. The Recall notification of the expanded recall was sent to consignees on November 17, 2009. Customers were instructed to destroy all affected product remaining in their inventory. For questions regarding this recall call 770-587-7400. |  | 79941 | 79,941 units | Nationwide Distribution including AL, FL, GA, KY, LA, MS, NC, OH, SC, TN, and VA | Terminated | July 19, 2012 | 3 | 79941 | Not_Computer
Z-1992-2012 | 62115 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Infusomat(R) Space Pump (US Version) with software version 686G030103 (commonly termed 'G03'), 686E030003, 686F030005, 686F030007, 686G030002, and 686G030102.Intended for use with adults, pediatrics and neonates and is intended to provide infusions of parenteral fluids/medications, blood and blood products indicated for infusion through FDA approved routes of administration. | Infusomat(R) Space Pump | 1 | July 16, 2012 | 2012 | B. Braun Medical, Inc. | Firm became aware of the potential for breakage of the anti free flow clip catch, located on the inside of the pump door, when the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed. If the clip catch is broken and the door opened, free flow protection is still ensured. However, after the set is manually removed from the pump without the closure of t | DESIGN: Device Design | B. Braun Medical, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify the affected product in their inventory and to complete and return the Product Acknowledgement Form via fax to (610)849-1197. Contact the Customer Support Department at (800) 626-7867 for assistance with returning any product or if you have questions regarding this notice. |  | 38225 | 38,225 pumps | Nationwide Distribution-including Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Open |  |  | 38225 | Not_Computer
Z-2022-2012 | 62337 | tube, tracheal (w/wo connector) | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | LMA Esophageal Intubation Detector (EID), Catalog Number EID100The device is used to assist verification of placement of the endotracheal tube or esophageal/tracheal double lumen tube. This device is to be used as an adjunct to assess intubation. Its purpose is not to eliminate clinical judgment | LMA Esophageal Intubation Detector (EID), Catalog Number EID100 | 2 | July 17, 2012 | 2012 | LMA North America Inc | The recall was initiated because LMANA has received a complaint that the Esophageal Intubation Detector (EID100) has a potential to create a false positive if used on a patient during confirmation of intubation. | PRODUCTION CONTROLS: Process Control | LMA North America Inc. sent a recall letter dated May 23, 2012, for US customers, and May 24, 2012, for International customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to fill out the bottom portion of the recall letter and promptly fax it back to LMA North America at (858) 622-4130. If customers have units of EID100 from the affected lots remaining in stock, customers were instructed to arrange for immediate return by contacting LMANA's customer service department at (800) 788-7999, menu selection #2. For further questions please contact Customer Service Representative at (800) 788-7999. |  | 18799 | 18,799 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, United Kingdom, ZA, and Australia. | Terminated | December 26, 2012 | 162 | 18799 | Not_Computer
Z-2023-2012 | 62020 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic MACH4 Endovascular Device (Catalog #500-55106, 500-55112; 500-55118; 500-55124; 500-55130; 500-55140; 500-55150; 500-54106; 500-56130; 500-56140; 500-56150) is a catheter, continuous infusion. it consists of the Intelligent Drug Delivery Catheter (IDDC) and the MicroSonic Device (MSD). It is intended for the controlled and selective infusion of physician-specific fluids, including thrombolytics, into the peripheral vasculature. All therapeutic agents utilized with the EkoSonic Endovascular System should be fully prepared and used according to the instruction for use of the specific therapeutic agent.The Kit MACH4 is labeled in parts: "***EkoSonic MACH4 Endovascular Device***MicroSonic Device and Intelligent Drug Delivery Catheter***".The Pouch MSD is labeled in parts: "***EKOS***MicroSonic Device***Sterile***".EkoSonic Endovascular System is a catheter, continuous infusion. It is intended for the controlled and selective infusion of physician-specific fluids, including thrombolytics, into the peripheral vasculature. The system is also intended for the infusion of solutions into pulmonary arteries. | EkoSonic MACH4 Endovascular Device | 2 | July 17, 2012 | 2012 | EKOS Corporation | EKOS Corporation have recently received several complaints from users where the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and report of MicroSonic Device (MSD) breaking during insertion or operation. | DESIGN: Device Design | Starting May 25, 2012 to May 31, 2012, EKOS Corporation sent "New warnings in EkoSonic Endovascular Device Instruction for Use" letter, dated May 24, 2012, to all affected customers.Attached with the notification letter is the revised EKOS MACH4 Endovascular Device Instructions for Use in Peripheral Vasculature. Customers are informed of the complaints about the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and reports of the MicroSonic Device (MSD) breaking during insertion or operation.Customers with questions about the revised Instruction for Use, please contact EKOS Corporation at 425-415-3132. |  | 21892 | 21,892 units | Worldwide Distribution - USA including OH, KY, and TN and the countries of Switzerland, Germany, Austria, Czech Republic, Slovakia Republic, Poland, Netherlands, Belgium, Luxembourg, United Kingdom, Republic of Ireland, Slovenia, Costa Rica, Argentina, Brazil, Dominican Republic, Taiwan, Korea, Greece, Mexico, Italy, Uruguay, Paraquay, Yugoslavia, New Zealand, India, Venezuela, South America, Norway, Sweden, Denmark, Turkey, Great Britain, Israel, Spain, Portugal, Greece, Korea, China, and Japan. | Terminated | November 20, 2012 | 126 | 21892 | Not_Computer
Z-2028-2012 | 62464 | coagulator-cutter, endoscopic, unipolar (and accessories) | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(TM) Jarit(R) Take-Apart Endoscopic Instruments.Integra(TM) Jarit(R) Take-Apart Endoscopic Instruments are intended for use in laparoscopic gynecological surgery and other operative procedures under endoscopic observations. For use when a rigid endoscopic instrument for grasping, dissecting and/or other manipulation of soft tissue is determined to be appropriate by the surgeon. | Integra(TM) Jarit(R) TakeApart Endoscopic Instruments | 2 | July 18, 2012 | 2012 | Integra Limited | The Integra(TM) Jarit(R) Take-Apart Laparoscopic Instrument shafts were released with internal component out of specification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra sent a Urgent Field Corrective Action letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers and sales reps to return the instruments identified on the attached part number list and complete a "Field Corrective Action Acknowledgement and Return Form" to be completed and returned to the firm.We apologize for any inconvenience this issue may cause you. Please call Integra Surgical Customer Service at 800-431-1123 if you have any questions. |  | 108 | 108 | USA (nationwide) including the states of AZ, CA, IL, MA, MN, NJ, OK, TN, TX, VA, WA and WI. | Terminated | March 19, 2013 | 244 | 108 | Not_Computer
Z-2032-2012 | 61769 | stains, microbiologic | JTS | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Gram Crystal Violet 250 mL, catalog #212525, in cartons of 4 bottles. Gram Stain Kits and Reagents are used to stain microorganisms from cultures or specimens by the differential Gram method. | Gram Crystal Violet 250 mL | 2 | July 18, 2012 | 2012 | Becton Dickinson & Co. | Potential for increase level of artifacts to appear as large, spherical and easy to distinguish or small & more difficult to distinguish from gram positive cocci. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard the affected product. An Acknowledgement Form was attached for customers to complete and return via fax to 410-316-4258. Contact the BD Customer Service Department for questions regarding this recall. |  | 878 | 878 cartons/ 3512 bottles | Nationwide Distribution. | Terminated | December 10, 2012 | 145 | 1417 | Not_Computer
Z-2025-2012 | 62404 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9900 Elite, OEC 9900 Elite MD Motorized C-arm System, OEC 9900 Elite NAV used as an aid to locate anatomical structures during open or percutaneous surgical procedures.The system is an aid to locate anatomical structures during open or percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgery technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. | OEC 9900 Elite | 2 | July 18, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC recalled imaging devices OEC 9900, OEC 9800, and OEC 8800 as a result of an FDA inspection identifying that the vertical lift column power supply in the mainframe C-arm is defective and subject to early life failure. | DESIGN: Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent an"URGENT RECALL NOTICE" dated February 22, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 800-874-7378 for questions concerning this recall. | Contact the recalling firm for information | 726 | 726 units | Worldwide Distribution. | Terminated | January 28, 2013 | 194 | 13175 | Not_Computer
Z-2029-2012 | 62408 | light, surgical, fiberoptic | FST | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink. GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink is used to transmit light. | GMFRC65 Replacement Surgical Headlight Cable for Cogent Microlink. | 2 | July 18, 2012 | 2012 | Gulf Fiberoptics, Inc. | The GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink falls outside of the firm's 510k due to light transmitting fiber material difference. | DESIGN: Device Design | The firm, Gulf Fiberoptics, Inc., sent an "RECALL NOTIFICATION-PLEASE DO NOT DISCARD" letter dated April 09, 2012 to its customer via (traceable letter & email). A notification was put on the firm's website. The letter described the product, problem and actions to be taken. The distributors were instructed to determine the status of the product; remove existing product from service immediately and return to their company in a timely manner; send all returned GMF-RC65 cables to Gulf Fiberoptics; forward copies of the recall letters to their customers, ultimately the end user, and complete and return the attached Recall Worksheet Form for each GMF-RC65 cables regardless of its status via fax to: +1 813.855.6627; email to: difigueoa@gulffiberoptics.com or mail to : Gulf Fiberoptics, 448 Commerce Blvd, Oldsmar, FL 34677. The customers were instructed to remove the cables from service immediately, if they think they may have one of the listed cables; contact customer service at +1 813.855.6618 for serial number verification and for instructions to return this cable to Gulf Fiberoptics.If you have any questions regarding this voluntary recall, please contact your Customer Service representative at: +1813.855.6618 or by email at: difigueroa@gulffiberotpics.com. |  | 169 | 169 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, FL, MI, NC, NJ, NY, OH, PA, TN, TX and WI and country of: South Africa. | Terminated | March 07, 2013 | 232 | 169 | Not_Computer
Z-2034-2012 | 61147 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc Blood Analysis BGEM Test Cards for use with the epoc(R) Blood Analysis SystemProduct Usage: Usage:epoc Blood Analysis BGEM Test Cards are used with the epoc(R) Blood Analysis System which is a POC testing system | epoc Blood Analysis BGEM Test Cards | 2 | July 19, 2012 | 2012 | Epocal Inc | Customer reported problem while performing incoming QC of newly received lot of test cards. Liquid quality control results were found to be out of range. Further investigation by Epocal determined that a portion of the test card lot (approximately 3000 cards) could potentially report low glucose results. | DESIGN: Device Design | Epocal Inc. sent a Voluntary Product Recall letter dated June 27, 2011 to the affected consignee. The letter identified the affected products, problem and actions to be taken. The letter ask that all products affected by this recall be sent back to Epocal and a replacement may be obtained. For questions call 1-613-738-6912. |  | 8650 | 8,650 | US nationwide Distribution in the state of CA. | Terminated | December 18, 2013 | 517 | 8650 | Not_Computer
Z-2039-2012 | 62293 | enzyme immunoassay, methadone | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Alere Triage Tox+MTD Drug Screen, PN 94400Product Usage:The Alere Triage TOX Drug Screen is a fluorescence immunoassay to be used with the Alere Triage Meters for the qualitative determination of the presence of drug and/or the major metabolites above the threshold concentrations of up to 10 distinct drug classes, including assays for acetaminophen/paracetamol, amphetamines, methamphetamines, barbiturates, benzodiazepines, cocaine, methadone, opiates, phencyclidine, THC and tricyclic antidepressants in urine. The acetaminophen/paracetamol assay will yield positive results when acetaminophen/paracetamol is ingested at or above therapeutic doses. This test provides only a preliminary result. Clinical consideration and professional judgment must be applied to any drug of abuse test result, particularly in evaluating a preliminary positive result. In order to obtain a confirmed analytical result, a more specific alternate chemical method is needed. Gas Chromatography/Mass Spectrometry (GC/MS), Liquid Chromatography/ Mass Spectrometry/ Mass Spectrometry (LC/MS/MS) and High Performance Liquid Chromatography (HPLC) are common confirmatory methods. A quantitative serum acetaminophen/paracetamol measurement is the common confirmatory method for preliminary positive acetaminophen/paracetamol results. | Alere Triage ToxMTD Drug Screen | 2 | July 19, 2012 | 2012 | Alere San Diego, Inc. | A recall was initiated because Alere San Diego has confirmed that the Triage TOX Drug Screen may have significantly decreased precision and accuracy than claimed. An increased frequency of false positive and/or false negative results has been observed for Triage TOX Drug Screen. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Alere sent an Urgent Medical Device Recall letter dated June 11, 2012 to all affected customers. The letter informed the customers of the affected product, problem identified and actions to be taken. Customers were instructed to discontinue all use of the recalled lots, immediately inform clinicians of this recall at the testing site(s) and complete and fax the enclosed customer verification form and list of affected lots within 10 days to confirm receipt of the letter. Customers with questions about the information contained in the recall letter, can contact Alere San Diego at 9975 Summers Ridge Road, San Diego, CA 92121 or call (877) 308-8287. |  | 9997 | 9,997 kits | US Nationwide Distribution | Terminated | December 04, 2012 | 138 | 9997 | Not_Computer
Z-2035-2012 | 62498 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Brilliance 6, 16, 40, 64, Big Bore, iCT and iCT SP are Computed Tomography X-Ray Systems.The devices are Computed Tomography Systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes | Brilliance CT 6slice,16slice, 40channel, 64channel,Big Bore Oncology & Brilliance iCT and iCT SP | 2 | July 19, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare was notified that several of the sixteen bolts securing the brackets which support the carbon table top on various models of their Brilliant CT scanner have been found to be defective and the bolt heads may break off. | PRODUCTION CONTROLS: Process Control | Philips sent a Urgent Medical Device Correction letter dated June 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will visit each site to inspect and correct the issue, and Philips will be tracking the progressof the correction. For further information, Customers were asked to contact their local Philips representative or Philips Healthcare Office. North American and Canadian consignees should contact the Philips Customer Care Solutions Center at 1-800-722-9377. |  | 524 | 524 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, MJ, NM, NV, NY, OH, OK, OR, PA, PR, TN, TX, VA, VT, and WI. and the countries of Canada, Argentina, Australia, Austria, Azerbaijan, Brazil, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Iceland, India, Indonesia, Iraq, Italy, Jordan, Kazakhstan, Kazakhstan, Libyan Arab Jamahiriya, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philipines, Poland, Portugal, Russia, Russia Fed, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Syrian Arab Republic, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | December 24, 2013 | 523 | 524 | Not_Computer
Z-2033-2012 | 61723 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker AVS Navigator Trials 6mm-15mmManufactured by:Stryker Spine SASZ.1 Marticot-33610-Cestas -France+33 (0) 5.57.97.06.30http;//www.stryker.comDistributed in the USA by Stryker Spine2 Pearl Court. AllendaleNJ 07401-1677 USA+1-201-780-8000Cestas France 33610.Intended Use: The trials are designed and intended to measure the height of the disc space intra-operatively. | Stryker AVS Navigator | 2 | July 19, 2012 | 2012 | Stryker Spine | Beginning in February 2011, Stryker Spine began receiving reports related to the shaft of the AVS Navigator trial failing at the junction of the trial head. | DESIGN: Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was included for customers to complete and return via fax to Regulatory Compliance at 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this recall. | Contact the recalling firm for information | 273 | 273 units | Nationwide. | Terminated | July 29, 2013 | 375 | 273 | Not_Computer
Z-2057-2012 | 62351 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediLoc Locking Plate System, 4.5mm Contour Locking Compression Femur 8 Hole Plate Right.Used for pediatric patients as indicated for pelvic, small and long bone fractures. | OrthoPediatrics Pediloc Locking Plate System | 2 | July 20, 2012 | 2012 | OrthoPediatrics Corp | The device is a Left 8 hole plate but is incorrectly labeled and etched as a Right device. | PRODUCTION CONTROLS: Error in Labeling | On 15-June-2012, a follow-up, URGENT Recall of MIS-BRANDED PediLoc Locking Plate e-mail was sent to all Distributors and direct representatives to provide them with a written notification of the voluntary recall. The e-mail identified the affected product and the reason for the recall. It also instructed them to quarantine all identified devices from the affected lot number immediately. A certified Voluntary Recall notice was also sent as a follow up along with a return prepaid pouch for the return of the recalled device(s) and Mandatory Reply Form. | Contact the recalling firm for information | 15 | 15 devices | Worldwide Distribution -- USA, including states of IN, GA, MI, FL, MO, OH, and SC and country of Australia. | Terminated | April 30, 2013 | 284 | 15 | Not_Computer
Z-2054-2012 | 62324 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Universal Spinal System Lamina HooksSynthes Medium TI Lamina Hook - Right 498.320Non sterileLOT 6876785Noncervical spinal fixation devices intended for posterior pedicle screw fixation, posterior hook fixation, or anterolateral fixation. Limited to skeletally mature patients with the exception of the small stature USS, which includes small stature and pediatric patients. These devices are indicated as an adjunct to fusion for all of these indications: degenerative disc disease, spondylolisthesis, trauma, deformities or curvatures, tumor, stenosis and failed previous fusion. | Universal Spinal System Lamina Hooks | 2 | July 20, 2012 | 2012 | Synthes USA HQ, Inc. | USS Lamina Hooks, Medium, Right are being recalled because the hooks are etched with the wrong part number - 498.321 (Lamina Hook, Medium, Left). The Lamina Hooks should be etched with part number 498.320 (Lamina Hook, Medium, right). | TRAINING: Employee Error | SYNTHES sent a Medical Device Recall letter dated May 31, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees and sales reps remove recalled product from their inventory, call the firm at 1-800-479-6329 to receive an authorization number and complete the enclosed verification form and return to:Credit/Returns, Synthes1101 Synthes AvenueMonument, Colorado 80132.If customers do not have any of the affected product they were instructed to complete the Verification Section at the end of the letter by checking the appropirate box indicating that no affected porudct has been located. Return the document to Synthes by fax to 610-251-9005 or Scan/Email: FieldAction@synthes.com.For any questions please call 610-719-5450. |  | 12 | 12 units | Nationwide Distribution including GA, PA, SC, TX, and UT. | Open |  |  | 12 | Not_Computer
Z-2058-2012 | 62520 | lubricant, patient | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(k) | Oto Ease by Westone, labeled as a bacteria free, greaseless lubricant. Sold in 0.5 oz consumer sized semi transparent flexible plastic bottles.Product Usage:The product is used as a patient lubricant which eases insertion of earmolds and hearing instruments such as hearing aids. | OTO EASE | 2 | July 20, 2012 | 2012 | Westone Laboratories, Inc. | Westone is recalling all Oto Ease ear lubricant in all packaging configurations and sizes due to potential microbial contamination of the product. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Westone Laboratories sent an Urgent Recall letter and Press Release to distributors and customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine product subject to recall and discard or return the product to Westone. In addition, if the affected product have been further distributed please identify your customers and notify them at once of this product recall. Consumers with questions may contact Westone Laboratories at 1-800-357-3240 between the hours of 8:00am and 5:00pm MST or email at otoeaserecall@westone.com. |  | 115 | Approx 100k sold annually x 15 years. | Worlwide Distribution - US (nationwide) and to countries of: Australia, Belgium, Brazil, Canada, France, Germany, Greece, Hong Kong, Ireland, Japan, Korea, Mexico, Panama, Philippines, Russian Federation, and Singapore, Sweden, Switzerland and Thailand and United Kingdom. | Terminated | December 14, 2012 | 147 | 115 | Not_Computer
Z-2024-2012 | 62393 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Trilogy 100, 200, and 202 VentilatorsThe Respironics Trilogy systems provides continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. Trilogy 100 is intended for pediatric through adult patients weighing at least 5 kg (11 lbs.). The devices are intended to be used in home, institution/hospital, and portable applications such as wheelchairs and gurneys, and may be used for both invasive and non-invasive ventilation. They are not intended to be used as a transport ventilator. | Respironics Trilogy 100, 200, and 202 Ventilators | 1 | July 23, 2012 | 2012 | Respironics, Inc. | The specified Trilogy units are being recalled because of the potential for a faulty component in the power supply. | DESIGN: Component Design/Selection | Respironics, Inc. began contacting its affected US customers by phone beginning on Friday, April 27, 2012, and requested that they quarantine any affected devices within their possession and that they retrieve any devices with patients and quarantine those as well. Replacement devices were to be provided to all affected customers. For questions customers should call 724-387-7651.For questions regarding this recall call 877-387-3311.. |  | 173 | 173 | Worldwide Distribution - USA including AZ, CA, IL, IN, LA, MI, MN, MO, NJ, NY, NC, PA, SC, TN, TX, and WI and the countries of Canada, Denmark, Estonia, Finland, France, Great Britian, India, Italy, Japan, Korea, Saudi Arabia, and Spain. | Terminated | November 01, 2012 | 101 | 173 | Not_Computer
Z-2059-2012 | 62307 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 4.5mm TI Multiloc Screw Length 38MM-SterileExpiration: 04/29/2012REF 04.019.038SManufactured or Distributed by Synthes (USA) 1101 Synthes Avenue, Monument CO 801324.5MM TI Multiloc Screw Length 38MM-SterileRef 04.019,038S. The Multiloc Screw 4.5 is used with a Multiloc PHN for treatment of proximal humerus fractures and is intended to be inserted through the Multiloc PHN until it engages the subchondral bone below the relative distal cortex (cortex closest to the humeral head, below the articular surface). | 4.5mm TI Multiloc Screw Length 38MM | 2 | July 23, 2012 | 2012 | Synthes USA HQ, Inc. | Two lots of 4.5mm multi-loc screws were recalled due to incorrect labeling of screw size. | PRODUCTION CONTROLS: Error in Labeling | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated May 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this recall. |  | 20 | 20 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, LA, MD, MN, NJ, NY, OH, PA, TN, TX, and VA. | Open |  |  | 20 | Not_Computer
Z-2070-2012 | 61969 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part: "***OPTETRAK***TIBIAL TRAY***TRAPEZOID CEMENTED***Ti Alloy*** SIZE 1F/1T, 2F/1T***FOR CEMENTED USE ONLY***Exactech, Gainesville, FL 32653-1630.***REF 204-04-21***".The Optetrak Comprehensive Knee Systems are indicated for use in skeletally mature individuals undergoing primary surgery for total knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis and/or post-traumatic degenerative problems. They are also indicated for revision of failed previous reconstructions where sufficient bone stock and soft tissue integrity are present. In the USA, the Optetrak Comprehensive Knee System is indicated for cemented use only. | OPTETRAKTIBIAL TRAY | 2 | July 24, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. Gainesville, FL is recalling their Optetrak Cemented Trapezoid Tibial Tray, Size 1F/1T, 2F/1T and the Optetrak Rotating Bearing Knee Tibial Tray, Finned, Cemented Tibial Tray, Size 2F/1T. Both trays have the potential to be mislabeled. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Exactech, Inc. sent an "IMPORTANT MARKET WITHDRAWAL NOTICE" on September 23, 2011, to all affected customers. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that has possession of the products. 3. Verify if they have any of the affected products listed. 4. Fax back the attached form. Customers were instructed to contact Exactech inventory representative to confirm quantities at their location. For questions regarding any inventory restocking, the customers were directed to contact Kaya Davis 1-800-392-2832. |  | 16 | 16 | Worldwide Distribution-USA-including the states of CA, MA, OH, and OK, and the countries of Brazil, Colombia, France, Italy, Japan, Luxembourg, and Spain. | Open |  |  | 28 | Not_Computer
Z-2074-2012 | 62450 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BridgePoint Medical, Stingray Catheter, REF M - 1000, Sterilized with Ethylene Oxide Gas, Prescription Only, Manufacturer BridgePoint Medical Inc. 13355 10th Ave N, Suite 110, Plymouth, MN 55441.Stingray" Catheters are intended to facilitate the intraluminal placement of conventional guidewires beyond stenotic coronary lesions (including chronic total occlusions) prior to PTCA or stent intervention. When used as part of the BridgePoint Medical System, the Stingray Catheter is intended to facilitate the intraluminal placement of conventional guidewires beyond stenotic coronary lesions (including chronic total occlusions [CTOs]) prior to PTCA or stent intervention. . | Stingray Catheter | 2 | July 24, 2012 | 2012 | Bridgepoint Medical | BridgePoint Medical Inc. is voluntarily recalling Stingray Catheter, Model M1000 lots BP20121440078 and BP20121560089. BridgePoint Medical, Inc. recently determined the potential of compromised sterility assurance of the product due to a faulty pouch seal. In no case was there any reported adverse clinical event related to the problem. | PRODUCTION CONTROLS: Packaging Process Control | The firm, BridgePoint Medical, sent an "Urgent Field Safety Notice" letter dated June 26, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory, quarantine the affected lots and complete and return the enclosed "Verification Form" via fax to 763-225-8718 ,attn: BridgePoint Medical, Inc., Upon receipt of form, BridgePoint will supply the customers with the appropriate materials to return the products.If you have any questions regarding this notice, please call BridgePoint Medical, Vice President of Quality Assurance, at 763-225-8511 or e-mail at cschlawin@bridgepointmedical.com. | Contact the recalling firm for information | 95 | 95 (57 USA, 38 OUS) | Worldwide distribution: USA (nationwide) including: AL, AZ, CA, CO, GA, IL, KY, MD, MN, MO, NC, NY, OR, PA, WA, and WI; and countries of: Australia and Sweden. | Terminated | September 06, 2012 | 44 | 95 | Not_Computer
Z-2038-2012 | 51222 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800; OEC FluoroTrak 9800 Plus; OEC 9800 Plus; OEC 9800 MD Motorized C-arm System, 892.1650 Image-intensified fluoroscopic x-ray system.The OEC 9800 is designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical and interventional procedures. The system includes features specifically designed for use in diagnostic and interventional cardiac imaging procedures and is also intended for cholangiography, endoscopic, urologic orthopedic, neurologic, vascular, critical care and emergency room procedures. It may be used for other imaging applications at the physician's discretion. | GE OEC 9800 | 1 | July 24, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The vertical lift column power supply in the mainframe C-arm of the device is defective and subject to early life failure. | DESIGN: Software Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The firm, GE Healthcare, Surgery, sent an "URGENT RECALL NOTICE" dated November 8, 2006 via certified mail with return receipt request to consignees/customers with specific instructions. Customers were advised "to stop using the system WITH NAVIGATION until this replacement has been successfully completed. OEC@ FluoroTrak 9800 Plus system may be used as a standard C-Arm providing you review the other issues as listed below." GEHC OEC is actively working on solutions that will permanently resolve each of the noted issues.If you have any questions or concerns regarding these issues, please do not hesitate to contact the service team for further information at 800-874-7378 option 8. Information is available at this number 24 hours per day, 7 days a week. | Contact the recalling firm for information | 8214 | 8214 units | Worldwide distribution. | Terminated | February 19, 2013 | 210 | 8214 | Not_Computer
Z-2077-2012 | 62326 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | .***ANSPACH***Sterile. REF SHD-4820DS***QTY: 1, 4.8 mm x 2mm Twist Drill with Depth Stop. Use with SHORT-HD Attachment ***RX OnlyProduct Usage:Cutting and shaping bone including bones of the spine and cranium. | Anspach Single Use, Sterile Bone Cutting Burrs (Model SHD4820DS) | 2 | July 25, 2012 | 2012 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Single Use Sterile Bone Cutting Burrs used with XMax, microMax, microMax Plus, eMax, eMax2 and eMax2 Plus High Speed Drill Systems. It is possible for the cutter to come apart. | DESIGN: Device Design | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated December 15, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the batches listed in the letter.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply form indicating if the item(s) are being returned and confiming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. | No consumer action necessary | 5 | 5 each | US Nationwide Distribution including the states of: Arizona. and California. | Terminated | July 25, 2013 | 365 | 5 | Not_Computer
Z-2075-2012 | 62582 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 100 mL Burette Set, Convertible Pin, 77 inch with 2 Prepierced Injection Sites and Option-Lok Microdrip Soluset with Calibrated Burette and Precision Drip Chamber; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 12722-65. Intended use: for the administration of fluids. | Hospira LifeShield LatexFree 100 mL Burette Set | 2 | July 25, 2012 | 2012 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | TRAINING: Employee Error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated July 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-877-722-7019 or e-mail to: hospira4921@stericycle.com; return affected product to Stericycle using the label provided with the letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. | No consumer action necessary | 4420 | 4,420 sets | Worldwide Distribution-USA (nationwide) including the states of California, Colorado, Connecticut, Florida, Illinois, Kentucky, Maryland, Massachusetts, Minnesota, Oklahoma, Texas, Utah, Virginia, Washington, West Virginia and Wyoming, and the countries of Barbados and Canada. | Open |  |  | 4420 | Not_Computer
Z-2081-2012 | 54724 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | The Handi-Fil Disposable Syringe Fill Tube is a sterile plastic tube, approximately 10 inches in length, with a "J" shaped curve that projects upward approximately two inches. The Handi-Fil Disposable Syringe Fill Tube is used to transfer radiographic contrast media from a sterile vial to an injection syringe. | P/N 302050 HANDI FIL DISPOSABLE SYRINGE FILL TUBE | 2 | July 26, 2012 | 2012 | Mallinckrodt Inc | A report was received from a Japanese distributor regarding a complaint from a hospital customer stating that ten packages of Handi Fil were not sealed. | PRODUCTION CONTROLS: Packaging Process Control | Covidien send an Urgent Device Recall letter dated February 1, 2010, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and discontinue use of the product immediately. Customers were asked to complete the attached business form and report the quantity of affected product currently in their stock. Once the form is completed customers should fax the form to 314-654-8206. Once the business notification form is received, an RGA number will be sent to the customer for return of the product. Customers will receive credit for all returned product. For questions regarding this recall call the Product Monitoring Department at 800-778-7898. For questions about credit for the product, please call Customer Service at 888-774-1414, menu option 1, then menu option 2.For questions regarding this recall call 513-948-5719. |  | 75 | 75 cases (100 units per case) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, IL, KS, KY, LA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, & WV and countries including: AUSTRALIA, BELGIUM, BRAZIL, CANADA, JAPAN & SINGAPORE. | Terminated | August 03, 2012 | 8 | 75 | Not_Computer
Z-2084-2012 | 62365 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Punch Tower Assembly. Non-Sterile. CAT NO. 8000-1089. Howmedica Osteonics Corp., 59 Route 17, Allendale, NJ 07401. During primary Scorpio TKA procedures, the Scorpio Tibial Punch Tower (8000-1089) is used in conjunction with Scorpio Tibial Keel Punches (3760-0305/-0709/-1113 and 3761-0305/-0709/-1113) and Tibial Templates (3750-0003/-0005/-0007/-0009/-0013) to prepare the tibia for baseplate implantation. Per the surgical protocol for Duracon and Scorpio Total Knee Systems with Xcelerate Instrumentation (LSPK31 03/03), tibial templates are used to determine the appropriate coverage of the tibial plateau. Once a template size is selected, the device is secured to the plateau with fixation pins, and the tibial punch tower is assembled by placing the tower onto the two small locating pins on the tibial template. A keel punch is then selected, assembled into the channel of the tower, and impacted with a mallet. Depending on the intended implant size, several keel punches of increasing size may be utilized to incrementally enlarge the cavity. Each keel punch is advanced until it fully seats on the tibial template. Tibial preparation is considered complete when the final keel punch fully seats on the tibial template. | Howmedica Osteonics Corp. | 2 | July 26, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report indicating that a Tibial Keel Punch could not be inserted into the slot of the Tibial Punch Tower. The slot that guides the stem of the keel punch into the channel of the tower was later determined to be undersized. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker Orthopaedics sent an Urgent Product Recall/Recall Acknowledgement letter dated January 10, 2012, to all affected customers. The letter identified the product(s) the problem and the action needed to be taken by the customer.Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For questions customers should call 201-972-2100.For questions regarding this recall call 201-831-5970. | Contact the recalling firm for information | 152 | 152 units | Worldwide Distribution | Terminated | June 13, 2013 | 322 | 152 | Not_Computer
Z-2083-2012 | 62481 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501.The device is labeled in part: "***Neusoft***NeuViz16 MULTI-SLICE CT SCANNER SYSTEM***MANUFACTURED FOR NEUSOFT MEDICAL SYSTEMS CO., LTD***MANUFACTURED BY PHILIPS AND NEUSOFT MEDICAL SYSTEMS CO., LTD***ADDRESS: NEUSOFT PARK, HUNNAN INDUSTRIAL AREA, SHENYANG***CHINA***". | NeuViz 16 MultiSlice CT Scanner System | 2 | July 26, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | One of the four screws that secure the Oil-accumulator to the X-Ray Tube pulled out in the NeuViz 16 CT scanner, which caused the Oil-accumulator to detach. | DESIGN: Device Design | Philips Healthcare North America Company sent an "Urgent Device Correction NeuViz 16 X-Ray Tube Oil-accumulator Detached" letter and attachments of Annex A Method to Identify Affected Products and Annex B Instruction for Mechanic, dated June 14, 2012, to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to follow the instructions on Annex A and Annex B.Philips and Neusoft Medical Systems planned to release the mandatory Field Change Order (FCO) and FCO kits to all customers by the end of June 2012.If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems, Co., LTD by e-mail to NMS-SERVICE@NEUSOFT.COM. | Contact the recalling firm for information | 15 | 15 units distributed in the US | Nationwide Distribution -- North Carolina, Ohio, and Texas. | Terminated | September 17, 2012 | 53 | 15 | Not_Computer
Z-2082-2012 | 62453 | table, obstetric (and accessories) | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Affinity Four Birthing Bed P3700 Product Usage:The Affinity, Four Birthing Bed is intended to be used as birthing beds for women of child bearing age in an LDR (Labor, Delivery, Recovery ) or LDRP (Labor, Delivery, Recovery, Postpartum) setting within the acute care Labor and Delivery market. They are not intended for use as general hospital beds. | Affinity Four Birthing Bed P3700 | 2 | July 26, 2012 | 2012 | Hill-Rom, Inc. | During review of our last braking system via that field corrective action, the system still did not meet our expectations, which is to eliminate the hazards of patient falling. Bed movement immediately prior to, or during the birthing process may still create distractions and delays, resulting in patient or caregiver falls, or contribute to a deviation in aprocedure, (e.g. IV insertion, epidural | DESIGN: Device Design | Hill-Rom sent an URGENT FIELD SAFETY NOTICE letter dated April 2, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were informed that Hill-Rom will visit their facility to modify the affected products, however, until then follow the actions to be taken outlined. For questions concerning this request or the procedure as outlined contact Hill-Rom Technical Support at 800-445-3720 (U.S.). |  | 12726 | 12,726 devices | Worldwide Distribution - US Nationwide and countries including: Argentina , Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Guyana, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, Turks and Caicos Islands, United Arab Emirates, United Kingdom, United States, US Virgin Islands, Venezuela and Vietnam | Open |  |  | 12726 | Not_Computer
Z-2087-2012 | 62117 | processor, cervical cytology slide, automated | MKQ |  |  | Pathology |  | PrepStain Slide Processor.The PrepStain System is a liquid-based thin layer cell preparation process. The PrepStain System produces SurePath Slides that are intended for use in the screening and detection of cervical cancer, pre-cancerous lesions, atypical cells and all other cytologic categories as defined by The Bethesda System for Reporting Cervical/Vaginal Cytology Diagnoses. The 1.2 DiTi is the tip on the PrepStain to which a pipette tip attaches and dispenses patient samples onto a glass microscope slide. | PrepStain System | 2 | July 26, 2012 | 2012 | Tripath Imaging, Inc. | Some Prep Stain Kit, 1.2 DiTi Cones have been found to be leaking and the PrepStain Preparation kits contain incorrect preventive maintenance (PM) instructions. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostics sent a letter dated March 2012 to all affected customers. The letter identifies the product, problem and actions that will be taken by the firm. Contact BD Diagnostics Technical Support at 1-877-822-7771, option 2 for questions regarding this letter. |  | 201 | 201 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WY, and WV and the countries of Belgium, Brazil, Canada, China, Columbia, Hong Kong, India, Japan, Korea, Peru, Singapore, Taiwan and Victoria (Australia). | Terminated | November 08, 2013 | 470 | 2659 | Not_Computer
Z-2090-2012 | 62456 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Rejuvenate Modular Stems Stryker Ireland Carrigtwohill Industrial EstateCarrigtwohill County Cork, IrelandRejuvenate Modular Necks Stryker Howmedica Osteonics Corp. Mahwah, NJ 07430.Stem Bodies Rejuvenate Modular consists of 6 modular stems bodies in sizes 7-12. The Modular Stem bodies are made of TMZF alloy, developed specifically for orthopaedic applications. Rejuvenate Modular Necks The Rejuvenate Modular Necks are made of Cobalt Chrome. The necks in this system consist of 4 lengths in 4 mm increments. They are 30 mm, 34 mm, 38 mm and 42 mm. They come in multiple neck angles and versions in each length. They are 127, 130 and 132. The versions are 0, 8, 16 degrees. In total there are 16 unique neck implants in the Rejuvenate Modular System. | Rejuvenate Modular Stems | 2 | July 26, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has updated the instruction for use (IFU) (product correction)for the ABGII Modular and Rejuvenate Modular Hip Systems. This is based on a reported rate of less than one percent for revisions potentially associated with fretting and or corrosion at or about the modular neck junction. On June 29, 2012, Stryker has decided to remove the affected product from the field since there may be an | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Stryker sent an "URGENT PRODUCT CORRECTION" letter dated April 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Product Correction Acknowledgment Form for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-831-5970 for questions regarding this notice.Stryker sent and "Urgent Update Product Recall" Letter dated January 3, 2013 to all affected customers. The new letter includes updated corrective actions for patients. |  | 23193 | (NLS/NLV) - 23193 units and (SPT) - 20,743 units | Nationwide Distribution | Open |  |  | 32196 | Not_Computer
Z-2079-2012 | 61974 | disinfector, medical devices | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge brand --- large capacity, floor loading jet spray, washers for mechanical washing, intermediate level, thermal disinfection and drying of moisture and temperature stable Hospital case carts, sterilization containers, material handling carts, stands and utensils. --- Model numbers Getinge 9120, 9122, 9125, and 9128 | Getinge | 2 | July 26, 2012 | 2012 | Getinge USA Inc | Getinge Disinfection AB is voluntarily correcting Washer Disinfectors, Getinge Electrically Heated 9100-Series manufactured between 2009-05-04 through 09-26-2011. This action is based on a report from the ground fault breaker manufacturer of a production error. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Getinge USA issued a Device Correction Notice dated June 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to perform a trip test to determine if they have a functioning ground fault breaker. Customers will be contacted by a Getinge representative to schedule an appointment to exchange the earth fault breaker. Customers were provided a Customer Response Form to complete and fax back to 585-272-5033. For questions customers should call 800-475-9040, ext. 5270.For questions regarding this recall call 585-475-1400. |  | 8 | 8 washers | Worldwide Distribution - USA including Maryland, Indiana, New Mexico, Massachuesetts, Texas, Kansas and the country of Canada | Open |  |  | 8 | Not_Computer
Z-2100-2012 | 33854 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI), P/N 14324 are sold in cases, packaged as part of STERRAD System Validation kits and STERRAD test packs.The STERRAD CYCLESURE 24 Biological Indicator (BI), P/N 14324 is intended to be used as a standard method for frequent monitoring of the STERRAD Sterilization System cycles. | STERRAD Cyclesure Biological Indicator (BI), | 2 | July 27, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling certain lots of the STERRAD CYCLESURE 24 Biological Indicator (BI) product because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | EXPIRATION DATING: Incorrect or No Expiration Date | ASP has posted the STERRAD CYCLESURE Biological Indicator customer letter on their external facing website at: wwww.aspjj.com/us/news/cyclesureletter on July 3, 2012. The letter informed customers who purchased STERRAD CYCLESURE 24 Biological Indicator (BI) product that Advanced Sterilization Products (ASP) is recalling certain lots because ASP has determined that it does not have adequate data to support the entire duration of the product's labeled shelf-life. The letter also provided the customers with what actions are needed for the recall. Customers were instructed to return all unused STERRAD CYCLESURE 24 BI product that exceeds the newly determined expiration dated using the enclosed prepaid UPS return label to Stericycle, 2670 Executive Drive Suite A, Indiapolis, IN 46241. Customers were instructed to count their inventory of product that exceeds the newly determined expiration date and record the data on the enclosed Business Reply Card and packing slip that are included with the customer recall letter. Customers were instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers were instructed to contact Stericycle at (877) 907-7030 if they need additional packing labels. If customers further distribute the affected products to other locations, then they were instructed to contact Stericycle to arrange for return of the product. Customers were instructed that new shipments of STERRAD CYCLESURE 24 BI product will be shipped on July 9, 2012 that has accurate expiration dates that do not require conversion. Customers wil be able to differentiate between current product and the newly shipped product by the lot number. Current product has a lot number containing 6 characters, eg. 123456.Newly shipped product will have a lot number containing 8 characters, eg. 123456EE.Customers with any additional questions were instructed to contact (877) 907-7030 or contact their A. |  | 38693 | 38,693 units | Worldwide Distribution - USA (nationwide) | Terminated | August 13, 2013 | 382 | 38693 | Not_Computer
Z-2092-2012 | 61764 | prosthesis, knee, patello/femorotibial, semi-constrained, uncemented, porous, coated, polymer/metal/polymer | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Nextgen Complete Knee Solution Replacement Jaw for Use with Femoral and Provisional Impactor/Extractor, REF 00-5901-026-40, Non-Sterile Zimmer, Warsaw IndianaThe instrument is attached to the femoral implant or provisional to aid in positioning and inserting the components. | NEXGEN Complete Knee Solution | 2 | July 27, 2012 | 2012 | Zimmer Inc. | Postmarket reports indicate that the spring clip may fracture or fall off of the instrument during use, posing the risk that a device fragment may be left in the surgical site. The risk of disassembly or fracture appears to be happening during use of this instrument where the instrument is typically impacted with a mallet to aid in insertion. The applied forces arecontributing to the instrument | DESIGN: Component Design/Selection | Zimmer sent an Urgent Medical Device Recall letter dated April 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Requests for alternative instruments should be requested by calling (574) 372-4470. Fax Back Inventory Return Certification Form is to be sent back to Zimmer at (574) 372-4265. All recalled instruments available at the distributor are to be returned to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. The Distributor notice included instructions to notify surgeons immediately of the recall and requested a list of hospitals and surgeons receiving the device be sent back to Zimmer via E-mail. Zimmer is continuing distribution of the replacement jaw ( 005901-026-40) until alternative tools are available. The distributors are asked to identify and deliver preferred alternative instruments to affected accounts. Beginning May 7, 2012 Urgent Recall notices are being sent to hospitals and surgeons identified with instructions to acknowledge the recall communication.For questions regarding this recall call 574-372-4807. |  | 65 | 65 units | Worldwide Distribution - USA (nationwide) and the countries of Singapore, China, Hong Kong,Canada, Australia, Japan, Europe, Germany, India, and Finland. | Open |  |  | 1727 | Not_Computer
Z-2101-2012 | 62553 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | INSTRUMENT ARM DRAPE IS2000, 20 PACK:Product Usage:The intended use of the Instrument arm drape is to protect equipment from contamination and to maintain a sterile field during surgical procedures using the da Vinci Surgical Systems. | INSTRUMENT ARM DRAPE | 2 | July 27, 2012 | 2012 | Intuitive Surgical, Inc. | Specific lots of the Instrument Arm Drapes were manufactured with a sterile adaptor that may have difficulty engaging an instrument. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letters dated July 16, 2012 via Fed Ex to all affected customers in the US. The customers outside of the US were notified starting July 17, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to Identify affected lots, order replacement products and return both full and partial boxes of the affected products for credit as soon as unaffected products are available. All staff should be made aware of the issue. Customers were instructed to complete and return the acknowledgement form immediately.Questions and concerns should be directed to Intuitive Customer Service at 1-800-876-1310.EXPANSION:Intuitive discovered an additional lot (Lot DA121075) of Instrument Arm Drape IS2000, 20 pack that should have been included in the recall and expanded the recall on September 4, 2012. Letters to customers were sent by Fed Ex on September 5, 2012. |  | 92390 | 92390 total instrument Drapes | Worldwide Distribution - US Nationwide and the countries of: Australia, Spain, India, Italy and Saudi Arabia. | Terminated | June 12, 2013 | 320 | 92390 | Not_Computer
Z-2108-2012 | 62590 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Maquet Getinge Group Operating Table System Magnus Carbon-Fibre Table Top Version 1180.16The carbon-fibre MAGNUS table top (1180-16XX) serves to support and position patients for surgical procedures (e.g. cardiovascular surgery) and interventional examination immediately before, during and after the surgical phase. The table top is radiotranslucent and enables intraoperative use of X-ray equipment. | Magnus CarbonFibre Table Top | 2 | July 30, 2012 | 2012 | Maquet Medical Systems, Usa | As a result of a small number of customer complaints, MAQUET conducted an investigation and identified a potential problem with specific MAGNUS Carbon-Fibre Table Tops manufactured between February 23, 2010 and February 28, 2012 and distributed between November 11, 2010 and April 6, 2012. There is a potential issue with the motor brakes for the longitudinal shift drive, and the tilt drive. | DESIGN: Device Design | MAQUET GETINGE Group sent An "Urgent Device Field Correction" notification, dated 13 April 2012 via Federal Express to all affected customers. An acknowledgement form is included with the customer letter which can be mailed or faxed back to Maquet Quality by the notified customers. If the completed Acknowledgement form is not received within 15 days, up to three follow ups will be made and documented via telephone call to each customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and notify their MAQUET Service Representative immediately. The MAQUET Service Representative will contact customers to make arrangements to replace both te logitudinal shift and tilt drives of their table top(s). For questions customers should call 888-627-8383, press option 2, followed by option 1.For questions regarding this recall call 201-995-8968. | Contact the recalling firm for information | 164 | 164 Units located in 27 Countries | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Germany, Denmark, Spain, France, Hong Kong, Ireland, Israel, India, Italy, Japan, South Korea, Netherlands, Poland, Russia, Saudi Arabia, Sweden, Singapore, and Thailand. | Terminated | September 04, 2013 | 401 | 164 | Not_Computer
Z-2113-2012 | 62493 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Padded Swing Away Arm Rest with Receiver;Uses on wheelchair models:Quickie Q7 (EIR4), GT (EIR3), GTi (EIT4) and Zippie Zone (EIRZ1).Product Usage:To empower physically challenged persons by providing a means of mobility. This includes temporary and permanent conditions. | Padded, swingaway armrests on Quickie Q7, GT, GTi and Zippie Zone wheelchairs | 2 | July 31, 2012 | 2012 | Sunrise Medical (US) LLC | Reports of injury when the Padded Swing Away Armrest is used for full body weight. Owner's manual states that armrest is only designed for resting of forearm. | DESIGN: Labeling Design | Sunrise Medical sent a Safety Notice letter dated June 25, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter request that clinicians/suppliers post this safety notice in a conspicuous location, making customers and therapists aware of the issue. The letter also encourage end users to read the owner's manual and follow the recommenations for proper usage of the recalled product. Customers are instructed to complete and return the Supplier Notification and Response Form. For questions contact Sunrise Medical Regulatory Affairs at 888-208-4901. |  | 49 | 14, 035 | Worldwide Distribution - US (nationwide) and the countries of Argentina, Brazil, Columbia, Costa Rica, Mexico, Panama, Uruguay and Venezuela. | Terminated | November 01, 2012 | 93 | 14035 | Not_Computer
Z-2115-2012 | 62566 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe Anatomic Replicator PlateIntended for Primary and Reverse Total Shoulder Arthroplasty | Equinoxe Anatomic Replicator Plate | 2 | July 31, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. initiated a recall of Equinoxe Replicator Plates, Equinoxe Locking Screws, Equinoxe Compression Screw Kits, Equinoxe Fracture Adapter Tray Screws, and Octane-C Intervertebral fusion devices due to the firm determining that the device labeling identifies the affected parts as having a ten year shelf life; however, the packaging configuration that was used it qualified to a five year | MISBRANDING: Labeling False and Misleading | Exatech sent an Important Voluntary Recall Notice dated June 18, 2012, to all affected customers. The Notice identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to immeidately case distribution or use of the affected product. Verify that they have the affected product. Contact their Exatech Inventory Representative for further instruction on how to return their products. The firm sent out recall notification to consignees. Customers were asked to complete and return the Voluntary Recall Fax Notice to 1-352-378-2617. Customers with questions should call 1-800-392-2832.For questions regarding this recall call 352-377-1140. |  |  |  | Worldwide Distribution - USA including FL, MO, MA, AL, KY, IN, NC, LA,PA, VA, OH, NY, ME, MI, TX, KS, TN, OK, NJ, WA, CO, GA, CA, MN, IA, OR, and IL and Internationaly to Austria, Luxenbourg, Sweden, Germany, Spain, France, Greece, Italy, Nederland, United Kingdom, Canada, India, Australia, Switzerland, and Sao Palo. | Terminated | July 03, 2013 | 337 |  | Not_Computer
Z-2114-2012 | 62626 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap SIBD 4 degree Trial Implant (SJ768T-US) and Aesculap SIBD 9 degree Trial Implant (SJ780T-US)Product Usage:The Aesculap Implant Systems SIBD Trial Implant instruments assist the surgeon in selecting the correct size implant for the patient | Aesculap SIBD 4 degree Trial Implant (SJ768TUS) and Aesculap SIBD 9 degree Trial Implant (SJ780TUS | 2 | July 31, 2012 | 2012 | Aesculap Implant Systems LLC | The 4 degree and 9 degree trial instruments (SJ768T-US/SJ780T-US) are correctly dimensioned, but the part number etched on the instrument may be incorrect. The 4 degree trial instrument (SJ768T-US) may be marked with the incorrect part number of the 9 degree trail instrument (SJ780T-US) and the 9 degree trial instrument (SJ780T-US) may be marked with the incorrect part number of the 4 degree tria | TRAINING: Employee Error | Aesculap sent an Important Product Removal Notification letter dated July 11, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product by contacting their local sales representative or the customer service department at 1-866-229-3002 for a Returned Goods Request )RGR) number. Customers were asked to complete and return the Distribution Inventory Sheet. For questions call 610-984-9260 or 610-984-9274. |  | 5 | 5 | US Nationwide Distribution - including the states of CO, OH, OK and TX. | Terminated | February 22, 2013 | 206 | 5 | Not_Computer
Z-2123-2012 | 61655 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | IntegraÂŽ Advansys Medial Lisfranc Plate (MLP), Integra, 311 Enterprise Drive, Plainsboro, NJ 08536The Advansys plates are intended for bone fixation such as: Arthrodesis of the 1st metatarsocuneiform joint to reposition and stabilize a metatarsus primus varus, Lisfranc arthrodesis, Mono or bi-cortical osteotomies or fractures near the 1st metatarsocuneiform joint. | Integra Advansys Medial Lisfranc Plate (MLP) | 2 | August 01, 2012 | 2012 | Integra LifeSciences Corp. | Incorrect technique applied may cause potential rupture of the tibialis anterior tendon during surgery. | DESIGN: Labeling Design | INTEGRA sent an Urgent Medical Device Correction Notice dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to identify members of surgical team that ues the affected product and make them aware of the Medical Device Correction Notice. Customers were asked to sign and return the enclosed Medical Device Correction Notice Acknowledgement and Return Form and to maintain a copy of the notification and signed copy of the acknowledgment form for their records. For any questions regarding this recall call 609-936-6832.. | For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. | 2437 | 2,437 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NM, NJ, NY, NV, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA & WI and internationally to Belgium; Danemark; France; Germany; IsraÂŽl; Italy; Lebanon; Luxemburg; Netherlands; Norway; Portugal; Saudi Arabia; Spain; Switzerland; Turkey & United Kingdom. | Terminated | February 19, 2013 | 202 | 2437 | Not_Computer
Z-2125-2012 | 62449 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Plus Digital; MDL number A349855.Product Usage: Is designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical and interventional procedures. Clinical applications may include cholangiography, endoscopic, urologic, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedures. The system may be used for other imaging applications at the physician's discretion. | OEC 9800 Plus | 2 | August 01, 2012 | 2012 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC had discovered an issue with the hard disk drives that were installed on OEC device models: UroView 2800; 9800 Plus Digital; MiniView 6800; and FlexiView 8800 Mobile C-Arm. The devices were manufactured between October 15, 2004 and | DESIGN: Software Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notices dated October 11, 2006 and February 21, 2007 via certified mail delivery letter to affected customers. The notices identified the affected products, problem and actions to be taken. Customers were informed that a GEHC OEC field service engineer would be contacting their facility to arrange to have a replacement hard disk drive installed on the affected system(s). Customers were instructed to ensure that images on the hard disk drive were stored via an alternate permanent media as the hard disk drive approaches the 170 image capacity until a GEHC OEC field service engineer had replaced the drive. For questions regardiug this issues contact the service team at 800-874-7378. | Contact the recalling firm for information | 82 | Part of a total of 82 units distributed. | Worldwide Distribution - US (nationwide) and the countries of: Canada, France, Japan, and China. | Terminated | August 02, 2012 | 1 | 82 | Not_Computer
Z-2124-2012 | 62555 | urinalysis controls (assayed and unassayed) | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | IRISpec" CA/CB two-part urine chemistry control twin set, Iris Diagnostics part number 800-0074.Intended Use: IRISpec CA/CB is an assayed QC material for monitoring of the urine chemistry analytes and devices as listed on the package insert. For in vitro diagnostic (IVD) use only. | IRISpec" CA/CB twopart urine chemistry control twin set | 3 | August 01, 2012 | 2012 | Iris Diagnostics | The firm recalled after it received complaints of IRISpec CA/CB two part Urine Chemistry Control, 101131-12, where the nitrite results read negative for the part B control, which should be positive. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm initiated the recall on 07/09/12 by sending out a Medical Device Recall notification to customers via e-mail. The recall notification identified the affected product, reason for recall, and instructions on the recall. Customers were to confirm how many sets of unopened or partially opened IRISpec CA/CB that was in their inventory, dispose of the affected product, and provide written confirmation of the product's disposition. The letter also stated that Iris Diagnostics would replace any unused inventory that was returned at no charge to the customer. Contact information for the recall was also provided. The communications also included a fax response sheet for return/replacement. |  | 862 | 862 | Nationwide Distribution | Terminated | January 27, 2014 | 544 | 862 | Not_Computer
Z-2121-2012 | 62512 | radioimmunoassay, thyroid-stimulating hormone | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer Genetic Screening Processor (GSP)Product Code: 2021-0010Product Usage: The GSP instrument is a fully automated, high throughput batch analyzer for time- resolved and prompt fluorescence analysis of samples in microtitration plates. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. The reagent kits running on the GSP instrument are used to test blood specimens dried on filter paper as an aid in screening newborns for congenital disorders. Currently assays for the determination of T4, TSH, 17OHP, IRT and GALT are cleared and available for use with the GSP instrument. | Perkin Elmer Genetic Screening Processor (GSP) | 2 | August 01, 2012 | 2012 | Perkinelmer | Contamination from GSP Shaker unit results in decreased signal activity of assay measurements for Genetic Screening for Neonatal IRT, TSH and IRT | DESIGN: Component Design/Selection | PerkinElmer sent an Urgent: Medical Device Recall letter dated June 27, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement units of the recalled product will be installed as soon as they are available. Customers were instructed to complete and return the response form attached. For further information contact your local PerkinElmer representative or GS. Technical.Support@PerkinElmer.com. | Contact the recalling firm for information | 78 | 78 units | Worldwide Distribution - US (nationwide) and the countries of Australia, Belgium, Brazil, Egypt, France, Germany, Netherlands, Norway, Saudi Arabia,Switzerland, | Open |  |  | 78 | Not_Computer
Z-2122-2012 | 62557 | disinfector, medical devices | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | The Reliance 130 and 130L Cart Washer/Disinfector is a high capacity mechanical washer intended for use in the washing, low-level disinfecting, and drying of carts, utensils, beds and other miscellaneous reusable items used in the care of patients at healthcare facilities / Crated for shipping. | RELIANCE 130 and 130L CART WASHER | 2 | August 01, 2012 | 2012 | Steris Corporation | To ensure that users are properly operating the Reliance 130 and 130L Cart Washers, a door close confirmation button and door open delay will be installed on your cart washer. | DESIGN: Labeling Design | Steris Corp. sent a Urgent Field Correction Notice dated July 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To ensure users are properly operating the Reliance 130 and 130L Cart Washers a 'door close' confirmation button and a door open delay will be installed on the affected cart washer(s). A STERIS Service Representative will contact the affected facilities to arrange for STERIS to upgrade their Reliance 130 and 130L Cart Washer(s). For further information or if affected customers have any questions regarding the STERIS visit to their facility, they can contact STERIS Field Service Dispatch at 1800-333-8828. |  | 367 | 367 | Worldwide Distribution -- USA (nationwide) including the states of : AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY. and the countries of : Australia, South Korea, France, Italy, Lebanon, Russia, Saudi Arabia, Spain, UAE, United Kingdom and Canada. | Open |  |  | 367 | Not_Computer
Z-2129-2012 | 62538 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Cardiac Catheterization Waste Management System w/1000mL Waste Bag, 2 Fluid Admin SetsProduct Usage:For waste managment. | Cardiac Catheterization Waste Management System | 2 | August 02, 2012 | 2012 | ICU Medical, Inc. | The firm initiated the recall due to a molding defect in a high pressure stopcock where a smallpercentage of the side ports could become totally occluded. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ICU Medical sent an Urgent Medical Device Recall letter dated April 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to inspect inventory, destroy affected products according to hospital protocol, complete and return verification form and contact ICU Customer Service to order replacement products. The letter states that ICU Medical will credit customer for any product destroyed once the completed verification form is received. This recall should be carried out to the user level and passed on to all those who need to beaware within your organization or any other organization the device may have been transferred to.Should you have any questions or require assistance relating to this recall please contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com. | Contact the recalling firm for information |  |  | Nationwide distribution | Terminated | August 23, 2012 | 21 |  | Not_Computer
Z-2135-2012 | 62504 | probe, radiofrequency lesion | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | RF Cannula, 100 mm,Curved 10 MM, 20 G, REF 0406-630-125 RX only, Disposable ( 10/pk)Product Usage:The Stryker RF Cannulae, in combination with the Stryker RF Generator/Multi- Gen and electrodes, are intended for coagulation of soft tissue in orthopedic, arthroscopic, spinal, and neurosurgical applications. | Stryker Radiofrequency (RF) cannula | 2 | August 02, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received 6 complaints that the 5mm needle was placed in 10 mm package and 10 mm needle was placed in the 5 mm package. If the clinician does not recognize that the active tip on the cannula is 10 mm instead of 5mm and the cannula is used in a cervical procedure there is a potential for a thermal skin lesion due to a larger lesion creation. If the 10 mm active tip cannula, as opposed | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker sent an Urgent MEDICAL DEVICE Recall Notifications dated July 16, 2012 to all consignees. The letter identified the affected products, problem and actions to be taken. Customers/distributors were instructed to check their inventory and quarantine any affected product found, complete and fax the business reply form to Stryker Instruments Regulatory Department, 866-521-2762 and return all affected product using the pre-paid shipper provided. Upon receipt of the recalled product, a replacement will be issued.For questions regarding this recall please contact Stryker Instruments (Recall Coordinator) Monday-Friday 8am-5pm (EST) at 269-389-4354. | Contact the recalling firm for information | 226 | 226 | Worldwide Distribution - US Nationwide and the countries of England, France, Netherlands, and Switzerland. | Terminated | May 02, 2013 | 273 | 436 | Not_Computer
Z-2132-2012 | 62528 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Extraneal and/or DIANEAL LOW CALCIUM Peritoneal Dialysis Solution in Plastic Container, PL-146 used with Automated Peritoneal Dialysis System/Automated PD Set with Cassette; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. The HomeChoice dialysis system is a personal cycler automated peritoneal dialysis system for pediatric and adult renal patients with fill volumes ranging from 60 mL to 3000 mL. DIANEAL Low Calcium peritoneal dialysis solutions are indicated for use in chronic renal failure patients being maintained on peritoneal dialysis. - and- The Personal Cycler Peritoneal Dialysis System [cycler and ancillary sets] is used for automatic control of dialysate solution exchanges in treatment of renal failure patients undergoing peritoneal dialysis. The Personal Cycler System automatically cycles a prescribed volume of dialysis solution into and out of the peritoneal cavity during the dialysis treatment. | Extraneal & Dineal PD Solution w/Automated PD Set w/Cassette | 2 | August 02, 2012 | 2012 | Baxter Healthcare Renal Div | Baxter discovered that a Baxter technical service representative was providing incorrect instructions for inspection of solution bags for a blocked frangible prior to connection to the Automated PD Set with Cassette. | TRAINING: Employee Error | Baxter Healthcare Corp. sent "Important Product Information" letters dated July 6, 2012 to the affected home patients and their clinicians, via first class mail on the same date. The letter identified the product, problem and actions to be taken. The customers were instructed to refer to the Extraneal and/or Dineal package insert, Directions For Use - Preparation for Administration and the HomeChoice APD Systems At-Home Guide for proper procedures to connect to a solution line. Customers are instructed to retain a copy of the letter. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Global Technical Services, Renal Technical Support at 1-800-553-6898, option 1, 24 hours a day.For questions regarding this recall call 1-800-422-9837. |  |  |  | Nationwide Distribution includingArizona, California, Ohio, Pennsylvania and Virginia. | Terminated | October 15, 2012 | 74 |  | Not_Computer
Z-2133-2012 | 62371 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | I-STAT EG7+ Cartridge Abbott Point of Care, Princeton, NJ.The i-Stat EG7 cartridge is intended to be used for quantitatively measure sodium, potassium, ionized calcium, oxygen, carbon dioxide, partial pressure in arterial, venous, or capillary whole blood. | ISTAT EG7 Cartridge | 2 | August 02, 2012 | 2012 | Abbott Point Of Care Inc. | i-Stat EG7 cartridges may generate falsely depressed PCO2 results for samples above 26 mm hg when stored at room temperature for four weeks or longer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Point of Care, Inc sent a Urgent Recall Notice letter dated January 14, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of Cartridges from the identified lot. All unused cartridges will be credited and are to be returned as instructed on the enclosed Business Reply Card.If you have any questions regarding this information please contact your Abbott Point of Care Technical Support at 800-366-8020 Option 1, or your Abbott Point of Care representative. | Contact the recalling firm for information | 52750 | 52,750 cartridges | USA ( nationwide ) | Terminated | May 22, 2013 | 293 | 52750 | Not_Computer
Z-2138-2012 | 62514 | tank, holding, dialysis | FIN | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TANGO3 Water Storage Tank with Ozone Disinfection System, Tango3, LLC, 132 Citizens Blvd., Simpsonville, KY 40067. Models are WOZ-250-PE, WOZ-240-SS, WOZ-180-SS. The TANGO3 Water Storage Tank with Ozone Disinfection System is intended to be used for disinfection of the water distribution system of a dialysis facility. | TANGO3 Water Storage Tank with Ozone Disinfection System | 2 | August 03, 2012 | 2012 | Tango3 LLC | During an FDA inspection, the firm was notified of Quality System deficiencies, which included lack of procedures for design control and customized modifications that were either not reviewed or received proper clearance from FDA. There was no assurance that all systems in the field were working properly and operating at a minimum of 0.2ppm of ozone. | PRODUCTION CONTROLS: Process Control | On 3/26/2012, the firm sent "Urgent: Customer Notification" letters and Acknowledgement Forms to their customers. The issue was described and recommended actions were provided. Customers were given the number 866-782-7286 to call if questions. |  | 44 | 44 units | Product was shipped to the following states: CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, NH, NJ, NM, OH, PA, RI, TX and VA. | Terminated | March 01, 2013 | 210 | 44 | Not_Computer
Z-2144-2012 | 62617 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Smith & Nephew Dyonics Electroblade Resector, 4.5mm Full Radius Blade, BONECUTTERPart Number: 72202213Product Usage:The Dyonics Electroblade Resector is indicated for use in arthroscopic surgical procedures of large and small articular cavities. The Dyonics Electroblade Resector is effective in soft and osseous tissue resection and in hemostasis of bleeding vessels. It is intended for arthroscopic procedures using saline solution, Ringers lactate or other conductive solutions as an irrigant under direct or video assisted fiberoptic visualization. | Smith and Nephew Inc. | 2 | August 03, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | Devices labeled with an expiry period of 48 months instead of 36 months; an additional 12 months for which product and package reliability has not been validated. | PRODUCTION CONTROLS: Error in Labeling | Smith & Nephew sent an Urgent - Product Recall 1st Notification Letter dated July 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to examine inventory, quarantine immediately, complete the form attached indicating the lot numbers and quantities. The letter also instructed customers to contact Smith & Nephew Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. They will provide customers with instructions for returning affected products and receiving a credit. Each OUS affiliate shall be notified via (email or equivalent) of the recall by Smith & Nephew Inc.; the affiliate must then notify each facility affected by the recall.For question concerning this recall call 405-917-8585 or 978-749-1630. |  | 340 | 340 units | Worldwide Distribution - US Nationwide and the countries of : Australia, Canada, Korea, Turkey, UAE, and Japan. | Open |  |  | 448 | Not_Computer
Z-2137-2012 | 62583 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | SPIDER Surgical Instrument: Flexible Clip ApplierThe SPIDERÂŽ Surgical Instruments are intended for use in minimally invasive surgical procedures for grasping, mobilizing, dissecting, retracting, cutting, cauterizing, ligating, suction/irrigation and other manipulation of tissues and vessels during a laparoscopic procedure under direct and/or endoscopic visualization. | SPIDER Surgical Instrument: Flexible Clip Applier | 2 | August 03, 2012 | 2012 | Transenterix, Inc | Some clip applier units may be inadequate to achieve clip closure. | PRODUCTION CONTROLS: Process Control | TransEnterix sent an Urgent Device Recall letter date July 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product, remove the affected product from their inventory and either return it or dispose of it to prevent usage. Customers were instructed to complete the attached Inventory Reconciliation Formand return the form even if they no longer have any affected units in their inventory. Customers were instructed to ship the affected product to TransEnterix using the enclosed shipping label. If returning a form only, they should e-mail to adunn@transenterix.com. For questions customers were instructed to call 1-888-879-4111.For questions regarding this recall call 919-765-8420. |  | 89 | 89 boxes(6 units per box) | Worldwide Distribution - USA (nationwide) and the countries of Austria, Germany, Italy, Qatar and Switzerland | Terminated | February 27, 2013 | 208 | 89 | Not_Computer
Z-2151-2012 | 62665 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PopLok Punch labeled in part:"PKL-35M***3.5mm***QTY 1***PopLok Punch***ConMed Linvatec Biomaterials, Ltd. Hermiankatu 6-8L 33720 Tampere, FINLAND***Rx Only***ConMed Linvatec 11311 Concept Blvd Largo, FL 33773-4908 USA***Made in USA***"Intended Use: Bone Punch. | PopLok Punch | 2 | August 06, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled one lot of their PopLok bone Punch 3.5MM because the metal shaft diameter is 4.5mm when it should have been 3.5mm. | TRAINING: Employee Error | The recall is being conducted to the end user level. ConMed Linvatec Urgent Medical Device Recall Notifications were sent by FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with the reason for the recall and health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form, even if there is no affected product available. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169. |  | 2 | 2 | Worldwide Distribution -- US and Korea. | Terminated | November 21, 2013 | 472 | 2 | Not_Computer
Z-2149-2012 | 62563 | pump, blood, extra-luminal | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Plasma Pump Model PP-04aUsed with Asahi Plasmaflo Plasma Separators and plasma exchange tubesets to perform therapeutic plasma exchange. | Pump, blood, extraluminal | 2 | August 06, 2012 | 2012 | Apheresis Technologies, Inc. | On 07/09/2012, Apheresis Technologies, Inc. initiated a recall on the Plasma Pump Model PP-04a due to a complaint received regarding a rotor of the PP-04a pump that stopped during a therapeutic plasma exchange treatment causing interruption of flow of both the plasma discard line and replacement fruit line. | PRODUCTION CONTROLS: Process Control | APHERESIS TECHNOLOGIES, INC. sent a letter dated July 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the device for repair. For questions customers should call 727-787-5616. |  | 13 | 13 | Natonwide Distribution | Terminated | December 16, 2013 | 497 | 13 | Not_Computer
Z-2148-2012 | 62683 | camera, cine, surgical, without audio | FWH | Surgical camera and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Chromovision High Definition (HD) Wireless Camera, Pt # CC-9715900.Visual recording of surgical procedures. | Chromovision High Definition (HD) Wireless Camera | 3 | August 06, 2012 | 2012 | Berchtold Corp. | Defective camera: There is a possibility that the connection of the wires to the connector may become loose, causing image distortion or loss. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BERCHTOLD Corporation sent a letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. BERCHTOLD will replace the customer's cameras at no cost to the customer. Customers with questions should call 843-569-6100, ext 322.For questions regarding this recall call 843-569-6100. |  | 17 | 17 cameras | Nationwide Distribution including FL, MI, MS, and NH. | Terminated | June 06, 2013 | 304 | 17 | Not_Computer
Z-2150-2012 | 62556 | system, measurement, blood-pressure, non-invasive | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Finapres Medical Systems Portapres with NiMh battery packProduct Usage:Continuously measure blood pressure, The Portapres can be worn on the body with a protective waist belt. The device can be used without mains; in that case a battery pack is used. | Finapres Portapres Ambulatory Continuous NonInvasive Blood Pressure Monitor | 2 | August 06, 2012 | 2012 | FMS, Finapres Medical Systems BV | Battery may become hot and leak | PRODUCTION CONTROLS: Process Control | Finapres Medical Systems BV sent a Field Safety Notice letter dated July 6, 2012 to notified US accounts of the recall. The letter identified the affected product, description of the problem and actions to be taken. The letter states that the recalling firm recommends not using the affected battery packs. A replacement battery was included in shipment. Customers were requested to return recalled battery. For questions call +31 (0)20 609 0974 or email info@finapres.com. | Contact the recalling firm for information | 3 | 3 units US | US Nationwide Distribution - including the states of: OH and NY | Terminated | September 12, 2012 | 37 | 3 | Not_Computer
Z-2146-2012 | 62561 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | djo surgical Screw 6.5 mm, Sz 25mm, low profile.Product is intended for the fixation of the acetabular shell or fracture repair | djo surgical6.5 mm cancellous bone screw | 2 | August 06, 2012 | 2012 | Encore Medical, Lp | One lot of the 010-55-030, 6.5 mm Cancellous Bone Screw was recalled because a product complaint identified a mislabeling of the device. Although the package label indicates Size 30mm, the screw dimension is actually 25mm. | PRODUCTION CONTROLS: Packaging | DJO Global sent an Urgent Field Safety Notice dated July 16, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the notice to all thoe who need to be aware within their organization or to any other organization where the potentially affected product was transferred, contact customers to place a replacement order and receive an RMA number, and return all affected devices using the RMA number. For questions customers should call 512-834-6255.For questions regarding this recall call 512-832-9500. | Contact the recalling firm for information | 18 | 18 devices | Nationwide Distribution including AR, CA, FL, ID, OK, PA, TX, and UT. | Terminated | December 04, 2012 | 120 | 18 | Not_Computer
Z-2147-2012 | 62395 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CrossFT Punch, Product Number PFT-45M. Manufacturer: CONMED LINVATEC. The intended use is to create a hole to facilitate insertion of the CrossFT suture anchor into bone. | CrossFT Punch | 2 | August 06, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled lot number #156920, PFT-45M, CrossFT Punch because there is a possibility that the device was etched incorrectly listing the item number as PFT-00M and includes incorrect depth lines for insertion. | TRAINING: Employee Error | The firm sent an "URGENT: Medical Device Recall Notification" letter dated August 12, 2011, to its customers after initial phone notifications. The letter describes the product, problem and actions to be taken. The customers were instructed to: 1) Immediately check their facility's inventory for the product number and lot number listed in the letter. If the product is at their facility, they were instructed to discontinue the use of the product. 2) To return the products to ConMed Linvatec for credit and return the products and Reply Form. 3) If the product was transferred to another facility, complete the Reply Form documenting the transfer of the product, and return i by mail or fax. For questions, contact 727-399-5205. |  | 2 | 2 | Nationwide Distribution, including the states of Colorado and Pennsylvania. | Terminated | November 21, 2013 | 472 | 2 | Not_Computer
Z-2152-2012 | 62535 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 731 EMV+ Uni-Vent, Impact Instrumentation Inc., 27 Fairfield Place, West Caldwell NJ 07006. Indicated for use in the management of acute or chronic respiratory failure or during resuscitation by providing continuous positive -pressure ventilation. | Model EMV UniVent | 3 | August 07, 2012 | 2012 | Impact Instrumentation, Inc. | Ten units of Model EMV+ Devices were shipped out in error before the verification and validation was completed for a new integrated circuit (IC) on the CPU PCB. | PRODUCTION CONTROLS: Process Control | Impact Instrumentation Inc. notified their sole customer via email on 3/15/12 and then emailed an "Urgent Device Product Recall" letter and tracking form dated 3/19/2012. The letter described the reason for the recall and provided recommended actions and the contact number for Customer Service. |  | 10 | Ten Units | US Distribution to the state of Illinois. | Terminated | August 15, 2012 | 8 | 10 | Not_Computer
Z-2155-2012 | 62648 | retractor | GAD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | LUXOR RETRACTOR BASE; Catalog Number: 48250030; Manufactured by: STRYKER Spine SAS; Z.I MARTICOT-33610-CESTAS-FRANCE; Distributed in the USA by: Stryker Spine. 2 Pearl Ct. Allendale, NJ 07401.Provides access to the thoracic and lumbar spine from a posterior approach via a small incision. | Stryker LUXOR RETRACTOR | 2 | August 07, 2012 | 2012 | Stryker Spine | Stryker received three reports involving blades of the Luxor Retractor becoming disassembled from the base causing the retractor to collapse insitu. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker sent Urgent Product Recall letters/return response forms dated April 23, 2012 via FedEx to Stryker branches/agencies and hospitals where product might have been distributed. The letter identified the affected product, discussed the issue, potential hazards, risk mitigations, and the return process. Customers are to examine their inventory and reconcile any of the affected product with the information provided. In addition, the Customer Response Form should be completed and faxed back to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should also be returned following the instructions provided. Questions should be directed to Dan Roberts, Regulatory Affairs Compliance Specialist at 201-760-8298. |  | 30 | 30 devices | Worldwide Distribution -- US, Switzerland, Netherlands, Brazil, Hong Kong, Germany, and Taiwan. | Terminated | November 13, 2013 | 463 | 30 | Not_Computer
Z-2154-2012 | 62595 | tubes, vacuum sample, with anticoagulant | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | 3mL 9NC Coagulation Sodium Citrate 3.2%, Item #454334,Greiner Bio-One North America Inc., 4238 Capital Drive, Monroe, NC 28110.Intended Use: Venous blood collection tubes. | 3mL 9NC Coagulation Sodium Citrate 3.2 | 2 | August 07, 2012 | 2012 | Greiner Bio-One North America, Inc. | Part of the lot received a lower concentration of citrate solution which may cause falsely lower values for common coagulation tests. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Distributors were notified by letter on 07/02/2012 and customers were sent a letter dated 07/05/2012. The letters identified the affected product and the reason for the recall. Distributors were asked to stop distributing the affected product and to complete the attached Product Disposition form. The form along with the affected product was to be returned to GBO. The firm will provide replacement stock or issue a credit once product return and receipt of the Product Disposition form has been confirmed. Customers were also asked to complete the Product Disposition Site Confirmation form and fax it back to Greiner Bio-One North America, Inc. at the number provided. The letter states that the firm will replace any affected product once the form is received. Questions should be directed to 800-515-8112. | Contact the recalling firm for information | 838.9 | 838.9 cases (1,006,700 pieces) | Nationwide Distribution | Terminated | June 11, 2013 | 308 | 838.9 | Not_Computer
Z-2158-2012 | 62375 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | ***REF IFNT611***Full OSSEOTITE Tapered Certain Implant 6 x 11.5mm***Sterile by radiation***Rx Only***BIOMET 3i Dental Iberica S.L. WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain. Lot #2011110798.Intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment in single tooth restorations and in partially or fully endentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridge work and to retain over dentures. | Full OSSEOTITE Tapered Certain Implant | 2 | August 08, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i is recalling their product Full Osseotite Tapered Certain Implant IFNT611. The depth of the implant's internal hex is too shallow. The condition, if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | PRODUCTION CONTROLS: Process Control | BIOMET 3i sent an Urgent Medical Device Recall letter on March 19, 2012, to all affected customers via email or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers are to return any unopened and unused product for replacement at no cost. In addition, customers should contact Customer Service at 1-800-342-5454 for assistance if they encounter issues with the mating component, or if they have any questions or concerns. Customers are to respond to the notification regardless of whether they have product to return. |  | 240 | 240 | Worldwide Distribution -- United States, UK, France, Canada, Mexico, Poland, Lebanon and Hong Kong. | Terminated | December 16, 2013 | 495 | 240 | Not_Computer
Z-2161-2012 | 62575 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Straumann Narrow CrossFit (NC) Screw for CARES Abutment ZrO2. Endosseous dental implantArticle Number: 027.2650Product Usage:Dental implants are intended for surgical placement in maxillary and/or mandibular arches to provide support for prosthetic restorations using a Straumann Narrow CrossFit CARES ZrO2 Abutment in edentulous or partially edentulous patients. Dental implants are intended for immediate placement and function on single-tooth and/or multiple tooth applications when good primary stability is achieved, with appropriate occlusal loading, in order to restore chewing function a Straumann Narrow CrossFit CARES ZrO2 abutment is attached by screw to the dental implant. Multiple tooth applications may be rigidly splinted. In the case of edentulous patients, four or more implants must be used. | Straumann | 2 | August 08, 2012 | 2012 | Straumann USA, LLC | Incorrect Screw (RC) included with NC CARES Abutment Set was not a Narrow CrossFit (NC) Screw for CARES NC Abutment ZrO2 | PRODUCTION CONTROLS: Packaging Process Control | Straumann notiifed accounts by telephone on July 10, 2012 and follow-up Field Safety Correction Action letter dated July 11, 2012, via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to discontinue use of the affected product and follow the instruction provided until replacement is received. For questions call 978-747-2514. | Contact the recalling firm for information | 14 | 14 Distributed in US | US Nationwide Distribution including the countries of FL, IN, LA, NY and TX. | Terminated | September 05, 2012 | 28 | 14 | Not_Computer
Z-2157-2012 | 62536 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | I-STAT PT/INR cartridge Abbott Point of Care Inc. Abbott Park, IL 60064 USAThis cartridge is useful for monitoring patients receiving oral anticoagulation therapy such as Coumadin or warfarin. | ISTAT PT/INR cartridges | 2 | August 08, 2012 | 2012 | Abbott Point Of Care Inc. | The i-STAT PT/INR cartridges have the potential to exhibit incorrectly elevated results. | PRODUCTION CONTROLS: Process Control | Abbott Point of Care Inc, sent a URGENT RECALL NOTICE dated March 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were provided in the letter with instructions in the letter on how to handle the affected product. Please complete and return the enclosed Business Reply card to acknowledge receipt of this communication and installation of the interium software CLEW and also 1) Update all handhelds 2) download appropriate assignment sheets. For further questions please call (609) 454-9000. |  | 3106704 | 2,950,296 cartridges US; 156,408 cartridges OUS | Worldwide Distribution -- (USA) nationwide | Terminated | February 08, 2013 | 184 | 3106704 | Not_Computer
Z-2159-2012 | 62500 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas KRAS Mutation Test for In Vitro Diagnostic UseProduct Usage: Usage:The Cobas KRAS Mutation Test, for use with the Cobas 4800 System, is a real-time PCR test intended for the identification of mutations in codons 12, 13 and 61 of the KRAS gene in DNA derived from formalin-fixed paraffin-embedded human colorectal cancer (CRC). | cobas KRAS Mutation Test | 3 | August 08, 2012 | 2012 | Roche Molecular Systems, Inc. | During the real-time stability testing for the cobas KRAS Mutation Test on the commercial lot PO6778, it was identified that the kit was just within specifications for the codon 61 peak height ratio, and that the kit has trended towards the upper specification limit at the 6 month time point but did passed the 6 month stability testing. The kit is labeled for 9 months shelf life stability. The d | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Roche issued a Safety Board Notice (Prooduct Support Notice 2012-01) and Product Bulletin 2012/03 dated February 24, 2012 to all affected customers. The notice idenfied the affected product, description of situation and actions to be taken. Customers were instructed to discontinue use of lot P06778 by March 4, 2012 and to quarantine and discard any remaining kits from lot P06778. The notice stated that a new lot will be made available as a replacement. | Contact the recalling firm for information | 264 | 264 kits to 13 EU countries; 2 kits to Canada and 33 kits to 11 countries in the Rest of the World | Worldwide Distribution - including the countries of Austria, Belgium, Czech Republic, Denmark, Germany, Italy, Luxemburg, Netherlands, Poland, Slovakia, Slovenia, Spain, United Kingdom, Canada, Chile, Greenland, Hong Kong, Israel, Malaysia, New Zealand, Philippines, S. Africa, Switzerland, Taiwan and Turkey. | Terminated | July 16, 2013 | 342 | 264 | Not_Computer
Z-2160-2012 | 62645 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | The finger cots are made out of natural rubber. They are disposable products used to cover a fingertip bandage when working with liquids and are intended to keep the fingertip bandage dry. Cots are packed in clear plastic boxes, with 60 cots in a boxThey are disposable products used to cover a fingertip bandage when working with liquids and are intended to keep the fingertip bandage dry | Afassco Finger Cots | 2 | August 08, 2012 | 2012 | Afassco Inc | The firm is initiating a field correction because the product(s) were missing the Latex Warning Label. Afassco Inc. is implementing a voluntary field correction for MULTI-COLORED FINGER COTS because product lacked the warning label:"CAUTION: This product contains Natural Rubber Latex which may cause allergic reactions." | PRODUCTION CONTROLS: Packaging Process Control | Afassco sent a Field Correction Notice to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Labels have been included with the notices that the firm will be sending to their distributors. The distributors were asked to examine their inventory and affix the caution label to the boxes of affected product. If additional warning labels are needed, they may contact Afassco (800-441-6774) for additional labels.If they have further distributed these products, they should notify consignees of the field correction. Their notification to their customers may be enhanced by including a copy of the notice sent to them. The recall should be carried out to the user level. For any questions customers may call 1-880-441-6774.For questions regarding this recall call 775-783-3555. |  | 1956 | 1956 Boxes | Nationwide Distribution including CA, FL. GA. HI, IL, IN, MI, NH, NV, TN, TX, and WV. | Terminated | January 14, 2014 | 524 | 1956 | Not_Computer
Z-2181-2012 | 62596 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly Clark Patient Warming System - Model 1000 Control Unit. Kimberly Clark Patient Warming System - Model 1000 Control Unit and Energy Transfer Pads is a thermal regulating system, indicated for monitoring and controlling patient temperature. | Kimberly Clark Patient Warming System Model 1000 Control Unit. | 2 | August 09, 2012 | 2012 | Kimberly-Clark Corporation | A defect of the "fuse holder" can lead to a Control Unit failure that may cause the device to lose its ability to warm and maintain set temperature and may also result in excessive heat with the fuse assembly that could lead to the generation of smoke. | DESIGN: Device Design | Kimberly-Clark sent an "URGENT: FIELD CORRECTIVE ACTION NOTICE" dated April 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Corrective Action Response Form was attached for customers to complete and return. Contact the Field Corrective Action Coordinator at 770-587-7400 for questions regarding this notice. |  | 129 | 129 units | Nationwide Distribution-including the states of AZ, CA, CO, DE, FL, GA, IA, ID, IN, IL, KS, KY, LA, MA, MD, ME, MO, MS, NE, NM, OH, OR, PA, RI, SC, TN, TX, VA, WI and WV. | Terminated | September 05, 2012 | 27 | 129 | Not_Computer
Z-2167-2012 | 62525 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith & Nephew 5.5 mm TWINFIX Ultra Ti Anchor with two #2 sutures (COBRAID-blue, COBRAIDblack), sterilePart Number: 72202896Product Usage: intended for use for thereattachment of soft tissue to bone | Smith & Nephew | 2 | August 09, 2012 | 2012 | Smith & Nephew, Inc. Endoscopy Division | Distal part of the anchor may break on insertion into bone during surgery | DESIGN: Device Design | Smith & Nephew sent an Urgent Product Recall 1st Notification letters dated June 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to inspect inventory and quarantine them immediately, complete and return the attached form indicating the Lot# and quantities that need to be returned and contact Smith & Nephew Returns Group at 800-343-5717 (option 3) or email to endo.andreturns@smith-nephew.com. for instructions for returning affected products and receiving a credit. For questions call 1-508-261-3731. |  | 2204 | 2,204 units | Worldwide Distribution - US Nationwide and the countries of: Argentina, AUSTRALIA, AUSTRIA, Belgium, CANADA,CHILECOLOMBIA, Denmark, DUBAI, FINLAND, FRANCEGDC BAAR Distribution,GERMANY, GREECE, INDIAITALY, KOREA, MALAYSIA, MEXICO, NETHERLAND, NORWAYPORTUGAL, PUERTO RICO, SHANGHAI, SOUTH AFRICASPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Open |  |  | 12189 | Not_Computer
Z-2179-2012 | 62747 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medstream 81" (206 cm) 20 Drop Standard bore Universal Administration Set with Y-Site, SterileManufactured for MSD (Medical Specialties Distributors LLC, Stoughton, MA. Product of ChinaProduct Number: MS721 Disposable Infusion set | Medstream 81" (206cm) 20 Universal Administration Set with YSite | 2 | August 09, 2012 | 2012 | Medical Specialties Distributors, Inc. | Devices leaking at the Y-site | PRODUCTION CONTROLS: Process Control | Medical Specialties Distributors LLC sent an Urgent Medical Recall letter dated August 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and remove all affected product from use immediately. Further distribution of the affected product should be discontinued immediately. Customers were instructed to contact their customers and inform them of the recall; complete and return the acknowledgement form and fax to 781-344-8320 to arrange for a return and to contact the Customer Service Department at (800) 967-6400 with questions. For questions regarding this recall call 781-344-6000. | No consumer action necessary | 472 | 472 cases (50/case) | Nationwide Distribution | Open |  |  | 472 | Not_Computer
Z-2176-2012 | 62415 | instrument, manual, surgical, general use | MDM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker Reliance AL Implant Blocker (Various sizes)The SLIDE Implant Blocker is a reusable instrument, sterilized prior to use. It allows the surgeon to distract the disc space in the lumbar region of the spine using an anterior or anterolateral approach while also allowing them to insert the implant into final position. | Stryker Reliance AL Implant Blocker | 2 | August 09, 2012 | 2012 | Stryker Spine | Stryker Spine has received six reports regarding the breakage of one or more prongs at the extremity of the instrument. | DESIGN: Component Design/Selection | Stryker sent an Urgent Product Recall letters/return response forms dated April 5, 2012, via FedEX to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital location to identify the product. Reconcile any SLIDE Implant Blockers with catalog and lot numbers referenced above. Retrieve and return the affected product using the enclosed pre-paid mailing label to Stryker Spine 59 Route 17 South, Allendale, New Jersey, 07401. For questions customers were instructed to call 201-760-8298. | Contact the recalling firm for information | 109 | 109 Lots | Nationwide Distribution including NC, AR, TN, MA, IL, GA, NY, LA, MN, MO, OH, KS, NM, TX, KS,and IN. | Terminated | March 21, 2013 | 224 | 109 | Not_Computer
Z-2180-2012 | 62341 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Angiographic catheter --- Soft-Vu (JB-1) Non-braided. --- Size: 5F; Length: 100cm; designed to accept guidewire: .035 in. --- Catalog No. [REF]: 10734201. Packaged in boxes of 5 catheters each, with manufacturer listed as AngioDynamics, Inc., Queensbury, NY USA.Intended use: angiographic diagnosis. | SoftVu Angiographic Catheter | 2 | August 09, 2012 | 2012 | Angiodynamics Worldwide Headquarters | AngioDynamics is recalling one lot of Soft-Vu Angiographic catheters because one of the catheter pouches within the outer box may be mislabeled. | MISBRANDING: Labeling False and Misleading | The firm, AngioDynamics, Inc., sent an "Urgent Medical Device Recall - Immediate Action Required letter dated June 15, 2012 to all consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to carefully read this recall notification in its entirety; IMMEDIATELY DISCONTINUE USE OF AND SEGREGATE RECALLED PRODUCT in a secure location for return to AngioDynamics, Inc.; immediately remove recalled product from their inventory; immediately forward a copy of this recall notification to all sites to which they have distributed affected product; and promptly complete, sign and return the Reply Verification Tracking Form (even if you do not have any product to return) via fax to Soft-Vu Angiographic Catheter Recall Coordinator at 1-518-798-1360 or email rdenino@agniodynamics.com. If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday - Friday: Eastern Standard Time) to obtain a replacement. The consignees/customers were also instructed to package and return the recalled product, write the RMA number on the box and send to AngioDynamics, Inc., 603 Queensbury Avenue, Queensbury, NY 12804; Attn: Soft-Vu Angiographic Catheter Recall Coordinator.If you have any questions about this recall action, contact your local Sales Representative or AngioDynamics, Inc. Customer Service Manager at 800-772-6446 or by email at customerservice@angiodynamics.com. |  | 7 | 7 boxes/35 units | Worldwide distribution: USA (nationwide) including states of: CA, KS, MO, and NC and countries of: Czech Republic and Switzerland. | Open |  |  | 7 | Not_Computer
Z-2182-2012 | 62611 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Reagent Management System (RMS) - Power Cord.The Siemens Healthcare Diagnostics Dimension(R) clinical chemistry analyzers are floor model, fully automated, microprocessor-controlled, integrated instrument systems that use prepackaged Flex(R) reagent test cartridges to measure a variety of analytes in human body fluids. | Dimension(R) Reagent Management System (RMS) Power Cord. | 2 | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that partial insertion of the power cord on the Dimension Reagent Management System (RMS) power input module can cause the plug to overheat and potentially create a fire hazard. RMS without a power cord retention clip may be affected. | DESIGN: Device Design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an attached Field Correction Effectiveness Check Form and fax it to (302) 631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 2554 | 2,554 | Distributed in the US including Washington, DC, Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WI, WV, and WY and the Virgin and Pacific Islands. | Open |  |  | 2554 | Not_Computer
Z-2175-2012 | 62609 | radioimmunoassay, vancomycin | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Vancomycin Flex(R) Reagent Cartridge (VANC, Catalog # K4086).For the quantitative measurement of vancomycin in human serum and plasma on the Dimension Vista system. | Dimension Vista(R) VANC Flex(R) Reagent Cartridge | 3 | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has received complaints of customers experiencing slopes | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm initiated their recall of this product on July 17, 2012 by sending an Urgent Field Safety Notification via FedEx next day delivery to all consignees. The Notice informed consignees of the high calibration slopes with the affected device. Affected consignees were requested to immediately discontinue use and discard any remaining Dimension Vista(R) VANC Flex(R) reagent cartridges, lot 12095AD. Affected customers were provided with replacement cartridges. Customers were to also complete the attached form and fax it back to the number provided. The notification letter should be forwarded to anyone, whom customers may have further distributed the affected lot. Questions and concerns should be directed to Siemens Technical Solutions Center at 800-441-9250. |  | 1894 | 1,894 cartons | Worldwide Distribution -- US, including Washington, DC and the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WV, and the countries of Canada, New Zealand, Germany, and Puerto Rico. | Terminated | January 14, 2014 | 523 | 1894 | Not_Computer
Z-2170-2012 | 62533 | test,natriuretic peptide | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat BNP cartridgesAbbott Point of Care Inc. The i-Stat BNP test is an in vitro diagnostic test for the quantitative measurement of B-type Natriuretic Peptide (BNP) in whole blood or plasma samples using EDTA as the anticoagulant. BNP measurements can be used as an aid in the diagnosis and assessment of the severity of congestive heart failure. | iStat BNP cartridges | 2 | August 09, 2012 | 2012 | Abbott Point Of Care Inc. | The frequency of suppressed results for i-Stat cTnl, BNP and CK-MP cartridges is affected by atmospheric pressure. | DESIGN: Device Design | Abbott Point of Care (APOC) sent an URGENT RECALL NOTICE dated February 2012 to all affected customers via e-mail and also contacted by phone. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that if their facility was performing testing at more than approximately 7500 feet above sea level, they should have an alternate test method available should they experience an issue. If they have forwarded any cartridges to another facility, they should provide a copy of the Urgent Recall Notice to them. For questions customers were instructed to call 800-366-8020, Option 1 or contact their Abbott Point of Care representative.For questions regarding this recall call 609-454-9000. |  | 1058400 | 1,004,200 cartridges US; 54,200 cartridges OUS | Worldwide Distribution - USA including AZ and CO and internationally to Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Chile, China, Colombia, Costa Rica, Dominican Republic, El Salvador, Finland, France, Georgia, Germany, Greece, Honduras, Hong Kong, India, Ireland, Italy, Korea, Kuwait, Mauritius, Mexico, Netherland (Antilles), Nigeria, Norway, Oman, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, R.O.C., Thailand, UAE, United Kingdom, and Yemen. | Terminated | February 08, 2013 | 183 | 9289349 | Not_Computer
Z-2186-2012 | 62598 | catheter, peritoneal | GBW | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | TAUT, Cholangiogram Catheter Intraducer, Peritoneal Catheter, Teleflex Medical, Research Triangle Park, NC.An introduction/drainage catheter accessory is intended to aid in the manipulation of or insertion of the device into the body. | TAUT, Cholangiogram Catheter Intraducer | 3 | August 10, 2012 | 2012 | Teleflex Medical | Product is missing the label on the internal packaging. | PRODUCTION CONTROLS: Packaging Process Control | Consignees were notified by letter on 07/20/2012. |  | 611 | 611 units | AK, AL, AZ, CA, CO, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, MS, NC, NY, OH, PA, SD, TX, UT, VA, WI, WV. | Open |  |  | 611 | Not_Computer
Z-2187-2012 | 62499 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The NCB Polyaxial Locking Plate System NCB screws 26mm: 02.03155.026, non-sterileProduct is implantable screws indicated for temporary internal fixation and stabilization of fractures and osteotomies of long bones. These screws are intended for bicortical anchorage. | NCB Polyaxial Locking Plate System NCB screws | 2 | August 10, 2012 | 2012 | Zimmer, Inc. | Zimmer investigation following a complaint found one lot of NCB Polyaxial Locking Plate System NCB screws, 26mm, distributed in the United States contained incorrect 32 mm screws. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 13, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the implicated lot in their inventory and user accounts; return all packaged units of this lot; and provide Zimmer with the name, complete mailing address, and phone number of each user facility and surgeon that have used the product via E-mail to dawn.kindle@zimmer.com. Zimmer will send a follow-up Urgent Recall notification to these sub-accounts with instructions to verify the screw length of any screws already in the implant rack inventory prior to use in surgery. All implicated product is to be returned. The customers were also instructed to complete and return the Inventory Return Certification form to Zimmer via fax at +1-574-372-4265. If you have any questions about this correction and/or product removal, contact Zimmer at +1-877-946-2761. | Contact the recalling firm for information | 31 | 24 to US Distributor with 7 further distributed | Nationwide distribution: USA including states of: AZ, IL, IN, ME and NJ. | Terminated | September 24, 2013 | 410 | 31 | Not_Computer
Z-2174-2012 | 62559 | tube tracheostomy and tube cuff | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Reusable Cannula Low Pressure Cuffed Tracheostomy Tubes, model numbers (Worldwide) 8LPC, 8FEN; (Outside U.S.) 8FENJ, 8FENJ-S, 8LPC-S, 8FEN-S. Product Usage: The product is intended for use to bypass upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. The product is sold as sterile. The packaging may indicate either Covidien or Tyco as the manufacturer as the firm has transitioned from Tyco to Covidien over the past 4 years. | Shiley | 1 | August 10, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is recalling certain cuffed Shiley tracheostomy tubes due to consumer complaints related to the connections and functions of the inner and outer cannula which were observed during mechanical ventilation of patients. | DESIGN: Component Design/Selection | Covidien sent an Urgent Medical Device Recall letter dated July 18, 2012 to all affected consignees via mail (FedEx) notification. The customer communication identified the affected product , problem and actions to be taken. Customers were instructed to examine their inventory, segregate product with the identified affected product codes and lot numbers and complete the enclosed verification form and fax to 925-463-4600. Customers were asked to return unused products from the affected product codes and lots. The firm also recommended the replacement of affected tubes that were already in use as soon as clinically appropriate. The letter stated to return the affected product for credit or question contact Technical Services group at 800-635-5267 option 3, option 1 and reference the Return Goods Authorization number. |  | 337227 | 337,227 units | Worldwide Distribution - US Nationwide | Open |  |  | 337227 | Not_Computer
Z-2183-2012 | 61855 | [6 empty] | BeautyAngel Health Fitness Natural Collagen Stimulation Anti AgingFor human use. | Beauty Angel | 2 | August 10, 2012 | 2012 | JK Products & Services, Inc | The Beauty Angel was marketed without a 510k. |  | JK Products sent a Mandatory Field Correction Notice - Immediate Action Required letter dated February 10, 2012, to all affected customers. The letter identified the product the problem, and the action to be taken by the customer. The firm instructed customers that they would be replacing the operating instructions for the affected product with new ones that meet the requirements defined by the regulatory authorities. Customers were instructed not to make claims that indicate that using the affected product will result in medical or health benefits, or that directly or indirectly imply that the affected product is intended to cure, treat, diagnose or pevent an illness, or disease. Customers were asked to complete the new owner information request on the return post care enclosed and return to the firm. If product was further distributed those accounts should be contacted immediately and should be advised of this Mandatory Field Correction Notice. For questions customers should call 870-935-1130. For questions regarding this recall call 870-268-2702. | Contact the recalling firm for information | 598 | 598 units | Worldwide Distribution - USA (nationwide) and Canada | Open |  |  | 598 | Not_Computer
Z-2188-2012 | 62278 | catheter, flow directed | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Thermodilution VIP Catheters Model 131F7J.For use in patients who require hemodynamic monitoring. | Edwards Lifesciences SwanGanz Thermodilution VIP Catheters | 2 | August 12, 2012 | 2012 | Edwards Lifesciences, LLC | The device is being recalled because the firm received multiple complaints related to passage of a J-tip guidewire when inserted through the distal lumen hub of the catheter. | PRODUCTION CONTROLS: Process Control | Edwards Lifesciences sent a "PRODUCT RECALL" letter dated May 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Confirmation Form was attached for customers to complete and return via fax to 1-800-422-9329. Contact Customer Service at 1-800-424-3278 for questions regarding this notice. | Contact the recalling firm for information | 32145 | 32,145 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of Canada, Japan, EMEA (Europe, Middle East and Africa), Asia Pacific, (Hong Kong, Indonesia, Malaysia, Australia, and Taiwan, Thailand), Latin America, (Chile, Colombia, Panama, Mexico, and Honduras). | Open |  |  | 32145 | Not_Computer
Z-2200-2012 | 62588 | retractor, self-retaining, for neurosurgery | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | 13.5 mm x 50 mm BrainPath SheathThe NICO BrainPath is a family of products that consists of multiple sized reusable and re-sterilizable obturators with coordinating single patient use, disposable sheaths (three lengths available for each: 50mm, 60mm, and 75mm). The obturator and sheath are assembled in the OR immediately prior to use. These two components are held together by an interference fit. The device when used as intended is designed to provide access to neurological tissues. | 13.5 mm x 50 mm BrainPath Sheath | 2 | August 13, 2012 | 2012 | Nico Corp. | BrainPath sheaths were mislabeled. Sheaths which measure 50mm were labeledas 60mm. The length callout on the label does not match the actual length of the sheath within thepackaging resulting in an improper fit as further explained below. This mislabeling wasdiscovered on 29-June-2012 during a clinical case. | PRODUCTION CONTROLS: Labeling Mix-Ups | NICO Neuro and Spine sent an Urgent Recall letter to all affected customers on July 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the recall response form.Affected sheaths distributed within the US will be obtained and discarded by NICO. Affected sheaths shipped to Canada will be obtained and discarded by NICOs distributor, Galen Medical.For questions customers should call 888-632-7071. |  | 20 | 20 devices (4 boxes, 5 devices per box) | Worldwide Distribution - USA including MI and Canada | Terminated | November 07, 2012 | 86 | 35 | Not_Computer
Z-2199-2012 | 62589 | monitor, carbon-dioxide, cutaneous | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec, V-Sign"2 Sensors The SenTec Digital Monitoring System (SMDS) is used for continuous and noninvasive real-time monitoring of ventilation (tcPCO2) and oxygenation (SpO2). | SenTec, VSign"2 Sensors | 2 | August 13, 2012 | 2012 | Sentec AG | SenTec is recalling a number of V-Sign"2 Sensors due to a problem regarding configuration parameter writing during manufacturing. | DESIGN: Device Design | SenTec sent a Field Safety Notice dated July 12, 2012, via email to all affected customers. It identified the product the problem and the action needed to be taken by the customer. Customers were instructed the affected Sensors in the field must be replaced.Please locate the affected Sensors and return them to SenTec for analysis immediately. SenTec will provide you with replacement Sensors within the next three months.This notice needs to be passed on to all those who need to be aware within your organization. Please transfer this notice to other organizations on which this action has an impact.Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. Further questions please call (262 ) 594-2114. | Contact the recalling firm for information | 8 | 8 | Worldwide Distribution -- Nationwide including the state of Missouri,. and the countries of Germany, Spain, France, Netherlands, Australia, Belgium, Thailand. | Terminated | October 11, 2012 | 59 | 8 | Not_Computer
Z-2185-2012 | 55121 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module model 8100.Subsequent product code: FPAThe Pump module is intended for facilities that utilize infusion for the delivery of fluids, medications, blood, and blood products using continuous or intermittent delivery through clinically acceptable routes of administration; such as, intravenous (IV), intra-arterial (IA),subcutaneous, epidural, enteral, or irrigation of fluid spaces. The Pump module is indicated for use on adults, pediatrics, and neonates. | Alaris Pump Module model 8100 | 1 | August 15, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion identified a potential risk associated with the Alaris Pump module model 8100. Carefusion has received reports of customers experiencing motor stalls during infusion with Alaris Pump Module (model 8100) manufactured between August 2010 and July 2011. |  | CareFusion sent an Urgent Medical Device Recall Notification letter dated July 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required action for users instructed high risk patients undergoing infusions at high rates, consider having additional devices as back up. Clinicians should weigh the risk/benefit to the patients before continuing to use the device. Customers with recall related questions were instructed to contact Carefusion support center at (888) 562-6018. Customers with adverse event report questions were instructed to contact customer advocacy at (800) 854-7128, option 1, option 1, option 3 or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system were instructed to contact Technical support at (888) 812-3229.For questions regarding this recall call 858-617-4000. |  | 102248 | 102,248 units | Worldwide Distribution - USA (nationwide) and the countries of: South Africa and Canada.Additional Countries: Netherlands. | Open |  |  | 102248 | Not_Computer
Z-2209-2012 | 62586 | analyzer, chemistry, micro, for clinical use | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Sample Tube Input Barcode Clips, catalog number 03137040001. Input S-tube barcode clips hold sample input tubes in the COBAS AmpliPrep Instrument (CAP) sample racks and have a barcode to identify the sample. | COBAS AmpliPrep Sample Input Barcode Clips | 2 | August 15, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain lots of COBAS AmpliPrep (CAP) S-tube Input (Box of 12 bags 12x24 S-tube and 12x24 S-clip) can cause the sample input tubes to stick in the barcode clips. As a consequence, the CAP instrument will lift both the sample input tube together with the barcode clip from the sample rack, resulting in an aborted run. No erroneous results would be produced if the issue occurs, but samples in proce | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 1/17/2012 an "Analyzer Bulletin" notification was sent to Roche affiliates and on 2/1/2012 an "Analyzer Bulletin" notification was sent to all affected US Consignees. The notification described the product issue and provided recommended actions. Customers were referred to Roche Technical Support at 800-526-1247 for questions. | Contact the recalling firm for information | 2132 | 2,132 boxes | Nationwide Distribution (USA). | Terminated | July 16, 2013 | 335 | 2132 | Not_Computer
Z-2211-2012 | 62612 | enzyme immunoassay, digitoxin | LFM | Digitoxin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products DGXN Slides --- [REF] 834 3386. Product is packaged in a box with multiple slides. The lot and expiration codes are found on the side of the package. Product Usage: Intended use: For in vitro diagnostic use only. VITROS Chemistry Products DGXN Slides quantitatively measure digoxin (DGXN) concentration in serum and plasma using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. | VITROS Chemistry Products Digoxin (DGXN) Slides | 2 | August 15, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics revised the VITROS Chemistry Products DGXN Slides Instructions for Use (IFU) and removed heparin plasma as a recommended specimen type for VITROS DGXN Slides | PRODUCTION CONTROLS: Process Control | Ortho Clinical Diagnostics sent an Urgent Field Safety Notice dated June 28, 2012, via US Postal Service Priority Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Immediately discontinue using heparin plasma samples and use only serum samples when processing VITROS DGXN Slides; Notify their laboratory and medical staff of this change and forward the information if the product have been distributed outside of your facility; Replace the pages of their VITROS Instructions for Use Manual with the enclosed Instructions for Use and Table of Contents; and Discontinue using any copies of the VITROSÂŽ DGXN Slide IFU (Version 7.0).--- On 28 June 2012 an Urgent Field Safety Notice was sent via US Postal Service Priority Mail to all affected Distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS DGXN Slides. --- Foreign affiliates were informed by email on 28 June 2012 of the issue and instructed to notify their consignees of the issue and actions. The letters included a Confirmation of Receipt form to be completed and return. For questions contact Customer Technical Servives at 1-800-421-3311. |  | 13338 | 13338 sale units | Worldwide Distribution - US (nationwide) | Open |  |  | 13338 | Not_Computer
Z-2210-2012 | 62471 | cannula, arterial, cardiopulmonary bypass (cpb), embolism protection | NCP | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Getinge Group Pediatric Arterial Cannula Maquet Cardiopulmonary AG Kehler Strabe 3176437 Rastatt Germany Phone +49 7222 932-0Fax: +49 7222 932-1888info.cp@maquet.comwww.maquet.comUsed to introduce blood and sterile solution to the circulation system during extracorporeal circulation procedures | Maquet Getinge Group Pediatric Arterial Cannula | 2 | August 15, 2012 | 2012 | Maquet Inc. | Certain pediatric arterial cannulas may have limited flexibility due to an increased length of the "hardened" cannula tip. | PRODUCTION CONTROLS: Process Control | Maquet issued Urgent Device 2, Recall letters/Urgent Product Recall Forms dated May 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their current stock and identify any affected product and remove from their inventory. Customers were instructed to contact the customer service department at 800-777-4222 (option #2, followed by option #2). Customers were also instructed to completed the enclsoed Recall Form to acknowledge receipt of the notification and to document the current status of the devices shipped to their facility. For questions regarding this recall call 973-709-7660. |  | 142 | 20 devices US; 122 devices foreign | Worldwide Distribution - USA including MI, MN and Internationally | Open |  |  | 142 | Not_Computer
Z-2063-2012 | 61747 | apparatus, exhaust, surgical | FYD | Air-handling apparatus for a surgical operating room. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Stryker Neptune Bronze Rover, Part 0700-007-000Neptune Bronze Rover Waste Management System Instructions For Use.Intended for use in the surgical environment or physician's office to collect and dispose of surgical fluid waste. | Neptune Waste Management System | 1 | August 15, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received two reports of serious injury as a result of tissue damage associated with the use of the Neptune 2, including an event in which one customer connected the Neptune 2 System to a passive chest drainage tube post operatively, resulting in a fatality. The current IFU did not specifically warn against connecting all Neptune Waste Management devices, which is a high vacuum/high fl | MISBRANDING: Labeling False and Misleading | On June 5, 2012, Stryker notified customers that it was recalling the IFUs for the above products. The current IFU did not specifically warn against connecting the Neptune Rover, which is a high vacuum/high flow device, to a passive drainage tube. Customers were instructed to review the revised IFU, distribute to affected departments, and educate users of the Neptune on this warning. Customers must confirm with Stryker via business reply form that they have completed these actions. Customers who have the Neptune 1 Gold, Neptune 1 Gold International or Neptune 2 Bronze will receive a follow up mailing containing warning labels for the device and instructions detailing how to apply them. Customers may continue to use the Neptune 1 Gold, Neptune 1 Gold International, and the Neptune 1 Bronze without any further actions.On September 18, 2012, Stryker notified customers via letter delivered by FedEx overnight delivery that it is expanding the recall on the Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) because FDA has also advised Stryker that these devices require, but do not currently have, 510(k) clearance. FDA is therefore unable to determine whether these devices are as safe and effective as their legally marketed predicate, the Neptune 1 (Gold) Waste Management System (510(k) K012992). As such, Stryker has ceased distribution of the Neptune Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) devices until FDA clears these devices.At this time, FDA does not consider the Neptune Silver, the Neptune 2 Ultra (120V) or the Neptune 2 Ultra (230V) to be legally marketed devices because their safety and effectiveness have not yet been determined. As such, FDA advises that the devices not be used. However, customers who do not have an alternative device to use should weigh the risks and benefits associated with continued use of these devices. If customers choose to continue use of the Neptune Silver, Neptune 2 Ultra (120V) or Nept. | Contact the recalling firm for information | 1075 | 1075 | Worldwide Distribution -- USA (nationwide) and country of: Canada.. | Open |  |  | 9933 | Not_Computer
Z-2207-2012 | 62650 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, X-Ray System, Diagnostic, general-purposeModel: 712070.The Essenta DR is a multifunctional stationary X-ray Intended Use From Labeling): system, in which the swivel arm rotation, the vertical movement of the swivel arm and the SID adjustment are motorized. It is designed for radiographic examination of the recumbent, standing or seated patient and also for lateral exposures. It enables the operator to quickly and safely move the unit to all the pre-programmed basic positions at the press of a button, using the autopositioning feature. Fine positioning of the tube/collimator and detector on the patient is easy to carry out using the command arm. | Philips Essenta DR, XRay System | 2 | August 15, 2012 | 2012 | Philips Healthcare Inc. | Essenta DR C-Arm may fall | PRODUCTION CONTROLS: Equipment Maintenance | Philips Medical sent an "URGENT FIELD SAFETY NOTICE" dated July 09, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to inspect the C-Arm unit and discontinue use if not parallel to the column, do not use the system and call for a Service Technician. A Field Change Order (FCO) is being released that requires an inspection of C-ARM connection and, if necessary, repair of systems. The bolts will be disconnected and thread inserts will be checked. If not the thread inserts are no longer secured, additional holes will be drilled to use in place of the original ones. For further information or support concerning this issue, contact your local Philips representative at 1-800-722-9377. | Contact the recalling firm for information | 7 | 7 units | Worldwide Distribution-USA including the states of CA, NC, and TN and the countries of China, Germany, and Hungary. | Terminated | March 01, 2013 | 198 | 7 | Not_Computer
Z-2213-2012 | 62540 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | 00-4309-028-01 - Trabecular Metal Reverse Shoulder Liner 40mm 65ÂŽ Neck AnglePer surgical technique 97-4309-003-00 Revision 4: "Complete Poly Liner insertion by snapping the appropriate 60ÂŽ (Standard) or 65ÂŽ (Retentive) Poly Liner Impactor to the Poly Liner Impactor Handle. Then, place the Poly Liner Impactor on the Poly Liner articular surface and forcefully strike the Poly Liner Handle with a mallet." | Trabecular Metal Reverse Shoulder Liner 40mm 65 Neck Angle | 2 | August 16, 2012 | 2012 | Zimmer, Inc. | Lot specific Trabecular Metal (TM) Reverse Shoulder Liner Impactors are being recalled due to reports of alignment peg fractures. All lots of TM Reverse Shoulder Liner Impactors manufactured from 455 stainless steel are being retrieved, as all reported fractured instruments were manufactured with thismaterial. The TM Reverse Shoulder Liner Impactor units that were manufactured with 13-8stain | DESIGN: Device Design | Zimmmer sent an Urgent Medical Device Recall letter dated July 18, 2012, to all affected customers. The letter identified the product, the problem, and thea action to be taken by the customer. Customers were instructed to inspect all TM Reverse instrument sets for affected lot(s)(instruments. If affected product was found customers were instructed to immediately place an order to backfill the identified product with Customer Service. If an instrument was was further distributed customers were to provide a copy of the Notice to ensure their customers were aware of the removal. Customers were also instructed to provide a list of all hospitals that currently use the affected product. For questions customers were instructed to call 800-348-2759.For questions regarding this recall call 574-371-1689. | No consumer action necessary | 252 | 252 (210 in USA) | Worldwide Distribution - USA (nationwide) and internationally to Germany, Belgium, Denmark, Finland, France, Italy, Netherlands, Spain, Sweden, Switzerland, and the United Kingdom. | Open |  |  | 724 | Not_Computer
Z-2217-2012 | 62766 | [6 empty] | DePuy Custom Implant DevicesPackaging: Packaging was not standardized on customs, it was dependant on the geometry of the productProduct Usage:Description: Devices meeting the regulatory definition of a custom medical device are those which are provided by prescription at a physician s request to meet a specific patient need and have been exempt from FDA premarket review under the Federal Food, Drug, and Cosmetic Act since 1976. Intended Use: See information contained in the Information for Use for Custom Implant Device (IFU-0902-00-500 rev C) | DePuy Custom Implant Devices | 2 | August 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | In connection with a Warning Letter received from the U.S. Food & Drug Administration on December 9, 2011, DePuy Orthopaedics is taking steps to retrieve custom device implants that surgeons may have ordered from DePuy Orthopaedics, but not implanted. DePuy Orthopaedics received a Warning Letter from the FDA on December 9, 2011 regarding custom device implants. DePuy Orthopaedics has discontinue | PREMARKET APPROVAL: No Marketing Application | DePuy sent a written communication letter dated July 31, 2012 to all consignees. The communication to the hospital customers who received the devices includes a "Surgeon Letter" with instructions to provide the surgeon letter to the Surgeon users. All DePuy Distributors will be notified via email. The sales representatives notified the hospital and surgeons either by mail or in person with a written communication. The letters identified the affected product, problem and actions to be taken. Customers are instructed to complete the attached Roponse form and return to fax number 574-371-4939 or email: kseppa@its.jnj.com. The sales representatives are expected to aid customers in returning of unused Custom Device implants, if needed. The devices can be returned through the normal DePuy Returns process, to attention of "Returns" and marked "Custom Returns" on the outside of the box.Questions Regarding the RemovalQuestions regarding the removal may be directed to the following DePuy representatives -Questions about removal information provided: Katie Seppa, Manager of Post-Market Surveillance, 574-372-7333 (M-F; 8 am  5 pm EDT)-Clinical related questions from surgeons: Direct to DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am-5 pm EDT). |  | 8338 | 8338 | US Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, HI, IA, ID, IL, IN, KS, KY, LA, MA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 13, 2014 | 515 | 8338 | Not_Computer
Z-2206-2012 | 62592 | wax,bone | MTJ |  |  | General & Plastic Surgery | 510(k) | Synthes(R) Hemostatic Bone Putty, Part numbers 08.901.001.97S, 08.901.001.98S, 08.901.001.99S, 08.901.001D, and VB1025.10S.Indicated for use as a water soluble implant material and for use in the control of bleeding from bone surfaces. | Synthes(R) Hemostatic Bone Putty | 1 | August 16, 2012 | 2012 | Synthes USA HQ, Inc. | There is the potential for Hemostatic Bone Putty to ignite if contacted with electrosurgical cautery systems under certain conditions during surgery. |  | The firm initiated their recall of this product on July 5, 2012 by sending a "Notice: Medical Device Recall" letter to their consignees. The letter identified the affected product along with the reason for recall; and, requested that consignees immediately cease using the product, examine their inventory, and remove the product from inventory. The letter further instructed consignees to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, and to complete and return the verification form that was included with the letter. Questions should be directed to 610-719-5450. |  | 8853 | 8,853 packs | Worldwide Distribution -- US, including Washington, DC and states of AK, AZ, CA, CT, FL, GA, ID, KS, LA, MD, MA, MI, MN, MO, NE, NH, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI, and the country of Canada. | Open |  |  | 8853 | Not_Computer
Z-2218-2012 | 62732 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 8000 BaseProduct Usage:The Sarns Modular Perfusion System 8000 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. | Sarns" Modular Perfusion System 8000 | 3 | August 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During a retrospective review of complaint events, TCVS identified on 06/15/12 that three (3)Sarns" System 8000 bases were manufactured and distributed from 11/13/06-11/29/06 with anincorrect incorrect amperage rating identified on the name plate label. The mislabeling wasidentified in January 2007 a set up process of the automark label printer. All units were correctedin the field in Marc | PRODUCTION CONTROLS: Labeling Mix-Ups | Terumo field correction/service records were completed in 2007 and all affected units were corrected in the field in March 2007 by replacing the incorrect label at the customer location. For questions about this report call 734-741-6056. |  | 3 | 3 | US Nationwide Distribution - including the state of New York | Terminated | November 07, 2012 | 83 | 3 | Not_Computer
Z-2220-2012 | 62649 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Altrx" Altralinked" Polyethylene linersDescription: An acetabular liner is one of the implants used during total hip arthroplasty to replace the damaged hip joint to help increase patient mobility and to reduce pain. Packaging: Product is placed into a PETG tray, sealed with a Tyvek lid. This is placed into a second PETG tray and sealed with a Tyvek lid. This package is placed into a paperboard carton, labeled, and shrink-wrapped. Intended Use: An acetabular liner is used in combination with other implants for total hip arthroplasty is intended to provide increased patient mobility and reduce pain by replacing the damaged hip joint articulation in patients where there is evidence of sufficient sound bone to seat and support the components. | Altrx" Altralinked" Polyethylene liners | 2 | August 17, 2012 | 2012 | DePuy Orthopaedics, Inc. | In the recalled lot, the products were not machined according to the production router, a size 52 cup (122136052) was machined instead. This resulted in a size 52 liner being etched and labeled as a size 54 liner. | TRAINING: Employee Error | Depuy sent an URGENT MEDICAL DEVICE RECALL notification dated July 24, 2012, to all affected consignees. The recall is extended to the Depuy Distributor level and hospital level. All DePuy Distributors were notified via email. The sales representatives notify the hospital either by mail or in person with a written communication.Distributors and customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-16 on the outside of the box. Reconciliation forms should be returned to the DePuy Sales Representative or faxed to 574-372-7567. For clinical questions from Surgeons contact DePuy's Scientific Information Office at 1-888-554-2482. For questions regarding recall information call 574-372-7333. | Contact the recalling firm for information | 20 | 20 liners | Worldwide Distribution - USA including PA and the country of Canada | Terminated | April 30, 2013 | 256 | 20 | Not_Computer
Z-2222-2012 | 62576 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Conformis iTotal CR-Cruciate Retaining Knee Replacement System Model Numbers:M57250600010 iTotal CR, Left KneeM57250600020 iTotal CR, Right KneeProduct Usage : Usage:The iTotal CR Knee Replacement System (KRS) is intended for use as a total knee replacement in patients with knee joint pain and disability whose conditions cannot be solely addressed by the use of a prosthetic device that treats only one or two of the three knee compartments, such as a unicondylar, patello-femoral or bi-compartmental prosthesis. | Conformis | 2 | August 17, 2012 | 2012 | Conformis Inc | Addendum Labeling iTotal CR Surgical Technique with Detailed Instructions of Total CR Poly Insertion Technique | DESIGN: Device Design | ConforMIS sent an Urgent Field Safety Notice letter dated July 5, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. An Addendum is included that provided updates to the Surgical Technique Guide. The updated technique guide will be delivered to all iTotal surgeons along with an on-site vist from ConforMIS sales representative. Customers were instructed to place the Surgical Technique Guide Addendum alongside the surgical plan, it puts the information in a highly visible location for the surgeon to review prior to the surgical procedure. Customers were asked to sign the communication verification form attached. For any questions regarding this notice, please contact your local ConforMIS representative or MDSS GmbH, Tel.: +49-511-62628630, info@mdss.com. |  | 1427 | 1427 units | Worldwide Distribution - US Nationwide and the country of Germany | Terminated | February 27, 2013 | 194 | 1427 | Not_Computer
Z-2205-2012 | 62642 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module Model 8100 (formerly Medley Pump Module)Subsequent Product Code: FPA.The Pump module is intended for facilities that utilize infusion for the delivery of fluids, medications, blood, and blood products using continuous or intermittent delivery through clinically acceptable routes of administration; such as, intravenous (IV), intra-arterial (IA),subcutaneous, epidural, enteral, or irrigation of fluid spaces. The Pump module is indicated for use on adults, pediatrics, and neonates. | Alaris Pump Module Model 8100 | 1 | August 17, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion has identified a potential risk associated with the Alaris Pump module model 8100. The pump module door keypad overlay may separate from the keypad assembly |  | CareFusion sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 20, 2012. The letter identified the product, problem, and actions to be taken by the customers. The potential risk identified in the letter was that when the pump module door keypad overlay has delaminated (loose, peeled away or separated from the door assembly) this could cause a potential for fluid ingress which could lead to a keypad malfunction causing termination of infusion with alarm. Termination of infusion could result in serious injury or death. Customers were informed that CareFusion will contact their facility with in 60 days of receiving the notification to schedule a visit to replace the door assembly of their affected pump module. Customers with recall related questions were instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports questions were instructed to contact Customer advocacy at (800) 854-7128, option 1, option 1, option 3 or e-mail at customer feedback @carefusion.com. Customers with technical questions were instructed to contact technical support at (888) 812-3229. Customers were instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process. |  | 43862 | 43,862 units total (40,780 units in the US) | Worldwide Distribution-USA (nationwide) and the countries of Australia, European Union, and South Africa. | Open |  |  | 43862 | Not_Computer
Z-2224-2012 | 62745 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Oxford Gap Gauge SML 3/4mm.The instrument is used for an intermediate check of the flexion and extension gaps when preparing the femur to accept the femoral component. | Oxford Gap Gauge | 2 | August 17, 2012 | 2012 | Biomet U.K., Ltd. | A drafting error during an unrelated design change has led to 4 misleading dimensions on the drawing. The CAD modelremained the same thickness but the problem dimensions were accidentally struck from an incorrect point result in anincrease of 1.0mm in the value of each dimension | DESIGN: Device Design | Biomet sent a "FIELD SAFETY CORRECTIVE ACTION" letter dated July 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Response Form was attached for customers to complete and return. Contact the firm for questions regarding this notice. |  | 119 | 119 units | Worldwide Distribution-USA-including the states of IN, SD, and TX and the countries of Austria, China, Japan, Netherlands, Sweden, Turkey, and UK. | Terminated | February 15, 2013 | 182 | 119 | Not_Computer
Z-2221-2012 | 62798 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC 8800 Mobile Fluoroscopic X-ray System, GE Healthcare, Surgery, Salt Lake City, UT.Mobile fluoroscopic examination of human anatomy. | Fluoroscopic XRay System | 2 | August 17, 2012 | 2012 | GE OEC Medical Systems, Inc | X-ray units did not meet manufacturer's specification for linearity. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE updated their user manual and changed their internal specifications to match those of the CFR In addition, a software code correction was to be made to forward productions in order to correct the reported defect.*** This recall issued was previously classified under an incorrect recall number. It is being reclassified with a new number. ***. |  | 175 | 175 units | Worldwide Distribution - USA - Australia, Brazil, Cameroon, China, Colombia, Cyprus, The Czech Republic, Finland, France, Germany, Greece, Italy, Japan, Korea, Malaysia, The Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Taiwan, Turkey, UAE, UK, and Canada. | Terminated | August 17, 2012 | 0 | 175 | Not_Computer
Z-2225-2012 | 62755 | radioimmunoassay, acth | CKG | Adrenocorticotropic hormone (ACTH) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxH Slidemaker Stainer Instrument System, PN 775222.The DxH Slidemaker Stainer is a fully automated slide preparation and staining device that aspirates a whole blood-sample, smears a blood film on a clean microscope slide, and delivers a variety of fixatives, stains, buffers, and rinse solutions to that blood smear. The DxH Slidemaker Stainer allows for adaptation of the smear appearance and stain methodology according to user preferences. Blood smears produced by the Slidemaker portion of the DxH Slidemaker Stainer are moved to baskets for transfer to the Stainer portion by a robot arm. Microscopic examination of the stained blood smears can be used to help determine the hematologic status of a patient. A stained blood film: " Allows for the differentiation of white blood cells " Facilitates the characterization of red blood cells and platelets " Aids in the identification of blood components and cellular abnormalities The optional floor stand has four drawers for storing the reagent and waste containers currently in use in the DxH Slidemaker Stainer. | UniCel DxH Slidemaker Stainer Instrument System | 2 | August 20, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the DxH Slidemaker Stainer Floor Stand shipped with Slidemaker Stainer instruments were missing vinyl caps, exposing sharpt edges on the DxH Slidemaker Stainer Reagent Cabinet (Floor Stand) drawer slides. The sharp edges on the drawer slides may expose operators to physical injury when replacing reagents and waste containers whic | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Beckman Coulter initiated a "PRODUCT CORRECTIVE ACTION (PCA)" which was communicated via phone calls starting the week of December 19, 2011 to all customers who purchased UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System. The phone calls provided the customers with an explanation of the problem identified, potential impact should the problem occur and an action to be taken. Customers were instructed to share the PCA information with their laboratory staff and if they had any questions, they should contact Beckman Coulter at (800) 526-7694. | Contact the recalling firm for information | 41 | 41 units (15 units in the US) | Worldwide Distribution-USA (nationwide) including the states GA, NY, IL, TX, WA, PA, IN, MS, NC, and MO and the countries of Belgium, France, Hong Kong, Italy, Spain, and Switzerland. | Terminated | August 20, 2012 | 0 | 41 | Not_Computer
Z-2227-2012 | 62791 | radioimmunoassay, total triiodothyronine | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Total T3 Reagent, PN: 33830The Access Total T3 assay is used for the quantitative determination of triiodothyronine (T3) levels in human serum and plasma in the assessment of thyroid function and diagnosis of thyroid disorders when T3 levels are between 0.1 and 8.0 ng/mL. | Access Total T3 Reagent | 3 | August 20, 2012 | 2012 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has identified a negative bias in test results when comparing elevated Total T3 patient samples diluted with Sample Diluent A to samples diluted with Access Total T3 Calibrator S0. | DESIGN: Device Design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The PCA letter informed the customers of the problem identified and the actions to be taken. Customers were informed to stop SDA usage as an alternative dilutent and that the S0 dilutent is to be used for manual dilution of the patient sample for Total T3. They will also be informed to no longer use the On Board Dilution (OBD) option for Total T3. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions regarding this notification were instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-4941. | Contact the recalling firm for information | 117385 | 117,385 units total (33,918 units in US) | Worldwide Distribution - USA (nationwide), Canada, and internationally. | Terminated | March 14, 2013 | 206 | 117385 | Not_Computer
Z-2228-2012 | 62640 | enzyme linked immunoabsorption assay, treponema pallidum | LIP | Treponema pallidumtreponemal test reagents. | Microbiology | Microbiology | 510(k) | Siemens ADVIA CentaurÂŽ Syphilis (SYPH) Assay , 200 ReadyPack Reference Number 10492493Product Usage: The ADVIA Centaur Syphilis (SYPH) assay is an in vitro diagnostic immunoassay for the qualitative determination of antibodies to Treponema pallidum in human serum or plasma (EDTA, lithium or sodium heparinized, citrate) using the ADVIA Centaur and ADVIA Centaur XP systems as an aid in the diagnosis of syphilis | Siemens | 2 | August 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Syphilis Assay Interference with HBsAg Assay | DESIGN: Component Design/Selection | Siemens Healthcare sent Urgent Device Recall Notice letters dated July 2012 via e-mail to all affected customers inside and outside the US. The letters identifed the affected product, reason for voluntary recall and risk to health. The letters provided actions to be taken by the customers and if there are any remaining inventory of the affected product, follow the guidelines outlined. Customers were instructed to complete and return the Field Correction Effectiveness Check form. For questions contact your local Technical Support Provider or Distributor. | Contact the recalling firm for information | 2050 | 2050 Kits | Worldwide Distribution - US Nationwide and the countries of:Foreign: Brazil, China, Denmark, Hong Kong, Malaysia, and New ZealandViet Nam | Terminated | November 22, 2013 | 459 | 2050 | Not_Computer
Z-2234-2012 | 62134 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC Proximal Tibial Spacer & Screws (2) 2.5 LT BLUE Size 2.5, 8mm***STERILE***Match insert/spacer/tray color. Use with LOGIC Tibial Components. REF 02--012-42-2508.The Optetrak Comprehensive Knee Systems are indicated for use in skeletally mature individuals undergoing primary surgery for total knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis, and/or post-traumatic degenerative problems. They are also indicated for revision of failed previous reconstructions where sufficient bone stock and soft tissue integrity are present. The Optetrak Comprehensive Knee Systems are contraindicated in patients with active infection, patients without sufficient bone stock to allow appropriate insertion and fixation of the prosthesis, patients without sufficient soft tissue integrity to provide adequate stability, patients with either mental or neuromuscular disorders that do not allow control of the knee joint, and patients whose weight, age, or activity level might cause extreme loads and early failure of the system. | LOGIC Proximal Tibial Spacer Size 2.5 | 2 | August 21, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL recalled their Optetrak Proximal Tibial Spacer after the devices were determined to be dimensionally incompatible to mate as intended with the Logic Tibial Tray. | DESIGN: Device Design | The firm, Exactech, Inc., sent a "IMPORTANT PRODUCT RECALL NOTICE" date October 17, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately cease distribution or use of these products; extend this information to their accounts that may have possession of this product; verify if they have any of the subject Logic size 2.5 Proximal Tibial Spacer products, complete and return the attached PRODUCT RECALL FAX NOTICE form within five (5) working days via fax to: 352-337-3915, and to contact the firm inventory representative to confirm quantities at the location. For questions regarding inventory restocking, please call 1-800-392-2832. |  | 70 | 70 | Nationwide distribution: USA including states of: AL, AZ, CO, FL, GA, IL, LA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TN, VA, and WA. | Terminated | November 07, 2013 | 443 | 70 | Not_Computer
Z-2246-2012 | 61671 | transport, patient, powered | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | Powered Mobility Offers Versatility (MÂŽV)TM ChairThe Powered MOV Chair models EMS, EMS24, HHC, and INO Powered MOV Chair, enable two operators to transport a seated patient weighing up to 500 Ibs up or down stairs. Key features include infinite positioning system, thermal protection system, and locking casters. In addition to the aforementioned features, the battery powered track reduces operator effort and risk of back injury while transporting a patient up or down a set of stairs. | Powered Mobility Offers Versatility (MV)TM Chair1 | 2 | August 22, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Paramed Corporation (d/b/a, Descent Control Systems), a manufacturer and distributor of patient transport devices, merged with an affiliate of Stryker Corporation in August 2009. As part of the due diligence process, Paramed's powered Mobility Offers Versatility (MÂżV) chair was reviewed to identify potential compliance issues in relation to product documentation and classification. | PREMARKET APPROVAL: No Marketing Application | Stryker sent a Urgent Medical Device Recall notice dated June 18, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to " Return the enclosed post card to confirm receipt of this notification." Locate the chairs listed in this notice." Call 269-389-6750 to arrange shipment of the chairs to Stryker Medical." If you have loaned or sold any of the chairs listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location." If you have disposed of any of the affected chairs and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.We apologize for any disruption that this advisory notice may cause. We strive to make products that meet our customer expectations for quality and durability.For further questions please call (269) 389-6927. |  | 159 | 159 units | Worldwide Distribution -- USA (nationwide) including the states of AK, AL, AZ, CA, FL, ID, IN, KS, KY, LA, MA, MD, MN, MO, NC, NJ, NY, OR, PA, TN, TX, UT, UT, VA, WA, WI, and WV, and the country of Canada. | Terminated | December 30, 2013 | 495 | 159 | Not_Computer
Z-2241-2012 | 62729 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE HC OEC 9900 Elite and GE OEC 9800 Fluoroscopes Snubber Boards are a PCB with a 70 amp fuse to provide fused protection of internal system components from the high voltage system components, such as the X-ray tube or High Voltage Transformer. System model number A349855, Snubber Board number 00-880405-03 rev 5. This product is used for image intensified fluoroscopic x-rays for interventional radiologic procedures. | GE HC OEC 9900 Elite and GE OEC 9800 Fluoroscopes | 2 | August 22, 2012 | 2012 | GE OEC Medical Systems, Inc | GE OEC recalled certain 9800 Fluoroscope & 9900 Elite Models because the Snubber Board did not conform to the IEC 60601-1 standard for clearance and creepage of voltage isolation. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | No customer communication took place. On April 26, 2010, GE began opening service cases for field engineers to begin replacing the snubber boards. This action has been completed as of October 29, 2010. | No consumer action necessary | 845 | 845 units | Worldwide and Nationwide (USA) distribution. | Terminated | August 22, 2012 | 0 | 845 | Not_Computer
Z-2237-2012 | 62410 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT cTnI cartridgesProduct Usage:The i-STAT cTnl test is an in vitro diagnostic test for the quantitative measurement of cardiac troponin I in heparinized whole blood or plasma samples. Cardiac I troponin measurement cane be used as an aid in the diagnosis and treatment of myocardial infarction and in the risk stratification of patients with acute coronary syndromes with respect to their relative risk of mortality. | iSTAT cTnI cartridges | 2 | August 22, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care has determined that i-Stat CTNL cartridges have the potential to exhibit incorrectly elevated cardiac troponin I results. This recall was expanded on July 19 2012 to include additional information regarding user documentation. | DESIGN: Device Design | Abbott Point of Care (APOC) sent an Urgent Recall Notice letters, via Fed Ex on March 07, 2012 to all affected customers. APOC sent an important product information update on July 19, 2012 to all of their affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to review the information included as it may pertain to their cartridges. Customers were instructed that if the affected product have been forwarded to another facility, provide a copy of this letter to them. For additional information, contact APOC at 800-366-8020 option 1 or your Abbott Point of Care representative. |  | 7830075 | 7,057675 cartridges US; 772,400 cartridges OUS | Worldwide Distribution - US Nationwide, International | Terminated | February 19, 2013 | 181 | 7830075 | Not_Computer
Z-2235-2012 | 62663 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | The Instructions For Use for the following:Stryker TRIO+ 4.5 x 45 mm STANDARD POST SCREWStryker TRIO+ 4.5 x 50 mm STANDARD POST SCREWStryker TRIO+ 4.5 x 55 mm STANDARD POST SCREWStryker TRIO+ 5.5 x 45 mm STANDARD POST SCREWStryker TRIO+ 5.5 x 60 mm STANDARD POST SCREWManufactured by: STRYKER Spine SAS, Z.I. Marticot - 33610 - Cestas-France +33 (0)5.57.97.06.30 http://www.stryker.comDistributed in the USA by: Stryker Spine, 2 Pearl Ct., Allendale, NJ 07401-1677 USA +1-201-750-8000.The system is intended for posterior, non-cervical pedicle and non-pedicle pedicle fixation of the spine. It is indicated for: degenerative disc disease (DDD), spondylolisthesis, fracture, dislocation, spinal stenosis, scoliosis kyphosis, lordosis, spinal tumor, pseudarthrosis, and failed previous fusion. | Stryker TRIO/TRIO | 2 | August 22, 2012 | 2012 | Stryker Spine | The front cover of Instructions For Use (IFU) #NOLI135B03Rev02 was improperly marked with a "Sterile" logo in the lower right corner. The products these IFUs were packed with are sold as "Non-Sterile" and require sterilization before use in surgery. | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was attached for customers to complete and return via fax to 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this notice. |  | 40 | 40 units. | Distributed in Atlanta, Georgia. | Terminated | February 27, 2013 | 189 | 40 | Not_Computer
Z-2245-2012 | 62790 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 6, 10, 16, 40, 64 and CT Big Bore are Computed Tomography X-Ray Systems These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories.Product Usage:The Brilliance 6, 10, 16, 40, 64 and CT Big Bore are Computed Tomography X-Ray Systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories | Brilliance 6, 10, 16, 40, 64 and CT Big Bore | 2 | August 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the foot switch cover becomes damaged and impinges on the foot switch. | DESIGN: Device Design | Philips Medical Systems (Cleveland), Inc. sent an Urgent - Medical Device Correction letter dated August 1, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taked. The letter contained important information for the continued safe and proper use of your equipment. The letter stated that, a Field Service Engineer will visit each affected site to install the hardware upgrade. For further information or support concerning this issue contact your local Philips representative or local Philips Health office. For North America and Canada contact Customer Care Solution Center (1-800-722-9377), option 5: Enter Site ID or Follow the prompts). | Contact the recalling firm for information | 126 | 126 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, DC, FL, GA, KY, MD, MI, MN, MS, ND, NM, NY, OH, OK, SC, TN, TX, VT, WA, and WI and the countries of: Austria, China, Denmark, France, Germany, Iceland, Japan, Netherlands, South Africa, Sweden, Switzerland, United Kingdom. There were two government consignees. There were no Canadian consignees. | Open |  |  | 126 | Not_Computer
Z-2236-2012 | 62744 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC InstaTrak 3500, picture archiving and communications system intended as an aid to locate anatomical structures during open percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. | OEC InstaTrak 3500 | 2 | August 22, 2012 | 2012 | GE OEC Medical Systems, Inc | GE Healthcare had recalled certain OEC InstaTrak 3500 Carts due to the potential for the cart to tip over when the arm of the imaging device is extended during use. | DESIGN: Component Design/Selection | GE Healthcare sent an "URGENT SAFETY ADVISORY NOTICE" dated October 11, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter stated that GEHC OEC Field Service Engineers would be contacting their facility to place a caution label on the articulating monitor arm of the cart and to provide a revised operators manual that will contain additional information on the safe usage of the cart. Until a GEHC Engineer had taken the actions described, they advised their consignees to exercise caution when fully extending the articulating monitor arm on the device to avoid potentially tipping over the cart. Contact the firm at 800-874-7378 for questions regarding this notice. | Contact the recalling firm for information | 99 | 99 units | Nationwide Distribution. | Terminated | August 22, 2012 | 0 | 99 | Not_Computer
Z-2261-2012 | 62871 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Visual-ICEÂŽ Cryoablation System, Model FPRCH 6000Product Usage: The Galil Medical Visual-ICE Cryoablation System is intended for cryoablative destruction of tissue during minimally invasive procedures; various Galil Medical accessory products are required to perform these procedures. The Visual-ICE System is indicated for use as a cryosurgical tool in the fields of general surgery, dermatology, neurology (including cryoanalgesia), thoracic surgery, ENT, gynecology, oncology, proctology, and urology. This system is designed to destroy tissue (including prostate and kidney tissue, liver metastases, tumors, skin lesions, and warts) by the application of extremely cold temperatures. | VisualICE Cryoablation System | 2 | August 23, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling a number of Visual Ice Cryoablation Systems because the gas regulators may fail to regulate the gas pressure, rendering the devices inoperable. | DESIGN: Device Design | Galil Medical sent a "Voluntary Medical Device Field Correction" letter dated July 20, 2012, via UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Galil Field Service team will be contacting customers to schedule a service visit the week of July 23rd. Customers were instructed to forward and communicate this field correction notice and instructions to individuals in their organization and/or to other organizations affected by this action. For questions call Galil Medical Customer Service. | Contact the recalling firm for information | 4 | 4 | Worldwide Distribution - US nationwide including the state of MN and the countries of United Kingdom and Netherlands | Terminated | September 13, 2012 | 21 | 4 | Not_Computer
Z-2262-2012 | 62746 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC Injection Cable for MedRad Provis V Contrast Injector. In certain fluoroscope X-ray modes, the cable provides a signal to the auto-injector to trigger a contrast injection. Part numbers include 00-902260-01 and 00-883654-01. | GEHC OEC Injection Cable for MedRad Provis V Contrast Injector. | 2 | August 23, 2012 | 2012 | GE OEC Medical Systems, Inc | GE initiated a correction to the instructions/use due to the potential for erroneous auto injection of contrast dye. | DESIGN: Device Design | An Urgent Safety Advisory Notice was sent to consignees on October 10, 2008. The letter identified the affected product and notified customers of the situation, along with instructing them to discontinue use until proper validation and instruction update for cable use was completed. GEHC OEC was to have contacted customers when the solution became available and resolved it at no cost to the customer. |  | 149 | 149 units | Nationwide distribution. | Terminated | January 10, 2013 | 140 | 149 | Not_Computer
Z-2271-2012 | 62578 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Low Profile Abutment, Non-Hexed Castable Cylinder Model LPCCC1. Product is labeled in part: "***REF LPCCC1***Low profile Abutment Hexed Castable Cylinder***Cylinder; Cylindre; Zylinder; Cilindro; Cilindro; Cilindro***RX only***NON STERILE***Caution: Law prohibits dispensing without prescription***Consult instructions for use***Castable cylinders are intended to be used to produce a definitive casting for dental restoration onto which porcelain can then be cast. | Low Profile Abutment | 2 | August 24, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i conducted a recall on the device Low Profile Abutment, Non-Hexed Castable Cylinder due to some packages labeled LPCCC1, Hexed Castable Cylinder possibly containing LPCCC2, Non-Hexed Castable Cylinders. | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet 3i sent a "MEDICAL DEVICE ADVISORY NOTICE" dated May 31, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this recall. |  | 611 | 611 units | Worldwide Distribution-Distributed in the USA and the countries to Spain, France, Germany, Switzerland, and Australia. | Terminated | November 21, 2013 | 454 | 611 | Not_Computer
Z-2267-2012 | 62864 | disinfectant, medical devices | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviCide, Part Numbers: 13-1025, 13-1000, 13-1024, MC-1000.Alternate brands:EnviroCide, part number: 13-3325.Private labeled products of Cavicide:Backscratchers Cavicide, part number: 13-4800.MaxiSpray Plus, part numbers: 13-7400 and 13-7405.pdCARE Surface Disinfectant, part numbers: 13-7525 and 13-7500.Z3 Surface Disinfectant, part number: 13-7900.CaviCide is a multi-purpose disinfectant/decontaminant cleaner for use on hard, non-porous inanimate surfaces. CaviCide is a bactericidal, virucidal, fungicidal, and tuberulocidal which is available in a ready-to-use, liquid form typically contained in 1 gallon bottles, 2.5 gallon bottles, 25 oz. spray bottles, or 55 gallon drums (depending on the item number). CaviCide may also be branded as a private label (see (1)(vi) below for names and Table 1 above for additional details). The product may be used (1) for precleaning non-instrument surfaces prior to disinfection; (2) for precleaning instruments prior to disinfection; (3) as a virucide on non-instrument surfaces; (4) as a disinfectant on non-instrument surfaces; (5) as a virucide on precleaned non-critical medical devices, instruments and implements; and (6) as a disinfectant on precleaned non-critical medical devices, instruments, and instruments. It is useful in health care settings such as hospitals, emergency medical settings, surgical centers, isolation areas, neonatal units, patient care areas, ophthalmic and optometric facilities, dental offices, dental operatories, animal care facilities, salon settings, emergency vehicles, schools, health club facilities, kitchens, bathrooms, prisons, morgues, and other critical care areas. | CaviCide | 2 | August 24, 2012 | 2012 | Metrex Research, LLC. | Metrex Research is recalling Cavicide containing the Spring Fresh Fragrance because the affected lot of product containing this fragance has caused an increase in reported health reactions in some individuals such as headache, nausea, and respiratory discomfort. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Metrex sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Mandatory Recall Response Form was attached for customers to complete and return via fax to 734-947-6725.Contact Customer Care Center at (800) 841-1428 for questions regarding this recall. |  | 17690 | 17,690 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, and YE. | Terminated | August 24, 2012 | 0 | 17690 | Not_Computer
Z-2265-2012 | 62828 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Canon Solid State X-ray Imager (Flat Panel/Digital Imager), Digital Radiography, Models CXDI-55C, CXDI-55G and CXDI-70C. These devices are intended to replace radiographic film/screen systems in all general purpose diagnostic procedures. | Digital Radiography, CXDI | 2 | August 24, 2012 | 2012 | Canon Inc. | Screws holding the 55REF circuit board are loosening and roaming inside of the sealed sensor. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Canon Inc. notified sales companies including Canon USA on 9/15/2010 with Service Bulletin 10-017, and on 8/25/2011 with Service Bulletin 11-011, and on 1/20/2011 with Service Bulletin 11-001. Sales representatives were to visit customers and check and replace the screws. |  | 468 | 468 units | Nationwide Distribution (USA) | Terminated | January 07, 2014 | 501 | 468 | Not_Computer
Z-2273-2012 | 51437 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TRESTLE LUXEÂŽ Anterior Cervical Plating System, Product code KWQ. 510(k) K102820.The TRESTLE LUXE Anterior Cervical Plating System is a temporary device used to stabilize the cervical spine during bone fusion development. Device implants include a range of plate sizes and bone screws to provide the versatility required for the specific indications noted. Fixation is achieved by means of a rigid plate that is surgically attached to the spine with bone screws. The TRESTLE LUXE Anterior Cervical Plating System is intended for use in the anterior cervical spine (C2-C7). Patients with trauma (including fractures), spondylolisthesis, and pseudoarthrosis, degenerative disc disease (defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies), spinal stenosis, tumors, and failed previous fusion. | TRESTLE LUXE Anterior Cervical Plating System | 2 | August 27, 2012 | 2012 | Alphatec Spine, Inc. | On May 9th 2012 Alphatec Spine received a complaint which reported that the distal tipof the Trestle Luxe Slide Alignment Tool (Part Number 71721), does not properly fit intothe hexalobe screw on the Trestle Luxe Plate. | DESIGN: Device Design | Alphatec Spine sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A form was attached for customers to complete and return via fax to 760-431-0289. Contact Alphatec Spine's Customer Service Department at customerservice@alphatecspine.comfor questions regarding this recall.). |  | 38 | 38 | Nationwide Distribution. | Terminated | September 04, 2012 | 8 | 38 | Not_Computer
Z-2274-2012 | 62564 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Apelo Modular Screwdriver labeled in part: "APELO PEDICLE SCREW SYSTEM***Modular Screwdriver***Atlas Spine, Inc.***www.AtlasSpine.com*** 1555 Jupiter Park Drive #4***Jupiter, Florida 33458***561-741-1108***Tissue Sleeve***MFG Date: XX/XX/XXXX***Lot #: XXXX***Weldment***MFG Date: 03/23/2009***Sleeve Assembly***MFG Date: XX/XX/XXXX***Lot #: XXXX***Qty: 1 Rx Only***Non-Sterile. Sterilize Prior to Use*** Rev. A***. The Modular Screwdriver is used to insert modular pedicle screw implants which are part of the Apelo System. | Apelo Modular Screwdriver for Apelo Pedicle Screw System | 2 | August 27, 2012 | 2012 | Atlas Spine, Inc. | Atlas Spine, Inc. recalled their Apelo Modular Screwdriver, a component of the Apelo Pedicle Screw System Instrumentation, because of complaints received about the tip breaking. If the tip is fractured inside of the modular screw, the tip could become implanted in the body or if the entire screw is removed, it could cause a delay in the surgery. | DESIGN: Component Design/Selection | Consignees were notified by Telephone and visit. All products will be returned, quarantined and then destroyed. |  | 4 | 4 pieces | USA Distribution to the state of Florida only. | Open |  |  | 4 | Not_Computer
Z-2276-2012 | 62919 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX OPA, High Level Disinfectant Solution, Glutaraldehyde-free (0.55% ortho-phthaladehyde) solution. The CIDEX OPA solution is intended for automated and manual cleaning and disinfection. | CIDEX OPA | 2 | August 27, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products has discovered instances of missing labels on bottles of CIDEX OPA Solution, Part # 20390 manufactured between May 2011 and April 2012. | PRODUCTION CONTROLS: Process Control | An "Urgent Medical Device Recall" letter dated August 10, 2012 was sent to all customers who purchased the CIDEX OPA solution, part #20390 (manufactured between May 2011 and April 2012). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to notify their customers if they have distributed bottles of CIDEX OPA Solution to other locations. Customers were instructed to call Stericycle at (866) 629-6182 to arrange for return of the product. |  | 1310844 | 1,310,844 units | Nationwide (USA) and Puerto Rico. | Terminated | September 11, 2013 | 380 | 1310844 | Not_Computer
Z-2277-2012 | 62780 | catheter, hemodialysis, implanted | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 16 Fr HemoStar XK Long-Term Hemodialysis Catheter, implanted Standard kit containing an AirGuard introducer. The device is used for patients requiring hemodialysis.The HemoStar XK catheter is made of radio-opaque polyurethane and allows for flow rates as high as 500 ml/min during hemodialysis treatment. | HemoStar XK LongTerm Hemodialysis Catheter | 3 | August 27, 2012 | 2012 | Bard Access Systems | Bard Access Systems is recalling one lot of 16 Fr HemoStar XK Long Term Hemodialysis Catheters because they were identified with an incorrect expiration date, 03/2015 rather than 10/2013. | EXPIRATION DATING: Incorrect or No Expiration Date | Bard Access Systems sent an "URGENT: *16 FR HEMOSTAR* LONG-TERM HEMODIALYSIS CATHETER RECALL" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form and Inventory Reconciliation Form was enclosed for customers to complete and return via fax to 1-801-522-5674. Contact BAS Customer Service at 1-800-290-1689 for questions regarding this recall. |  | 120 | 120 units | National US distribution only: MO, KS, AZ, OK, OR, AL, MA, TX, FL, MI, OH, LA, NJ, NC, WA, AR. | Terminated | January 07, 2013 | 133 | 120 | Not_Computer
Z-2293-2012 | 61982 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Product is labeled in part: "***MEDLINE***Extended PTFE Coated Blade with Extended Insulation 4.0 (10.16cm) Length***Reorder: ES0014AM***Contents: 1 Each***SINGLE USE ONLY**latexfree***RX Only***STERILE EO***www.medline.com Manufactured for Medline Industries, Inc. MEDLINE is a registered trademark of Medline Industries, Inc. 1-800-MEDLINE***Mundelein, IL 60060 USA Made in USA.***" Part # ES0014AM -- Lot # 050409-05, 050609-05.Product Usage: Intended for cutting and coagulation of soft tissue. | Mediline Industries, Inc. | 2 | August 28, 2012 | 2012 | Unimed Surgical Products, Inc. | On 6/16/2010 Unimed Surgical Products, Inc., 10401 Belcher Road, Largo, FL 33777 initiated a recall of the Medline Industries, Inc. Electrosurgical electrodes used for cutting and coagulation of soft tissue due to concerns of sterility. | PRODUCTION CONTROLS: Process Control | Unimed Surgical Products, Inc. issued a Medical Device Removal letter dated June 4, 2010 to Medline Industries, Inc. The letter identified the affected product, problem and actions to be taken. Medline was instructed to quarantine any undistributed product and notify its customers of the recall. For further information, contact Unimed Surgical Product, Inc. by phone 800-531-2627 or FAX 800-886-1225. | No consumer action necessary | 2400 | 2400 each | US Nationwide Distribution - including the state of IL. | Terminated | August 30, 2012 | 2 | 25440 | Not_Computer
Z-2281-2012 | 61171 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | The Eon Mini Neurostimulation (IPG) System (Model 3788) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with 4, 8, and 16 electrode leads.Product Usage:The Eon and Eon Mini Neurostimulation Systems are indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. | Eon Mini | 2 | August 28, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | As of 11/30/2011, St. Jude Medical has received 110 patient complaints of warmth or heating at the implantable pulse generators (IPGs) implant site during charging for the Eon IPG and 116 reports of similar symptoms for the Eon Mini IPG, respectively. | DESIGN: Labeling Design | ST. Jude Medical sent an Important Medical Device Correction letter dated December 19, 2011 to all affected customers. The letter identified the affected product, problem and recommendations to be followed. For questions contact your St. Jude Medical Neuromodulation Division Representative. | No consumer action necessary | 30560 | 30,560 units | Nationwide within the US, and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Isreal, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Open |  |  | 56948 | Not_Computer
Z-2251-2012 | 62715 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC IT3000; IT2500 (EnTrak); IT2500 Plus (EnTrak Plus); IT3500 (InstaTrak); IT3500 Plus (InstaTrak Plus); ConneCTstat; and ConneCTstat Plus picture archiving and communications systems. The systems are an aid to locate anatomical structures during open or percutaneous surgical procedures. It is intended for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. | OEC IT3000; IT2500 (EnTrak); IT2500 Plus (EnTrak Plus); IT3500 (InstaTrak); IT3500 Plus (InstaTrak P | 1 | August 28, 2012 | 2012 | GE OEC Medical Systems, Inc | GE Healthcare Surgery had discovered that using the Inverted Headset Placement is not a validated configuration with the IT 3000, 2500, 2500 plus, 3500, 3500 plus, ConneCTstat, and ConneCTstat Plus Surgical Navigation equipment. | MISBRANDING: Labeling False and Misleading | GE Healthcare sent an Urgent Safety Advisory Notice dated June 22, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised not to use the ENT registration Headset in the inverted position for any CT scans or surgical procedures and to remove all references in their Naviation System Operator's Manual that discuss an inverted headset placement, and replace it with the enclosed updated instructions for the Navigation System Operator's Manual. For any questions customers should contact the service team for further information at 800-874-7378, option #8.For questions regarding this recall call 801-536-4516. |  | 768 | 768 units | Worldwide Distribution - USA (nationwide) and the countries of India, Japan, Canada, Switzerland, Italy, United Kingdom, Germany, Austria, Hong Kong, France, and Denmark. | Terminated | September 28, 2012 | 31 | 768 | Not_Computer
Z-2283-2012 | 62712 | fibrinogen and fibrin split products, antigen, antiserum, control | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Stratus(R) CS Acute Care(TM) D-dimer Calibrator (DDM CalPak) lots 502058002 and 501318002The Stratus(R) CS Acute Care(TM) D-dimer Calibrator (DDMR CalPak), Catalog No. CDDMR-C is an in-vitro diagnostic product intended to be used for the calibration of the Stratus(R) CS Acute Care(TM) D-dimer )DDMR) method | Siemens Stratus(R) CS Acute Care(TM) Ddimer Calibrator (DDM CalPak) | 3 | August 28, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Stratus(R) CS Acute Care(TM) D-Dimer CalPak calibrator lots 502058002 and 501318002 may exhibit SLOPE or CALCV error messages (instrument errors), which indicate calibration failures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | SIEMENS sent an Urgent Field Safety Notice dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included a response form to be faxed to the recalling firm at 302-631-8467 by the consignee. If customers experienced calibration failures, they were instructed to contact the Siemens Technical Solutions Center at 800-405-6473 for assistance. Customers were also asked to forward the notification to anyone to whom they may have distributed the affected product. For questions regarding this recall call 302-631-6311. |  | 786 | 786 cartons | Nationwide Distribution - USA including AK, AL, AR, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, WA, WI, WV, and WY. | Terminated | January 16, 2014 | 506 | 786 | Not_Computer
Z-2253-2012 | 62769 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | ON-Q pump with ONDEMAND bolus button, ON-Q C-bloc with ONDEMAND; ON-Q C-bloc with Select a Flow and ONDEMAND Model/Reference #: 7000180, 7000181, 7000182, 7000367, 7000397, CB003, CB005, CB006, P400X1-7+5/60, P400X2-14+5/30, P400X5+5/60, and PMB01. Part Numbers: 101347100, 101347300, 101347400, 101360200, 101368400, 101368600, 101368700, 102987000, 102987200, 102987300, 103161800, 103534300, 5001438, 5001470, 5001651, 5001780, 5001808, 5001809, 5001812, 7000180, 7000181, 7000182, & 7000367.Indicated for continuous and or intermittent delivery of medication to or around surgical wound sites and or close proximity to nerves. | ONQ pump with ONDEMAND bolus button | 1 | August 28, 2012 | 2012 | I-Flow Corporation | It was determined that in a small quantity of ON-Q pumps with ONDEMAND, the bolus button may not lock in the down position when depressed and/or the orange bolus refill indicator may stay in the lowermost position. As a result, the patient may receive continuous infusion at a rate greater than expected. | DESIGN: Device Design | An Important Voluntary Recall Notice letter dated 5/8/12 was sent to all their customers who purchased the ON-Q pump with ONDEMAND bolus button. The letter provided customers with the problem indicated and the actions to be taken. Customers were instructed to complete and sign the attached Recall Response Form and fax to (920)969-4565 or scan and email to IflowProductInquiry@kcc.com.Once I-Flow receives the completed FAX BACK form, an I-Flow Customer Service Representative will immediately contact the customers to provide instructions on product return and credit. Customers with questions regarding the recall letter, were instructed to contact I-Flow Customer Service between 8am to 5pm (PST) at (800) 443-3569. Or customers can contact I-Flow 24 hour Technical Support line at (800) 444-2728. |  | 80646 | 80,646 units total (65,926 units in US) | Worldwide Distribution -- USA, Canada, Australia, the EU (Germany), United Arab Emirates, and Cyprus. | Terminated | April 25, 2013 | 240 | 80646 | Not_Computer
Z-2285-2012 | 62572 | system, test, anticardiolipin immunological | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | PHADIA Varelisa Cardiolipin IgMCardiolipin autoantibody immunological test system (IgM class antibodies) | PHADIA VarelisA Cardiolipin IgM Antibodies Assay | 3 | August 28, 2012 | 2012 | Phadia US Inc | Customers were not following the labeled assay procedure. This was due to a letter issued to them in 2004 from Phadia GmbH, which indicated that customers could omit the pre-wash step. The omission of the pre-wash step, when combined with a change in material source in kit lot #0067, caused elevated results. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | On 8/2/2012 an URGENT PRODUCT CORRECTION INFORMATION FOR PHADIA VARELISA CARDIOLIPIN IGM, Article Number 15696 letter was sent to all consignees the received the Direction for Use modification letter in 2004. | Contact the recalling firm for information | 837 | 837 kits (96 assays per kit), total of 80,352 individual assays | USA Nationwide. | Terminated | December 12, 2012 | 106 | 837 | Not_Computer
Z-2284-2012 | 62776 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | QKC1692E The QKC1692E Quick Connect device connects endoscopes to the sterilization tray for C1220E liquid chemical sterilization processor. | QKC1692E | 2 | August 28, 2012 | 2012 | Steris Corporation | STERIS has been actively seeking devices for quick connect revalidation projects. STERIS was unable to obtain two (2) of the needed devices; these two devices represent the design of seven (7) models in one quick connect. Due to the unavailability of these two validation devices, STERIS will remove 7 models from the labeling of QKC1692E and will re-label all QKC1692Es in the field.. | DESIGN: Device Design | The firm, Steris, sent a "Voluntary Field Correction: Re-labeling of QKC1692E" letter on August 10, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter to re-label the affected QKC1692E(s) and return the completed form by September 30. 2012 to: Regulatory Affairs Specialist via Kate_Beveridge@STERIS.com or via fax at (440) 350-7082.If you have any questions, contact Director, Regulatory Affairs at 440-392-7519. |  | 64 | 64 units | Nationwide distribution: USA including states of: AR, AK, CA, CO, DE, FL, IL, MA, MD, NC, NE, NH, NY, OH, PA, SD, TN, TX, WI, and WV. | Terminated | July 05, 2013 | 311 | 64 | Not_Computer
Z-2300-2012 | 49060 | kit, plastic surgery and accessories | FTN | Plastic surgery kit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Various SRI Surgical Disposable Accessory Packs.SRI's Disposable Accessory Packs (also commonly known as convenience kits and custom procedure trays) are an assembly of medical devices packaged together according to the customer specifications, which are then wrapped in a sterilization wrap and sterilized. This allows the customer to open one package for a surgical procedure instead of many individual packages. | Various SRI Surgical Disposable Accessory Packs. | 2 | August 29, 2012 | 2012 | SRI Surgical | Distribution of a product that did not meet specifications. (labeling error) | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | SRI sent a "MEDICAL DEVICE CORRECTION" letter dated July 16, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 813-891-9550 ext. 3124 for questions regarding this recall. |  | 802135 | 802,135 | Nationwide Distribution-including the states of CA, FL, GA, KY, MD, MI, NC, NJ, OK, PA, SC, TN, TX, and UT. | Terminated | September 10, 2012 | 12 | 802135 | Not_Computer
Z-2301-2012 | 62705 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DGPHP RFA High-Power Single Use Grounding Pads and Cool-Tip RFA Electrode Kits (containing DGPHP RFA High-Power single use grounding pads)The product is used during ablation surgical procedures. | Cooltip RFA Electrode Kits | 2 | August 29, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a recall of various production lots of DGHP RFA High Power single use grounding pads and Cool-tip RFA Electrode Kits due to the potential for foil degradation on the DGP-HP single use grounding pad which may result in an electrical burn to the patient at the application site. | PRODUCTION CONTROLS: Process Control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 6, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine and discontinue use of the device; immediately advise all DGPHP RFA High-Power single use grounding pads and Cool-tip RF Electrode kit users of this recall; complete and return the attached Recalled Product Return Form in its entirety via fax to (800) 895-6140; return affected product to Covidien Attn: Field Returns Department, 195 McDermott Road, North Haven, CT 06473; if purchased directly from Covidien include completed form via fax or by email to : feedback.customerservice@covidien.com; if purchased from distributor complete form and contact your Distributor directly (completed form and all affected units must be returned through the Distributor); and reply to Covidien WHETHER OR NOT you have affect product at your site. If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance, Monday through Friday, 8:30-5:00, directly at 203-492-5232. |  | 184934 | 184,934 units | Worldwide distribution: USA (nationwide) and countries including: Australia, Brazil, Canada, Columbia, Deutchland, France, Japan, Mexico, Panama, Puerto Rico, Singapore, and Uruguay. | Terminated | September 06, 2013 | 373 | 184934 | Not_Computer
Z-2303-2012 | 62865 | system, measurement, blood-pressure, non-invasive | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VSi/VS2+ patient monitors ModelPart NumbersVSi/ 863275, 863276, 863277, VS2+ 863278, 863279.The SureSigns VS2+ vital signs monitor is for monitoring, recording and alarming of multiple physiological parameters of adults, pediatrics, and neonates in healthcare environments. Additionally, the monitor is intended for use in transport situations within a healthcare facility. The SureSigns VSi vital signs monitor is for measurement of multiple physiological parameters of adults, pediatrics, and neonates in healthcare environments. Additionally, the monitor is intended for use in transport situations within a healthcare facility | Philips Healthcare SureSigns VSi/VS2 | 2 | August 30, 2012 | 2012 | Philips Healthcare Inc. | Philips SureSigns VSi/VS2+ patient monitors may not be properly secured and may fall from the wall mount | DESIGN: Component Design/Selection | Philips sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers using the VSi/VS2+ with the GCX wall mount are instructed to complete and fax a Customer Reply Wall Mount Order Card to the Business Unit or Key Market indicating the number of wall mounts in use. Upon receipt of the Customer Reply Wall Mount Order Card, Philips will send a replacement of the mounting hardware. Customers should contact their local Philips representative for questions regarding this recall. |  | 2314 | 2314 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, DENMARK, EL SALVADOR, FRANCE, GERMANY, ICELAND, INDIA, IRELAND ISRAEL, ITALY, JORDAN, KENYA, LEBANON, MALAYSIA,NAMIBIA NETHERLANDS, NORWAY, POLAND, PORTUGAL, RWANDA SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open |  |  | 2314 | Not_Computer
Z-2302-2012 | 62801 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer + additive/metal/polymer + additive | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert - HXLPE-VE, Sterile, Material: Highly cross linked vitamin E PE, Sz 2/9mm.The insert is intended to more closely match the kinematics of the knee, allowing some rotation along the medial condyle and increased congruency along lateral condyle. This insert is to be used along with the other components that comprise the 3DKnee System. | 3DKnee e Tibial Insert | 2 | August 30, 2012 | 2012 | Encore Medical, Lp | The firm is recalling one lot of 392-09-706, 3DKnee e+ Tibial insert, 9 mm, Size 6, Right. The shelf life labeling of specific lots for a component were found to be incorrect and were found to be expired. | EXPIRATION DATING: Incorrect or No Expiration Date | DJO Global sent an Urgent Field Safety Notice dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Notice to all those who need to be aware within their organization or to any organization where the potentially affected product has been transferred. Customers were instructed to contact the firm to place a replacement order and receive an RMA number. For questions customers were asked to call 512-834-6255.For questions regarding this recall call 512-832-9500. | No consumer action necessary | 7 | 7 units | Worldwide Distribution including Venezuela. | Terminated | January 25, 2013 | 148 | 7 | Not_Computer
Z-2306-2012 | 62537 | antigens, cf (including cf control), influenza virus a, b, c | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Orasure QuickFlu Rapid A+B TestManufactured for OraSure Technologies, Inc Bethlehem, PA 18015 USA.The BioSign Flu A + B test is an in vitro rapid qualitative test that detects influenza type A and type B nucleoprotein antigens directly from nasal swab, nasopharyngeal swab and nasal pharyngeal aspirate/wash specimens obtained from patients with signs and symptoms of respiratory infection. It is intended to aid in the rapid differential diagnosis of influenza A and B viral infections. A negative test result is presumptive and it recommended that these results be confirmed by viral culture. Negative results do not preclude influenza virus infection and should not be used as the sole basis for treatment or other management decisions. The test is intended for professional and laboratory use. | Orasure QuickFlu Rapid AB Test | 2 | August 31, 2012 | 2012 | Princeton Biomeditech Corp | There have been reports of dual positive (Flu A + Flu B positive) results being encountered. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | OraSure Technologies, Inc. sent an "IMPORTANT NOTICE FOR ORASURE QUICKFLU RAPID FLU A + B TEST" letter dated March 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 610-882-1820 for questions concerning this recall. | Contact the recalling firm for information | 370 | 370 (12 tests per kit) ;2800 kits (22 tests per kit) | Nationwide Distribution-including the states of NJ, NY, and PA. | Open |  |  | 563 | Not_Computer
Z-2260-2012 | 62848 | ventilator, continuous, minimal ventilatory support,facility use | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V60 Ventilator, Continuous Ventilator, Model# V60, PN 1056490Product Usage:The V60 Ventilator is an assist Ventilator and is intended to augment patient breathing. It is intended for spontaneously breathing individuals who require mechanical ventilation: patients with respiratory failure, respiratory insufficiency, or obstructive sleep apnea in a hospital or other institutional settings under the direction of a physician. The V60 Ventilator is intended to support pediatric patients weighing 20kg (44 Ibs) or greater to adult patients. It is also intended for intubated patients meeting the same selection criteria as the noninvasive applications. It is intended to be used by qualified medical professionals, such as physicians, nurses, and respiratory therapists. | V60 Ventilator | 1 | August 31, 2012 | 2012 | Respironics California Inc | Respironics is recalling the V60 ventilator because certain blower motor assemblies in the V60 ventilators may not meet specifications. Specifically, the impeller may not be properly seated on the blower motor shaft. This could subsequently cause the blower to cease functioning which would result in the failure of the V60 ventilator to deliver therapy to the patient due to the loss of ventilatio | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Respironics California a division of Philips Healthcare sent a Medical Device Correction letter dated July 31, 2012 to all affected customers. The letter identified the affected product, problem identified and actions to be taken. The letter states that a Phillips Respironics Field Service Engineer or distributor will be contacting customers to schedule a no-cost replacement of the blower motor assembly on all affected V60 ventilators. For information or support concerning the recall issue, contact Philips Respironics at (800) 722-9377. |  | 197 | 197 units total (116 in the US) | Worldwide Distribution - US (Nationwide) and the countries of: Australia, Belgium, Canada, China, Germany, Italy, Japan, Kuwait, Netherlands, Norway. Qatar, Saudi Arabia, Turkey, and United Kingdom. | Terminated | March 22, 2013 | 203 | 197 | Not_Computer
Z-2309-2012 | 61451 | stent, iliac | NIO |  |  | Cardiovascular |  | Zilver 518 Vascular Self Expanding StentThe Zilver Vascular Stent is intended for use as an adjunct to percutaneous transluminal angioplasty (PTA) in the treatment of symptomatic disease of iliac arteries up to 100 mm in length with a reference vessel diameter of 5 to 9 mm. Patients should be suitable candidates for PTA and / or stent treatment. | Zilver 518 Vascular Self Expanding Stent | 2 | September 01, 2012 | 2012 | Cook Ireland Ltd. | Cook Ireland received a product complaint from Japan on December 22nd, 2011. The complainant reported that a 5 French Zilver Vascular Stent (part number ZIV5-18-125-8-80 from Lot# CF694429) was inserted into a sheath over a wire guide where it became lodged in the sheath hub. After further examination, the user determined that the product was a 6 French Zilver Vascular Stent (part number ZIV6-35 | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Cook Medical, sent a letter dated February 13, 2012 to foreign consignee. The letter described the product, problem and actions taken. The firm stated that re-training of the relevant personnel has been carried out; they will continue to monitor to identify potential emerging trends and further corrective and preventive action will be considered in an effort to avoid the recurrence of similar events.If you have any question contact in US: Customer Relations Manager at 800.346.2686.ext 2222 or email: rita.harden@cookmedical.com; OUS: Customer Quality Supervisor at +353 61 239343. | Contact the recalling firm for information | 15 | 15 units of ZIV5-18-125-8-80 LOT# CF694429 | Worldwide distribution: USA (nationwide) and country of: Japan. | Terminated | February 04, 2013 | 156 | 15 | Not_Computer
Z-2310-2012 | 61984 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC TIBIAL INSERT POSTERIOR STABILIZED***Match Component Colors 6 YELLOW, 11mm, Use with LOGIC Femoral and Tibial Components***REF 02-012-35-6011.There are three product sizes being recalled, 11mm, 13mm, 15mm. All other information is the same.The Optetrak Logic Total Knee System is indicated for use in skeletally mature individuals undergoing primary surgery for total knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis, and/or post-traumatic degenerative problems. They are also indicated for revision of failed previous reconstructions where sufficient bone stock and soft tissue integrity are present. | OPTETRAK LOGIC TIBIAL INSERT POSTERIOR STABILIZED | 2 | September 04, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL is recalling their Optetrak Logic Posterior Stabilized Tibial Insert, Size 6 after the device was determined to be out of dimensional specification. | PRODUCTION CONTROLS: Process Control | Exactech, Inc., sent an "IMPORTANT PRODUCT MARKET WITHDRAWAL NOTICE" letter dated August 19, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that are in possession of the products. 3. Verify if they had any of the affected products on the list. 4. Fax back the attached form. In addition, they were instructed to contact the Exactech inventory representative to confirm quantities at their location. If you have any questions regarding inventory restocking issue, please call -800-392-2832. |  | 193 | 193 | USA ( nationwide) WA, OH, OK, TX, NC, NY, CO, VA, MA, IL, TN, MN, NJ and FL. | Terminated | November 07, 2013 | 429 | 193 | Not_Computer
Z-2311-2012 | 62593 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Acrobat 3000 (AC3000) Spring Arm as support arm for iLED operating lights, TRUMPF Medizinsysteme GmbH + Co KG 82178 Puchheim.The TRUMPF surgical lights are intended to locally illuminate an operating or examination area of the patients body with a high intensity light. | Acrobat 3000 (AC3000) Spring Arm | 2 | September 05, 2012 | 2012 | Trumpf Medical Systems, Inc. | Cracks can arise on the welded seam on the rear joint. | PRODUCTION CONTROLS: Process Control | TRUMPF sent an Urgent Safety Information letter dated February 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that within their organization, all users of the affected product be informed about the Urgent Safety Information. If affected product was further distributed, customers should forward a copy of the information or inform the necessary parties. Customers were asked to confirm receipt of the Urgent Safety Information by completing and returning the confirmation of receipt.For questions regarding this recall call 888-474-9359. | No consumer action necessary | 402 | 402 units | Worldwide Distribution - USA including CA, OH, OK, NY, PA, TX, WY and the country of Canada | Terminated | April 11, 2013 | 218 | 402 | Not_Computer
Z-2315-2012 | 62928 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Toshiba Ultra-Short Magnetic Resonance Imaging System, EXCELART Vantage. MRT-1503/P5-Vantage MRI System.MR imaging. | Toshiba UltraShort Magnetic Resonance Imaging System | 2 | September 06, 2012 | 2012 | Toshiba American Medical Systems Inc | The firm recalled due to a potentail problem with the position of the hold in the lug teminals on the power cables connected to the terminal blocks on the gradient coil. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Toshiba America Medical Systems, Inc, sent a Urgent Medical Device Correction letter dated February 21, 2012, via USPS return receipt mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed if you subsequently find that images are abnormally noisy compared to images acquired under the same conditions in the system before this modification was performed, contact your Toshiba Service Representative.Please share this information with all users and reviewing radiologist as well as clinical engineering or biomedical group at your facility. Please complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email toraffairs@tams.com.Contact information was also provided (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US (nationwide) distribution including the states of Ohio and Texas. | Terminated | September 07, 2012 | 1 | 3 | Not_Computer
Z-2316-2012 | 62505 | reagents, clostridium difficile toxin | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm MRSA ACP Assay Catalog #441637, box 48 tests labeled in part***GeneOhm Sciences Canada, Inc., 2555 boul, du Parc-Technologique, Quebec, QC, Canada G1P 4S5***BD GeneOhm MRSA ACP Assay is a qualitative in vitro diagnostic test for the direct detection of methicillin-resistant Staphylococcus aureus (MRSA) DNA from nasal swabs in patients at risk for nasal colonization. Test is intended to aid in the prevention and control of MRSA infections in healthcare settings. It is not intended to diagnose MRSA infections nor to guide or monitor treatment for MRSA infections. | BD GeneOhm MRSA ACP Assay | 2 | September 06, 2012 | 2012 | Becton Dickinson & Co. | Leakage in Cepheid SmartCycler reaction tubes. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS overnight delivery on May 22, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Discontinue use and discard any affected product. BD would send replacement tubes for the discarded material.Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification.Fax the form to:Attention: Regulatory ComplianceFAX: 410-316-4258Email: RegulatoryComplianceFax@bd.comIf further assistance is needed, customers should contact BD Technical Services Department at 1-800-838-8663.For questions regarding this recall call 410-316-4054. |  | 38 | 38 kits | Nationwide Distribution including CA, DE, FL, GA, IL, IN, KS, KY, ME, MT, NC, NJ, NM, NY, OH, OK, PA, SC, TX, VA, WA, WI, and WV. | Terminated | November 02, 2012 | 57 | 200 | Not_Computer
Z-2322-2012 | 63042 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira.Mobile X-Ray System. | Siemens Mobilett Mira | 2 | September 06, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens has become aware of a potential issue with the Mobilett Mira system. During normal clinical operation the switch on the motor drive handle of the unit may become loose and slide horizontally to the right and become trapped under the cover. This may cause the system to continuously move. In response, Siemens has issued Update Instruction XP038/12/S to r | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens initiated their correction of the MOBILETT Mira on August 21, 2012 by sending an "UPDATE INSTRUCTIONS XP038/12/S to all affected customers. The Instruction identifies the product, problem, and work steps for customers to follow. A XP038/12/S Completion Protocol letter was included for customers to complete and file with the instruction once the update is completed. |  | 20 | 20 | Nationwide Distribution- including the states of GA, ID, IL, IA, KS, NC, OH, SD, TX, and WI. | Open |  |  | 20 | Not_Computer
Z-2326-2012 | 58247 | tape and bandage, adhesive | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General & Plastic Surgery | 510(K) Exempt | Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg, PN 10630Used with cold therapy pads to provide insulation barrier between the pad and patient's skin. | INTELLIFLO DRESSING | 2 | September 06, 2012 | 2012 | Breg Inc, An Orthofix Company | This recall has been initiated due to confirmation that the Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg does not meet the required thickness as specified. The dressing in the sterile pouch measures a thickness of 0.010" instead of 0.020". This change in thickness makes the dressingthinner, less durable and subject to tearing. The product is mislabeled. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Breg Orthofix, called their customers on August 11, 2010 sent a letter dated August 12, 2010 and emails to all customers. The call and letter described the product, problem and actions to be taken. The customers were instructed to take immediate action to identify and quarantine the product at their facility. Breg established a return materials authorization and is contacting their customers to provide a Return Authorization to return the dressings. Breg is also providing replacement dressings. If you have any questions regarding communication or the Intelli-Flo Sterile Dressing, contact Breg Customer Relations at (800) 321-0607. |  | 502 | 502 units | Worldwide distribution: USA (nationwide) and country of: Nova Scotia. | Terminated | September 06, 2012 | 0 | 502 | Not_Computer
Z-2325-2012 | 62852 | indicator, sterilization | LRT | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 100 ft roll of tubing inside a cardboard box. Medical devices are sealed into the tubes prior to ethylene oxide sterilization. Tube maintains sterility of the enclosed medical devices until opened. | VisUAll | 3 | September 06, 2012 | 2012 | Steris Corporation | On 7/20/2012 the firm became aware that a lot of tubing was distributed by the supplier after it was rejected by Steris due to failing process indicators. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Steris sent a Urgent Voluntary Recall Notice letter dated August 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Steris has learned that the VIS-U-All tubing is not performing to product specifications. Upon receipt of replacement product we ask that you destroy any remaining inventory of the affected lot in your possession. In the interim period, should you receive a failing result following a sterilization cycle, the results must be handled in accordance with your internal hospital procedures.We apologize for any inconvenience this matter may cause you. For further questions please call 1-(800) 548-4873. |  | 77 | 77 packs | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, IL, IN, KY, LA, MO, MS, NJ, OH, PA, SD, TX, and VA., and the countries of Japan, Peru, and Suriname. | Terminated | September 04, 2013 | 363 | 77 | Not_Computer
Z-2338-2012 | 62925 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassette, Batch # 12A078.The STERRAD 100S and STERRAD NX sterilizations systems are low-temperature general purpose sterilizers used to sterilize heat and moisture sensitive reusable medical devices. The sterilant used is hydrogen peroxide supplied in cassette form as a separate accessory. The STERRAD Cassettes are used in conjunction with the STERRAD Sterilizers, and each cassette model is for use only in the matching sterilizer model. | STERRAD NX Cassette | 3 | September 07, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling (removing) the Sterrad 100S and Sterrad NX System cassettes because the camera used to verify printing on the Sterrad System cassettes was inadvertently disabled. The STERRAD 100S and STERRAD NX System cases were shipped without verification of proper printing. | PRODUCTION CONTROLS: Error in Labeling | Advanced Sterilization Products sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated July 16, 2012 to all their customers who purchased the STERRAD 100S and STERRAD NX System cassettes. The letter informed the customers of the problem identified and the actions to be taken. For the STERRAD NX System cassettes, there should be no impact to the functionality of these affected cassettes. Although printing may be smeared, illegible, or missing, the packaging has multiple references to the information, and the cassettes should be usable despite any printing error.For the STERRAD 100S System cassettes, there is a chance that the barcode will be unreadable and the STERRAD 100S Sterilizer will not accept the cassette.Customers were instructed to complete and return the business reply card and packing slip that are included in the letter. Customers were instructed to contact Stericycle at (888) 406-9303 for additional labels. Customers with questions about the recall letter were instructed to contact Stericycle at (888) 406-9303. |  | 7550 | 7,550 units | Nationwide Distribution. | Terminated | November 08, 2013 | 427 | 15045 | Not_Computer
Z-2335-2012 | 62895 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | The SYSTEM 1E Liquid Sterilant Processing System.The SYSTEM 1E Liquid Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned immersible and reusable critical and semi-critical heat sensitive medical devices, including endoscopes and their accessories. | The SYSTEM 1E Liquid Sterilant Processing System. | 2 | September 07, 2012 | 2012 | Steris Corporation | The firm initiated a recall after complaint investigations revealed that when the units are left on and not in use at night and/or on weekends the hose disconnects/separates resulting in water leakage that can cause damage to property. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | STERIS sent a Urgent Voluntary Field Correction Notice letter dated August 1, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS has received no reports of injuries associated with the disconnection of a SYSTEM IE water hose.STERIS Corporation will instill new hoses and connections on your SYSTEM IE Liquid Chemical Sterilant Processing System. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 6051 | 6,051 units | Worldwide Distribution -- USA (nationwide) and the countries of Bahrain and Hong Kong. | Open |  |  | 6051 | Not_Computer
Z-2334-2012 | 61869 | booth, sun tan | LEJ | Ultraviolet lamp for tanning. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | SUN¦DASH RADIUS 252 tanning system; 52 lamps, Ultraviolet Radiation.The device is intended to be used for tanning of the human skin. | SUNDASH RADIUS 252 | 2 | September 07, 2012 | 2012 | JK Products & Services, Inc | The firm has tracked a systematic failure in the door cord on the Sun¦Dash Radius 252 product line. In some instances when the door cord is installed with little or no slack outside the upper decor piece, the cord has been known to fail causing a short in the wiring at the pivot point in the cord. | PRODUCTION CONTROLS: Equipment Maintenance | JK Global Services sent out Mandatory Inspection/Modification letters to their Distribution Service Teams dated 07/08/2005, 01/12/2006, 09/15/2010, 10/04/2010 and 03/01/2012. A letter dated September 15, 2010, was sent to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to contact a JK representative or their Distributor as soon as possible to schedule an inspection to the affected product in their possession. For any questions customers were instructed to call 1-800-445-0624 or send an e-mail to support@jkamerica.com. |  | 1948 | 1948 units | Worldwide Distribution - USA (nationwide) and Canada | Open |  |  | 1948 | Not_Computer
Z-2336-2012 | 62970 | prosthesis, laryngeal (taub) | EWL | Laryngeal prosthesis (Taub design). | Ear Nose & Throat | Ear Nose & Throat | 510(k) | BLOM-SINGER indwelling TEP Occluder, catalog IN2004-TOThe Blom-Singer Indwelling (TEP(tracheoesophageal Puncture) Occluder is a non-valved silicone prosthesis designed for placement in and maintenance of the TE (tracheoesophageal) puncture following total laryngectomy. The Blom-Singer Indwelling TEP (tracheoesophageal Puncture) Occluder is indicated for placement in and maintenance of the TE puncture following total laryngectomy when placement, or replacement, of an indwelling prosthesis is performed by a qualified, trained medical professional. | BLOMSINGER indwelling TEP Occluder | 3 | September 07, 2012 | 2012 | Helix Medical LLC | Helix Medical LLC is recalling the BLOM-SINGER indwelling TEP Occluder due to an invalid expiration date. | PRODUCTION CONTROLS: Labeling Mix-Ups | A recall letter dated March 8, 2012, was sent to customers who purchased the Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO). The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to contact Customer Service Center at (800) 477-5969 if they have products to be returned. Customers were instructed to return Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO) to:Helix Medical, LLCAttn to: Customer Service-IN2004-TO.1110 Mark AvenueCarpinteria, CA 93013For questions regarding this recall call 805-684-3304. |  | 2 | 2 devices | Worldwide distribution: USA (nationwide) including state of: Ohio and country of United Kingdom | Terminated | April 02, 2013 | 207 | 2 | Not_Computer
Z-2331-2012 | 62875 | tray, surgical | LRP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | CMS-2981-R3 Arthrogram Tray containing 2 oz. bottle PVP Prep SolutionCMS-2981-R3 Arthrogram Tray containing 2 oz. bottle PVP Prep Solution | PVP Prep Solution | 2 | September 07, 2012 | 2012 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer due to no microbial testing in place at the manufacturer and product did not meet proper finished goods specifications. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated August 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory for the affected product and to discontinue the use of any affected product. Customer were instructedto complete and return the Return Response form, acknowledging they have received the letter and checked their inventory. Any product still in the customers inventory should be returned to the firm. Customers should call (919) 202-8462, ext 205 to obtain a Return Authorization Number and the shipping number.For questions regarding this recall call 919-202-8462. |  | 15 | 15 cases | Nationwide Distribution including GA, MI, NY, and PA. | Terminated | September 10, 2012 | 3 | 38 | Not_Computer
Z-2333-2012 | 62842 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Aesculap S4 Element 4.5mm Polyaxial Screws (ST240T, ST241T, ST242T, ST243T, ST244T, and ST245T).The S4 Spinal System is intended for anterior/anterolateral and posterior, non-cervical pedicle and non-pedicle fixation. Fixation is limited to skeletally mature patients and is intended to be used as an adjunct to fusion using autograft or allograft. The device is indicated for treatment of the following acute and chronic instabilities or deformities: 1 - degenerative disc disease, 2 - spondylolisthesis, 3 - trauma, 4 - spinal stenosis, 5 - deformities or curvatures, 6 - tumor, 7 - pseudoarthrosis, and 8 - failed previous fusion. | Aesculap S4 Element 4.5mm Polyaxial Screws | 2 | September 07, 2012 | 2012 | Aesculap Implant Systems LLC | There is a possibility that the screw channel that holds the rod in place may be out of tolerance. When the depth of the channel is not deep enough, the rod may not be able to provide ample compression on the insert preventing the screw head from locking down to the bone screw. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Aesculap sent an "IMPORTANT PRODUCT REMOVAL NOTIFICATION" letter dated July 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Distribution Inventory Sheet was attached for customers to complete and return with the affected product. Contact the firm at (610) 984-9275 for questions regarding this recall. |  | 5 | 5 | Nationwide Distribution-including the states of IL, MN, NV, OK, and TX. | Open |  |  | 5 | Not_Computer
Z-2327-2012 | 62774 | catheter, coude | EZC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Foley procedural trays packaged in breather bags. The product is packaged ten units in each shipper box, and are sold sterile for single use.Used to catheterize patients. The finished trays are sold in Tyvek breather bags. | Bard Foley procedural trays packaged in breather bags | 2 | September 07, 2012 | 2012 | C.R. Bard, Inc., Urological Division | Incomplete seal on sterile packages of product. |  | C. R. Bard Inc, sent a Urgent Product Recall Notification letter dated June 27. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This notification is being sent because there is a potential for a limited number of units within the subject lots (noted above) that may contain an open seal of the sterile packages. The issue is visually obvious even before the package is opened. No patient injury or adverse events have been reported. Do not use or further distribute any of the affected products.Once the product affected by tlle recall has been removed from your inventory 1.Fill out the Recall & Effectiveness Check Form. Be sure to state the quantities affected of the recalled product that you have in stock. It is extremely important that we receive this information. 2. Fax the Recall & Effectiveness Check Form to BMD at 1-770-784-6469. The Recall Coordinator will issue you with a Return Authorization (RCL) to facilitate the expedient return and replacement of the product. 3. A mailing label is enclosed for your convenience to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and include the RCL number. All products should be retumed to the following shipping address that is in the letter. | Contact the recalling firm for information | 56694 | 56,694 units | Nationwide distribution: USA | Terminated | February 07, 2013 | 153 | 56694 | Not_Computer
Z-2346-2012 | 62909 | pump, infusion, pca | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Sterile Empty Vial and Injector (30mL PCA vial), Rx only, One/List No. 6021-03, Hospira, Inc., Lake Forest, IL 60045.For use with a compatible Hospira PCA pump set with injector and a compatible Hospira infusion device. | Sterile Empty Vial and Injector (30mL PCA vial), | 2 | September 10, 2012 | 2012 | Hospira, Inc. | Reports of leaking during filling and administration. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated August 3, 2012, to all affected customers. The letter identifed the product, the probem, and the action to be taken by the customer. Customers were instructed to first consider the use of Prefilled Morphine PCA Vials wherever possible. In some instances, the treating physician may require an alternative to morphine.Customers may continue using the affected product by following the instructions provided in the Correction Letter. Customers were asked to complete the attached Reply Form and fax it to 1-888-879-8144.For questions regarding this recall call 877-946-7747. |  |  | unknown | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, NK, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY and the countries of Canada, Costa Rica and the Virgin Islands. | Open |  |  |  | Not_Computer
Z-2340-2012 | 62915 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Multileaf CollimatorTo be used for single or multiple fractions, delivered as static and/or dynamic beams of radiation, in all areas of the body where such treatment is indicated. | Agility | 2 | September 10, 2012 | 2012 | Elekta, Inc. | Information released with the product Agility did not originally contain information on compatibility with the add-on Dynamic Multileaf Collimator Apex. | PRODUCTION CONTROLS: Error in Labeling | ELEKTA sent an Important Field Safety Notice, FCO Ref: 20001405000 dated July 7, 2012, to current Agility users and a separate notice FCO Ref: 20001407026 for current Apex users who may be in the process of considering an upgrade to Agility. Both notices made it clear that Apex has not been integrated or classified as compatible for use with Agility. Users are advised not to use Apex with Agility.For questions regarding this recall call 770-300-9725. |  | 2 | 2 | Nationwide Distribution including DC & MD | Open |  |  | 2 | Not_Computer
Z-2339-2012 | 62807 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Tip Cover for 8mm Monopolar Curved Scissors (Hot Shears)Accessory, Disposable, 8mm MCS, Box of 10.The intended use of the Tip Cover Accessory is to provide dielectric insulation over a section of the Monopolar Curved Scissor instrument so that RF energy is only available at the scissor blades. The Monopolar Curved Scissors are intended for endoscopic manipulation of tissue, including: cutting, blunt and sharp dissection, electrocautery. This accessory is to be used with the Intuitive Surgical Endoscopic Instrument Control Systems (da Vinci, da Vinci S and da Vinci Si Surgical Systems Models IS1200, 12000, ISSOQO). | Tip Cover, Accessory | 2 | September 10, 2012 | 2012 | Intuitive Surgical, Inc. | There is potential for the sterility of the product to be compromised. | TRAINING: Employee Error | Intuitive Surgical sent a "MEDICAL DEVICE CORRECTION" letter dated August 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return via fax to (408) 716-3040. Contact Customer Service at 1-800-876-1310 for questions regarding this recall. | Contact the recalling firm for information | 11121 | 11,121 Boxes | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Canada, and Lebanon | Terminated | December 31, 2012 | 112 | 11121 | Not_Computer
Z-2341-2012 | 62988 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Uroskop Access.Fluoroscopic, image intensified x-ray system | Siemens Uroskop Access | 2 | September 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens has discovered a possible hazard to patients or hospital personnel when tilting the Uroskop Access system starting with serial number 6076. When tilting the table on the Uroskop Access table, the motor may become disconnected and the entire table assembly may fall on the floor. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent a Update instructions letter dated April 2012, to all affected customers. The safety instruction letter identified the product the problem and the action needed to be taken by the customer.The update has been completed successfully. The customer has been informed of the benefits of improvements resulting from this update.Further questions please call (610) 219-4834. | Contact the recalling firm for information | 5 | 5 | Nationwide distribution: USA including states of: IN, MI, MS, TN and TX. | Terminated | September 20, 2012 | 10 | 5 | Not_Computer
Z-2367-2012 | 62989 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira mobile x-ray systemProduct Usage:Mobile x-ray system | Siemens Mobilett Mira | 2 | September 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, firm became aware of a potential issue with the Mobilett Mira system, (serial numbers 1034, 1038, 1041, 1042, 1043, 1050, 1053, and 1054. During normal clinical operation the center screw on the driving wheel may become loose resulting in system immobilization. There is no risk of the system tipping over or collapsing on the floor, however the system may tilt t | DESIGN: Device Design | Siemens has issued an Update Instruction XP018/12S in April 2012 to all affected consignees notifying them of the problem. The update instruction provides a correction of the wheel fixation. For questions call Siemens at 610-219-4834 or 610-448-4634. | Contact the recalling firm for information | 9 | 9 | US Nationwide Distribution including the states of: GA, IA, KS, NC and TX. | Terminated | September 25, 2012 | 14 | 9 | Not_Computer
Z-2362-2012 | 62140 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 101-05-30***NOVATION MODULAR DRILL BIT LENGTH: 30mm, DIAMETER 3.2mm***Use with 30mm Screw***STERILE***EXACTECH 2320 NW 56th Court, Gainesville, FL 32653***.To assist the surgeon in the implantation of Novation Hip system components in according to a conventional technique for hemi & total hip arthroplasty. Specifically, these instruments are reusable, surgically invasive, and intended for transient use. | NOVATION MODULAR DRILL BIT | 2 | September 11, 2012 | 2012 | Exactech, Inc. | Exactech Inc. of Gainesville, FL is recalling their Novation Modular Drill Bits after the devices were reported to have been used by end user while in a non-sterile condition. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Finished Device Change Control | The firm, Exactech, Inc., called the lone customer on January 24, 2012 to discuss the product, problem and actions to be taken. The customer was instructed to return all unused devices and associated packaging. |  | 2 | 2 | Nationwide distribution: USA state of: Arkansas only | Terminated | September 26, 2012 | 15 | 2 | Not_Computer
Z-2368-2012 | 62926 | spinal vertebral body replacement device | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Enduramesh Corpectomy Spacer SystemProduct Usage: The Enduramesh is a vertebral body replacement system intended for use in the thoracic and lumbar spine (T1-S1) to replace a collapsed, damaged or unstable vertebral body due to tumor or trauma/fracture. The device is intended for use with either allograft or autograft. The device may be used with a cleared supplemental fixation system based on surgeon discretion and preference. | Enduramesh Corpectomy Spacer System | 2 | September 11, 2012 | 2012 | Lucero Medical LLC | The firm voluntarily recalled the device after they discovered that the sterilization recommendations on the product insert did not pass sterilization validation. | PRODUCTION CONTROLS: Error in Labeling | Lucero Medical sent a Field Correction Notice letter dated July 15, 2012 to it's sole consignee. The notice identified the affected product., problem and actions to be taken. The customer was instructed to complete the response card and submit back to their local Lucero Medical Representative. | No consumer action necessary | 169 | 169 | US Nationwide Distribution | Terminated | February 20, 2013 | 162 | 169 | Not_Computer
Z-2348-2012 | 62321 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Safety monitor included with System 8000 BaseThe Sarns" Modular Perfusion System 8000 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. The safety monitor is indicated to warn when the blood level in the reservoir reaches the operator-set limit, and to stop the arterial and cardioplegia pumps. | Sarns System 8000 | 2 | September 11, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received complaints from our Affiliate, TerumoDubai that the safety monitors were tripping off. Terumo CVS' investigation determined that 8monitors were manufactured with an underrated circuit breaker due to an error where an associatepulled an underrated circuit breaker from an incorrect inventory location. Since this time, inventorylocations hav | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | TERUMO Cardiovascular Systems communicated with their affiliate, Terumo Dubai, in 2009 at the time the complaints were reported. No further communication was warranted.For questions regarding this recall call 734-741-6173. | Contact the recalling firm for information | 8 | 8 | Worldwide distribution: International only-country of: Terumo Dubai, UAE | Terminated | January 29, 2013 | 140 | 8 | Not_Computer
Z-2369-2012 | 54659 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 V1.05(FG-5400-00).The CARTO 3 V1.05 System is intended for catheter-based atrial and ventricular mapping. The mapping system allows real-time display of cardiac maps in a number of different formats. Maps may be displayed as cardiac electrical activation maps, cardiac electrical propagation maps, cardiac electrical potential maps, impedance maps and cardiac chamber geometry maps. The acquired patient signals, including body surface ECG and intracardiac electrograms may also be displayed in real time on the System's display screen. The CARTO 3 V1.05 System is also intended to support EP procedures, maintaining CARTOÂŽ XP System capabilities, in the presence of a high metallic environment and provide a data communication channel to the Stereotaxis Niobe Catheter Guidance System as well as in a regular EP lab. The CARTO 3 V1.05 System includes the CartoMerge capability to import, register and merge CT or MRI structural images with CARTO Maps physiological information and real time catheter navigation, the Fast Anatomical Mapping (FAM) functionality, enabling creation of cardiac maps using'catheters with magnetic proximal and distal location sensors and the CartoSound image integration functionality, that enables integration of intracardiac echo (ICE) visualization to provide 3D combined maps. In addition to the use of specialized navigation catheters, the system is also intended for use with non-navigational electrophysiology catheters. | CARTO 3 V1.05 System | 2 | September 12, 2012 | 2012 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has become aware of information regarding a potential hazard that may occur during the operation of a Bloom EP Stimulator device in conjunction with RF generators and other electrophysiology (EP) equipment, including the CARTO 3 Electropysiology Mapping System. |  | Biosense Webster sent an " URGENT FIELD SAFETY NOTICE" dated November 20, 2009 to all affected customers who purchased the Carto 3 Electrophysiology Mapping System. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to continue use of the CARTO 3 System and that it was safe and no further action is needed. Customers with questions about the notice were instructed to contact Biosense Webster representatives.On January 12, 2010 another letter was sent from Biosense to all their customers recommending to avoid any connection of an EP stimulator's pacing pins to the direct ports on the CARTO 3 System when applying RF energy, unless the set up is indicated by the stimulator's instructions for use. The letter also informed the customer that the user manual for the CARTO 3 System has been formally revised to include a warning. Customers were instructed to complete an acknowledgement form to confirm receipt of the revised user manual. | Contact the recalling firm for information | 71 | 71 units total (30 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Denmark, France, Germany, Israel, Italy, Netherlands, and United Kingdom. | Terminated | September 12, 2012 | 0 | 71 | Not_Computer
Z-2372-2012 | 62912 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnilife Science Apex Knee System Modular Tibia Augment, Size 3 x 4mmRef: KC-54030.The Apex Knee System is intended for use as a primary or revision total knee replacement. | Omnilife Science Apex Knee System Modular Tibia Augment | 2 | September 13, 2012 | 2012 | Omni Life Science | The Apex Knee System Modular Tibia Augment may breach the inner and outer sterile pouches, compromising the sterility | PRODUCTION CONTROLS: Packaging | Omnilife Science contacted affected customers by phone and sent an "URGENT MEDICAL DEVICE RECALL" e-mail on August 9, 2012. The e-mail identifies the product, problem, and actions to be taken by the customers. Customers with questions regarding product return should contact Customer Service at 800-448-6664. Customers with questions regarding the recall should contact the firm at 774-226-1845. | Contact the recalling firm for information | 310 | 310 units | Worldwide Distribution-USA (nationwide including the states of CO, NY, OK, PA, UT, VA, and WI, and the countries of Spain and the Netherlands. | Terminated | January 10, 2013 | 119 | 1396 | Not_Computer
Z-2392-2012 | 62894 | instrument, coagulation, automated | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | Instrumentation Laboratory ACL TOP 700 CTSPN 0000280050 ACL TOP is a bench top, fully automated, random access analyzer designed specifically for in vitro diagnostic clinical use in the hemostasis laboratory for coagulation and/or fibrinolysis testing in the assessment of thrombosis and/or hemostasis. The systems provide results for both direct hemostasis measurements and calculated parameters | Automated Coagulation Analyzer | 2 | September 14, 2012 | 2012 | Instrumentation Laboratory Co. | Low Abnormal Fibrinogen Sample Reports as Normal and High Fibrinogen Sample Reports as Low | DESIGN: Software Design | The firm, Instrumentation Laboratory, sent a "Mandatory Test Parameter Upgrade" letter dated June 7, 2012 with upgrade kits P-14.8 via USPS on August 3 thru August 10 to US and Canadian customers and notified International Affiliates via e-mail on August 2, 2012. The customers were ask to complete and return the Mandatory Test Parameter Upgrade Tracking Form via fax: 781-861-4207 or email: ra-usa@ilww.com. If you have any questions, call 781-861-4204. | No consumer action necessary | 140 | 140 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, REPUBLIC OF LUXEMBOURG, MACAU, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SLOVAKIA (Slovak Republic), SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED KINGDOM, URUGUAY and VIETNAM. | Open |  |  | 4297 | Not_Computer
Z-2384-2012 | 62884 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost Systems with M Cabinet Model: 712025Stationary radiographic system, digital-Universal RAD Applications | Philips DigitalDiagnost Systems | 2 | September 14, 2012 | 2012 | Philips Healthcare Inc. | Creepage Distance between the two primary phases in the GEO Transformer is 4.5 mm.does not meet (IEC) International standard of 5.5 mm. | DESIGN: Device Design | Philips Healthcare Imaging issued an Urgent - Field Safety Notice dated August 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The unit will be exchanged with a new unit that complies with the IEC requirements. Please reference Field Change Order (FCO) 71200081 when contacting I your local Philips representative. 1-800-722-9377.Customer were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter and to maintain a copy of the letter with the equipment instruction for Use. For further assistance, customers should contact their local Phillips representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 978-687-1501. | No consumer action necessary | 4 | 4 systems | Worldwide Distribution - USA including OH and WA and the country of Canada | Open |  |  | 4 | Not_Computer
Z-2381-2012 | 62728 | guide, surgical, instrument | FZX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | The 2.7mm Drilling Guide is a reusable instrument used in conjunction with a drill to prepare oriented holes in bone for insertion of a Standard SurfixÂŽ Locking Screw with plates for the Tiblaxis, Uni-CP, Advansys MLP and Advansys DLP compression plate systems.The Drilling Guides are initially supplied to users in kits containing the devices that comprise the Tibiaxis, Uni-CP, Advansys MLP and Advansys DLP systems. Should the Drilling Guides need to be replaced, they can be replenished by the customer placing an order. The replacement is shipped to the customer in sealed plastic bags. Uni-CP" Compression Plate The Uni-CP" Compression Plate is indicated for fixation of bone fractures or for bone reconstruction: - Arthrodesis in foot and ankle surgery - Fracture management in the foot and ankle - Mono or bi-cortical osteotomies in the foot and ankle TIBIAXYS" Plates The NewdealÂŽ TIBIAXYS" Plates are indicated For fixation of bone fractures or for bone reconstruction. Examples include Osteotomies and fractures of ankle joint, distal tibia and fibula. The NewdealÂŽ TIBIAXYS" Plates must be fixed with the Newdeal SurfixÂŽ 3.5mm diameter fixation screws and lock-screws. ADVANSYSÂŽ Mid-Foot Plating System: Dorsal Lisfranc Plate (DLP) and Medical Lisfranc Plate (MLP) The Dorsal Lisfranc Plates are intended for fractures, fusions, osteotomies and replantations of small bones at the tarsometatarsal joints The Medial Lisfranc Plates are intended for bone fixation such as: arthrodesis of the 1st metatarsocuneiform joint to reposition and stabilize a metatarsus primus varus: - Lisfranc arthrodesis - Mono or bi-cortical osteotomies or fractures near the 1st metatarsocuneiform joint | Integra 2.7 mm Drilling Guide | 2 | September 14, 2012 | 2012 | Integra Life Sci. | Integra' s quality system noted an adverse trend in complaints for breakage of the drilling guides during use. None of the complaints resulted in patient injuries but in some cases resulted in an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that any potential for risk would be transient, self-limiting or minor, or result | DESIGN: Device Design | Integra sent a "SUBJECT: MEDICAL DEVICE RECALL" letter dated June 14, 2012 to all affected customers. The letter identified the product, problem, and action to be taken by the customers. Customers were asked to complete a Recall Acknowledgement and Return Form whether or not they were affected by the recall, and return all their affected inventory to Integra LifeSciences.Customers were asked to contact their regional managers or the firm at 609-936-6832 for questions regarding the recall. | No consumer action necessary | 222026.7 | 2,389 units, Catalogue No. 219635ND; 2.7mm Drilling Guides | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Austria, Czech republic, Denmark, France, Germany, Ireland, Israel, Italy, Kingdom of Saudi Arabia, Lebanon, Norway, Portugal, Switzerland, South Africa, Spain, Sweden, and United Kingdom. | Terminated | February 20, 2013 | 159 | 2389 | Not_Computer
Z-2386-2012 | 63014 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System, models 728131 and 728132. The MX 16-slice CT Scanner System can be used as a whole body computed tomography X-ray system. | MX 16slice | 2 | September 14, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On August 24, 2012 the firm recalled the MX 16-slice after discovering that a bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | DESIGN: Process Design | An "Urgent Medical Device Correction" Field Safety Notice dated August 24, 2012 was sent to all US consignees via certified mail. The problem was described and recommended actions were provided. Philips Key Market representatives distributed the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Orders. US Customer contact number: 1 800-722-9377. |  | 328 | 328 units | Worldwide distribution, Nationwide (US) including the states AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, and VA and the countries of Albania, Argentina, Australia, Belgium, Brazil, Chile, China, Colombia, Dominican Republic, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Italy, Jordan, Lebanon, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Romania, Russia, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Turkey, Uganda, and Venezuela. | Terminated | August 28, 2013 | 348 | 328 | Not_Computer
Z-2383-2012 | 63033 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Trilogy, Novalis Tx, Unique, linear accelerators; Model Number: H14, H18, H29, Reference/FSCA Identifier: CP-08881; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA.1) The Varian Low Energy Linear Accelerator is indicated for precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment i is indicated. (2) The Varian High Energy Linear Accelerator is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Varian brand Clinac, Trilogy, Novalis Tx Unique, linear accelerators | 2 | September 14, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Some X-jaws (lower collimator jaw) carriers were manufactured with the incorrect metal alloy, which can potentially cause the jaw carrier to crack. If both jaw carriers on the same X-jaw were to crack, the jaw would be able to move freely in the closed direction. In this case, the deviation of the jaw position will not be detected by the position readout interlock circuitry; Treatment after such | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Varian sent an "URGENT MEDICAL DEVICE CORRECTION URGENT FIELD SAFETY NOTICE" dated August 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-827-4265 for questions regarding this recall. |  |  |  | Worldwide Distribution. | Terminated | December 04, 2013 | 446 |  | Not_Computer
Z-2385-2012 | 62899 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sterile Distraction Screws of 3 different sizes: 12mm, 14mm and 16mm sold under TSI, 3COR, DB Surgical, and American Surgical Products brands and Sterile XL Screws of 7 different sizes: 90mm, 100mm, 110mm, 120mm, 130mm, 140mm, 150mm sold under TSI brandSterile Distraction Screws are intended for the distraction in the anterior approach to the cervical spine. Sterile XL Screws are intended for the fixation of the blade to the vertebral body to prohibit drifting in the extreme lateral approach to the lumbar spine. | Sterile Distraction Screws of 3 different sizes | 2 | September 14, 2012 | 2012 | Tedan Surgical Innovations Llc | This letter is to notify you that several of our sterile screw products are being voluntarily recalled. Product sterility may be compromised due to a mechanical failure of the protective packaging tube. Tube end plugs may dislodge creating the potential for the screw to come out of the protective tube and then the sharp end of the screw pierce the outer sterile barrier pouch. | PRODUCTION CONTROLS: Packaging | The firm, TeDan Surgical Innovations (TSI), sent an "URGENT: Medical Device Recall" letter dated June 11, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine all products subject to recall and return product to TSI. In addition, if they have further distributed this product, identify all customers and notify them of this product recall at once in writing; and complete and return the enclosed Mandatory Response Form as soon as possible and no later than June 25, 2012 via mail: 12675 W. Airport Blvd., Suite 200, Sugar Land, TX 77478; fax: (713) 726-0846 or email: thernandez@tedansurgical.com.If you have any questions, contact the Business Manager at 713-726-0886, Monday thru Friday, from 8am - 5pm or via email at THERNANDEZ@TEDANSURGICAL.COM. | Contact the recalling firm for information | 384 | 12,065 Sterile Distraction Screws (2413 boxes); 384 Sterile XL Screws (128 boxes) in distribution | Worldwide distribution: USA (nationwide) and countries of : Australia and Taiwan. | Terminated | July 31, 2013 | 320 | 12449 | Not_Computer
Z-2394-2012 | 63105 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, Pre-Filled Flush Solution, 10 mL in 12 mL Single Use Syringe, Rx only, List no. 1078-20, made in Nashville, TN, manufactured for Hospira, Inc., Lake Forest, IL 60045 USA.Flush syringe. | 0.9 Sodium Chloride Injection, USP, PreFilled Flush Solution, 10 mL in 12 mL Single Use Syringe | 3 | September 15, 2012 | 2012 | Amsino Medical USA | One lot of 0.9% Sodium Chloride Flush syringe, 3 mL in 12 mL may be mislabeled as 0.9% Sodium Chloride Flush syringe, 10 mL in 12 mL. | TRAINING: Employee Error | The firm, Amusa, notified their sole consignee of the mislabeled product via e-mail and telephone on August 15, 2012. An "URGENT-CLASS II MEDICAL DEVICE RECALL" letter dated August 22, 2012 was sent to the customer. The letter identified the product, problem, and actions to be taken by the customer. A Response Form was included for the customer to complete and return. Contact the firm at 615-521-2723 for questions regarding this recall. |  |  | unknown | Distributed to one customer in Illinois. | Terminated | August 13, 2013 | 332 |  | Not_Computer
Z-2395-2012 | 63078 | container, i.v. | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Empty Evacuated Container 250 mL and 1000 mL.Empty Evacuated containers are used for mixing compatible medications or blood collections. | Empty Evacuated Container 250 mL and 1000 mL. | 2 | September 17, 2012 | 2012 | Hospira Inc. | Embedded iron oxide glass defect which may have the potential to break off into the solution. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Hospira sent an "URGENT DEVICE RECALL" dated August 28, 2012 to all affected customers. The letter identified the product, problem,and actions to be taken by the customers. The letter included an Urgent Device Recall Reply Form for customers to complete and return via fax to 1-888-237-5582. Contact Hospira Customer Care at 1-877-946-7747 for questions regarding this recall. |  | 191094 | 191,094 units | Worldwide Distribution-USA (nationwide) and the countries of Saudi Arabia, Canada and Barbados. | Terminated | September 18, 2013 | 366 | 191094 | Not_Computer
Z-2402-2012 | 54705 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | A mobile, electro-hydraulically operated, fluoroscopic/C-Arm compatible general surgical table. Product Usage:Product provides a flexible, easy-to-use, articulated posturing of a patient for either general, I.A., or specialized surgical procedures. There are movable components and accessories intended for use during diagnostic examinations or surgical procedures to support and position a patient. | Amsco 3085SP Surgical Table | 2 | September 17, 2012 | 2012 | Steris Corporation | Customers may be storing objects on the base and/or around the column of the surgical table. This prohibited practice and misuse known by the customers can damage the override switch assembly located at the top of the column cover. Depending upon the nature and severity of the damage to the electronics of the table, it is possible that the table could move without prompting by an operator . | TRAINING: Employee Error | Steris sent an Urgent Field Correction Notice letter dated February 17, 2010 to their customers. The letter identified the affected product, description of the problem and actions to be taken. Customers were instructed to examine their units for any structural damage. If found they are to discontinue use and arrange for a field representative to visit the facility to assess the damages and arrange for repair of the unit. For information or questions regarding STERIS visit to your facility contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 11044 | 11,044 Units | Worldwide Distribution - US Nationwide including District of Columbia and Puerto Rico and the following countries: American Samoa, Argentina, Australia, Bahrain, Bangladesh, Bermuda, Brazil, Canada, China, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, Germany, Greece, Guam, Hong Kong, India, Indonesia, Israel, Italy, Japa, Jordan, Kazakhstan, Korea, Kuwait, Lebanon , Libya , Malaysia , Mexico, Morocco, New Zealand, Panama, Peru, Philippines , Poland, Portugal, Qatar, Russia, Saudi Arabia , Singapore, South Africa, Spain, Sri Lanka, Taiwan, Thailand, Trinidad and Tobago, Turkey & United Kingdom. | Terminated | March 13, 2013 | 177 | 11044 | Not_Computer
Z-2396-2012 | 63047 | endoscope and/or accessories | KOG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | InfinityÂŽ sampling device, Cytology Brush, United States Endoscopy Group, Inc., 5976 Heisley Road, Mentor, OH 44060The Infinity sampling device is a disposable cytology brush used to retrieve cytological cell samples in the gastrointestinal tract. | Infinity sampling device, Cytology Brush | 2 | September 17, 2012 | 2012 | US Endoscopy Group Inc | US Endoscopy received five complaints which noted difficulty in deploying the cytology brush through the catheter. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | US Endoscopy sent an Addendum and Correction to the Instructions for Use letter dated December 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to rinse the brush head with sterile water immediately prior to use (reference the attached Addendum. Customers were also instructed to retract and deploy the brush to ensure that is is functioning properly prior to use. Customers were asked to review the instruction for use for this product along with the Addendum to ensure proper usage and performance.For questions customers should call 1-410-639-4494 or their local Endoscopic Product Specialist.For questions regarding this recall call 440-639-4494, ext 378. |  | 393 | 393 Devices | Worldwide Distribution - USA including CA, FL, GA, IA, IL, MA, ME, MN, MO, NC, NH, OH, OK, PA, TN, TX, VA, WA & WV and the countries of Australia, Canada, Columbia, Finland, Germany, India, Israel, Italy, Netherlands, New Zealand & United Kingdom. | Terminated | March 22, 2013 | 186 | 393 | Not_Computer
Z-2399-2012 | 62391 | punch, surgical | LRY | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | JUGGERKNOT INÂŽGUIDE PUNCH, 2.9MM.The JuggerKnot 2.9mm In-Guide Punch is intended to be used to prepare a 2.9mm hole into bone in which the JuggerKnot 2.9mm Suture Anchor is to be placed. The JuggerKnot 2.9mm Suture Anchor is intended to be used to attach soft tissue to bone. | Biomet Juggerknot in Guide Punch | 2 | September 17, 2012 | 2012 | Biomet, Inc. | Investigation found that the current design of the product allows the press-fit of the knob/shaft interface to be overcome by the impaction force of a mallet, causing the knob to potentially unfasten and slide down the shaft. | DESIGN: Device Design | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return via fax to 574-372-1683. Questions related to the recall should be directed to 574-371-3758, Monday through Friday, 8 a.m. to 5 p.m. | Contact the recalling firm for information | 62 | 62 | Nationwide Distribution-including the states of WA, IN, AK, PA, GA, SD, KS, NY, VA, AR, AL, OH, WI, MA, and CA. | Terminated | April 18, 2013 | 213 | 62 | Not_Computer
Z-2400-2012 | 63098 | adaptor, stopcock, manifold, fitting, cardiopulmonary bypass | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(k) | Guardian II Hemostasis Valve, Model 8210 and Guardian II NC Hemostasis Valve, Model FH102.Intended to maintain hemostasis during the use of diagnostic/interventional devices. The device is indicated for maintaining a seal around diagnostic/interventional devices with outside diameters up to 8.0F (2.67 mm or 0.105") during diagnostic/interventional procedures. | Guardian II NC Hemostasis Valve with Guidewire Introducer | 2 | September 17, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions, Inc., became aware of a potential problem involving pouches missing a seal from Guardian II (Model 8210) and Guardian II NC (Model FH102) Hemostasis Valve which poses a risk of product contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Vascular Solutions, sent an "URGENT - Medical Device Field Safety Notice" dated August 24, 2012 and a follow-up "Urgent Medical Device Removal - Lot Specific" letter dated August 30, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check their inventory and to follow steps outlined in the letter which included to complete and return the "VSI Account Inventory Form" via email to: customerservice@vasc.com or fax: to (763) 656-4251.If you have specific questions or concerns regarding this matter, please contact your local Vascular Solutions Account Manager; call (763) 656 4300 or email: Senior Regulatory Operations Associate at SOuellette@vasc.com and/or email: www.vascularsolutions.com. | Contact the recalling firm for information | 3500 | 3500 units | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, CT, DE, FL, GA, ID, IL IN , KY, LA,MD, MN, MO, NV, NM, NY, OH, PA, SC, TN, TX, VA, WA, and WV; and countries of: Italy and Korea.Italy, Korea | Terminated | December 04, 2012 | 78 | 3500 | Not_Computer
Z-2398-2012 | 62850 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | The Bio-Rad MONOLISA Anti-HBs EIA is for In Vitro Diagnostic Use and it has the catalog number 25220. It is a qualitative and quantitative enzyme immunoassay for the detection of antibody for hepatitis B surface antigen in human serum and EDTA or citrated plasma. The Bio-Rad MONOLISA Anti-HBs EIA is a qualitative and quantitative enzyme immunoassay for the detection of antibody for hepatitis B surface antigen in human serum and EDTA or citrated plasma. The assay results may be used as an aid in the determination of susceptibility to hepatitis B virus (HBV) infection in individuals prior to or following HBV vaccination or where vaccination status is unknown. Assay results may be used with other HBV serological markers for the laboratory diagnosis of HBV disease associated with HBV infection. | MONOLISA ANTIHBs EIA | 2 | September 17, 2012 | 2012 | Bio-Rad Laboratories | Prolonged exposure of the MONOLISA Anti-HBs EIA Conjugate Diluent component to light may result in elevated Optical Densities (OD) in some samples such that a negative sample could elevate to a grey zone or low positive result. | PRODUCTION CONTROLS: Packaging | The firm, Bio-Rad Laboratories, Inc., sent an "Important Product Information Bio-Rad MONOLISA Anti-HBs EIA Catalog #25220", letter dated 27 July 2012, to their domestic consignees/customers via FedEx. The International subsidiaries will be notified by e-mail. The consignees/customers were informed of the important storage and handling information regarding the MONOLISA Anti-HBs EIA Conjugate Diluent component . The consignees/customers were instructed to follow the storage and handling instructions as stated in the Package Insert, and compare their laboratories practices to the storage and handling instructions stated by Bio-Rad.Customers who have questions can call Bio Rad Technical Support at 1-800-224-6723, option 2, then option 3. |  | 2007 | 7,400 (approximately)- All lots since distribution date May 9, 2007 | Worldwide distribution: USA (nationwide) and countries of: Canada, South Africa, and Zimbabwe. | Terminated | December 14, 2012 | 88 | 7400 | Not_Computer
Z-2405-2012 | 62917 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Precise Digital AcceleratorDelivery of radiation to defined target volumes | Precise Digital Accelerator | 2 | September 19, 2012 | 2012 | Elekta, Inc. | There have been a small number of instances in hospitals where the joint between the drive gear and the gantry base drive wheel has become loose, leading to fatigue failure of some of the securing bolts. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Elekta sent an Important Field Safety Notice dated August, 2012 to inform users that inspection and possibly bolt replacement needs to be carried out. Users should return the signed acknowledgement to Elketa, and an inspection should be carried out under FCO20002103053.For questions customers should call +44(0)1293 654200.For questions regarding this recall call 770-300-9725. |  | 110 | 110 | Nationwide Distribution including AL, AZ, CA, CO, CT, FL, GA, IL, IA, MD, MA, MI, MN, NE, NV, NJ, NM, NY, ND, OH, OK, OR, PA, VA, SC, TN TX, WA, WV, WI and Puerto | Open |  |  | 110 | Not_Computer
Z-2410-2012 | 62897 | radioimmunoassay, follicle-stimulating hormone | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur FSH Assay: ADVIA Centaur Calibrator BUS: Calibrator B 00652707 (6 pack) or 00649625 (2 pack )OUS:OUS: 2 pack REF: 00650577 and Ref: 00654130 (6 pack)Product Usage:ADVIA Centaur CP FSH assay: For in vitro diagnostic use in the quantitative determination of follicle-stimulating hormone (FSH) in serum using the ADVIA Centaur and ADVIA Centaur XP systems | Siemens Healthcare ADVIA Centaur systems Calibrator B for FSH | 3 | September 19, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Prolactin Assay and FSH, Calibrator B fails the analytical sensitivity stated in the IFU for lots below 40 | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens Healthcare issued an Urgent Field Safety Notice letter dated August 2012 by e-mail to all affected Siemens Healthcare Diagnostics Branch offices on August 7, 2012. Customers were instructed to discontinue using the ADVIA Centaur Calibrator B lots 40 and below following their obtaining Calibrator 8 lot 41 or higher. Customers were asked to complete and fax-back the response form attached. |  | 15445 | 15,445 kits | Worldwide Distribution - US Nationwide and the countries of: Canada, Argentina, Australia, Bangaldesh,Brazil, Chile, China, Columbia, Croatia, Ecuduar, Egypt, Hong Kong, India, Israel,Japan, Korea, Maylasia, New Zealand, Mexico, Peru, Philipines, Pakistan, Saudia Arabia, Singapore, Thailand, Taiwan, Viet Nam, South Africa and UK. | Terminated | February 26, 2014 | 525 | 15915 | Not_Computer
Z-2415-2012 | 63001 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Kimberly-Clark MIC Safety PEG Kit - Pull TypeThe Kimberly-Clark MIC PEG Kits - Pull type are Percutaneous Endoscopic gastrostomy (PEG) tubes intended for a non-surgical "pull" endoscopic placement technique. The PEG tube is held in place during use with an external retention disk. | KimberlyClark MIC Safety PEG Kit Pull Type | 2 | September 20, 2012 | 2012 | Kimberly-Clark Corporation | The placement guidewire diameter is too large and does not pass-through the introducer needle during the placement procedure of the PEG Feeding Tube. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Kimberly-Clark Health Care sent an Urgent Voluntary Medical Device Withdrawal letter dated June 22, 2010, to all affected customers. The letter included a Product Withdrawal Response Sheet which was to be returned to Kimberly-Clark.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine and destroy all affected products and complete the attached Product Withdrawal Response Sheet and Fax the form to 920-225-4182.For questions customers should call 770-587-7362.For questions regarding this recall call 770-587-8393. |  | 277 | 277 units | Nationwide Distribution including AL, AK, AZ, CA, CO, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, ME, MA, MI, MN, MS, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | February 07, 2013 | 140 | 1043 | Not_Computer
Z-2423-2012 | 63146 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Asahi KASEI, REXEED(TM)- 15 LX, HEMODIALYZERdialysis filter | Asahi KASEI, REXEED(TM) 15 LX, HEMODIALYZER | 2 | September 20, 2012 | 2012 | Asahi Kasei Medical Co., Ltd., Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products, though there is no sufficient information to date to conclusively link these adverse events to the products. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | AsahiKASEI Medical Co., LTD sent a Medical DeviceRecall letter on September 21, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm that they have no inventory and to forward the letter to all clinics who may have received the product requesting they check their stock and arrange for the return of any recalled product. Non-responding consignees will be contacted by telephone.For questions regarding this recall call Asaki Kasei Medical America, Inc. at 888-362-6105. |  | 24 | 24 units | REXEED SX: Worldwide Distribution - USA including TN and the countries of Japan and GermanyREXEED LX: Worldwide Distribution - USA including TN and Germany | Terminated | February 05, 2013 | 138 | 325008 | Not_Computer
Z-2417-2012 | 63129 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | VerifyÂŽ Bowie-Dick Test Card, Model #S3098, Steris Biological Operations, 9325 Pinecone Drive, Mentor, OH 44060The Verify Bowie-Dick Test Card consists of a card printed on one side with bars of chemical indicator ink. On the opposite side of the card, space is provided for critical cycle information to be recorded. The card is laminated inside a thin clear plastic enclosure. The enclosure has two holes in it to allow air removal and steam penetration. The preassembled test is used to evaluate the effectiveness of air removal from the sterilizer chamber during prevacuum pulse steam sterilizer cycle. Following a prevacuum cycle, the chemical bars of the Bowie-Dick Test card uniformly darken indicating that residual air has been sufficiently removed to allow complete steam penetration into the test card. If air is trapped in the card during the exposure phase of the cycle, the color change of the bars will be incomplete or uneven. Thus the card can provide an immediate indication of inadequate removal of air during a cycle. The Bowie-Dick Type test is performed in an empty chamber each day the sterilizer is to be used, usually before the first sterilization cycle. | Verify BowieDick Test Card | 2 | September 20, 2012 | 2012 | Steris Corporation | STERIS has learned that the Verify Bowie Dick Test Cards are not performing to product specifications. Specifically, test cards have resulted in false fail results when sterilizer performance is within acceptable ranges. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Steris Corporation sent a Urgent Voluntary Recall Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine their inventory and compare the lot numbers printed on the box with the listed lots. If they have products from the affected lot numbers to contact STERIS Customer Service for account credit. After notifying STERIS of the amount of recalled product on hand they are to immediately destroy the product. Distributors are to also notify their customers of the recall. They are to send a copy of the Customer Recall Notice. They are to inventory their stock and notify STERIS of any remaining recalled product for replacement product. All recalled product should be destroyed.Please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 629 | 629 cards | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, IA, IN, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK,OR, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. and the country of Canada. | Terminated | March 27, 2013 | 188 | 629 | Not_Computer
Z-2413-2012 | 63069 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Bard Ventralight ST Mesh. - Ellipse 7" x 9" (17.8 cm x 22.9 cm) Product Code 5954790. The Bard Ventralight ST Mesh is a dual-component (absorbable and nonabsorbable) sterile prosthesis. The device classification name is "Mesh, Surgical, Polymeric." It is in tented for use in the reconstruction of soft tissue deficiencies, such as for the repair of hernias. | Bard Ventralight ST Mesh | 2 | September 20, 2012 | 2012 | Davol, Inc., Subs. C. R. Bard, Inc. | Mislabeled: The 8" x 10" package contains 7" x 9" VENTRALIGHT ST Mesh product and the 7" x 9" package contains 8" x 10" VENTRALIGHT ST Mesh | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Davol issued an "URGENT: MEDICAL DEVICE RECALL" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter advises customers to examine their inventory and quarantine affected product for return to the firm. Customers were asked to complete and return a Recall Effectiveness Check Form. Contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services & Support Department at 1-800-562-0027 if you have any questions. | Contact the recalling firm for information | 49 | 49 units | Nationwide Distribution. | Terminated | September 03, 2013 | 348 | 49 | Not_Computer
Z-2411-2012 | 63011 | perimeter, automatic, ac-powered | HPT | Perimeter. | Ophthalmic | Ophthalmic | 510(K) Exempt | Octopus 101 Perimeter; PeriTrend Software AnalysisA perimeter is a device intended to determine the extent of the peripheral visual field of a patient. The device projects light on various points of a curved surface, and the patient indicates whether he or she sees the light. | Octopus 101 Perimeter | 2 | September 20, 2012 | 2012 | Haag-Streit USA Inc | The firm recalled the device after learning of a possible deviation in the stimulus luminance of the Perimeter OCTOPUS 101. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, HAAG-STREIT USA, Inc., sent a "Field Safety Notice 2012-01/01" dated August 10, 2012 via certified mail or onsite visit to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make all users of the product and relevant personnel aware of this Field Safety Notice; forward a copy to third parties, if affected product has been passed on; keep this notice at least until the completion of the corrective measure, and complete and return the enclosed Confirmation of Receipt Form via fax to: 513-229-3867 or scanned and emailed to : fieldsafetynotice@haag-streit-usa.com.If you have any questions, call 513-336-7255. |  | 64 | 64 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, CT, FL, GA, IA, IL, LA, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, SD, WA, WI, and WV; and to the countries of Jamaica, Puerto Rico, and Mexico. | Open |  |  | 64 | Not_Computer
Z-2430-2012 | 62736 | apparatus, suction, ward use, portable, ac-powered | JCX | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Model 324JL On-Board Suction SystemIntended use: On-board suction system for Ambulance manufacturers. | Model 324JL OnBoard Suction System | 3 | September 20, 2012 | 2012 | Impact Instrumentation, Inc. | A shipment mix-up took place. | PRODUCTION CONTROLS: Process Control | The firm, Impact Instrumentation Inc., contacted their customers via phone and email notification on April 9, 2012. Customers were provided the correct packing slip and asked to acknowledge the correction.If you have any questions, contact the Regulatory Affairs Manager at 973-882-1212 or email: aglordano@impactii.com. |  | 2 | 2 devices | Nationwide distribution: USA including states of: FL and OH. | Open |  |  | 2 | Not_Computer
Z-2425-2012 | 62577 | drape, surgical | KKX | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ULTRA Fabric Reinforced Surgical Gown, X-Large, Sterile, Test File.The Kimberly-Clark ULTRA Fabric Reinforced Surgical Gown is a sterile, single use surgical gown intended to protect surgical patients and operating room personnel from the transfer of microorganisms, body fluids, and particulate material. | ULTRA Fabric Reinforced Surgical Gown | 3 | September 20, 2012 | 2012 | Kimberly-Clark Corporation | The impacted product was not intended for commercial distribution to end users. | TRAINING: Employee Error | Kimberly-Clark sent a Urgent Medical Recall Notice dated on December 31, 2010. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to: Evaluate their Distribution Facilities Inventory of Ultra Fabric Reinforced surgical Gown to determine if any of impacted product remained within their facility. If present, to discontinue distribution of the impacted product and place in a quarantined location. Complete the attached Distributor Product Withdrawal Response Form and fax it to KC within five days. Destroy the impacted product according to their facility procedures. And to review their distribution records to identify all customers who were shipped any of the impacted products. Notify each of their customers who were shipped any of the impacted products by sending them a copy of the attached Customer Letter with the Customer Withdrawal Response Form. A minimum of three (3) attempts to contact each non-responding customer within interval of approximately four to six weeks will be performed. Once all customers who were notified have responded or after approximately four weeks after the third attempt to contact the last non-responder, a recall Termination will be initiatedIf your Kimberly-Clark Representative has not already contacted you, please contact the Ultra Fabric Reinforce surgical Gown Coordinator at (770) - 587-7700. |  | 32 | 32 cases | US (nationwide) including the states of CA, CO, LA, NM, OK, TX and WA. | Terminated | November 01, 2012 | 42 | 32 | Not_Computer
Z-2431-2012 | 63016 | test, qualitative and quantitative factor deficiency | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Spectrolyse PAI-1, In-Vitro Diagnostic for Quantitative determination of Plasminogen Activator Inhibitor Type- 1 (PAI-1) activity in human plasma. Catalog Number: 101201For the determination of Plasminogen Activator Inhibitor Type 1 (PAI-1) activity in human plasma. The test is for in Citro diagnostic use. | Spectrolyse PAI1 | 2 | September 20, 2012 | 2012 | Sekisui Diagnostics Llc | SPECTROLYSE PAI-1 activity assay producing lower than expected absorbance values with "0" standard. This will reduce the slope of the standard curve requiring the user to make more frequent sample dilutions. | PRODUCTION CONTROLS: Process Control | The firm, Sekisui Diagnostics, LLC, sent an "URGENT PRODUCTS RECALL NOTICE" dated August 28, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to immediately stop using the product; to discard any remaining inventory and complete and return the Customer Acknowledgement Form by September 4, 2012 documenting the # of kits destroyed via FAX to Quality Assurance at 203-602-5553. The customers were also instructed to return the form even if they do not have the affected product currently in inventory. Sekisui will replace kits.If you have any questions contact, Customer Service Department at 800-637-3375, ext 10, 12, or 31, or email to recall coordinator at MKHAN@AMDIAG.COM. | Contact the recalling firm for information | 699 | 699 kits | Nationwide distribution: USA including states of: FL, IN, IL, and WA. | Open |  |  | 699 | Not_Computer
Z-2427-2012 | 63089 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 12/14 TAPER HIGH OFFSET NECK USE WITH SMITH & NEPHEW STEMS AND 12/14 FEMORAL COMPONENTS ONLY CO-CR, REF 71352112, Smith & Nephew, Inc. Memphis, TN 38116 USA.USE WITH SMITH & NEPHEW STEMS AND 12/14 FEMORAL COMPONENTS. | 12/14 TAPER HIGH OFFSET NECK | 2 | September 20, 2012 | 2012 | Smith & Nephew Inc | Modular necks were manufactured out of specification by a vendor; the taper provide is undersized. | TRAINING: Employee Error | Smith & Nephew sent an "URGENT:CLASS II RECALL-HIGH OFFSET MODULAR NECK R1213 letter dated August 22, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to cease use and distribution of the product and quarantine for return. An Inventory Return Certification Form was included for customers to complete and return via fax to 901-566-7975. |  | 70 | 70 units | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, IL, IN, MA, NC, NJ, NY, OH, OK, PA, TN, TX, and WA, and the country of Canada. | Terminated | February 27, 2013 | 160 | 70 | Not_Computer
Z-2426-2012 | 63142 | table, operating-room, electrical | GDC | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085 SRT Surgical Tables, Product Usage:The 5085 is a general surgical table with high patient weight capacity, extended width capability, and lower minimal table top elevation. It accommodates all general surgical procedures including but not limited to, cardiac and vascular, endoscopic, gynecology, urology, nephrectomy, neurology, ophthalmologic, orthopedics and other procedures requiring intraoperative fluoroscopic C-arm imaging and also supports laparoscopic surgical technique for the largest surgical patients. The 5085SRT table is a variation of the 5085 table that enables patient transport on hard level surfaces within the surgical suite (from pre-operative areas to the operating room and again from the operating room to post operative recovery). | Surgical Tables | 2 | September 20, 2012 | 2012 | Steris Corporation | STERIS has learned from Customer feedback and field service experience that damage from external impact to the table hydraulic column can cause the stainless steel shroud sections covering the table hydraulic column to become misaligned; if the damage is severe, the up/down movement of the table may be impeded. In addition, the table hand control may indicate that the table floor locks are unlock | DESIGN: Device Design | Steris sent an Urgent Voluntary Field Correction Notice dated April 6, 2012 to all affected customer. The letter identified the affected product, description of the problem and actions to be taken. Customers were informed that a STERIS Service Representative will contact affected facility to arrange for the replacement of the stainless steel shroud covering the hydraulic column and relocation of the floor lock pressure switches on their 5085 and/or 5085SRT table(s). For further information or if you have questions regarding the visit to your facility, contact STERIS Field Service Dispatch at 1-800-333-8828.If you have questions regarding this matter, call the Director of Marketing Tables and Case Goods at 440-392-7148 or STERIS Customer Service at 1-800-548-4873. | Contact the recalling firm for information | 379 | 379 tables | Worldwide Distribution - US Nationwide including the states of: AL, AR, CA, CO, CT, FL, HI, ID, IL, IN, IA, KY, LA, MA, MD, ME, MI, MN, MO, MS, NV, NYH, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WI & WV and countries of: Australia, Canada, China, India, Korea, Mexico, Panama and Thailand. | Terminated | September 04, 2013 | 349 | 379 | Not_Computer
Z-2360-2012 | 62132 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | BR980-9600 Hysteroscopy Sterile Procedure Kit containing 2 oz. bottle PVP Prep SolutionProduct Usage: Usage:The PVP contained in the kit is used in skin prep | PovidineIodine prep solution/Providine prep pads | 1 | September 20, 2012 | 2012 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer of the Povidone Iodine Prep Solution due to no microbial testing and product did not meet proper finished goods specifications. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated September 27, 2011 and a follow-up letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter requested that customers inspect inventory, discontinue use and return any remaining affected products. A Return Authorization Number and shipping number will be given to customer by calling 919-202-8462 ext 205. Customers were instructed to complete and return the attached Acknowledgement and Return Product form. | Contact the recalling firm for information | 135 | 75 cases & 60 Bottles for reps | US Nationwide Distribution including the states of CO, FL, KS, MA, MI, NJ and PA. | Terminated | April 29, 2013 | 221 | 994 | Not_Computer
Z-2429-2012 | 62978 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra HRAIM Head Ring Assembly with Intubation MountsThe CRW stereotactic system is used in conjunction with the HRAIM head ring is a target-centered system that can be configured to be CT-only or CT/MR compatible. The CRW stereotactic system is a multi-purpose system used for localizing intercranial targets for precisely directing instruments such as: Biopsy forceps RF lesioning electrodes Deep brain electrodes Recording and stimulating electrodes | Head Ring Assembly with Intubation Mounts (HRAIM) | 2 | September 20, 2012 | 2012 | Integra LifeSciences Corporation | T-handle screw is used on a complementary product to the CRW, the Integra HRAIM, which is the CT-compatible intubation Head Ring Assembly. When the T-handle on the HRAIM intubation hoop is tightened, the T-handle could stop in a vertical position that prevents the CRWPRECISE and certain models of the CRW-ASL from seating properly on the head ring although it appears to be fully seated | DESIGN: Device Design | The firm, Integra, sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" dated August 10, 2012 to US & Canada consignees/customers via traceable courier service and Non-US & Non-Canada consignees/customers - through written notification delivered either by traceable courier, email, facsimile, or through a directed visit by an Integra Sales Representative. The notification describes the product, problem and actions to be taken. The Integra sales representative will replace the screws according to the attached rework instructions. If for some reason it is not possible to replace the screws during the visit, the sales representative will leave the screws and customers can then replace them by following the attached rework instructions and discard the old screws immediately after replacement. The customers were instructed to complete and return the Product Recall Acknowledgement Form via the sales representatives visit; fax at 1-609-750-7999 or email to: csnirecalls@integralife.com and include the number of devices to be identified and replaced . Note:The T -Handle screws that are replaced will be disposed of at the consignee's/customers location.Please feel free to contact our service hotline at 888-772-7378 should you have any additional questions. |  | 188 | 188 units | Worldwide distribution: USA (Nationwide) and countries of: Argentina, Australia, Belgium, Canada, China, Czec, Denmark, France, Hong Kong, Hungary, Italy, Japan, Malaysia, Philippines, Singapore, Spain, South Africa, and UK. | Open |  |  | 363 | Not_Computer
Z-2414-2012 | 62765 | prosthesis, hip, constrained, cemented or uncemented, metal/polymer | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | S-RomÂŽ Total Hip System Poly LinerPackaging: The insert is placed into an inner foil pouch and vacuum sealed; a ring is placed into a Tyvek/PET pouch and sealed; the foil pouch and Tyvek/PET pouch are placed into a nylon outer pouch and sealed. This is then placed into a folding carton, labeled, and shrink-wrapped.Description: Total Hip System Poly DialÂŽ Liner is part of the S-ROMÂŽ Total Hip System. The Liner is fitted inside the acetabular shell; and is the bearing surface for the Head on the S-ROMÂŽ stem. The S-ROM liner is used for total hip arthroplasty is intended to provide increased patient mobility and reduce pain by replacing the damaged hip joint articulation in patients where there is evidence of sufficient sound bone to seat and support the components. Intended Use: Total Hip System Poly Dial Liner is part of the S-Rom Total Hip System. The Liner is fitted inside the acetabular shell; and is the bearing surface for the Head on the S-RomÂŽ stem. The S Rom liner is used for total hip arthroplasty is intended to provide increased patient mobility and reduce pain by replacing the damaged hip joint articulation in patients where there is evidence of sufficient sound bone to seat and support the components. | SRom Total Hip System Poly Liner | 2 | September 20, 2012 | 2012 | DePuy Orthopaedics, Inc. | Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of four lots of the S-ROM" liner because the Carton label has an Incorrect Color Code Series Designation. We have received one complaint regarding this issue. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | DePuy sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated August 22, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Reconciliation Form was included in the letter for customers to complete and return via fax to 574-371-4939. Contact your local Depuy Orthopaedics sales representative for questions related to this recall. |  | 12 | 12 liners | Nationwide Distribution including the states of MA, TX, OH, CA, and NC. | Terminated | October 09, 2013 | 384 | 12 | Not_Computer
Z-2442-2012 | 62927 | catheters, transluminal coronary angioplasty, percutaneous | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Sleek ***OTW ***Ref ***426-1501X***Inflated Balloon Dia/mm 1.5 Inflated Balloon Length/mm 15 Usable Catheter Length/cm 150***Lot 50026688***2013-09***Cordis***PTA Dilatation Catheter***Cordis***Johnson Johnson Company***Sterile***EO***Method of Sterilization: ETO***ClearStream***Clearstream Technologies, Ltd.***Moyne Upper, Enniscorthy Co Wexford, Ireland***Phone: +353-53-92-37111***Fax: +353-53-92-37-100***Made in Ireland***Distributed by: Cordis Cashel***Cahir Road, Cashel, Co. Tipperary, Ireland***".The SLEEK OTW Percutaneous Transluminal Angioplasty (PTA) Peripheral Catheter Family is a non-reusable semi-compliant coaxial design catheter with a balloon mounted on its distal tip. The hub/"Y" connector consists of a through lumen, allowing the catheter to track over a guidewire, and a balloon port, used to inflate the balloon. The catheter shaft is small beneath the balloon achieve low profile. The distal tip is further tapered to accept the appropriate guidewire. Platinum iridium bands serve to locate the balloon under fluoroscopy. A 0.014" (0.356 mm) guidewire is recommended for use with the SLEEK OTW Catheter. | Sleek OTW Ref 4261501XInflated Balloon | 2 | September 21, 2012 | 2012 | Cordis Corporation | On 1//20/2011 Cordis Corporation , Miami Lakes, FL initiated a recall of their Sleek OTW PTA Dilatation Catheter Model #4261201X Lot #50027149 and Model #4261501X Lot #50026688. Product is incorrectly labeled. | PRODUCTION CONTROLS: Labeling Mix-Ups | Cordis (Johnson & Johnson Company) sent an Urgent Medical Device Field Corrective Action Notice dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review, sign, and return the enclosed Acknowledgment Form directly to Cordis at the fax number on the form, or to their local sales representative. Pass the revised labeling notice to all who need to be aware within their facility or organization. Also, customers were asked to maintain awareness of this notice until all affected product in their facility has been consumed. In addition to their local Cordis representative, they may contact the normal HCS Customer Service number at 1-800-327-7714, option 1, option 3.For questions regarding this recall call 786-313-2365. |  | 1 | 1 unit | Worldwide Distribution - USA and Germany | Terminated | November 26, 2013 | 431 | 6 | Not_Computer
Z-2445-2012 | 62440 | antiserum, cf, varicella-zoster | GQX | Varicella-zoster virus serological reagents. | Microbiology | Microbiology | 510(k) | Trinity Biotech, Captia VZV IgG. Product code 2325600: 96 Tests; product code: 2325601: 480 tests. For in vitro diagnostic use. Intended for the detection and quantitative determination of IgG antibody to VZV in human sera. | Captia VZV IgG | 3 | September 24, 2012 | 2012 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | The Positive Control for the Captia Varicella-Zoster Virus (VZV) IgG Kit has been reported as running high out of its assigned ISR range. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Trinity Biotech sent a "PRODUCT RECALL LETTER" dated May 29, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Fax Back Form was included in the letter for customers to complete and return via fax to (716) 488-1990. Customers were instructed to discard any of the affected product. Contact the Trinity Biotech Help Desk at 800-325-3424, option 2 for questions regarding this recall. | Contact the recalling firm for information | 810 | 810 kits | Nationwide Distribution-including Puerto Rico and the states of AL, CA, IL, IN, MA, MD, NJ, OH, PA, TN, TX. | Open |  |  | 810 | Not_Computer
Z-0573-2 | 3529 | [6 empty] | Brand name: Olympus Bronchoscopes15 Model Numbers: BF-40 (Z-0573-2), BF-P40 (Z-0574-2), BF-1T40 (Z-0575-2), BF-3C40 (Z-0576-2), BF-XP40 (Z-0577-2), BF-XT40 (Z-0578-2), BF-240 (Z-0579-2), BF-P240 (Z-0580-2), BF-1T240 (Z-0581-2), BF-6C240 (Z-0582-2), BF-160 (Z-0583-2), BF-P160 (Z-0584-2), BF-1T160 (Z-0585-2), BF-3C160 (Z-0586-2), BF-XT160 (Z-0587-2).Common name: Video BronchoscopeClassification name: Bronchoscope (flexible or rigid) BF-40, BF-P40, BF-1T40, BF-3C40, BF-XP40, BF-XT40, BF-240, BF-P240, BF-1T240, BF-6C240, BF-160, BF-P160, BF-1T160, BF-3C160, BF-XT160. |  | 1 | September 25, 2012 | 2012 | Olympus America Inc. | Loosened bronchoscope biopsy channel port & microbial contamination of the port. |  | Please be aware that this is not a new recall. The firm took action in 2001; but, due to changes in FDA's recall reporting system this recall is now being posted by the Agency. |  | 5135 | 5135 units (revised final #): 4663 units (revised final #) in USA, 472 units to Canada & South America. | Medical facilities nationwide & Puerto Rico. The devices were also distributed to U.S. Government/military accounts and 53 foreign accounts. Distribution was provided to CDRH & OE on 3/4/02.Foreign distribution: Canada, Mexico, Dominican Republic, Argentina, Brazil, Panama, Chile, Peru, Ecuador, Colombia, Venezuela, Paraguay, Uruguay, Costa Rica, El Salvador. | Open |  |  | 5135 | Not_Computer
Z-2451-2012 | 63044 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray Systems, (Brilliance 6, Brilliance 16, MX8000 Dual v. Exp, GEMINI Dual), Philips Healthcare Systems, Cleveland, OH.The Brilliance 6, 16 and MX8000 Dual v. Exp are Computed Tomography X-Ray Systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. GEMINI Dual is an imaging system that combines positron emission tomography (PET) and X-ray computed tomography. The GEMINI produces attenuation and non-attenuation corrected images of the distribution of PET radiopharmaceuticals in the head, body and total body as well as X-ray transmission images of these areas. The PET and CT images are registered and displayed in a fused format to provide combined PET and anatomical data at different angles for interpretation by trained health professionals. The PET and CT portions of the system can be used either as an integrated system or as a stand-alone PET or CT system. GEMINI can provide CT data suitable for use in attenuation correction. | Brilliance 6, Brilliance 16, MX8000 Dual v. Exp, GEMINI Dual | 2 | September 25, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare sent an Urgent Medical Device Correction letter dated August 22, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that no action needed to be taken at that time b the customer or user. A Philips Field Service Engineer would schule a time to inspect and repair all the affected systems within approximately 6 months. For further information customers were instructed to contact their local Philips representative or contact the Philips Healthcare Customer Care Solution Center at 1-800-722-9377 and follow the recorded menu options toreach a Customer Solutions Engineer.For questions regarding this recall call 440-483-7000. |  | 112 | 112 Units | Worldwide Distribution - USA including AL, AZ, CA, CO, FL, GA, IL, KS, KY, MI, MS, NJ, NM, NV, NY, OH, TX, WA & WI and internationally to: Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, India, Italy, Japan, Jordan, Mexico, Saudi Arabia, South Korea, Spain, Thailand, Turkey & United Kingdom. | Terminated | December 24, 2013 | 455 | 112 | Not_Computer
Z-2455-2012 | 63009 | light source, endoscope, xenon arc | GCT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ***REF LS7600***Qty 1*** EC REP*** ConMed Linvatec Xenon Light Source | Light Source, Endoscopic, Xenon Arc | 2 | September 25, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec, Largo, FL initiated a recall of their products LS7700 Xenon Universal Light Source and LS7600 Xenon Light Source are intended to be used with an endoscope to provide illumination during endoscopic procedures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Linvatec Corporation sent an "URGENT Medical Device Recall Notification letter dated May 20, 2011 to its customers via FedEx service. Firm advised their suppliers of the Light Sources potentially has unapproved components and does not meet the finished product requirement. The distributors and end user level are advised if any affected product is in their possession can be used until returned for advance exchange. The customers were instructed to: 1) locate the identified product by serial number at your facility and complete the Reply Form attached; 2) Contact ConMed Linvatec Customer Service (800-237-0169) to receive SR Number Advance Exchange affected product only. Once Customer Service authorizes the SR Number your Advance Exchange Light Source will be shipped to your facility; 3) Once replacement unit is received, utilize the packaging material to return the recalled Light Source to ConMed Linvatec using the enclosed FedEx Airway Bill. Mail or FAX copy of Reply Form to ConMed Linvatec; 4) Complete and return a copy of the Reply Form. For questions contact your local ConMed Linvatec office or email Custseryl@linvatec.com, Customer Service at 800-237-0169 or fax to 727-319-5701. | No consumer action necessary | 1 | 1 | Worldwide Distribution - US Nationwide and the countries of Canada, DK, HK, IN, IT, KR, MY, RU, CO, VN, ZA. | Terminated | September 26, 2012 | 1 | 58 | Not_Computer
Z-2450-2012 | 63092 | catheter, angioplasty, peripheral, transluminal | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Armada 35/ Armada 35 LL PTA Catheters. The device is intended for dilatation of lesions in the vasculature. | Armada Catheters | 2 | September 25, 2012 | 2012 | Abbott Vascular | Abbott Vascular is recalling the Armada 35 and Armada 35 LL PTA Catheters because they have discovered that some devices may exhibit difficulty inflating and/or deflating. | DESIGN: Process Design | An "Urgent Device Recall" letter dated 8/20/2012 was sent to all customers who purchased the Armada 35 and Armada 35 LL PTA Catheters. The letter informed the customers of the problem identified and the action to be taken. A list of lot numbers and an effectiveness check form was sent to customers with the customer notification letter. Customers were instructed to call Abbott Vascular Customer Service at (800) 227-9902. |  | 61799 | 61,799 units | Nationwide(US) Distribution and Puerto Rico. | Terminated | July 15, 2013 | 293 | 61799 | Not_Computer
Z-2446-2012 | 63100 | catheter, embolectomy | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Embolectomy Catheters (Models A44XX, A4FXX, CE0XXXST, and CE0XXX):Model Numbers & Description:A4402, 3F-40CM, SYNTEL RT-EMB;A4403, 3F-80CM, SYNTEL RT-EMB;A4404, 4F-40CM, SYNTEL RT-EMB;A4405, 4F-80CM, SYNTEL RT-EMB;A4406, 5F-80CM, SYNTEL RT-EMB;A4407, 6F-80CM, SYNTEL RT-EMB;A4408, 7F-80CM, SYNTEL RT-EMB;A4F00, 2F-60CM (PREM) SYNTEL CATHETER;A4F01, 2F-80CM, (PREM) SYNTEL CATHETER;A4F02, 3F-40CM, (PREM) SYNTEL CATHETER;A4F03, 3F-80CM, (PREM) SYNTEL CATHETER;A4F04, 4F-40CM, (PREM) SYNTEL CATHETER;A4F05, 4F-80CM, (PREM) SYNTEL CATHETER;A4F06, 5F-80CM, (PREM) SYNTEL CATHETER;A4F07, 6F-80CM, (PREM) SYNTEL CATHETER;A4F08, 7F-80CM, (PREM) SYNTEL CATHETER;CE0260ST, L2F-60CM PREM SYNTEL CATH;CE0280ST, L2F-80CM PREM SYNTEL CATH;CE0340ST, 3F-40cm (PREM) SYNTEL CATHETER;CE0380ST, 3F-80CM PREM SYNTEL CATH;CE0440ST, 4F-40CM PREM SYNTEL CATH;CE0480ST, 4F-80CM PREM SYNTEL CATH;CE0580ST, 5F-80CM PREM SYNTEL CATH;CE0680ST, 6F-80cm (PREM) SYNTEL CATHETER;CE0780ST, 7F-80CM PREM SYNTEL CATH;CE0340, 3F-40cm, SYNTEL RT-EMB;CE0380, 3F-80cm, SYNTEL RT-EMB;CE0440, 4F-40cm, SYNTEL RT-EMB;CE0480, 4F-80cm, SYNTEL RT-EMB;CE0580, 5F-80cm, SYNTEL RT-EMB;CE0680, 6F-80cm, SYNTEL RT-EMB;CE0780, 7F-80cm SYNTEL RT-EMB.BARD Embolectomy Catheters are indicated for the removal of emboli and thrombi from the peripheral arterial system. These devices consist of a catheter body with a latex-free balloon on the distal end and an inflation hub on the proximal end. | Embolectomy Catheters | 2 | September 25, 2012 | 2012 | Applied Medical Resources Corp | Applied Medical is recalling specific lot numbers of its Vascular Catheters because they may have a potential for packaging particulate matter to reside on the product. | PRODUCTION CONTROLS: Process Control | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to complete the attached Recall Notification Confirmation Form and to return it to the firm via e-mail to recall60476587@appliedmedical.com or fax to (949) 713-8832. For product return questions, customers were instructed to contact Sales Operations Administrator at (949) 713-8652.For questions regarding this recall contact the firm at (949) 713-8041. | Contact the recalling firm for information | 40238 | 40,238 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Germany, Ecuador, Estonia, Spain, France, United Kingdom, Hong Kong, Ireland, Israel, Italy, Republic of Korea, Libyan Arab Jamahiriya, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Taiwan, and South Africa. | Terminated | March 22, 2013 | 178 | 40238 | Not_Computer
Z-2454-2012 | 63156 | sealant,polymerizing | NBE |  |  | General & Plastic Surgery |  | Progel Pleural Air Leak Sealant, Model #: PGPS002Lot #: 111202-002Product Usage:The Progel Pleural Air Leak Sealant is indicated for application to visceral pleura during an open thoracotomy after standard visceral pleural closure with, for example, sutures or staples, of visible air leaks (;::2mm) incurred during open resection of lung parenchyma. | Progel Pleural Air Leak Sealant | 2 | September 25, 2012 | 2012 | Neomend Inc | Product was distributed with an incorrect expiration date. Product was labeled with anexpiration date of 2013-04 and should have been labeled with 2012-09. | EXPIRATION DATING: Incorrect or No Expiration Date | Neomend sent a Product Recall Notice letter dated August 13, 2012 to all affected customers. The notification letter identified the product affected, problem and actions to be taken. Customers were instructed to check their stock supply and ask the Central Supply Department to determine whether they have any unused product from the affected lot. If you do have unused units of this product follow the instructions provided. If you do not have any product with the affected lot number complete and return the enclosedTrackingVerification Form. For questions contact your Neomend sales representative. | Contact the recalling firm for information | 95 | 95 units | US Nationwide Distribution | Terminated | September 25, 2012 | 0 | 95 | Not_Computer
Z-2464-2012 | 63197 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Skyra, Verio, Spectra and Biograph MMRNuclear Magnetic Resonance Imaging System and Tomographic Imager | Siemens Magnetom Skyra, Verio, Spectra and Biograph MMR | 2 | September 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | When positioning the Flex Large 4 Coil (part No. 8625761) off-center in left-right direction as it is used for example for hip and shoulder imaging and, in addition at an off-center position with respect to the head-foot direction, the coil can heat up in the area of the electronic housing (white plastic lids). | DESIGN: Device Design | Siemens Healthcare sent a Customer Safety Advisory Notice dated September 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included updated instruction to all customers.For questions regarding this recall call 610-219-6300. |  | 370 | 370 | Nationwide Distribution - USA including Washington, DC and AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI and Puerto Rico. | Terminated | January 14, 2014 | 475 | 370 | Not_Computer
Z-2460-2012 | 63091 | colorimetry, acetaminophen | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Alere Triage TOX Drug Screen 9 Panel, P/N: 94402.The Alere Triage TOX Drug Screen is a fluorescence immunoassay intended to be used with the Alere Triage Meters for the qualitative determination of the presence of drug and/or the major metabolites above the threshold concentrations of up to 10 distinct drug classes, including assays for acetaminophen/paracetamol, amphetamines, methamphetamines, barbiturates, benzodiazepines, cocaine, opiates, phencyclidine, THC and tricyclic antidepressants in urine. The acetaminophen/paracetamol assay will yield positive results when acetaminophen/paracetamol is ingested at or above therapeutic doses. | Alere Triage TOX Drug Screen 9 Panel | 2 | September 26, 2012 | 2012 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage TOX Drug Screen because there is a possibility that after installing the code chip provided with the affected lots: AMP, BAR, and PCP, the threshold ranges may not be consistently set correctly in all meters. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Alere sent a 'NOTICE OF CORRECTION" letter dated July 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers with questions were instructed to contact Technical Services at (877) 308-8287. Customers were instructed to complete the Customer/Distributor Verification Form Notice of Correction even if they do not have any involved product and fax back to Technical Services at (858) 805-8457 or scan and email to responses.ts@alere.com. |  | 591 | 591 kits | Nationwide Distribution. | Terminated | February 01, 2013 | 128 | 591 | Not_Computer
Z-2458-2012 | 62711 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | The MAYFIELDÂŽ Horseshoe Gel Pads are used in conjunction with the MAYFIELDÂŽ Base Units and MAYFIELDÂŽ Swivel Adaptors for neurosurgeries in the prone or supine positions. Those devices provide cranial stability and immobility for various craniotomies and spinal procedures.The Gel Pads are supplied to users in labeled plastic bags, 1 per bag.Product Usage:The MAYFIELDÂŽ Horseshoe Gel Pads are used in conjunction with the MAYFIELDÂŽ Base Units and MAYFIELDÂŽ Swivel Adaptors for neurosurgeries in the prone or supine positions. They are plastic gel filled devices that provide cushioning between the head and other metallic part of the MAYFIELDÂŽ system. The Left and Right Pads are mirror images. | MAYFIELD Horseshoe Gel Pads Right Gel Pad | 2 | September 26, 2012 | 2012 | Integra Limited | An investigation of customer complaints identified that for two product lots, Left Horseshoe Gel Pads were packaged as a Right Horseshoe Gel Pads and vice versa. None of the complaints resulted in a report of patient injuries or an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that there is no potential harm to patients, onl | PRODUCTION CONTROLS: Packaging Process Control | Integra Lifesciences sent an "URGENTURGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 27, 2012 to all affected customers. The letter identifies the affected products, problem and actions to be taken by the customers. Customers were instructed to check their inventory for affected products and quarantine them for shipment back to Integra. Customers were asked to complete the Recall Acknowledgement and Return Forms ASAP whether or not they intended on returning any recalled product. Contact Integra Customer Service at 1-877-444-1114, Option 3 for Technical Support. | No consumer action necessary | 94002 | 18 from Lot 093984 for Right Gel Pads | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DC, GA, IA, IL, IN, NC, and NY and the countries of: Canada and Belgium. | Terminated | February 20, 2013 | 147 | 18 | Not_Computer
Z-2459-2012 | 63141 | catheter, steerable | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Z Flex 270, Steerable Sheath, Sterile EO, Rx only, Distributed by Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, Manufactured by Greathbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441. To facilitate transvenous introduction of diagnostic / therapeutic catheters into the vasculature and into the chambers of the heart. | Boston Scientific, Z Flex 270, Steerable Sheath | 2 | September 26, 2012 | 2012 | Greatbatch Medical | During the investigation of four returned Z Flex 270 Steerable Sheaths that were used with cardiac cryoablation balloon catheters, the internal PTFE sheath liners were found to display varying degrees of damage that could pose an embolic risk to a patient. | DESIGN: Device Design | Greatbatch Medical sent a "SAFETY ALERT-Z-FLEX-270 STEERABLE SHEATH" letter dated September 6, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. The letter advised the customer to communicate with their customers to assure the that the recalled sheath model is not used in conjunction with cryoablation catheters. Contact the firm at 763-951-8214 for questions regarding this issue. | No consumer action necessary | 787 | 787 | Distributed in MA. | Open |  |  | 787 | Not_Computer
Z-2461-2012 | 62965 | chair and table, medical | KMN | Medical chair and table. | General Hospital | General Hospital | 510(K) Exempt | InvacareÂŽ I-Fit" Shower Chair (Models 9780, 9781, 9781-1), Invacare Corporation, Elyria, OH 44035.These shower chairs are designed for individuals who require seating support when bathing or showering and provide comfort and safety in the bathtub. | IFit" Shower Chair | 2 | September 26, 2012 | 2012 | Invacare Corporation | On August 6, 2012, Invacare voluntarily recalled the InvacareÂż I-Fit" Shower Chair (Models 9780, 9781, 9781-1) based on the return, quality and adverse event data received. | DESIGN: Device Design | Invacare sent an "URGENT: RECALL INFORMATION" letter dated August 31, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of any affected product and to call the Recall Support Line at 877-496-5041 for assistance. |  | 21168 | 21,168 chairs | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of AUSTRALIA, CANADA, GUAM, ISRAEL, NEW ZEALAND, THAILAND, TRINIDAD AND TOBAGO. | Terminated | February 11, 2014 | 503 | 21168 | Not_Computer
Z-2470-2012 | 62419 | cannula, trocar, ophthalmic | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | This voluntary medical device correction is related to Alcon products that present a 23G nonvalved trocar cannula and infusion cannula together. These items are available as components within Accurus 23G TotalPLUS Pak, Accurus 23G Combined Procedure Packs (0.9mm and 1.1mm), Constellation 23G Total PLUS Pak, Constellation 23G Combined Pack (0.9 mm and 1.1 mm) and 23G Small Parts Kit. Single use 23G non-valved trocar cannulas are used to provide internal access to the ocular cavity during ophthalmic surgery. Single use infusion cannulas are used for the infusion of fluid or air during ophthalmic surgery to maintain intraocular pressure during posterior segment surgeries. | Alcon 23G NonValved Trocar Cannula and Infusion Cannula | 2 | September 27, 2012 | 2012 | Alcon Research, LTD. | This medical device correction is related to Alcon products that present a 23G non-valved trocar cannula and infusion cannula together because of an issue regarding the loose fit between the 23G non-valved trocar cannula and infusion cannula that could potentially cause a detachment between the non-valved trocar cannula and infusion cannula during surgery. | DESIGN: Device Design | Alcon sent an Urgent a Voluntary Medical Device Correction letter on May 9, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed NOT to use the affected product, instead customers should use a 1-Count Valved Trocar Cannula and to complete and immediately return the Medical Device Correction Response Form to Alcon via fax at: (817) 916-9087 or email to: ReplacementCannulas@alconlabs.com. For questions customers should call 1-800-862-5266. | Contact the recalling firm for information | 413 | 93,319 (30, 413 distributed domestically & 62,906 internationally) | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Greece, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | February 12, 2013 | 138 | 93319 | Not_Computer
Z-2471-2012 | 63188 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | UltraFlow HPC Flow Directed Micro Catheter, Model # 105-5065.Product Usage:The UltraFlow HPC Flow Directed Micro Catheter is intended to access peripheral and neuro vasculature for the controlled selective infusion of physician-specified therapeutic agents such as embolization materials and of diagnostic materials such as contrast media. | UltraFlow HPC Flow Directed Micro Catheter | 2 | September 27, 2012 | 2012 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The recall was initiated because ev3 Neurovascular has identified one lot of UltraFlow Flow Directed Micro Catheters and Marathon Flow Directed Micro Catheters may have been mislabeled. | PRODUCTION CONTROLS: Error in Labeling | Micro Therapeutics Inc, Dba Ev3 Neurovascular sent a recall notification letter dated September 23, 2011 to all their affected customers. The letter identified the affected product, problem, and actions to be taken. The notification letter also provides the customer (physician) with appropriate information regarding the affected products. The letter instructs customer to stop using the affected product, segregate from inventory and return for replacement with conforming product. The notice requests the customer to complete and return the reply form indicating use or return of devices in the affected lot(s). Customers with questions regarding the recall letter were instructed to contact Kim Salceda at (949) 680-1305. |  | 36 | 36 units | Worldwide Distribution - US Nationwide including the states: MN, FL, OR, TX and the countries of: Brazil, Belgium, Czech Republic, Finlind, France, Germany, Netherlands, Poland, and United Kingdom. | Terminated | September 28, 2012 | 1 | 36 | Not_Computer
Z-2467-2012 | 62948 | lubricant, personal | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY LIQUIBEADS Vaginal Moisturizer; Distributed By: McNeil-PPC, Inc. Skillman, NJ 08558-9418; UPC 380041087704; 380040087958.Personal lubricant | KY LIQUIBEADS Vaginal Moisturizer | 2 | September 27, 2012 | 2012 | Mc-NEIL-PPC, Inc. | Johnson & Johnson Healthcare Products (JJHP) is voluntarily recalling this product to address FDA concerns that the product should have been cleared under a separate medical device application. This voluntary recall is not being conducted on the basis of any consumer safety or product performance issues related to K-Y LIQUIBEADS. JJHP received a Warning Letter on May 23, 2012 indicating that thi |  | Johnson & Johnson sent Recall Letters/Response forms dated June 29, 2012 on July 2, 2012. The recall notification included affected lots and instructions for wholesale distributors and retailers that handle the product. The recalled product is being handled by Stericycle. The notification included phone and fax numbers in the event assistance is required. Customers were asked to complete the enclosed business Reply Form and return immediately; identify all warehouse/wholesale inventories of products and return immediately per the attached "Return Authorization Form." For any questions regarding the form, customers were asked to call Stericycle at 888-406-9309. Customers could also fax questions to 877-877-0318. Once customers have collected all returns, they should send back immediately per the enclosed instructions.For questions regarding this recall call 908-874-1000. |  | 818712 | 818,712 | Nationwide Distribution | Terminated | September 27, 2013 | 365 | 818712 | Not_Computer
Z-2465-2012 | 63119 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost-Mobile Detector Holder for Digital Diagnost (Single Detector, Dual Detector, Release 3)System codes: 712020, 712022, 712025This system is used for making X-ray exposures for diagnostics. | System, Xray Stationary | 2 | September 27, 2012 | 2012 | Philips Healthcare Inc. | The Wireless Portable Detector (WPD) is not fixed securely inside the Mobile Detector Holder and may falldown | DESIGN: Device Design | Philips Healthcare sent a Urgent Field Safety Notice dated August 13, 2012, to all affected customers. Notifications began the week of September 6, 2012, for US and Worldwide Customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips field service engineer will Secure the 4 countersunk screws, which are used to attach the lock mechanism, with thread locker (loctite). Further information or support concerning this issue, contact your local Philips representative. 1-800-722-9377. | Contact the recalling firm for information | 93 | 93 Units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, Chile, Finland, France, Germany, Hungary. Italy, Japan, Latvia, Netherlands, New Zealand. Oman, South Korea, Switzerland and United Kingdom. | Open |  |  | 93 | Not_Computer
Z-2469-2012 | 63107 | catheter, irrigation | GBX | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Catheter Irrigation Set labeled in part: ***24K Arthroscopy Outflow/Suction Tubing Set***This Arthroscopy Tubing Set is intended for use, in conjunction with the ConMed Linvatec 24k Pump, for pump-induced outflow and/or suction of fluid from the joint during arthroscopic procedures. | 24K Arthroscopy Outflow/Suction Tubing Set | 2 | September 27, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled their 24k Arthroscopy Outflow/Suction Tubing Set (Product Number 24k100)because there is a possibility that the product may have a breach in the Tyvek seal that could potentially compromise the sterility of the contents. | TRAINING: Employee Error | ConMed Linvatec sent a Urgent Medical Device Recall Notification via FedEx to all affected customers. This notice should be received July 5, 2011. The letter identified the product the problem and the action needed to be taken by the customer. All returned product under this removal action will be identified, segregated, and quarantined per ConMed Utica, N.Y. procedures. Returned product will be dispositioned in accordance with ConMed Utica, N.Y. procedures. This can include repackaged and resterilization with new lot number or product scrap. ConMed Utica, N.Y. will maintain control of all individual lot numbers returned and dispositioned.Firm is striving for 1005 effectiveness. If a Reply form is not received within approximately 30 days of the first mailing, ConMed Linvatec will make two additional contacts by mail approximately 30-35 days apart.Should you have any additional questions please contact our Customer Service at 800-237-0169 or fax at 727-319-5701. | No consumer action necessary | 30064 | 30064 units | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA ,WI and Hawaii., and the countries of Australia Canada, China, Switzerland, Germany, Czech Republic, Denmark, Egypt, Spain, France, South Africa, United Kingdom, Israel, Italy, Netherlands, New Zealand, Poland, Portugal, Russia and Saudi Arabia. | Terminated | July 25, 2013 | 301 | 30064 | Not_Computer
Z-2493-2012 | 63125 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cemented Tibial DrillProduct Usage:The cemented tibial drill is used to prepare the tibial bone prior to the impaction of the applicable stemmed tibial broach. The cemented tibial drill has engraved lines that are size specific to guide the user as to the depth to insert the drill into the tibial bone. | Persona Personalized Knee System | 2 | September 28, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a recall of the Cemented Tibial Drill for the Zimmer Persona Personalized Knee System due tothree reports being received as of August 29. 2012 for tibial cortex perforation during preparation of the tibial bone.The investigation into the reports that have been received found that the perforations are occurring while using theCemented Tibial Drill. The perforations have o | DESIGN: Device Design | Zimmer sent an Urgent Medical Device Recall dated August 2012 to distributors via electronic mail and Hospital risk managers and surgeons had notifications hand delivered by sales representatives. The recall notice describes the reason for the recall, emphasizes surgical technique to be used, includes instructions for delivering notifications to surgeons and risk managers, instructions for replacing the affected product, and instructions for returning the affected product. Customers were instructed to locate the affected product and contact their sales representative for replacement. Consignees are asked to fill out the Certificate of Delivery, scan it and send back per the instruction provided. |  | 60 | 60 | Worldwide Distribution - US Nationwide including the states of: AZ, CA, CO,FL, MN, IL, NJ, NY, OH, OR and the countries of Austrailia, Utd. Arab Emir., France, Germany, India, and Italy. | Open |  |  | 60 | Not_Computer
Z-2494-2012 | 63165 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost with patient carriage Model: 712050 with serial number range from 00 000 00 to 09 000 543.Product Usage:Stationary radiographic System | Philips Healthcare | 2 | September 28, 2012 | 2012 | Philips Healthcare Inc. | Hook does not securely hold the footplate in vertical position, causing the footplate to disengage and may fall on patient or operator | DESIGN: Device Design | Philips Healthcare sent an Urgent Field Safety Notice letter dated August 15, 2012 to all affected customers. The letter identified the affected product, problem and the actions to be taken to avoid or minimize the occurrence of the issue. Customers were advised to follow the instruction for Use and that Philips plans to replace the old hook with one that has a new design. For questions contact your local Philips representative 1-800-722-9377. |  | 536 | 536 units | Worldwide Distribution - US Nationwide and the countries of :Australia, Austria, Bahrain, Belgium, Canada, Chile, China, Czech Republic, Denmark, Finland, , France, Germany, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Qatar, Romania, Russia, Saudi Arabia, Sweden, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan. Thailand, Turkey,United Arab Emirates and United Kingdom. | Terminated | July 26, 2013 | 301 | 536 | Not_Computer
Z-2485-2012 | 63073 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Optilock 2.7 mm Screw Inserter.Intended Use: The OptiLock Upper Extremity Plating System is intended for fixation of fractures and osteotomies. The System is intended for fractures and fracture dislocations, osteotomies and non-unions of the proximal humerus, particularly in osteopenic bone. The System is intended for fixation of fractures, osteotomies and non-unions of the olecranon, humerus, radius, ulna, particularly in osteopenic bone. The OptiLock VL Distal Radius Plating System is intended for use in distal radius and ulna (and other small bones), including use in osteopenic bone. | Optilock | 2 | September 28, 2012 | 2012 | Biomet, Inc. | Field complaints indicate that recent shipments of the 14-400795 and 36505 drivers do not properly mate("lock") into some of the specified screw heads. Drivers near MMC (Maximum Material Condition) or that are misshapen may not properly mate with screws at MMC. Design dimensions and tolerances of the drivers are not properly specified to prevent an incorrect fit between the drivers and the screw | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE'' dated August 23, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. A Fax Back Response Form was included for customers to complete and return via fax to 574-372-1683. Contact the firm at 574-372-1570 for questions related to this recall. | Contact the recalling firm for information | 268 | 268 units total | Worldwide Distribution-USA (nationwide) including the states of AL, GA, MT, TA, AZ, WA, CA, TA, IN, NY, NJ, VA, MA, MI, MD, PA, KY, NC, UT, NH, FL, OK, CO, LA, OH, MO and SD, and the countries of Japan, Netherlands, Australia, and Costa Rica. | Terminated | August 27, 2013 | 333 | 631 | Not_Computer
Z-2487-2012 | 62594 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Restoration Modular Cone Body TrialsStryker Ireland Carrigtwohill Industrial Estate Carrigtwohill County Cork, Ireland.The Cone Body Trials simulate the respective Cone Body implants and mate with the stem trials to assist the surgeon in assessing the fit of the proposed Implants. The cone body trials is an assembly of the Body and 8 mm Hex Locking Bolt. The trial stem is attached to the trial Cone body via the distal thread of the Locking bolt. | Stryker Orthopaedics | 2 | September 28, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of a burr, or a metal shaving, observed in the internal threaded portion of the Cone Body Trial. The typical size of the shaving observed was 1-2 mm in length and 0.037 mm squared in cross section. | DESIGN: Device Design | Stryker Orthopaedics sent a Urgent Product Field Action - Product Inspection Branch Acknowledgement Form on April 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All inspections must be completed immediately as per the provided Product Correction Bulletin. Please document the results of your inspection on the Tracking Log, indicating conforming devices as "Not Affected" and non-conforming devices as "To be returned." If the devices listed on your tracking log have previously been returned or are otherwise not on hand, indicate the appropriate disposition following our usual process.Please assist us in meeting our regulatory obligation by faxing back the attached Product Acknowledgment Form and Tracking Log within 10 business days to 855-251-3635. Any further questions please call 201-831-5523. |  | 993 | 10, 983 units | Nationwide distribution | Terminated | June 13, 2013 | 258 | 10983 | Not_Computer
Z-2480-2012 | 62110 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | 00780302601 MIS 12/14 Slotted Femoral Head Provisional -3.5X2600780302602 MIS 12/14 Slotted Femoral Head Provisional +0X2600780302603 MIS 12/14 Slotted Femoral Head Provisional +3.5X2600780302604 MIS 12/14 Slotted Femoral Head Provisional +7X2600780302605 MIS 12/14 Slotted Femoral Head Provisional +10.5X26The provisional heads replicate final femoral implant geometry and are used to trial range of motion, joint stability, and determine appropriate femoral head size and offset during total hip arthroplasty. | Slotted Femoral Head Provisional | 2 | September 28, 2012 | 2012 | Zimmer, Inc. | Complaints have been received reporting sub-optimal fit between provisional femoral heads (VerSys and MIS) and the mating components (ie stems and provisional femoral necks), in some cases alleging the fit was too loose and in other cases, alleging the fit was too tight. Some reports of intraoperative dislodging of the provisional head have also been received. 13 of the complaints allege that th | DESIGN: Device Design | Zimmer, Inc. sent an E-mail letter dated April 21, 2011, to all affected customers in the US. A June 2011 E-mail was sent to all customers outside the US. The letters requested Inventory & Operations Managers to check product for proper fit /function, complete the return form with worn provisional heads to Zimmer Inc, Attention Product Service, 1777 West Center St. Warsaw, IN 46580.For questions regarding this recall call 574-372-4753. |  | 231641 | 231,641 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Mexico, South America, Europe, Puerto Rico, China, and Korea. | Open |  |  | 231641 | Not_Computer
Z-0001-2013 | 62773 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | MIRS for MATRIX Spine System. Minimally Invasive Reduction Screw (MIRS) for percutaneous instrumentation of the thoracolumbar spine.The Synthes USS are non-cervial spinal fixation devices intended for posterior pedicle screw fixation (T1-S2/ilium), posterior hook fixation (T1-L5), or anterolateral fixation (T8-L5). | Synthes Minimally Invasive Reduction Screw System (MIRS) | 2 | October 01, 2012 | 2012 | Synthes USA HQ, Inc. | Synthes is initiating a Medical Device Labeling Correction following a detailed review of the current labeling in response to reported complaints. Synthes has identified technique measures and has updated the labeling associated with this system. The labeling reflects updated technique recommendations and emphasizes the importance of avoiding rod mis-alignment. | DESIGN: Labeling Design | SYNTHES sent an Urgent Medical Device Labeling Correction letter dated June 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to replace the previous version of the labeling with the new labeling provided. In addition, they can access an electronic version of the labeling by accessing Product Literature at the following link: http://us.synthes.com. The new version can be identified by code J11031-B located at the bottom right corner of the back cover. To confirm their receipt and understanding of the information contained in the letter, the recall package also included a response form to be completed by consignee and returned to Synthes.by Fax: 610-251-9005 or Scan/email: FieldAction@Synthes.com.Customers with questions were asked to call 1-800-620-7025, ext 5403 or contact their Synthes Spine Sales Consultant. For questions regarding this recall call 610-719-5377.The firm initiated this recall action on June 18, 2012 by sending an Urgent Medical Device Labeling Correction letter to all accounts. The updated labeling was Included with the notification letter. | Contact the recalling firm for information | 1602 | 1602 | Nationwide Distribution including CA, FL, ID, IN, MD, MI, MT, LA, NY, OH, PA, TX, VA and WA. | Open |  |  | 1602 | Not_Computer
Z-0003-2013 | 63152 | system, test, rheumatoid factor | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, PN 447070The IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) reagent, in conjunction with Beckman Cal 5 Plus, is intended for use in the quantitative determination of human rheumatoid factor concentrations in human serum and plasma samples by rate nephelometry. This assay is designed for use with the IMMAGE Immunochemistry System. | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent | 2 | October 01, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of QC shifts, lot-to-lot variation, and an increase in the percent of patient samples recovering | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated January 27, 2012 with an attached PCA Response Form to all customers who purchased the IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use of IMMAGE RF reagent lots M012376 and M101865; call Order Assistance at (800) 526-3821 in the United States, (800) 463-7828 in Canada or contact your local Beckman Coulter representative for replacements; and complete and return the enclosed FAX BACK RESPONSE FORM within 10 days via mail at: Beckman Coulter, Inc., 250 S. Kraemer Blvd, Mail Stop A2.0246, Brea, CA 92822-8000; Attn: Regulatory Affairs or fax to: (714) 961-4234. Note: Beckman Coulter has added additional release testing that will assure future lots RF will meet the reference interval claim of the IMMAGE RF Instructions for Use (IFU).If you have any questions concerning this letter, contact Beckman Coulter Support Center (call Center) on the Beckman Coulter Website at: http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative. |  | 925 | 925 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | May 30, 2013 | 241 | 925 | Not_Computer
Z-0002-2013 | 63134 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COULTER PrepPlus, Part Number: 286600Product Usage: The COULTER PrepPlus, is a microprocessor-controlled pipetting and diluting system, designed for automating sample preparation or assay methods. It is capable of aspirating and dispensing liquid samples. | COULTER PrepPlus | 2 | October 01, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Coulter PrepPlus may experience intermittant reagent level sense failures where the instrument may not produce a warning that a reagent vial is low or out of reagent. | DESIGN: Device Design | Beckman Coulter sent an Urgent Product Correction letter dated January 20, 2012 to all their customers who purchased the affected product. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to complete and return the Response Form with in 10 days. Customers with questions regarding the PCA letter were instructed to contact Beckman Coulter Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter Representative. | Contact the recalling firm for information | 36 | 36 units total (15 units in the US) | Worldwide Distribution - US Nationwide and the countries of: Australia, Botswana, Canada, France, Germany, Japan, Malawi and Sweden. | Terminated | October 02, 2012 | 1 | 36 | Not_Computer
Z-0005-2013 | 61780 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | IR1200, IR1250(IR1200 Plus), 2020 and One Touch(R) Ping(R) Insulin Pump glucose management system. The device is intended to deliver insulin therapy. | IR1200, IR1250(IR1200 Plus), 2020 and One Touch(R) Ping(R) Insulin Pump | 2 | October 02, 2012 | 2012 | Animas Corporation | Wear of the keypad that may lead to damage such as tearing or peeling of the edge of the keypad from the pump. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Animas Corporation, sent an "IMPORTANT PATIENT NOTIFICATION" letter dated July 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to examine their insulin pump for wear or damage on a regular basis, as part of their routine process of inspecting their pump. If keypad is worn or damaged, the buttons may not respond as intended. If consignees/customers see any damage or have issues with the keypad, they are to contact Animas through their website at www.animas.com/contact-animas or telephone Animas Customer Technical Support at 1-855-230-7574, or the phone number provided on the back of their insulin pump. Animas carefully investigates all complaints related to the insulin pump keypad.If you have any questions contact Animas Customer Technical Support at 1-855-230-7574. |  | 236470 | 236,470 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, V, VT, WA, WI, WV, WY and Washington, DC and countries of: Austria, Australia, Canada, Czech Republic, France, Finland, Germany, Hungary, Israel, Italy, Mexico, New Zealand, Puerto Rico, Spain, and United Kingdom. | Terminated | January 15, 2014 | 470 | 236470 | Not_Computer
Z-0004-2013 | 63241 | catheter, straight | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Male, 16Fr, Straight-tip Intermittent Urinary Catheters M16. The product is distributed individually or in intermediate boxes containing 30 individual, sterile packaged catheters (PA/PE film and dialyzing paper pouches).There are 30 catheters packaged in sterile pouches that are placed into an intermediate box, 10 of these intermediate boxes packaged into one case. K072539The Cure Catheter" device, M16, is an intermittent urinary catheter intended to be used by males for the purpose of bladder drainage. It is manufactured with conventional medical grade, latex-free, biocompatible materials. The straight-tip has been designed to eliminate trauma to the urethra and is provided in 16Fr size in an easy-to-open, sterile, single-use package. The coudÂŽ-tip catheter, M16C, is designed with a curved tip that makes it easier to pass through the curvature of the prostatic urethra in men. It is also used to pass through urethral strictures. This product is prescribed for people who have difficulty or are unable to pass a straight catheter through their urethra and into their bladder. | Male, 16Fr, Straighttip Intermittent Urinary Catheters M16 | 2 | October 02, 2012 | 2012 | Cure Medical LLC | The firm initiated the product recall because labeling on some of the intermediate boxes of catheters identify the contents to be M16C, a coudÂż-tip catheters while the actual contents are M16Fr male, 16Fr, straight-tip catheters with the pouch and the case box labeled as M16Fr male, 16Fr, straight-tip catheters. | PRODUCTION CONTROLS: Labeling Mix-Ups | Cure Medical LLC sent an Urgent Medical Device Correction and Removal Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter instructed the customers to:a) Check their shelves for M16C with lot number 120425 and quarantine all affected product.b) Verify the contents (pouch label) are M16C and correspond with the intermediate box label.c) If all match, have the dealer sign and return a supplied verification record.d) If any mismatch is found, contact Cure Medical for assignment of a return goods authorization.e) Cure Medical will replace any incorrect product along with processing costs for this situation to the dealer.If affected product was further distributed customers were asked to identify those customers who received them and notify them of the problem. For questions customers were instructed to call 1-800-570-1778. | Contact the recalling firm for information | 420 | 420cases (= 4,200 boxes) | Nationwide Distribution | Terminated | February 07, 2013 | 128 | 420 | Not_Computer
Z-0008-2013 | 63253 | electrocautery, gynecologic (and accessories) | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile, single-use, disposable electrosurgical electrode holder. It connects the electrode to the foot activated electric generator. Label reads in part "***GYNEX*** Electrode Handpiece, Foot Activated*** Manufactured for GYNEX***"Gynex catalog number: 950 (Gynex REF 950, Electrode Handpiece)It connects the electrode to the food activated electric generator. Electrodes are used in gynecological procedures, in a physician's office, for cutting to remove tissue and coagulation to control bleeding by the use of high frequency current. | Electrode hand piece, foot activiated | 3 | October 03, 2012 | 2012 | Gynex Corporation | The wrong electrical connector pin was attached to the cord during assembly by the manufacturer. this electrical connector is too big to fit into the port on the foot activated generator. this makes the device unusable. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were first notified by phone with a follow up e-mail or fax. The consignees were asked to complete the response section regarding the destruction/return of the device.For questions regarding this recall call 425-882-1179. | Contact the recalling firm for information | 19 | 19 boxes (containing 12 electorde hand pieces each) | Nationwide Distribution | Terminated | October 22, 2012 | 19 | 19 | Not_Computer
Z-0007-2013 | 62966 | catheter and tip, suction | JOL | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | BriteField illuminated Suction Device;Model numbers 8Fr, 10Fr and 12Fr, Kit (included 8Fr, 10Fr and 12Fr and a sterilization tray.The Invuity Brite Field Illuminated Suction Device is intended to aspirate from the surgical site and to provide illumination from a high intensity light source. | BriteField illuminated Suction Device | 2 | October 03, 2012 | 2012 | Invuity, Inc. | Invuity is notifying customers to return all obsolete BriteField Illuminated Suction Devices from the field. It has not been distributed since March 2011 as the next generation product was introduced at that time. | OTHER/UNDETERMINED: Pending | Invuity sent E-mails to those locations thought to still be in possession of the old model of the device in July 2012. Distributors had been e-mailed in May 2012 to make sure that they had returned any of the device that was in stock.For questions regarding this recall call 415-655-2100. |  | 197 | 197 devices | Nationwide Distribution including AR, GA, SC, CA, OR, UT, and WA. | Terminated | April 23, 2013 | 202 | 197 | Not_Computer
Z-0006-2013 | 63196 | dislodger, stone, flexible | FGO | Ureteral stone dislodger. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Endoscopic retrieval device.The disposable Raptor grasping device in intended to retrieve foreign bodies or resected tissue specimens in endoscopic procedures of the upper and lower gastrointestinal tract or biliary tract. | Raptor | 2 | October 03, 2012 | 2012 | US Endoscopy Group Inc | When the outer catheter is either coiled or in a contorted configuration outside of the endoscope and the patient, and excessive force is applied to the handle, the jaws may become inoperable and may become stuck in the open condition. |  | US Endoscopy initiated a Voluntary Product Recall letter initiated November 11, 2011. The letter identified the product the problem and the action needed to be taken by the customer.US Endoscopy will provide a free replacement or grant your facility a credit toward a future purchase for every device that is returned to US Endoscopy or verified as having been used or destroyed.For further questions please call 440-639-4494 ext. 378. | Contact the recalling firm for information | 1153 | 1153 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, FL, GA, HI, IA, IL, KY, LA, MD, ME, MI, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, and WI; and to the following countries: Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Caribbean, Chile, Columbia, Croatia, Czech Republic, Denmark, Ecuador, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, India, Ireland, Israel, Italy, Jordan, Lithuania, Malta, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, and United Arab Emirates. | Terminated | July 02, 2013 | 272 | 1153 | Not_Computer
Z-0010-2013 | 63136 | catheter, urological | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 8 Fr. Bardex Pediatric All-Silicone Foley Catheter with a 3cc balloon. The product is packaged individually with 12 units per corrugated case. The product is sold sterile for single use and is used to provide continuous urinary drainage from the bladder. | 8 Fr. Bardex Pediatric AllSilicone Foley Catheter | 2 | October 04, 2012 | 2012 | C.R. Bard, Inc., Urological Division | The 8 Fr. Bardex All-Silicone Foley Catheter with 3cc balloon had an incorrect inflation volume (10 ml instead of 5 ml) printed on the valve cap. | TRAINING: Employee Error | An "Urgent: Product Recall Notification" was sent to all customers and distributors in September 2012 via Federal Express. The letter described the product issue and provided recommended actions. Customers can contact Bard at (770) 784-6471 if questions. | Contact the recalling firm for information | 4032 | 4,032 units | Nationwide distribution: USA and Puerto Rico | Terminated | March 21, 2013 | 168 | 4032 | Not_Computer
Z-0016-2013 | 63140 | radioimmunoassay, angiotensin i and renin | CIB | Angiotensin I and renin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Renin IRMA Kit, Part Number DSL25100Product Usage:The Renin IRMA; DSL25100 immunoradiometric assay kit provides materials for the quantitative measurement of Renin in plasma. This assay is intended for in vitro diagnostic use. | Active Renin IRMA Kit | 2 | October 05, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Renin IRMA Test Kit may exhibit a decrease in stability of the coated tubes. As a result, quality control and patient sample results may be falsely elevated by up to 72%. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter sent an Urgent Production Correction letter dated January 10, 2012 to all their customers who purchased the Manual Immunoassay Active Renin IRMA Test Kit. The letter provides the customers with an explanation of the problem and actions to be taken. Customers were instructed to discontinue use and discard all remaining affected inventory. Customers were asked to complete and return the enclosed Response Form with in 10 days. Customers with questions contact Customer Service Support at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633. | Contact the recalling firm for information | 216 | 216 units total (6 units in the US) | Worldwide Distribution - US Natiowide including the state of: OH, CT, CA and the state of Argentina, Austria, Belgium, Brazil, Canada, China, Columbia, Czech Republic, France, Germany, Greece, Hungary, Italy, Kuwait, Netherlands, Poland, Portugal, Slovakia, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2012 | 0 | 216 | Not_Computer
Z-0013-2013 | 63059 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 117" (3.0 m), with 150 mL Burette (Ball-Valve Drip Chamber) (Non-DEHP); a sterile Rx IV fluid pathway device; Product Usage: The device is used for the administration of fluids from a container into the patient's vascular system through a vascular access device. | Baxter Interlink System Buretrol Solution Set with 150 mL Burette NonDEHP | 1 | October 09, 2012 | 2012 | Baxter Healthcare Corp. | Baxter has determined that the ball-valve feature of the Buretrol Solution Sets may not function as expected, allowing air to flow past the valve and into the tubing at the completion of dose within the burette. |  | Baxter Healthcare Corp. sent an Urgent Product Recall letter dated September 7, 2012, via first class mail to all affected customers, informing them of the problem with the ball-valve, and instructing them to discontinue use of the affected product and to contact Baxter for instructions to return the affected product.. The customers were also requested to complete the attached customer reply form, confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Medical Information Services at Baxter at 1-800-933-0303. |  | 1995040 | 1,995,040 units | Worldwide Distribution - US Nationwide and Internationally to Canada, Costa Rica, Australia, New Zealand, Hong Kong, and the United Kingdom. | Open |  |  | 5027417 | Not_Computer
Z-0025-2013 | 63154 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | The Eon Mini Implantable Pulse Generator (IPG) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with four, eight, and 16 electrode leads. It is powered by a hermetically sealed battery within a titanium case. It uses microelectronic circuitry to generate constant current electrical stimulation. Stimulation programs can be delivered as either single stimulation or MultiStim programs depending on patients need.Product Usage:Eon Mini system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. | Eon Mini 3788 (IPG) | 2 | October 09, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Eon Mini IPGs manufactured in April 2012 could potentially exhibit a sudden, brief surge in stimulation that would be felt by the patient as uncomfortable or painful. SJMs routine testing, monitoring, and analysis of manufacturing processes identified three lots of Eon Mini IPGs that the internal battery has the potential to come into contact with the internal microcontroller board. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ST. Jude Medical sent an important Medical Device Advisory letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem and recommended actions to be taken. For questions contact your St. Jude Medical Neuromodulation Division Representative or call the Technical Services Department at 1-800-727-7846. |  | 56 | 56 units | US Nationwide Distribution | Open |  |  | 56 | Not_Computer
Z-0024-2013 | 63211 | tomography, optical coherence | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Reference materials "RETINAL OCT Analysis and Interpretation Method" shipped with iVue System (K091404) and the "Retina Poster" shipped with iVue System (K091404) and RTVue Systems (K101505); Models: iVue-100, RTVue-100; The physical Product is manufactured and distributed by Optovue, Inc. Fremont, CAThe product is an optical coherence tomography system indicated for the in vivo imaging and measurement of the retina, retinal nerve fiber layer, optic disk, and for cornea and anterior eye scans as an aid in the diagnosis and management of retinal disease. | Optovue RTVue, iVue | 2 | October 09, 2012 | 2012 | Optovue Inc. | Devices labeled for an intended use not included in the existing 510(k) or PMA. Specifically, "The Retina" poster provided with the systems, and the book "RETINAL OCT Analysis and Interpretation Method" by Bruno Lumbroso and Marco Rispoli which are not labeled within 51O(k) cleared product description. |  | Optovue Inc., sent a Voluntary Field Correction and Removal letter dated September 6, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer .Please sign and return the attached form indicating your acknowledgement that that the Affected Materials have been returned or destroyed. If you have any further questions please contact Optovue, at 1-866-344-8948. |  | 795 | 795 units | Worldwide Distribution -- USA (nationwide) and Internationally. | Terminated | March 04, 2013 | 146 | 795 | Not_Computer
Z-0020-2013 | 63174 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Newport Portable Package, Model #50000, Serial # 13591, 13592, 13604, 13606, 13614, 13662, 13663, 13664, 13726, 13731, 13733, 13736, 13737, 13738, 13740, 13768, 13770, 13771, 13772, 13784, 13787.Product Usage:Portable Analgesia Gas Machine. | Newport Portable Package | 1 | October 09, 2012 | 2012 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner. | DESIGN: Device Design | Accutron sent an Urgent Advisory Notice letter dated April 6, 2012 to all their customers who purchased the Ultra PC% Cabinet Mount Flowmeter. The letter informed the customers of the problem identified and actions to be taken. Customers were instructed to discontinue use or to continue use of the device only with extreme care until a replacement can be installed. Extreme care includes the use of the titration method for nitrous oxide administration coupled with close monitoring of both the patient and the actual gas flow through the flowtubes. For question contact Accutron at (800) 531-2221 or by email at jheadrick@accutron-inc.com. | Contact the recalling firm for information | 220 | 220 units total | US Nationwide Distribution | Terminated | February 01, 2013 | 115 | 220 | Not_Computer
Z-0027-2013 | 54725 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoHelix Part #HCS-056-32 3.8MM Cannulated Hex Driver NON STERILE MAT: 465 SSThe drivers are used to drive screws in the Mini MaxTorque Screw System, which is intended to be used to stabilize and aid in the fusion of fractures and osteotomies involving small bones of the hand, wrist, afoot, and ankle. | Mini MaxTorque HCS05632 Drivers | 2 | October 09, 2012 | 2012 | Orthohelix Surgical Designs Inc | The tips of the drivers are breaking while being used to drive the screws into patients. | DESIGN: Software Design (Manufacturing Process) | The firm, OrthoHelix Surgical Designs, Inc., sent an "IMPORTANT RECALL NOTICE: Mini MaxTorque HCS-056-32 & HCS-056-38 Drivers" dated January 29, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue the use of these drivers immediately and return them to OrthoHelix Surgical Designs Inc, 1065 Medina Rd., Suite 500, Medina, OH 44256; ATTN: Driver Recall.Should you have any questions about system, call 330-247-1145 or the Director of Regulatory Affairs & Quality Assurance at 330-247-1444. |  | 120 | 120 Pieces | Nationwide distribution: USA including the states of CA, CO, FL, GA, IL, MO, NC, NH, NJ, NM, NY, OH, OR, PA, SD, TN, TX, VA, WA, and WI. | Terminated | February 20, 2013 | 134 | 167 | Not_Computer
Z-0023-2013 | 63019 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PREVI Isola System, Catalog numbers 29500 and 29500R, bioMerieux, Inc., Durham, NC.Intended to perform the inoculation and streaking of Prepoured Media (PPM) with liquid microbiological samples. | PREVI Isola System | 2 | October 09, 2012 | 2012 | Biomerieux Inc | The instrument may not dispense the specimen onto the pre-poured media resulting in a false negative result. | DESIGN: Labeling Design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" letter dated August 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return. Contact Customer Service at 1-800-682-2666 for questions regarding this notice. | Contact the recalling firm for information | 350 | 350 systems | Worldwide Distribution-USA (nationwide)-including DC and the states of CA, CO, CT, FL, GA, IL, IN, KY, LA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TX, and WI, and the country of Canada. | Open |  |  | 350 | Not_Computer
Z-0021-2013 | 63158 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System, Part Number 629029.The UniCel DxH 800 Analyzer is a quantitative, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides a: " Complete Blood Count (CBC), Leukocyte 5 Part Differential (Diff), Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood. " Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal, serous and synovial) (BF). | Unicel DxH 800 Coulter Cellular Analysis System | 2 | October 09, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH 800 Coulter Cellular Analysis System because the NRBC, Differential or Reticulocyte mix chamber drain ports may become plugged by rubber pieces from specimen tube caps, and result in biohazard or chemical exposure and/or the potential for generating incorrect results. | DESIGN: Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated April 3, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Customers were notified that Beckman Coulter is in the process of developing and testing a design change that will prevent the mix chamber waste port from plugging. Customers with questions were instructed to contact Beckman Coulter Customer service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative. | Contact the recalling firm for information | 1281 | 1281 units total (693 units in the US) | Worldwide Distribution-USA (nationwide) and the counties of Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Kuwait, Macao, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Philippines, Puerto Rico, Qatar, Romania, Russian, Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | June 20, 2013 | 254 | 1281 | Not_Computer
Z-0028-2013 | 63070 | instrument, biopsy | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needles, Soft Tissue Biopsy Needle, NAC-1820M, Sterile, Remington Medical, Inc., 5830 Meadowridge Court, Alpharetta, GA 30005.Prostate Biopsy Needles shall be single use only and are to be used by a Urologist in a healthcare setting to obtain needle biopsies of the prostate. The needles are compatible with the Bard Magnum biopsy instrument. The needles have an echogenic tip for accurate placement under ultrasound guidance and centimeter markings along the cannula to facilitate depth placement. | Automatic Cutting Needles | 2 | October 09, 2012 | 2012 | Remington Medical Inc. | Sterility of the product may be compromised. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Remington Medical, Inc., sent an "URGENT - MEDICAL DEVICE RECALL" letter dated August 29, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products; complete and return the FaxBack form via Fax to: Quality Engineer Remington Medical, Inc., at 770-887-0732 or email: tracyw@remmed.com; and call the Quality Engineer or customer service at (800) 989-0057 ext 233 to obtain a RGA number and have a call tag issued.Should you have any questions, please do not hesitate to call the Quality Engineer at (800) 989-0057 ext 233. | Contact the recalling firm for information | 800 | 800 needles | Nationwide distribution: USA including states of: AL, AZ, CA, FL, IN, MA, NC, NY, OH, PA and TX. | Terminated | January 15, 2013 | 98 | 800 | Not_Computer
Z-0022-2013 | 63214 | sterilizer, ethylene-oxide gas | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M Steri-Vac Gas Sterilizer/Aerator models 5XL and 8XL.3M Steri-Vac Gas Sterilizer/Aerator is a compact unit designed to sterilize heat- and/or moisture-sensitive devices. This gas sterilizer/aerator is intended for indoor use only. | SteriVac | 2 | October 09, 2012 | 2012 | 3M Company - Health Care Business | 3M determined that the electrical grounding mechanism for the Steri-Vac 5XL and 8XL gas Sterilizer/Aerator did not meet all of the requirements of the standard in that the electrical wiring, which is insulated and contained within the sealed door should have a stronger connection to the protective electrical grounding mechansim. | DESIGN: Component Design/Selection | 3m Healthcare Service Center sent a Urgent Medical Device Field Correction letter dated September 19. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3Ms goal is to service the affected units with minimal disruption to your facility.We apologize for any inconvenience this field correction may cause. Please be assured that 3M is committed to the highest level of safety and customer service and we are proud of our longstanding safety record for 3M" Steri-Vac" Gas Sterilizers.For Further Information If you have any questions, please contact the 3M Health Care Service Line at 1-800-688-5888 or your local 3M representative.. |  | 707 | 707 units (61 US, 646 OUS) | Worldwide Distribution--USA (nationwide) including the states of AL, CO, CT, GA, IN, KY, MA, MI, MN, MO, MS, NJ, OH, PA, SC, TN, VA, WA and WI., and the countries of Argentina, Barbados, Chile, China, Colombia, Costa Rica, Ecuador, Germany, India, Indonesia, Japan, Korea, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Pakistan, Paraguay, Peru, Philippines, Poland, Republic of Guinea, Russia, South Africa, Taiwan, Thailand, Trinidad, UAE, Venezuela, Vietnam and Zimbabwe. | Terminated | January 29, 2013 | 112 | 707 | Not_Computer
Z-0042-2013 | 62968 | system, test, osteocalcin | NEO | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IDS-iSYS N-MidÂŽ Osteocalcin, Catalog Number IS-2900. Packaged in cardboard Box containing plastic bottles.The IDS-iSYS N-MidÂŽ Osteocalcin Assay is intended for the quantitative determination of Osteocalcin in human serum or plasma on the IDS-iSYS Multi-Discipline Automated Analyser (Analyser). Results are to be used in conjunction with other clinical and laboratory data to assist the clinician. Osteocalcin is an indicator of osteoblastic activity in human serum and plasma and is intended to be used as an aid in the prevention of osteoporosis. | IDSiSYS NMid Osteocalcin | 2 | October 10, 2012 | 2012 | Immunodiagnostics Systems Ltd | It has been noted by staff at IDS Inc. that stock of the affected products held at IDS Inc. are labeled with IVD labels although they are not registered for sale as IVDs in the United States. The products should have been labeled as Research Use Only or Investigational use only dependent on their usage. | MISBRANDING: Labeling False and Misleading | Immunodiagnostics Systems sent Customer notification letters via email and notified customers of the recall. The notification included the reason for recall, product information with codes, lot number, manufacturing/distribution date, and exp date. The notification included instructions to customers: Actions to be taken by the customer:Customers should complete the attached Recall and Acknowledgement Form and take the following actions:1) Check your stock for the products included within the scope of the recall. Cease use and distribution and quarantine all affected product lots immediately.2) Complete the enclosed Recall & Acknowledgement Form and immediately fax back to Immunodiagnostic Systems. This will allow us to document your receipt of this letter and the amount of product you have on hand for return.3) Return any affected product freight collect, along with the original completed Recall & Acknowledgement Form to your local IDS office.Contact information was also provided.For questions regarding this recall call +44 (191) 5195212. | Contact the recalling firm for information | 5 | 5 | Nationwide Distribution including NY, CA, ME and CT. | Terminated | December 17, 2012 | 68 | 51 | Not_Computer
Z-0037-2013 | 63157 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ATLANTIS(R) Anterior Cervical Plate System FIXED ANGLE SCREW MODULE, REF 176-504, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132.Orthopedic.This instrument is a precision device which may incorporate a measuring function and has uses as described on the label. Unless labeled forsingle use, this instrument may be re-used. | ATLANTIS(R) Anterior Cervical Plate System FIXED ANGLE SCREW MODULE | 2 | October 10, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. An incorrect gauge could measure a screw to be approximately 1 mm shorter than its actual length. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic conducted a conference call on August 30, 2012 and sent an "URGENT VOLUNTARY RECALL" notice dated September 3, 2012 to all affected customers. The notice identified the product, problem, and actions to be taken by the customers. A Recall Questionnaire Form was attached to the letter for customers to complete and return. Contact your Medtronic Sales representative or the firm's Global Quality Department at 800-876-3133, ext. 3197 for questions regarding this notice. | Contact the recalling firm for information | 1014 | 1014 units | Worldwide Distribution-USA (nationwide) and the countries of India, Netherlands, Hong Kong, S. L Korea, Singapore, Canada, Australia, Japan, and Mexico. | Terminated | August 15, 2013 | 309 | 2308 | Not_Computer
Z-0032-2013 | 63085 | chair, with casters | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Caremor Cliner is labeled in part: "***Manufactured By Winco***Ocala, FL 34474 www.wincomfg.com***Model No. 535 Serial No. 535A-104522***"Literature: "***Caremor Cliner 5351/5361***Winco***ISO 9001-2008 Certified***Our most comfortable and versatile recliner***Ease of operation, outstanding durability, and exceptional value have made this one of our bestselling recliners. The Caremor Cliner provides patient comfort and support as well as user friendly features.***KEY FEATURE'S***Infinite positioning back***Padded armrests w/arm guard***Webbed elastic suspension back and seat***Removable side panels**Continuous seat & legrest cushion***5351 w/tray***5361 w/out tray***1.800.237.3377***5516 SW 1st Lane***Ocala, Florida 34474***www.wincomfg.com***Proudly Made in the USA***"Patient recliner chair | Caremor Cliner | 2 | October 10, 2012 | 2012 | Winco Mfg., LLC | Winco Mfg., LLC in Ocala, FL is recalling the Caremor Cliner model numbers 5351 and 5361. The recall was initiated due to a report that a back frame did not "snap" onto a chair frame properly. | PRODUCTION CONTROLS: Process Control | Winco notified consignees via phone except for the one located in Saudi Arabia, who was contacted by e-mail. All consignees have been sent a follow-up e-mail. The notification identified the product, the problem, and the action to be taken to be taken by the customer. Customers should remove the chair from use and the upholstered back will not be required to be returned. Winco will send out new fully assembled upholstered backs. For questions customers should call 352-854-2929, ext 123. | Contact the recalling firm for information | 20 | 20 chairs | Worldwide Distribution - USA including Arkansas, California, Maryland, Mississippi, New York, Pennsylvania and the country of Saudi Arabia | Terminated | May 23, 2013 | 225 | 43 | Not_Computer
Z-0046-2013 | 63169 | acid, folic, radioimmunoassay | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Folate Calibrators, Part Number: A14207.The Access Folate assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of folic acid levels in human serum, plasma (heparin) or red blood cells using the Access Immunoassay Systems. The serum folate level is an indicator of recent folate intake. A low RBC folate value can indicate a prolonged folate deficiency. The Access Folate Calibrators are intended to calibrate the Access Folate assay using the Access Immunoassay Systems. | Access Folate Calibrators | 2 | October 10, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access Folate Calibrators because it may produce falsely elevated results that could mask a folate deficiency. |  | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated July 24, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return the enclosed Response Forms within 10 days. Customers with questions were instructed to call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter representative. |  | 4635 | 4,635 units total (3,214 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Austria, Australia, Bangladesh, Belgium, Botswana, Brazil, Canada, China, Colombia, Czech Republic, Croatia, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Sweden, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zimbabwe. | Terminated | October 17, 2013 | 372 | 4635 | Not_Computer
Z-0035-2013 | 63161 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | TetraCXP Software System, Part Number A40051.Allows simultaneous identification and enumeration of total CD3+, total CD4+, total CD8+, dual CD3+/CD4+, dual CD3+/CD8+ and/or total CD3+, CD19+ and CD3+/CD56+ lymphocyte percentages and absolute counts in whole blood by flow cytometry. The systems also provide the CD4/CD8 ratio when using CD45-FITC/ CD4-RD1/CD8-ECD/CD3-PC5, and total lymphocyte percentage when using CD45-FITC/CD56-RD1/ CD19-ECD/CD3-PC5. | TetraCXP Software System | 2 | October 10, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has determined that the specimen and prepared sample stability claims stated in the product labeling for tetraCXP software and CYTO-STAT reagents (CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 and CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5) need to be updated for alignment with current standards and clinical decision-making associated with lymphocyte immunophenotyping | PRODUCTION CONTROLS: Process Control | Beckman Coulter sent a Product Corrective Action (PCA) letter dated May 22, 2012, with attached PCA Response Form to all customers who purchased the tetraCXP Software, CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that the tetraCXP SYSTEM Guide and CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 & CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5 product labeling will be revised to reflect the updated claims. Customers were instructed to complete and return the enclosed Response form within 10 days. Customers with questions were instructed to contact Beckman Coulter Customer Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative.For questions regarding this recall call 714-961-4483. |  | 146 | 146 units total (37 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:AlbaniaAustraliaAustriaBahrainBangladeshBarbadosBelarusBelgiumBosnia and HerzegovinaBulgariaBurundiCanadaChileChinaColombiaCroatiaCzech RepublicDenmarkEcuadorEl SalvadorFinlandFranceFrench GuianaFrench PolynesiaGermanyGreeceGuadeloupeHong KongIndiaIndonesiaIraqIrelandIsraelItalyJapanKenyaKorea, Republic ofKuwaitLibyan Arab JamahiriyaLithuaniaMacaoMalaysiaMexicoMonacoMoroccoNamibiaNetherlandsNew ZealandNigeriaNorwayOmanPakistanPanamaPolandPuerto RicoQatarReunionRomaniaRussian FederationSaudi ArabiaSingaporeSlovakiaSloveniaSouth AfricaSpainSudanSwedenSwitzerlandTaiwanTurkeyUnited Arab EmiratesUnited KingdomUruguay | Terminated | October 10, 2012 | 0 | 146 | Not_Computer
Z-0030-2013 | 62832 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL)Product Usage: Enzyme n Immunoassay, Phenobarbital | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL) | 2 | October 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed highter imprecision with the recalled lots on Beckman Coulter AU(R) Systems. Also, if the Syva(R) Emit(R) 2000 Phenobartbital Assay (4D019UL) lots D4, D5, D6, and/or D7 as an application on AU(R) Systems is used, the issue may be experienced. Firm has found that within-run imprecison of the recalled lots when run on an AU(R) System may be higher that what would normally be obs |  | Siemens sent an Urgent Field Safety notice letter dated July 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 79744 | 79,744 cartons | United States Nationwide Distribution in the states of: CA, FL, KS, MA, MN, MS, NM, NY, NC, OR, PA, TN, TX, UT, VA, and WA. | Open |  |  | 79744 | Not_Computer
Z-0029-2013 | 63175 | enzyme immunoassay, methadone | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen Panel, Model# 94400.Product Usage: The Triage TOX Drug Screen is a fluorescence immunoassay intended to be used with the Triage Meters for the point-of-care qualitative determination of the presence of drug and/or the major metabolites above the threshold concentrations of up to 10 distinct drug classes, including assays for acetaminophen/paracetamol, amphetamines, methamphetamines, barbiturates, benzodiazepines, cocaine, methadone, opiates, phencyclidine, THC and tricyclic antidepressants in urine. | Triage TOX Drug Screen | 2 | October 10, 2012 | 2012 | Alere San Diego, Inc. | Biosite is recalling the Triage TOX Drug Screen kits where the box containing test devices because it was mislabeled. | PRODUCTION CONTROLS: Error in Labeling | Biosite sent a Notice of Correction letter dated May 28, 2009 to all their customers who purchased the Triage TOX Drug Screen. The letter informed the customers of the problem identified and actions to be taken. Customers are instructed that performance of Triage TOX Drug Screen test devices is not affected in any way by the misprinted box flap, provided the test is used with a urine sample as per the instructions for use. Customers are instructed to complete and fax the enclosed verification form within 10 days. In the US, customers were instructed to contact Biosite at (877) 441-7440 or email productadvisory@invmed.com. Customers in UK are instructed to contact European Representative, Lesley Paice, lesley.paice@invmed.com. | Contact the recalling firm for information | 45638 | 45,638 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Brazil, Canada, Colombia, Czech Republic, Finlind, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, India, Indonesia, Italy, Lebanon, Mexico, Netherlands, Norway, Puerto Rico, Romania, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, and Uruguay. | Terminated | October 10, 2012 | 0 | 45638 | Not_Computer
Z-0055-2013 | 63362 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | 3 PORT "OFF" MANIFOLD w/WASTE BAG, SYRINGE andDyePodÂŽ, Part Number (Item #) AG8049-NS, Lot # 2528655.K914299.Waste system. | 3 PORT "OFF" MANIFOLD w/WASTE BAG, SYRINGE and DyePod | 2 | October 11, 2012 | 2012 | ICU Medical, Inc. | ICU Medical, Inc. has identified a potential comingling of components where a 60 drop micro dripchamber may have been included on a 20 drop macro drip set. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | ICU Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated September 28, 2012 to its customer. The letter identified the product, problem, and actions to be taken by the customer. The customer has been instructed to inspect their inventory for the specific lot numbers, examine for the presence of a metal cannula, and certify destruction of any affected devices. The customer letter included a Verification Form to be completed and returned. Contact ICUCustomer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com" for questions regarding this notice. |  | 60 | 60 | Nationwide distribution: Virginia only. | Terminated | February 19, 2013 | 131 | 60 | Not_Computer
Z-0054-2013 | 63111 | twister, wire | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device. The common name is TWISTER, WIRE. The model number is TQ01A. The Catalog numbers are:17-00001: TORQ Sternal Closure Device Shipping Carton (6-10 packs), Sterile and,15-00001 (or 17-00002): TORQ Sternal Closure Device 10-pack - Sterile.The device is supplied double-pouched, inside a shelf box containing 10 devices each. Shelf boxes are shipped to customers in either a small shipping carton containing one (1) shelf box (for a total of 10 devices), or large shipping carton containing six (6) shelf boxes (for a total of 6 x 10 = 60 devices).This product is used to close the sternal wires during cardiac surgery, and it is not an implantable device. The TORQ Sternal Closure device is used during sternal closure to tension and twist standard stainless steel surgical sutures of USP size 6 or 7 (metric size 8.0 or 9.0). The TORQ device is designed for use as a standard surgical suture wire twister for sternal closure. (Surgical sutures are not provided with the device.) The TORQ device is disposable, single-use, non-implantable and supplied sterile. | TORQ Sternal Closure Device | 2 | October 11, 2012 | 2012 | Kardium | The TORQ Sternal Closure Device, Lot 062711, is being recalled due to probability one of the metal parts may be dislodged during use. | TRAINING: Employee Error | Kardium Inc. sent an "URGENT: TORQ STERNAL CLOSURE DEVICE RECALL" e-mailed dated August 30, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the TORQ Sternal Closure Device lot 062711 immediately. Contact the firm at 604-248-8891, x285 for any questions about this recall. |  | 480 | 480 units | Nationwide Distribution-including the states of FL, GA, IL, and NH. | Terminated | February 05, 2013 | 117 | 480 | Not_Computer
Z-0059-2013 | 63191 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Closed Tube Aliquotter (UCTA) for UniCel DxC 660i, UniCel DxC 680i, UniCel DxC 860i, UniCel DxC 880i Synchron Access Clinical Systems. Part Number A45853.Product Usage: The UniCel DxC Synchron Clinical Systems are fully automated, computer-controlled clinical chemistry analyzers designed for the in-vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | UniCel Closed Tube Aliquotter for UniCel DxC 660i, 680i, 860i and 880i Synchron Access Clinical Sys | 2 | October 12, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the current sample syringe pump drive assemblies installed on DxC Integrated ystem with UniCel Closted Tube Aliquotter (UCTA) serial numbers between 547 and 574 may cause a hardware motion error (ID0x08000501). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recall notification was initiated on July 6, 2012 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form to all customers who purchased the UniCel DxC 880i, UniCel DxC 860i, UniCel 680i, UniCel DxC 660i Clinical System Integrated Workstations UCTA Sample Syringe Pump Drive Assembly. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Beckman Coulter will schedule a Service Call to replace the affected sample syringe pump drive assemblies. Customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any questions regarding the PCA, were instructed to call Customer Support at (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact Beckman Coulter Representative. |  | 18 | 18 units | Worldwide Distribution and US (nationwide) and the countries of China, France, Italy and Taiwan. | Terminated | November 15, 2012 | 34 | 18 | Not_Computer
Z-0057-2013 | 62879 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Hemorrhoidal Circular Stapler and Accessories 33mm Product Usage: The PROXIMATE HCS Hemorrhoidal Circular Staplers, (PPH01 and PPH03), and accessories have application for general surgical treatment of anorectal wall defects by means of transanal stapling and resection of mucosal and musculomucosal tissue. The TRANSTARTM CIRCULAR STAPLER, (STR10), and accessories have application for general surgical treatment of anorectal wall defects by means of transanal stapling and resection of mucosal and musculomucosal tissue. | PROXIMATE and Transtar | 1 | October 13, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated this voluntary global recall for specific production lots of PROXIMATEÂż PPH Hemorrhoidal Circular Stapler and Accessories 33mm (PPH01 and PPH03) and TranstarÂż Circular Stapler Procedure Set (STR10) due to difficulty firing the device which may result in incomplete firing stroke that may result in an incomplete staple formation. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated August 3, 2012 via overnight, Federal Express UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return per the instructions provided. Customers were advised to complete the Business Reply Form (BRF) and fax it to Stericycle at 1-888-628-0733 within 3 business days, even if you do not have affected product. The letter states that Ethicon Endo-Surgery will issue credit for recalled products returned by November 5, 2012. Any product returned after November 5, 2012 will not be eligible for credit. For clinical or product support call 1-800-873-3636, option 6 or contact your local sales representative and reference Event 7283. | Contact the recalling firm for information | 157881 | 157,881 devices | Worldwide Distribution - US (nationwide) in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the countries of: ARGENTINA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MALAYSIA, MEXICO, NEPAL, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, and YEMEN. There was one Canadian consignee and 43 U.S. Government consignees. | Terminated | July 12, 2013 | 272 | 157881 | Not_Computer
Z-0061-2013 | 63124 | [6 empty] | CPT Hip System Femoral Stem Petite, Item No. 00-8011-000-05.CPT stems are primary cemented femoral hip stems used by surgeons in total hip arthroplasty. | CPT Modular Stem | 2 | October 15, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the CPT Modular Stem (item 00-8011-001-00), CPT Stem Petite (item 00-8011-000-05; ) and CPT Femoral Stem 12/14 Neck Taper Extended Offset,(item 00-8114-001-10 ) because devices were packaged in incorrect cartons and distributed internationally and domestically. Zimmer received five complaints of this packaging issue. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | Zimmer Inc. sent an "URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC" letter dated September 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Consignees were asked to return all recalled product with a completed Inventory Return Certification Form. Contact the firm at 1-800-348-2759 for questions regarding this notice. |  | 20 | 20 | Worldwide Distribution-USA (nationwide) including the states of Indiana and Oklahoma and the countries of Australia, Thailand, Malaysia, Germany, India, Bulgaria, Egypt, UK, Russian FED, Slovak, and Brazil. | Terminated | January 15, 2014 | 457 | 138 | Not_Computer
Z-0066-2013 | 63115 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***PressFT***2.1 w/ Two #0 (3.5 metric) Hi-Fi***Sutures***NP212***2009-02***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***To reattach soft tissue to bone in orthopedic surgical procedures | PressFT2.1 w/ Two 0 (3.5 metric) HiFiSutures | 2 | October 15, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling PressFT HI-FI Sutures (items NP211, NP212, NP261, NP262, NP211H, NP261H-all lot numbers) due to the potential for the distal tip of the driver to become separated from the driver shaft after successful insertion of the implant. | PRODUCTION CONTROLS: Process Control | ConMed Linvatec sent an Urgent Medical Device Recall Notification letter on June 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility for the recalled products, segregate them for return to ConMed Linvatec. The customers were also told to return the REPLY FORM to the Regulatory Department at the firm. For questions customers were instructed to call 800-237-0169 or email at Custserv1@linvatec.com.For questions regarding this recall call 727-392-6464. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA including AZ, CA, CO, FL, GA, LA, MA, ME, MI, MN, NJ, NM, NY, PA, TX, UT, and WA. Internationally to BE, CA, ES, MY, NZ and PL | Open |  |  |  | Not_Computer
Z-0071-2013 | 63225 | accessory equipment, cardiopulmonary bypass | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Backplate Replacement CXFVR 1500 HolderThe Flexible venous Reservoir is mounted on the Backplate/Holder and secures the flexible reservoir by strategically placed mounting arms. Additionally, there are three port supports the user incorporates to position the lines that enter and exit the venous bag. The Backplate/Holder is designed to allow the user to move the port support location to accommodate different tubing circuit designs to permit mounting flexibility. | Terumo Backplate Replacement CXFVR 1500 Holder | 2 | October 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Backplate reservoir holders were manufactured with an incorrect configuration of connector supports. The devices were manufactured and configured with two 3/8 inch and one Âż inch connector supports instead of one 3/8 inch and two Âż inch connector supports (See location of supports as # 7 and # 8 on backplate attachment). The backplates were manufactured as described in manufacturing instructions a | DESIGN: Process Design | The firm, Terumo, interviewed the Sr. Product Manager from the Marketing Department of the company on August 29, 2012. The interview included the description of the product, problem and action taken. Replacement connectors were provided to users and all units were corrected.If you have any questions, call 734-663-4145 or Manager, Regulatory Affairs at 734-741-5576. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, NC, NJ, PA, and WI;and country of: EUROPE. | Terminated | January 29, 2013 | 105 | 33 | Not_Computer
Z-0114-2013 | 63027 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare innova 3100-IQ, Mobile X-ray systemThe innova systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography, diagnostic and interventional procedures, and optionally , rotational angiography procedures. They are also indicated for generating fluoroscopic images of human anatomy for cardiology, diagnostic, and interventional procedures. | GE Healthcare innova 3100IQ | 2 | October 17, 2012 | 2012 | GE Healthcare, LLC | It was discovered of a potential failure of the spectral filter mechanism inside the collimator of the innova 3100, innova 3100-IQ and innova 3131-IQ systems. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. A GE Healthcare Service Representative will replace the collimator to address the issue.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. Further questions please call 262-513-4122. |  | 199 | 199 installed in the US | Nationwide distribution | Terminated | September 18, 2013 | 336 | 199 | Not_Computer
Z-0072-2013 | 63359 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ***This recall is being conducted due to an incomplete recall for RES 57829*** Curlin Infusion Administration Set, REF 340-4165, Sterile EO, The Curlin Ambulatory Infusion Pump System with disposable PVC Administration Tubing Sets and user accessories, provide a means for the volumetric delivery of fluids used in Parenteral, Enteral, Epidural, Subcutaneous and Intravenous applications. | Curlin Ambulatory Infusion Administration Set | 2 | October 17, 2012 | 2012 | MOOG Medical Devices Group | This recall is the same issue for the FDA processed recall initiated by Moog Medical Devices on 1/25/11 which was related to customers reporting unclearable Air-In-Line alarm with specific lot numbers of Curlin Administration Sets used with Curlin Ambulatory Infusion Pumps. The lots associated with this recall were inadvertently omitted from the original recall which has since been terminated by | PRODUCTION CONTROLS: Process Control | MOOG sent a Medical Device Recall Notification letter dated September 26, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove product from their inventory and return to the firm. Customers were also instructed to contact their distributor to arrange for replacement product. For questions customers were instructed to call 1-800-970-2337, prompt #7.For questions regarding this recall call 801-264-1001. |  | 12000 | 12,000 sets | Nationwide Distribution including California | Terminated | April 05, 2013 | 170 | 12000 | Not_Computer
Z-0081-2013 | 63336 | electrode, electrosurgical | JOS | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Leveen Coaccess Introducer Set.Disposable, monopolar, electrosurgical devices, intended for use in conjunction with BSC radiofrequency generator for the thermal coagulation necrosis of soft tissues, including partial or complete ablation of nonresectable liver lesions. | LeVeen Coaccess Introducer Set | 2 | October 18, 2012 | 2012 | Boston Scientific Corporation | Leveen Standard and Leveen Coaccess Needle Electrode System, Directions for Use(DFU) in Portuguese, are incorrect. all other languages in the DFU are correct. Consignees in Brazil, Portugal,and Hong-Kong are being notified. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Boston Scientific sent notifications, dated September 17, 2012, to each account in Brazil, Portugal and Hong Kong affected by the recall. The affected accounts were provided the corrected pages of the Portuguese DFU as part of the communication to ensure that those users that rely on the Portuguese DFU have the correct information. |  | 19716 | 19,716 units distributed total all products | Distributed in Brazil, Portugal and Hong Kong. | Terminated | July 17, 2013 | 272 | 19716 | Not_Computer
Z-0087-2013 | 62569 | filler, bone void, osteoinduction (w/o human growth factor) | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM CrunchMultiple SizesGrafton is a human bone allograft product containing demineralized human bone tissue combined with inert additive to yield a demineralized bone matrix (DBM) allograft. It is intended for use as a bone graft extender, bone graft substitute and bone void filler in bony voids or gaps of the skeletal system (ie spine, pelvis and extremeties) not intrinsic to the stability of the bony structure. The voids or gaps may be surgically created defects or effects created by traumatic injury to the bone. | Grafton DBM Crunch | 2 | October 18, 2012 | 2012 | Osteotech Inc | Multiple lots of GRAFTON Âż and XPANSE Âż, which are used for bonegrafting procedures are being recalled because there is a possibility that sterility of the outer surface of the inner pouch may becomecompromised. | DESIGN: Device Design | Osteotech sent out recall letters/questionairres dated April 20, 2012, to Risk Managers/service representatives. The recall was expanded on August 1 2012. The letter identified the product, the problem, and the action to be taken by the customer. Sales Representatives would facillitate the return of any affected product they may have or provide cautionary labels for any affected product and/or support in the completion of the attached questionnaire. Customers were asked to disseminate this information to additional personnel within their facility as appropriate and return the questionnaire to Medtronic.. Customers with questions should contact their Medtronic Sales Representative at 800-876-3133, extension 3197.For questions regarding this recall call 732-542-2800.Osteotech sent an additional letter stating that The OPTIUMTM (DBM) product line was added in error. |  | 3541 | 3,541 units | Worldwide Distribution - USA (nationwide) and Internationally. | Terminated | January 17, 2014 | 456 | 37249 | Not_Computer
Z-0078-2013 | 63365 | clip, implantable | FZP | Implantable clip. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Hem-o-lok Large Polymer Ligating Clips, REF #544240, Rx only, Sterile, Teleflex Medical, Research Triangle Park, NC 27709.Teleflex Medicals Hem-o-lok polymer ligating clips are single use, sterile, non-absorbable medial devices that have been designed to ligate vessels and tissue structures during various types of surgical procedures. Vessels and tissue structures in the range of 0.5 to 16 mm can be effectively ligated with Hem-o-lok clips. | Hemolok | 2 | October 18, 2012 | 2012 | Teleflex Medical | Product is being recalled due to the possibility that the tray may contain pinholes, compromising product sterility. | PRODUCTION CONTROLS: Packaging Process Control | Teleflex sent a Urgent Medical Device Recall Notification dated September 25, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory and quarantine any affected inventory found. They were to complete the enclosed Field Corrective Action Acknowledgement Form and return it to the Teleflex Medical sales representative who will then remove the affected product from the facility and ensure its return to Teleflex Medical.We sincerely apologize for any inconvenience this action may cause your operations. If you have any questions please contact your local sales representative or Customer Service at 1-866-246-6990. |  | 126 | 126 units | Nationwide Distribution including the state of Nevada. | Terminated | November 08, 2012 | 21 | 126 | Not_Computer
Z-0082-2013 | 63223 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Columbus 4-IN-1 Femoral Cutting Guide M3 Instrument (NQ083R) | Columbus 4IN1 Femoral Cutting Guide M3 Instrument | 2 | October 18, 2012 | 2012 | Aesculap, Inc. | The manufacturer informed the recalling firm of a potential laser marking error in size identification on two Columbus 4-IN-1 Femoral Cutting Guide Instruments. The guides are correctly dimensioned, but the laser markings have been incorrectly etched. | TRAINING: Employee Error | Both consignees were notified verbally on May 15, 2012 and asked to inspect their inventory for the recalled items, remove the mislabeled instruments and return them to the recalling firm. Both mislabeled instruments were returned to the firm. Customer questions were directed to (610) 984-9074. |  | 2 | 2 | Within the US, product was distributed to PA and TN. | Terminated | January 15, 2014 | 454 | 2 | Not_Computer
Z-0075-2013 | 63086 | drills, burrs, trephines & accessories (simple, powered) | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | "***Electric Systems Foot Control with Direction Only***Rx Only***"Product Usage:Cutting and shaping bone including bones of the spine and cranium. | ANSPACH Electric System Foot Control with Direction Only | 2 | October 18, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. in Palm Beach Gardens, FL is recalling the foot control/foot pedal due to the Foot Control Cable lacking the shield connection from the cable to the connector body. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Anspach sent an Urgent Product Removal letter dated January 9, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact their sales representative to facilitate the removal action by providing a replacement for the affected product. For question contact your local sales representative or the Product Support department. | Contact the recalling firm for information | 41 | 41 units | Worldwide Distribution - US (nationwide) and the countries of: Italy, Japan, Switzerland, Turkey and Taiwan. | Terminated | July 25, 2013 | 280 | 67 | Not_Computer
Z-0077-2013 | 63255 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Helion and Xenion surgical lights, 501/01 series KLM lights. The surgical lights are intended to locally illuminate an operating or examination area of the patients body with a high intensity light. | Helion and Xenion surgical lights | 2 | October 18, 2012 | 2012 | Trumpf Medical Systems, Inc. | Incidents regarding fractures of the front joint of the spring arm have been reported. | DESIGN: Device Design | Trumpf sent an "URGENT RECALL NOTICE" dated September 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Confirmation of Receipt Form was attached for customers to complete and return. Contact the firm at 888-474-9359 for questions regarding this notice. |  | 275 | 275 units | Nationwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, IL, KS, MA, MI, MO, NC, NJ, NY, OH, SC, TX, UT, VT and WI. | Terminated | January 21, 2014 | 460 | 275 | Not_Computer
Z-0096-2013 | 63276 | arterial blood sampling kit | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Pulset Arterial Blood Gas (ABG) sampling kit.Modified Kit, Pulset, 1 cc, w/Crickett 25g x 5/8" A Needle, 25U Balanced Heparin, part number: 3165-95.An arterial blood sampling kit is a device, in kit form, used to obtain arterial blood samples from a patient for blood gas determinations. | Pulset Arterial Blood Gas (ABG) sampling kit. | 2 | October 20, 2012 | 2012 | Westmed Inc | Westmed is recalling Westmed ABG convenience kits containing TRIAD brand alcohol pads because TRIAD is recalling TRIAD brand alcohol pads due to potential contaminants. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Westmed sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 18, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were also given the TRIAD GROUP: "Urgent Drug Recall" and "Facts: Alcohol Recall" both dated January 3, 2011. Westmed customers were instructed to remove the TRIAD alcohol pad from the kit and dispose of the alcohol pad only. Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with questions were instructed to contact the firm at (520) 264-7987 x1233. | Contact the recalling firm for information | 1259200 | 1,259,200 units total | Worldwide Distribution-USA (nationwide) and the countries of Canada, El Salvador, United Arab Emirates, Holland, Costa Rica and Lebanon. | Terminated | October 22, 2012 | 2 | 1259200 | Not_Computer
Z-0110-2013 | 63130 | calibrator, multi-analyte mixture | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Synchron Systems Synchron Calibrator Multi Calibrator Diskettes, Part Number 442600.The SYNCHRON MULTI-CALIBRATOR, used in conjunction with SYNCHRON reagents, is intended for use on SYNCHRON Systems for the calibration of Albumin, Blood Urea Nitrogen (Urea), Calcium, Cholesterol, Glucose, Inorganic Phosphorus, Lactate,Magnesium, Total Protein, Triglycerides, and Uric Acid. | Synchron Systems Synchron Calibrator Multi Calibrator Diskettes | 3 | October 22, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed a mislabeling error, some Synchron Systems Multi Calibrator Diskettes (REF 442600, lot M002642G) may actually be Thyroxine (T4) Reagent Calibrator Diskettes. | PRODUCTION CONTROLS: Labeling Mix-Ups | Beckman Coulter sent an Urgent Product Correction letter dated December 15, 2011, withan attached Fax Back Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. If customers are currently using Multi Calibrator lot M002642, then they were instructed to verified diskette and discard all Multi Calibrator diskettes lot M002642 and M002642G without a dark green dot. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers were instructed to contact Beckman Coulter Support Center at http:www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US or Canada, customers were instructed to contact their local Beckman Coulter Representatives.For questions regarding this recall call 714-961-4483. |  | 5746 | 5746 units total (2729 units in US) | Worldwide distribution: USA (Nationwide) and Internationally to: Afghanistan Angola Australia Bangladesh Belgium Bhutan Brazil Brunei Darussalam Bulgaria Canada China Colombia Costa Rica Cote d'Ivoire Czech Republic Denmark Ecuador Egypt El Salvador Eritrea Ethiopia Finland France French Polynesia Germany Ghana Greece Guatemala Hong Kong India Indonesia Ireland Israel Italy Japan Jordan Kenya Kuwait Lebanon Lesotho Libyan Arab Jamahiriya Lithuania Macao Malawi Malaysia Mexico Morocco Namibia Netherlands New Zealand Nigeria Oman Pakistan Panama Paraguay Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Rwanda Saudi Arabia Serbia Singapore Slovakia South Africa Spain Swaziland Sweden Switzerland Taiwan Thailand Tunisia Turkey Uganda United Arab Emirates United Kingdom Viet Nam Zambia Zimbabwe | Terminated | December 28, 2012 | 67 | 5746 | Not_Computer
Z-0112-2013 | 63215 | kit, identification, enterobacteriaceae | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx GNR Traffic Light PNA FISH, In Vitro Diagnostic for Culture Identification Kit for Escherichia coli,and/or Klebsiella pneumoniae and/or Pseudomonas aeruginosa Catalog number: KT011.GNR Traffic Light PNA FISH is a multicolor, qualitative nucleic acid hybridization assay intended for the identification of Escherichia coli,and/or Klebsiella pneumoniae and/or Pseudomonas aeruginosa on smears from positive blood cultures containing Gram-negative rods observed on Gram stain. Sub-culturing of positive blood cultures is necessary to recover organisms for susceptibility testing and/or differentiation of mixed growth. The GNR Traffic Light PNA FISH assay is indicated for use as an aid in the diagnosis of coli, and/or K. pneumoniae, and/or P. aeruginosa bacteremia. | AdvanDx GNR Traffic Light PNA FISH | 2 | October 22, 2012 | 2012 | Advandx Inc | GNR Traffic Light PNA Fish may have false positive results with Pseudomonas putida and Pseudomonas fulva | DESIGN: Device Design | AdvanDx sent a "GNR Traffic Light PNA FISH (KT011) Package Insert Replacement" letter dated September 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received an Acknowledgement Form to complete and return to AdvanDx so that the effectiveness of the replacement can be evaluated. Contact AdvanDx Technical Support at 1-866-376-0009 for questions regarding this notice. | Contact the recalling firm for information | 187 | 187 kits | Worldwide Distribution-USA (nationwide) and the countries of Austria, Denmark, Italy, Germany, Finland, and UK. | Terminated | March 05, 2013 | 134 | 187 | Not_Computer
Z-0117-2013 | 63353 | implant, cochlear | MCM |  |  | Ear Nose & Throat | PMA | HiRes 90K devices with Helix Electrodes Model CI 1400-02HThe HiResolution Bionic Ear System is a cochlear implant system designed to provide useful hearing to individuals with severe to profound hearing loss. It consists of internal and external components. The internal components include a stimulator (HiRes90K ICS) and electrode array (HiFocus 1j or Helix). The HiRes90K ICS is implanted surgically under the skin behind the ear. The electrode array is inserted into the cochlea. The external components include a sound processor (body-worn or ear-level), a headpiece, and a cable. The system converts sound into electrical energy that activates the auditory nerve. The auditory nerve then sends information to the brain, where it is interpreted as sound. The Helix consists of a fantail, electrode lead, and Helix electrode array. The electrode, comprised of platinum-iridium wires, is housed in a silicone carrier and extends from the titanium case. The Helix electrode array is 24.5 mm in length and is designed to be inserted approximately 18-21 mm into a normal patent cochlea. It consists of 16 planar contacts arranged along the medial surface of the electrode array for stimulation of discrete segments of the cochlea. The Helix electrode insertion tool is designed to assist the surgeon with the insertion of the Helix electrode. Only the Helix insertion should be used for insertion of the Helix electrode; the use of other instruments may result in damage to the electrode. The insertion tool is a single-use instrument. | HiRes 90K | 2 | October 22, 2012 | 2012 | Advanced Bionics Corporation | Advanced Bionics is recalling the the HiRes 90K cochlear implant with HiFocus Helix Electrode because of loose-fitting insertion tool assemblies. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Advanced Bionics, sent a " RECALL NOTIFICATION" letter dated September 14, 2007 to all their customers who purchased the HiRes 90K devices with Helix Electrodes Model CI 1400-02H. The letter identified the product, problem, and actions to be taken. The customers were instructed to immediately locate and return the affected product to Advanced Bionics. The firm's Service representative will contact the customers within 48 hours to retrieve the affected product and send a replacement as needed. The customers were also instructed to complete and return the enclosed Notification Acknowledgement form via fax to: (661) 362-1511.If you have any questions regarding this letter, call Advanced Bionics at (877) 829-0026 between the hours of 5 AM and 5 PM Pacific Time, Monday through Friday. | Contact the recalling firm for information | 17 | 17 units | Worldwide distribution: USA (nationwide) including state: IL and country of: Europe. | Terminated | October 22, 2012 | 0 | 17 | Not_Computer
Z-0116-2013 | 63199 | kit, identification, enterobacteriaceae | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH, In Vitro Diagnostic Culture Identification Kit for Escherichia coli and Pseudomonas aeruginosa Catalog Number : KT007Usage: E. coli/P. aeruginosa PNA FISH is a multicolor, qualitative nucleic acid hybridization assay intended for identification of Escherichia coli and Pseudomonas aeruginosa on smears of positive blood cultures containing Gram-negative rods. The E.coli/P. aeruginosa PNA FISH assay is indicated for use as an aid in the diagnosis of E. coli and/or P. aeruginosa bacteremia. | E. coli/P. aeruginosa PNA FISH | 2 | October 22, 2012 | 2012 | Advandx Inc | Some strains of Pseudomonas putida cross-react with the P. aeruginosa PNA probe used in E. coli/P. aeruginosa PNA FISH (KT007) due to a sequence similarity resulting in a cross-reaction that may produce a red-positive signal | DESIGN: Process Design | AdvanDx notified customers on September 14, 2012 by email along with a PDF file of the new package insert to each contact followed by written notification and a paper copy of the new package insert via FedEx courier. The notification described the product, problem and actions to be taken. The customers were instructed to destroy the package insert in any existing inventory of the kit and replace it with the enclosed package insert. "Note: All shipments of E. coli/P.aeruginosa PNA FISH as of September 14, 2012, will contain this new package insert." The customers were also instructed to complete and return the Acknowledgement of Receipt form to AdvanDX: Regulatory Affairs; 400 TradeCenter, Suite 6990, Woburn, MA 01801-7476; Fax to: +1 (781) 376-0111; bsc@advandx.com. If you have any questions or requests for further information, please contact AdvanDx Technical Support: US (Toll-Free): 1-866-376-0009, techsupport@advandx.com. | Contact the recalling firm for information | 33 | 33 units | Worldwide distribution: USA (nationwide) including: CA, FL, IL, MI, and NJ; and country of: Great Britain. | Terminated | March 05, 2013 | 134 | 33 | Not_Computer
Z-0119-2013 | 63313 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite Tapered Certain Implant 5 X 10mm Model INT510Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Osseotite Tapered Certain Implant | 2 | October 23, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled their Osseotite Tapered Certain Implant, Model # INT510 due to a small percentage of the devices are missing the internal threading. | PRODUCTION CONTROLS: Process Control | Biomet 3 i sent an Urgent Medical Device Recall letter dated August 3, 2012, to all affected customers via e-mail, facsimile, telephone, or postal mail. Each customer was instructed to check their inventory for the affected product and provide feedback to Biomet 3i. Product is to be returned to:Biomet 3iP.M. Regulatory Compliance4555 Riverside DrivePalm Beach Gardens, Florida 33410For questions customers were instructed to call 1-800-342-5454. For questions regarding this recall call 561-776-6700. |  | 227 | 227 | Worldwide Distribution - USA including PA, GA, OR, KS, NY, VA, CT, NH, IL, OK, and MA and internationally to Australia, Brazil, Columbia, France, Germany, Ireland, Italy, Japan, Paraguay, Portugal, and Spain. | Terminated | December 20, 2012 | 58 | 227 | Not_Computer
Z-0120-2013 | 63296 | bracket, metal, orthodontic | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | The brand name of the device is Damon 3MX Bicuspid Bracket, an orthodontic bracket. Part Number 494-4491, Description -Damon 3MX Bicuspid Bracket, Lot Number - 011265039 Product Usage:Damon 3MX Bicuspid Brackets are intended to be a fixed attachment on a tooth which holds an archwire during orthodontic treatment. | Damon 3MX Bicuspid Bracket | 3 | October 23, 2012 | 2012 | Ormco/Sybronendo | The firm initiated the recall as a result of a misidentification of the brackets. Damon 3MX Bicuspid Brackets are produced with a purple colored ID dot to identify them as maxillary bicuspid brackets; however the affected lot of product listed below was produced with dark blue colored ID dots,which identified them as maxillary central incisor brackets. | PRODUCTION CONTROLS: Labeling Mix-Ups | Ormco/Sybronendo sent an Urgent Medical Device Recall letter via USPS 1~1 class mail, on April 24, 2012 to all consignees. Consignees were sent an updated customer notification via USPS lSI class mail on June 28, 2012. The letters identified the affected product, problem and actions to be taken. Consignees were instructed to complete the Acknowledgement/Return Form and to return any affected product in their inventory. Customers were instructed to contact Ormco Customer Care at 1-800-854-1741 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. | Contact the recalling firm for information | 148 | 148 | Worldwide Distribution - US Nationwide and the countries of Canada, China, Costa Rica, Ecuador, French Polynesia, India, Kuwait, Malaysia, Philippines, Singapore, South Korea, Thailand and UAE. | Terminated | December 12, 2012 | 50 | 148 | Not_Computer
Z-0132-2013 | 63316 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line. The MobiCath Bi-Directional Guiding Sheath is intended for introducing various cardiovascular catheters into the heart. Including the left side of the heart through the interatrial septum. | MobiCath BiDirecctional Guiding Sheath. | 2 | October 24, 2012 | 2012 | Greatbatch Medical | We are notifying you of a potential interaction between therapeutic catheters and MobiCath Bi-Directional Guiding Sheaths. We became aware of these interactions while investigating recent complaints regarding catheters "self-advancing" while being used in conjunction with the MobiCath Sheaths. | DESIGN: Device Design | The firm, Greatbatch, Ltd., sent an "URGENT CUSTOMER NOTIFICATION Medical Device - Voluntary Field Safety Notice" dated Monday, September 24, 2012 to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to read the "Description of the Problem" and "Recommendations for Clinical Use" sections carefully; complete and return the attached Field Action Acknowledgement Form via mail to: Biosense Webster, Incl, 15715 Arrow Highway, Irwindale, CA 91706; Attn: Recall Coordinator, or Fax to: (909) 839-8514; pass on this notice to anyone in your facility that needs to be informed; maintain awareness of this communication until information has been incorporated into the device labeling, and to avoid the catheter inside the MobiCath Bi-Directional Guiding Sheath from "self-advancing", physicians should reposition the catheter handle when manipulating the catheter in and out of the sheath. Greatbatch, Ltd. is updating the MobiCath Bi-Directional Guiding Sheath Instructions For Use (IFU) accordingly. Supplemental information will be provided with future shipments of these devices, until the revised IFU is available.For questions related to the issue please contact your Biosense Webster representative or call (800) 729-9010, Monday through Friday from 9:00 AM to 6:00 PM EST. For questions related to the Acknowledgment Form and its return, please contact Field Action Coordinator at (909) 839-8468. |  | 9040 | 9040 | Nationwide distribution: CA only. | Open |  |  | 9040 | Not_Computer
Z-0133-2013 | 63283 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DePuy Mitek VAPR Hook Electrode 3.5mm Hook Electrode with Integrated Handpiece Catalog Number: 227305The DePuy Mitek VAPR Electrodes for use with the VAPR Radiofrequency System are intended for resection, ablation, excision, hemostasis of blood vessels, and coagulation of soft tissue in patients requiring arthroscopic surgery of the knee, shoulder, ankle, elbow, and/or wrist. | VAPR Hook Electrode 3.5mm Hook Electrode with Integrated Handpiece | 2 | October 24, 2012 | 2012 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Product may have wrong large product identification label with symbols on the box indicating a S90 VAPR Electrode instead of the VAPR 3.5 hook electrode | PRODUCTION CONTROLS: Error in Labeling | The firm, Depuy Mitek, sent a "Medical Device Labeling Alert" letter dated March 14, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect your inventory to determine if any inventory of product code 227305 lot with the number M1108086 that is incorrectly labeled; review the following label information and the attached examples of correct and incorrect labels; and complete and return the Business Reply Form via fax to: 1-508-828-3750. To return the product call 1-877-379-4871 for an RGA and replacement product and return to DePuy Mitek, Attn: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324.Contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-48. |  | 376 | 376 units | Nationwide distribution: USA including states of: AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MD, MI, MN, NJ, NY, OH, OR, TN, TX, UT, VA, and WA. | Open |  |  | 376 | Not_Computer
Z-0137-2013 | 63302 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | AquaLiner Hydrophilic Guidewire 0.035" X 260cm, Straight/ Stiff.AquaLiner Hydrophilic Guidewire is designed for use in angiographic procedures to introduce/ position catheters and interventional devices within the vasculature. | AquaLiner Hydrophilic Guidewire | 2 | October 25, 2012 | 2012 | Nipro Medical Corporation | Nipro Medical Corporation, Miami, FL recalled their AquaLiner Hydrophilic Guidewire because the wires contain insufficient lubricity which causes them to get stuck in the housing and become nonfunctional. | PRODUCTION CONTROLS: Equipment Maintenance | Nipro Medical Corp. sent notifications to their consignee via e-mail and fax on July 27, 2012 and September 13, 2012 to their affected customer. The letters identified the product, problem, and actions to be taken by the customer. The letter instructs the customer to place all affected product in stock on hold. A Aqualiner-Recall Response form was attached for the customer to complete and return. Contact the firm at 305-599-7174 ext. 249 for questions regarding this notice. |  | 300 | 300 | Worldwide Distribution-Distributed in the state of New York and the countries of Spain, Germany, Argentina, Panama, Mexico, Chile, Colombia, Ecuador, Peru, and Brazil. | Terminated | May 25, 2013 | 212 | 3000 | Not_Computer
Z-0146-2013 | 63121 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SpiraLok 6.5mm with Ethibond without needlesProduct Number: 222966SPIRALOK Anchors are primarily used to re-attach soft tissue to bone in rotator-cuff repair surgery, but are also indicated for use in the foot/ankle, knee and elbow | SpiraLok 6.5mm with Ethibond without needles | 2 | October 25, 2012 | 2012 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Fracture of SPIRALOK Anchors post op, requiring patient revision | DESIGN: Device Design | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated September 6, 2012, to all affected customers. Dr Doctor letters were included with the recall letter. Recall is coordinated with Stericycle , contact 1-888-202-3694 The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the enclosed Business Reply Form, even if they have no product remaining in their inventory. Customers were instructed to isolate all inventory of the affected product and return to Stericycle at the following address or fax to 1-888-912-2189 with the completed Business Reply Form:Stericycle2670 Executive Drive, Suite AIndianapolis, IN 46241Attn: Event #8874Customers with questions were instructed to call 1-888-202-3694.For questions regarding this recall call 508-880-8100. | Contact the recalling firm for information | 94 | 94 units | Worldwide Distribution - USA (nationwide) and the countries of AE, AR, AT, AU, BE, BR, CA,, CH,CL .CN, CO, CZ,DE, DK, EC, ES, FI, FRGB, GP, GR, IE, IL, IN, IT, KR, LU, LY, MX, NC NL, NO, NZ, PE, PLPT, RE, RU, SG, SI, TH, TN, TR, UG ,UY , VE, and ZA. | Open |  |  | 78594 | Not_Computer
Z-0157-2013 | 63361 | adaptor, stopcock, manifold, fitting, cardiopulmonary bypass | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(k) | Gish 3/8 x 3/8" Straight Connector with Luer fitting, Item Numbers: EC2160B, EC2160S, Lot Numbers: 1210100010, 1215100026, 1217000138, Exp. Dates: 4/16/17, 5/31/15, 7/31/15.Product Usage: The connectors are indicated for use in connecting flexible tubing used in extracorporeal blood or fluid circuits. The connectors are distributed as sterile stand alones and bulk non-sterile. | Gish Connector with Luer fitting | 2 | October 25, 2012 | 2012 | Sorin Group USA, Inc. | Due to a molding irregularity, certain lots of connectors have the potential to leak fluid or entrain air at the interface of the luer fitting and mating cap. Entrained air could result in patient injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sorin Group USA, Inc. sent a Product Recall letter to all affected customers by certified mail on September 14, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to discontinue use of the affected product, return all unused product to the firm and to complete and return a Response Form via fax or email. For further questions, contact the firm at 1-800-650-2623. | Contact the recalling firm for information | 1640 | 1,640 units | US Nationwide Distribution in the states of: AL, AR, CA, CO, CT, FL, IL, MI, MN, MO, MS, NC, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | April 23, 2013 | 180 | 3020 | Not_Computer
Z-0156-2013 | 63379 | system/device, pharmacy compounding | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Repeater Pump, Class II device used for facilitating repeatable drug dosage distribution and reconstitution in hospital pharmacies. The device are Fluid Transfer Tube Sets, Sterile/EO, Std Volume Trifurcated, BAXA Corporation, Englewood, CO.The Repeater Pump is a device that provides a peristaltic pump driven fluid transfer that facilitates repeatable drug dosage distribution and reconstitution in hospital pharmacies. The device is for use with IV bags, syringes, elastomeric infusers, and other drug administration containers. | Repeater Pump | 2 | October 25, 2012 | 2012 | Baxa Corporation | The firm is issuing a field correction to the operators manual included with the 099 Repeater Pump because the Technical Manual was distributed with the device as opposed to the Operator Manual and the instruction for use contradict each other with regards to multiple source container use. | PRODUCTION CONTROLS: Packaging Process Control | Baxa Corporation sent a "REPEATER PUMP PRODUCT NOTICE" to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received a CD containing a copy of the correct Operator Mannual, 5300-0754. Contact Technical Support at 800-678-2292 for questions regarding this notice. |  |  |  | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Columbia, Guam, Japan, Malaysia, Mexico, Panama, Puerto Rico, Taiwan, UK, and Venezuela. | Terminated | February 15, 2013 | 113 |  | Not_Computer
Z-0139-2013 | 63321 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.5mm Cancellous Bone ScrewIntended to be used for the fixation of the acetabular shell or fracture repair. | 6.5mm Cancellous Bone Screw | 2 | October 25, 2012 | 2012 | Encore Medical, Lp | A product complaint was received where a size 20 screw was labeled as size 30. After receipt, measurement of the screw determined it to be a size 25. Further investigation also determined that there were other screws labeled as size 25, but a size 30 screw was inside of the package. It was thought that all of these screws were found in-house and were contained, but it was determined later that on | PRODUCTION CONTROLS: Labeling Mix-Ups | Encore Medical sent a Urgent Field Safety Notice dated September 28, 2012, to all affected customers.This notice needs to be passed on all those who need to be aware withinyour organization or to any organization where the potentially affecteddevices have been transferred.Please contact Customer Service to place a replacement order and receive an RMA number. Return all affected devices using the RMA number.For further questions please call (512) 834-6255. | No consumer action necessary | 1 | 1 unit | Nationwide Distribution only to CA. | Terminated | January 25, 2013 | 92 | 1 | Not_Computer
Z-0160-2013 | 62947 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000Product Usage:The QUADROX oxygenators are intended for the use in extracorporeal circulation during cardiopulmonary bypass in cardiac surgery. Within the specified flow rate range, the device oxygenates the blood, eliminates carbon dioxide and regulates blood temperature. The devices utilization period is limited to six hours. | QUADROXi(D) Oxygenator | 2 | October 26, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | It has come to firm's attention that, in some rare cases, the blood outlet connector on some models of MAQUET Oxygenators may detach from the oxygenator. This event has occurred in about 0.01% (1:10,000) of units shipped. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Maquet sent a Field Safety Notices and Confirmation Forms dated June 13, 2012 to all affected customers. A Follow-up communication dated July 20, 2012 was mailed to all affected US customers. The Field Safety Notice identifies the problem, affected product, risk factors and actions to be taken by the customer. Customers were instructed to complete the Confirmation form and return in the self-addressed prepaid envelop. For questions call 201-995-8968.Maquet issued an updated Field Safety Notice/Field Safety Notice Confirmation Form dated 2/14/13 via Fed Ex on 2/18/2013.Maquet sent out an Updated Field Safety Notice/field Safety Notice Confirmation Form dated 4/4/2013 in order to notify customers they would be providing safety clamps. |  | 23203 | 23,203 units | US Nationwide Distribution | Open |  |  | 23203 | Not_Computer
Z-0159-2013 | 62386 | dissector, surgical, general & plastic surgery | GDI | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. Product Usage:The product is used for pushing tissue from the incision or wound during medical procedures. Per the IFU, the indications for use are for gynecological laparoscopy, laparoscopic cholecystectomy, and other laparoscopic procedures. | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm) | 2 | October 26, 2012 | 2012 | Progressive Medical Inc | The tip of the product could come loose or unravel. |  | The recalling firm sent a letter dated October 28, 2011 to the customer who reported the problem and calls were initiated 6/26/12 to the remainder of the customers. The reason for recall was explained and the customer were asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For questions call 1-800-969-6331. | Contact the recalling firm for information | 124 | 124 boxes | US Nationwide Distribution in the states of: CO, IL, IN, KS, LA, MA, MN, MO, NE, OH, OK, TN, TX, VA, and WI. | Terminated | October 26, 2012 | 0 | 124 | Not_Computer
Z-0162-2013 | 63167 | system, test, immunological, antigen, tumor | MOI | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA Centaur BR Assay: Catalog Number:US: REF 02419937 (10214722) (50 Tests);REF 03896216(10334837)(250 Tests) NotCE Marked; OUS:REF03204829(10310271)(50 Tests); REF01157807(1 0314722) (250 Tests) CEMarked; REF 09010686 (10340081)(250 Tests).The ADVIA Centaur BR assay is an in vitro diagnostic test for the quantitative serial determination of cancer antigen CA 27.29 in human serum using the ADVIA Centaur and ADVIA Centaur XP systems. The test is intended for use as an aid in monitoring patients previously treated for Stage II or Stage Ill breast cancer. Serial testing for CA 27.29 in the serum of patients who are clinically free of disease should be used in conjunction with other clinical methods used for the early detection of cancer recurrence. The test is also intended for use as an aid in the management of breast cancer patients with metastatic disease by monitoring the progression or regression of disease in response to treatment. | Siemens ADVIA Centaur BR Assay | 2 | October 28, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | BR Assay for CA 27.29 Calibration Interval Change due quality control (QC) results outside of the expected ranges. |  | Siemens sent two "URGENT FIELD SAFETY NOTICE" letters dated August and September 2012 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was included for customers to complete and return. |  | 15455 | 15,455 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Columbia, Japan, Korea, Mexico, New Zealand, Peru, Saudi Arabia, and Spain. | Open |  |  | 15455 | Not_Computer
Z-0165-2013 | 63112 | electrode, pacemaker, temporary | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Myocardial Heart WiresIntended for temporary atrial and ventricular pacing and sensing during or after cardiac surgery in conjunction with an external pulse generator. | Myocardial Heart Wires | 2 | October 29, 2012 | 2012 | Oscor, Inc. | Oscor, Inc. in Palm Harbor, FL is recalling myocardial heart wires, models TME60S, TME64S, TME64C, TME64S-3, TME65S, and TME66S. The recall was initiated due to the spacing between anchoring zig-zag to electrode is out of specification. | TRAINING: Employee Error | Oscor sent a customer notification letter dated July 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to send the listed models and lots in their inventory back to the recalling firm for immediate replacement. They were also instructed to call Oscor Customer Service at 727-937-2511 to obtain a Return Goods Product (RGA) number.For questions regarding this recall call 727-937-2511, ext. 133. |  | 1425 | 1425 | Worldwide Distribution - USA including OR, FL, Al, TX, NM, CA, AL, PA, KY, IN, OH, and IL and internationally to Canada, Russia, and New Zealand | Terminated | December 17, 2013 | 414 | 1425 | Not_Computer
Z-0164-2013 | 63090 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***All in one CrossFT***Punch***PFT-00M***4.5mm to 6.5mm***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***"To create pathway for insertion of CrossFT Suture Anchor into bone. | All in one CrossFT Punch | 2 | October 29, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling the product due to the punch being mis-assembled with the incorrect handled. The correct handle is a clear anodized silver colored handle. | PRODUCTION CONTROLS: Process Control | Linvatec Corp sent a Urgent Medical Device Recall Notification letter dated April 9, 2012, to all affected customers. The letter was sent via FEDEx and by e-mail. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory for the recalled product and segregate them for return to the recalling firm. Furthermore, return the products along with the "REPLY FORM".Should you have any further questions please contact our Customer Service department at (800) 237-0169. |  | 2 | 2 units | Worldwide Distribution--US (nationwide) including state of Minnesota and the country of Thailand | Terminated | November 21, 2013 | 388 | 2 | Not_Computer
Z-0168-2013 | 63385 | orthosis, spondylolisthesis spinal fixation | MNH | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | ILLlCOÂŽ MIS Posterior Fixation System, Bone-screw Internal Spinal Fixation System,Orthosis, Spinal Pedical FixationOrthosis, Spondylolisthesis Spinal FixationThe ILLICO MIS Posterior Fixation System is intended to facilitate the surgical correction of noncervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. When used for a minimally invasive posterior approach ILLlCOÂŽ MIS Instrumentation is used in conjunction with polyaxial screw components. The ILLICO MIS Posterior Fixation System is intended for use as a posterior spinal fixation device to aid in the surgical correction of various spinal deformities and pathologies of the spine. It is intended to provide stabilization during the development of fusion utilizing a bone graft. Specific indications for the ILLICO MIS Posterior Fixation System are dependent in part on the configuration of the assembled device and the method of attachment to the spine. | Illico Ti Cannulated Polyaxial Screw | 2 | October 29, 2012 | 2012 | Alphatec Spine, Inc. | The firm initiated the recall because the stainless steel guide-wire of theIllico MIS Fixation System did not completely fit/pass through the tip of certain Illico TiCannulated Polyaxial Screws, therefore rendering such screws nonfunctional. | PRODUCTION CONTROLS: Process Control | Alphatec Spine sent an Urgent Medical Device Recall Notification letter dated September 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Alphatec Spine's Customers Service Department (customerservice@alphatecspine.com) immediately for instructions on how to return the affected product. Customers were asked to fill out the lower portion of the letter and fax to (760) 431-0289 to confirm that they have read the notification and will take all necessary removal actions. For questions customers were instructed to call 1-800-922-1356. For questions regarding this recall call 760-431-9286. | Contact the recalling firm for information | 209 | 209 screws | Nationwide Distribution including AZ, IL, CA, SC, TN, TX, GA, CT, NY, FL, OK, LA, MO, NV,and UT. | Terminated | March 29, 2013 | 151 | 209 | Not_Computer
Z-0163-2013 | 63272 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aviator Drill Bit (10 mm, 12 mm, 14 mm, 16 mm, Manufactured by: Stryker Spine SAS Z.1 Marticot -33610- Cestas France+33 (0) 5.57.97.06.30http://www.stryker.comDistributed in the USA by Stryker Spine2 Pearl Ct., Allendale NJ 07401-1677 USA+1-201-750-8000Drill bits are reusable instruments, sterilized prior to use, used in conjunction with an Aviator Drill Guide to provide a pathway for screw insertion. The aviator Anterior Cervical Plating (ACP) System is intended for use as an aid in cervical spinal fusion and is intended for unilateral fixation. | Stryker Spine | 2 | October 29, 2012 | 2012 | Stryker Spine | Aviator drill bits are 10 mm longer than the Reflex Hybrid drill bits and are not interchangeable which could lead to over-drilling by as much as 10 mm. | DESIGN: Labeling Design | Stryker Spine sent an Urgent Medical Device Field Correction Notification letter beginning July 5, 2012, to all affected customers. Teh letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the drill bit length in each respective drill guide prior to surgery to ensure the correct length protrudes through the guide. Customers with questions should call 201-760-8298. Customers were asked to report any adverse events or product quality problems to 1-888-457-7463. |  | 6154 | 6154 units | Nationwide distribution. | Terminated | January 28, 2013 | 91 | 6154 | Not_Computer
Z-0161-2013 | 60030 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro Elastomeric Infusion Pump.AccuFlux, Model # CT-0020-100H. Product Usage:The devices are intended for patients requiring intravenous, percutaneous, subcutaneous, intra-operative sites or epidural administration of medications. It is the responsibility of the user to ensure that the medication is prepared and administered accordance with the drug manufacturers package insert. The devices deliver controlled amounts of medication directly to the intraoperative site for pain management and or antibiotic administration. The devices infuse the medication at an hourly flow rate. Medications are infused intraoperatively and postoperatively through intramuscular or subcutaneous routes. The devices are also intended for controlled delivery of local anesthetics in close proximity to nerves for post operative regional anesthesia and pain management. Routes of administration may be intraoperative or percutaneous. It is for continuous infusion of medications for general infusion use, including antibiotic delivery and chemotherapy. | WOLFPAK, Medpro, Vipat | 1 | October 30, 2012 | 2012 | First Medical Source LLC | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. The recall was initiated because First Medical Source has confirmed that these lots may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to the patients. The pr | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. First Medical Source received an email from Medpro International with the following lnstruction to Customer - Informed customers to remove product from sales/discontinue use, cease distribution, remove the products from inventory and quarantine, inform their customers to retum the unused products to them, and destroy the product. Distributors' customer service will contact the customer to coordinate the return of the goods in question where required.First Medical Source emailed to their customer, Williams Medical Technologies, to informed them of the recall of the Medpro Elastomeric infusion pumps. Customer was instructed to inform them of the date and method of disposal of recalled products. |  | 500 | 500 units | US Nationwide distribution in the state of IL | Terminated | October 30, 2012 | 0 | 500 | Not_Computer
Z-0182-2013 | 63387 | elisa, antibody, west nile virus | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgG DxSelect" kit Catalog No. EL0300GKit Lots 122150, 122155, 122641, 122650 and 122859 that includes West Nile Virus IgG AntigenWells, Catalog Number EL0351, Lot Number 121117.510(k) No. K031953The Focus Diagnostics West Nile Virus ELISA IgG is intended for qualitatively detecting IgG antibodies to West Nile virus in human serum. In conjunction with the Focus Diagnostics West Nile Virus IgM Capture ELISA, the test is indicated for testing persons having symptoms of meningioencephalitis, as an aid in the presumptive laboratory diagnosis of West Nile virus infection. | West Nile Virus IgG DxSelect | 3 | October 31, 2012 | 2012 | Focus Diagnostics Inc | The firm recalled due to higher reactivity with some samples in the effected lots, which may reflect ahigher false positive rate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Focus Diagnostics Inc., sent a Urgent Medical Device Recall letter dated October 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you received at least one of the effected lots in the month of August 2012. Please document the number of kits in your inventory on the table below, even if zero (0), and fax to 562-240-6526.For further questions please call (714) 220-1900. |  | 425 | 425 units | Worldwide Distribution--US (nationwide) and the country of: Canada, Deuche, France, Greece, Mexico, Israel, and Japan. | Terminated | December 04, 2012 | 34 | 425 | Not_Computer
Z-0185-2013 | 63338 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Extremity Components:PROTHESE A EXPANSION COMPLETE D=6 LG 12 MM, Part Number PPO90000, Lot Numbers: U11113596, W12372884, W12376849, X01383482, X06391313, X07401148, X07401445, and X09418985PROTHESE A EXPANSION COMPLETE D=8 LG 15 MM, Part Number PPO90100, Lot Number X09418986PROTHESE A EXPANSION COMPLETE D=10 LG 18 MM, Part Number PPO90200, Lot Numbers: X06391315, X07401150, X07420387, and X09401447PROTHESE A EXPANSION COMPLETE D=12 LG 20 MM, Part Number PPO90300, Lot Numbers: W12383045, W12383485, X06391316, X06396622, X07401448, and X09418988TIGE HUMERALE PROXIMALE "NGR" LONG, Part Number PPP70088, Lot Number X01359005Orthopaedic | PROTHESE A EXPANSION COMPLETE D6 LG 12 MM | 2 | October 31, 2012 | 2012 | Wright Medical Technology Inc | There may be inaccuracies on the labeled shelf life on all products packaged at the Toulon, France facility. | PRODUCTION CONTROLS: Packaging Process Control | Wright sent an Urgent Product Recall Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and immediately return to Wright Medical Technology, Inc. Customers were asked to complete and return the enclosed notification form by fax to 901-867-7401. For questions customers should call 901-867-4324.For questions regarding this recall call 901-867-4655. |  | 840 | 840 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Austria, Belgium, Bulgaria, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Latvia, Lithuania, Norway, Poland, Portugal, Serbia, Slovakia, Slovenia, Spain, The Netherlands, Argentina, Brazil, Chile, China, Japan, Taiwan, Turkey, United Arab Emirates, Australia, South Africa, Russia, and United Kingdom. | Terminated | January 07, 2014 | 433 | 33135 | Not_Computer
Z-0175-2013 | 63347 | ventilator, emergency, manual (resuscitator) | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | AirFlow Resuscitator, Adult Resuscitator, AF1000 Series, Single use, Manufactured by: Ventlab Corporation, 155 Boyce Drive, Mockville, NC 27028The Disposable Resuscitators with optional built-in pressure monitor is intended to provide respiratory support in the presence of reversible apnea, commonly associated with cardiopulmonary arrest. | AirFlow Resuscitator, Adult Resuscitator | 1 | November 01, 2012 | 2012 | Ventlab Corporation | Possible volume leakage through the inlet valve during compressions of the resuscitator bag. | DESIGN: Component Design/Selection | VentLab Corporation sent an Attention: Recall Notification letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete and return the enclosed response form by fax to 336-753-5002, indicating whether or not their facility has any of the affected product. Customers can also scan and email the response form to CSR@Ventlab.com. For questions regarding this recall call 336-753-5000. | Contact the recalling firm for information |  |  | Nationwide Distribution including IL, OH, NY and TN. | Terminated | June 30, 2013 | 241 |  | Not_Computer
Z-0189-2013 | 63096 | mixture, hematology quality control | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(k) | Boule Con-Diff Hematology ControlFor in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. | Boule Con Diff US Tri Pack | 3 | November 02, 2012 | 2012 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled Boule Con Diff US Tri Pack (Product Part 501-605, Lot #1102-575 and 1102-576; Product Part 501-607, Lot #1102-576; and Product Part 502-012, Lot #1102-574) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was initiated due to the receipt of a complai | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Clinical Diagnostic Solutions, Inc. sent a letter dated March 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm istructed the accounts to replace the assay sheet they were currently using with the new assay sheets enclosed. For an unused product they would need to discard. To correct the ranges in their instrument, they would need to scan in the new bar codes specific for their instrument type. For any questions customers were instructed to contact CDE Technical Support atl 800-453-3328.For questions regarding this recall call 954-791-1773. | Contact the recalling firm for information | 623 | 623 vials | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MS, ND, ND, NE, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WI, WV,and WY. Foreign: Brazil and Colombia | Open |  |  | 1721 | Not_Computer
Z-0188-2013 | 63310 | somatic gene mutation detection system | OWD |  |  | Immunology |  | cobas 4800 BRAF V600 Mutation Test for In Vitro Diagnostic Use Roche Molecular Systems, Inc., Product Usage:The primary use of the cobas 4800 BRAF V600 Mutation Test is the detection of the BRAF V600 mutations in DNA extracted from formalin-fixed, paraffin-embedded human melanoma tissue. It is intended to be used as an aid in selecting melanoma patients whose tumors carry BRAF V600 mutations for treatment with Zelboraf (vemurafenib). | Roche cobas | 3 | November 02, 2012 | 2012 | Roche Molecular Systems, Inc. | It was identified that BRAF controls dispositioned as "Internal Use Only" were incorrectly packaged into cobas 4800 BRAF V600 Mutation Test CE-IVD kit lot R03607 and released for distribution. Two control batches manufactured during two process validation studies are involved:PV7711: BRAF Mutant (MUT) ControlPV0853: BRAF Wild Type (WT) Control | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Roche sent a Safety Board Notice - Product Bulletin 2012-013 on 31-Jul-2012 to affected customers. The notice identifed the affected product, description of situation and actios to be taken. Customers were instructed to discontinue use, quarantine and discard any remaining kits of the affected product in inventory. Customers were advised to contact their affiliate sales office for replacement. The notice states a new cobas 4800 BRAF V600 Mutation Test lot (R07749) is anticipated to be available on 03-Aug-2012. For questions call +01.908.253.7569. |  | 44 | 44 kits | Worldwide Distribution - Australia, Austria, Belgium, China, Hungary, Turkey and United Kingdom. | Terminated | September 11, 2013 | 313 | 44 | Not_Computer
Z-0192-2013 | 63309 | enzyme controls (assayed and unassayed) | JJT | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Cliniqa Liquid QCTM Cardiac Marker Control for i-STAT Level 1 (List Number 06F12-09; Lot Number 071039)Cliniqa Liquid QCTM Cardiac Marker Control for i-STAT Level 2 (List Number 06F12-10; Lot Number 081039)Cliniqa Liquid QCTM Cardiac Marker Control for i-STAT Level 3 (List Number 06F12-11; Lot Number 091039)Usage: Assayed human plasma which may be used to verify the integrity of newly received i-STAT CK-MB cartridges. | Cliniqa Liquid QCTM Cardiac Marker Control | 3 | November 02, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care (APOC) has determined that for the Cliniqa Liquid QC Cardiac Marker Control for i-STAT Level 1 (List #06F12-09; Lot Number 071039); i-STAT Level 2 (List #06F12-10; Lot Number 081039); and i-STAT Level 3 (List #06F12-11; Lot Number 091039) results may be generated below the lower value assignment range for some vials. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Abbott Point of Care, Inc. (APOC), sent an "URGENT RECALL NOTICE" dated July 2012 via FED-EX on July 19, 2012. The recall notice identified the product, problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of the control lot numbers; return all unused controls on the enclosed Customer Reply Form; If another institution was provided the controls, provide a copy of this notice to them. Note: if any used vials gave acceptable product performance, a review of the cartridge performance and associated patient results is not necessary)Note: APOC is shipping replacement products to customers facilities to minimize any disruption this might cause.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care representative. | Contact the recalling firm for information | 154 | 154 (06F12-09 (74) boxes; 06F12-10 (35) boxes and 06F12-11 (45) boxes.) 1 box contains six, 3ml plastic bottles. | Worldwide distribution: USA (nationwide) including states of: AZ, GA, ID, IL, LA, MD, MO, MN, NC, NV, OK, PA, TN, TX, VA and WI; and country of: Italy. | Terminated | June 20, 2013 | 230 | 154 | Not_Computer
Z-0202-2013 | 62945 | blade, scalpel | GES | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | "***Sharpoint IQ Geometry Slit Knife Angled, Bevel Up***Surgical Specialties Corporation Reading, PA***REF 72-2840***"Ophthalmic surgery | SHARPOINT IQ Geometry Slit Knives | 2 | November 05, 2012 | 2012 | Surgical Specialties Puerto Rico Inc. dba Angiotech PR Inc. | Potential for bent tips. | PRODUCTION CONTROLS: Process Control | Angiotech sent an Urgent Product Recall Notification letter dated August 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the current inventory, segregate the affected lot numbers, and to return the identified product. Customers were asked to complete the attached Customer Acknowledgement Form and fax to the Quality Assurance department at 787-658-1073. If product was further distributed customers should notify their customers to whom they sold the affected product. Customers with questions should call 787-658-1800 option 2 ext. 1775 or 787-658-1800 option 2 ext 1807.For questions regarding this recall call 787-658-1800. |  | 117948 | 117,948 units (total) | Worldwide Distribution - USA including DE,IL, NC, CA, CO, GA, KY, PA, MD, MN, NY, and MA and internationally to Austria, UAE, China, Argentina, Italy, Taiwan, Czech Republic, Australia, UK, India, Germany, Spain, France, South Africa, Malaysia, Japan, Lebanon, Korea, Greece, Saudi Arabia, Switzerland, Pakistan, Tunisia, Serbia, Morocco, Netherlands, Portugal, Truk, Belgium, and New Zealand | Terminated | January 13, 2014 | 434 | 117948 | Not_Computer
Z-0205-2013 | 63412 | instruments, dental hand | DZN | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Vector TAS Modular Driver, a manual driver for Vector TAS dental screws. | Vector TAS Modular Driver | 2 | November 05, 2012 | 2012 | Ormco/Sybronendo | The Instructions For Use (IFUs) for Vector TAS Modular Driver for some foreign language sections incorrectly lists the sterilization temperature as 130ÂżC. The correct sterilization temperature should be listed as 132ÂżC and/or 270ÂżF. | PRODUCTION CONTROLS: Error in Labeling | "Urgent - Medical Device Correction" letters dated 12/12/2011 were sent via USPS 1st class mail to all customers and again on 2/21/2012 for recall expansion. The device issue was described and recommended actions were provided. Customers were instructed to acknowledge and complete the Return Form. Questions were referred to (800) 854-1741. | Contact the recalling firm for information | 4543 | 4,543 | Worldwide distribution, including Nationwide (USA) and the countries of Canada, Austria, Belgium, Bulgaria, Switzerland, Germany, Denmark, France, Estonia, Spain, Finland, Iceland, United Kingdom, Greece, Croatia, Hungary, Ireland, Italy, Luxembourg, Netherlands, Monaco, Poland, Romania, Serbia, Slovenia, Turkey, Jersey, and Martinique. | Terminated | November 05, 2012 | 0 | 4543 | Not_Computer
Z-0207-2013 | 63396 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Trochar Scalpel Instrument (Surgical Punch)OrthoPediatrics intramedullary rods (nails) are generally rod-shaped devices, with screw holes at either end for fixation to bone. This device is intended to be inserted into the medullary canal of the femur for fixation of fractures by aligning and stabilizing the bone fragments. Additional stabilization may be realized by installing transverse screws through holes in the rod. These devices are made of medical grade stainless steel. The OrthoPediatrics PediNail" system is used for pediatric and small stature adult patients as indicated to stabilize fractures of the femoral shaft; subtrochanteric fractures; ipsilateral neck/shaft fractures; prophylactic nailing of impending pathologic fractures; nonunions and malunions; fixation of femurs that have been surgically prepared (osteotomy) for correction of deformity. | Trochar Scalpel Instrument (Surgical Punch) | 2 | November 05, 2012 | 2012 | OrthoPediatrics Corp | A complaint received from sales representative that a replenishment instrument was too short. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | PRODUCTION CONTROLS: Reprocessing Controls | OrthoPediatrics Corp sent an Urgent Recall e-mail dated September 14, 2012, to all affected customers. The e-mail identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all identified devices from the affected lot numbers immediately. Customers were also asked to complete the Mandatory Reply Form. Customers with questions should call 574-268-6379. |  | 78 | 78 | Nationwide Distribution including AL, IL, NY, FL, TX, MI, KY, NC, CO, and GA. | Terminated | September 18, 2013 | 317 | 78 | Not_Computer
Z-0193-2013 | 63094 | mixture, hematology quality control | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(k) | Boule Con Diff US Tri Pack, 6 Vials.For in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. | Boule ConDiff Hematology Control | 3 | November 05, 2012 | 2012 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled the Boule Con Diff US Tri Pack (Product part 501-605, Lot #1005-549 and Product part 501-607, Lot #1005-549) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was imitated due to receipt of customers stating the MCV values and RDW% recovered for Boule | PRODUCTION CONTROLS: Process Control | Clinical Diagnostic Solutions, Inc. sent a "PRODUCT NOTIFICATION:BOULE CON DIFF TRI-LEVEL LOT 1005-549" letter dated August 4, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact CDS Customer Service at 1-800-453-3328. |  | 532 | 532 packs | Nationwide Distribution. | Open |  |  | 1222 | Not_Computer
Z-0253-2013 | 62833 | urinary homocystine (nonquantitative) test system | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) Homocysteine Flex(R) Reagent CartridgeThis is an in vitro diagnostic test for the quantitative measurement of total homocysteine (HCYS) in human serum, heparinized plasma, and EDTA plasma on the Dimension Vista(R) System. Measurements of homocysteine aid in the diagnosis and treatment of patients suspected of having hyperhomocysteinemia and homocystinuria. | Siemens Dimension Vista(R) Homocysteine Flex(R) Reagent Cartridge | 3 | November 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Cartridges are leaking. The leaking is observed when opening the wrapper that holds the reagent cartridge. | PRODUCTION CONTROLS: Process Control | Siemens sent a Urgent Field Safety Notice letter dated July 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.At this time Siemens Healthcare Diagnostics is conducting a voluntary recall for Dimension VistaÂŽ HCYS FlexÂŽ reagent cartridge lots 12062MA and 12066MA. Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂŽ HCYS (K7044) lots 12062MA and 12066MA. Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information. We ask that you forward this notification to anyone to whom you may have distributed this lot.Contact your Siemens Customer Service Center to report your no-charge replacement needs at 888-588-3916.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 583 | 583 cartons | USA (nationwide) including the states of AL, CA, CO, CT, DC, DE, FL, IA, IL, IN, LA, MA, MD, MI, MO, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | January 15, 2014 | 435 | 583 | Not_Computer
Z-0213-2013 | 63103 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Wagner SL Revision Stem, uncementedSterile | Wagner SL Revision | 2 | November 06, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc., is initiating a correction to the patient labels of products manufactured before March 20I0 by the Zimmer GmbH production site in Winterthur, Switzerland, after receiving four complaints of a discrepancy between product patient labels and product labels with respect to the reference and lot numbers of the product. While the information printed on the product label is correct, the info | PRODUCTION CONTROLS: Error in Labeling | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Certificate of Acknowledgement Form was attached for customers to complete and return via fax to Zimmer, Inc. at 574 372-4265. |  | 87 | 87 | Nationwide Distribution. | Open |  |  | 23110 | Not_Computer
Z-0254-2013 | 62580 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Sysmex XE-2100 Automated Hematology System; an automated differential cell counter; Sysmex Corporation, Kobe, Japan; Sysmex America, One Nelson C. White Parkway, Mundelein, IL 60060; catalog #983-1341-1 The Sysmex XE-2100 is an automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The XE-2100 performs hematology analyses according to the RF/DC detection method, HydroDynamic Focusing (DC Detection), flow cytometry method (using a semiconductor laser), and SLS-hemoglobin method. | Sysmex XE2100 Automated Hematology System | 2 | November 06, 2012 | 2012 | Sysmex America, Inc. | There is a variation of reticulocyte counts between XE Series (XE-2100, XE-5000) and XT-Series (XT-2000i, XT-4000i) hematology analyzers that may result in a retic counts that are an average of 21% lower on the XE-Series because of a low bias. | DESIGN: Device Design | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated May 2012 to all affected customers on June 1, 2012. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to review patient results to determine if their Reference Intervals need adjustment, and to review QC results via Insight and document any change that may be linked to the date of calibration if appropriate. Any questions concerning the information contained in the notification were directed to the Sysmex Technical Assistance Center at 1-888-879-7639. | Contact the recalling firm for information | 1099 | 1,099 units | Nationwide distribution: USA including Puerto Rico. | Open |  |  | 1931 | Not_Computer
Z-0259-2013 | 63540 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | AMS AdVance" Male Sling System720088-01AMS AdVance" Male Sling System (US)720088-02AMS AdVance" Male Sling System (Intl)72404224AMS Male TO Sling System (Intl).The AMS AdVance Male Sling System is intended for the placement of a suburethral sling for the treatment of male stress urinary incontinence (SUI). | AdVance Male Sling System | 2 | November 07, 2012 | 2012 | American Medical Systems, Inc. | American Medical SystemsÂż (AMS) is issuing a voluntary recall on a number of lots of the AdVanceÂż Male Sling System (AdVance).AMS has recently received a small number of complaints regarding the resorbable suture in the AdVance sling not lasting for the entire 15 month shelf life of the product. An internal investigation has been conducted and confirmed that some sutures degraded prior to the | DESIGN: Device Design | American Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" dated October 18, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to return any existing stock of the recall product and to complete and return the Product Recall Acknowledgment Form. Contact AMS Customer Service at 1-800-328-3881 for additional information or questions.. |  | 4872 | 4872 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of Austria, Australia, Brazil, Canada, Germany, Estonia, Spain, Finland, France, Great Britain, Hong Kong, Croatia, India, Morocco, Mexico, Malaysia, Netherlands, New Zealand, Panama, Poland, Russian Federation, Solomon Islands, Singapore, Thailand, Turkey, Venezuela, South Africa, Chile, Colombia, Costa Rica, Slovenia, Argentina, China, and Taiwan. | Terminated | April 11, 2013 | 155 | 4872 | Not_Computer
Z-0258-2013 | 63598 | gown, surgical | FYA | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MicroCool Breathable Gown with Towels, X-Larg, X-Long, SterileThe product is a sterile, single use, surgical gown intended to protect surgical patients and operating room personnel from the transfer of microorganisms, body fluids, and particulate material. | MicroCool Breathable Gown with Towels | 2 | November 07, 2012 | 2012 | Kimberly-Clark Corporation | Package seam of some units may not remain properly sealed to ensure the contents might not be sterile as labeled. | PRODUCTION CONTROLS: Packaging Process Control | Kimberly Clark sent an Urgent Voluntary Product Recall letter dated July 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to determine if they have any affected product in their inventory. Any affected product should be quarantined and distribution discontinued. Distributors were also asked to conduct a sub-recall for any affected product further distributed. Customers were asked to complete the attached Distributor Product Recall Response Form and fax to 1-920-225-3522. Upon completion of the response form customers were asked to destroy any affected product.For questions regarding this recall call 770-587-7835. |  | 933 | 933 ((30 units per case) | Nationwide Distribution | Terminated | June 13, 2013 | 218 | 90789 | Not_Computer
Z-0268-2013 | 63297 | material, impression | ELW | Impression material. | Dental | Dental | 510(k) | The brand name of the device is AlgiNot, a dental impression material.Product Part Numbers AlgiNot Intra Kit 34683 AlgiNot Volume Refills 34682 Lot Number 1-1017AlgiNot is a Class II Medical Device. The device listing number this product is0136496 and the establishment registration number for Kerr Corporation is 1815757.Intended as an alternative to traditional alginate materials. | AlgiNot | 2 | November 08, 2012 | 2012 | Kerr Corporation | The firm recalled the product because it may cure harder than expected within theshelf life, which could result in a material that may be more difficult to remove. Use of this product onpatients with loose teeth may result in possible tooth fracture or tooth extraction. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Kerr Corporation sent "URGENT: MEDICAL DEVICE RECALL" letters to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. US, Canadian, and Rest of World consignees were notified via USPS 1st class mail on June 19, 2012, and Australianconsignees on July 5, 2012. Consignees were instructed to complete and return the Acknowledgement Form. Notification of the recall expansion was sent, via USPS 1st class mail, on February 21, 2012 to US, Canadian, Australia and European, consignees; however, once further investigation determined that all unexpired lots of the product were affected, the Alginot recall was expanded a second recall letter was sent to clarify that all lots of AlgiNot FS Volume recall and AlgiNot recall expansion are affected by this expandedrecall. Letters were mailed to US, Canadian and Rest of World (ROW) customers on June 19, 2012; Australian and European consignees were notified of the expansion on July 5, 2012.All consignees were instructed to complete the Return Form to acknowledge the updated instructions for use and the added cautionary statement. Consignees not responding are to be contacted at least three times via telephone/fax/e-mail as follow up on the recall. |  | 132 | 45 kits and 87 refills | Worldwide Distribution-USA (nationwide) and the countries of Canada, Europe, and Australia. | Terminated | November 15, 2012 | 7 | 132 | Not_Computer
Z-0263-2013 | 62271 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | 0.7/1.3U163CS-36In combination with below systemsMobileArt Evolution, MobileDaRt EvolutionProduct Usage:The device is a mobile X-ray system in order to take general radiography of patients in the hospital who cannot move and/or of outpatients in emergency. | MobileArt Evolution, MobileDaRt Evolution | 2 | November 08, 2012 | 2012 | Shimadzu Medical Systems | Shimadzu is recalling Diagnostic X-ray Beam-Limiting Device, Mobile x-ray system and Diagnostic x-ray tube assembly because they do not comply with the Product Performance Standards, 21CFR 1020.30 (m) Beam quality --(1) Half-value layer (HVL), which was revised in 2006. Certain devices might not comply with the requirements for them in some x-ray generating condition after the revision in 2006. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Shimadzu sent an Urgent Voluntary Medical Device Recall - Field Safety Notice letters dated September 19, 2012 to all the affected customers. The letter identifed the description of non-compliance, how to identify the affected products, corrective action plan by Shimadzu and actions to be taken customers. Customers were instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at Shimadzu Medical Systems National Technical Support (800) 228-1429. | Contact the recalling firm for information | 133 | 133 units | US Nationwide Distribution | Open |  |  | 1402 | Not_Computer
Z-0266-2013 | 63273 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Telescopic Strut ShortStrykerStryker Trauma AGCH-2545 SeizachDistributed in the USA by Howmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430.The use of external devices provides the surgeon with a means of bone fixation for the management of fracture and reconstructive surgery. These devices are intended only to assist healing and are not intended to replace normal bone structures. External fixation devices are intended to hold fractured bone surfaces in apposition to facilitate normal healing. While proven successful, the devices are manufactured from metal, plastic, and composite materials. No such fracture fixation device, subject to fatigue, can be expected to withstand activity levels in the same way as would a normal health bone. The external fixation system, therefore, will not be as strong, reliable or durable as a normal human bone. | Stryker Orthopaedics Hoffman LRF Telescopic Strut | 2 | November 08, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Three complaints have been filed where the Hoffman LRF Telescopic Struts have broken during load bearing application by patients having a body weight greater than 250 LBS. | DESIGN: Device Design | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE FIELD REMOVAL" letter dated July 2, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Business Reply Form was included for customers to complete and return to via fax to 1-865-252-3635. Contact the firm at 1-866-626-7747 for assistance. |  | 165 | 165 units | Nationwide Distribution. | Open |  |  | 165 | Not_Computer
Z-0269-2013 | 63599 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TI Standard Polyaxial Screw Assembly 6.5MM x 55MM. Part # 62065-55.Subsequent Product Codes: MNH, MNI.The Zodiac Polyaxial Pedicle Screw System is intended for use as a posterior spinal fixation device to aid in the surgical correction of various spinal deformities and pathologies in the thoraco-lumbo-sacral portion of the spine. It intended to provide stabilization during the development of fusion utilizing a bone graft. Specific indications for the Zodiac Polyaxial Pedicle Screw System are dependent in part on the configuration of the assembled device and the method of attachment to the spine. | Zodiac Pedicle Screw System | 2 | November 08, 2012 | 2012 | Alphatec Spine, Inc. | The recall was initiated because Alphatec Spine confirmed that the TI Standard Polyaxial Screw Assembly was laser marked with the incorrect length of 45 MM. | TRAINING: Employee Error | Alphatec Spine distributors who received the TI Standard Polyaxial Screw Assembly were contacted by phone and by email on December 9, 2010. A phone script was used as a recall notification to contact their three distributors who received the recalled products. The recall notification informed the customers of the problem identified and the action to be taken. Customers were informed that they will receive replacements within 2-3 days. Customers with questions were instructed to call (800) 922-1356.For questions regarding this recall call 760-431-9286. |  | 11 | 11 units | Nationwide Distribution including TN, WI, and FL. | Terminated | November 09, 2012 | 1 | 11 | Not_Computer
Z-0267-2013 | 53337 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Quadra S Offset Broach HandleThe offset broach handle is intended to attach to the broaches of different sizes as the surgeon broaches the femoral bone in preparation for the implantation of the femoral stem. | Broach Handle, manual orthopedic surgical instrument | 2 | November 08, 2012 | 2012 | Medacta Usa Inc | Medacta USA Inc has initiated a recall on the Quadra S Offset Broach Handle because of the potential of the rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medacta USA, sent a Urgent Safety Notification letter dated October 13, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.The letter informed the customers of the lot numbers on recall and the reference numbers. The recall was initiated because the offset broach handles contain a lever mechanism that is connected to the locking axis by means of a metallic wire. The rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. For further questions please call (858) 705-0350. |  | 12 | 12 units | Worldwide Distribution -- US including the states of Texas, Nevada and Colorado., and the country of Belgium. | Terminated | November 09, 2012 | 1 | 12 | Not_Computer
Z-0276-2013 | 63445 | dissector, surgical, general & plastic surgery | GDI | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. Exclusively Distributed in the U.S. Market by: Progressive Medical, Inc., St. Louis, MO, Made in China.The product is used for pushing tissue from the incision or wound during medical procedures. Per the IFU, the indications for use are for gynecological laparoscopy, laparoscopic cholecystectomy, and other laparoscopic procedures. | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm c 40cm) | 2 | November 09, 2012 | 2012 | Progressive Medical Inc | Some of the sterile packages had broken seals. | PRODUCTION CONTROLS: Packaging | Progressive Medical Inc contacted their customers via telephone on December 9, 2011, through December 15, 2011.Customers were informed the reason for recall was explained and the customer was asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For further questions please call (314) 961-5786. |  | 70 | 70 boxes | USA (nationwide) including the states of CO, IN, KS, MA, MN, MO, NE, OH, SC, TN, TX, VA, and WI. | Terminated | November 14, 2012 | 5 | 70 | Not_Computer
Z-0287-2013 | 53336 | bath, flotation, tissue | IDY | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | PT Link Waterbath, Model#PT100.Lab Vision PT Module Plus provides an automated standarized alternative to labor intensive and time consuming dewaxing and epitote retrieval procedures. | PT100/PT Link Waterbath | 2 | November 13, 2012 | 2012 | Dako North America Inc. | Dako is recalling the PT Link Waterbath because of a possible electrical hazard. Specifically, if the heater pad under the reagent tank becomes wet, it may present an electrical shock hazard when the lid is open. | DESIGN: Device Design | The firm, Dako, sent a "DAKO PT LINK ADVISORY NOTICE" dated May 1, 2008 to customers who purchased the PT Waterbath along with a customer acknowledgement document informing the users of the potential electrical hazard. The notice described the product, problem and actions to be taken. The customers were instructed to follow the attached Safety Instructions; place the enclosed warning labels directly on the instrument as shown in the attached instructions and adhere to the warning labels as described. Customers with questions were instructed to call Dako Products' Technical Support at (800) 424-0021. |  | 136 | 136 units | Nationwide distribution: USA including states of: AK, CA, CO, CT, DC, FL, GA, HI, IA, IL IN, KS, KY, MA, MD, MI, MN, MO, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, and WA. | Terminated | November 13, 2012 | 0 | 136 | Not_Computer
Z-0288-2013 | 63586 | system, endovascular graft, aortic aneurysm treatment | MIH |  |  | Cardiovascular |  | TriVascular brand Fill Polymer Kit, Endovascular Graft, Aneurysm Treatment System, European Model Number: TVFP14-B, (in European distributor inventory only); PRODUCT Usage:The TriVascular Ovation Abdominal Stent Graft System is indicated in subjects diagnosed with an aneurysm in the abdominal aorta having vascular morphology suitable for endovascular repair, including: (1) Adequate iliac/femoral access compatible with vascular access techniques, devices, and/or accessories, (2) Non-aneurysmal proximal aortic neck: - with a length of at least 7 mm proximal to the aneurysm, - with an inner wall diameter of no less than 16 mm and no greater than 30 mm, and - with an aortic angle of d 60 degrees if proximal neck is e 10 mm and d 45 degrees if proximal neck is | TriVascular Fill Polymer Kit | 2 | November 13, 2012 | 2012 | Trivascular, Inc | Selected lots of the Fill Polymer Kit (European Model Number: TVFP14-B) in European distributor inventory only that potentially contain leaking stopcocks due topossible cracks in the stopcock component. TriVascular indicated that these Fill Polymer Kits are for use with the Ovation or Ovation Prime Abdominal Stent Graft Systems that may display leaks upon initiation of mix but prior to use in | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Trivascular sent an Urgent Field Safety Notice letter dated October 19, 2012 to all affect consignees. A verbal (telephone) communication was also initiated. The forms communication identified the affected product, problem and actions to be taken. Consignees were instructed to immediately quarantine and return affected products. A return response form was included for consignees to complete and returned via the e-mail address or fax number provided. Consignees were instructed to further notify its (sub-level) consignees, and have them return affected product directly to the recalling firm. For questions call 1-707-543-8732. |  | 41 | 41 units | Worldwide Distribution - US Nationwide and the countries of Belgium, Cyprus, Italy, Spain, Switzerland and Turkey. TriVascular stated that no purchased devices are in hospital inventory in Europe or the U.S. | Terminated | December 17, 2012 | 34 | 41 | Not_Computer
Z-0277-2013 | 63446 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation System Tube to Tube Clamp, Part Number 390.007.Intended for use to provide treatment for long bone and pelvic fractures that require external fixation. Specifically, the components can be used for: Stabilization of soft tissues and fractures; Polytrauma/multiple orthopedic trauma; Vertically stable pelvic fractures, or as treatment adjunct for vertically unstable pelvic fractures; Arthrodeses and osteotomies with soft tissue problems - failures of total joints; Neutralization of fractures stabilized with limited internal fixation; Non-unions/septic non-unions; Intra-operative reductions/stabilization tool to assist with indirect reduction; and Unilateral rectilinear bone segment transport or leg lengthening. | Synthes External Fixation System Tube to Tube Clamp, Part Number 390.007 | 2 | November 13, 2012 | 2012 | Synthes USA HQ, Inc. | The insert provided with the product for the specific lot number indicates that this product is MR safe. This product is MR Conditional and is etched as MR Conditional. The insert is incorrect. | PRODUCTION CONTROLS: Labeling Mix-Ups | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated September 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Verification Form via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this notice. |  | 8 | 8 | Nationwide Distribution-including the states of CA, MN, NC, and TN. | Open |  |  | 8 | Not_Computer
Z-0278-2013 | 63344 | controller, temperature, cardiopulmonary bypass | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HX2 Temp Manage SYS 120VThe Terumo HX2 Temperature Management System is a selfcontained, dual channel system for supplying temperature controlled water to heat exchangers used in an extracorporeal circuit. | HX2 Temp Manage SYS 120V | 2 | November 13, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Based on a retrospective review of quality data, AC Control boards used in the manufacture of theHX2 Temperature Management System were identified as being manufactured using the incorrectflux material. Use of the incorrect flux may contribute to premature board failure. A search ofTCVS complaints revealed 3 complaints received related to the HX2 failure. Each failure occurredduring set up | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | TERUMO sent an Urgent Medical Device Recall letter dated October 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo has determined that a field correction is required to notify affected customers that the boards may fail, instruct the users that in a case of failure to swich to the second channel on the device or use a back-up unit, and that TCVS will replace the affected boards. Customers with questions were instructed to call 1-800-621-2818. For questions regarding this recall call 734-741-6173. |  | 19 | 19 units | Nationwide Distribution including CA, HI, KS, LA, MI, MO, OR, PA, UT, and WI | Terminated | July 11, 2013 | 240 | 34 | Not_Computer
Z-0286-2013 | 63275 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Special Orthopaedic Solution Peri-Acetabular ProsthesisV40 Femoral Extension ComponentProduct Usage:The PAR extension piece is a component of the PAR assembly. The Peri-Acetabular Reconstruction (PAR) endoprosthesis was developed to meet the unique needs of patients who require reconstruction of large acetabular defects after a Type I and II pelvic resection. | Stryker Orthopaedics | 2 | November 13, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that insufficient taper torsional strength may result when a PAR extension piece is used with a forged CoCr stem. | DESIGN: Device Design | Stryker Orthopaedics sent an Urgent Product Recall letter dated July 16, 2012 via Fed Ex with return receipt to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were instructed to fax back the attached Product Recall Acknowledgement Form to 201-821-6069. For questions call 201-972-2100. |  | 301 | 301 units | Worldwide Distribution - US Nationwide | Open |  |  | 301 | Not_Computer
Z-0289-2013 | 63536 | alkaline picrate, colorimetry, creatinine | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Creatinine 120 CP, Model Number Creatinine Part # A11A01868, Lot Numbers: 203273706, 203273707, 203273808, 203273709. The ABX PENTRA Creatinine 120 CP is an in vitro diagnostic assay for the quantitative determination of creatinine in human serum, plasma and urine based on a kinetic method using alkaline picrate (Jaffe method). It is composed of a 27 ml monoreagent cassette. Reagent is a chemical solution with additives.ABX PENTRA Creatinine 120 CP reagent, with associated calibrator and controls, is a diagnostic reagent for quantitative in vitro determination of Creatinine in human serum, plasma and urine based on a kinetic method using alkaline picrate (Jaffe method). Creatinine measurements are used in the diagnosis and treatment of renal diseases, in monitoring renal dialysis, and as a calculation basis for measuring other urine analytes. This reagent is intended for use on the ABX PENTRA 400 Clinical Chemistry Analyzer. | ABX PENTRA Creatinine 120 CP | 3 | November 14, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | This recall was initiated because some customers were experiencing an alarm during Creatinine calibration "REAG RANGE HIGH". This alarm caused the calibration to fail and prevented the user from using the Creatinine assay. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | HORIBA Medical sent customer communication letters dated October 24, 2011, to all affected customers via Fax with the exception of two communications that were emailed. The response or confirmation of receipt was conducted via telephone and also documented on attached Notification List. Customers were given a list of procedures to follow and asked to complete and fax the Acknowledgement of Receipt Form to +33 (0)4 67 14 15 17.For questions regarding this recall call 949-453-0500, ext 208. | Contact the recalling firm for information | 8503 | 8,503 units | Nationwide Distribution | Terminated | November 29, 2012 | 15 | 8503 | Not_Computer
Z-0295-2013 | 63603 | vinyl patient examination glove | LYZ | Patient examination glove. | General Hospital | General Hospital | 510(k) | Select Medical Products Vinyl Powder-Free Exam Gloves, Medium Model # 6370 and Large Model # 6371, 100 per box; 10 boxes per case.Usage: Patient Examination. | Vinyl PowderFree Exam Gloves | 2 | November 15, 2012 | 2012 | Gulf South Medical Supply Inc | Gulf South Medical Supply recalled their Select Medium and Large size Vinyl Exam Gloves because they were inadvertently shipped prior to U.S. Customs clearing them for sale. | TRAINING: Employee Error | The firm, PSS World Medical, sent an "URGENT RECALL NOTICE" dated June 20, 2012 via USPS to all customers that received the product. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, quarantine the affected lot numbers listed in this recall and return the product; If the product was further distributed, to inform their customers of the recall and request they return the affected product directly to them; and contact their Gulf South Customer Care team via telephone at 1-800-347-2456 or email to customercare@gsms.com to request a return goods authorization.If you have any questions, contact the Product Integrity Leader Quality Assurance at 904-380-4501. |  | 292 | 275 cases and 17 boxes | Nationwide distribution: USA including states of: DE, KY, LA, MD, NJ, NY, PA, TX, VA, and WV. | Open |  |  | 292 | Not_Computer
Z-0292-2013 | 63322 | agent, absorbable hemostatic, collagen based | LMF | Absorbable hemostatic agent and dressing. | General & Plastic Surgery | General & Plastic Surgery |  | ETHICON SURGIFLO HEMOSTATIC MATRIX KIT Plus FlexTip with Thrombin distributed by ETHICON ETHICON Inc. Somerville, NJ 08876 USA Ref 2993SURGIFLO Hemostatic Matrix, mixed with sterile saline or thrombin solution, is indicated for surgical procedures (except ophthalmic) for hemostasis, when control of capillary, venous and arteriolar bleeding by pressure, ligature, and other conventional procedures is ineffective or impractical. This is a topical absorbable hemostat consisting of a flowable gelatin matrix (paste) and a lyophilized Thrombin component. This kit also contains components necessary for preparing and applying the matrix during surgery. | ETHICON SURGIFLO HEMOSTATIC MATRIX KIT Plus FlexTip with Thrombin | 1 | November 16, 2012 | 2012 | Ethicon, Inc. | SURGIFLO Hemostatic Matrix Kit & SURGIFLO Hemostatic Matrix Kit with Thrombin have an issue within the packing process where a cut could potentially breach the double Tyvek pouch of the packaging. | PRODUCTION CONTROLS: Packaging | Ethicon Inc. sent an Urgent Product Information letter dated August 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory to determine if they have any of the affected product. Affected product should be returned for replacement to Stericycle using the enclosed pre-paid shipping label. Customers with questions should call 1-877-384-4266 or contact their Ethicon Biosurgery Sales Representative.For questions regarding this recall call 877-384-4266. |  | 28276 | 28,276 | Worldwide Distribution - USA (nationwide) and the countries of Russia, China, and the European Union. | Open |  |  | 73476 | Not_Computer
Z-0343-2013 | 63581 | agent, absorbable hemostatic, collagen based | LMF | Absorbable hemostatic agent and dressing. | General & Plastic Surgery | General & Plastic Surgery |  | Superstat Modified Collagen Hemostatic Sponge, 50mm Disc, 3.0%, Part# 9200-30.For use when hemostasis is desired along suture lines, in diffusely bleeding sites, in incisions, dissections, and around arterial or venous anastomoses. Indicated in oozing bleeding, either capillary or venous, in almost all types of incisions or wounds, and as an adjunct to conventional closure techniques | Superstat Modified Collagen Hemostatic Sponge | 2 | November 16, 2012 | 2012 | Superstat Corp | The recall was initiated because Superstat Corporation has confirmed that no documentation are available to verify that quarterly dose sterility audits were performed for the Superstat Modified Collagen Hemostatic Sponge. | PRODUCTION CONTROLS: Process Control | Superstat initially contacted customers by telephone and sent an Urgent - Device Recall Notification Letter on October 24, 2012, to all affected customers. The recall letter provided the reason for recall was due to inadequate documentation of quarterly dose sterility audits. The recall includes all lots and part numbers of Superstat produced between 2002 and 2012. Customers are instructed to complete and return the enclosed response form as soon as possible. Customers with questions are instructed to call customer service at (800) 487-3786.For questions regarding this recall call 310-605-1655. |  | 26422 | 26,422 units total | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, Thailand, Mexico, Canada, Pakistan, Israel, Greece, Serbia, Canda, Isle of Man, Nova Scotia, and Singapore. | Terminated | February 06, 2013 | 82 | 26422 | Not_Computer
Z-0330-2013 | 63481 | electrode, circular (spiral), scalp and applicator | HGP | Fetal scalp circular (spiral) electrode and applicator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Corometrics Qwik Connect Plus Spiral Electrode- Model #7000AA0A fetal scalp circular (spiral) electrode and applicator is a device used to obtain a fetal electrocardiogram during labor and delivery. | Corometrics Qwik Connect Plus Spiral Electrode | 2 | November 16, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | Thirty fetal scalp electrodes were found to have an incomplete package seal and lacked the printed lot number stamped on during sealing. | TRAINING: Employee Error | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter on October 16, 2012 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: DO NOT use Qwik Connect Plus Spiral Electrodes part number 7000AAO with lot number S03123 or product with missing lot numbers; isolate all affected product; complete and return the Urgent Medical Device Correction Confirmation form via fax at 800-535-7923, return form even if no recall product is in inventory; to return only unused product, and if they have forwarded any affected lot numbers of this product to any other healthcare institutions, forward a copy of this letter to those institutions.If you have any questions or concerns regarding this notification, contact GEMSIT Customer Service at +1-800-588-7044 (Domestic) or Vital Signs Customer Service at +1-800-932-0760 (International). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 126 | 126 cases (50 pouches per case) | Worldwide distribution: USA (nationwide) including states: AZ, CA, CT, FL, GA, IL, KS, KY, MA, MD, MN, MO, NJ, NM, NY, OH, PA, SC,TX, WA, and WY; and countries of: Australia and Canada. | Open |  |  | 126 | Not_Computer
Z-0297-2013 | 63618 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Product is Full OSSEOTITE 2 Tapered Certain Implant, Model Number XIFNT413, and size 4 x 13 mm. The product is sterile.Product provides a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Dental implant | 2 | November 16, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomet 3i sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 28, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this notice. |  | 191 | 191 devices | Worldwide Distribution-USA (nationwide) including the states of VA, SC, HI, LA, OH, FL, IL, NJ, NC, CA, CT, MI, TN, AK, NH, NY, MA, GA, PA, and KY and the countries of United Kingdom, Spain, Germany, Canada, Japan, Singapore, Taiwan, Israel, Poland, Hong Kong, Ireland, The Netherlands, Austria, and France. | Open |  |  | 267 | Not_Computer
Z-0344-2013 | 62916 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | IO Fix IntraOsseous Fixation System; Extremity Medical. www.extremitymedical.com; Customer Service: 888.499.0079.The Extremity Medical Screw and Washer System is intended for reduction and internal fixation of arthrodeses, osteotomies, intra- and extrarticular fractures and nonunions of the small bones and joints of the foot, ankle, hand and wrist. | IO FIX IntraOsseous Fixation | 2 | November 18, 2012 | 2012 | Extremity Medical LLC | Product Event Reports revealed that lag screw passes through washer (e.g. X-Post) intra-operatively. | TRAINING: Employee Error | Extremity Medical sent an e-mail notification (Extremity Medical- Sales Bulletin- IO Fix) dated April 13, 2011 to distributors notifying them of a change made to the 4.0 mm Lag Screw. |  | 120 | 120 Units | Worldwide Distribution-USA (nationwide) including the states FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE, and the country of Switzerland. | Terminated | March 21, 2013 | 123 | 120 | Not_Computer
Z-0348-2013 | 52634 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Screw Tulip , 6.25mm RodSubsequent Codes: KWP, KWQ, MNH, MNI | SpheRx Posted Pedicle Screw System | 2 | November 19, 2012 | 2012 | NuVasive Inc | The tulip portion of the screw assembly can disengage from the shank after implantation. | DESIGN: Device Design | Telephone communications to the sales representatives of surgeons actively using the Posted Pedicle Screw System were made on 01/05/2009. The calls were unscripted and requested that all inventory be forwarded to NuVasive's San Diego office for quarantine. Surgeons who were actively using the system were also contacted and notified of the situation. Follow-up communications were sent via mail on 01/19/2009 to all surgeons who participated in the Alpha release. The letter informed the customers of the problems identified. Customers were instructed that the recalling firm does not believe any further action is required and they were instructed to monitor the progress of their patients and notify NuVasive if any issues are detected.Customers with questions or concerns are instructed to contact the firm at (858) 909-1964. |  | 240 | 240 units total | Nationwide Distribution-including the states of OR, MO, FL, WI, CA, and NC. | Terminated | November 19, 2012 | 0 | 240 | Not_Computer
Z-0379-2013 | 62987 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Thermo Scientific MAS Omni IMMUNE and Omni IMMUNE PRO;Liquid Assayed Integrated Immunoassay Control.Intended for use as an assayed control for monitoring assay conditions in many clinical laboratory determinations. Assay values are provided for the specific systems listed. The user can compare observations with expected ranges as a means of assuring consistent performance or reagent and instrument.. | MAS Omni IMMUNE and Omni IMMUNE PRO | 3 | November 19, 2012 | 2012 | Microgenics Corporation | In house testing confirmed that CK-MS and Insulin claimed in both Omni IMMUNE and Omni IMMUNE PRO products are displaying reduced stability. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Microgenics Corporation sent a Recall Notification letter dated August 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed that immediate attention is required and identify the two affected constituents. The letter contains a list of the affected product with their catalog number, lot number and expiration date and request that the CK-MB and Insulin use be discontinued. The remaining components of the product may be used. Evaluate quality controls according to your laboratory procedures.Retain a copy of the letter for your laboratory records. If the product was forwarded to another laboratory, provide them with a copy of the letter.Complete the attached Recall Response form and return within 5 days.Questions may be directed to 1-800-232-3342. |  | 2076 | 2076 total in US , 596 total Foreign | Worldwide Distribution--USA (nationwide) and the countries of Germany, Canada and Australia. | Terminated | March 04, 2013 | 105 | 2076 | Not_Computer
Z-0384-2013 | 63077 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***SENTINEL***Cannulated Drill Bit***S8585***8.5 X 229mm (9")***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***LOT***2009-02***".This drill bit is an eccentric mono-fluted cutting edge for drilling bone tunnels in ACL and PCL reconstruction. | Surgical Bit | 2 | November 19, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling the S8580, 8.0MM SENTINEL Cannulated Drill Bit, lot numbers 322225 and 322226; and the S8585, 8.5MM SENTINEL Cannulated Drill Bit, lot number 322287. Recall was initiated due to the incorrect labeling on package. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | ConMed/Linvatec sent an Urgent Medical Device Recall Notificaiton letter via phone and FedEx on April 9, 2012 and again on May 11, 2012, to all affected customers. They also sent the same Notification on May 14, 2012, to international accounts. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete a reply formand fax it to 727-319-5701. Customers were also instructed to return all affected product to the firm. All returned product under this removal action will be identified, segregated and quarantined per firm procedures, etc.Customers with questions should contact Customers Service at Custserv@linvatec.com.For questions regarding this recall call 727-392-6464. |  | 70 | 70 total | Worldwide Distribution - USA including NJ, UT, GA, NY, PA, IL, LA, OH, CA, MS and Internationally to Canada, Japan, and Romania. | Terminated | November 21, 2013 | 367 | 70 | Not_Computer
Z-0345-2013 | 63505 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel Dxl 800 Access Immunoassay System, Part Numbers: 973100, A71456, and A71457.The UniCel DxI Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. The UniCel DxC series of instruments is a DxI Immunoassay System integrated with a chemistry analyzer. | UniCel Dxl 800 Access Immunoassay System | 2 | November 19, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the UniCel Dxl and DxC Access Immunoassay Systems had field modifications to the aspirate and dispense probe assemblies. | DESIGN: Process Design | Beckman Coulter sent a field modification was released in August 2011 which included:- Installing a new aspirate and dispense tubing routing bracket.- Replacing existing aspirate and dispense probe tubing with modified tubing to help eliminatemisrouting of aspirate and dispense probe tubing.During installation the FSE informed the customer:-That aspirate and dispense tubing are different lengths.- That probes are not interchangeable with other instruments once tubing has been modified.- To route probe tubing by referring to the labels on the tubing.An updated service document, Routine Maintenance Changes, A82294B was provided to every customer.For questions regarding this recall call 714-961-4483. |  | 1756 | 1,756 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, AU, BG, BW, CH, CN, CZ, DE, ES, FR, GB, HK, HR, HU, IN, IT, JO, JP, LB, MA, MX, NL, PL, QA, RO, RU, SK, TN, TR, ZA, AO, AR, AT, BE, CO, DZ, FI, IE, IL, KR, KW, LI, NA, NO, NZ, PF, PH, PR, SE, SG, SZ, TH, TW, VE, VN, and YT. | Terminated | November 19, 2012 | 0 | 2188 | Not_Computer
Z-0387-2013 | 63358 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Torx 6.5mmCancellous Bone ScrewRef 2030-6525-1The 6.5mm Cancellous Bone Screws are used in THA when the surgeon elects to use a cluster/multi-hole shell in the acetabulum. The screws are intended to provide initial cup fixation and help prevent spin-out of the cup due to physical loading from the patient. They are intended as supplemental fixation. | Torx 6.5mm Cancellous Bone Screw | 2 | November 19, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received reports from the field indicating that certain lots of 6.5mm Cancellous Bone Screws are either associated with a product mix or potentially associated with a product mix. Packaging associated with a 35 mm screw may contain a 25mm screw. Conversely, packaging associated with a 25mm screw may contain a 35mm screw. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The firm, Stryker Orthopaedics, notified Domestic locations of this action by e-mail on August 6, 2012 and foreign locations of this action by e-mail on August 7, 2012. Stryker follow-up by sending an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated August 17, 2012 to their consignees/customers via Fed Ex with return receipt. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and return affected product to Attn: Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430 (attach fluorescent orange sticker and mark box with words "Product Recall"), or contact your Stryker Sales Representative to arrange for return of the product ,and complete and return the attached Product Recall Acknowledgment Form within 5 days via fax to 855-251-3635.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 72 | Approximately 72 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Belgium, Finland, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Poland, Portugal, Spain, Turkey and UK. | Open |  |  | 430 | Not_Computer
Z-0382-2013 | 63386 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SPECIALTY AUDIBLE TORQUE WRENCH, XIAIS2217XLPMFG BY: STRYKER SPINE2 PEARL COURTALLENDALE NJ 07401-1677The Xia 3/Mantis Redux/Specialty Audible Torque Wrenches are used to final tighten the blocker of the Xia 3/ Mantis Redux Polyaxial Screws, Hooks, Monoaxial Screws, Rod to Rod Connectors and Off-set connectors. The Torque Wrench is a t-handle deflection beam torque wrench with a 5 mm hex tip. The instrument allows the controlled application of 12 Nm torque via a visual indicator on the bell shaped cylinder just below the handle. The Specialty Audible Torque Wrench has an extra feature where it creates and audible "click" when the desired final tightening torque is reached. | SPECIALTY AUDIBLE TORQUE WRENCH | 2 | November 19, 2012 | 2012 | Stryker Spine | Multiple complaints have reported the hex tip of the torque wrench fracturing during final tightening. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker Spine sent an Urgent Medical Device Removal Notification letter/Product Accountability Forms dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 201-760-8150. | Contact the recalling firm for information | 28 | 28 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, France, Germany, Netherlands, Japan, Spain, Denmark, Switzerland, and the United Kingdom. | Open |  |  | 101 | Not_Computer
Z-0395-2013 | 63066 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | 00780302201 MIS 12/14 Slotted Femoral Head Provisional -2X2200780302202 MIS 12/14 Slotted Femoral Head Provisional +0X2200780302203 MIS 12/14 Slotted Femoral Head Provisional +3.5X22The provisional heads replicate final femoral implant geometry and are used to trial range of motion, joint stability, and determine appropriate femoral head size and offset during total hip arthroplasty | MIS Slotted Femoral Head Provisional | 2 | November 19, 2012 | 2012 | Zimmer, Inc. | Zimmer is in the process of enhancing the surgical technique instructions for the Versys and Slotted Provisional Femoral Heads. These enhancements inform users that the head may become difficult to visually locate if it becomes displaced deep or medial to the iliopsoas tendon and provide non-invasive imaging methods that can be used to locate the head. The observed rate of occurrence for this typ |  | The firm, Zimmer, sent notification by email to consignees beginning October 10, 2012 with a follow-up letter titled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete a spreadsheet of current users and contact information and immediately provide this information via email to: CorporateQuality.Postmarket@zimmer.com, and review and become familiar with the content and recommendations of the attached letter sent to the customers. For questions call 1-877-946-2761. | Contact the recalling firm for information | 56516 | 56,516 all lots, all sizes | Worldwide distribution: USA (nationwide) and countries of: Asia, Barbados, Canada, Europe, Mexico, Puerto Rico, San Juan, and South America. | Open |  |  | 56516 | Not_Computer
Z-0404-2013 | 63596 | system, therapeutic, x-ray | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Balloon Applicator Set;3.0 cm,Carl Zeiss Meditec AG, Oberkochen, GermanyINTRABEAM System is intended for radiation therapy treatments | INTRABEAM Balloon Applicator Set | 2 | November 20, 2012 | 2012 | Carl Zeiss Meditec, Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carl Zeiss Meditec sent an Important Field Corrective Action letter dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return the affected product to Carl Zeiss Meditec, Inc. A Carl Zeiss Meditec representative will contact the customer to schedule a visit to support the removal of the affected product. For questions customers were instructed to call 925-580-5371 or by email at v.venturo@meditec.zeiss.com.For questions regarding this recall call 925-557-4616. |  | 36 | 36 tootal, all lots | Nationwide Distribution including IL and MD | Terminated | November 20, 2012 | 0 | 36 | Not_Computer
Z-0409-2013 | 62805 | bed, flotation therapy, powered | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Kinetic Therapy SystemINDICATIONS: Treatment and prevention of pulmonary complications as a result of immobility, Thoracic or lumbar fracture, Cervical traction, Skeletal traction | RotoRest Delta Kinetic Therapy System | 2 | November 20, 2012 | 2012 | KCI USA, Inc. | The firm is initiating a voluntary medical device correction for the RotoRest Delta Therapy System head pack. KCI has received complaints alleging patients experiencing skin breakdown on the head and face due to the positioning of the head pack. In a few instances, the patient received injuries ranging from Stage III to Stage IV pressure ulcers. | DESIGN: Component Design/Selection | KCI sent an Urgent - Voluntary Medical Device Correction letter dated June 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that KCI would contact them soon to coorinate a shipment of a replacement of the affected product. For any questions customers were instructed to call 1-800-275-4524, Selection Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. |  | 282 | 282 units | Nationwide Distribution including AL, CA, CO, OK and GA. | Terminated | April 22, 2013 | 153 | 282 | Not_Computer
Z-0386-2013 | 63126 | prosthesis, hip, constrained, cemented or uncemented, metal/polymer | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trilogy II Constrained Liner The Trilogy Longevity Constrained Liner serves as the articular surface for a femoral head in total hip arthroplasty. | Trilogy | 2 | November 20, 2012 | 2012 | Zimmer, Inc. | Trilogy Longevity Constrained Liner devices were packaged in incorrect cartons. The product is packaged in the incorrect carton. Product labels and patient labels are correct, matching the product inside the carton. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Zimmer contacted their customer beginning the week of August 27, 2012, via telephone. Zimmer informed the customer the reason of the recall. The customer was instructed to quarantine and return the device to Zimmer Product Service, 1777 W. Center St. Warsaw, IN 46580. The customer was asked to provide information if the product was further distributed. For further questions please call 574-372-4316. | Contact the recalling firm for information | 1 | 1 | USA (nationwide) including the states of Ohio, Carolinas and New England. | Terminated | August 27, 2013 | 280 | 1 | Not_Computer
Z-0401-2013 | 63557 | bone grafting material, synthetic | LYC | Bone grafting material. | Dental | Dental | 510(k) | BioPlant, a dental bone grafting material. BioPlanthas been assigned the product code LYC (Bone Grafting Material, Synthetic) by theFDA and is classified as a Class II Medical Device. BioPlant is a Class II Medical Device. The Device Listing Number is 0007254.BioPlant is a synthetic, multipurpose bone-grafting material intended to be used in dental applications to maintain and restore the bony alveolar ridge in post-extraction tooth-root sockets or after oral surgery, ridge preservation, ridge augmentation, augmentations in the maxillary sinus, repair of periodontal defects and bone voids associated with dental implant placement. | BioPlant | 2 | November 20, 2012 | 2012 | Kerr Corporation | The firm initiated the recall becaue disturbing the BioPlant product soon after it is placed may halt or delay the ossification process. | DESIGN: Labeling Design | Kerr Corporation sent an Urgent Medical Device Correction letter and acknowledgement form dated October 25, 2012, to all affected customers via USPS First Class. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: "...review the table with product description , part numbers and lot numbers to determine if they have any of the affected product in their inventory and confirm that they were aware of the additional placement instructions necessary to achieve desired ossification results detailed in the paragraph in the letter. Customers were instructed to contact Kerr Customer Care at 1ÂŽ800ÂŽ537ÂŽ7123 if they have any questions. Customers were also instructed to complete and fax back the enclosed Acknowledgement Form to 714-288-4609 in order to confirm their receipt of this correction regardless of whether they have any affected product in their inventory.For questions regarding this recall call 714-516-7400. | Contact the recalling firm for information | 22410 | 22,410 units | Worldwide Distribution - USA (nationwide) Canada, Europe, Australia and ROW consignees. | Terminated | September 23, 2013 | 307 | 22410 | Not_Computer
Z-0407-2013 | 63532 | probe, ent | KAK | Ear, nose, and throat manual surgical instrument. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Vital Signs Disposable General Purpose 9 French Temperature ProbeFor monitoring of core temperature | VITALTEMP General Purpose Probe | 2 | November 20, 2012 | 2012 | Vital Signs Colorado Inc. | There is a potential for epistaxis during use of the Disposable General Purpose 9 French Temperature Probe. | DESIGN: Device Design | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter and Confirmation form on August 17, 2012 via Fed Ex to its consignees/customers. The letter describes the product, problem and action to be taken. The customers were instructed to do the following: Do not use the Disposable General Purpose 9 French Temperature Probe P/N 1016 or 1016EU; return all unopened product(s) to Vital Signs Devices (VSD) 11039 East Lansing Circle, Englewood, CO 80112-note: replacements and/or refunds discussed in letter are available only in exchange for the return of any unopened product(s); complete and return the URGENT MEDICAL DEVICE CORRECTION CONFIRMATION via fax to: 800-535-7923; forward a copy of the letter to any institutions they may have forwarded the product and for customers using the GXXXXX64 kit that includes the Disposable General Purpose 9 French Temperature Probe follow the directions listed in the letter.Note: VSD is removing p/n 1016 and 1016EU, Disposable General Purpose 9 French Temperature Probe, from the market and offering a replacement. If you have any questions or concerns regarding this notification, contact Customer Service at +1-800-932-0760 (Domestic & International). Hours of Operation 8:00 am EST to 6:00 pm EST. | Contact the recalling firm for information | 243685 | 243,685 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL,GA, IL, LA, MA, MI, MO, NC, NH, NY, OH, OK, PA, RI, SE,TX, and WA; and countries of: AUSTRALIA, BARCELONA, GERMANY, NEW ZEALAND, NETHERLANDS, SWITZERLAND, and UNITED KINGDOM. | Terminated | June 14, 2013 | 206 | 243685 | Not_Computer
Z-0410-2013 | 63122 | prosthesis, hip, hemi-, femoral, metal/polymer, cemented or uncemented | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | MultipolarÂŽ Bipolar Cup Liner 22 mm LD.; Item 00-5001-040-22Product Usage:The liner snaps into the MultipolarÂŽ Bipolar Cup and serves as the articulation surface for the femoral head. The split poly locking ring helps contain the head within the liner and also provides a popping sound as conformation that the femoral head has completed seated, as noted in surgical technique 97-5003-003-00 | Multipolar | 2 | November 20, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the MultipolarÂż Bipolar Cup Liner 22 mm J.D. (item 00-5001-040-22) due to reports that the split poly ring was incorrectly assembled. | DESIGN: Process Design | Zimmer sent an Urgent Medical Device Recall letter to all affected consignees during the week of September 3, 2012. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, quarantine the recalled product and return remaining inventory. Customers were advised to complete Inventory Return Certification form attached. |  | 195 | 195 | Worldwide Distribution - USA nationwide including the states of: AK, CA, IL, FL, HI, MI, MO, NC, NJ, NY, OH, PA, TX and the countries of Germany, Brazil, Canada, Japan, Taiwan, Thailand, India and Korea. | Terminated | September 23, 2013 | 307 | 195 | Not_Computer
Z-0402-2013 | 63512 | light, surgical headlight | EBA | Dental operating light. | Dental | Dental | 510(K) Exempt | Freedom Cordless LED Light System. Freedom System, Standard Length, Part Number 910863-1, Serial Numbers 77001000-77003470 and Freedom System, Long Length, Part Number 910863-2, Serial Numbers 77001000-77003470.Freedom Cordless LED Light System is a Class I Medical Device. The devicelisting number this product is E115952.The intended use of this device is to illuminate oral structures and operating areas. | Surgical Headlight | 2 | November 20, 2012 | 2012 | Kerr Corporation | The firm recalled the Freedom Cordless LED Headlight because some units were assembled with an incorrect component in thecontrol board electronics that may cause some Freedom Cordless LED Headlight units to become warm to the touch; however, no skin damage will occur. | PRODUCTION CONTROLS: Labeling Mix-Ups | Orascoptic sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 17, 2011. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return to the firm. Contact Orascoptic Customer Care at 608-831-2555 for assistance regarding this notice. |  | 1330 | 1330 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, and Europe. | Terminated | November 20, 2012 | 0 | 1330 | Not_Computer
Z-0411-2013 | 63629 | tube, tympanostomy | ETD | Tympanostomy tube. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | PARPARELLA-TYPE VENT TUBE, 1.02 MM, SILICONE, REF 240044, QTY 6, Rx Only, STERILE EO, GYRUS ACMI INC.,Tympanostomy tube | PARPARELLATYPE VENT TUBE, 1.02 MM | 2 | November 20, 2012 | 2012 | Gyrus Acmi, Incorporated | One lot of tympanostomy tubes were manufactured with silicon that was found to be contaminated with ethylene vinyl acetate copolymer. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Olympus, sent an "URGENT: Medical Device Recall" letter dated October 15, 2012 to their consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect their stock to confirm if any of the affected lot of product remains in their possession; immediately cease any further use of any affected product; remove it from their stockroom and quarantine it until it is shipped back to Gyrus ACMI, Inc.; call their Gyrus ACMI customer service representative (800-773-4301) to obtain a Returned Goods Authorization and promptly return any affected product to Gyrus ACMI, Inc., 2925 Appling Rd, Bartlett, TN 38133., and complete and return the RECALL REPLY FORM to Attn: Regulatory Affairs via fax to: (508) 804-2624, even if there is no affected stock left at your facility.If you have any questions, contact the customer service at 800-773-4301. |  | 10 | 10 boxes/60 units | Worldwide distribution: USA state of Kansas and country of Korea. | Terminated | April 09, 2013 | 140 | 10 | Not_Computer
Z-0408-2013 | 62835 | enzyme immunoassay, theophylline | KLS | Theophylline test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Systems Vista(R) Theophylline THEO Flex(R) Regent Cartridge (K4071)The THEO method is an in vitro diagnostic test for the quantitative measurement of Theophylline in human serum and plasma on the Dimension Vista(R) system. | Dimension Vista(R) Systems Vista(R) Theophylline THEO Flex(R) Reagent Cartridge (K4071) | 2 | November 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received complaints from customers of "Abnormal Reaction" errors occurring on calibrations, QC and patient samples. The Abnormal Reaction flag is being triggered by an internal check in the method for atypical particle reagent aggregation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens Healthcare Diagnostic sent a Urgent Field Safety notice dated July, 20. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂŽ Theophylline FlexÂŽ reagent cartridge (THEO, Catalog # K4071) lot 12135AE and/or 12171BB.Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information and to facilitate your additional no charge replacement.We ask that you forward this notification to anyone to whom you may have distributed this lot.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. For questions concerning the status of your no charge replacement order please contact your Siemens Customer Service Center at 888-588-3916.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 959 | 959 cartons | USA (nationwide) including the states of AL, AZ, CA, CO, DC, DE, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, MS, MT, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, VA, VT, WA and WI. | Terminated | January 08, 2014 | 414 | 959 | Not_Computer
Z-0427-2013 | 63580 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LAD Flexible ClearSeal SINGLE USE, King Systems ETO Sterilie, 10 devices/case"Size 4 / Adult ClearSeal Flexible King LADÂŽ Patient Size: 50-70kg "Product Usage:Laryngeal airway | KING LAD SILI CONE | 2 | November 21, 2012 | 2012 | King Systems Corp. | The firm received two complaints from a user facility in New Jersey regarding color-coded EtO sterilization indicators on the primary device packaging that did not undergo the expected color change. The complainant reported devices that had pink color indicators when the product packaging specified that the color should change from pink to yellow (or brown) if the product had subjected to an EtO |  | King Systems sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated October 31, 2012 to all affected consignees via email. The letter identified affected product, problem and actions to be taken. The letter states that further use and/or distribution of any remaining product affected by this recall should cease immediately. The letter instructs customers to destroy all of the affected products and indicate the destruction by signing and returning the attached forms using the enclosed self-addressed envelope. For questions contact your King Systems customer service representative at 800-642-5464. | Contact the recalling firm for information | 263 | 263 cases of product. (2,630 devices) | US Nationwide Distribution - WI, TX, NY, MO, IN, KY, FL, NJ, PA, TN, and NC | Terminated | May 28, 2013 | 188 | 263 | Not_Computer
Z-0412-2013 | 63144 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spine Frontier Inspan compressorsPart Number: 11-60004.Inspan compressors provide pressure to drive spikes of Inspan implants into the spinous process bone, resulting in fixation of the implant. | Spine Frontier Inspan Compressor | 2 | November 21, 2012 | 2012 | SpineFrontier, Inc. | Inspan Compressors may break at weld causing a loss of compression | DESIGN: Device Design | Spine Frontier issued e-mail notices to customers on May 14, 2012 advising of the potential for breakage. Units were replaced with weld free design. |  | 17 | 17 units | Worldwide Distribution-USA (nationwide) including PR and the states of FL, MA, MO, OK, and TX, and the country of Mexico. | Open |  |  | 17 | Not_Computer
Z-0413-2013 | 63633 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | CapLOX II Final Torque Driver, Indicator is a Pedicle screw spinal system.The CapLOX II Spinal System is a posterior, non-cervical pedicle fixation system intended to provide immobilization and stabilization of spinal segments in skeletally mature patients as an adjunct to fusion in the treatment of the following acute and chronic instabilities or deformities of the thoracic, lumbar, and sacral spine including degenerative spondylolisthesis with objective evidence of neurologic impairment, fracture, dislocation, scoliosis, kyphosis, spinal tumor, pseudoarthrosis and failed previous fusion. | CapLOX II Pedicle Screw System | 2 | November 21, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine recalled their CapLOX II Final Torque Driver, Indicator due to the device prematurely breaking | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Captiva Spine contacted the customer by phone to have them return the device.For questions regarding this recall call 561-277-9480. |  | 10 | 10 | Nationwide Distribution (Distributed to one customer located in MN) | Terminated | January 11, 2013 | 51 | 10 | Not_Computer
Z-0428-2013 | 63500 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | IMRIS ORT100 (Part Number 110470-000) and ORT200 (Part Number 112216-00), are tables used in both the Neuro II-SE and Neuro III-SV Magnetic Resonance Imaging Systems. | IMRIS ORT100 and ORT 200 | 2 | November 21, 2012 | 2012 | Innovative Magnetic Resonance Imaging Systems Inc. | IMRIS Inc. initiated a recall of their OR Tables ORT100 and ORT200. When proper procedure is not followed, the foot end of the IMRIS intra-operative OR table may be attracted to the MRI magnet as the magnet approaches the table. This poses a potential risk of injury to both the patient and operators standing near the table. | DESIGN: Labeling Design | Consignees were sent on 10/19/12 a IMRIS "Urgent Medical Device Correction" letter dated October 16. 2012. The letter described the product and the problem involved in the recall. It also provided changes to the Operator's Manual and labeling which addresses the issue. Customers may contact 1-866-475-0525. |  | 23 | 23 | Nationwide Distribution, including the states of FL, GA, IN, MD,MA, MI, MN, MO, TN, UT, and VA. | Terminated | December 11, 2012 | 20 | 23 | Not_Computer
Z-0416-2013 | 63541 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trilogy Acetabular System Shell with Cluster Holes Porous 56 mm O.D. Sterile The Trilogy Acetabular System is indicated for either cemented or uncemented use In skeletally mature Individuals undergoing primary or revision surgery for rehabilitating hips damaged as a result of noninflammatory degenerative joint disease (NIDJD) or its composite diagnoses of osteoarthritis, avascular necrosis, protrusio acetabuli, traumatic arthritis, slipped capital epiphysis, fused Hip, fracture of the pelvis, and diastrophic variant. The shells with screw holes permit the use of titanium alloy screws to provide additional fixation and security, particularly in those cases where acetabular bone stock is deficient. | Zimmer Trilogy | 2 | November 21, 2012 | 2012 | Zimmer, Inc. | Zimmer TrilogyÂż Acetabular System Shell with Cluster Holes, Porous,56 MM O.D., sterile; is missing polar hole threads. Zimmer received one complaint in which no injury was reported. | PRODUCTION CONTROLS: Process Control | Zimmer sent a Urgent Medical Device Recall Notice dated September 20, 2012, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customer. Customers were instructed to locate quarantine and return affected product to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 for product in the USA. Customers were asked to complete the Inventory Return Certification Form and follow the instructions where to return this form.For further questions please call (574) 372-4807. |  | 95 | 24 USA; 71 worldwide | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CT, KS, MA, MD, MN, MO, NY, SD, TX and WI and the countries of Bulgaria,Canada, Germany, Italy, Netherlands, Russian Fed, Sweden and UK | Open |  |  | 95 | Not_Computer
Z-0429-2013 | 63539 | mixture, hematology quality control | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(k) | Minotrol 16, Model Numbers: 2043748, 2043752, 2043756, Lot Number: MX090(Levels: Low, Normal, and High).Minotrol 16 is a tri-level control intended for in vitro diagnostic use in monitoring the accuracy and precision on HORIBA Medical hematology blood cell counters using HORIBA Medical reagents. Minotrol 16 is an in vitro diagnostic reagent composed of human erythrocytes, mammalian leukocytes, and platelets in a plasma-like fluid with preservatives. Minotrol 16 is a stable material that provides a means of determining the accuracy and precision of hematology blood cell counters. It is handled in the same manner as a patient specimen. The assay tables are determined on validated instruments using the appropriate reagents | Minotrol 16 | 3 | November 21, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical Irvine Technical Support received reports from some customers that have experienced high recover values for Hemoglobin when using Minotrol 16 lot #MX090 (Level High). Customer communication was sent 10/01/2010. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | HORIBA Medical sent a Product Information Notice letter dated November 5, 2010 via Fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instructions included with the Product Information Notice. For questions customers were instructed to call the Technical Support Hotline at 1-888-903-5001.For questions regarding this recall call 949-453-0500, ext 208. |  | 22954 | 22,954 | Nationwide Distribution | Terminated | April 22, 2013 | 152 | 22954 | Not_Computer
Z-0414-2013 | 63516 | loupe, diagnostic/surgical | FSP | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | The brand name of the device is Laser Loupes, a dentalloupe with laser eye protection. Laser Loupes have been assigned product code 79 FSP and are classified as a Class I Medical Devices.Laser Loupes areavailable with four (4) different filters of different colors. Each colored filter isassociated with specific laser frequency ranges. The device listing number for this product is E160359.The intended use of this device is to magnify the operatory field and provide laser eye protection. | Laser Loupes | 2 | November 21, 2012 | 2012 | Kerr Corporation | The firm initiated the recall for Laser Loupes because the ink used to mark which laser frequency each laser Loupe protectsagainst may wipe off. The loupes work according to specification, but the marking may not permanently adhere tothe device. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Finished Device Change Control | Customers were notified via USPS 1st class mail on April 10, 2012 to US and Canadian, and Rest of World (ROW) consignees, and on April 27, 2012 to European consignees. The notification letter identifies the product, reason for recall, and instructions to customers:"Please review the table below to verify that the laser loupes being used in your office has the correct filter color. ORASCOPTIC KINDLY REQUESTS YOUR COOPERATION IN COMPLETING AND FAXING BACK THE ENClOSED ACKNOWLEDGEMENT FORM IN ORDER TO CONFIRM YOUR RECEIPT OF THIS FIELD CORRECTION.If you are an authorized Orascoptic distributor, please inform your affected customers of this notification within forty-eight (48) hours of receipt.". Customer contact information was also provided:"Please contact Orascoptic Customer Care at 608.831.2555 or toll free in the United States and Canada at 1-800369-3698 for more information.". |  | 639 | 639 units | Worldwide Distribution--USA (nationwide) and the countries of Canada and EU. | Terminated | November 21, 2012 | 0 | 639 | Not_Computer
Z-0415-2013 | 63613 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | Ranawat/Burnstein Acetabular Shell 54mm/Limited Holes/Porous Coated is a metallic, full-hemisphere, acetabular shell component. Intended Use: The shell utilizes a modular polyethylene liner and a femoral head component that is taper fit onto a femoral stem at the time of surgery. Cemented or non-cemented total hip replacement . The device is intended for general use in skeletally mature individuals undergoing primary and/or secondary revision surgery at high risk of hip dislocation due to a history of prior dislocation, joint or bone loss, soft tissue laxity, neuromuscular disease, or intra-operative instability and for whom all other options to constrained acetabular components have been considered | Ranawat/Burnstein Acetabular Shell | 2 | November 21, 2012 | 2012 | Biomet, Inc. | The color code on label conflicts with the liner size (liner size 24 should be green, not yellow) | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | BIOMET called distributors who received the affected product to return it to Biomet with follow-up Urgent Recall Notices sent via E-mail October 22, 2012.. The notice included the identity of the recalled device, reason for the recall, and directions to complete the FAX back form and return it to 574-372-1683. Product is to be returned to Return Goods, Recalls, Biomet, Inc.56 East Bell Drive, Warsaw, IN 46580. Questions related to the recall should be directed to Field Action Coordinator (574) 372-1570, Monday through Friday. 8 a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. | Contact the recalling firm for information | 16 | 16 (5 USA) | Nationwide Distribution including CA, IN, NC, OH, and UT. | Terminated | March 18, 2013 | 117 | 16 | Not_Computer
Z-0417-2013 | 63075 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | S-Lift Implant Guides (also known as S-LIF)Compressor Part Number: 11-50223Product Usage:Surgical Guides allow a surgeon to place implants into tight spaces, such as a vertebral disc space, without having the implant contact the surrounding bone or tissue. In the case of S-LIFT Implant Guides, they are used to place vertebral bodies into the disc space without the body contacting the superior or inferior vertebrae. | SLift Implant Guides | 2 | November 21, 2012 | 2012 | SpineFrontier, Inc. | When using the optional S-LIFT (also known as S-LIF) Implant Guides without stops, there is a known risk of moving the Implant Guides, and possibly the implant, too far into the disc space creating the potential for injury to the patient. The instructions for Use for the S-Lift Implant Guides will be amended to add the following warning: "Care is to be taken to not over insert the Implant Guides | DESIGN: Labeling Design | Spine Frontier issued a Correction Notice Letter August 22, 2012 to all affected customers to update the Instructions for Use. The letter identified the affected product, problem and actions to be taken. Customers were Instructed to remove the Instructions for Use they currently have and replace them with the update. The letter instructs customers to complete and return the Response Form to 978-232-3990. |  | 19 | 19 units | US Nationwide including the states of: MA and TX. | Open |  |  | 19 | Not_Computer
Z-0432-2013 | 63554 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 4.5mm PLEO Osteotomy Guide.OrthoPediatrics Corp. markets a pediatric, extension osteotomy bone plate system. This system is marketed under the brand name OrthoPediatrics PediLock Extension Osteotomy Plate ( PLEO ). The product included in this recall is a Class I, 4.5mm PLEO Osteotomy Guide Instrument, common name: Guide, Surgical, Instrument. The PediLocÂŽ Extension Osteotomy Plate is part of the OrthoPediatrics~PediLocÂŽ system. The PediLocÂŽ Extension Osteotomy Plate consists of 3.5mm and 4.5mm plates, left and right side specific. The PediLocÂŽ Extension Osteotomy Plate is designed to fit the anatomy of the distal femur in children and adolescents, above the distal femoral physis. The distal section of the plate is designed so that three locking screws can be placed in the distal bone segment at an angle parallel to the distal femoral physis. The proximal portion of the plate is designed to accept locking or non-locking screws. The subject instrument of this recall is an adjunct instrument for the convenience of some surgeons who choose to use it. This instrument is not illustrated in the Surgical Technique, nor required to perform the procedure. | 4.5mm PLEO Osteotomy Guide. | 2 | November 22, 2012 | 2012 | OrthoPediatrics Corp | Complaint received from sales representative that the drill guide tubes would not screw into replenishment instrument. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | OrthoPediatrics sent an "URGENT RECALL" e-mail of the Trochar Scalpel Instrument to all affected customers. The e-mail identifies the product, problem, and actions to be taken by the customers. The e-mail states that the instrument has been manufactured incorrectly and will not work (function as intended). The threads in the instrument are faulty and the guide tubes will not thread in correctly. Customers will receive follow up information. |  | 24 | 24 | Nationwide Distribution-including the states of OH, NC, CO, FL, IN, MI, and TX. | Terminated | September 18, 2013 | 300 | 24 | Not_Computer
Z-0430-2013 | 63479 | [6 empty] | 0.9% Sodium Chloride Injections, USP, 1000ml, VWR/BaxterProduct Usage:The concentrated sodium chloride is used to prepare inoculum that is analyzed by the VITEK 2 instrument. The VITEK 2 user manual instructs all users to prepare inoculum with 0.45%-0.50% sodium chloride injections, and the system is designed to use 0.45%-0.50% sodium chloride concentration only. | Sodium Chloride (NaCl) Injection | 2 | November 22, 2012 | 2012 | Biomerieux Inc | The firm inventoried and distributed 0.9% Sodium Chloride Saline bags as 0.45% Sodium Chloride Saline bags for use with the VITEK 2 system. | TRAINING: Employee Error | Biomerieux sent an Urgent Product Removal Notice letter dated October 16, 2012 to all affected customers via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected product and immediately discard any remaining inventory. The letter included instructions on receiving product replacement or credit. For questions or concerns contact your local BioMerieux representative. |  | 199 | 199 cases/14-1000mL bags | Worldwide Distribution - USA (nationwide) including the states of: AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, PA, RI, TN, TX, VA and WI. Distribution was also made to Canada, military facilities and other foriegn countries. | Terminated | December 05, 2013 | 378 | 199 | Not_Computer
Z-0434-2013 | 63631 | monitor, carbon-dioxide, cutaneous | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec V-Sign Membrane Changer Insert (VS-MCI), 5 pieces. The SenTec V-Sign Membrane Changer Insert is indicated for use with the V-Sign Membrane Changer only.The V-sign Sensor is indicated for use with the SenTec Digital Monitor when continuous non-invasive monitoring of carbon dioxide tension, oxygen saturation, and pulse rate are required for adult and pediatric patients. | VSign Membrane Changer | 2 | November 22, 2012 | 2012 | Sentec AG | A manufacturing defect was observed for the membrane changer insert that does not allow electrolyte to be present under the new membrane. The PCO2 performance of a sensor without electrolyte will be deteriorated. | DESIGN: Component Design/Selection | SenTec notified the one US distributor via e-mail on November 5, 2012. The distributor was advised to stop distribution of affected lot numbers. The communication to the distributor described the issue, identified affected product, and stated that SenTec will replace current stock at the end-customer level. The end user can discard the affected products. The distributor is asked to respond back with an e-mail to SenTec that they took appropriate action. |  | 265 | 265 boxes (5 per box) | Distributed only in MO. | Terminated | December 18, 2012 | 26 | 460 | Not_Computer
Z-0431-2013 | 62834 | agent, absorbable hemostatic, collagen based | LMF | Absorbable hemostatic agent and dressing. | General & Plastic Surgery | General & Plastic Surgery |  | Styker Orthobiologics Vita Plasma Separator, a component of Vitagel Surgical Hemostat.Sterile, single use device designed to prepare plasma from the patient's blood using a tabletop centrifuge. | Styker Orthobiologics Vita Plasma Separator, a component of Vitagel Surgical Hemostat | 2 | November 22, 2012 | 2012 | Stryker Orthobiologics, Inc. | When preparing plasma using VitaPrep, it was reported that a small amount of blood collected in the bottom of the centrifuge cup following the centrifugation step. Firm has since discovered that if blood is present in the Luer connection of the VitaPrep following the blood draw step, this blood could be displaced to the outside of the syringe when the VitaPrep cap is reattached to the syringe. T | DESIGN: Device Design | Stryker sent a "MEDICAL DEVICE FIELD CORRECTION" dated July 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included a Business Reply Form for customers to complete and return via fax to 1-484-323-8803. Contact the firm at 1-888-774-8870 for questions regarding this notice. | Contact the recalling firm for information | 3268 | 3,268 | Nationwide Distribution-including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV | Open |  |  | 3268 | Not_Computer
Z-0437-2013 | 63544 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen CompleteKnee Solution Stemmed Nonaugmentable Tibial componentThe NexGenÂŽÂŽ system is a semiconstrained, nonlinked condylar knee prostheses. The device is indicated for patients with severe knee pain and disability due to rheumatoid arthritis; osteoarthritis; traumatic arthritis; polyarthritis; collagen disorders and/or avascular necrosis of the femoral condyle; post-traumatic loss of joint configuration (particularly when there is patellofemoral erosion, dysfunction or prior patellectomy); moderate valgus, varus or flexion deformities; salvage of previously failed surgical attempts or for a knee in which satisfactory stability in flexion cannot be obtained at the time of surgery. | NexGen CompleteKnee Solution | 2 | November 26, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the NexGenÂż Stemmed Nonaugmentable Tibial Component due to the devices being processed through a manufacturing cleaning operation that was operating outside of the validated parameters.As a result, the devices may contain residual particulate from the manufacturing process.Risks include: Allergic reaction due to the foreign material or late onset | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Zimmer sent an Urgent Medical Device Recall letter dated September 21, 2012, to all affected customers. The letter identifiedthe product, the problem, and the action to be taken by the customer. Customers were instructed to locate and quarantine the affected product. Customers were to return the affected product along with the completed Inventory Return Certificate Form. If product was further distributed the Recall Letter should be forwarded to their customers. Product was to be returned to USA accounts - Zimmer Product Service Department, 1777 West Center St. Warsaw, IN 46580. If product has been implanted, affected surgeons will be notified. For questions regarding this recall call 574-372-4807. |  | 123 | 123 worldwide (33 USA) | Worldwide Distribution - USA including AZ,CA, FL, IA, IL, IN, MO, NE, NJ, NY, OH, PA, WA. Internationally to Thailand, Canada, Germany, France, UK, Hungary, Netherlands, Russian Fed, and Sweden, | Open |  |  | 123 | Not_Computer
Z-0440-2013 | 63681 | cannula, manipulator/injector, uterine | LKF |  |  | Obstetrics/Gynecology | 510(k) | Cooper Surgical Advincula Arch , Reusable Uterine Manipulator HandleRef: UMH750The intended use of the RUMI Arch is in an operative endoscopy (laparoscopy) where a uterus is present and where positioning of the uterus, fallopian tubes and ovaries is desirable. These types of surgeries include laparoscopic tubal ligation, and/or operative laparoscopy | Cooper Surgical Advincula Arch | 2 | November 26, 2012 | 2012 | Cooper Surgical, Inc. | Excessive pressure exerted on the uterine tip when attached to the Arch may, under certain circumstances, cause the internal steel rod of the tip to protrude through the cleaning hole on the shaft of the handle | DESIGN: Device Design | CooperSurgical issued notification on October 31, 2012, to all affected customers Fed X . The notification identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return units for replacement. Customers with questions should contact Customer Serivce at 1-203-601-9818 or 800-243-2974. Customers were asked to fill out the attached Replacement of Recalled Product Form and fax it to 1-800-262-0105.For questions regarding this recall call 203-601-9818. |  | 1215 | 1215 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Czech Republic, Germany, Italy, Spain, France, United Kingdom, Japan, New Zealand, South Korea, Saudi Arabia, Sweden, Singapore, Slovenia, and Taiwan. | Open |  |  | 1215 | Not_Computer
Z-0438-2013 | 62918 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | IO FiX X-Post Reamer, Extremity Medical, www.extremitymedical.com; Customer Service: 888.499.0079The X-Post Reamer is an accessory to the IO FiX Screw and Washer System that is used for reaming bone most common to orthopedic surgical procedures. | IO FiX XPost Reamer | 2 | November 26, 2012 | 2012 | Extremity Medical LLC | Product Event Reports revealed a broken reamer experienced intra-operatively. Pre-drilling was not performed prior to reaming. | TRAINING: Employee Error | Extremity Medical sent a notification by e-mail that describes the "Slight modification to the Surgical Technique" on April 25, 2011. For questions regarading this recall call 973-588-8980. |  | 156 | 156 units | Worldwide Distribution - USA including FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE. Internationally to Switzerland. | Terminated | March 21, 2013 | 115 | 156 | Not_Computer
Z-0444-2013 | 63583 | cover, barrier, protective | MMP | Surgical drape and drape accessories. | General & Plastic Surgery | General Hospital | 510(k) | Pentron Clinical Metrex Research; Pinnacle Cure Sleeve, a protective barrierThe intended use of these devices is to protect both patient and healthcare worker from transfer of microorganisms, body fluids and particulate matter. | Pinnacle Cure Sleeve | 2 | November 27, 2012 | 2012 | Metrex Research, LLC. | The firm initiated the recall due to a labeling error. The boxes labeled as Pinnacle Cure Sleeve 4511, which fits 11mm light guides, actually contain Pinnacle Cure Sleeves that fit 7mm or 8mm light guides (part number 4500). | PRODUCTION CONTROLS: Error in Labeling | The firm, Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 20, 2011 to its consignees/customers. The consignees/customers were instructed to determine if they have any of the affected product in their inventory; if they have the affected product to contact Metrex Research Customer Care at 1-800-841-1428 to receive an RMA number; return any affected product in their inventory; and complete and return the Recall Response Form via mail to: Metrex Research, 29210 Wick Road, Romulus, Michigan 48174, Attn: QA Lab or fax to: 734-947-6725.If you have any questions regarding this recall, you can contact our Customer Care Center at 1-800-841-1428. |  | 3348 | 3,348 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, and Europe. | Terminated | November 29, 2012 | 2 | 3348 | Not_Computer
Z-0442-2013 | 63711 | prosthesis, hip, semi-constrained (metal uncemented acetabular component) | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | 510(k) | Echo Bi-Metric Press-Fit Stems/Echo Bi-Metric 11mm x 135mm Collarless Porous Stem (Orthopedic Implant)Indications: 1) Noninflammatory degenerative joint disease including osteoarthritis and avascular necrosis, 2) rheumatoid arthritis, 3) correction offunctional deformity, 4) treatment of non-union, femoral neck fracture, and trochanteric fractures of the proximal femur with head involvement, unmanageable using other techniques, 5) revision of previously failed total arthoplasty. | Echo BiMetric PressFit Stems/Echo BiMetric 11mm x 135mm Collarless Porous Stem | 2 | November 27, 2012 | 2012 | Biomet, Inc. | It was reported that the surgeon noticed that the Echo 11mm hip stem had an etching indicating it was 13mm stem. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 8, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to take the following actions:ÂŽ Immediately locate and remove the identified device(s) listed below from circulation.ÂŽ Carefully follow the instructions on the enclosed FAX Back Response Form .ÂŽ Fax a copy of the Response Form to 574-372-1683 prior to return of product.ÂŽ Use priority carrier for your shipment.Questions related to notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 4 | 4 | Nationwide distribution: USA including states of: AZ, MD, and KS. | Terminated | August 28, 2013 | 274 | 4 | Not_Computer
Z-0446-2013 | 63675 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The brand name is MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001. The generic name is infusion administration set.Labeling indicates ~10 ml, Latex-free Infusion Set with Needle-free Port, One injection Port. Device is single use only, do not resterilize, and is available only by prescription. Device made in U.S.A. by Iradimed Corporation located in Winter Park, FL 32792. Lot LBG001, Exp 07/2016The MRidium MRI Infusion Pumps and Infusion Sets are intended for general hospital or clinical use by medical professionals whenever it is required to infuse patients with intra-venous or intra-arterial fluids before, during, or after Magnetic Resonance Imaging (MRI) scans. The device is available for sale only upon the order or a physician or other related licensed medical professional, and not intended for any home use applications. | MRIdium MRI Infusion Sets | 2 | November 28, 2012 | 2012 | Iradimed Corporation | Iradimed Corporation is recalling MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001 because a section of the device does not meet specifications. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Iradimed Corporation contacted all affected customers directly by phone, and followed up with an email notification. Customers were instructed to return the product to Iradimed Corporation. For questions regarding this recall call 407-677-8022. |  | 719 | 719 sets | Nationwide Distribution including AZ, CA, FL, KY, MD, MA, NJ, NY, PA, and TX. | Terminated | July 19, 2013 | 233 | 719 | Not_Computer
Z-0441-2013 | 63722 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Vertex Max Drill Tube is a component within the Navigated Vertex Max Drill Guide Set (item 9733370). The Navigated Vertex Max Drill Guide Set is intended to be used with a Medtronic Computer Assisted Surgery System to help assist with screw placement during spine surgical procedures. The device is reusable and requires sterilization according to product labeling.. | Medtronic Navigation | 2 | November 29, 2012 | 2012 | Medtronic Navigation, Inc. | Medtronic is recalling certain Vertex Max Drill Tubes which are used with the StealthStation System due to a defect related to the internal threading for component set up. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic Navigation sent a Urgent Medical Device Recall letter dated November 13 -14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed do not use the Vertex Max Cannulated Drill Tube clinically. Medtronic Navigation will remove the Vertex Max Cannulated Drill Tube. Medtronic Navigation will provide a replacement Vertex.Max Cannulated Drill Tube manufactured from a different lot which was inspected to verify the defect has been corrected.If you have any questions about this notification, please contact a Sales Representative in your area. For further questions please call (720) 890-3200 Ext. 3303 . |  | 12 | 12 Units | Worldwide Distribution--US (nationwide) including the states of CA, FL, MA, LA, and MI and the countries of Switzerland, Saudi Arabia, Lebanon, Canada and Korea. | Terminated | January 04, 2013 | 36 | 12 | Not_Computer
Z-0454-2013 | 63634 | apparatus, suction, ward use, portable, ac-powered | JCX | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | NeptuneÂŽ 2 Rover - ULTRA ~120V ~ 60 Hz 12AREF 0702-001-000, Rx, Caution, consult accompanying documentsConsult instructions for use, Max Vacuum; 21.0 in-Hg, 530 mm-Hg, The Neptune Waste Management System consists of a mobile rover unitused to suction and collect fluid waste, small debris and electrocauterysmoke from a surgical site. Mobility allows the rover unit to berelocated to a waste disposal area where the rovers collection canisterscan be emptied, via the docking station.Intended to be used in the Operating Room, Surgical Centers and Doctors Offices to collect and dispose of surgical fluid waste as well as collect smoke from electrocautery or laser devices. | Neptune 2 Rover Ultra | 2 | December 03, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Beginning May 24, 2011 Stryker Neptune 2 (N2) Waste Management Systems castors were replaced with a new castor design because investigation found the old design had the potential for the casters to bind compromising the Rover's ability to roll. This can result in additional strain on the operator, due to multiple exposures resulting in back pain. |  | Stryker Instruments sent a "STRYKER NEPTUNE 2 ULTRA 120V & 230V WASTE MANAGEMENT SYSTEM" letter dated June 22, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Reply Form via fax to Stryker at 269-906-9121 or email to sales.service@stryker.com. Technical questions should be addressed to Neptune Field Service at 800-800-4236 ext. 5052. |  | 5865 | 5865 | Worldwide Distribution-USA (nationwide) and Pacific countries, and the countries of EEMEA, Canada, Latin America, and Japan. | Open |  |  | 6019 | Not_Computer
Z-0464-2013 | 62749 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Total Hip Pack, K12T-05537, Merit's Total Knee Pack, K12T05538A custom procedural tray is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. | Custom Procedural Trays | 2 | December 03, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Stryker Hytrel Togas which are being recalled because the clear tape that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | PRODUCTION CONTROLS: Process Control | The firm, Merit Medical, sent an "URGENT PRODUCT RECALL NOTICE" letter dated June 8, 2012 to its customers by letter sent by US Postal Service - Certified Mail Receipt describing the problem, instructions as to actions that need to be taken, and requesting the customer to complete and sign the Customer Response Form and immediately fax a copy to this form to 1-804-416-1031. Customers were instructed to mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Should the customer choose to return the product for rework, they are to return the affected products by shipping them back to Merit Medical Systems via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department RGA #NC-53417, Merit Medical Systems, Inc., Attn: Manager of Quality Assurance and Regulatory Affairs, 12701 Kingston Ave., Chester, VA 23836. For questions customers were instructed to contact Manager of Quality Assurance and Regulatory Affairs at 1-804-416-1048 or email: jbraxton@merit.com. |  | 83 | 83 packs | Nationwide distribution: US including states of: NE and SD. | Terminated | December 10, 2012 | 7 | 83 | Not_Computer
Z-0459-2013 | 63407 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 18 Fr 6mm Sft Flow Angled WireThe Sarns Aortic Cannula is indicated for use in perfusion of the ascending aorta during cardiopulmonary bypass surgery. These devices are indicated for up to 6 hours of use. | Sarns Soft Flow Aortic Cannulae | 2 | December 03, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | The presence of plastic flash was identified at the tip of certain lots of Sarns Soft-Flow Aortic Cannula. The plastic flash has the potential to detach and contribute to an adverse patient effect. While this potential failure has not been reported from the field as an observed defect, it was determined that the product does not meet specifications and this condition could result in an unacceptabl | PRODUCTION CONTROLS: Process Control | TERUMO sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to remove this cannulae from circulation and return them for credit. Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. | Contact the recalling firm for information | 5760 | 5760 total | Worldwide Distribution - USA including AL , CA , CO , DC , DE , FL , GA , IL , IN , KS , KY , LA , MA , MI , MO , MS , MT , NC , ND , NE , NH , NY , OH , PA , TN , TX , VA , and WI. Internationally to Australia, United Arab Emirates (UAE), Singapore, Malaysia, BELGIUM, Japan, and CANADA. | Terminated | April 23, 2013 | 141 | 5760 | Not_Computer
Z-0463-2013 | 63615 | test, qualitative and quantitative factor deficiency | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Diagnostics, Spectrolyse PAI-1 In-Vitro Diagnostic Use for the quantitative determination of Plasminogen Activator Inhibitor Type 1 in human plasma.Catalog # 101201.Spectrolyse PAI-1 is intended for the quantitative determination of Plasminogen Activator Inhibitor Type-1 (PAI-1) activity in human plasma. The test is for in vitro diagnostic use. | Sekisui Diagnostics, Spectrolyse PAI1 catalog 101201 | 2 | December 03, 2012 | 2012 | Sekisui Diagnostics Llc | Spectrolyse PAI-1 activity assay, producing lower than expected absorbance values 0 IU/ml standard reducing the slope of the standard curve requiring more frequent sample dilutions. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sekisui Diagnostics issued an "URGENT PRODUCT RECALL NOTICE" dated October 24, 2012 to all consignees. The letter instructed consignees to immediately stop using the specified lot of product, discard any remaining inventory and return the enclosed response form. Contact the Customer Service Department at 800-637-3375, ext. 10, 12, or 31 for assistance with this recall. |  | 63 | 63 kits | Nationwide Distribution-including the states of FL, IN, NC, OH, UT, and WA. | Terminated | January 08, 2013 | 36 | 63 | Not_Computer
Z-0453-2013 | 63642 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | OmniGuide BP-ROBOTIC ( BeamPath Robotic), Single UseCat. # BP-ROBOTICThe OmniGuide Fiber is indicated for the incision, excision, ablation, vaporization and coagulation of body soft tissues including intra-oral tissues. It is indicated in the medical specialties of general and plastic surgery, oral/maxillofacial surgery, dentistry, dermatology, gynecology, otorhinolaryngology, gastroenterology, neurosurgery, urology, and pulmonology, and can be used in open surgical procedures as well as endoscopic minimally invasive procedures in conjunction with rigid or flexible endoscopes, such as in laryngoscopy, gastroscopy, colonoscopy, laparoscopy, thoracoscopy, hysteroscopy and bronchoscopy | OmniGuide BPROBOTIC | 2 | December 03, 2012 | 2012 | OmniGuide, Inc. | Failure of glue joint, allowed a portion of the single use device to shift inside, may separate and fall into the patient | PRODUCTION CONTROLS: Process Control | OmniGuide issued letter notifiction on November 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that an OmniGuide sales representative would contact them immediately to recover the affected product. For any questions customers were instructed to call 617-551-8561. For questions regarding this recall call 617-551-8410. |  | 45 | 45 units | Nationwide Distribution including AZ, FL, MA, NY, OH, and OR. | Open |  |  | 45 | Not_Computer
Z-0452-2013 | 63429 | photometric method, iron (non-heme) | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Dimension Vista(R) IRON Flex(R) reagent cartridge (K3085)The IRON method is an in vitro diagnostic test for the quantitative determination of iron in human serum and plasma on the Dimension Vista(R) System. Measurements of iron are used in the diagnosis and treatment of diseases such as iron deficiency anemia and other disorders of iron metabolism. | Siemens Dimension Vista(R) IRON Flex(R) reagent cartridge (K3085) | 2 | December 03, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the potential for falsely elevated results with IRON Flex(R) reagent cartridges when tests are processed from a well set contaminated with iron from environmental sources. All lots of IRON Flex(R) reagent cartridges have the potential for this issue. If a well set is contaminated with trace amounts of iron, the entire well set is impacted, and the mag | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm intiated their recall of the specified lots on October 8, 2012 by issuing a letter to all consignees informing them of the problems with the reagent cartridges. Firm recommended batching the processing of the tests on one instrument and provided instructions to disable normal flow of IRON and to run the QC on each IRON Flex(R) reagent cartridge well set. In addition, a a reply form and Frequently Asked Questions attachment were enclosed with the communication.For questions regarding this recall call 302-631-6311. |  | 19310 | 19,310 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to Canada, AU, NZ, JP, DE, SA, and MY | Terminated | January 14, 2014 | 407 | 19310 | Not_Computer
Z-0470-2013 | 63442 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker OrthopaedicsTIBIAL OFFSET BUSHING TRIATHLON REVISION INSTRUMENTSREF 6543-2-600;NON-STERILE Howmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430 USAA Subsidiary of Stryker Corp.Stryker FranceZAC Satolas Green PusignanAv de Satolas Green 69881 MeyzieuCedex FranceMade in USAThe Offset Bushings are parts of the Triathlon TS system. They are used during trialing to assess the appropriate magnitude and orientation of offset between the Implant and the intramedullary Stem. The feedback from the Offset Bushing allows for the selection of the corresponding Offset Adaptor, which allows for optimal bone coverage, compensating for differences between the Implant and bone canal axes. | Stryker Howmedica Osteonics Corp. | 2 | December 04, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Offset Bushings associated with the lots identified were mismarked: the rotational reference numbers on the Tibial Offset Bushing were marked counterclockwise instead of clockwise; the Femoral Offset Bushing did not include rotational markings. | PRODUCTION CONTROLS: Error in Labeling | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated July 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgement Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for assistance with this recall. |  | 8 | 8 | Worldwide Distribution. | Open |  |  | 15 | Not_Computer
Z-0472-2013 | 63612 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | OSS Reamer Sleeve REF132-472662 KNEE REAMER SLEEVE 13.0 thru 20.0 STAINLESS STEELThe OSS Knee Reamer Sleeves are primarily used to guide the bullet tip of cylindrical reamers during the tibial preparation in OSS procedures. They are reusable. | OSS Reamer Sleeve | 2 | December 04, 2012 | 2012 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the thin shaft reamer may get stuck and not slide into the sleeve. | DESIGN: Process Design | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated November 1, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately locate, discontinue use of and remove the product from circulation; and return recalled devices with directions to distribute notices, if product was further distributed; and to complete and return the enclosed FAX Back Response Form within three (3) business days to FAX # 574-372-1683. Distributors were charged with locating and returning products to Biomet ATTN: Return Goods-Building B RECALLS RGA#: _, Biomet, Inc., 56 East Bell Drive, Warsaw, IN 46580. Questions should be directed to the Field Action Coordinator at (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 36 | 36 | Worldwide distribution: USA (nationwide) and countries of: Japan and Netherlands. | Terminated | April 29, 2013 | 146 | 82 | Not_Computer
Z-0468-2013 | 63643 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression MRI Patient Monitoring System.The Expression MRI Patient Monitoring System is intended for use by healthcare professionals to monitor vital signs for patients undergoing MRI procedures. The Patient Monitoring System also provides signals for synchronization to the MR scanner. | Invivo Corporation Expression MRI Patient Monitoring System | 2 | December 04, 2012 | 2012 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling the Expression MRI Patient Monitoring System because during transport of the systems cart with a docked Display Controller Unit (DCU), a caster (wheel) may become detached. If this occurs, the cart can become unbalanced and may fall. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Healthcare sent an "URGENT-MEDICAL DEVICE RECALL" letter dated October 15, 2012 to all affected customers. The letter identified the product, problem, and actons to be taken by the customers. Customers were instructed to inspect all casters of the unit to ensure that they are all secured. If a caster is loose, customers were told to lock the caster in place, limit movement of the cart and contact their local Phillips Invivo Representative. Nevertheless, a Philips Invivo representative will contact the customer regarding their affected device. All affected devices will have new casters installed in order to correct the problem. Contact your local Philips Invivo Representative at 1-800-722-9377 for further information and support. |  | 543 | 543 units | Worldwide Distribution-USA (nationwide) including the states of MO; MS; MA; IL; OH; NY; KS; PA; WI; VA; SD; AZ; FL; OR; TX; MA; MO; FL; MA; MI; CT; CA; NJ; NC; ME; IL; MA; MA; WV; MA; MN; CA; AL; AL; WI; DC; OH; WV; CA; NC; NY; MT; NJ; CA; WA; IN; MI; PA; TX; NC; TN; TX; ID; NM; IL; PA; IN; IL; WA; CO; MN; CA; IN; MD; MI; DC; GA; DE; CO; CT; WA; NJ; MD; TX; TN; MN; WA; CA; FL; OH; PA; NY; IN; OK; ID; IA; NY; WI; VA; CA; HI; MS; TX; HI; CT; TN; OR; WA; PA; NH; CA; MA; LA; IN; PA; AZ; GA; UT; ND; GA; TX; TX; IN; CA; MI; CT; CA; NC; MS; VA; IN; MI; NE; NV; MI; WA; IL; KY; LA; OR; OK; LA; TN; KY; TX; CA; PA; NM; NC; CO; DC; AL; MI; TX; SD; PA; OR; ID; NM; SD; CO; MN; NJ; MI; ND; MN; MT; LA; CT; WA; TX; IA; NJ; WI; SD; and UT, and the countries of Canada, Italy, Thailand, Russian Federation, Saudi Arabia, Slovenia, Singapore, Republic of Korea, Uzbekistan, Iceland, Israel, Austria, United Kingdom, Afganhistan, Japan, South Africa, Germany, Qatar, Netherlands, Switzerland, Denmark, and Sweden. | Open |  |  | 543 | Not_Computer
Z-0466-2013 | 63553 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Sliced CT Scanner System, part number 989605858501. NeuViz Dual series CT Scanner System, part number 989605651321The NeuViz 16 CT Scanner System can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. | NeuViz 16 MultiSliced CT Scanner System and NeuViz Dual series CT Scanner System | 2 | December 04, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | The four screws that secure the patient table top to the carrier pulled out and caused the table top to detach. This issue may affect NeuViz Dual and NeuViz 16 scanners. | DESIGN: Device Design | Philips and Neusoft Medical Systems Co., Ltd. sent an Urgent Device Correction NeuViz Dual, NeuViz 16 letter, dated Oct. 18, 2012, to all affected US customers and distributors. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to:1. Have a mechanic check the patient table top screws fixing. Detailedinstructions are attached as Annex with the letter.2. If it is found that table top screws are improperly assembled, stop usingthe product and contact local Service personnel for repair.Philips and Neusoft Medical Systems Co., Ltd is in the process of upgrading devices with aforementioned kits and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems Co. Ltd. at nms-service@neusoft.comFor questions regarding this recall call 425-487-7665. | Contact the recalling firm for information | 20 | 20 units of NeuViz 16 CT Scanner Systems and 25 units of NeuViz Dual series | Nationwide Distribution including North Carolina, Ohio, Texas, and Puerto Rico. | Terminated | July 08, 2013 | 216 | 20 | Not_Computer
Z-0474-2013 | 63708 | prosthesis, shoulder, non-constrained, metal/polymer cemented | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-Modular Humeral Head 48mm x 24mm Ti 6AI4V/Ion Implanted (Orthopedic Implant)Product Usage:The OSS Knee Reamer Sleeves are primarily used to guide the bullet tip of cylindrical reamers during the tibial preparation in OSS procedures. They are reusable. | OSS Knee Reamer Sleeves | 2 | December 05, 2012 | 2012 | Biomet, Inc. | The items in this lot are missing the ion implantation feature. The surface of the head may be softer and more prone to scratching, which in tum could potentially cause higher wear of the poly bearing. Increased wear could potentially lead to a shorter implant life. | PRODUCTION CONTROLS: Process Control | Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE dated November 1, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The notification letter instructed the surgeon to montior the patient as their is a potential for increased wear which have an unknown effect on implant life. For questions call 574-372-1570. | Contact the recalling firm for information | 4 | 4 | USA Nationwide Distribution in the state of WI. | Terminated | August 29, 2013 | 267 | 4 | Not_Computer
Z-0478-2013 | 63635 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | The GlideScope Video Laryngoscope (GVL) is a video laryngoscope that incorporates a high-resolution color camera, an LED light source, a rechargeable lithium battery, and video output to a remote monitor. It consists of a video laryngoscope and a portable video monitor.The affected GVL and Part numbers are GVL 3  0574-0007, GVL 4  0574-0001, and GVL 5  0574-0030.The GlideScope Video Laryngoscopes are intended for use by qualified medical professionals to obtain a clear, unobstructed view of the vocal cords for medical procedures. | The GlideScope Video Laryngoscope (GVL) GVL 3, GVL4, GVL5 | 1 | December 07, 2012 | 2012 | Verathon, Inc. | Re-usable video laryngoscope blades of the GlideScope GVL3, GVL4, and GVL5 were recalled due to potential cracking. | DESIGN: Device Design | The firm, Verathon, sent an "URGENT - Medical Device Recall" letter dated October 25, 2012, to their consignees/cutomers. The letter described the product, problem, and actions to be taken. The customers were instructed to return the recalled blades back to Verathon and complete and return the Device Recall Notification Reply Form via fax to (425) 883-2896 or email: cservice@verathon.com. All returned products will be quarantined and will be destroyed later. Verathon will replace the affected video laryngoscope blade at no cost to customers. Customers in Canada and US with questions can call 800-331-2313. If you are located outside of the US or Canada, please call 425-867-1348. |  | 2710 | total 2710 units | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NV, NJ, NY, OK, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and countries of: Argentina, Australia, Canada, China, France, Germany, Indonesia, Japan, The Netherlands, Pakistan, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, United Kingdom, and Venezuela. | Terminated | September 27, 2013 | 294 | 2710 | Not_Computer
Z-0504-2013 | 63728 | bronchoscope (flexible or rigid) | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | WANG Transbronchial Aspiration Needles Cat. No. SW-121, WANG Cytology Needles, Central, Carinal and Mediastinal Regions, 21 g x 15 mm long needle. --- Common Name: Bronchoscope (flexible or rigid) and accessories. The ConMed WANG Transbronchial Aspiration Needles are designed for use with a bronchoscope. They are used to puncture the tracheobronchial wall and to aspirate sufficient tissue and/or cell specimens to stage bronchogenic carcinoma. These are disposable sterile, single-use devices. | WANG Transbronchial Aspiration Needles | 2 | December 10, 2012 | 2012 | ConMed Corporation | ConMed received complaints from a user facility reporting that the WANG Transbronchial Aspiration Needles would not insert properly through the instrument channel of a 2.0 mm bronchoscope. If the device will not fit through the instrument channel of a 2.0 mm bronchoscope, it must be replaced with an alternate device and could result in minor inconvenience to the user. | DESIGN: Component Design/Selection | ConMed Corporation sent an Urgent Field Safety Notice letter dated November 30, 2012 with Response Form , to all affected customers via USPS Priority Mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete Attachment II and return it with any unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502United Sates of AmericaFor information regarding this recall call 315-624-3225. | Contact the recalling firm for information | 452 | Total 452 | Worldwide Distribution - USA including CA , FL, IA, IN, MA, MD, ME, MO, MS, NC, NM, OH, PA, TX. Foreign distribution to Spain, Portugal, Japan and Hungary. | Open |  |  | 452 | Not_Computer
Z-0501-2013 | 61209 | electrode, cutaneous | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Empi STIMC RE Self Adhesive electrodes. For use on TENS, NMES, FES. a DJO brand A cutaneous electrode is an electrode that is applied directly to a patient's skin either to record physiological signals (e.g., electroencephalogram) or to apply electrical stimulation. | Empi STIMC RE Self Adhesive electrodes | 2 | December 10, 2012 | 2012 | Covidien LLC dba Uni-Patch | Covidien is conducting a voluntary recall regarding specific lots of TENS electrodes because it was determined that during the manufacturing process the surface of the wires in these lots may have been exposed to an employee who had a minor cut. | TRAINING: Employee Error | Covidien sent An Urgent Product Recall letter, dated February 9, 2012, to all affected consignees. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the affected product if found. The letter asked customers to fill in teh quantity remaining in their inventory and fax it to 651-565-5888. A Covidien Customer Service rep would contact the consignee for returning product. Consignees were also requested to notify their customers of the recall.For questions regarding this recall call 508-261-8532. | Contact the recalling firm for information | 28138 | 28,138 pouches | Nationwide Distribution including CA, IL, MA, MN, and SD. | Terminated | January 03, 2013 | 24 | 37036 | Not_Computer
Z-0515-2013 | 63496 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | MAVIG PORTEGRA 2 system operated in conjuction with a Siemens AX systemCeiling mounted surgical light | MAVIG PORTEGRA 2 system | 2 | December 12, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A particular component may not have been fitted during the installation of some MAVIG PORTEGRA 2 systems. | TRAINING: Employee Error | Siemens sent a Customer Safety Advisory Notice on October 3, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were reminded that they should contact the service department in the event of damage or other irregularity. Customers were instructed to forward the information contained in the notice to their employees responsible for the operation of the affected product. Customers were also asked to forward the safety information to other organization affected by the product and to any users of the product.For questions regarding this recall call 610-219-4834. | Contact the recalling firm for information | 1315 | 1315 | Nationwide Distribution | Open |  |  | 1315 | Not_Computer
Z-0520-2013 | 63758 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI) product, P/N 14324.Validation kits with Sterrad Cyclesure 24 BI products: Codes/Serial Numbers, Product Description:14324 Sterrad Cyclesure 24 Biological Indicator.14324-02 Sterrad Cyclesure 24 Biological Indicator.10134 Sterrad 100NX Express Cycle Kit.10135 Sterrad 100NX Express Cycle Upgrade Kit.10137 Sterrad 100NX Duo Cycle Upgrade kit.14325 Cyclesure Challenge Pack Kit.20103 Sterrad 100NX Test Pack Kit.20228 Sterrad 100NX Intl Validation Kit.20229 Sterrad 100NX Validation Kit, USA.20232 Sterrad 50/100S Validation Kit.20236 Sterrad 50/100S Validation Kit.20248 Sterrad 100NX Express Validation Kit.20253 Sterrad NX Validation Kit.20254 Sterrad NX Validation Kit Domestic, US.The Sterrad Cyclesure 24 Biological Indicator (BI), P/N 14324 is intended to be used as a standard method for frequent monitoring of the Sterrad Sterilization System cycles. | STERRAD CYCLESURE 24 Biological Indicator | 2 | December 12, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling all lots of Sterrad Cyclesure 24 Biological Indicator (BI) manufactured between February 2008 and December 2011 because it may not have adequate data to support the entire duration of the labeled shelf-life. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated November 2, 2012 to all customers who received the Sterrad Cyclesure 24 Biological Indicator (BI) product P/N 14324 manufactured between February 2008 and December 2011. The letter informed the customers of the problem identified and the actions to be taken. ASP determined that the 24 BI may not have adequate data to support the entire duration of the labeled shelf life of affected product manufactured during the above timeframe. The letter also informed customers that this is a different recall from the previous July 3, 2012 STERRAD CYCLESURE 24 BI recall. Customers with questions or need additional packing labels are instructed to call Stericycle at (877) 650-7692. |  | 38532 | 38,532 units | Worldwide Distribution-USA (nationwide) and the countries of AR, AU, BE, BR, CA, CL, CN, CO, EC, GT, HK, ID, IN, JP, KR, MX, MY, PA, PE, PH, SG, TH, VE, and RU. | Terminated | July 19, 2013 | 219 | 38532 | Not_Computer
Z-0519-2013 | 63530 | lamp, infrared, therapeutic heating | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(k) | Quasar Power Pack contains a Baby Quasar and a Baby Blue. The Baby Quasar is an infrared LED lamp. The Baby Blue uses blue LED light to be used in combination with the Baby Quasar to treat acne.The Quasar Power Pack is labeled in part: "***BABY QUASAR & BABY BLUE***QUASAR POWERPACK***FOR YOUTHFUL, RADIANT, CLEAR SKIN***SKIN REJUVENATING SYSTEM WITH SEQUEPULSE PLUS ADVANCED ACNE CLARIFYING SYSTEM***Manufactured by: Quasar Biotech, Inc., BPP101 LAB-003-E, ***Made in the USA with highest quality components and durable aluminum. Five Year Warranty.***" | Quasar Power Pack | 2 | December 12, 2012 | 2012 | Quasar Bio-Tech, Inc. dba Silver Bay LLC | Quasar Bio-Tech is recalling their Baby Quasar, Baby Quasar Pink, Quasar Power Pack, and Quasar MD because the device is unapproved. |  | The firm, Quasar Bio-Tech, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 16, 2012 to its customers. All domestic customers were notified by Certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to cease distribution and remove the products from sale; to complete and return the enclosed Response Form via e-mail to DeNette@quasarbiotech.com or fax to: Quasar Bio-Tech, Attn: Quality/Recall Coordinator: at 941-306-5816, and return all remaining stock to Quasar Bio-Tech at Quasar Bio-Tech Inc., Attn: Quality/Recall Coordinator, 1431 Tallevast Road, Sarasota FL 34243.If you have any questions, contact the Quality/Recall Coordinator at 941-306-5812, Monday-Friday from 9:00 am to 5:00 pm Eastern Standard time. | Contact the recalling firm for information | 7734 | 7,734 devices total (included in batches identified for the Baby Quasar ) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, FL, GA, HI, ID, IL, KS, MD, MI, MN, MO, NC, NJ, NY, NV, OH, OR, TN, TX, UT, VA, and WA; and countries of: Aruba, Canada, Cayman Islands, China, Estonia, Ireland, Malaysia, New Zealand, Singapore, Sweden, Taiwan, and United Kingdom. | Terminated | May 21, 2013 | 160 | 7734 | Not_Computer
Z-0522-2013 | 63680 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 123 POC systemcobas b 123 POC system is fully automated POC system for whole blood in vitro measurement of pH, blood gases (BG), electrolytes Na+ K+, iCa2+ (ISE), hematocrit (Hct), metabolites (Gluc, Lac), total hemoglobin (tHb), hemoglobin derivatives (O2Hb, HHB, COHB, MetHb), and oxygen saturation (SO2). In addition, the cobas b 123 POC system calculates derived parameters. It is dedicated for use in a Point-of-Care environment and laboratory. | Cobas b 123 POC system | 2 | December 13, 2012 | 2012 | Roche Diagnostics Operations, Inc. | It has been determined that cobas b 123 Fluid Pack COOX cuvettes for lot numbers 21426121 and 21426171, exhibit a significant increase in the out-of-box failure rate caused by a change in the thickness in the cuvette layer. Internal investigations found the altered thickness of the cuvette was due to production equipment settings. Cuvettes were produced within specifications, but within the upper | DESIGN: Software Design (Manufacturing Process) | Roche sent an Urgent Medical Device Removal (UMDR12-268) notification and faxback form on November 14, 2012, to all affected consignees. Customers were requested to discontinue the use of cobas b 123 Fluid Pack COOX, 200 lots 21426121 and21426171 and discard the product per local regulations. Customers were instructed to use the enclosed faxback form to order replacement product at 1-800-722-7222. For questions customers were instructed to contact Roche Diagnostics Technical Support at 1-800-526-2272.For questions regarding this recall call 317-521-3966. | Contact the recalling firm for information | 23 | 23 units of Fluid Pack COOX, 200 were distributed | Nationwide Distribution including NE and AZ. | Terminated | August 07, 2013 | 237 | 23 | Not_Computer
Z-0523-2013 | 63723 | antisera, fluorescent, all types, hemophilus spp. | GRO | Haemophilusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Taxo XV Factor Strips, Catalog number 231104, in shelf pack vials, fifty strips/vial labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152 USA 800-638-8663, www.bd.com/ds*** | BD BBL Taxo XV Factor Strips | 2 | December 13, 2012 | 2012 | Becton Dickinson & Co. | In vitro diagnostic test kit was not manufactured according to specifications and may cause false negatives or misidentification of Hemophilus species present in patient specimens. | TRAINING: Employee Error | BD Diagnostic Systems sent an Urgent Product Recall letter dated October 22, 2012, to all affected customers.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue distribution and discard affected inventory for credit. . Customers were requested to return a response form by fax to 410-316-4258. For questions customers were instructed to contact BD Customer Service Department at 1-800-638-8663. For assistance with replacements, customers were instructed to call 1-800-675-0908.For questions regarding this recall call 410-316-4258. | Contact the recalling firm for information | 650 | 650 vials | Worldwide Distribution - USA (nationwide) to international customers in Belgium, India, Japan, South Korea, Thailand and Mexico | Terminated | January 24, 2013 | 42 | 650 | Not_Computer
Z-0497-2013 | 63712 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂŽ II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. | SynchroMed II implantable infusion pump | 1 | December 13, 2012 | 2012 | Medtronic Neuromodulation | Medtronic notified Healthcare Professionals of the impact of unapproved drugs on the performance of the SynchroMed infusion pump system. Use of unapproved drugs with SynchroMed pumps can result in an increased risk of permanent motor stall and cessation of drug infusion. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm Medtronic, sent an "Urgent medical Device Safety Notification" dated November 2012, to its customers beginning November 09, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to follow the recommendations provided in the letter and complete and return the PHYSICIAN REPLY FORM via mail at Medtronic Neuromodulation Quality Assurance, 7000 Central Ave NE, Mailstop RCC150, Minneapolis, MN 55432; fax to: 800-897-3899 or email to: neuro.quality@medtronic.com.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 174700 | 174,700 devices total | Worldwide distribution: USA (nationwide) and countries including: Canada and Mexico. | Open |  |  | 174700 | Not_Computer
Z-0521-2013 | 63694 | immunohistochemistry reagents and kits | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Ready-to-Use Mouse Monoclonal Antibody Terminal Deoxynuceotidyl Transferase Staining Reagent; an in vitro staining reagent for the qualitative identification by light microscopy of Terminal Deoxynucleotidyl Transferase molecules in paraffin sections; Catalog No. RTU-TdT-339; packaged in 7 mL vials; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. RTU-TdT-339 is intended for the qualitative identification by light microscopy of Terminal Deoxynucleotidyl Transferase molecules in paraffin sections. The clinical interpretation of any staining or its absence should be complemented by morphological studies using proper controls and should be evaluated within the context of the patient's clinical history and other diagnostic tests by a qualified pathologist. | Novocastra ReadytoUse Mouse Monoclonal Antibody Terminal Deoxynuceotidyl Transferase | 2 | December 13, 2012 | 2012 | Leica Microsystems, Inc. | RTU-TdT-339 is not stable up to the expiry date on the product labeling, affecting the staining intensity. |  | Leica Microsystems, Inc. sent a Product Recall Notification letter dated November 14, 2012, via first class mail on the same date, to all affected customers. The letter identified the product ,the problem, and the actions to be taken by the customers. Customers were instructed to discontinue use of the product and to return any unused/part used reagents to the Richmond, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-236-3747. Additional questions should be addressed to the account's Leica Representative. | Contact the recalling firm for information | 322 | 322 vials | Worldwide Distribution -- USA including Arkansas, California, Florida, Illinois, Maryland, Michigan, Minnesota, Missouri, North Carolina, Oregon, Rhode Island, Tennessee, Texas and Wisconsin, and the country of Canada. | Open |  |  | 322 | Not_Computer
Z-0514-2013 | 63710 | ventilator, high frequency | LSZ |  |  | Anesthesiology |  | Patient Circuit used with the Life Pulse High Frequency Ventilator. The Patient Circuit is a required disposable component of the Life Pulse High Frequency Ventilator.Model #(s):Individual Patient Circuit - catalog # 902Patient Circuit Kit - catalog # 937: contains 2 circuits and 2 x 2.5 & 2 x 3.5 LifePort Adapters.Labeling states manufactured by: Bunnell Incorporated 436 Lawndale Drive, Salt Lake City, Utah 84115.Used for ventilating critically ill infants with pulmonary interstitial emphysema and infants with respiratory distress syndrome complicated by pulmonary air leaks, who are, in the opinion of their physicians, failing on conventional ventilation. The Patient Circuit is to provide a conduit for the humidification, warming, and temperature monitoring of the pressurized gas. The Patient Circuit is indicated for seven day single use. | Life Pulse High Frequency Ventilator | 1 | December 14, 2012 | 2012 | Bunnell, Inc. | Customer complaints received indicate the heater wire insulation can melt, causing a short which creates sparking and smoke, in the circuit close to the humidifier cartridge. Bunnell is recalling all lots of circuits distributed between March 12, 2012 and November 30, 2012. | DESIGN: Component Design/Selection | On 12/5/12, the firm, Brunnell, Inc., issued a public press notification to their consignees. A. On or before 12-12-2012 a certified letter will be sent to all customer listed in the customer database. Specifically each customer will receive two certified letters; one directed to Director of Respiratory Therapy and a second to the Director of Biomedical Engineering. The certified letter will contain a copy of the Recall Notice and a copy of the Certificate of Medical Necessity. The Certified Letter receipts will be maintained. On the mailing date a copy of the Recall Notice will be placed on the Bunnell website. B. An Excel database will be maintained for each customer/lot #. The database will contain the current status for each. Included will be: 1. Customer Information 2. Date Certified Letter Sent 3. Date receipt received and filed 4. Date product returned or that no stock remains 5. Lot number & Quantity C. On or before 02-01-2013, a second letter will be sent to those customers who have not responded to the original certified letter. D. Data will be maintained in the database E. On or before 04-01-2013, a phone call will be placed to those customers who have not responded to either of the letters. F. Data will be maintained in the database G. On or before 04-15-2013, a second phone call will be placed to those customers who have not responded to the letters or previous phone call. H. Data will be maintained in the database I. On or before 06-07-2013, a Final Recall Notification certified letter will be sent to those customers who have not responded to either of the previous letters or phone calls. J. The Certified Letter receipt will be maintained K. Data will be maintained in the database L. When PMA for alternate heater wire is approved, steps A-K will be repeated. Dates will be determined. M. Following completion of Step L, a final review of the database will occur. Final tabulation of the status for each custo. |  | 5743 | 5743 patient circuits distributed in commerce | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Malaysia, and Uruguay (animal study only). | Terminated | February 21, 2014 | 434 | 5743 | Not_Computer
Z-0535-2013 | 63549 | spinal vertebral body replacement device | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Titan Spine Endoskeleton TA Implant products. 2107-0118 18mm Standard Endoskeleton Implant VBR 2107-012020mm Standard Endoskeleton Implant VBR2107-021818mm Large Endoskeleton Implant VBR2107-022020mm Large Endoskeleton Implant VBREndoskeleton TA VBR is intended for use in the thoracolumbar spine (T1-L5) to replace all or part of a collapsed, damaged, or unstable vertebral body due to tumor or trauma. | Titan Spine Endoskeleton TA Implant | 2 | December 17, 2012 | 2012 | TITAN SPINE, LLC | Titan Spine LLC, is conducting a recall on the Endoskeleton TA VBR products. The labeling accompanying this device contains incorrect indications for use. | MISBRANDING: Labeling False and Misleading | Titan sent an Urgent Medical Device Recall letter dated October 19, 2012, to all affected customers. The letter described the problem , the product affected and the corrective measures to be taken. Customers were instructed to review any surgeries that were performed at the L5-51 level with the affected device. Customers were asked to communicate any complications associated with those cases to Titan Spine. Customers with questions should call 1-262-242-7801.For questions regarding this recall call 262-242-7801. |  | 92 | 92 pcs. | Nationwide Distribution including NV | Terminated | January 14, 2013 | 28 | 92 | Not_Computer
Z-0536-2013 | 63653 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Polaris 5.5 Button Lock Screw Inserter. Catalog number 14-500178 | Biomet Polaris 5.5 | 2 | December 17, 2012 | 2012 | Ebi, Llc | Broken screw insertor hardness was measured and found to be below the acceptable limits. | PRODUCTION CONTROLS: Process Control | Biomet mailed Urgent Medical Device Recall Notice/Fax Back Response Form dated 9/25/2012 to the distributors on 9/29/2012. |  | 15 | 15 units | Nationwide Distribution, including the states of MO, TX, and IN. | Open |  |  | 15 | Not_Computer
Z-0527-2013 | 63725 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | EBI XFIX DFS Rail System (Adult Rail Gradual Swivel Clamp). The Biomet Rail System consists of a unilateral external fixator rail, modular bone screw clamps, and compression and distraction units. The device is used in the treatment of bone fractures. | Biomet Rail System | 2 | December 17, 2012 | 2012 | Biomet, Inc. | It was reported that the number scale/graduation markings were missing on 14 pieces of the swivel clamp. | DESIGN: Process Design | On 11/9/2012 an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter was sent to all consignees (distributors and Doctor notices) informing them of the issue and providing recommended actions. Customer contact number (574) 372-1570 Monday through Friday from 8 am to 5 pm. | Contact the recalling firm for information | 49 | 49 | Worldwide Distribution, including the states of TX, TN, PA, CA, NY, FL, and MI and the countries of Japan and Columbia. | Terminated | August 28, 2013 | 254 | 49 | Not_Computer
Z-0526-2013 | 63824 | device, biofeedback | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | EEG NeuroAmp, Model Number: CS 10090; CS 10137.Biofeedback and Relaxation | EEG NeuroAmp | 2 | December 17, 2012 | 2012 | EEG Info | EEG Info, Inc. is recalling the EEG Neuroamp device because they have identified a potential risk associated with the mislabeling of certain devices. | MISBRANDING: Labeling False and Misleading | EEGInfo sent a recall notification letterdated November 29, 2012, to all affected customers. The letter informed the customers they have not received approval from the FDA to market and distribute a device for therapeutic applications other than that of its FDA intended use-biofeedback relaxation. Customers are informed that the safety and effectiveness for other indications has not been established. The risk or likelihood of serious adverse health consequences is remote. Customers were instructed to forward this notice if affected product was further distributed. Customers with questions about the notification letter are instructed to contact EEG Info's customer support line during normal business hours, Monday-Friday, from 8am-5pm PST. For questions regarding this recall call 818-456=5965. | Contact the recalling firm for information | 1308 | 1308 units | Nationwide Distribution | Terminated | June 11, 2013 | 176 | 1308 | Not_Computer
Z-0532-2013 | 63520 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Medline Total Knee CDS; a single patient prescription procedure pack and disposables, including a back table cover;Product Usage: This Complete Delivery Systems (CDS) custom pack includes components necessary for a total knee replacement surgical procedure. | Medline Total Knee CDS | 2 | December 17, 2012 | 2012 | Medline Industries Inc | The Back Table Cover component in the custom pack was placed in the non-sterile portion of the pack when it should have been in the sterile portion. | PRODUCTION CONTROLS: Packaging | Medline Industries Inc notified the sole direct account on September 20, 2012 via telephone that the affected custom pack were recalled. Instruction was provided to the staff at the facility not to use the pack until it had been reworked with a sterile back table cover. |  | 18 | 18 packs | US (nationwide) in the state of West Virginia | Terminated | December 17, 2012 | 0 | 18 | Not_Computer
Z-0529-2013 | 63732 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek 11 MM Fully Fluted Reamer, SterileProduct Number: 232423DePuy Mitek Sterile Reamers are intended to cut through cortical and cancellous bone during Cruciate Ligament Repair/Reconstruction of the knee. | DePuy Mitek Sterile Reamers | 2 | December 17, 2012 | 2012 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Specific units of the DePuy Mitek Sterile Reamers are not labeled correctly, 7 units were labeled as 11MM while the box actually contained a 10 MM reamer | PRODUCTION CONTROLS: Packaging Process Control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refer to the attached instructions to report their inventory status and for returning the affected product. Customers with questions should call 1-508-880-8477. For questions regarding this recall call 508-880-8100. | Contact the recalling firm for information | 20 | 20 | Nationwide Distribution including IL and NJ | Terminated | December 21, 2012 | 4 | 20 | Not_Computer
Z-0533-2013 | 63781 | prosthesis, knee, patello/femorotibial, semi-constrained, uhmwpe, pegged, cemented, polymer/metal/polymer | MBV | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray 2.5 mm Adaptor, REF 141490Sterile, QTY.1Product Usage:Knee joint replacement prostheses intended for application with or without bone cement. Knee joint replacement components include femoral, tibial, and patellar components, intended for both primary and revision applications. Indications: 1) painful and disabled knee joint resulting from osteoarthritis, rheumatoid arthritis, traumatic arthritis when one or more compartments are involved, 2) correction of varus, valgus, or posttraumatic deformity, 3) correction or revision of unsuccessful osteotomy, arthrodesis, or failure of previous joint replacement procedure | Biomet Offset Tibial Tray Adaptor | 2 | December 17, 2012 | 2012 | Biomet, Inc. | Biomet investigation identified that the square end of the offset tibial tray adaptor locking insert is oversized and may not engage insert tool screwdriver. This could cause a possible delay in surgery greater than 30 minutes, potentially exposing the patient to the increased risks from being under anesthesia for a longer duration. | TRAINING: Employee Error | Biomet Recall sent an Urgent Medical Device Recall Notices letter dated November 15, 2012 to distributors, hospital operating room(OR)managers and Surgeons receiving the affected product. The letter identified the affected product, the reason for the recall and actions to be take, including directions to immediately locate and discontinue use of the affected product. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with affected product to be returned are asked to obtain RGA # by calling 574-372-6677 or email: rgarequest@biomet.com. A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Return product to the Attention of Audrey Daenzer, Field Action Coordinator, Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m. |  | 200 | 200 | Worldwide distribution - USA (nationwide) and the countries of Australia, Canada, Chile, Europe, Japan, Korea, Panama, Puerto Rico and Thailand. | Terminated | November 08, 2013 | 326 | 200 | Not_Computer
Z-0534-2013 | 63808 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Arobella Medical 6mm Qurette used with the Arobella Medical AR1000 Wound Therapy System. Major components of the AR1000 Wound Therapy System include an ultrasonic generator, converter hand piece, done-shaped applicator tip (qurette), and a saline supply reservoir. Product Usage:The Arobella Medical AR1000 Wound Therapy System is interned for the selective dissection and fragmentation of tissue, wound debridement ( of acute and chronic wounds, burns, and diseased or necrotic tissue), and cleansing irrigation of the wound site for the removal of debris, exudates, fragments, and other matter. | AR1000 Qoustic Qurette" | 2 | December 17, 2012 | 2012 | Arobella Medical, LLC | This recall has been initiated because the Arobella Medical 6mm Qurette Q07 Lot Numbers 0001 and 0002 can fracture during use and may release a fragment or fragments. The 6 mm Qurette is a component of the AR1000 Quostic Wound Therapy System. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Arobella Medical sent an " Urgent: Medical Device Recall" letter dated November 2, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine and return any product subject to the recall. If the affected product have been further distributed, customers were advised to notify their customers of the product recall. For questions call 952-288-2810. |  | 99 | 99 | Worldwide distribution - US (nationwide) and the states of AL, CA, ID, IN, IA, MI, MN, MS, NJ, NM, OH, PA. SC, TX, UT, VA, WI. including the country of SAUDI ARABIA. | Open |  |  | 99 | Not_Computer
Z-0538-2013 | 63641 | prosthesis, elbow, constrained, cemented | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/Morrey New Ulnar Revision Kit for Use with Non-Interchangeable Regular Humerals STERILE Rx, Tivanium QTY/PKG:1Product Usage: The ulnar and humeral components are connected by placing the hollow, outer axis pin across the two components and securing It with the solid Internal axis pin. The Inner and outer pin connection Is articulation point of the placement elbow. | Coonrad/Morrey | 2 | December 18, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall because affected lots could include an extra Small Inner Pin instead of the required Regular Inner Pin in the Coonrad/Morrey Elbow Replacement Set and the New Ulnar Revision Kit . | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Zimmer sent "URGENT MEDICAL DEVICE RECALL" letter dated November 19, 2012 to Distributors, Risk Managers/ Implanting Surgeons, and direct accounts. The notice describes the product, problem, risks and actions to be taken. Consignees were instructed to immediately locate, quarantine and return affected product with the completed Inventory Return Certification form to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 (US accounts) or Zimmer International Logistics GmbH, ATTN: Fao Tanja Herold (Recall Warsaw), Max-Immelmann-Allee 12 79427 Eschbach Germany for international accounts. If an implant is removed from customer inventory, they were directed to provide a copy of the Risk Manager letter upon retrieval of that implant(s) to ensure all facilities are aware of the removal. Surgeons are asked to consider risk information in the letter should a patient present with unexplained symtoms. Hospitals that received direct shipments from the Warsaw Distribution Center will be sent a copy of the Risk Manager letter directly. Questions and concerns along with information on any implanted product is to be addressed to CorporateQuality.PostMarket@zimmer.com using the provided spreadsheet or contact Zimmer at 1-800-447-5633. |  | 17 | 17 | Worldwide distribution - USA nationwide and the countries of: Canada, Germany, Finland, Japan, Italy, San Juan, Switzerland, and UK. | Open |  |  | 40 | Not_Computer
Z-0546-2013 | 63690 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona, The Personalized Knee System, Tibial Drill Guide, Cemented, 15.7 mm Diameter, Nonsterile, NonsterileThis drill guide is used with the Persona Knee System for patients with severe knee pain and disability due to: Rheumatoid arthritis, osteoarthritis, traumatic arthritis, polyarthritis. o Collagen disorders, and/or avascular necrosis of the femoral condyle. o Post-traumatic loss of joint configuration, particularly when there is patellofemoral erosion, dysfunction or prior patellectomy. o Moderate valgus, varus, or flexion deformities | Persona, The Personalized Knee System | 2 | December 18, 2012 | 2012 | Zimmer, Inc. | All lots of the Persona Cemented Tibial Drill, 5 Degree Cemented Stemmed Natural Tibia implants, and associated tools including the Tibial Drill Guide, Broach, and Sizing Plate are being removed and replaced with enhanced products and instruments to reduce risks of tibial perforation which could result in tibial bone fracture due to the stress riser created by a perforation; tibial plate loosening | DESIGN: Device Design | Zimmer sent distributors beginning the week of November 14, 2012, E-mail information requesting all equipment be collected in distributors office before shipment back to Zimmer and replacement. A telephone conference discussing the recall and amended surgical technique was held the week of November 14, 2012, and Urgent Recall notices were sent to Distributors describing the devices; risks; updated surgical technique with directions to hand deliver surgeon/risk manager letters to all affected surgeons and risk managers in hospitals receiving the product. The notices included directions to remove all affected product and ensure risk manager letter is provided. The letter stated a Zimmer Inc. employee will conduct onsite reconciliation and asked return of distribution Spreadsheet to CorporateQuality.Postmarket@zimmer.com within 24 hours of receiving notification. Product is to be returned with certificate of sterilization to Zimmer Product Service Department 1777 West Center St. Warsaw, IN 46580 or Zimmer International Logistics, Eschbach Germany for International Accounts.For questions regarding this recall call 574-372-4807. |  | 94 | 94 | Worldwide Distribution - USA (nationwide) and internationally to Germany and Europe. | Open |  |  | 4035 | Not_Computer
Z-0556-2013 | 63775 | cane, safety walk | KHY | Cane. | Physical Medicine | Physical Medicine | 510(K) Exempt | Invacare Height Adjustable (walking) Canes, Model Numbers: 8916, 8917, and 40918-6. The device is packaged six (6) canes per shipping carton. The cane is used as an aid in walking or ambulation. | Cane | 2 | December 18, 2012 | 2012 | Invacare Corporation | Following their receipt of several customer complaints, Invacare recalled their height adjustable walking canes. The device was recalled from distribution based on reports indicating that there is a possibility that the button which holds the cane at the desired height, may suddenly and unexpectedly compress, causing the cane to collapse and cause possible injury to the user. | DESIGN: Device Design | The firm, Invacare, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated November 6, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to examine their existing stock using the enclosed detailed 'Product Tracking Sheet' and quarantine any affected canes; Acknowledge receipt and understanding of the urgent recall notice and indicate the number of canes remaining in inventory on the enclosed, 'Provider Response Card;'. Fax the completed Provider Response Card to Invacare Regulatory Affairs at: 1- 440-326-3544; disassemble each cane and discard both pieces in the trash; (Disassembly instructions are also provided); and lastly, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 8 p.m. EST, to order replacement merchandise at no charge.If you have any questions concerning these instructions, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 6 p.m. ET. |  | 38280 | 38,280 canes | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN KS, KY, LA,MA, MD ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH,OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Australia, Puerto Rico, and Taiwan Province of China. | Terminated | August 20, 2013 | 245 | 38280 | Not_Computer
Z-0557-2013 | 62826 | transport, patient, powered | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | ThyssenKrupp Accessibility LEVANT Stairlift.Intended to mechanically transport one mobility impaired person in a fold-down seat up and down straight stairs, indoors and outdoors. | ThyssenKrupp Accessibility LEVANT Stairlift. | 2 | December 18, 2012 | 2012 | ThyssenKrupp Access Corp | The seat could break as a result of damage resulting from a seat belt being trapped underneath the foldable part of the seat when the seat belt is not buckled. | DESIGN: Device Design | The firm, ThyssenKrupp Access sent an "IMPORTANT SAFETY AND PRODUCT USE INFORMATION" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact the firm at 1-800-925-3100 ext. 5484 for assistance regarding this recall. |  | 5413 | 5,413 stairlifts | Worldwide Distribution-USA (nationwide) including the states of IL, IN, MI, OH, WI, IA, KS, MN, MO, NE, SD, AR, LA,OK, TX. AZ, CO, ID, MT, NM, UT, MT, AK, CA, HI, OR, WA, AL, FL, GA, KY, NC, SC, TN, MD, PA, VA, WV, CT, MA, ME, NH, NJ, NY, RI and VT and the countries of Canada, Nova Scotia and British Columbia | Terminated | January 29, 2014 | 407 | 5413 | Not_Computer
Z-0568-2013 | 63739 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Medline Angio Drape Pack Radiology-LF; a sterile custom single patient prescription procedure pack, including a waste bag, for use by the Interventional Radiology Department; Packaged for Medline Industries, Inc.; Reorder #DYNJ35993DThis custom pack includes components necessary for a specific procedure. | Medline Angio Drape Pack RadiologyLF | 2 | December 19, 2012 | 2012 | Medline Industries Inc | The non-sterile Waste Bag component in the custom pack was attached to the outside of the pack in its own packaging. The product was not labeled or identified as sterile, but was assumed to be sterile by the customer. | PRODUCTION CONTROLS: Packaging | The firm, Medline Industries, Inc., telephoned the consignee/customer on November 8, 2012. The phone call informed the customer of the product, problem and actions to be taken. The customer was instructed to inform the staff at the facility not to use the pack until it had been reworked to replace the non-sterile waste bag with the sterile equivalent. It was also noted that the product does not need to be returned as it is being replaced directly at the facility.If you have any question call 847-837-2759 or email lsimmons@medline.com. |  | 105 | 105 packs | Nationwide distribution: New York only. | Terminated | December 20, 2012 | 1 | 105 | Not_Computer
Z-0561-2013 | 63744 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | AutoTac Delivery Handle, REF 400-200, Rx Only, non-sterile, BIOHORIZONS, BIRMINGHAM, AL 35244.Dental. | AutoTac Delivery Handle | 2 | December 19, 2012 | 2012 | BioHorizons Implant Systems Inc | The cleaning instructions provided may not be sufficient to remove residual contaminants after surgical use. | DESIGN: Component Design/Selection | BioHorizons sent an "URGENT DEVICE CORRECTION" letter on November 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact the Customer Care Department at (888) 246-8338 for questions regarding this recall. | Contact the recalling firm for information | 174 | 174 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Bulgaria, Canada, Cayman Islands, Chile, Colombia, Costa Rica, Ecuador, Egypt, Fiji, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Panama, Peru, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | September 24, 2013 | 279 | 1420 | Not_Computer
Z-0566-2013 | 63604 | stretcher, wheeled | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | The Stryker Model 6390, Power-LOAD, is a power-loading cot fastener system designed to lift, lower, or steer ambulance cots into and out of the ambulance. | Power Load | 2 | December 19, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Model 6390 Power-LOAD systems built between September 12, 2012 and September 27, 2012 may have been built with a load switch slide that does not meet specifications causing the lifting ability of the arms to be negatively impacted. | DESIGN: Component Design/Selection | All in process and Load Switch Slides within Stryker control were quarantined on 09/27/2012. Direct accounts were telephoned beginning 10/26/2012 with "Urgent Medical Device Correction" notices sent beginning on 11/30/2012. The notices explained the reason for the recall, hazards involved, and recommended actions. Stryker field service was to contact consignees to schedule replacement of the affected load switch slide. Consignees were asked to forward a copy of his notice to the new users and advise Stryker of their new location. Questions or concerns should be directed to (269) 389-6927 or (269) 389-6954. | Contact the recalling firm for information | 24 | 24 | USA, Nationwide including the states of PA, NM, MI, IA, CA, CO, WA, VA, IL, WI, NY, WY, MO, TX, and CT. | Terminated | August 30, 2013 | 254 | 24 | Not_Computer
Z-0569-2013 | 62764 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Cancellous ScrewPackaging: Product is housed within a polyurethane protector; this is placed inside an uncoated Tyvek/PET peel pouch and sealed. This package is placed into a paperboard carton, labeled and shrink-wrapped.PINNACLEÂŽ CANCELLOUS Bone Screws are ancillary fixation devices used in total hip arthroplasties. These screws are used to secure PINNACLE ACETABULAR Shell implants to the acetabulum via screw holes designed in the shell. Intended Use: This product is indicated for use in total hip replacement procedures for patients suffering severe pain and disability due to structural damage in the hip joint from rheumatoid arthritis, osteoarthritis, post-traumatic arthritis, collagen disorders, avascular necrosis, and nonunion of femoral fractures. | Pinnacle Cancellous Screw | 2 | December 19, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary recall of ten lots of the PinnacleÂż Cancellous Screw due to an incorrect description on the label. The label description reads Pinnacle Cannulated Screw and should read Pinnacle Cancellous Screw. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, DePuy Orthopedics, Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated July 2012 to its customers via email and mail. The notice described the product, problem and actions to be taken. The customers were instructed to determine whether they would like to retain and use the product or return affected product to DePuy in Warsaw, IN, and to complete and return the Reconciliation forms to your DePuy Sales Representative or by fax to: 574-371-4939. DePuy sales representatives are expected to aid customers in returning of recalled products.For questions about recall information provided, please contact the Manager, Post-Market Surveillance, 574-372-7333 (M-F; 8 am - 5 pm EDT.). |  | 118 | 118 units | Nationwide distribution: USA states of: GA, MA, MN, SC and VA. | Terminated | March 20, 2013 | 91 | 118 | Not_Computer
Z-0563-2013 | 63807 | slide stainer, automated | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | CoverStainer Slide RackThe CoverStainer Slide Rack is an accessory to the CoverStainer instrument. The purpose is to hold microscopic slides, during transfer through the staining process. The rack can hold up to 10 slides at a time, and is necessary for optimal CoverStainer instrument performance. | CoverStainer Slide Rack CS 10330 | 2 | December 19, 2012 | 2012 | Dako Denmark A/S | The expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. | DESIGN: Device Design | Dako sent a Voluntary Product Recall Notification dated November 13, 2012, to all affected customers. This voluntary recall notification was initiated because the expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. The result can be that slides are falling out of the rack during the staining process. While this has been observed at only a few customer sites, the corrective action includes all CoverStainer installations. Customers were notified that Dako would correct the rack situation by using a new manufacturer and has imroved the rack design. Customers with questions were instructed to contact their local Dako representative.For questions regarding this recall call 805-566-3036. |  | 25 | 25 Units US (50-100 racks per unit) 106 total | Worldwide Distribution - USA (nationwide) and the following countries:AustraliaFrance Rep. of KoreaBelgium Germany RussiaBrazil Israel Saudi ArabiaCanada Italy SpainChile Japan SwedenChina Mexico TaiwanDenmark Netherlands TurkeyEcuador Norway UKFinland Poland | Terminated | May 14, 2013 | 146 | 25 | Not_Computer
Z-0567-2013 | 63777 | stimulator, electrical, implanted, for parkinsonian tremor | MHY |  |  | Neurology |  | Medtronic, DBS Extension Kit for Deep Brain Stimulation, Catalog # 7482, 37085, 37086. Rx Only, Manufactured in: Medtronic, Inc. Villalba, Puerto Rico. MedtronicÂŽ DBSTM Therapy for Parkinson's Disease - Bilateral stimulation of the internal globus pallidus (GPi) or the subthalamic nucleus (STN) using Medtronic DBS Therapy for Parkinson's Disease is indicated for adjunctive therapy in reducing some of the symptoms of advanced, levodopa-responsive Parkinson's disease that are not adequately controlled with medication. STN/GPi target system components: ÂŽ KinetraÂŽ Model 7428 or SoletraÂŽ Model 7426 Neurostimulator ÂŽ ActivaÂŽ PC Model 37601, Activa RC Model 37612, Activa SC Model 37602, or Activa SC Model 37603 Neurostimulator ÂŽ Model 7482, Model 7482A, Model 7483, Model 37085, or Model 37086 Extension ÂŽ Model 3387 or Model 3389 Lead Medtronic DBS Therapy for Tremor - Unilateral thalamic stimulation of the ventral intermediate nucleus (VIM) using Medtronic DBS Therapy for Tremor is indicated for the suppression of tremor in the upper extremity. The system is intended for patients who are diagnosed with essential tremor or parkinsonian tremor not adequately controlled by medications and where the tremor constitutes a significant functional disability. VIM target system components: ÂŽ Soletra Model 7426 Neurostimulator ÂŽ Activa PC Model 37601, Activa RC Model 37612, Activa SC Model 37602, or Activa SC Model 37603 Neurostimulator ÂŽ Model 7482, Model 7482A, Model 7483, Model 37085, or Model 37086 Extension ÂŽ Model 3387 or Model 3389 Lead. | DBS Extension Kit for Deep Brain Stimulation | 2 | December 19, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has identified 140 Restore DBS Extension Kits (model numbers 7482, 37085, and 37086), which may contain a damaged winged connector boot. The connector boot may not have been molded properly and may be cracked. Approximately 14% of the 140 affected extension kits may contain a damaged radiopaque winged connector boot. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic beginning November 15, 2012 , initiated a retrieval of unused inventory in the field. In the US, three customers were identified. Medtronic field representatives visited these US accounts to facilitate product identification and removal. If product was removed a letter was provided to the facility describing the recall. The letter dated November 2012, "Medtronic Medical Device recall" was addressed to the Hospital. Beginning on December 3, 2012, Medtronic mailed healthcare professional letters to physician who have used the affected product. The Medtronic " Medical Device Recall" letter was dated November 2012. The letter identified the problem, the product involved in the recall, and provided advise if the product was used. A reply card was provided to each physician to confirm they have received and understood the communication. For further questions please call (763) 514-2000. |  | 140 | 140 units | Worldwide Distribution -- USA (nationwide) including the states of CT, MI, NH, NY, OH, PA and VT., and the countries of AUSTRALIA, JAPAN, BELGIUM, FRANCE, GERMANY, ITALY, NETHERLANDS, POLAND, SPAIN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | November 14, 2013 | 330 | 140 | Not_Computer
Z-0558-2013 | 63776 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | MAYFIELD Composite Series Base Unit, Model A3101, a Swivel Adaptor (A3018) is an integral component of the Base Unit, Model A3101.The MAYFIELD Base Unit (A3101) is an operating room table accessory intended to be used to support a patient during diagnostic examination and/or surgical procedures where a rigid support between surgical table and headrest, or skull clamp is necessary, positional freedom is required and where X-ray imaging modalities will be used. A Swivel Adaptor (A3018) is an integral component of the Base Unit A3101. | MAYFIELD Base Unit | 2 | December 19, 2012 | 2012 | Integra LifeSciences Corporation | An investigation of an adverse trend of complaints conducted by Integra LifeSciences identified various product performance issues which could impact the functionality of the firm's Mayfield Composite Series Base Units and Mayfield Skull Clamps. In order to prevent additional complaint reports, the firm decided to recall and replace certain lot codes of the Base Units and Skull Clamps that were d |  | Integra issued an 'Urgent: Voluntary Medical Device Recall Notification with 'Recall Acknowledgment and Return Form' to all Integra Sales Representatives and Customers via visit and E-mail on November 5, 2012. The recall notice informs the customers of the firm's receipt of complaints pin-pointing quality issues with the MAYFIELD Base Units and MAYFIELD Skull Clamps which may impair their functionality. The customers are asked to sign and return the enclosed Recall Acknowledgement and Return Form so as to document their receipt of the recall notification and replacement products. The replacement products are to be delivered by the Integra NeuroSpecialist who is also responsible for returning the signed forms to Integra as documentation that the recalled devices have been replaced. Should the customer(s) have questions or concerns, they are directed to contact Integra Customer Service at: 1-877-444-1114,Option 3 for Technical Support. |  | 5 | 5 base units were distributed | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA TN, TX, UT, VA, WA and WI. | Terminated | February 21, 2014 | 429 | 13 | Not_Computer
Z-0570-2013 | 63737 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix PMIC-108 Panels, Catalog number 448418, labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, USA.The BD Phoenix" Automated Microbiology system is intended for the rapid identification (ID) and antimicrobial susceptibility testing (AST) of clinically significant bacteria. | BD Phoenix PMIC108 Panels | 2 | December 19, 2012 | 2012 | Becton Dickinson & Co. | An antibiotic susceptibility test kit may produce false intermediate or false resistant results with clinically significant bacteria. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, BD, sent an "URGENT PRODUCT RECALL" letter dated November 2012 via the telephone and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and discard any remaining packages for replacement and to complete and return the Acknowledgement Receipt Form, whether or not you have any inventory remaining, via fax to: BD Regulatory Compliance at 410-316-4258. BD will issue replacements for the discarded materials.If further assistance is needed regarding replacements please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. | Contact the recalling firm for information | 12250 | 12,250 tests | Nationwide distribution: USA including states of: AR, CA, KY, OK, RI, WI, and WY. | Terminated | January 09, 2013 | 21 | 12250 | Not_Computer
Z-0564-2013 | 62394 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Case Carton:BD Oral Dispensing Syringe1 mL Clear with Tip CapBD, Franklin Lakes, NJ 07417 USA Made in USA www.bd.comBD, Laagstraat 57, B-9140 Temse, BelgiumOn 100 unit plastic bag:BD oral syringes are intended to dispense oral medications. | Becton Dickinson and Company | 2 | December 19, 2012 | 2012 | Becton Dickinson & Company | BD received a complaint report for mixed syringe tips (oral and non-oral syringe tips) within the same lot. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Becton Dickinson sent a Urgent Product Recall letter/Customer Packing list dated October 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review inventory and determine if they have any of the affected product. If so, it should be removed from inventory. The enclosed customer list should be completed as well. Any questions please call 1-201-847-4267. | Contact the recalling firm for information | 2786000 | 2,786,000 units | (USA) nationwide | Open |  |  | 2786000 | Not_Computer
Z-0593-2013 | 63057 | unit, phacofragmentation | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Bausch + Lomb Stellaris PC Vision Enhancement System, Posterior Fluidics Module (PFM). The system is intended for the emulsification and removal of cataracts, anterior, and posterior segment vitrectomy. | Bausch Lomb Stellaris PC Vision Enhancement System, Posterior Fluidics Module (PFM) | 2 | December 20, 2012 | 2012 | Bausch & Lomb Inc | Posterior fluidics modules installed in the system may need to be recalibrated. | PRODUCTION CONTROLS: Process Control | "Urgent Product Recall" notification letters dated August 31, 2012 were issued via first class mail, e-mail, and/or phone calls to affected customers. The product issue and the field correction was provided. Customers in the US may contact 1-800 338-2020 for questions. | Contact the recalling firm for information | 124 | 124 units | Worldwide Distribution was made in the U.S., Canada, Australia, Singapore, Korea, Switzerland, France, Germany, United Kingdom, Portugal, Italy, Spain, Ireland, Africa, UAE, Jordan, Turkey, Ukraine, and Lebanon.There was no government or military distribution. | Terminated | July 11, 2013 | 203 | 124 | Not_Computer
Z-0572-2013 | 63784 | radioimmunoassay, luteinizing hormone | CEP | Luteinizing hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Access hLH assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of luteinizing hormone (LH) levels in human serum and plasma using the Access Immunoassay Systems.The Access hLH Calibrators are intended to calibrate the Access hLH assay for the quantitative determination of luteinizing hormone (LH) levels in human serum and plasma using the Access Immunoassay Systems. | Access hLH Calibrators | 2 | December 20, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter Inc. (BEC) has demonstrated that real time stability testing of hLH Calibrator of certain lots have failed the 13-month and 14-month time points. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these lots did not meet BEC's procedural requirements for a 12-month shelf-lif | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Beckman Coulter Inc. (BEC) sent an "URGENT FIELD SAFETY NOTICE" dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these the affected lots did not meet BEC's procedural requirements for a 12-month shelf-life claim.Contact the firm at 952-448-4848 for assistance with this recall. |  | 14409 | 14409 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Georgia, Germany, Ghana, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Moldova Republic, Morocco, Myanmar, Namibia, Netherlands, Netherlands Antilles, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, Yemen, Zambia, and Zimbabwe | Open |  |  | 14409 | Not_Computer
Z-0607-2013 | 63667 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cervix Applicator Set, part number GM11004310, a component used with the tube for segments and cervical sleevesThe guide tube for segments and cervical sleeves in the Segmented Cervix Applicator Set was developed to treat cancer of the cervix, vagina, endometrium, vaginal stump and uterus. It is used for Brachytherapy. | Segmented Cervix Applicator | 2 | December 20, 2012 | 2012 | Varian Medical Systems, Inc. | The guide tube which is part of the Segmented Cervix Applicator Set is difficult to disconnect from cervical sleeve. | DESIGN: Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE" letter dated September 28, 2012 to all affected consignees/Customers. The letter identified the affected product, description of problem and recommended user actions to be taken. The consignees/customers were instructed to immediately remove all affected guiding tubes from clinical use, and sequester the components for pending return to Varian Brachy Therapy. Consignees/Customers were requested to complete the attached proof of notification form and return it to Varian Medical Systems.For questions contact-AT&T; international direct dialing network (http://www.business.att.com/bt/dial_guide.isp. Then after hearing the AT&T; signature noise, and. "Please enter the number you are calling", dial 800-360-7909; UK:+44 1293-601-327 or US:+1-434-979-1540. To submit a support request or general inquiry, compose an email with a blank subject line to: brachyhelp@varian.com. |  | 15 | 15 units | Worldwide Distribution - U.S. state of Georgia and the countries of Canada, Germany, India, Poland, and Thailand. | Terminated | April 02, 2013 | 103 | 15 | Not_Computer
Z-0577-2013 | 63545 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SOL SYS 8IN CALC 1.5/18.0Packaging: Product is packed within a poly protector, with protection added as needed for fragile points. The entire stem is then placed within a foam sleeve for added protection. The assembly is then placed and sealed in an inner and outer tray for aseptic transfer. The sealed package is placed into a folding carton, labeled, and shrink wrapped. Intended Use: SOLUTION System Calcar Stems are used in total hip arthroplasty (THA) to provide increased patient mobility and reduce pain by replacing the damaged hip joint articulation in patients where there is evidence of sufficient sound bone to seat and support the components. | SOL SYS 8IN CALC 1.5/18.0 | 2 | December 20, 2012 | 2012 | DePuy Orthopaedics, Inc. | An Engineering Analysis has determined the Solution Calcar Stem has the potential to result in impingement with a variety of heads. This impingement may not allow the stem and head to lock as intended. The amount of possible impingement varies depending on the femoral head. Worst case condition: Calcar stem with -2mm offset articuleze maximum of 0.937mm radial impingement. It is evident that a | DESIGN: Device Design | Deputh Orthopaedics sent an Urgent Information - Medical Device Recall Notice dated December , 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The notification letter instructed consignees to please cease further distribution or use of products immediately. The purpose of the communication was to inform them of this recall and request acknowledgement of receipt of this letter by their facility. Reconciliation forms should be returned to your DePuy Sales Representative or faxed to the following Fax Number: 574-371-4939.For questions regarding this recall call 574-267-8143. |  | 1732 | 1732 total product codes | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR | Terminated | November 13, 2013 | 328 | 1732 | Not_Computer
Z-0571-2013 | 63788 | exerciser, powered | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | SpineSix consists of a motorized treatment table that is controlled by an integral computer. The system is operated by means of 12-inch Color LCD touch screen that displays treatment time, positioning and magnitude and direction of motion.There are eight fully customizable Patient Protocol settings for commonly usedtreatment settings.The SpineSix BioMotion Spinal System is intended to reproduce normal spine mechanics and physiological movements so that a patient can comfortably and safely experience functional range of motion in a non-weight-bearing environment. | SpineSix01/SpineSix02 | 2 | December 20, 2012 | 2012 | BioMotion Medical Systems, LLC | BioMotion is issuing a field correction for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to rectify all regulatory deficiencies involving the former owners of the SpineSix product and concurrently preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. |  | BioMotion issued a field correction letter dated October 18, 2012 for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to all affected customers. The letter identified the product, problem, and actions to be taken. BioMotion is preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. Contact the firm at 772-405-7899 for questions concerning this letter. | Contact the recalling firm for information | 28 | 28 units | Distributed in Florida, Georgia, Tennessee, New Jersey, New York, Pennsylvania, and Wisconsin. | Terminated | November 05, 2013 | 320 | 28 | Not_Computer
Z-0596-2013 | 61734 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medrad(R) Continuum MR Infusion System Standard Administration KitPeristaltic Pump System designed for patients who require medications and other fluids during an MR procedure. It is intended to provide infusion therapy directly prior to, during and immediately after the MR procedure functioning while either stationary or mobile. | Medrad(R) Continuum MR Infusion System Standard Administration Kit | 2 | December 20, 2012 | 2012 | Medrad Inc | Bayer HealthCare is recalling all MEDRAD(R) Continuum MF Infusion System tubing sets (macrobore and microbore), catalog numbers MIK200A, MIL 200B, 3015153, 3015155, 3015156, 3015157, 3015158, 3015159, and 3015161. The firm is initiating this recall because this tubing exhibits a variation that, when used in conjunction with any Continuum pump, may result in flow rate accuracy performance issues. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Bayer HealthCare, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 17, 2012 via the USPS certified mail to its consignees/customers. The firm expanded the recall on May 11, 2012 and sent a second letter entitled "EXPANDED URGENT MEDICAL DEVICE RECALL" to its affected consignees/customers notifying them of the additional issue. The letters described the product, problem and actions to be taken. The customers were instructed to return pump now; contact customer support at 1-877-229-3767 for a return goods authorization (RGA); complete and return Response Form, include RGA numbers, via Fax to Bayer Compliance at 412-4060942, and/or wait to return the pump when they begin contacting customers with this subset of pumps in August to make arrangements for a replacement pump.If you have questions, please contact our customer support team at 1-877-229-3767 or visit www.medrad.com for more information. |  | 1229 | 1,229 Pumps total | Worldwide distribution: USA (nationwide) and countries of: Australia, Bermuda, Brazil, Canada, China, Chile, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Great Britain, Guam, Hong Kong, India, Israel, Italy, Japan, Korea, Lebanon, Mexico, Malaysia, Netherland, Poland, Portugal, Puerto Rico, Sweden, Singapore, Taiwan, Thailand, and Venezula. | Open |  |  | 1229 | Not_Computer
Z-0608-2013 | 37530 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Qualicheck 1+, Aqueous Quality Control, Level 2, S7140, REF 944-042. The solution is used for quality check of blood gas analyzers. | Qualicheck 1, Aqueous Quality Control, Level 2, S7140, REF 944042. | 3 | December 21, 2012 | 2012 | Radiometer America Inc | The barcode for the ABL77 on the insert incorrectly enters a control range for pH of 7.90-7.95 into the analyzer. The insert correctly states that the control range is 7.40-7.45. | PRODUCTION CONTROLS: Process Control | Radiometer sent Field Action Notes regarding the affected product. The note identified the product, the problem and the action to be taken by the customer. Instead of scanning the barcorde on the insert to enter the QC data into the anyalyzer, customers were instructed to manually key in the QUALICHECK type, lot number, and correct control range.For questions regarding this recall call 440-871-8900. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 840 | 840 Ampules | Distributed to KY, NC, and NY. | Terminated | February 05, 2013 | 46 | 840 | Not_Computer
Z-0525-2013 | 63729 | regulator, pressure, gas cylinder | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Praxair Vantage Grab n Go - Gas Cylinder pressure regulatorCatalog Number: OX-M-AEGNGVNTG Usage: Regulator reduces pressure in oxygen-containing cylinder. | Praxair Vantage Grab n Go | 1 | December 21, 2012 | 2012 | Praxair Inc. | Isolated incidents of ignition inside Grab n' Go cylinders that had been knocked over or otherwise subject to significant physical impact. | DESIGN: Device Design | The firm, PRAXAIR, sent a "MEDICAL DEVICE CORRECTION" letter dated November 26, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Grab'n Go operating instructions provided in the letter which include: storing the Grab'n Go systems indoors in a safe area; never store a system in an elevated location; handle systems safely and avoid jarring or dropping the system and locate at the following web link: www.praxair.com/GNGinstructions. The customers were also instructed to ensure that all of the Grab'n Go units that were sent by Praxair are sent to Praxair for servicing and refilling after use, in accordance with their normal process.If you have any questions, contact your sales representative at 630-320-4452 or email: Thelma_Brantley@Praxair.com. |  | 8877 | 320,451 units (255,690 US ; 55,884 Canada; 8877 OUS) | Worldwide distribution: USA (nationwide) and countries of: Brazil, Canada, China, Columbia, Costa Rica, India, Paraguay and Uruguay. | Open |  |  | 320451 | Not_Computer
Z-0611-2013 | 63676 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line.The MobiCath Bi-Directional Guiding Sheath is intended for introducing various cardiovascular catheters into the heart, including the left side of the heart through the interatrial septum. | MobiCath BiDirectional Guiding Sheath | 2 | December 22, 2012 | 2012 | Greatbatch Medical | Greatbatch Medical observed some anomalies of the inner lumen of the MobiCath Bi-Directional Guiding Sheaths upon implementation of a new visual inspection tool. The anomalies may include loose, string-like liner material, liner abrasion (fraying), and/or attached string-like liner material (from scratches or damage on the inner wall of sheath lumen). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biosense Webster sent an "Urgent Field Safety Notice" dated November 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Voluntary Field Removal Certification Form was attached for customers to complete and return. Contact the firm at 800-729-9010 for questions related to this issue. |  | 10201 | 10,201 units | Worldwide Distribution-USA including the states of AL, AK, AZ, ALR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KY, MD, MA, MI, MN, MO, NE, NV,NM, NJ, NH, NY, NC, OH, OK OR, PA, RI,SC,SD, TN, TX, UT,VT, VA, WA, and WI, and the countries of AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HUNGARY, IRELAND, ISRAEL, ITALY, THE NETHERLAND, NORWAY, SWEDEN, AUSTRALIA, and CANADA | Open |  |  | 10201 | Not_Computer
Z-0610-2013 | 63457 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics; Stryker T2 Humerus SystemTeflon Tube, sterile; REF 1806-0073SThe T2 Nailing System is designed to treat complex fractures in Humeral, Femoral, Tibial and Ankle applications. The Teflon Tube is an optional instrument used to exchange the Ball Tip Guide Wire with the Smooth Tip Guide Wire once reaming is completed.Stryker Stryker Trauma GmbH 24232 Schonkirchen, Germanydistributed in USA by: Howmedica Osteonics Corp., 235 Corporate Drive, | T2 Humerus Sterile Teflon Tube | 2 | December 22, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | During the manufacturing of the affected lot a manufacturing error has led to embrittlement of the tube material of the T2 Humerus Teflon Tube. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter & Acknowledgement form dated August 7, 2012 via FedEx Return Receipt to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to inform their users of this Medical Device Removal and pass this notice to all those individuals who need to be aware within their organization; return all affected products available at their location to: Stryker Osteosynthesis, 325 Corporate Drive, Mahwah, NJ 07430 REF: PFA#2012-124 or contact Stryker customer service; complete and sign the enclosed Product Recall Acknowledgement Form and Fax to 1-855-251-3635 or email to Recall Coordinator at janice.pevide@stryker.com and keep a copy of their records.If you have any questions, contact the Manager, Regulatory Compliance, at 1-201-972-2100 or email: colleen.omeara@stryker.com. |  | 100 | 100 units (73 units US; 27 units OUS) | Worldwide distribution: USA (nationwide) including states of: AL, CO, CT, IL, IN, MO, MS, NC, NJ, NY, OR,and TX; and countries of: Australia, Canada, France, Spain and United Kingdom. | Terminated | September 11, 2013 | 263 | 100 | Not_Computer
Z-0613-2013 | 61212 | calibrator, primary | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Total T4 Calibrators, Part Number: 33805.The Access Total T4 Calibrators are intended to calibrate the Access Total T4 assay for the quantitative determination of total thyroxine (T4) levels in human serum and plasma using the Access Immunoassay Systems. Total T4 assay results are used for diagnosis and confirmation of thyroid disorders such as hyper- and hypothyroidism. | Access Immunoassay Systems Total T4 Calibrators | 2 | December 22, 2012 | 2012 | Beckman Coulter Inc. | A recall communication was initiated because Beckman has identified that Access Total T4 calibrator lots 019818, 021654, and 024072 do not meet their 12 month expiration date claim. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A recall communication was initiated on July 27, 2011 with Beckman forwarding a Product Corrective Action (PCA) letter with attached PCA response form to all their customers who purchased the Access Immunoassay Systems Total T4 Calibrators. The letter provides the customers with an explanation of the problems identified and actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers can call for product replacement in the US, by contacting Client Services at (800) 526-3821, option 1. Customers in Canada can contact Customer service at (800) 463-7828. Customers with questions regarding this notice can contact Customer Support Center at http:www.beckmancoulter.com or call (800) 854-3633 in the US and Canada. |  | 6634 | 6,634 units total (2,742 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Egypt, El Salvador, Ethiopia, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Kazakhstan, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Macao, Malaysia, Mexico, Netherlands, Nigeria, Pakistan, Panama, Philippines, Puerto Rico, Romania, Russian Federation, Serbia, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zambia. | Terminated | December 26, 2012 | 4 | 6634 | Not_Computer
Z-0594-2013 | 63683 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Ardis Inserter, Rx, non-sterile | Ardis Interbody System Inserter | 1 | December 22, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine has received reports of the PEEK Ardis Interbody Spacer breaking into fragments when too much lateral force is applied to the Ardis Inserter. If the PEEK Ardis Interbody Spacer breaks during implantation, surgical delays of up to 60 minutes could occur. Sharp edges of the broken implant may contact and damage adjacent structures, requiring surgical intervention. Observed health risks | DESIGN: Device Design | Zimmer Spine, sent a "MEDICAL DEVICE CORRECTION" memo dated November 12, 2012 to all consignees / customers describing product, problem and actions to be taken including revised surgical technique. Beginning the week of December 17, 2012, Zimmer began the immediate retrieval of the affected units as a Class I recall. December 20, 2012 nation wide press outlining the recall was issued. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer Spine by calling 866-774-6368 between the hours of 8AM and 5PM CST or by email at qualitycompliance@zimmer.com. |  | 315 | 315 (172 in USA) | Worldwide distribution: USA (nationwide) including states of: AR, AZ, CA, CO, CT, DC, FL, GA, ID, LA, MD, MI, MN, MO, NC, NV, NY, OH, OK, PA, TN, TX, WI, and WV; and countries of: Australia, El Salvador, France and Singapore. | Terminated | February 25, 2014 | 430 | 315 | Not_Computer
Z-0612-2013 | 63632 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay System,Access 2 Immunoassay System,Synchron LXi 725 ClinicalSystem, UniCel DxC 600iSynchron Access Clinical SystemPart Numbers: Access  81600;Access 2  81600N, 386220, A69186, A25656, A12900;Access 2 (LXi)  A15642;Access 2 Section, DxC 600i  A25640.The Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. The Synchron LXi and the UniCel DxC 600i are Access systems integrated with a chemistry analyzer. | Access Immunoassay System | 2 | December 22, 2012 | 2012 | Beckman Coulter Inc. | It has been determined that, in some cases,pulley drive components have broken parts or have become loose, mispositioned, ordisconnected from their drive shafts. When these defects occur, the pulley may not turn at all, ormay turn at an incorrect speed, resulting in inadequate washing or mixing of the contents of thereaction vessel. | DESIGN: Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 1, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Contact Technical Support for questions regarding this notice. |  | 6858 | 6,858 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 23, 2013 | 366 | 6858 | Not_Computer
Z-0616-2013 | 63497 | [6 empty] | Brookstone Heated Body Bean, electrically heated SKU numbers: 734863 (Blue), 734864 (Plum), 734865 (Red).Brookstone 3 in 1 wrap ,SKU734397 (Brown).The Products are liquid filled bean shaped pouches that are electrically heated and then detached from the charger prior to use. | Brookstone Heated Body Bean | 2 | December 27, 2012 | 2012 | Brookstone Company, Inc. | Additional Warning Statement to prevent burns to the skin | DESIGN: Labeling Design | Brookstone sent a "MEDICAL DEVICE CORRECTION" letter dated December 10, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Recall Information is on the website link from the homepage, www. Brookstone.com. Recall posted: http://www.brookstone.com/recall-information. Questions call Brookstone Customer Care at 800.846.3000 or visit any Brookstone Store. |  | 109956 | 109,956 units | Nationwide Distribution. | Terminated | June 13, 2013 | 168 | 109956 | Not_Computer
Z-0615-2013 | 63857 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Flowtron Trio DVT Pump; an Rx intermittent pneumatic compression deep vein thrombosis pump; ArjoHuntleigh AB, Eslov, Sweden; Model 512003; a non-invasive prophylaxis system for reducing the incidence of deep vein thrombosis. | Flowtron Trio Pump | 2 | December 27, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | The Flowtron Trio DVT Pump has been marketed without proper FDA Approval. | PREMARKET APPROVAL: No Marketing Application | ArjoHuntleigh sent an Urgent Field Correction Recall letters dated December 10, 2012, to all affected customers, informing the accounts that the pumps were marketed without proper FDA approval, and that ArjoHuntleigh is removing the pumps from the market and replacing them with a suitable alternative. The accounts were requested to disseminate the information to all organizations where the devices may have been transferred, and to complete the Customer Response Form and return it to ArjoHuntleigh by mail, fax or e-mail. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. | Contact the recalling firm for information | 2859 | 2,859 unts | USA (nationwide) including the states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Dakota, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin and Wyoming. | Open |  |  | 2859 | Not_Computer
Z-0619-2013 | 63825 | turbidimetric method, protein or albumin (urinary, non-quant.) | JIQ | Urinary protein or albumin (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Analyzing system consists of a small and portable analyzer and plastic microcuvettes. The microcuvette contains reagents deposited on its inner walls. The urine sample is drawn into the cavity by capillary action. The filled cuvette is inserted into the device where the contents of the cuvette are mixed through vibration. Within 90 seconds, the immunochemical reaction is completed and the turbidity is measured photometrically at 610 nm. The albumin concentration is proportional to the turbidity. When the end point is reached, the result is displayed in mg/L.The system can be used for the quantitative determination of low levels of albumin in urine for the purpose of screening for, diagnosing, monitoring and to supplement the clinical evidence in the treatment of microalbuminuria. | HemoCue Albumin 201 Analyzer | 3 | December 28, 2012 | 2012 | HemoCue AB | Incorrect version of Quick Reference Guide (QRG) was delivered with products in shipment made from April1 st 2011 and onwards. In addition, a warranty statement was missing in the analyzer package. | PRODUCTION CONTROLS: Labeling Mix-Ups | HemoCue sent an "IMPORTANT MEDICAL DEVICE CORRECTION NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Correction Notification Verification Form was included for customers to complete and return. Contact the firm at 562-668-5741 for questions regarding this notification. |  | 48 | 48 affected analyzers | Distributed in the states of FL, MA, CA, OR, NY, NC, MO, ID, IN, KY, PA, TX, OK, IL, GA, KS, and UT. | Terminated | March 14, 2013 | 76 | 48 | Not_Computer
Z-0618-2013 | 63769 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack with or without coating ( Carmeda BioActive Surfact, Carmeda BioActive Surfact and Trillium Biosurface,or Balance Biosurface) with the following Model numbers:0E27R16, 1A30R6, 2493R24, 2493R25, 5D56R5, 5Z93R4, 5Z93R5, 7E64R2, 7J53R2, 7M13R1, 7M14R1, 7N16R, 7P25R1, 7P93R1, 7Q11R1, BB7L63R2, BB7N26R, CB5174R11, CB175R13, CB5Q03R6, CB5Q03R7, CB6C53R6, CB7C15R2, CB7C59R2, CB7C60R2, CB7C74R2, CB7D91R5, CB7E35R2, CB7E38R1, CB7G21R4, CB7L48R1, CB7L72R1, CB7P82R1, HY6U96R2, HY6Y52R2, HY7E87R1, SS7J91R3, SSCB7L48R, SSTL7G78R1,TL5S33R7, TL6VTTR1, TL7B51R1, TL7G20R3, TL7G78R3, TL7R87R1. Sterilized Using Ethylene Oxide, Nonpyrogenic, Assembled in Mexico, Manufacturer Medtronic, Inc, Minneapolis, Mn 55432Product is used by perfusionists as part of the extracorporeal circuit during cardiopulmonary bypass procedures, and is configured specifically as designated by each customer. | Medtronic Intersept Tubing Packs | 2 | December 28, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is initiating an Urgent Medical Device Customer Notification. We have confirmed that certain Medtronic Intersept Tubing Packs contain packaging trays that are susceptible to damage that can compromise product sterility. | DESIGN: Packaging Design/Selection | Medtronic sent a "Urgent Medical Device Customer Notification" letter dated November 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately identify and quarantine all affected product and to return affected product, but if they have concerns about replacement than they recommend to inspect and return any damaged product..For further questions please contact your local Medtronic sales representative or Lifeline Technical Services at 1-877-526-7890. Outside the US, please call1+763-526-7890. We apologize for any inconveniences this issue may have caused you or your institution. |  | 1552 | 1552 | Worldwide Distribution--USA (nationwide) including the states of AZ, CA, CT, FL, GA, IL IA, MD, MA, MI, MS, NJ, NY, OH, OK, SC, TN, TX and WI. and the countries of CANADA and MALAYSIA. | Terminated | April 04, 2013 | 97 | 1552 | Not_Computer
Z-0620-2013 | 63915 | pacemaker battery | DSZ | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Companion External Battery, Part Number: 293001-001The Companion External Battery is a component of the Companion 2 Driver System and the Companion Driver System. Both Driver Systems, with the temporary Total Artificial Heart (TAH-t) are indicated for use as a bridge to transplantation in cardiac transplant-eligible candidates at risk of imminent death from biventricular failure. | Syncardia | 2 | December 28, 2012 | 2012 | SynCardia Systems Inc. | Syncardia Systems is recalling the Companion External Batteries used with the Companion 2 Driver System because they may potentially be at risk of thermal events and/or other physical damage such as melting, charring and/or burning. | PRODUCTION CONTROLS: Process Control | Syncardia Systems sent a" Urgent Field Safety Notice" dated December 10, 2012, to all affected customers. The letter identified the product the problem and what action the customer needs to take.. The letter informed the customers that there were 3 reported thermal battery events (2 of which occurred on-site at either SynCardia or at the battery manufacturer, Dow Kokam, and one incident that occurred in the field) which caused physical damage to the batteries including burning and/or charring. The batteries were not in use with a patient at time of the reported issue, and there were no reports of any adverse impact to any patient or user. Customers are instructed to follow the directions on the attached Return page, and return their customer response form to SynCardia as soon as possible.Customers with questions are instructed to call (520) 545-1234 or (520) 440-7593. |  | 81 | 81 units | Worldwide Distribution -- US (nationwide) | Terminated | June 07, 2013 | 161 | 81 | Not_Computer
Z-0617-2013 | 63855 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Fast-Cath Transseptal Guiding Introducer, 90, 8F, REF 406586, Length 60 cm, Sterile EO, Rx only, St. Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55343Indicated for introducing various cardiovascular catheters into the left side of the heart through the interatrial septum. | FastCath" Transseptal Guiding Introducer | 3 | December 28, 2012 | 2012 | St Jude Medical Inc | St. Jude Medical is recalling a single batch of Fast Cath Transseptal Guiding Introducers which was packaged as a 180 degree curve style introducer instead of a 90 degree curve style introducer as indicated on the labeling. | PRODUCTION CONTROLS: Process Control | St Jude Medical hand delivered the" Product Recall" letter dated March 27, 2012, to the affected customers on April 5, 2012, and March 28, 2012. The letter identified the product, the problem and the action needed to be taken by the customer.The returned devices and those not distributed were quarantined at the SJM Minnetonka distribution center until their destruction.Customers were requested to to complete and return the reconciliation form.Further questions please call (651) 756-6526. |  | 26 | 26 | US Distribution only including the states of FL, KY and OK. | Terminated | July 12, 2013 | 196 | 26 | Not_Computer
Z-0623-2013 | 63862 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Space Pump Pole Clamps for the Infusomat Space Pump, model number 8713130.The Pole Clamp is an accessory to Space Pumps designed to support 1 to 3 pumps on a pole. | B. Braun Space Pole Clamps for the Infusomat Space Pump | 2 | December 31, 2012 | 2012 | B. Braun Medical, Inc. | Firm has become aware of the potential for breakage of the pole clamp locking grids, or rasters, in the Space Pole Clamps. | DESIGN: Device Design | B. Braun Medical, Inc sent an Urgent Medical Device letter dated March 23, 2012 , to all affected customers. The letter identified the product the problem and the action needed by the customer to take.B. Braun has implemented an improvement to the pump locking mechanism of t he older style pole clamps and will be upgrading any older style pole clamps that may remain within your facility.Please return the completed form from attachment 1 in the self-addressed envelope provided or fax the form to (610) 849-1197 within two (2) weeks of receipt\.If you require any additional information or have any questions, comments or concerns please contact your B. Braun Medical Customer Service at 800-627-7867. |  | 13078 | 13,078 | Worldwide Distribution--US (nationwide) including the states of AL, AZ, NC and TN. | Open |  |  | 13078 | Not_Computer
Z-0626-2013 | 63772 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list number 11005 Product Usage: The Plum A+ Hyperbaric Infusion System is exclusively designed for hyperbaric therapy in monospace and multiplace chambers. | Plum A Hyperbaric Single Channel Infusion Pumps | 2 | December 31, 2012 | 2012 | Hospira Inc. | The door roller assembly on the Plum A+/A+3 infuser family of infusers has the potential to break which can lead to possible unrestricted flow and/or over-delivery during the removal of the IV administration set's cassette from the pump. | DESIGN: Device Design | Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters, dated December 7, 2012, to all affected customers. The letters identified the products, problem and action(s) to be taken until doors have been replaced. The customers were instructed to read the information provided carefully, including instructions to inspect the door roller assembly prior to loading a cassette for signs of damage. Customers are to complete the attached Reply Form and return it to the firm via fax at 1-877-650-8359. Customers were instructed to notify their consignees if they have further distributed the affected product. If customers require additional assistance, they are to contact Stericycle at 1-877-650-0362.. |  | 180 | 180 units | Worldwide Distribution -- USA, including Puerto Rico, and the countries of Austria, Australia, Argentina, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, Finland, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates, and Vietnam. | Open |  |  | 326612 | Not_Computer
Z-0622-2013 | 62413 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Standard Slingbar 450.The Standard Slingbar 450 is a versatile sling bar which can be used for most lifting situations. It can be used with different Liko ceiling and mobile lifts. | Liko Standard Slingbar 450. | 2 | December 31, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom has received reports from facilities which indicate that serious injuries have been received related to the Standard Sling bar 450. There have been 6 complaints of injury related to the product over the past 32 years (6 out of 838,687,856 estimated patient uses). One of the injuries was associated with a patient during use, four injuries occurred when the slingbar was not in use wher | DESIGN: Device Design | Hill-rom sent an URGENT MEDICAL DEVICE RECALL notification letter dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.. Customers were requested to please read, fill in completely, and return the Response Form to Hill-Rom immediately. This information will be compiled and reported to the applicable regulatory governing bodies. If we do not receive your response we will be required to contact you asking for your response.Continue to follow safe lifting practices, and do not allow unauthorized persons, especially children, to play around or operate the lift.If you have any questions concerning this Customer Safety Notification, please contact Hill-Rom Technical Support at 800-445-3720. |  | 261 | 261 units | Worldwide Distribution -- US (nationwide) including the states of AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, WA, WV, and WY., and the country Canada,We have not yet defined our European or International facilities at this time. | Open |  |  | 261 | Not_Computer
Z-0630-2013 | 63616 | system,network and communication,physiological monitors | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM Central StationMindray DS USA, Inc. 800 MacArthur Boulevard, Mahwah, NJ 07430The Central Monitoring Station (CMS) network transfers information between DOM Central Station and other network devices. It also allows information transfer between several CMS. Network connections consist of hardwired network cables and or WLAN connections. The DPM Central Station is used for remote monitor management, storing, printing, reviewing, or processing of information from net worked devices, and it is operated by medical personnel in hospitals or medical institutions. | DPM Central Station | 2 | January 03, 2013 | 2013 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is a software anomaly with the DPM Central Monitoring System where the trend data of one patient maybe replaced with the trend data of a second patient. | DESIGN: Software Design | Mindray sent a Product Corrective Action letter dated May 12, 2012, to all affected customers via certified mail return receipt requested to two medical facilities. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the issue would be corrected through a softweare upgrade of the affected product. A Mindray Service Representative would contact their facility to arrange for the software upgrade at no cost to them.For questions regarding this recall call 1-800-288-2121. |  | 2 | two units | Nationwide Distribution including OH and GA | Terminated | February 19, 2013 | 47 | 2 | Computer
Z-0634-2013 | 63746 | clinical sample concentrator | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated fluid handling system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14The Abbott m2000 System is intended for use in performing nucleic acid testing in clinical laboratories. | Abbott m2000sp Instrument | 2 | January 03, 2013 | 2013 | Abbott Molecular | An issue with the m2000sp Instrument software has been identified whereby, after a series of steps, if changes are made to the control or calibrator information on the Sample Extraction Assay Details screen, the changes are not saved. | DESIGN: Software Design | Abbott Molecular Inc. sent a Field Correction Recall/Urgent Field Safety Notice letter dated December 4, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to their m2000sp software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2013. They were requested to complete the enclosed reply sheet and fax it to (847) 775-6728 at prior to December 18, 2012 to acknowledge receipt and understanding of the letter.For questions customers were instructed to contact their local Abbott Molecular representative.For questions regarding this recall call 224-361-7274. |  | 908 | 215 units - U.S., 693 units - ROW | Worldwide Distribution - USA (nationwide) and internationally to Algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Burkina Faso, Burundi, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ethiopia, Estonia, Finland, France, Gaza & Jericho, Germany, Greece, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mexico, Mozambique, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam and Zambia | Open |  |  | 908 | Computer
Z-0633-2013 | 63838 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PENTRA C200, Part Number: A11A01629.The Pentra C200 system is a fully automated chemistry analyzer using colorimetry, turbidmetry and potentiometry technologies. It is mostly dedicated to in vitro diagnostic analyses based on homogeneous samples such as serum, plasma and uring. | PENTRA C200 | 2 | January 03, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA AST CP because during the AST application validation on the Pentra C200 using an Audit Systems linearity kit, it was determined that two higher values for AST (expected at 800 U/L and 1000 U/L) were recorded at 280 and 199 U/L without alarm. | DESIGN: Software Design | Horiba Medical sent a Field Safety Notification and Acknowledgement Letter dated November 20, 2012, to all affected customers. The letter notified the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 calendar days via mail, email or fax. Customers with questions were instructed to contact their local representative.For questions regarding this recall call 714-273-9254. |  | 384 | 384 units | Nationwide Distribution including CA, OK, NY, FL, AZ, NY, VA, and TX. | Terminated | September 12, 2013 | 252 | 67270 | Computer
Z-0645-2013 | 63829 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA ML, Software Versions: V8.0.x and Below; V.9.0.1 and V9.0.2 | ABX Pentra ML | 2 | January 07, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX PENTRA ML because they have confirmed an issue with the ABX PENTRA ML data management and validation system. An incorrect patient result is transmitted to the Laboratory Information System. This result is from a previous run of the same patient that is stored in the ML database. This issue occurs in a very rare sequence of specific events. The affected software | DESIGN: Software Design | .A recall letter dated 11/09/2012 was sent to all their customers who received the ABX PENTRA ML products. The letter informed the customers of the problem identified and the action to be taken. Customers are instructed to complete and return the enclose response form within 10 days. Customers with questions about the letter are instructed to contact their local HORIBA Medical representative.Recall Letter, as produced per FCD-0377, Field Safety Notice & Acknowledgement will be mailed via FedEx to the ten (10) customers noted on the list the week of November 12, 2012. Effectivity will be evaluated by receipt of the Field Safety Notice Acknowledgment Form. Follow-up communications with customers will be conducted, as needed. | Contact the recalling firm for information | 10 | 10 units | Nationwide in the US, including the states of CA, MI, IN, FL, and GA. Center Recall Depth is Retail. | Terminated | March 12, 2013 | 64 | 10 | Computer
Z-0674-2013 | 63817 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Scanning Systems: Brilliance 64 - Model #728231; and Ingenuity CT - Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance 64 and Ingenuity Computed Tomography (CT) | 2 | January 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Software anomalies. Philips Medical Systems recalled their Brilliance 64 running software v4.0.0xx516 and their Ingenuity CT running software v4.0.0xx518 due to software issues detected in these CT scanning systems. If these issues were to re-occur they could pose the risk of injury to patients undergoing scans; device users; or device service personnel. | DESIGN: Software Design | The firm, Philips Healthcare, sent an 'URGENT- Medical Device Correction" letters' dated November 15, 2012, and December 3, 2012, to their customers. Both letters describe the product, problem and actions to be taken. The letters state that Philips is notifying all affected customers of this information via a 'Field Safety Notice and will be installing a software update on the affected Brilliance 64 and Ingenuity CT system in order to address these issues. A Philips' Field Service Engineer will be contacting each customer site and make arrangements to install the necessary software update on these device systems. The customers should contact their local Philips representative or local Philips Healthcare office. In North America and Canada needing further information or support concerning this update, should contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5, and enter their site ID or follow the prompts. | Contact the recalling firm for information | 6 | A total of nine (9) units were distributed (3 units in the U.S & 6 units overseas) | Worldwide distribution: USA including IN, OH and OR; and the countries: Australia, Belgium, China, Israel, and Sweden. | Terminated | June 18, 2013 | 154 | 9 | Computer
Z-0675-2013 | 63393 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System Brilliance Workspace Portal, Model #728269 Philips Medical Systems, Highland Heights, OH.The Brilliance Workspace Portal is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. The portal (together with the scanner) provides image processing and display through software applications that process, analyze, display, quantify and interpret medical images/data. | Computed Tomography XRay System Brilliance Workspace Portal | 2 | January 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips was notified that when performing post processing with the Brilliance Workspace Portal platform with software versions 2.6.1.4, the coronary arteries in the rendered image are segmented automatically. However, the center line is not extracted or highlighted upon hovering and making it impossible to activate the vessel and choose the appropriate name (label) from the drop-down menu. | DESIGN: Software Design | Philips Healthcare sent a "CUSTOMER INFORMATION -URGENT-MEDICAL DEVICE CORRECTION" letter dated September 13, 2012. The letter identified the product, problem, and actions to be taken by the customers. Philips Service Engineers will contact customers for implementation of the software update. Contact your local Philips representative or local Philips office for further information concerning this notice. |  | 108 | 108 units | Worldwide Distribution-USA including the states of AZ, CA, CO, IL, IL, IN, MA, ME, MI, MO, NC, NJ, NY, OR, TN, TX, VA and WI and the countries of Austria, Canada, China, France, India, Ireland, Italy, Japan, Netherlands, Norway, Spain, Switzerland, Thailand and United Kingdom. | Terminated | February 14, 2014 | 394 | 108 | Computer
Z-0694-2013 | 63931 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Smart Disclosure System, Model 92810, is an integrated component of the Intesys Clinical Suite (ICS) G2, Version 4.03.The Smart Disclosure product collects patient waveforms, alarms, vital signs, and 12-lead reports. A 24-hour waveform acquisition is standard, with an upgrade option to 72 hours.The product provides basic trends in both tabular and graphical format, along with a retrospective ECG analysis. The analysis may be defined by patient, and the data may be presented in multiple formats, including a rhythm, a histogram, and a summary view.Standard reports are available from Smart Disclosure, including saved events, disclosure, trends, histogram, and summary.The Spacelabs Smart Disclosure System, Model 92810 is indicated for use in clinical situations where there is a need for review of physiological waveform information and alarm events up to 72 hours after the fact. Smart Disclosure is also indicated in those situations where a retrospective analysis of monitoring patients' ECG waveform data, that can be annotated and edited, is desired. The intended use of the Spacelabs Smart Disclosure is to interface with the Spacelabs monitoring network, providing the user with a means of recalling waveform information and performing retrospective analysis. The most recent 72 hours of monitored patient ECG waveform data can be analyzed, with each analysis limited to 24 hours or less. | Spacelabs Smart Disclosure System, Model 92810 | 2 | January 17, 2013 | 2013 | Spacelabs Healthcare, Llc | Displaying or printing of an incorrect waveform for the time indicated in the printout or display in the Spacelabs Healthcare Intesys Clinical Suite (ICS) G2, Version 4.03, Clinical Access portion of Smart Disclosure, Model 92810. | DESIGN: Software Design | Spacelabs has sent the Urgent -Medical Device Correction Spacelabs Healthcare ICS G2 Version 4.03 Model 92810 Smart Closure letter, dated 19 December 2012, to all U.S. customers. The firm plans to send a customer letter (translated as necessary) via e-mail to all international subsidiaries and distributors of record on December 27, 2012.The letter is addressed to Hospital Administrator/Biomedical Manager/Office Manager.Users are advised on the following actions:1) Advise your staff of this potential failure and advise them to always click the"Refresh" button when scrolling to new waveform data prior to printing or viewing waveform information or storing saved events.2) We will contact you at the earliest possible date to schedule a convenient time for your hospital to upgrade your affected software at no cost.Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. | Contact the recalling firm for information | 27 | 27 in US and 17 International | Worldwide Distribution-USA including Puerto Rico and the states of California, Florida, Iowa, Illinois, Louisiana, Michigan, Minnesota, North Carolina, South Carolina, Texas, and Washington, and the countries of Argentina, Mexico, France, and India. | Terminated | July 08, 2013 | 172 | 27 | Computer
Z-0696-2013 | 63617 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA 400 and 400 plus Analyzer, an in vitro diagnostic analyzer. | Roche COBAS INTEGRA 400 and 400 plus Analyzer | 3 | January 18, 2013 | 2013 | Roche Diagnostics Operations, Inc. | A software security issue with Oracles TNS-Listener component has been identified. TNS-Listener is supplied with the Roche COBAS INTEGRA 400/400 plus Analyzers. The Oracle TNS-Listener can be considered as an authentication and redirection component between the client and the database inside the Roche product. This means an attacker could register their own non-Roche database via remote access | DESIGN: Software Design | Roche sent an Urgent Medical Device Correction letter on November 19, 2012, to all affected customers via UPS Ground (receipt required). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached form and return by fax and file this Medical Device Correction for future reference. For technical support customers were instructed to contact Diagnostics Technical Support at 1-900-428-2336.For questions regarding this recall call 317-521-3911. |  | 845 | 845 Units | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR. | Terminated | January 13, 2014 | 360 | 845 | Computer
Z-0700-2013 | 63972 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GreenLight HPS (High Performance System)GreenLight XPS (Xcelerated Performance System)Manufactured by AMS, Solutions for LifeInnovation Center.GreenLight HPS and XPS Surgical Lasers Systems and accessories are intended for use in endoscopic 532 nm laser resection of the prostate for the treatment of benign prostatic hyperplasia. | GreenLight HPS (High Performance System); GreenLight XPS ( Xcelerated Performance System) | 2 | January 22, 2013 | 2013 | AMS Innovative Center - San Jose | During electrical recertification testing it was found that the Voltage Scaling Board was non-conforming to the IEC 60601-1-2 standard. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | AMS sent "NOTICE OF DEVICE FIELD CORRECTIVE ACTION" letters dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-408-456-3320 for questions about this notification. |  | 389 | 389 devices, all model numbers. | Worldwide Distribution. | Terminated | February 13, 2014 | 387 | 389 | Computer
Z-0742-2013 | 63800 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM5 Monitor, a vital signs monitor used on human patients in healthcare facilities. | Mindray DS USA Inc. | 2 | January 27, 2013 | 2013 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the DPM5 monitors where ST and Arrhythmia features may not be available. | DESIGN: Device Design | Mindray DS USA Inc. sent an Urgent Product Corrective Action letter on September 5, 2012 via Fed Ex with delivery confirmation to all users. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday 8:30-5:30 pm, to schedule a software upgrade. |  | 13 | 13 units | US Nationwide Distribution including the states of: VA, TN, FL, & TX | Terminated | October 22, 2013 | 268 | 13 | Computer
Z-0702-2013 | 63868 | x-ray controls - fluoroscopic | RCA |  |  |  |  | Juno DRF , X-ray system, diagnostic, Fluoroscopic, general-purposeModel: 70920Product Usage:Universal R/F, general use | Juno | 2 | January 28, 2013 | 2013 | Philips Healthcare Inc. | Values of fluoroscopy time, Air Kerma Rate andCumulated Air Kerma may not be displayed "atthe fluoroscopist's working position" as per 21 CFR 1003.21 | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Philips Medical issued a notification by letter dated December 13, 2012. via Certified Mail to affected customer. The letter identified the affected product, problem and actions to be taken. Philips will contact affected customers to schedule an upgrade of the equipment . For questions or support concerning this issue, please contact Dominic Siewko at 978-659-7936. |  | 16 | 16 units | US Nationwide Distribution including the states of AL, CA, CO, CT, IL, MD, MI, MO, NY, OH, PA | Open |  |  | 16 | Computer
Z-0754-2013 | 61611 | system, simulation, radiation therapy | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092-2 Version 7.4, 5160092-3 Version 7.5, 5160092-4 version 7.6. | AdvantageSim MD | 2 | January 30, 2013 | 2013 | GE Healthcare, LLC | It was reported by a customer site that when using GE AdvantageSim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | DESIGN: Software Design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 11,2012, to all consignees on May 24, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The letter addressed the Safety Issue, Safety Instructions, Affected Product,, Product Correction and Contact Information. GE indicated that a GE Healthcare service representative would contact they would correct all affected systems by providing a software upgrade at no cost to the consignee. For questions customers were instructed to call one of the following numers:USA 800-437-1171Japan 0120 - 055 - 919For questions regarding this recall call 262-513-4122. |  | 65 | 65 | Worldwide Distribution - USA including AL, CT, FL, IL, MA, MO, NJ, NY, PA, TX VA, WA, WI and Internationally to TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SOUTH AFRICA, NIGERIA, MALAYSIA, KOREA, JAPAN, INDIA, GERMANY, FRANCE, EGYPT, DENMARK, CHINA, CHILE, CANDA, BRAZIL, AUSTRIA, and GREAT BRITAIN. | Terminated | January 16, 2014 | 351 | 65 | Computer
Z-0750-2013 | 63875 | implant, hearing, active, middle ear, totally implanted | OAF |  |  | Ear Nose & Throat |  | Esteem Programmer Model 6001, US and ISA, Model 3001 Dell E54XX, part of the Esteem Totally Implantable Hearing System. The ISA and Commander are software packages that run on a laptop personal computer. | Esteem Programmers | 2 | January 30, 2013 | 2013 | Envoy Medical Corporation | Envoy Medical is conducting a voluntary correction of a limited number of Esteem Programmers, part of the Esteem Totally Implantable Hearing System, to reduce their susceptibility to noise interference. | DESIGN: Software Design | Envoy Medical Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-233-1204 for questions relating to this notice. |  | 44 | 44 | Worldwide Distribution-USA including the states of AL, AZ, CA, CT, FL, GA, IL, MA, MI, MN, NC, OK, RI, and TX, and the country of GERMANY. | Terminated | April 09, 2013 | 69 | 44 | Computer
Z-0762-2013 | 63997 | system, nucleic acid amplification, mycobacterium tuberculosis complex | MWA |  |  | Microbiology | PMA | COBAS TaqMan Mycobacteria tuberculosis (MTB) Test FOR IN VITRO DIAGNOSTIC USE. | Roche | 2 | January 31, 2013 | 2013 | Roche Molecular Systems, Inc. | The COBAS TaqMan MTB Test Definition Files (TDF) v3.0, for both CE-IVD and JPN-IVD tests, being used with AMPLILINK software version 3.3 series and COBAS TaqMan MTB Test Kits, both CE-IVD and JPN-IVD, have an incorrect Target Elbow Max cutoff value of 48 creating the potential for false negative results to be generated for patient samples with Target Elbows in the range of 48 to 50. |  | Roche issued Product Advisory Notice (PAN) 2012-11 and Urgent Field Safety Notice PAN 2012-11 by email to Local Safety Officers (LSO) and General/Country Managers from all countries on 28-Nov-2012. The notice identifed the affected product, problem and actions to be taken. LSOs are to provide the information in the Field Safety Notice (FSN) to each customer and return a Feedback Form to Roche (email address: rmdsafeb@roche.com) once actions in FSN are executed. |  | 383 | 335 CE-IVD TDFs; 48 JPN-IVD TDFs | Worldwide Distribution - Foreign (22 EU countries - Austria, Belgium, Bulgaria, Croatia, Cypress, Czech Republic, Denmark, Germany, Finland, France, Greenland, Hungary, Iceland, Italy, Luxembourg, Netherland, Norway, Sweden, Switzerland, Spain, Turkey, UK); (41 Rest of World countries - See Attachment for full country distribution + Japan, Portugal and Mexico) | Open |  |  | 383 | Computer
Z-0741-2013 | 64117 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-T1 Ventilator with software versions 1.1.2 or lower.The HAMILTON T-1 Ventilator is intended to provide positive pressure ventilatory support to adults and pediatrics | The HAMILTON T1 | 1 | January 31, 2013 | 2013 | Hamilton Medical Inc | Analysis of a customer complaint has shown that during ventilation of small pediatric patients with high airway resistance and low lung compliance, the oxygen consumption of a Hamilton T1 ventilator with software versions 1.1.2 and lower must be calculated using a larger margin than originally expected. | DESIGN: Software Design | Hamilton sent a Field Safety alert letter dated January 11, 2013 to all customers. The letter identified the affected product, explained the potential problem and a included Medical Device Safety Alert and Corrective Action document. Hamilton is working on a new software version for this issue. Customers are asked to update their operator manual with the update oxygen consumption formula page which provided revised information for calculating oxygen capacity. When the new software is available, Hamilton will provide loaner devices to each location while their unit is being upgraded.Questions should be addressed to Hamilton Medical at 800-426-6331 ext 215. |  | 97 | 97 total ventilators; 69 of part number 161005 and 28 of part number 161006 | US Nationwide Distribution including the country of Switzerland. | Open |  |  | 97 | Computer
Z-0767-2013 | 64127 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Phillips Easy Diagnost Eleva DRF. Model Number: 706034Product Usage:Universal Fluoroscopic Applications | Phillips Easy Diagnost Elva DRF | 2 | February 01, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | DESIGN: Software Design (Manufacturing Process) | Phillips issued an Urgent Field Safety Notice to all customers on January 14, 2013, to notify users that the Phillips EasyDiagnost Eleva systems, release 3.0 has a mirror icon that could be misinterpereted and could potentially lead to misdiagnosis. Philips will upgrade the system to version 3.1 which involves software and hardware upgrades of the system. This software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Additionally, Philips will provide updated Instructions for Use (IFU) for the Eleva workspot and the EasyDiagnost 3.1 system and an addendum to the EasyDiagnost system IFU. Users will be contacted by Phillips to schedule the installation. For questions contact your local Philips representative: 1-800-722-9377. |  | 14 | 14 units | US Nationwide Distribution including the states of FL, CA, WA, AZ, NC, KY, DE, AZ and MO. | Open |  |  | 14 | Computer
Z-0765-2013 | 64163 | file, pulp canal, endodontic | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Sybron Endo Lightspeed: LSX Files, 25MM, Part Numbers: 835-2525; LSX Files, 25MM Assorted Kit, Part Numbers: 835-2455; LSX Files, 30MM, Part Numbers: 835-2031. An engine-driven rotary endodontic instrument used during endodontic procedures for cleaning and shaping the root canal. | Lightspeed LSX Files | 2 | February 01, 2013 | 2013 | Ormco/Sybronendo | The length markings on the Lightspeed LSX files were wrong. | PRODUCTION CONTROLS: Process Control | Sybron Endo sent a Urgent: Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It has come to our attention that the length rings laser-etched on several lots of the Lightspeed LSX Files are inaccurate. We are asking for the product to be returned since use of the file could be a safety risk due to misadjustment of the file length. Please contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. |  | 740 | 740 units | Worldwide Distribution - USA and the country of Canada. | Terminated | November 07, 2013 | 279 | 740 | Computer
Z-0768-2013 | 64057 | generator, oxygen, portable | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Li-Ion battery used in the Oxus Portable Oxygen Concentrator (POC), RS-00400.Also branded as Reliability Plus byDrive. The Portable Oxygen Concentrator is intended to deliver concentrated oxygen for adult patients with chronic pulmonary diseases. | Reliability Plus by Drive | 2 | February 02, 2013 | 2013 | Oxus America, Inc. | Due to past battery inventory storage conditions and deficiencies in the maintenance procedure, it was possible that some batteries were over-discharged to point of damaging the cells. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Oxus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-475-1568 for information related to this recall. |  | 1724 | 1724 total units | Worldwide Distribution-USA including the states of New York, Michigan, Colorado, Texas, Iowa, Florida, Georgia, Ohio, Wisconsin, Oregon, Alabama, Illinois, Utah, Oklahoma, West Virginia and California and the countries of Canada, European Union, South Africa, Korea and Brazil. | Terminated | July 17, 2013 | 165 | 1724 | Computer
Z-0798-2013 | 64114 | system,network and communication,physiological monitors | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA, Inc., DPM Central Monitoring System with ambulatory Telepack. Part numbers: 300BF-PA00002, and 300BF-PA00003 (DPM Central Monitoring System) 115-006908-01, -03, -04 (DPM software upgrade kit). The DPM Central Monitoring System transfers information between DPM Central station and other networked devices. | DPM Central Monitoring System | 2 | February 11, 2013 | 2013 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified a software anomaly that may where, in the presence of an asystole, the Telepack ECG waveform may produce a unique step pattern (instead of a flat line) if Lead 1 or V is being utilized for detecting cardiac arrhythmia. | DESIGN: Software Design | Mindray North America/ Mindray DS USA Inc. sent an "URGENT PRODUCT CORRECTION ACTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your Mindray Service Representative at 1-800-288-2121 for questions relating to this issue. |  | 51 | 51 units | Nationwide Distribution. | Terminated | August 15, 2013 | 185 | 51 | Computer
Z-0802-2013 | 60695 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo CDI 101 Hematocrit/Oxygen Saturation Monitoring System. Monitors blood/oxygen saturation, hematocrit and hemoglobin levels. | CDI 101 Hematocrit/Oxygen Saturation Monitoring System | 2 | February 12, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Inaccurate readings. Software contains unused (turned off) accuracy correction factors for slope (gain) and offset (bias) for Hematocrit/Saturation calculations. The accuracy on any cuvette size did not meet the specifications. | DESIGN: Software Design | Terumo sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 1, 2010 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions or concerns. |  | 622 | 622 units | Worldwide Distribution. | Terminated | November 08, 2013 | 269 | 622 | Computer
Z-0813-2013 | 63654 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Physiological Patient Monitor and accessories. | Dash 3000, Dash 4000 and Dash 5000 Monitor and accessories. | 2 | February 13, 2013 | 2013 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor. These issues may occur when the Ethernet controller in the XPC823 CPU on the Dash and DashPort 2motherboard functions intermittently. GE Healthcare has identified that these issues occur most often at sites with large Unity Network installations. It should be noted that the Dash 3000/4000/5000 patient m |  | GE Healthcare sent an Important Product Information letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told the otherboard would be replaced, on request, should the customer continue to expeience the issue with the affected product. A second letter dated August 30, 2012, was sent on September 5, 2012, informing consignees that the motherboard replacement did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided when it is available. Customers with questions were instructed to call 1-800-558-7044.For questions regarding this recall call 262-513-4122. | No consumer action necessary | 119262 | 119,262 | Nationwide Distribution - ALL STATES IN CONTINENTAL USA INCLUDING PR and DC excluding AK, DE, ND, and VT.ALGERIA, AUSTRALIA, ANGOLA, ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM, URUGUAY.PUERTO RICO | Open |  |  | 119262 | Computer
Z-0822-2013 | 64111 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion TeleHealth Monitor Model Number V500; Product Codes 08832984 (US) and 02080468 (UK) - Physiological patient monitor (without arrhythmia detection or alarms). Viterion V100, V100-BGM and V500 Telehealth Monitors transfer medical information. | Viterion V500 Telehealth Monitors | 2 | February 14, 2013 | 2013 | Viterion TeleHealthcare Llc | Viterion telehealth monitors do not transmit "HI" (above the meter's measurable range) and "LO" (below the meter's measurable range) blood glucose readings from Bayer's Contour Elite XL blood glucose meters, or from Roche's Accu-Chek Advantage 03 and 04 meters to healthcare professionals. | DESIGN: Software Design | Bayer Healthcare sent "IMPORTANT RECALL INFORMATION" to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact Viterion Support at 1-800-866-0133 option 1 for questions relating to this issue. | No consumer action necessary | 369 | Domestic - 296 units, Foreign - 73 units | Worldwide Distribution-USA (nationwide) and the United Kingdom. | Open |  |  | 8334 | Computer
Z-0824-2013 | 64070 | system, measurement, blood-pressure, non-invasive | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare SureSigns VS2+Model 863278 SureSigns VS2+ NBP and SpO2Model 863279 SureSigns VS2+ NBP, SpO2 and TemperatureProduct Usage:The SureSigns VS2+ vital signs monitor is for monitoring, recording and alarming of multiple physiological parameters of adults, pediatrics, and neonates in healthcare environments. Additionally, the monitor is intended for use in transport situations within a healthcare facility. | Philips Healthcare SureSigns VS2 | 2 | February 14, 2013 | 2013 | Philips Healthcare Inc. | VS2+ screen may lock up and stop monitoring during patient use | DESIGN: Software Design | Philips issued an Urgent Medical Device Correction letter on January 14, 2013. The field action will take place in two phases: Phase 1 Philips sent customer an Urgent Medical Device Correction Notification/Field Safety Notice. This notification explains the issues and gives the customers instructions on actions to take while they await the hardware correction.Phase 2 Once the hardware correction Philips will contact customers and perform the correction free of charge. Until the correction is available.Affected products may continue to be used while waiting for the correction from Philips if the unit is not used in NBP interval mode.A Philips Healthcare representative will contact you to arrange for the correction of all affected devices. If you need any further information or support concerning this issue, please contact the Philips Customer Care Center at 1-800-722-9377.On 3/14/13, firm stated Initially the fix would be a hardware upgrade but have since changed that to a software update. Customers within the US will be sent a new Field Safety Notice as well as a USB software update package and will be required to notify us upon the upgrade of their affected units. They will be performing the updates themselves. | Contact the recalling firm for information | 3248 | 3248 units | Worldwide Distribution - US Nationwide and the countries of: AUSTRALIA, AUSTRIA, BELGIUM,CANADA, CHILE, DENMARK, FRANCEGERMANY, ICELAND, INDIA, IRELAND,ISRAEL,ITALY, JORDAN, KENYAKUWAIT, LEBANON, MALAYSIA, NAMIBIA, NETHERLANDS, NEW ZEALANDNORWAY,PALESTINE, POLAND,PORTUGAL, SAUDI ARABIA, SINGAPORESLOVENIA, SOUTH AFRICA, SOUTH KOREA,SPAIN,SWEDEN, SWITZERLANDTHAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESUNITED KINGDOM | Open |  |  | 3248 | Computer
Z-0814-2013 | 63662 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash Port 2 Docking Station.Product Usage: is a docking station (quick mount/dismount base) for a Dash patient monitor. The system is to monitor physiologic parameter data on adult, pediatric, and neonatal patients. | GE Healthcarr Dash Port 2 Docking Station | 2 | February 14, 2013 | 2013 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 and DashPort 2monitor. This report of correction encompasses the DashPort 2 device, in which the DashPort 2 stopscommunicating with the Dash, resulting in Loss of Monitoring at the CIC. The DashPort 2 status LEDsin the back of the unit indicate Unity Network and DashPort 2 to Dash communication has stopped. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent an Important Product Information Letter dated June 13, 2012 via FedEx to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter indicated that the motherboard would be replaced, on request, should the customer continue to experience the issue. A second letter dated August 30th 2012 was issued informing consignees that the correction that was initially proposed did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided at no charge when it is available. For questions contact Technical Support at 1-800-558-7044. |  | 10497 | 10,497 | Worldwide Distribution and US (nationwide) including DC and the countries of ALGERIA, ARAB EMIRATES, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, GERMANY, INDIA, ISRAEL, ITA LY, JAPAN, JORDAN, KUWAIT, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE SLOVENIA, SOUTH KOREA, SPAIN SWEDEN SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Open |  |  | 10497 | Computer
Z-0823-2013 | 64242 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650.The CARESCAPE Monitor B650 is a multi-parameter patient monitor intended for use in multiple areas and intra-hospital transport within a professional healthcare facility. The CARESCAPE Monitor B650 is intended for use on adult, pediatric, and neonatal patients and on one patient at a time. | GE Healthcare, CARESCAPE Monitor B650 | 2 | February 14, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of potential issues associated with the CARESCAPE Monitor B650. There have been no reported illnesses or injuries that have occurred with the use of the device. | DESIGN: Software Design | GE Healthcare sent an Important Product Notification letteron September 30, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to keep the patient under close observation or use an alternate monitoring device. Customers should not disconnect the affected USB devices from the patient monitor during an active patient case. If the license is enabled, the alarm limit adjustments from the iCentral Client has an L-NETM-00 license should be disabled. GE Healthcare would contact customers to schedule the update. Customers with questions should contact Technical Support or their local representative.For questions regarding this recall call 262-513-4122. |  | 2813 | 2813 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA,IL,IN, IA, KS, KY,LA, MD, MA, MI, M, MO, NE, NE,NJ, NM, NY, NC, OH, OK PA, SC, TN, TX, UT, VA, WI. Internationally to UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, ROMANIA, PORTUGAL, POLAND, PERU, NORWAY, NEW ZEALAND, NETHERLAND, MALAYSIA, LITHUANIA, KUWAIT, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, COLOMBIA, CHILE, CANADA, BELGIUM BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | January 27, 2014 | 347 | 2813 | Computer
Z-0836-2013 | 64307 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020, Animas(R) IR 1250 (1200 plus), Animas(R) IR 1200 Insulin Infusion PumpsIndicated for continuous subcutaneous infusion of insulin. | Animas(R) 2020, Animas(R) IR 1250 (1200 plus), Animas(R) IR 1200 Insulin Infusion Pumps | 3 | February 17, 2013 | 2013 | Animas Corporation | Due to an internal calendar date limit (December 31, 2015) these pumps will cease to operate on January 1, 2016, and will result in a call service alarm that is displayed on the pump display. |  | The firm initiated their recall of this product on December 17, 2012 by sending a letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Animas Customer Support Department at (866) 796-6373 for questions or concerns. |  | 37998 | 37,998 | Worldwide Distribution-USA (nationwide) including Washington, DC, Puerto Rico, US Virgin Island and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Austria, Australia, Bermuda, Canada, New Zealand, Czech Republic, Italy, France, Finland, Hungary, Israel, United Kingdom, Spain, Sweden, Norway, Denmark, Germany, Mexico and Ireland. | Terminated | December 24, 2013 | 310 | 37998 | Computer
Z-0841-2013 | 64328 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System.Radiation treatment planning. | Monaco RTP System | 2 | February 20, 2013 | 2013 | Elekta, Inc. | The beam is displayed at the "gantry" angle, but the dose is calculated at the "rotation_start" angle. |  | Elekta provided customers with an "Important Safety Notice - Dose Display Not Matching Beam Geometry - Monaco RFC 35748". The notice identified the product, problem, and actions to be taken by the customers. Contact the form at 800-878-4267 for questions relating to this notice. |  | 13 | 13 | Nationwide Distribution-including the states of AK, CA, FL, NJ, NC, OH, OR, TN, TX, WA, and WI. | Terminated | January 21, 2014 | 335 | 13 | Computer
Z-0867-2013 | 64319 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Eclipse Treatment Planning System for Radiotherapy; Model Number: H48. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | February 20, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | The calculation of dose dynamic plans containing different energy fields is incorrect when the Pencil Beam Convolution 11.0 algorithm (PBC 11.0) is used. | DESIGN: Software Design | Varian Medical Systems, Inc. Oncology Systems sent an Urgent Medical Device Correction Letter dated January 30, 2013 to all affected customers. The Letter identified the affected product, problem and actions to be taken. Varian Medical Systems is developing a technical correction for this issue, Its customers will be contacted by a Varian Customer Service representatives when this correction is available to schedule its installation. The letter provides customers with help desk contact information. |  | 1168 | 1168 units | Worldwide Distribution - USA Nationwide | Open |  |  | 1168 | Computer
Z-0871-2013 | 64201 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart XL+ Defibrillator/Monitor Model number: 861290The HeartStart XL+ is intended for use in a hospital setting by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced life support or defibrillation | Philips HeartStart XL Defibrillator/Monitor | 2 | February 21, 2013 | 2013 | Philips Healthcare Inc. | Philips HeartStart XL+ Defibrillator/Monitor may become locked out of clinical use | DESIGN: Software Design | Philips Healthcare sent an Urgent Medical Device Recall letter dated January 2013 on January 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if the Philips HeartStart XL+ Defibrillator/Monitor is first powered on and then there is disconnection or an interruption in AC power, the device may be locked out of clinical mode until the user enters the service mode password and performs an Operational Check. Users are provided with the Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Customers were advised to follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 3785 | 3785 units | Worldwide Distribution - USA (nationwide) and Internationally to ALGERIA, ANGOLA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BULGARIA, COLOMBIA, CROATIA (LOCAL NAME HYVATSKA), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MAURITIUS, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA (SLOVAK REPUBLIC), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VIET NAM. | Open |  |  | 3785 | Computer
Z-0877-2013 | 63957 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | QuietCare-Networked product, Facility Server/model: QC101601; The QuietCare system is intended to detect and analyze residents' movements. | Care Innovations QuietCareNetworked product | 2 | February 23, 2013 | 2013 | Intel-GE Care Innovations LLC | Care Innovations detected a software anomaly: During the engineering investigation into the cause of why a Wander Alert was not generated, a defect in the implementation of the Wander and Night Motion ADLs was discovered. | DESIGN: Software Design | Care Innovations sent a "DEVICE CORRECTION NOTIFICATION" letter dated 12/20/12 to all affected customers. The letter identified the product and problem. Correction was performed remotely by the firm. Contact the firm at 917-232-4775 for questions regarding this notice. | No consumer action necessary | 19 | 19 units | Nationwide Distribution. | Terminated | February 26, 2013 | 3 | 19 | Computer
Z-0761-2013 | 64226 | system, x-ray, film marking, radiographic | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Brivo XR385 DR-F/ X-Ray SystemX-Ray System | GE Healthcare Brivo XR385 DRF/ XRay System | 2 | February 26, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the Brivo XR385 DR-F/ Model 5271997 radiographic product related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the issue or bring the product intocompliance with each applicable Federal and IEC standard in accordance with a planto be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the software on the system to address the issue.Contact Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 9 | 9 units installed in US | Nationwide Distribution including NM, WY, FL, MO, MA, and MI. | Terminated | December 05, 2013 | 282 | 9 | Computer
Z-0884-2013 | 63674 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost X-Ray System with Eleva Software 2.1.3Model numbers 712020, 712022, and 712082Product Usage: Digital radiographic system | Philips Digital Diagnost | 2 | February 26, 2013 | 2013 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | DESIGN: Software Design (Manufacturing Process) | Philips Medical Systems issued on October 25, 2012 an Urgent Field Safety Notice for the Digital Diagnost System informing consignees of the software upgrade. The UFSN described the problem, the hazard involved, the action to be taken by the consumer/user and the action planned by Philips. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 551 | 551 | Worldwide Distribution - USA Nationwide and the Foreign countries of: Austria, Canada, Belgium, China,Denmark, France, germany, Italy, Netherlands,New Zealand, Norway, Philippines, Portugual, SOuth Korea, Spain, Switzerland, Thailand, and UK. | Open |  |  | 551 | Computer
Z-0786-2013 | 63392 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Mx8000 IDT Whole Body Computed Tomography X-ray system Distributed by Philips Medical Systems, Highland Heights, OH 44143The Brilliance Scanner are computed Tomography X-ray systems with multi-slice capability of up to 16 slices,which are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Mx8000 IDT Whole Body Computed Tomography Xray system | 2 | February 26, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The low contrast pins in the Image Quality (IQ) phantom deteriorates over time. This prevents the customers from doing the Daily Image Quality Checks as defined in the current Instructions for Use (IfU). | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Philips sent a Urgent Medical Device Recall Notice letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken.This issue does not effect the performance of the Scanner. The Customer Information is intended to inform you about:1. What the problem is and under what circumstances it can occur.2. The action you as a customer can take to minimize the effect of this problem.3 The action planned by Philips to correct this problem.If you need any further information or support concerning this issue please contact your local Philips Representative. In the US and Canada please contact the Philip's Healthcare customer care solutions at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer. In all other countries please contact your local Philips Healthcare office. | No consumer action necessary | 1474 | 1,474 Units | Worldwide Distribution - USA ( nationwide) including the states of AK, AL, AR, AZ, CA, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, WA, WV and WY, and the countries of Albania, Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, China, Colombia, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Ghana, Gibraltar, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Venezuela, Viet Nam & Yemen. | Open |  |  | 1474 | Computer
Z-0880-2013 | 63822 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore, Model #728243; and Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore Oncology, Model #728244. | Brilliance Big Bore and Brilliance Big Bore Oncology CT XRay Scanners | 2 | February 26, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Software anomaly. Philips determined that the TumorLOC software program for the referenced PET/CT X-Ray scanning systems has the capability to generate Intensity Projection data sets from respiratory gated data. A circumstance was identified by the firm whereby these generated datasets are 'flipped' and incorrectly labeled left-to-right when the are saved to disk. If the flipped Intensity Project | DESIGN: Software Design | Philips Medical issued an URGENT - Medical Device Correction letter to their customers concerning the Tumor LOC IPD issue on October 25, 2012. The letter informs the customer of the nature of this software problem and names the Brilliance and GEMINI scanning systems affected by this issue. The letter goes on to describe the hazard inherent to this software issue and provides 'Workaround procedures which the customers can use in order to help mitigate the effects of this problem. Customers with questions should contact their Customers Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 1-800-722-9377. |  | 765 | 765 units | Worldwide Distribution - USA including AK, AK, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, OA MA, MD MI, MN, MO, MS, NC, D, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV; Internationally to Canada and other countries.. | Open |  |  | 871 | Computer
Z-0885-2013 | 64213 | block, beam-shaping, radiation therapy | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Elekta, Apex add-on Micro MLC, collimator for radiation therapy. | Apex Addon microMLC | 2 | February 26, 2013 | 2013 | Elekta, Inc. | If the Apex collimator is incorrectly installed on the radiation head, a touchguard inhibit activates the view restrictions dialog box only shows the top 12 restrictions. | TRAINING: Employee Error | Elekta sent an Important Field Safety Notice Field Safety Notice FCO Ref: 20001407029 "Important information for the safe and satisfactory operation of Apex" to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obey all the applicable instructions, warning, and cautions in the user documentation for each piece of equipment.For questions regarding this recall call 770-300-9725. |  | 1 | 1 | Nationwide Distribution including TX | Open |  |  | 1 | Computer
Z-0703-2013 | 64134 | system, x-ray, film marking, radiographic | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Definium 5000 / 5220493 X-Ray System | GE Healthcare Definium 5000 | 2 | February 27, 2013 | 2013 | GE Healthcare, LLC | When quickly switching the protocols from AEC to Fixed under the same view, if the user selection process time is less than 0.5 seconds, the User Interface (UI) may display different exposure techniques than selected, and may not display AEC setting. The exposure can still be made as commanded, however it may not be properly indicated. | DESIGN: Software Design | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. A GE Healthcare Service Representative will update the software on the system to address the issue.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 312 | 312 installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 281 | 312 | Computer
Z-0897-2013 | 64403 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. Diagnostic imaging system | Ingenuity TF PET/CT | 2 | February 28, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The system may not default to the same gating trigger with which the data were acquired. This could cause a risk for patients. | DESIGN: Software Design | Philips Medical Systems (Cleveland), Inc. sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated January 21, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact your local Philips representative or Philips Healthcare office for further information concerning this issue. |  | 1 | 1 Unit | Distributed only in Ohio. | Terminated | July 19, 2013 | 141 | 1 | Computer
Z-0900-2013 | 63671 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Stationary radiographic system | Philips Digital Diagnost | 2 | March 01, 2013 | 2013 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | DESIGN: Software Design | Philips Healthcare issued an Urgent Field Safety Notice dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review the information contained in the Notice with all member of their staff who need to be aware of the contents. Also, customers were asked to retain a copy with the equipment Instruction for Use. For further information or questions customers were instructed to contact their local Phillips representative Technical Support Line at 1-800-722-9377.For questions regarding this recall call 978-659-4519. |  | 360 | 360 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Canada, Chile, Czech republic, Estonia,France, Germany, Hungary, Inida, Ireland, Italy, Japan, Libyan Arab Jamahari, Mexico, Netherlands, New Zealand, Norway , Oman, Panama, Peru, Poland, Qatar, Reunion, Russia, Saudi Arabia, South Africa, South Korea,Spain, Sweden, Swizerland, Thailand, UAE, and the UK. | Terminated | July 29, 2013 | 150 | 360 | Computer
Z-0901-2013 | 64071 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Volume Viewer Plus 5.x (Advantage Workstation 4.2) Volume Viewer 3.x (Advantage Workstation 4.1) Volume Analysis (Advantage Workstation 4.0) Volume Analysis (Advantage Workstation 3.1)Advantage Windows Volume Rendering Option is intended to provide fast, easy, volume visualization of three dimensional structures imaged with a computed tomography (CT) or magnetic resonance (MR) system. It is dedicated to clinical and diagnostic applications in radiology, surgery, and treatment planning. | GE Healthcare Volume Viewer Plus 5.x (Advantage Workstation 4.2) Volume Viewer 3.x (Advantage Wor | 2 | March 01, 2013 | 2013 | GE Healthcare, LLC | Incorrect views and measurements possible. GE Healthcare became aware of multiple potential safety issues involving the Advantage workstation, Volume Viewer Plus and Volume Analysis. | DESIGN: Software Design | GE Healthcare sent a "Urgent Medical Device Correction " letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the Edit Contour feature until this software functionality is corrected or replaced.GE Healthcare will correct all affected systems by providing a software upgrade at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 363 | 363 | Worldwide Distribution - USA including the states of AR, CA, CO, FL, GA, IL IN, IA, LA, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX and VA. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, ARGENTINA, BELARUS, CANADA, CHINA, CHILE, EGYPT CZECH REPUBLIC, FRANCE, FINLAND, GERMANY, GABON, HONG KONG, INDIA, IRELAND, ITALY KENYA, KOREA, JAPAN, MEXICO, NIGERIA, PHILIPPINES, PORTUGAL REPUBLIC OF MALTA, RUSSIAN FEDERATION SAUDI ARABIA, SPAIN SWITZERLAND, TURKEY, TAIWAN, TUNISIA, UNTIED KINGDOM and VENEZUELA. | Terminated | January 28, 2014 | 333 | 363 | Computer
Z-0907-2013 | 64214 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE syngo RT Therapist (RTT) version 4.3, ARTISTE Model, Part No. 08162815, Digital linear accelerators, MOSAIQ OIS system, Distributed by SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Manufactured by SIEMENS AG, MEDICAL SOLUTIONS, GERMANY. | ARTISTE" syngo RT Therapist (RRT) | 2 | March 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety issue when using the "In-session resumption" feature of the syngo RT Therapist (RTT) 4.3 software. The RTT will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | DESIGN: Software Design | The firm issued "Urgent Medical Device Correction" letters via mail to all its direct consignees, informing them of the affected product issue and providing instructions on the recall. The firm stated that the release of this UI TH005/13/S to distribute the Customer Safety Advisory Notice began on January 17, 2013. | Contact the recalling firm for information | 5 | 5 units | 5 devices, all located outside of the United States in Germany. | Terminated | March 05, 2013 | 0 | 5 | Computer
Z-0906-2013 | 64259 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Extended Brilliance Workspace (EBW), Computed Tomography X-Ray System. The Extended Brilliance Workspace (EBW) is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Philips Medical System Extended Brilliance Workspace (EBW), Computed Tomography | 2 | March 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Incorrect values may be displayed. Philips is issuing a software update (version 4.5.5) to the Extended Brilliance Workspace (EBW) systems currently running software version(s) 4.0.2, V4.5.2 or V4.5.3, addressing a number of non-conformances. | DESIGN: Software Design (Manufacturing Process) | Philips Healthcare sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.This Field Safety notification is to inform the actions that should be taken by the customer/user to prevent risks for patient or users, and the actions planned by Philips to correct the problem.If you need further information or supplies please contact your local Philips Healthcare office. For North American and Canada contact the Customer Care Solutions Center ( 1-800 722-9377) option 5 enter site or follow the prompts. |  | 2580 | 2,580 units | Worldwide Distribution - USA including the states of AL, AT, AZ, BAH, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NV, NY, OH, OR, PA, PR, PR, SC, TN, TX, VA, VT, WA, WI & WV. and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican, Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Greece, Guadeloupe, Guinea, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, LAO People's Dem, Rep., Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela & Viet Nam. | Terminated | December 26, 2013 | 296 | 2580 | Computer
Z-0913-2013 | 64000 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer. An automated chemistry analyzer that measures analytes such as Glucose, Magnesium, and Potassium. | AU5800 Clinical Chemistry Analyzer | 2 | March 06, 2013 | 2013 | Beckman Coulter Inc. | Due to a software data storage limitation which only occurs under certain conditions, measurements of the Reagent Blank (RB) or Calibration (CAL) sample results are not saved and the software reverts to a previously stored, acceptable RB/CAL result. | DESIGN: Software Design (Manufacturing Process) | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Call Center Hotline at (800) 854-3633 for questions regarding the notice. |  | 28 | 28 in U.S.; 4 Canada; 318 to other foreign countries. | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Brazil, China, Czech Republic, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea Republic, Malaysia, Netherlands, Russian Federation, Spain, Switzerland, Taiwan, Turkey, United Kingdom, and Viet Nam. | Terminated | February 24, 2014 | 355 | 28 | Computer
Z-0915-2013 | 64440 | sterilizer, steam | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers designed for sterilization of heat and moisture-stable materials used in healthcare facilities. | AMSCO C and AMSCO 400 Steam Sterilizers | 2 | March 06, 2013 | 2013 | Steris Corporation | The control boards alarm set points and temperature tolerances allowed for erroneous activation of the AD (analog/digital) alarm as temperature fluctuated during operation. When this alarm occurs, the sterilizer locks up and cannot be used until serviced, causing Customer nuisance. | DESIGN: Software Design (Manufacturing Process) | Steris Corporation initiated their recall on February 12, 2013. They will notify customers via onsite visits by STERIS Field Service Technicians to each affected Customer site.For further questions, please call (440) 392-7601. |  | 395 | 395 Units | Worldwide Distribution - USA including Canada and the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, SK, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | December 05, 2013 | 274 | 395 | Computer
Z-0926-2013 | 64256 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-arm Imaging SystemProduct Usage:image-intensified fluoroscopic x-ray system, mobile | Medtronic | 2 | March 08, 2013 | 2013 | Medtronic Navigation, Inc. | Potential failure of the braking system that controls the O-arm Imaging System gantry movement in the vertical direction | DESIGN: Component Design/Selection | Medtronic Navigation issued a Field Safety Notice letter dated January 18, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter instructed customers not to discontinue use and that a Medtronic Navigation Representative will be contacting them shortly to schedule a time to conduct the motion control box upgrade. For questions contact your local Medtronic Navigation representative or call Technical Services Group at 800-595-9709 or (720) 890-3200, or email at rs.navtechsupport@ medtronic.com. |  | 463 | 463 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, AUSTRIA, BELGIUM, BRAZIL, BULGARIA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK,FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN ,KOREA, REPUBLIC OF KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, NETHERLANDS, New Zealand, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, Taiwan(R.O.C.) , THAILAND, TURKEY, UNITED ARAB EMIRATES. and UNITED KINGDOM. | Open |  |  | 463 | Computer
Z-0939-2013 | 64447 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System (IACS) Monitoring Solution. Catalog Number MS20783 (Infinity MCable-Dual Hemo) Physiological patient monitoring system. | Draeger Infinity Acute Care System (IACS) Monitoring Solution | 2 | March 11, 2013 | 2013 | Draeger Medical, Inc. | The displayed value for the invasive blood pressure measurement of the IACS was different than the actual patient values. Liquid ingress in the cable connection may cause inaccurate measurements. This can happen if the Dual Hemo MCable is used in a non-vertical position, a vertical position but upside down, or laying it on the floor or on the patient bed. | DESIGN: Labeling Design | Drager sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-543-5047 for questions relating to this notice. |  | 260 | 260 | Distributed in the states of CA, IL, and MA. | Terminated | May 20, 2013 | 70 | 260 | Computer
Z-0935-2013 | 64291 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR450w, Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, Signa Vibrant, and GE 1.5T Signa HDe MR System. The GE Signa HDx MR system is a whole body magnetic resonance scanner. | Optima MR450w, Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, Signa Vibrant, an | 2 | March 11, 2013 | 2013 | GE Healthcare, LLC | An error on the SR Viewer Reporting Tool in which edits can be made on a report and can be saved into incorrect reports. | DESIGN: Software Design | GE Healthcare sent an Urgent Field Safety Notice dated May 5, 2011, to all affected customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to open only one patient report at a time when editing a report in SR Viewer.US Customers did not receive a customer letter, however all US systems have been updated.For questions customers were instructed to contact their local Sales/Service Representative.For questions regarding this recall call 262-513-4122. |  | 4026 | 4026 (1672 US; 2354 OUS) | Worldwide Distribution - USA (All States in continental USA including PR except RI.). Internationall to:ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BELARUS, BELGIUM, BULGARIA, BOSNIA & HERZEGOVINA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF IRAN, IRAQ, ITALY, IRELAND, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, KAZAKHSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MACEDONIA, MARTINIQUE, MOROCCO, NETHERLANDS, NORWAY, PAKISTAN, PERU, PHILIPPINES, REUNION, NEW ZEALAND, OMAN, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SINGAPORE , SLOVAKIA, SOUTH KOREA, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, URUGUAY, VIET NAAM, and YUGOSLAVIA. | Open |  |  | 4026 | Computer
Z-0945-2013 | 64401 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 6, 10, 16, 40, 64, Big Bore and the iCT, iCT SP, Ingenuity CT and MX8000 Dual v. EXP are Computed Tomography X-Ray Systems.These devices are Computed Tomography X-Ray Systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance 6, 10, 16, 40, 64, Big Bore and the iCT, iCT SP, Ingenuity CT and MX8000 Dual v. EXP | 2 | March 13, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field stating the patient table on a system had an unexpected downward movement while a patient was on the table. Philips has determined the cause was a malfunction of the vertical brake. | DESIGN: Component Design/Selection | Philips Healthcare sent a "Urgent Medical Device Correction" letter dated December 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed Philips received reports from the field stating that the screws holding the vertical brake together failed.A Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. If you need any further information please contact your local Philips representative or local Philips Healthcare office. For North America and Canada Contact the Customer Care Solutions Center at ( 1-800-722-9377, option 5: Enter Site ID or follow the prompts. | Contact the recalling firm for information | 4797 | 4797 Units | Worldwide Distribution - Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Open |  |  | 4797 | Computer
Z-0948-2013 | 64376 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Somatom CTs equipped with software syngo CT2010A, syngo CT2010B, syngo CT2011A and CT2012B.Computed tomography x-ray system. | Siemens Somatom CTs equipped with software syngo CT2010A, syngo CT2010B, syngo CT2011A and CT2012B | 2 | March 14, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Sporadic hang-up malfunctions concerning 3D reconstructions during activated "Preview Image" functionality. | DESIGN: Software Design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated November 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The correction for this problem will be available with Service Pack 1 for syngo CT2011A And Service Pack 4 for syngo CT2012B. Both Service pack are planned to be released starting February 2013. Systems still running on syngo CT2010A or syngo CT2010B are already scheduled to be upgraded to syngo CT2011A or syngo CT2012B respectively. Contact your local Siemens Application Specialist for additional information. |  | 896 | 896 | Nationwide Distribution. | Terminated | January 31, 2014 | 323 | 896 | Computer
Z-0949-2013 | 64580 | bed, flotation therapy, powered | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare SpO2RT2 bed. Product Usage: Bed, flotation therapy, powered bed may be used in a variety of settings including, but not limited to, acute care, including critical care, step down/progressive care, medical/surgical, high acuity sub-acute care, post anesthesia care unit (PACU), and sections of the emergency department (ED). The TotalCare SpO2RTÂŽ 2 Bed System is capable of being used with a broad patient population as determined appropriate by the caregiver or institution. | TotalCare SpO2RT 2 bed | 2 | March 14, 2013 | 2013 | Hill-Rom, Inc. | In certain situations a software problem with the TotalCare SpO2RTÂż 2 bed may occur that can lead to inadvertent loss of sufficient surface cushionpressure to adequately support the patient. This may increase the likelihoodof skin irritation for patients at higher risk of developing pressure-related skin conditions. A correction to the software program is currently being developed andtested, | DESIGN: Software Design | Hill-Rom sent an Urgent Field Safey Notice letter dated February 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter advised customers to follow steps provided until the updated software is available. Customers were requested to forward a copy of this letter to any other facility personnel you deem appropriate. Questions should be directed to Hill-Rom Technical Support at 800-445-3720. |  | 309 | 309 | USA Nationwide Distribution including the state of Canada | Open |  |  | 309 | Computer
Z-0950-2013 | 64436 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity Computed Tomography X-Ray SystemsThe Brilliance 64 and Ingenuity CT are whole body Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance 64 and Ingenuity Computed Tomography XRay Systems | 2 | March 14, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | This field change order is being released to update software and customer release notes to the affected installed base. | DESIGN: Software Design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 13, 2013, via certified mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter. Customers were instructed to retain a copy with the equipment Instruction for Use. For questions customers were instructed to contact their local Phillips representative or local Philips Healthcare office. For North Americal and Canada contact Customer Care Solutions Center at 1-800-722-9377, option 5: Enter Site ID or follow the promptsFor questions regarding this recall call 1-800-722-9377. | Contact the recalling firm for information | 42 | 42 Units | Worldwide Distribution - USA including AR, AZ, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, WY, and worldwide including the contries of Australia, Belgium, China, France, Germany, Israel, Japan, Norway, Sweeden, and Switzerland. | Open |  |  | 42 | Computer
Z-0951-2013 | 64320 | plasma, coagulation factor deficient | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Vial Label:1) LIATEST VWF:Ag Buffer - 5 ml Diagnostica Stago2) LIATEST VWF:Ag Latex + 4 ml R3 Diagnostica Stago3) LIATEST VWF:Ag Latex Diluent 4 mL Diagnostica StagoModel, catalogue, number 00518 - 58978 - 58979Kit Label:LIATEST VWF:AgDIAGNOSTICA STAGO S.A.S.92600 AsniÂŽres sur Seine (France)Coagulation Assay | Diagnostica Stago Inc. | 2 | March 14, 2013 | 2013 | Diagnostica Stago, Inc. | Diagnostica Stago has confirmed that sometimes the delta OD measured can be changed leading to a lowered percentage result. | DESIGN: Software Design (Manufacturing Process) | Diagnostica Stago Inc. sent e-mail notifications on 1/28/2013 to all consignees (laboratory supervisor). The written field safety notice, entitled Important Field Safety Notice, with its response form was attached. The letter identified the affected product and gave a description of the defect. Customers were asked to confirm all the results obtained in the abnormal range by either of the two methods provided. Users are to complete and return the enclosed response form as soon as possible to the Stago Hotline via fax, 973-644-9348. Questions should be directed to the Stago Hotline at 800-725-0607. |  | 2901 | 2901 kits | Worldwide Distribution - USA and Canada. | Terminated | July 19, 2013 | 127 | 2901 | Computer
Z-0965-2013 | 64546 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific/Cameron Health SQ-RX Model 1010 subcutaneous pulse generators. The S-ICD System is intended to provide therapy for the treatment of life-threatening ventricular tachyarrhythmias. | SQRX | 2 | March 18, 2013 | 2013 | Boston Scientific CRM Corp | An internal protective fuse can be unintentionally activated while the device is charging its capacitors for shock delivery or induction. Should this occur, the defibrillator would not be able to deliver therapy or communicate with the Q-TECH Model 2020 programmer, and would be unable to emit tones or otherwise respond to magnet application. | DESIGN: Software Design | Consignees were sent a Boston Scientific "Important Medical Device Information" letter dated March 1, 2012. The letter was addressed to "Dear Doctor". The letter described the product involved in the recall, the problem and provided Recommendations/Actions. US customers may contact Technical Services for more information at 800-227-3422. |  | 2090 | 2090 (570 USA; 1520 OUS) | Worldwide Distribution, USA including the states of AZ, CA, CO, DC, GA, IL, MD, MA, MN, NJ, NY, NC, OH, PA, SC, VA, WA and the countries of AUSTRIA, BELGIUM, CANADA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NEW ZEALAND, PORTUGAL, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, and the UNITED KINGDOM. | Terminated | May 23, 2013 | 66 | 2090 | Computer
Z-0968-2013 | 64535 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs BleaseSirius Anesthesia Workstation, BleaseFocus Anesthesia Workstation, and CAS I/II Absorber contained in Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation installed with the CAS I/II Absorber. Anesthesia workstation for use in the hospital environment and locations not requiring portability. | BleaseFocus Anesthesia Workstations | 1 | March 20, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | A defect in CAS I/II Absorbers in the Spacelabs Anesthesia Workstations and Service Kits may cause a condition leading to an increase in the carbon dioxide concentration within the inspired gas being delivered to the patient. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spacelabs sent an URGENT-MEDICAL DEVICE CORRECTION letter, dated 16 January 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers until a Spacelabs representative replaces the affected product: 1) Immediately cease the use of BleaseSirius and BleaseFocus Anesthesia Workstations that shipped after 31 October 2012. 2) Immediately cease use of Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation containing a CAS I/II Absorber from such Service Kit. 3) Advise your staff of this situation Customers can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical support. | Contact the recalling firm for information | 108 | 108 total (17 units in the US and 91 units outside the US) | Worldwide Distribution - USA Nationwide including the states of New Mexico, North Carolina and the countries of Argentina, China, Mexico, and Taiwan. | Terminated | May 14, 2013 | 55 | 108 | Computer
Z-0977-2013 | 64112 | hexokinase, glucose | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BS-200 Chemistry AnalyzerDesigned for clinical laboratory use. | BS200 Chemistry Analyzer | 2 | March 22, 2013 | 2013 | Mindray DS USA, Inc. d.b.a. Mindray North America | An issue has been identified with the BS-200 where clinical chemistry results may be inaccurate. | DESIGN: Device Design | Mindray North America /Mindray DS USA Inc. sent Urgent Product Corrective Action letters via certified mail, return receipt requested, signature required on 11/28/2012. The letters identified the affected product as well as the issue, effect on patients, and the corrective action. Customers are to remove the analyzer from service until the affected pump has been replaced. Also, customers are to contact Mindray Customer Service Representative Tony Li at 1-800-288-2121 x2112 for technical assistance, and to arrange for replacement pumps at no charge. | Contact the recalling firm for information | 3 | 3 units | Distributed in the states of NC and TX. | Terminated | June 12, 2013 | 82 | 3 | Computer
Z-0994-2013 | 64674 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i X-Ray System. A radiographic and fluoroscopic x-ray system for use in diagnostic and angiographic procedures. | Toshiba Infinixi XRay System | 2 | March 25, 2013 | 2013 | Toshiba American Medical Systems Inc | Lost and incorrect images. When using the "F-Rec" fluoroscopy image acquisition recording option in the "manual mode" an infrequent software error may cause images from another study to be displayed on the monitor. The acquired images may be lost requiring additional image acquisitions to be performed. | DESIGN: Software Design | In November, 2012, Toshiba America Medical Systems, Inc. sent an Urgent: Medical Device Correction letter to customers. The letter contains a return reply form that is to be faxed to TAMS for retention. | Contact the recalling firm for information | 95 | 95 units | USA nationwide. No international. | Terminated | March 26, 2013 | 1 | 95 | Computer
Z-0992-2013 | 64684 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis zeego systems with software version VC14, VC20 and VC21 in conjunction with specific technical configurationAngiographic x-ray system | Siemens Axiom Artis zeego systems | 2 | March 25, 2013 | 2013 | Siemens Medical Solutions USA, Inc | In the course of the firm's product monitoring activities, a potential risk for patient or operators was determined during the operation of an Artis zeego system with software revision VC14, VC20 and VC21 in conjunction with a specific technical configuration which cannot be completely executed. |  | Siemens sent an Urgent Customer Information letter dated November 8, 2012 to affected consignees. The letter notified consignees of the issue and informed them the update AX039/12/S would be started and affected systems would be checked and corrected if required. Customers were instructed that they would be contacted to arrange a date for checking their system.For questions regarding this recall call 610-219-6300. |  | 52 | 52 | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, LA, MD, MI, MO, NC, NH, OR, PA, SD, TN, TX, WA and WI and Puerto Rico. | Terminated | December 24, 2013 | 274 | 52 | Computer
Z-0997-2013 | 64662 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Systems with Software Artis VC20x/VC21A/VD10x that have a DSA license installed.Angiographic x-ray system | Artis Systems with Software Artis VC20x/VC21A/VD10x | 2 | March 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm has become aware of a potential issue on Artis systems with the software Artis VC20x/VC21A/VD10x that have a DSA license installed. Under certain preconditions, executing roadmap OGP (organ program) in the DSA overlay mode may lead to imprecise registration of the DSA mask image with the roadmap image displayed on the live monitor. | DESIGN: Software Design (Manufacturing Process) | Siemens sent a Customer Safety Advisory Notice on March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information on to all those who need to be aware of it and to instruct their personnel accordingly. Customers were asked to forward the safety information to other organizations that couald be affected by this action.For questions regarding this recall call 610-219-6300. |  | 81 | 81 | Nationwide Distribution including AL, CA, CO, FL, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, PA, SC, TX, UT, VA, WA, and WI and to Puerto Rico. | Open |  |  | 81 | Computer
Z-1002-2013 | 64461 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3The software is a medical diagnostic system that allows the processing, review, analysis, communication, and media interchange of multidimensional digital images acquired from a variety of imaging devices. It provides the ability to review digital images from original DICOM data for multiple modalities, including CT, MR, D, CR, XA, US, NM, PET, etc. The software is not meant for primary image interpretation in mammography. | Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to ve | 2 | March 27, 2013 | 2013 | Vital Images, Inc. | Vital Images has found a potential error in measurement generated through the software which affects all versions of Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3.This error occurs when images are rotated at the time of reconstruction by the scanner in non-90 degree increments of the transverse/axia | DESIGN: Software Design | Vital Images sent an "Urgent Field Safety Notice' letter dated February 5, 2013, to all affected customers. The letter identified the probelm the product and the action needed to be taken by the customer. Please immediately inform all potential users of Vitrea of the information contained in this letter. We recommend that all measurements associated with the error described in this letter be verified against results from other technologies.You will be contacted by Vital Images Customer Support or by the customer support department of your distributor to arrange for installation of this patch on your system. We recommend that you implement the patch. If you have an earlier version, Vital Images Customer Support or your distributor will contact you to discuss your options for replacement.Please sign this Notice in the space below to verify that you have received and understood it, and return your signed Notice by email to fieldnotices@vitalimages.com.If you have questions, direct customers of Vital Images should contact Vital Images Customer Support at 1-800-208-3005 or by e-mail to support@vitalimages.com; customers through a distributor of Vital Images products should contact that distributor.Updated 5/28/2013:A second consignee Vital Images, Inc "Urgent Field safety Notice" letter dated March 26, 2013 was sent to consignees on 3/25/13. The follow-up letter was sent via e-mail and mail to 4,616 already-notified consginees, who had received affected versions 6.3.0, 6.2.3, or 6.1.6. The letter described the problem and the affected product being recalled and provided a number of steps to follow which included to retun the Effectiveness Check form. | Contact the recalling firm for information | 7649 | 7649 (Expnaded 5/28/13 to 7654) | Worldwide Distribution - All states in continental USA including DC, PR and the countries of ALGERIA, ARGENTINA, ANGOLA, ARMENIA, AUSTRALIA, AUSTRIA, BENGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL BULAGARIA, CAMBODIA, CANADA, CENTRAL AFRICAN REPUBLIC, CHAD, CHILE, CHINA , COLOMBIA, CONGO, CROATIA, CYRPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, JORDAN, KAZAKHASTAN, KOREA, LEBANON, LITHUANIA, MACEDONIA, MALYASIA, MALI, MAURITIUS, MEXICO, MALDOVA, MOSCOW, NYNMAR, NAMBIA, NAMIRGIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGER, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILLIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, RUSSIAN FEDERATION, SAUDI ARABIA, SENEGAL, SERBIA, SEBIA AND MONROBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOMALIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBERKISTAN, VENEZUELA, VIETNAM and YEMEN. | Open |  |  | 7649 | Computer
Z-1014-2013 | 64543 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical, BCIÂŽ Remote Alarm Cable (Normally Open & Normally Closed), Reorder Numbers: WW3398NO & WW3398NC. For use with respiration carbon dioxide measurement equipment. Product Usage: The BCIÂŽ Remote Alarm Cable is designed for use with the 9004 CapnocheckÂŽ Capnography (Plus or Sleep) system to transfer alarm signals from the CapnocheckÂŽ Capnography system to a remote alarm system. The BCIÂŽ Remote Alarm Cable and the 9004 CapnocheckÂŽ Capnography systems are intended for use in all critical environments including ventilatory applications, patient transport, and anesthesia. | BCI Remote Alarm Cable Normally Open /Normally Closed | 1 | March 29, 2013 | 2013 | Smiths Medical ASD, Inc. | In rare instances, the BCIÂż Remote Alarm Cables (BCIÂż Cable) are not transferring alarms when used with some remote nurse alarm systems. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Smiths Medical sent an "Urgent Medical Device Field Safety Notice-Recall" letter dated February 96, 2013 via UPS to all affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to inspect their inventory and remove all affected devices from use, complete and return the attached "Confirmation Form", by fax to 1-800-237-8033 or email to BCIRemote.AlarmCable@smiths-medical.com. The letter stated that a Customer representative will contact customers with instructions for returning affected products for replacement or credit. For questions contact Smiths Medical's Customer Service Department at 10-800-258-5361. | Contact the recalling firm for information | 39 | 39 | USA Nationwide Distribution - including the states of: CO, IL, IN, MA, SC and WI. | Terminated | July 12, 2013 | 105 | 39 | Computer
Z-1024-2013 | 64448 | pump, infusion, insulin bolus | OPP | Infusion pump. | General Hospital | General Hospital | 510(k) | Z-800 Volumetric Infusion System.For intravenous infusion of parenteral fluids, blood and blood products to a human patient. | Zyno Z800 Volumetric Infusion System | 2 | March 29, 2013 | 2013 | Zyno Medical LLC | Update Instruction For Use for Preventative Maintenance (PM) Schedule and addition of Odometer software to monitor accumulative fluid volume infused by the pump |  | Zyno Medical initiated recall 6/13/11 to accounts advising of a Preventative Maintenance (PM) Schedule and a software upgrade Odometer to monitor counts of accumulative fluid volume infused by the unit. On 3/1/13, firm issued correction notice to users that have not had the software implemented on their unit. Units will be serviced at their scheduled PM. The notice identified the product, problem, and actions to be taken by the customers. |  | 2589 | 2589 units | Nationwide Distribution. | Open |  |  | 2589 | Computer
Z-1029-2013 | 64393 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT.Diagnostic imaging system that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. | Ingenuity TF PET/CT | 2 | April 01, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Issues have been detected in the Ingenuity TF PET/CT currently running software version 4.0.60605.0 that if they were to re-occur, could cause a risk for patients, users or service personnel, including inaccurate standard uptake values (SUV). Philips Medical Systems (Cleveland), Inc. is issuing a software update to address the issues. | DESIGN: Software Design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated December 19, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. A Philips Service Engineer will contact the customer for implementation of the software update on the affected system. | Contact the recalling firm for information | 1 | 1 | Distributed to one customer in Ohio. | Terminated | June 18, 2013 | 78 | 1 | Computer
Z-1046-2013 | 64570 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) EXL(TM) LM (Siemens Material Number 10486890).Intended to measure a variety of analytes in human body fluids. | Siemens Dimension(R) EXL(TM) LM | 3 | April 01, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Through investigation of customer complaints, Siemens Healthcare Diagnostics has confirmed an increase in System Check failures due to high System Check LOCI results. Upon further investigation, Siemens has determined that many of these errors are due to the current limit set for System Check LOCI results. The errors have no impact on results, but they may cause customer inconvenience. Siemens h |  | Siemens initiated their recall of this product on February 27, 2013 by sending an Important Customer Notification letter to consignees by FedEx. The letter identified the product, problem, and actions to be taken by the customers. |  | 2207 | 2207 total | Worldwide Distribution. | Terminated | January 14, 2014 | 288 | 2207 | Computer
Z-1031-2013 | 64497 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Medtronic Prostiva RF Therapy Generator.The Prostiva RF system is indicated for the treatment of symptoms due to urinary outflow obstruction secondary to benign prostatic hyperplasia (BPH) in men over the age of 50 with prostate sizes between 20 and 50 cm3. | Prostiva | 2 | April 02, 2013 | 2013 | Medtronic Neuromodulation | Medtronic has found through device testing that if the optional footswitch is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitch connector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equi | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic sent a "Urgent Medical Device Correction" letter dated January 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Medtronic has found through device testing that if the optional footswitch (Model 60883) is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitchconnector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equipment. The status messages on the display screen and associated audible alarms operate as intended, and would notify the clinician of any change in status as described in chapter 2 of the system manual.Enclosed you will find an insulating cap that should be placed over the footswitch connector port when the Model 8930 RF Generator is used without the footswitch. The footswitch connector port is located on the rearpanel of the Model 8930 RF Generator as shown below.For further questions In the United States, contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 or your local product representative. Outside of the United States, contact your local product representative. |  | 848 | 848 (696 US, 152 OUS) | Nationwide Distribution including Puerto Rico and all States in continental USA except ID, MT, and NH. | Terminated | November 21, 2013 | 233 | 848 | Computer
Z-1044-2013 | 64675 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model:Primeview 3i 2.0, Part No. 8139847, Current R630;Product Usage:The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | April 02, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device. | DESIGN: Software Design (Manufacturing Process) | Siemens Medical Solutions sent a Customer Information letter via mail to all affected customers. The letter identified the affected product, problem and provide instructions on the correction. Customers were instructed to include the Customer Information letter in the Owner Manual chapter "Safety Advisory Letters".9-16-13 UPDATE: - Control Console SW 9.2.18 was released to support the rol-out10/18/2013; Distribution of update instructions for TH017/13S and TH018/13S - Distribtuion of the Control Console SW 9.2.18 released. |  | 621 | 621 units (foreign account) in total | Worldwide Distribution | Open |  |  | 621 | Computer
Z-1050-2013 | 64386 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Micros IM2, Part# M60CSB110EN04 and D00A00202The IM2 Data Management System consists of the IM2 Data Manager Software, Desktop Computer, keyboard, mouse, and a unidirectional serial cable. The IM2 software is used to store, retrieve, and process laboratory data received via a unidirectional serial connection from the ABX Micros 60. | ABX Micros IM2 | 3 | April 03, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX Micros IM2 Data Management System because it may display incorrect patient test results under unanticipated operating conditions. | DESIGN: Device Design | Horiba Medical sent an Urgent Safety Notice dated February 15, 2013, to all affected customers. The letter informed the customers that Horiba Medical is informing all ABX Micros IM2 Data Management System customers of their intention to remove the ABX Micros IM2 product from the market. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to complete and return the enclosed response form within 10 days.For questions regarding this recall call 949-453-0500, ext 208. |  | 180 | 180 units | Nationwide Distribution | Open |  |  | 180 | Computer
Z-1051-2013 | 64189 | incubator, neonatal | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂŽ and Giraffe Incubator p/n 6600-0723-700 (rev A and B) labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. GE Healthcare Giraffe Omnibed is a combination of an infant incubator and an infant warmer. | Giraffe Omnibed and Incubator | 1 | April 05, 2013 | 2013 | GE Healthcare, LLC | Neonatal beds and incubator display boards may randomly overwrite stored information and radiant heater and oxygen set points upon power up or recovery from a power failure. | PRODUCTION CONTROLS: Process Control | GE Healthcare sent an "Urgent Medical Device Correction" letter, dated January 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were given the safety issue and instructions for continued use of the Giraffe OmniBed and Incubators, as well as given the affected product details. GE Healthcare will correct all affected Giraffe Incubators and Giraffe OmniBeds at no cost, and will contact customers to arrange for this correction. If customers have any questions or concerns regarding this notification, please call: United States: 1-800-345-2700. For other countries, please contact your local GE Healthcare Service Representative. | Contact the recalling firm for information | 679 | 679 units | Worldwide Distribution -- USA in the states of AR ,CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, MA, MD, MI, MO, MS, NC, NE, NJ, NY, NV, OH, PA, TN, TX, VA, WA, WI, and WV, and the countries of Arab Emirates, Albania, Australia, Belgium, Brazil, Canada, Chile, China, Croatia, Czech., Denmark, Ecuador, Estonia, Finland, France, Germany, Hungary, Italy ,Korea, Kurdistan, Lebanon, Macedonia, Malaysia, Mexico, Norway, Oman, Poland, Russia, San Salvador, Saudi Arabia, Singapore, South Africa, Sweden, Turkey, Uruguay, Venezuela, and United Kingdom. | Terminated | July 10, 2013 | 96 | 679 | Computer
Z-1070-2013 | 64656 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting Cardiac Catheterization module and IMPAX CV Outbound Reporting.Facilitates digital reports for catheterization, echocardiography, and congenital heart disease. | IMPAX CV Reporting Cardiac Catheterization module and IMPAX CV Outbound Reporting | 2 | April 08, 2013 | 2013 | AGFA Corp. | Baseline pulmonary capillary wedge (PCW) pressure values from IMPAX CV Reporting Cardiac Catheterization module were not correctly transferring into the hospital's electronic health record (EHR) system. | DESIGN: Software Design | AGFA sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013 to the consignees who use the IMPAX CV Reporting tool described for this issue. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 877-777-2432, referencing HQ_10020006 for questions regarding this notice. |  | 26 | 26 | Nationwide Distribution including the states of AK, GA, IL, KY, MA, NY, NC, RI, SC, TX, and VA. | Terminated | January 09, 2014 | 276 | 26 | Computer
Z-1067-2013 | 64667 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-IOI: AQUILION 32/64; System, X-Ray, Tomography, Computed.Designed to produce cross-sectional images of a human body. | TSXIOI: AQUILION 32/64; System | 2 | April 08, 2013 | 2013 | Toshiba American Medical Systems Inc | The firm initiated this recall due to a potential software issue. The DLP (Dose Length Product) value may be displayed incorrectly when the vHP option is used. | DESIGN: Process Design | Toshiba sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter contained a return reply form that is to be faxed to TAMS for retention. Contact the firm at (800) 421-1968 for questions relating to this letter. |  | 23 | 23 units | Nationwide Distribution including the states of WA, OH, CA, TX, WV, IL, AZ, CA, WI, ME, WI, NC, FL, MT, and HI. | Terminated | June 17, 2013 | 70 | 23 | Computer
Z-1069-2013 | 64654 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV 7.8 SU3 - OCR Service.The IMPAX CV Reporting component facilitates the quick , effective creation of digital structured reports for adult catheterization and echocardiography, pediatric echocardiography, congenital heart disease, nuclear cardiology and non-invasive vascular disease management. | IMPAX CV 7.8 SU3 OCR Service | 2 | April 08, 2013 | 2013 | AGFA Corp. | Software design error in IMPAX CV 7.8 SU3 | DESIGN: Software Design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to distribute this information within your facility to all individuals who need to be aware of this. Please fax back to (864) 421-1664 that the information was received and understood was requested from the customers.If you have any questions about this matter please contact your local AGFA Healthcare organization, or call (864) 421-1754. |  | 2 | 2 | Distribute in the state of FL and KY | Open |  |  | 2 | Computer
Z-1068-2013 | 64673 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-30IC AQUlLlON ONE System, Computed Tomography X-Ray. This system supports Whole body scanning. | TSX30IC: AQUlLlON ONE; System, XRay, Tomography, Computed. | 2 | April 08, 2013 | 2013 | Toshiba American Medical Systems Inc | Due to a problem with the control software of the x-ray high-voltage generator, if scanning is performed with a tube current less than 30mA, the control software may detect an error and perform error processing even though the x-ray output is normal. | DESIGN: Process Design | Toshiba American Medical Systems Inc. sent a Urgent Medical Device Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer was informed a New software to prevent the occurrence of this issue will be installed on your system. Toshiba Service Representative will contact you for an appointment to install the New Software on your Aquilion One System.Please fax the return reply form to (877) 349-3054.We thank you for your urgent attention to this matter. If you have any further questions regarding this letter please call (800) 521-1968. |  | 1 | 1 unit | Distributed in the state of MD. | Terminated | April 08, 2013 | 0 | 1 | Computer
Z-1077-2013 | 63974 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ELEKTA MOSAIQ SW version below 2.41Product Usage: The intended use of the SIEMENS branded ARTISTE and ONCOR family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | ELEKTA MOSAIQ SW | 2 | April 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A potential safety risk with the MOSAIQ OncologyInformation System that has been identified by the manufacturer IMPAC Medical Systems, Inc. | DESIGN: Software Design (Manufacturing Process) | Siemens sent an Urgent Medical Device Correction letter with update instructions to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your responsible Siemens Service Engineer or Application Specialist. |  | 101 | 101 | Worldwide Distribution | Terminated | July 19, 2013 | 100 | 101 | Computer
Z-1081-2013 | 64855 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON SC2000 Ultrasound System, Model 10433816, hardware that contains version 4 of the MPI board 10434200;The SC2000 ultrasound imaging system is intended for the following applications: Cardiac, Neo-natal and Fetal Cardiac, Pediatric, Transespohageal, Adult Cephalic, Peripheral Vessel, Abdominal, Abdominal Intraoperative, Intraoperative Neurological, Musculo-skeletal Conventional, and Musculo-skeletal Superficial applications. The system also provides the ability to measure anatomical structures and calculation packages that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | ACUSON SC2000 | 2 | April 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | When using the V5M transducer, the display of the temperature on the ACUSON SC2000 system can be lower than the actual temperature of the V5M transducer. This occurs only if the system hardware contains revision 4 of the MPI Board. There is a potential to cause esophageal burns in patients. | DESIGN: Device Design | Siemens Medical Solutions USA Inc. sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A return response form is included for customers to complete and mail to the recalling establishment. Siemens stated additional follow up action will occur, once material is available, where Siemens field service engineers will visit each unit and replace the affected MPI 4 board.For further questions please call (650) 694-5747. |  | 120 | 120 units | Worldwide Distribution - USA and countries of Australia, Austria, Bosnia Herzegovina, Belgium, Finland, France, Germany, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, P.R. China, Portugal, Rep.of Korea, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan and United Kingdom. | Terminated | February 14, 2014 | 310 | 120 | Computer
Z-1100-2013 | 64763 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w 3.0T System. Part Number: 5352293.Product Usage: It is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images, spectroscopic images, parametric maps, and/or spectra,dynamic images of the structures and/or functions of the entire body, including, but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate,blood vessels, and musculoskeletal regions of the body. | Discovery MR750w 3.0T System | 2 | April 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to image flip in frequency encoding direction associated with a specific workflow of your Discovery MR750w system.Image flip issue in frequency encoding direction (including S/I, A/P and R/L) can result in incorrect localization ofanatomy and pathological findings. | DESIGN: Software Design | GE Healthcare sent an Important Product Information letter dated April 23, 2012. The letter described the Issue, Instructions, Affected Product Details and Product Correction. For questions contact your local service representative. |  | 46 | 46 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, IL, MD, NJ, NY, OH, OK, TX. and the countries of Australia, Belarus, Canada,France, Germany, India, Japan, Lebanon, New Zealand, Saudi Arabia, Singapore, Spain, Sweden, Switzerland and United Kingdom. | Open |  |  | 46 | Computer
Z-1096-2013 | 64725 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare LOGIQ S8 (with software revisions R1.1.1 and R1.5.1) and LOGIQ S7 (with software revisions R1.0.1, R1.0.2 and R1.0.3) diagnostic ultrasound system with the S4-10-D Ultrasound Probe. The device is intended for use in ultrasound evaluation. | LOGIQ S8 and LOGIQ S7 Ultrasound System | 2 | April 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to the S4-10-D Ultrasound Probe used in conjunction with the LOGIQ S8 and LOGIQ S7 Ultrasound systems. It was also discovered that this issue also leads to a specification issue of the acoustic output of this probe. There is a potential for a skin burn when the flow model in the cardiac application is activated with the S4-10-D Ultraso | DESIGN: Software Design | Consignees were sent on 2/8/2013 a GE Healthcare " Urgent Medical Device Correction" letter dated February 8, 2013. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Managers of General/Diagnostic Imaging. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and Contact Information. US customer contact number is (1)800-437-1171. |  | 2061 | 2061 | Worldwide Distribution, USA, including the states of AL, AZ, AR, CA, CO,K CT, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT NE, NV N NJ, NM, NY, NC, ND, OH, OK, PA, RI ,SC, SD, TN, TX, VA, WA, WI and the countries of AUSTRALIA, AUSTRIA, BRAZIL, CANDA, CHILE, CHINA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY INDIA, INDONESIA ITALY JAPAN, KOREA, LATVIA, LEBANON, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA , SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM. | Terminated | November 21, 2013 | 224 | 2061 | Computer
Z-1101-2013 | 64670 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MRT-1504/A4 - Titan X MRI SystemMRT-30 I0A/5 - Titan 3T MRI SystemProduct Usage: MRI System. | MRI System | 2 | April 11, 2013 | 2013 | Toshiba American Medical Systems Inc | It was found that there is a problem with the LCD monitor used in the console. The LCD monitor maybecome blank during use. Specifically, the backlight of the LCD monitor may turn off while the systemis in use, causing the LCD monitor to become blank. However, the other parts of the system continue tooperate normally, and the risk of a serious health hazard is therefore very low. The LCD Moni | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Toshiba sent an Urgent Medical Device Correction letter dated October 1, 2012 via USPS return receipt mail to all affected customers . The letter identified the affected product, problem and actions to be taken. The letter contains a return reply form that is to be faxed to TAMS for retention. For questions contact your local Toshiba representative at 800-521-1968. |  | 27 | 27 units | USA Nationwide Distribution including the states of : IL, NH, MA, MI, MO, PR, NV, NM, MN, ID, NY, OH, NC, AZ, MS, LA, OR, PA, CA, and CT. | Terminated | November 06, 2013 | 209 | 27 | Computer
Z-1104-2013 | 64693 | transducer, ultrasonic, diagnostic | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Aplio 500/400/300 Diagnostic Ultrasound System (TUS-500/400/300).Device Listing Number: Dl53118.Diagnostic Ultrasound system. | Aplio 500/400/300 Diagnostic Ultrasound System | 2 | April 12, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba Medical Systems has received reports of startup issues with the Toshiba TUS-500 andTUS-300 ultrasound systems. The startup issue is predominant when attempting to bring thesystem up from standby mode when the system is used remotely. It has been determined that thismay potentially cause damage to the system. | DESIGN: Device Design | Toshiba sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated December 11, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (800) 421-1968 for questions regarding this letter. |  | 233 | 233 units | Nationwide Distribution. | Open |  |  | 233 | Computer
Z-1074-2013 | 64624 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | All OneTouch Verio IQ Blood Glucose Meters sold as:Verio IQ System Kits - Meter;Verio IQ Starter kits - Meter;Verio IQ Warranty Meter.Product Usage: The OneTouchÂŽ Verio IQ Blood Glucose Monitoring System is intended to be used for the quantitative measurement of glucose (sugar) in fresh capillary whole blood samples drawn from the fingertips. The system is intended to be used by a single patient and should not be used for testing multiple patients. The OneTouchÂŽ Verio IQ Blood Glucose Monitoring System is intended for self-testing outside the body (in vitro diagnostic use) by people with diabetes at home as an aid to monitor the effectiveness of diabetes control. | OneTouch Verio IQ Blood Glucose Monitoring System | 1 | April 12, 2013 | 2013 | Lifescan Inc | The Verio IQ meter will shut off and revert to set up mode at glucose values above 1023 instead of displaying EXTREME HIGH GLUCOSE. | DESIGN: Software Design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated March 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to identify and hold all affected product in inventory and discontinue distributing. The letter states that a Lifescan sales representative will contact customers to arrange for pick up and replacement. |  | 710413 | US 710,413 | USA Nationwide Distribution | Open |  |  | 710413 | Computer
Z-1105-2013 | 64737 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo WorkFlow SLR, Radiological Information System, Version VA31A SP1/SP2/SP3 The information system syngo(R) Workflow SLR is a digital radiology information system (RIS) with integrated modules for patient administration, examination, reporting, statistics and system administration, as well as multiple interfaces for patient registration, order entry, and external billing modules. It is not intended to treat disease or administer any medicinal substances or drugs. There is no contact between device and patient. syngo(R) Workflow SLR can be used as a separate RIS, interfacing to a Hospital Information System (HIS) or Picture Archiving and Communication Systems (PACS | Siemens syngo WorkFlow SLR, Radiological Information System | 2 | April 12, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Under certain conditions, when using the Search functionality, the potential exists for the wrong patient's Interactive Documents to display when viewing a signed report in the single patient view. This can occur when the Interactive Document is viewed or auto displayed, dismissed by the user, and then a different patient and signed exam is selected in single patient view. The Notes Viewer will | DESIGN: Software Design (Manufacturing Process) | Siemens Medical sent a Customer Safety Advisory Notice letter dated October 5, 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.There will be a Hot Fix available.If you have any further questions please contact Siemens Customer Product Support at (1-888) 767-8326. | Contact the recalling firm for information | 37 | 37 | Worldwide Distribution - USA including the states of AL, CA, CO, DC, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV, and WY. and the countries of Canada and Australia. | Terminated | July 16, 2013 | 95 | 37 | Computer
Z-1103-2013 | 63856 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra XKnifeÂŽ Stereotactic Radiosurgery and Radiotherapy Treatment Planning,Software Versions 5.0.1 and 5.0.2.Intended for use in stereotactic and non-stereotactic (frameless stereotactic), collimated beam, computer planned, linear accelerator (Linac) based treatment. | Integra XKnife Stereotactic Radiosurgery and Radiotherapy Treatment Planning | 2 | April 12, 2013 | 2013 | Integra LifeSciences Corp. | Depending on the system configuration, a software error message in versions 5.0.1 and 5.0.2. occurs if a beam plan is transmitted from XKnife using DICOM-RT. |  | Integra initiated a voluntary recall on December 10, 2012 by providing a written Urgent: Medical Device Recall Notification to US consignees who have experienced the software issue due to the configuration of their system. The letter identifies the affected product, as well as, advises customers of the nature of the issue and how it affects treatment records.- Advises customers that they will be contacted when V 5.0.3 software, where the software error has been addressed, is available to schedule an upgrade.- Requests that customer consignees complete and return the Product Recall Acknowledgment Form.If customer have additional questions, they should contact the firm's service hotline at 1-888-772-7378. | Contact the recalling firm for information | 12 | 11 XKnife V5.0.1 and 1 XKnife V5.0.2 | Distributed in the states of IL, ID, LA, MA, NJ, NY, OH, PA, and SD. | Terminated | August 01, 2013 | 111 | 12 | Computer
Z-1102-2013 | 64547 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal (ISP) Philips Medical Systems (Cleveland), Inc.,IntelliSpace Portal (ISP) is a multimodality (CT, NM, and MR) thin-client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. IntelliSpace Portal (ISP) is a medical software system that allows multiple users to remotely access IntelliSpace Workspace Portal (ISP) from compatible computers on a network. The system allows networking, selection, processing and filming of multimodality DICOM images. Both the client and server software are only for use with off the shelf hardware technology that meets defined minimum specifications. trained physicians may use the images as a basis for diagnosis upon ensuring that monitor quality, ambient light conditions and image compression ratios are consistent with clinical application. | Intellispace Portal (ISP) Philips Medical Systems (Cleveland), Inc., | 2 | April 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips was notified of a problem using the IntelliSpace Portal with software version 4.0.2 with the URL PACS integration. Philips was informed that when loading a study to an analysis application while using URL PACS integration configuration, after another study was opened by a review application, the Analysis monitor is updated with the study of the current patient, but the Review monitor conti | DESIGN: Software Design (Manufacturing Process) | Philips sent a Urgent Medical Device Correction letter dated February 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They advised their customers that when reviewing two disparate image datasets, not to only be focused on the actual image content but to also be alerted to the patient's study name and/or ID prior to interpretation.They recommend to manually exit the Review and Analysis applications before loading a new study to Intellispace Portal v4.0.2 from the PACS.For North America and Canada contact the Customer Care Solutions Center at (1-800) 722-9377. | Contact the recalling firm for information | 313 | 313 units | Worldwide Distribution - (USA) Nationwide Distribution including the states of AL, AR, AZ, CA, FL, GA, IA, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, TN, TX, WA, WI & WV. and the countries of Australia, Austria, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Monaco, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand & United Kingdom | Terminated | November 04, 2013 | 206 | 313 | Computer
Z-1112-2013 | 64653 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12Product Usage:The Alaris PC unit is the central programming, monitoring and power supply component for the Alaris System | Alaris PC Unit 8015 | 2 | April 15, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has identified potential risk associated with bolus programming with the Alaris PC unit (model 8015) software version 9.12. Automating infusion pump parameter input from either the electronic medical record (EMR) or the Alaris Auto-ID module for a continuous infusion with a bolus dose option may result in a loss of information from the Guardrails Data s | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carefusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are required to confirm receipt of the notification by returning the Recall response card to Carefusion by postage-paid, self-addressed mail, fax, or email. The letter stated that a Carefusion representatives will contact customers by telephone to schedule field remediation. For questions call Carefusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266, and/or Technical Support at 888-812-3229. |  | 4090 | 4,090 total units | Worldwide Distribution USA Nationwide and the countries of Australia and Canada. | Open |  |  | 4090 | Computer
Z-1098-2013 | 64728 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12Product Usage: The Alaris PC unit is the central programming, monitoring and power supply component for the Alaris System. | Alaris PC Unit 8015 | 1 | April 15, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has received reports of a communication error on the Alaris PC unit model 8015 with software version 9.12 when the Alaris EtCO2 module or Alaris SpO2 module is attached. | DESIGN: Software Design | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. Customers are required to confirm receipt of the notification by returning the Recall Response card to Carefusion by postage-paid, self-addressed mail, fax, or email. Carefusion representatives will contact customers by telephone to schedule field remediation. For questions contact CareFusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266 and/or Technical Support at 888-812-3229. |  | 9241 | 9,241 units | Worldwide Distribution - USA (nationwide) and the countries of Australia and China. | Open |  |  | 9241 | Computer
Z-1126-2013 | 64712 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife System Medical Physicist, Robotic Radiosurgery System, Medical charged-particle radiation therapy system, Model: Iris Variable Aperture Collimator P/Ns 028986 and 031777. The CyberKnife System is indicated for radiation treatment. | CyberKnife System Iris Variable Aperture Collimator | 2 | April 16, 2013 | 2013 | Accuray Incorporated | A defect on the master-side connector block of the pneumatic tool changer could introduce a tilt in the collimator housings, resulting in a potential shift to the central axis of the radiation beam. The issue could pose significant clinical impact to plans created using the smallest collimators. | DESIGN: Component Design/Selection | On 3/19/13, an "Urgent Device Correction Letter" was distributed to all affected users, with a description of the problem and user corrective action steps. US customers can contact 1-877- 668-8667 if questions. | Contact the recalling firm for information | 176 | 176 units | Worldwide distribution, including Nationwide in the US and the countries of Korea, Myanmar, Canada, Belgium, England, Czech Republic, Finland, France, Germany, Greece, India, Italy, Poland, Russia, Spain, Switzerland, Turkey, Columbia, Mexico, Japan and Saudi Arabia. | Terminated | October 23, 2013 | 190 | 176 | Computer
Z-1133-2013 | 64748 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Muse v7 cardiology information system. Model number 2026443-001: MUSE 7.0 app cd, Model Number 2026443-015: MUSE 7.0.2 app cd, Model number 2026443-017: MUSE 7.1 app cd, Model number 2026443-029:MUSE 7.1.1 A app cd, and Model Number 2026443-031: MUSE 7.1.1 app cd.Intended to store, access and manage cardiovascular information on adult and pediatrics patients. | Muse v7 | 2 | April 17, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare conducted a recall on the Muse v7 because incorrect medications were displayed on the MUSE, which could result in incorrect treatment. | DESIGN: Software Design | Consignee letters were not sent for this recall. GE Healthcare's recall strategy was to update all affected systems that were distributed. Updating of the devices was carried out from February 2012 to August 2012 by GE Field personnel. |  | 963 | 963 | Worldwide Distribution -- USA, including all states except ME, and the U.S. Virgin Islands; and, the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NETHERLAND, NEW ZEALAND, NORWAY, PHILIPPINES,S QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | July 19, 2013 | 93 | 963 | Computer
Z-1003-2013 | 64708 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Computed Tomography X-Ray System.These devices are whole-body computed tomography (CT) x-raysystems or sub-systems, each with a continuously rotating x-ray tubeand multi-row detectors enclosed by a gantry. X-ray transmissiondata acquired and taken at different angles can be reconstructed intocross-sectional images. Each device also includes signal analysisand display equipment, patient and equipment supports, components, and accessories. | Philips Healthcare Computed Tomography XRay System | 2 | April 17, 2013 | 2013 | Philips Healthcare Inc. | Philips Healthcare discovered the customer was performing CTDI measurements for facility accreditation and received results that were 15 to 25% higher than Philips protocol parameters stated. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Philips Healthcare proposed corrective action plan (CAP) is comprised of the following elements:* Updating procedures in manufacturing to verify, either through software calibration or through gage-pin measurement, that A-plane collimation is appropriate to the design specification. It is our understanding that these procedures have been implemented.* For refurbished PET/CT and SPECT/CT systems distributed by Philips, themanufacturing process now requires that CTDI testing be conducted in addition to air testing and that each system meet updated CTDI tolerance limits to be published in the product's instructions for use.* Reassessing and correcting the correlation between air dose and CTDI specifications for all CT systems or sub-systems, resulting in updates of the CTDI specifications (tolerances) to be published in the instructions for use. It is our understanding that this reassessment and correction have been implemented, that the updated CTDI specifications will be made available as new CTDI tolerance tables with the customer information letter to be distributed immediately, and that these updated CTDI specifications will also be distributed as an addendum to the instructions for use manualsto the installed base during the next 12 months.* For all Brilliance CT, Gemini PET/CT, and Precedence SPECT/CT systems, in the installed base as well as in forward production, instructions for field service will be updated to provide clear directions on how to verify that the A-plane collimator has been appropriately adjusted to ensure expected product performance after its replacement and/or the replacement of an x-ray tube.This plan appears to adequately address the problem, and it is hereby approved. For further questions please call (978) 687-1501. |  | 6125 | 2089 US 4036 ROW | Worldwide Distribution including USA (nationwide) | Open |  |  | 6125 | Computer
Z-1137-2013 | 64900 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, AW Server.Product Usage: . AW Server is a medical software system that allows multiple users to remotely access AW applications from compatible computers on a network. The system allows networking, selection, processing and filming of multimodality DICOM images. Both the client and server software are only for use with off the shelf hardware technology that meets defined minimum specifications. | GE Healthcare, AW Server. | 2 | April 17, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Results Viewer in the AW Server.After starting the Results Viewer with the selected series of a patient, the pixel and certain annotation data from a different patient is displayed in one of the viewports. This occurs on an occasional basis, when multiple users are simultaneously accessing the Results Viewe | DESIGN: Software Design | GE Healthcare sent an "Important Product Information" letter dated June 27, 2012. The letter was addressed to Hospital Administrators , Radiology Managers and Risk Managers. The letter described the Subject, Instructions and Affected Product Details. GE Healthcare will correct all affected product by providing a software upgrade at no cost. For questions call: US 800-437-1171, Japan 0120-055-919, and/or local GE Healthcare Service Representative. |  | 429 | 429 | Worldwide Distribution - USA Nationwide including the states of: AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, LA, MD, MI, ,N, ,S, MO, MT, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN,TX, UT, VA, WA, WI. and the countries of : AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA, BRAZIL, CANADA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KOREA, LIBYA, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, POLAND, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY and UNITED ARAB EMIRATES. | Terminated | February 04, 2014 | 293 | 429 | Computer
Z-1149-2013 | 64641 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0)Product Usage: The Cardiohelp System is a blood oxygenation and carbon dioxide removal system used to pump blood through the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). | CARDIOHELPi | 2 | April 18, 2013 | 2013 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). | Contact the recalling firm for information | 154 | 154 units affected | Worldwide Distribution - USA Nationwide and the following countries: Australia, Austria, Belarus, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Iceland, Iran, Ireland, Italy, Jordan, Kazakhstan, Korea, Kuwait, Luxembourg, The Netherlands, Norway, Oman, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | August 22, 2013 | 126 | 154 | Computer
Z-1144-2013 | 64683 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea CT Multi-Chamber Cardiac Functional Analysis application on Vitrea version 5.2 (Vitrea Enterprise Suite 1.3 and Vitrea fX 3.1) and subsequent versions of VitreaWorkstation, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite, through version 6.4.VitreaÂŽ CT Multi-Chamber Cardiac Functional Analysis provides a variety of tools for working with clinical CT images of the coronary arteries, heart, and surrounding tissue. | Vitrea CT MultiChamber Cardiac Functional Analysis application on Vitrea version 5.2 | 2 | April 18, 2013 | 2013 | Vital Images, Inc. | Vital Images has made a decision to perform a field correction to the Vitrea CT Multi-Chamber Cardiac Functional Analysis application. Vital Images has identified a defect in which numerical values may be incorrect in the Results Table under the Cardiac Analysis portion of the user interface. | DESIGN: Software Design | Vital Images sent an Urgent Field Safety Notice dated March 15, 2013, to all affected consignees. The letter described the problem and the product affected by the recall. Letter provided a series of steps to follow which included to cease use of the device and to send the signed enclosed Effectiveness Check form to fieldnotices@vitalimages.com or fax to 1-952-540-3717. For questions customers were instructed to call Customer Support at 1-800-208-3005.For questions regarding this recall call 952-487-9500. |  | 102 | 102 | Worldwide Distribution - USA including AL, AR, CA, CO, IA, MI,MN, MO, NY, NC, OH, OR, PA, TX, and UT.Internationally to AUSTRIA, BELGIUM, CANADA, CHINA, COLOMBIA, GEORGIA, JAPAN, KOREA, MACEDONIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SINGAPORE, SWEDEN, SWITZERLAND, TURKEY, UKRAINE, UNITED KINGDOM. | Open |  |  | 102 | Computer
Z-1141-2013 | 64723 | electrode, ion specific, sodium | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HORIBA ABX PENTRA 400 I.S.E. Module, Model Numbers: P400ISE110EN02 and P400ISE110US02, All Serial Numbers, Software Versions 5.0.7 and Below.Used to measure absorbance and ion electrodes. | ABX PENTRA 400 I.S.E. Module | 2 | April 18, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA 400 Analyzer with I.S.E. Module customers that there are noLinearity Limit Flags for the I.S.E. assays currently programmed into the ABX PENTRA 400 Analyzer software. Allother assays run on this analyzer do have Linearity Limit Flags to alert the user to values that fall below or above thevalidated linearity range for each assay. | DESIGN: Software Design | A recall letter, Field Safety Notice & Acknowledgement mailed to the customers March 20, 2013. The letter included the product information, reason for recall, instructions on what to do with the recalled product, and contact information:"If you have any questions or concerns, please contact your local HORIBA Medicalrepresentative. We sincerely apologize for any inconvenience this may cause your laboratory. Please contact your Customer Service Representative at 1-888-903-5001 (option 3) for any additional questions related to this customerletter.". |  | 905 | 905 units | Nationwide Distribution. | Open |  |  | 905 | Computer
Z-1155-2013 | 64809 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Positioning SystemProduct Usage: is intended for use in the support and positioning of a patient during radiosurgery and radiotherapy procedures and other medical procedures when precise positioning is required. | RoboCouch Patient Positioning System | 2 | April 20, 2013 | 2013 | Accuray Incorporated | A5 gearboxes responsible for pitch (head up and head down) and roll (patient left and right) motions of the RoboCouch manipulator may be defective. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Accuray issued an Urgent Advisory Notification letter dated March 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that Field Service Engineers will be replacing the affected product. Questions and concerns are to be directed to Accuray Customer Support at 1-877-668-8667. |  | 16 | 16 devices are subject to correction. | Worldwide Distribution - USA Nationwide including the states of: OH, MO. OK, CA and VA and the countries of: Finland, France, Luxembourg, Poland, Russia, Turkey and Japan. | Terminated | August 20, 2013 | 122 | 16 | Computer
Z-1080-2013 | 64594 | full-montage standard electroencephalograph | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software versions 5.1 uo to 6.2e :Harmonie-E Long Term Monitoring system (Models HSYS-REC-WD, HSYS-REC-WT; HSYS-REC0LT2; HSYS-REC-EAMP)Product in use with Harmonie-E Long Term Monitoring System : ICTA-S & ICTA-D Seizure Detector, Slow wave Seizure Detector Harmonie-S Sleep System: Rapid Eye Movement Detector, Spindle Detector, Heart Rate Processor, Periodic Leg Movement Detector and Oximetry/Pulse rate Threshold Notification VITA/ICU Neurological Monitoring System: Burst Suppression Detector, Cerebral Function Monitor Trend, Visual and Auditory Notifications on all trends graphs in PolyTrends Harmonie {-s}{-E} Readers , | Harmonie software versions 5.1 uo to 6.2e : | 2 | April 22, 2013 | 2013 | Stellate Systems | The firm is initiating this recall because the implied or stated indications in their labeling associated with Harmonie software have not yet been cleared by the FDA for sale in the USA. |  | Stellate sent an "DEVICE CORRECTION" letter dated April 6, 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were instructed to check the product code referenced in the service manual to verify if your system is affected and check in "Help" & "About Harmonie" in the menu to determine the software version on these systems. Stellate will shortly release updates to disable the noted features that were present in your particular version of the Harmonie software and revise labeling and Instructions for Use. All customers in possession of an affected system will be contacted by our Customer Support team to set up an appointment for installation of the updated version.If you have any questions or concerns regarding this notification, please do not hesitate to call our customer support line at 1-888-742-1306. | Contact the recalling firm for information | 2660 | 2660 Software licenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, DE, FL, HI, IA, IN, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, OR, PA, RI, SC, TX, VA, VT, WI, WV, and WY; and the countries of: Argentina, Australia, Bahrain, Brasil, Canada, China, Denmark, Finland, France, Germany, Greece, Haiti, Hong Kong, Hungary, Indonesia, Iran, Israel, Italy, Japan, Korea, KSA, Mexico, Netherlands, Norway, Pakistan, Portugal, Rep. DE Panama, Saudi Arabia, Spain, Switzerland, Taiwan ROC, Thailand, Turkey, and United Kingdom. | Terminated | April 22, 2013 | 0 | 2660 | Computer
Z-1164-2013 | 64804 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Xeleris 2 Processing & Review Workstation.The display, processing, archiving, and communication of data acquired by Emission Tomography cameras used in diagnostic radiology, including procedures for planar imaging, whole body imaging, tomographic (SPECT) imaging, positron imaging by coincidence, attenuation correction, and anatomical image registration | Xeleris 2 Processing & Review Workstation | 2 | April 23, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue in your Xeleris 2.1 workstation (that is used for processing Discovery 530/570 data), that may impact patient safety.On imported CT scans, Myocardial Perfusion Imaging (MPI) studies may create artifacts in attenuation corrected images. This may lead to a potential for misdiagnosis. | DESIGN: Software Design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 5, 2011, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to contact their local GE Healthcare Service Representative if they have questions or concerns. They may also contact GE Healthcare Call Center , United States or Canada at 800-437-1171.Other countries 262-896-2890.For questions regarding this recall call 262-513-4122. |  | 11 | 11 | Worldwide Distribution - USA including GA, and NY; and countries of: Canada, Israel, Italy and Switzerland. | Terminated | February 04, 2014 | 287 | 11 | Computer
Z-1163-2013 | 64828 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand RPM Respiratory Gating System, RPM Gating v1.7.5 with 3D Option, Model Number: HS2Product Usage: The RPM device is to be used to characterize the patients respiratory motion information to synchronize their operation with the respiratory motion. | Varian | 2 | April 23, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the RPM Respiratory Gating System version I.7.5 with 3D Option. While operating with the 6-dot marker block, this version of RPM Gating may produce an inverted respiratory waveform in amplitude and phase-gating sessions. | DESIGN: Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice Letter dated March 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain a copy of this document along with their most current product labeling, advise the appropriate personnel of the content of this letter, and for future reference this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. The letter stated that a Varian Medical Systems has developed a technical correction and a Customer Service representative will contact customers to schedule installation. For question contact your local Varian medical Systems Customer Support District or Regional Manager. | Contact the recalling firm for information | 2508 | 2508 units | Worldwide Distribution - USA (Nationwide) and countries of: Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Mexico Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab E, and United Kingdom. | Open |  |  | 2508 | Computer
Z-1168-2013 | 64694 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ventri Discovery NM 530c, Ventri with Premium Table Discovery NM 530c, models H3000ZW & H3000SA. The system is used to perform nuclear imaging procedures. | Ventri Discovery NM 530c | 2 | April 24, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident which caused a patient injury during an unload following an exam on a Ventri scanner. During a patient unload following a scan, the patients fingers may be pinched between the moving cradle edge and the stationary base of the table while the table is moving outward automatically. |  | Consignees were sent on 1/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated January 9, 2013. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers andManagers of Radiology / Head of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 556 | 556 | Worldwide Distribution - US, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, IL, IN, IA, KY, LA, ME, MD, MA, MI, ,N,MS, MO, NE, NV, NH, NJ, NM, NY,NC, OH, OK,OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI and the countries ofTAIWAN, TURKEY, SINGAPORE, SWD EN, PORTUGAL, POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MALAYSIA, MEXICO, SOUTH KOREA, JAPAN, ITALY, INDIA, ISRAEL GREAT BRITAIN, FRANCE, DENMARK, GERMANY, SWITZERLAND, CANADA, BRAZIL BELGIUM, AUSTRALIA, and ARGENTINA. | Terminated | August 06, 2013 | 104 | 556 | Computer
Z-1167-2013 | 64731 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry AnalyzerThe Beckman Coulter AU5800 Clinical Chemistry Analyzer is an automated chemistry analyzer that measures analytes in samples, in combination with appropriate reagents, calibrators, quality control (QC) material and other accessories. This system is for in vitro diagnostic use only. Applications include colorimetric, turbidimetric, latex agglutination, homogeneous enzyme immunoassay, and in selective electrode. | AU5800 Clinical Chemistry Analyzer | 2 | April 24, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter, Inc. initiated this recall because there is a software error that has the potential to cause the reporting of erroneous patient results because of misidentification by the AU5800 system. | DESIGN: Software Design | Beckman Coulter sent an Urgent Product Correction letter during the week of March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were informed of the software limitation and provides instructions to be taken (temporary workaround) until the software is upgraded and released. Customer contact information provided: "Call Center Hotline at 800 854-3633. New customers will receive a copy of the customer notification at the time of installation by the Field Service Personnel until revised software that corrects this issue is released. |  | 490 | 490 | Worldwide Distribution - USA, including Canada and Puerto Rico and the states of NJ, TX, CA, UT, KS, MA, NY, MI, GA, VA, IN, WI, IL, OH, NV, NC and PA. and the countries of Australia, Belgium, Czech Republic, Japan, Germany, Brazil, China, Italy, Ireland, Israel, India, Hong Kong, Malaysia, Netherlands, Spain, Switzerland, France, Taiwan, Turkey, United Kingdom and Vietnam. | Open |  |  | 490 | Computer
Z-1159-2013 | 64839 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Pump - BatteryProduct Usage: The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. | GemStar Infusion System | 1 | April 26, 2013 | 2013 | Hospira Inc. | Battery level lower then 2.4 volts results in corrupt history log and loss of data for preview therapy's parameters and causes an 11/004 error resulting in the pump not being able to be used. | DESIGN: Device Design | Hospira sent an Urgent Device Field Correction letter dated March 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. Lithium batteries that are older than three (3) years should be replaced. Customers should contact their Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, to determine if their battery need to be replaced and if necessary, arrange for the return of their device to perform battery replacement. Customers with questions should call 1-877-907-7516. For questions regarding this recall call 224-212-4892. |  | 185 | List Number 13000: 25,119 pumps; List Number 13100: 185 pumps; List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open |  |  | 38801 | Computer
Z-1182-2013 | 64944 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Software correction applies to customers who use the RT Therapist RTT4.2.108 and RT Therapist Connect RTT 4.2.108 on Siemens ARTISTE, ONCOR and PRIMUS systems.ARTISTE, ONCOR and PRIMUS family of linear accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens ARTISTE, ONCOR and PRIMUS systems. | 2 | April 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Software update to fix multiple safety related issues. | DESIGN: Software Design | Siemens sent updated instructions of TH004/13/S on April 4, 2013. The SW Update will be delivered by the Siemens Customer Service Group.For questions regarding this recall call 925-293-5442. |  | 245 | 245 active devices | Worldwide Distribution - USA (nationwide) and internationally to Poland, India, Canada, Germany, United Kingdom, New Zealand, Belgium, Malaysia, Republic Korea, Ireland, Australia, Norway, Japan, Turkey, Croatia, Saudi Arabia, Colombia, P.R. China, Brazil, Thailand, Romania, Lebanon, Mexico, Italy, Sweden, Czech Republic, Russian Federation, and South Africa. | Terminated | August 12, 2013 | 108 | 245 | Computer
Z-1181-2013 | 64833 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaÂŽ CT Brain Perfusion made available in various configurations under the names Vitrea, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite.Noninvasive post-processing application designed to evaluate areas of brain perfusion. | Vitrea | 2 | April 26, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. (Vital Images) received one report about potential user confusion when viewing a brain perfusion scan on Vitrea CT Brain Perfusion 2D when the scan was performed using a wide coverage shuttle acquisition technique. This may result in patient treatment delays, misdiagnosis or mistreatment. | DESIGN: Labeling Design | Consignees were sent on 3/15/2013 a VITAL Images, Inc., "Urgent Field Safety Notice" dated March 15, 2013. The letter described the problem and the product involved in the recall. The letter listed the limitations of the product and the steps to take. Consignees were requested to return the Product Recall Effectiveness Check form. Questions should be directed to Vital Images Customer Support at 1-800-208-3005 or via e-mail, support@vitalimages.com. |  | 1978 | 1,978 | Worldwide Distribution -- USA, excluding RI and SD and including DC and PR, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELARUS, BOLIVIA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GREECE, GUAM, HONDURAS, HONG KONG, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN. | Open |  |  | 1978 | Computer
Z-1177-2013 | 64645 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Getinge Group battery modules for use with Maquet SERVO-i ventilator systems (Part #6487180).Battery for use with the SERVO-i Ventilator System. The SERVO-i Ventilator System is intended for treating and monitoring patients ranging from neonates to adults with respiratory failure or insufficiency. The ventilator is equipped with at least 2, NiMH battery modules (Part No. 6487180) which automatically supply 12V DC Power in case of an AC power failure, ensuring that ventilator settings and stored data remain intact in the event of an AC power failure. The battery modules give the user a possibility to use the SERVO-i during intra-hospital transportation. | SERVOi ventilator system | 1 | April 26, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | When using the SERVO-i on battery power, in a very small number of cases, some battery modules distributed after January 31, 2010 have a shorter battery run time than expected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Maquet Getinge Group" sent "Field Safety Notices"/"Field Safety Notice Confirmation Forms" dated March 7 & 11 2013 via Federal Express and an updated Field Safety Notices dated April 15, 2013 to its customers. The Field Safety Notices identifies the problem, product, risk factors and actions to be taken by the customer. The customers were instructed to reconnect the ventilator to mains power immediately or add additional battery modules in case of 'low battery voltage' alarm; ensure that at least two fully charged batteries are installed at all times, and that sufficient resuscitation equipment is accessible to the ventilator; identify and immediately replace battery modules with date codes 1005-1243, that have generated either: a "low battery voltage" alarm or a "replace battery" message, prior to the expiry date; and complete and return the enclosed Field Safety Notice Confirmation Form by e-mailing a scanned copy to FieldActions@maquet.com or via fax at (973) 807-1832.If you have any questions or require replacement batteries, please contact your local MAQUET representative, or MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 90000 | 90,000 units (batteries) | Worldwide distribution: USA (nationwide) and countries of: ALGERIA, ANDORRA, ANTILLES (Netherlands), ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBADJAN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BRUNEI, BULGARIA, CANADA, CHILE, CHINA,COLOMBIA, CONGO, COSTA RICA, CROATIA, CUBA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, YUGOSLAVIA (Former), FIJI, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONDURAS, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LIBYA, LITHUANIA, LUXEMBURG, MACAU, MACEDONIA, MALAYSIA, MALDIVES, MARTINIQUE (French), MEXICO, MOLDAVIA, MOROCCO, NEPAL, NETHERLANDS, NEW CALEDONIA (French), NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, POLYNESIA (French), PORTUGAL, QATAR, REUNION (French), ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SEYCHELLES, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SUDAN, SURINAM, SWEDEN, SWITZERLAND, SYRIA, TADJIKISTAN, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VATICAN city state, VENEZUELA, VIETNAM and YEMEN. | Open |  |  | 90000 | Computer
Z-1171-2013 | 64851 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport Medical Instruments HT70and HT70 Plus ventilator Power Pac batteries Model number: HT70 and HT70 Plus. The HT70 and HT70 Plus ventilator systems are intended to provide continuous or intermittent positive pressure mechanical ventilatory support for the care of individuals who require mechanical ventilation through invasive or non-invasive interfaces. | Newport Medical Instruments HT70 and HT70 Plus ventilator Power Pac batteries | 1 | April 26, 2013 | 2013 | Newport Medical Instruments Inc | Newport Medical Instruments is conducting a voluntary recall on certain Newport Medical Instruments HT70and HT70 Plus ventilator Power Pac batteries due to customer reports of Newport HT70 and HT70 Plus ventilators alarming and going to internal backup battery sooner than expected while the ventilator is being operated on Power Pac battery. | DESIGN: Component Design/Selection | Newport Medical Instruments sent an Urgent Medical Device Recall letter dated April 4, 2013, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who sold the affected product to another party, or placed the affected product with a patient or healthcare provider were instructed to immediately inform the other party, patient, or healthcare provider of this action and provide them with a copy of the letter. If customers were using the affected product, they should ensure that an alternate source of power is available nearby. Customers were instructed to report any issues with Newport" HT70 and HT70 Plus ventilators or Power Pac batteries to the Technical Support group to ensure proper device reporting procedures are followed. Please call 800-451-3111, extension 500, for U.S. customers; and 714-427-5811, extension 500 for international customers and you will be given further instructions).For questions regarding this recall call 714-427-5811, ext 344.A press release was issued by the firm on 04/18/13. |  | 4392 | 2528 ventilators and 1864 power pac batteries | Worldwide Distribution - USA including US (AL, AR, AZ, CA, CO, CT, DE, FL, GA, lA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, WA, and WI. Internationally to Argentina, Australia, Bangladesh, Belgium, BoliVia, Brazil,Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, France, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Ireland, Italy, Japan, Kazakhstan, Kosova, Lebanon, Lithuania, Malaysia, Mexico, Nambia, Nepal, Norway, Oman, Paraguay, Peru, Philippines, Poland, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Sri Lanka, Switzerland,Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Open |  |  | 4392 | Computer
Z-1189-2013 | 64860 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | HeartStart MRx Monitor/Defribillator; Product Codes: M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6.The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. | HeartStart MRx Defibrillator | 2 | April 29, 2013 | 2013 | Philips Healthcare Inc. | Device may fail to deliver defibrillation therapy in either Manual Defib or AED mode. | DESIGN: Software Design | The firm, Philips Healthcare, sent an "Field Safety Notice-URGENT - MEDICAL DEVICE RECALL" letter dated April 2013 to its customers. The letter describes the product, problem and actions to be taken. In addition, letter provides instructions for using the MRx prior to receiving the software upgrade and informs the customer that Philips is initiating a correction consisting of a software upgrade for all units affected. The customers were instructed to follow the "ACTION TO BE TAKEN BY CUSTOMER/USER" section of the Field Safety Notice; send a copy of the attached package including customer letter and Field Safety Notice to any customer to whom you have distributed one of the affected device; and provide your local Philips organization with the names and addresses of customers to whom you have sold affected devices, so that arrangements can be made to perform the software updates.Should you have any questions or concerns about this device correction, please contact your local Philips representative at 1-800-722-9377.Philips updated the Field Safety Notice dated May 2013 to users who have not had the software upgrade performed. It provided additional information in the "Problem Description" and "Action to Be Taken By Customer/User". |  | 64972 | 64972 | Worldwide distribution: USA (nationwide) and countries of: Afghanistan, Albania, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegowina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic Of, Kuwait, Laos, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Malta, Mexico, Mongolia, Morocco, Namibia, Nepal, Netherlands (Antilles), New Caledonia, New Zealand, Nicaragua, Nigera, Norway, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Portugal, Puerto, Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, SRI Lanka, Sweden, Switzerland, Taiwan (Province of China), Tanzania, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Veit Nam, and Yemen. | Open |  |  | 64972 | Computer
Z-1185-2013 | 64868 | pacemaker, cardiac, external transcutaneous (non-invasive) | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Medi-Trace Cadence Pediatric Radiotransparent Defibrillation ElectrodeProduct ID: 22550P.Intended for use in cardioversion, pacing, ECG monitoring, and defibrillation procedures. | Covidien MediTrace Cadence Pediatric Radiotransparent Defibrillation Electrode | 2 | April 29, 2013 | 2013 | Covidien LLC | Defibrillation Electrodes may arc/spark result in thermal damage to the leadwire, which could render the device incapable of delivering the appropriate energy or shock to the patient | PRODUCTION CONTROLS: Process Control | Covidien sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 8, 2013 to all affected customers. The letter identifies the product, problem and requests that customers immediately stop using and segregate any affected product from inventory. Questions:Call 1-800-962-9888 option 8, then extension 2500. |  | 7485 | 7485 (sets of 2) | Worldwide Distribution-USA (nationwide) and the countries of Korea, Japan, Canada, Australia, Brazil, Belgium, Singapore, and South Africa. | Open |  |  | 140955 | Computer
Z-1187-2013 | 64812 | oximeter | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Rad-8 Pulse Oximeters The Rad-8 family of pulse oximeters and accessories are indicated for the continuous, noninvasive monitoring of functional oxygen saturation of arterial hemoglobin (SpO2) and pulse rate (measured by an SpO2 sensor). | Pulse Oximeter | 2 | April 29, 2013 | 2013 | Masimo Corporation | Masimo has identified a very small number (0.02%) of Rad-8 devices with 20-pin patient cable connector that, when moved, can power off without the operator pressing the power button. | DESIGN: Device Design | Masimo Corporation sent an Urgent Recall Notification letter dated March 8, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.1.Customers were instructed to please promptly remove all Rad-8 device(s) from use and your inventory and segregate these devices.2. Please call 1-800-326-4890 and select option 2 for Technical Services. Determine with Technical Services whether you will return the affected device(s) to Masimo for repair or if eligible, receive a repair kit. If you will return the device(s) for repair, obtain a Return Material Authorization Number (RMA number).3. Complete the Tracking/Verification Form in Attachment 2 and fax it to Masimo at 1-949-297-7700, even if you do not have or intend to use the affected device(s).4. If you are eligible and intend to repair units onsite fax the completed original form in Attachment 2 to Masimo at 1-949-297-7700.5. If you have elected to return the affected device(s), and have an RMA number include the completed original Form in Attachment 2 with the returned device(s) and indicate the RMA number on the outside of the return-shipping box.We apologize for the inconvenience. Please be assured that Masimo is committed to consistently providing high quality products and services to you, our customers. We thank you for your patience and cooperation while we actively work to resolve this issue... | Contact the recalling firm for information | 42908 | 42,908 units | Worldwide Distribution - USA (nationwide) including the countries of Africa, Asia, Australia/New Zealand, Europe, Canada, Latin America, and Middle East / SW Asia. | Open |  |  | 42908 | Computer
Z-1208-2013 | 64918 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, GE 1.5T SIGNA HDe MR SYSTEM.Indicated for use as a diagnostic imaging device. | GE 1.5T SIGNA HDe MR SYSTEM | 2 | May 01, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a Reformat Image Flip Issue associated with GE MR Products.The axial and coronal reformatted images from a sagittal 3D series (such as with applications of FSPGR-3D T1, BRAVO T1, or CUBE T1) of over 126 slices may display a left-to-right image flip .In addition, the acquired sagittal images may be annotated with left-right location annotation revers | DESIGN: Software Design | GE Healthcare sent an "Important Product Information" letter dated June 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their local service representative for questions regarding this notice. | Contact the recalling firm for information | 822 | 822 (154 US; 668 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, excluding the state of RI and the countries of Yugoslavia YEMEN Viet Nam Venezuela Uzbekistan Uruguay United Arab Emirates Tunisia Thailand Tanzania Taiwan Turkey Syria Switzerland Sweden Spain South Korea South Africa Slovenia Slovakia Singapore Saudi Arabia Russia Romania Reunion Qatar Portugal Poland Philippines Peru Pakistan Paraguay Philippines Oman Norway Nigeria New Zealand Netherlands Morocco Mexico Martinique Malaysia Macedonia Mauritania Luxembourg Lithuania Libya Lebanon Latvia Kuwait Kenya Korea Korea (Republic Of [South] Korea) KUWAIT Kazakhstan Jordan Japan Jamaica Italy Israel Ireland Iran Iraq Indonesia India Hungary Hong Kong Honduras Guadeloupe Greece Germany Georgia, Republic of French Polynesia France Finland Estonia Ecuador Egypt Denmark Dominican Republic Czech Republic Croatia Colombia China Chile Canada Bulgaria Brazil Bosnia And Herzegovina Belgium Belarus Barbados Bahrain Austria Australia Argentina Arab Emirates Algeria and United Kingdom. | Open |  |  | 6024 | Computer
Z-1000-2013 | 64589 | unit, x-ray, intraoral | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | CareStream Dental CS 2200 X-ray systemProduct Usage:The expected usage (intended use) for all of these products is for dental radiographic examination and diagnosis of the teeth, jaw and oral structures. | CareStream Dental | 2 | May 02, 2013 | 2013 | Trophy Sas | There have been approximately 2% of the x-ray generators manufactured between February 2012 and November 2012 returned due to errors. These generators showed the type of defect that indicates poor contact of HT transformer power pin. The defect is a bent connector associated with the generator. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Trophy, A Carestream Dental Company, - A Subsidiary of Carestream Health Inc. The Center for Devices (CDRH) understands that the corrective action plan (CAP) involving 167 units will consist of implementing rework procedure TL008, which describes the procedure for a bending the metallic strip the female connector to insure strong connection with the male counterpart.The CAP appears to adequately address the problem and is hereby approvedFor further questions please call (585) 781-1997. |  | 167 | 167 installed in the US | USA Nationwide Distribution | Open |  |  | 167 | Computer
Z-1277-2013 | 64935 | unit, electrosurgical, endoscopic (with or without accessories) | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SurgMaster UES-40 electrosurgical unit ("UES-40")Product Usage: Designed for use in a medical facility under the supervision of a trained physician. It has been designed for general (open) and endoscopic surgery including urology, gynecology, respiratory and gastroenterology in conjunction with Olympus designed electrosurgical accessories, endoscopes (fiberscopes, videoscopes and rigid scopes) applicable for electrosurgery (cutting and coagulation), light sources and other ancillary equipment. | Olympus SurgMaster UES40 electrosurgical unit ("UES40") | 2 | May 08, 2013 | 2013 | Olympus America Inc. | Olympus initiated a software upgrade and enhancement in 2009 and completed the action in 2010. The upgrade and enhancement were initiated due to complaints of overheating, smoke vapors and failure to operate associated with the Olympus SurgMaster electrosurgical unit, model UES-40 ("UES-40") when used in conjunction with the Olympus WA22557C electrode. The complaint investigations revealed that |  | Olympus America Inc sent a Field Corrective Action letter dated November 10, 2009 to all affected consignees. The letter identified the affected product, problem and actions to be taked. Consignees were directed to contact Gyrus at 888-524-7266 to schedule return of the UES-40 generator for the software upgrade. The letter also directed consignees to contact OAI Customer Service at 800-848-9024 if they needed a new copy of the UES-40 Instruction Manual. For questions call 484-896-5688. | Contact the recalling firm for information | 608 | 608 | Nationwide Distribution - USA Nationwide including the states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WY. | Terminated | July 05, 2013 | 58 | 608 | Computer
Z-1318-2013 | 65029 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ, Innova 2121 IQ, Innova 3131 IQ , Innova 3100, Innova 4100.For use in generating fluoroscopic images of human anatomy for diagnostic and interventional procedures. | GE Healthcare Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ, Innova 2121 IQ, Innova 3131 IQ , Innov | 2 | May 14, 2013 | 2013 | GE Healthcare, LLC | While performing a fluoroscopic examination on the GE Innova systems, there is a potential of nonrecoverableloss of imaging modes with no advance warning, due to a relay failure in the power distribution box (PDB). Although no injuries have been reported due to this issue to date, there is riskof complications during the procedure as a result of imaging functionality loss. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare did not send out customer letters. GE Healthcare completed the correction by replacing the faulty relay in all affected systems.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests recorded the repair of the affected units. Exceptions to completion of the service request were recorded as well. Records of field action completion will be retained per their processes and procedures.Should you have any questions or require additional information, please call (262) 548-2757. |  | 713 | 713 | Worldwide Distribution -- Nationwide including the states of AL, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, HI, ID, IL , IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WY and WI., and the countries of Philippines, Peru, Paraguay, Panama, Mexico, Malaysia, Korea, Japan, Guatemala, Ecuador, Dominican Republic , Columbia, China, Chile, Canada, Brazil.,Argentina and Venezuela | Open |  |  | 713 | Computer
Z-1324-2013 | 64986 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cylinder Applicator Set, Model # GM11004150, when used for Pulsed Dose Rate (PDR)Product Usage: Usage:Used during pulsed dose rate brachytherapy to treat vaginal and rectal cancer. | Segmented Cylinder Applicator Set | 2 | May 15, 2013 | 2013 | Varian Medical Systems, Inc. | Segmented cylinder applicator set may slip during treatment, causing the delivery of radiation to areas outside the target. | DESIGN: Device Design | Varian Medical Systems, Inc. notified customers of the recall with 'Urgent Medical Device Correction Urgent Field Safety Notice," dated 4/11/13, and delivered by mail to customers. The notice identified the product, problem, and instructions for the customers. Varian Medical Systems, Inc. informed customers of the possible malfunction of the GM11004150 Segmented Cylinder Applicator Set for pulsed dose rate brachytherapy treatment, with codes H640383, H640277, H640384, H640788, H640D126, H640385, H640372, H640303, H640387. The notice instructed customers to cease use of the Segmented Cylinder Applicator Set and return the devices along with a completed proof of notification form (included with the recall notice) to the recalling firm. |  | 21 | 21 devices | Worldwide distribution: USA (nationwide) and in the countries of: Belgium, France, Martinique, Slovenia, Sweden, and the United Kingdom. | Terminated | January 13, 2014 | 243 | 21 | Computer
Z-1325-2013 | 64997 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Clinacs, High Energy Accelerator, Radiation Treatment System, Model Numbers: H14, H27,H29, HCX, Product Usage:The Varian High Energy Accelerator is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. | Varian High Energy Clinacs | 2 | May 15, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified that under certain conditions, photon beams in High Energy Clinacs may experience a gradual change in beam symmetry, potentially reaching an asymmetry of up to approximately 7% before interlock occurs. At maximum asymmetry, this may result in no more than approximately 3.5% dose deviation from expected at any point in the beam, and may result in minor injury to the patient or | DESIGN: Process Design | Varian sent a Customer Technical Bulletin to all affected customers. The bulletin identified the affected product, problem and actions to be taken. For questions, issues, or concerns contact Varian at 1-888-827-4265 or go to http://my.varian.com. | Contact the recalling firm for information | 3943 | 3943 units | Worldwide distribution - USA (nationwide) including Puerto Rico, and countries of: Algeria, Angola, Australia, Austria, Bahrain, Bangladesh, Belarus, Belguim, Brazil, Bosnia and Hezregovina, Bulgaria, Canada, Chile, China, Columbia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, Frances, French Polynesia, Georgia, Germany, Greece, Gaudaloupe, Gaum, Hong Kong, India, Indonesia, Ireland, Italy, Israel, Latvia, Lithuania, Luxemburg, Japan, Kazarhstan, Kuwait, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morroco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Phillipines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudia Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Syria, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezula, Vietnam, and Zimbabwe. | Terminated | May 15, 2013 | 0 | 3943 | Computer
Z-1326-2013 | 65028 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm InSight, software versions 1.x, 2.x, 3.xThe device is a Mini C-arm Fluroscopic Imaging System designed to provide physicians with general fluoroscopic visualization of a patient, including but no limited to surgical orthopedic and podiatry use, critical and emergency procedures, and light anatomy imaging situation. | Fluoroscan Mini Carm InSight, software versions 1.x, 2.x, 3.x | 2 | May 17, 2013 | 2013 | Hologic, Inc. | The audible alarm to alert the operator when cumulative irradiation time reaches five minutes is not accurate and the total cumulative irradiation time is not accurate. | DESIGN: Software Design | The firm, Hologic, Inc. sent an "Urgent: Medical Device Recall" letter dated April 25, 2013, via FedEx to consignees/customers notifying them of the software defect and providing them with corrected user installable software version. The customers were instructed to complete and return the Enclosure to Customer Notification, Confirmation of Notification form via fax to Hologic at 866-652-8674 or mail the document to Hologic using the enclosed pre-addressed, pre-paid return envelope.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. | Contact the recalling firm for information | 287 | 287 | Worldwide distribution: USA (nationwide) and countries of: DE, RU, AU, IT, GB, ES, FR, KR, ZA, CA, AR, LT, TW, CN, BE, AT, JP, QA, SA, SG, HK, SI, NL, and SE. | Open |  |  | 3067 | Computer
Z-1340-2013 | 64647 | camera, surgical and accessories | KQM | Surgical camera and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Ref: VP1600 HD Still Capture System ConMed Linvatec .Captures, stores and transfers images generated by a medical endoscopic camera system during surgical procedures. | HD Still Capture | 2 | May 20, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The VP1600 HD Still Capture Image System is recalled due to an anomaly between the software code in the VP1600 console and the communications interface for various keyboard key strokes (Caps lock and Num Lock) which may lock up and requires the user to reboot to continue. | DESIGN: Software Design (Manufacturing Process) | The firm, ConMed, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/8/2013 to its customers after initial phone notifications. The letter identified the product, problem, and actions to be taken by the customers. |  | 267 | 267 units | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the states of CA, CT, FL, GA, IN, KY, KS, LA, MD, MI, MN, MO, MS, MT, NE, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, and WI, and the countries of Austria, Belgium, Canada, Denmark, Hong Kong and Zambia. | Terminated | November 06, 2013 | 170 | 267 | Computer
Z-1338-2013 | 64713 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 16-Slice Water, System Code: 728245; Brilliance CT 10-Slice Water, System Code: 728250; Brilliance CT 6-Slice Water, System Code: 728255; MX8000 IDT 16, System Code: 728120 & MX8000 IDT 10, System Code: 728125 Computed Tomography X-ray systems. Distributed by Philips Medical Systems (Cleveland), Inc., 595 Miner Road, Cleveland, OH 44143Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance CT | 2 | May 20, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a report from the field stating when they viewed the CTDIVol information from the PACS system, the calculations were too high. Instead of storing the CTDIVol value for each slice in the DICOM tag, the system took the sum of all CTDIVol values from all the images in the dataset and stored this value in the DICOM tag, resulting in incorrect information. |  | The firm, Philips Healthcare, sent an "Urgent - Medical Device Correction" letter dated March 15, 2013, to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to disregard unusually higher CTDIvol value for Type-3 DICOM tag 0018:9345 in the PACS system, if seen. The firm is implementing software update versions 4.2.2 or 2.2.2 to correct the issue. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 515 | 515 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, VT, WA, WI, WV and WY; countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Denmark, France, Germany, Gibraltar, India, Iraq, Ireland, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Philippines, Poland, Portugal, Singapore, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom and Venezuela. | Open |  |  | 515 | Computer
Z-1339-2013 | 65096 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | LATITUDEÂŽ Patient Management System, Model 6488, Version 7.3.The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | LATITUDE Patient Management System | 2 | May 20, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.For questions regarding this recall call 651-582-4000. |  | 1 | 1 | Nationwide Distribution including DC and PR | Terminated | November 26, 2013 | 190 | 1 | Computer
Z-1345-2013 | 56447 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TD-Synergy.Laboratory Information System. | TDSynergy | 2 | May 21, 2013 | 2013 | Technidata S.A. | Technidata has discovered that using the F8 key in the ERM session to add a combined test can result in the modification of the value of a result that has already been entered for a test of the request. | DESIGN: Software Design | The firm, Technidata, sent an "P1 DEFECT INFORMATION" notice dated July 2, 2010 to their consignee/customers. The notice describes the product, problem and actions to be taken. The notice directs them to find the correction to the software issue on the Technidata website where they are able to install the correction kit depending on the version and operating system used. Note "You must back up the database and the installation directory before installing this correction".Should you need any further information, do not hesitate to contact the Support team at +33 (0) 4 76 04 13 00. |  | 10 | 10 | Worldwide Distribution-USA (nationwide) including Washington DC and the states of California, Nevada, Kentucky, Alabama, Rhode Island, Pennsylvania, and Louisiana and the country of Italy. | Terminated | December 05, 2013 | 198 | 10 | Computer
Z-1346-2013 | 64930 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | cobasÂŽ 4800 system z480 instrument. Integrates fully automated total nucleic acid isolation directly from secondary sample tubes. | cobas 4800 | 2 | May 21, 2013 | 2013 | Roche Molecular Systems, Inc. | In rare instances, channel shifted results have been generated with cobasÂż 4800 assays run on v1.1.1 of the cobasÂż 4800 system. Raw data review indicates that, in these cases, all signals are shifted by one channel:channel 1 results report as channel 2,channel 2 results report as channel 3,channel 3 results report as channel 4, andchannel 4 results report as channel 1.In the reported | DESIGN: Software Design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification with response form dated April 1, 2013, to all affected customers via UPS Ground on April 1, 2013. The letter identified the product, the problem, and the action to be taken by the customer. A contact number (Roche Diagnostic Technical Support) is provided for technical support 24 hours a day, seven days a week at phone number: 1-888-912-7090. | Contact the recalling firm for information | 425 | 425 instruments | Nationwide Distribution. | Open |  |  | 425 | Computer
Z-1368-2013 | 64850 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CobasÂŽ 8000 modular analyzer series software version 03-01, 02, 03 and 04, impacting the cobas e 602 module only.Product Usage: Fully automated immunoassay analyzer intended for the in-vitro quantitative/qualitative determination of analytes in body fluids. This analyzer is optimized for high throughput workloads in the professional environment using a combination of ion selective electrodes and photometric analysis modules | Roche cobas e 602 module | 2 | May 23, 2013 | 2013 | Roche Diagnostics Operations, Inc. | When using the cobas e 602 module, Diluent Multiassay may be used for autodilution of certain assays instead of Diluent Universal. This may cause incorrect results under specific circumstances if all of the following conditions are met simultaneously: More than a total of 2000 reagent kits or diluents have been registered on the cobas e 602 modules within one system core, and two different diluent | DESIGN: Software Design (Manufacturing Process) | Roche sent an Urgent Medical Device Correction (UMDC) letter dated April 4, 2013 via UPS to all affected customers. The letter idenified the affected product, problem and instruction to be followed. Customers were instructed to consult with the physician or pathologist at their facility to determine specific clinical implications for patients. A future software version will correct this issue. Customers were asked to complete and fax back the UMDC attached fax form to 1-88-628-0730. Questions should be addresssed to Roche Diagnostics Technical Support 24 hours a day at 1-800-428-2336. |  | 72 | 72 units | USA Nationwide Distribution | Open |  |  | 72 | Computer
Z-1374-2013 | 65031 | system, environmental control, powered | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare-Networked.Intended for use in monitoring the environmental conditions and activity (motion) of an individual living in a senior housing community. | QuietCareNetworked | 2 | May 24, 2013 | 2013 | Intel-GE Care Innovations LLC | Customers at 4 facilities reported receiving duplicate alerts using the QuietCare-Networked product. | DESIGN: Software Design | Care Innovations contacted their customers by phone and e-mail regarding the potential problems and fixes on April 4, 17, and 23, 2013. The information provided stated the product, problem, and actions to be taken by the customers. Contact Technical Support at 855-885-2273 for questions regarding this issue. |  | 46 | 46 | Worldwide Distribution-USA Nationwide-list of countries not available. | Terminated | June 04, 2013 | 11 | 145 | Computer
Z-1404-2013 | 65082 | refractometer, ophthalmic | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | Canon Full Auto Ref-Keratometer RK-F2, Catalog Number 6937B002AA.Product Usage: This device is intended to be used to measure the refractive power and the radius of corneal curvature of the human eye. | Full Auto RefKeratometer | 2 | May 25, 2013 | 2013 | Canon Inc. | Canon Inc. had received 17 claims from our foreignsales companies between October to November2012 that measurement value of RK-F2s refractivepower (SPH value) sometimes indicated abnormalvalue which occurs approximately 5% in probability. | DESIGN: Software Design | Canon initiated recall by sending letters on November 22, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken and contact information. |  | 384 | 384 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Japan, Korea, Singapore and United Kingdom. | Terminated | September 20, 2013 | 118 | 384 | Computer
Z-1407-2013 | 65171 | stimulator, electrical, evoked response | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Natus neurology Dantec KEYPOINT.NET 2.20.Catalog/Part Numbers:9031A070103, 9031A006502, 9033A071103, 9033A072102, 9031A070102,9033A007402, 9033A071102, 9033A072103, 9033A0701, 9033A0711HA,9033A071101, 9031A070101, 9031A006401, 9031A006501, 9033A004701,9031A0032, 9033A0517, 9033A072101, 9031A0031, 9031A0701BV,9033A0044, 9033A0073AA, 9031A0063, 9033A0523, 9033A071143,9033A072143, 9031A006406, 9031A006506, 9031A070106, 9033A0046069033A004706, 9033A007406, 9033A0537, 9033A0544, 9033A071106,9033A072106, 9033A0524, 9033A0711BV, 9033A007401, 9033A0516,9031A006403, 9031A006503, 9033A056403, 9033A0721BV, 9031A006402,9031A006405, 9031A051306, 9031A070105, 9033A0072, 9033A007301,9033A0711AA, 9033A0721AAProduct Usage:Dantec KEYPOINT.NET Software 2.20 is installed on the Keypoint EMG System which is intended as an electro-physiological aid to assess diagnosis and prognosis and to monitor diseases of the central and peripheral nervous system. It can also be used to study functional aspects of nerves and muscles. | Natus neurology Dantec KEYPOINT | 2 | May 25, 2013 | 2013 | Alpine Biomed Aps | Alpine Biomed ApS is recalling Natus neurology Dantec KEYPOINT.NET ver. 2.20 Software due to the presence of a software bug. | DESIGN: Software Design (Manufacturing Process) | Alpine Biomed ApS sent a Recall Notification/ Field Safety Notice dated April 8, 2013 to all affected customers via e-mail and courier. Alpine Biomed ApS advised their customers to upgrade all installations of Keypoint.NET v2.20 to Keypoint.NET v2.21. Alpine BioMed ApS has recommended that their customers complete and return the verification form to QA/RA Engineer Michael Jacobsen at e-mail michael.jakobsen@natus.com. These actions and the return of the verification form to Alpine Biomed ApS should be completed within 3 months of receipt of the recall notifications. If any assistance is required to complete these actions, please contact Technical Support Engineer Michael Jensen via e-mail michael.jensen@natus.com, alternatively QA/RA Engineer Michael Jacobsen by e-mail michael.jakobsen@natus.com. |  | 871 | 871 licenses installed | Worldwide Distribution - US Nationwide and the countries of: Germany, France, Bulgaria, India, Colombia, Tunisia, Saudi Arabia, Finland, Denmark, Brasil, Italy, Estonia, South Korea, Russia, Greece, Thailand, Israel, Poland, Turkey, Slovakia, Portugal, Switzerland, Hungary, Great Britain, Netherlands, China, Indonesia, Australia, Sweden, Spain, Iran, Singapore, Norway, Austria and Czech Republic. | Terminated | January 06, 2014 | 226 | 871 | Computer
Z-1419-2013 | 65136 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment.RayStation is a software system designed for treatment planning and analysis of radiation therapy. | RayStation Therapy Treatment Planning System | 2 | May 28, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. | DESIGN: Software Design | RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.For questions regarding this recall call +46 8 545 06130. | Contact the recalling firm for information | 70 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA including Connecticut, Florida, Massachusetts, Missouri, New York, Ohio, and foreign distribution to(Canada, Italy, Australia, Germany, Netherlands, and Switzerland. | Open |  |  | 97 | Computer
Z-1441-2013 | 65191 | ophthalmoscope, ac-powered | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Optovue iVue with Normative Database (NDB) with Software Version 3.0 and 3.1.Intended for in vivo imaging. | Optovue iVue with Normative Database | 3 | May 31, 2013 | 2013 | Optovue, Inc. | Colors from the NDB comparison for Ganglion Cell Complex thickness in the singular OU report in iVue version 3.0 (US) and 3.1(international) are different from those in the individual iWellness ID and OS reports. | DESIGN: Software Design | Optovue sent a Field Correction Notice letter dated May 16, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 329 | ~329 users | Distributed Nationwide and in Canada. | Terminated | October 24, 2013 | 146 | 329 | Computer
Z-1450-2013 | 64128 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips BuckyDiagnost Radiographic. Catalog Numbers: 704031, 704032, 704035, 704060, 704062This system is used for making x-ray exposures for diagnostics | Philips Bucky Diagnost | 2 | June 17, 2013 | 2013 | Philips Healthcare Inc. | An inspection of IATD (Installation Acceptance Test Document)IA SID (Source Image Distance) records shows table height measurements out of specification with the consequence of potential system mis-alignment | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Philips Healthcare issued an Urgent Field Safety Notification on January 11, 2013, to advised users of issued and perform corrective action to the units. Field Service Engineer (FSE) will execute corrections and attach record to IATD file.Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare.Further questions please call (978) 687-1501. |  | 116 | 116 units | Worldwide Distribution - US Nationwide and the country of Canada. | Terminated | October 15, 2013 | 120 | 182 | Computer
Z-1537-2013 | 65376 | ventilator, continuous, minimal ventilatory support,facility use | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, Model #V60Respironics Material P/N (Philips 12 Digit P/N): 85008 (85008)1053613 (989805628251)1053614 (989805612101)1053615 (989805613391)1053616 (989805613661)1053617 (989805611761)1053618 (n/a)R1053618 (n/a)1076709 (n/a)1076715 (989805627411)1076716 (989805627431)1076717 (989805627441)DU1053617 (989805616411)U1053614 (989805636441)U1053617 (989805636631)Power Mgmt. Board, Respironics Material P/N: 1054358, Philips 12 Digit P/N: 4536561512581.The V60 Ventilator is an assist Ventilator and is intended to augment patient breathing. It is intended for spontaneously breathing individuals who require mechanical ventilation: patients with respiratory failure, respiratory insufficiency, or obstructive sleep apnea in a hospital or other institutional settings under the direction of a physician. The V60 Ventilator is intended to support pediatric patients weighing 20kg (44 lbs.) or greater to adult patients. It is also intended for intubated patients meeting the same selection criteria as the noninvasive applications. It is intended to be used by qualified medical professionals, such as physicians, nurses, and respiratory therapists. | V60 Ventilator | 1 | June 17, 2013 | 2013 | Respironics California Inc | Respironics California, Inc. has initiated a recall on the V60 ventilator because of an issue with the software on the V60 Power Management Board Assembly. If a component fails on the Power Management Board Assembly, it may cause ventilator support to be lost with potentially no audible alarm from the ventilator. | DESIGN: Software Design | The firm, Philips Healthcare, sent a "MEDICAL DEVICE CORRECTION" letter dated June 3, 2013 to all customers who received the V60 Ventilators. The letter describes the product, problem identified and the actions to be taken. The customers were informed that the V60 ventilator may continue to be used in accordance with its directions for use, pending the completion of the update. In addition, the customers were instructed to refer to their manual for additional information on warnings. A Philips Field Service Engineer, Approved Sevice Provider or Distributor will be contacting the customers to schedule a no-cost update and replacement of the Power Management Board Assembly Software on all V60 ventilators shipped from the manufacturer prior to April 1, 2013.If you need any further information or support concerning this issue, contact your local Philips Respironics representative at (800) 722-9377. |  | 19200 | 19,200 units total (12,060 units in US) | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt , Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Kenya, Kuwait, Lebanon, Libya, Malaysia, Malta, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Utd.Arab.Emir. and Vietnam. | Open |  |  | 19200 | Computer
Z-1448-2013 | 64130 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost Stationary Radiographic SystemModels: 712020, 712022, 712050, 712052, 712055, 712057, 712062Product Usage:Universal RAD Applications | Philips Digital Diagnost | 2 | June 17, 2013 | 2013 | Philips Healthcare Inc. | Installation Acceptance Test Document (IATD) non compliant with regulations. Records shows table height measurements out of specification with the consequence of potential system mis-alignment. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Philips Heathcare issued an Urgent - Field Safety Notice letter dated January 11, 2013 to all affected customers to inform them of the systems afffected by this Field Action. The letter identified what the issue is, under what circumstances it may occur, the actions to take to avoid or minimize the occurrence of the issue and the actions planned by Philips to correct the issue. Customers were inform that if they have multiple BuckyDiagnost systems, they will be contacted by a Field Service Engineer to clarify which systems are affected. The Field Service Engineer (FSE) performs a full IATD (Installation PHILIPS Acceptance Test Document) and file the IATD record. Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare. For questions contact your local Philips representative. 1-800-722-9377. |  | 101 | 101 units | Worldwide Distribution - USA Nationwide and the country ofCanada | Open |  |  | 278 | Computer
Z-1551-2013 | 65229 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, PET VCAR.PET VCAR application software is affected if installed on any of the following versions of the Advantage Workstation and AW Server: Advantage Workstation 4.6, Volume Share 5 enhanced  vxtl.11.3, vxtl.11.3 ext1, or vxtl.11.3 ext2 AW Server 2.0, Volume Share 5 enhanced  vxtl.11.3, vxtl.11.3 ext1, or vxtl.11.3 ext2PET VCAR (Volume Computer Assisted Reading) is a PET/CT software package which can be used by the clinician to assist in diagnosis, staging, treatment planning and monitoring treatment response. | GE Healthcare, PET VCAR | 2 | June 18, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the automated PERCIST response criteria calculation applied in the PET VCAR application on Advantage Workstation or AW Server. | DESIGN: Software Design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 14, 2013, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to call one of the following numbers:United Sates 800-437-1171Japan 0120-055-919For other countries they were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 407-475-4717. |  | 231 | 231 (38 US; 193 OUS) | Worldwide Distribution - USA including CA, CO, CT, FL, GA, IL, NY, NY NC OH, TX, VA, WV, WI.Internationally to: AUSTRALIA, AUSTRIA, BAHARIN, CANADA, CHILE, CHINE, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GEORGIA, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, JAPAN, KOREA, MEXICO, MOROCCO, NIGERIA, POLAND, ROMANIA, RUSSIA,SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | January 21, 2014 | 217 | 231 | Computer
Z-1559-2013 | 65438 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echelon, Echelon Oval and Oasis MR imaging device systems; one unit per box. The Echelon, Echelon Oval and Oasis MR systems are imaging devices, and are intended to provide the physician with physiological and clinical information, obtained non-invasively and without the use of ionizing radiation. The MR systems produce transverse, coronal, sagittal, oblique, and curved cross-sectional images that display the internal structure of the head, body, or extremities. | Echelon, Echelon Oval and Oasis MR imaging device systems | 2 | June 19, 2013 | 2013 | Hitachi Medical Systems America Inc | Hitachi discovered a software error that can occur when simultaneously scanning a patient while processing images previously acquired from a different patient. The software error can occur randomly. The error causes an image in the processed dataset to have the same unique identifier (UID) as in the scanning dataset. The processed image is stored in the place of the original image, which is lost. | DESIGN: Software Design | Hitachi sent a Device Correction Notice beginning on January 31,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Software correction is targeted to start by March 1, 2013. All sites will be visited by Hitachi Service to perform corrections. All letters are recorded by tracking number and date received by the customers.For further questions please call 800-800-3106 x 3720. |  | 45 | 45 system units | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, IL, IN, KS, KY, MD, NC, ND, NY, OH, OR, PA, TX, VA, WA, and WI. One (1) system unit was also distributed to Mexico. | Terminated | September 20, 2013 | 93 | 45 | Computer
Z-1560-2013 | 65331 | station, pipetting and diluting, for clinical use | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tecan Freedom EVO with Touch Tools Suite Version 3.0 The intended use of the Touch Tools Suite is to run previously defined scripts and processes on a Freedom EVO platform (Touch Tools Suite is an optional software add-on for Freedom EVOware (the base operational software for the Freedom EVO 2 instrument. | Freedom EVO 2 with TouchTools Suite Version 3.0 | 2 | June 19, 2013 | 2013 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. | Contact the recalling firm for information | 12 | 12 units | Worldwide distribution: US (nationwide) including states of: CA, KY, NC, OH, TX, and WI; and countries of: France, Germany and Italy. | Terminated | June 20, 2013 | 1 | 12 | Computer
Z-1567-2013 | 65379 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850Product Usage: The CARESCAPE TM Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility on adult, pediatric, and neonatal patients and on one patient at a time. The system is indicated for monitoring of Hemodynamics, Airway Gases (Fi/Et CO2, O2, N2) and Anesthetic Agent), Spirometry, Gas Exchange (O2 Consumption (VO2), CO2 production (VCO2), energy expenditure (EE), and respiratory quotient (RQ)) and neurophysiological (including electroencephalography (EEG), Entropy, Bispectral Index (BIS) and Neuromuscular Transmission (NMT) Monitoring) status. It is a stand-alone monitor or interfaced to other devices. is designed to provide configuration and troubleshooting capabilities to qualified service personnel. The CARESCAPE Monitor B850 is not intended for use during MRI. | CARESCAPE Monitor B850 | 2 | June 21, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare become aware of 4 potential issues associated with the CARESCAPE Monitor B850. These issues include CARESCAPE Monitor B850 uDOM failure, CARESCAPE Monitor B850 unstable USB communication, CARESCAPE Monitor B850 remote printing issue with currently printing PRN 50 recorder and CARESCAPE Monitor B850 reboots due to an X server freeze. | DESIGN: Software Design | GE Healthcare sent "Important Product Information" letter dated 12 July 2012 to all affected customers. The letter described the Subject (issues) , Instructions, Affected Product Details, Product Correction and actions to be taken. For questions contact Technical Support 1-800-558-7044 or your local Service Representative. |  | 12779 | 12,779 | Worldwide Distribution - USA Nationwide including the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, IA, ME, MD,MA, MI, MN, MS, MO, MT, NE,NV,NH, NJ, NM, NY, NC, OR, OK OH, PA, PR, SC, TN, TX, UT, VA, WA, WY, WI, WY, and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK,, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLAND, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Open |  |  | 12779 | Computer
Z-1562-2013 | 65435 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE with syngo RT Therapist version 4.3Product Usage:The intended use of the SIEMENS branded ARTISTE family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | SIEMENS ARTISTE | 2 | June 21, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The potential for a safety issue, for the image quality degradation, that is addressed by the Customer Safety Advisory Notice when verifying the patient position based on 2D images that were acquired as single-exposure MV or In-Line kView images using the syngo RT Therapist 4.3. | DESIGN: Device Design | Siemens sent an Urgent: Medical Device Correction Customer Safety Advisory Notice letter to all affected customers. The letter identified the affected product, problem, what preventive measures can the user take and actions to be taken. Customers were advised to perform the preventive measures provided and inform all affected personnel immediately. The letter states that Siemens is currently working on a new version of the software. Customers are asked to include the Customer Safety Advisory Notice in the System Owner Manual chapter "Safety Advisory Letters". For question call 925-602-8157. |  | 8 | 8 units | Worldwide distribution | Terminated | November 26, 2013 | 158 | 8 | Computer
Z-1583-2013 | 65246 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek XS Plus SystemThe CoaguChek XS Plus system for use by professional healthcare providers for quantitative prothrombin time testing for the monitoring of warfarin therapy. | CoaguChek XS Plus System | 2 | June 24, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche has confirmed the potential for an undetected elevated INR test result with the CoaguChek XS, CoaguChek XS Plus, and CoaguChek XS Pro meters. In rare cases, instead of a value, an ERROR 6 message is displayed. | DESIGN: Software Design | Rochet sent an Urgent Medical Device Correction letter dated May 16, 2013, via UPS Ground to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions: If you are using Warfarin therapy in combination with antibiotics and/or chemotherapy drugs, which could potentially lead to an extremely high test result ( | Contact the recalling firm for information | 3343 | 3343 meters | Nationwide Distribution | Open |  |  | 89896 | Computer
Z-1585-2013 | 64789 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. | Centricity Laboratory Core Lab; Versions 3.3, 4.0, 4.1 | 2 | June 24, 2013 | 2013 | GE Healthcare It | If environment variables MICRO_CORRECT_MSG_TYPE and CORRECT_MSG_TYPE are set to different values, in Microbiology result entry when released results (Final Report released) are changed and the Corrected Result messages are edited on the free-text form, some of the originally entered panel free-text is lost. For this issue to occur the original result must be amended and the environmental variabl | DESIGN: Software Design (Manufacturing Process) | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Software correction is available. Please contact your GE service representative to request or install this update. Please contact your GE Service Representative with any question regarding this issue. For other questions please contact GE Helpdesk.Help Desk: North America: 1-888-778-3375 Help Desk: Europe: + 44 1603 877595 Help Desk: Asia Pacific + 61 3 5327 4000. | Contact the recalling firm for information | 9 | 9 Units | Worldwide Distribution - US Distribution including the states of OH, NY, GA, TX, TN and MN and the countries of Australia, Qatar and England | Open |  |  | 9 | Computer
Z-1570-2013 | 65135 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic SynchroMed II, Model 8637, (The SynchroMed II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. Product Usage: The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II infusion system include the pump, catheter, and catheter accessories.The SynchroMed II Programmable Pump is indicated when patient therapy requires the chronic infusion of the following drugs or fluids: Chronic intraspinal (epidural and intrathecal) infusion of preservative-free morphine sulfate sterile solution in the treatment of chronic intractable pain, Chronic intrathecal infusion of preservative-free ziconotide sterile solution for the management of severe chronic pain, Chronic intrathecal infusion of Lioresal Intrathecal (baclofen Injection ) in the management of serve spasticity, Chronic Intravascular Infusion of Floxuridine (FUDR) or methotrexate for the treatment of primary or metastatic cancer. | SynchroMed II Implantable Infusion Pump | 1 | June 25, 2013 | 2013 | Medtronic Neuromodulation | Medtronic is providing Healthcare Professionals with important safety information and patient management recommendations regarding the unintended delivery of drug during the priming bolus function for the SynchroMed implantable infusion pump. The unintended delivery of drug can contribute to patient overdose or underdose symptoms which may be clinically relevant. | DESIGN: Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the background and nature of the issue and provided recommendations for patient management and monitoring after initiation of Intrathecal Therapy. The letter directed customers to continue to use the priming bolus precedure; to monitor all patients following start or restart of intrathecal therapy; consider priming the pump prior to implant; and educate caregivers and family members to recognize the signs and sypmtoms associated with intrathecal drug therapy complications. For questions contact your local Medtronic field representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. | No consumer action necessary | 261109 | 261,109 total devices estimated implanted Worldwide | Worldwide Distribution - USA Nationwide | Open |  |  | 261109 | Computer
Z-1588-2013 | 65137 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 64, Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are whole body Computed Tomography X-ray systems. Produce cross-sectional images of the body by computer reconstruction of x-ray transmission. | The Brilliance 64, Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners | 2 | June 25, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the operator selects a patient name from the worklist using the mouse and then selects a different patient name using the keyboard (arrow keys then press enter) the patient name that populates the exam information page is the patient name selected with the mouse. Also, it is possible to begin a normal clinical scan while the anti-virus scan is running. However, if the user is running a bol | DESIGN: Software Design (Manufacturing Process) | Philips Medical Systems sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated Mar 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 50 | 50 units | Nationwide Distribution-including the states of AR, AZ, DE, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, and WY. | Terminated | November 27, 2013 | 155 | 50 | Computer
Z-1579-2013 | 65443 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic SynchroMed II, Model 8637, (The SynchroMed II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II infusion system include the pump, catheter,r and catheter accessoriesThe SynchroMed II Programmable Pump is indicated when patient therapy requires the chronic infusion of drugs or fluids: | SynchroMed II Implantable Infusion Pump | 1 | June 25, 2013 | 2013 | Medtronic Neuromodulation | Medtronic Neuromodulation is providing Healthcare Providers with safety information and patient management recommendations related to the potential for electrical shorting internal to they SynchroMed infusion pump. An electrical short circuit in a feedthrough may present as a motor stall or low battery reset/alarm and lead to a loss of or reduction in therapy which may result in the return of und | DESIGN: Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the issue of the pump's potentially shorting, how to identify if a pump is affected, and provided ongoing patient management recommendations. The Medtronic Rep would complete a Physician Notification Confirmation Form. A copy of the recall notification letter and the completed Confirmation Form is to be left with the physician. Customers with questions were instructed to call 1-800-707-0933.For questions regarding this recall call 763-526-9738. | Contact the recalling firm for information | 261109 | 261,109 total devices estimated implanted Worldwide | worldwide | Open |  |  | 261109 | Computer
Z-1599-2013 | 65407 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Multilite Dye Handpieces, an optional accessory of the RevLite SI Laser.The Revlite SI laser is intended for General Dermatology, Dermatologic and General Surgical Procedures. Which includes treatments such as tattoo removal, pigmented lesion removal, removal or lightening of hair, skin resurfacing and reduction of red pigmentation in hypertropic and keloid scars The two Multilite Dye Handpieces (HPs) are optional handpieces that can be used with the Revlite SI Laser. Manufactured by:ConBio, A Cynosure Company47733 Fremont blvdFremont, CA 94538 | RevLite SI Laser | 2 | June 26, 2013 | 2013 | ConBio, a CynoSure Company | An inconsistency was identified related to electronic display identification of the optional Multilite Dye Handpieces of the RevLite SI Laser system that may confuse user. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Cynosure Inc., sent a "MEDICAL DEVICE FIELD ACTIONS NOTIFICATION" letter dated May 9, 2013 to its customers via UPS next day. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the Revlite SI Multilite Dye Handpieces upon receipt of this letter and complete and return the enclosed Field Corrective Action Acknowledgement Form via e-mail to: mflynn@cynosure.com or FAX to: 978-256-6556. A Cynosure representative will be contacting customers to schedule and arrange an exchange of the affected handpieces.If additional information is required please contact VP of Regulatory Affairs and Quality Systems at 978-367-2350 or via email at ikulinets@cynosure.com. | Contact the recalling firm for information | 12 | 12 (6 of model 659-1050; 6 of model 659-1100) | Worldwide distribution: USA (nationwide) including states of: IL, ME, NM and VA; and countries of: Australia and Chile. | Terminated | June 26, 2013 | 0 | 12 | Computer
Z-1597-2013 | 65361 | monitor, bed patient | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Code Alert Advanced 4-Way Care Solution, CA520 System,PN 9600-1600, RF Technologies. This product is used to alert caregivers when a patient at risk for falls is leaving a bed or chair. | Code Alert, Sensatec | 2 | June 26, 2013 | 2013 | RF Technologies, Inc. | RF Technologies, Inc. has initiated a recall involving CA520 Code Alert Integrated Care Management Advanced 4-Way Care Solution . Due to a firmware error, the alarm may not be transmitted to the system server immediately. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent on 5/17/2013 a "Medical Device Recall" letter dated 3 May 2013. The letter described the problem and the product involved in the recall. The letter advised consignees to discontinue use of the device and return to RF Technologies. |  | 167 | 167 | Nationwide Distribution in US including the states of AK, AZ, CA, GA, IA, KY, MT, NE, NJ, ND, and WI. | Terminated | December 03, 2013 | 160 | 167 | Computer
Z-1616-2013 | 65022 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 3100 and Innova 4100 systems.The Digital Fluoroscopic Imaging System is indicated for use in diagnostic and interventional angiographic procedures of human anatomy. It is intended to replace image intensifier fluoroscopic systems in all diagnostic or interventional procedures. | GE Healthcare, Innova 3100 and Innova 4100 systems | 2 | June 28, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare reported a potential safety issuedue to temporary loss of X-Ray Imaging Modes during Collimator Auto-Calibration which affects Innova 3100 and Innova 4100 Systems. | DESIGN: Software Design | A Customer Letter was not originally sent. GE Healthcare has retrospectively re-assessed the reportability of the field action and has determined that the action should be reported to the FDA.GE Healthcare has already updated all of the affected Innova 3100 and Innova 4100 systems that were distributed.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests have recorded the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion have been retained per their processes and procedures.If you have any questions, please call 262-548-2757.For questions regarding this recall call 262-513-4122. | Contact the recalling firm for information | 48 | 48 | Worldwide distribution - USA including AZ, CA, CO, FL, ID, IL, ME, NV,NY, NC, OK, TN, UT, and WI; and Internationally to Canada, Chile, France, Germany, Italy, Japan, Korea, Norway and Spain. | Open |  |  | 48 | Computer
Z-1652-2013 | 65410 | transmitters and receivers, physiological signal, radiofrequency | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-01 used with ER920 Wireless Series Arrhythmia Event Recorder.Non-rechargeable disposable battery pack. | Lithium Thionyl nonrechargeable disposable battery pack | 2 | July 02, 2013 | 2013 | BRAEMAR, INC. | Braemar has recently became aware of a battery related incident that occurred with EWT battery packs that are used in a Braemer ER920W wireless event monitor and Fusion device. The event did not result in any patient related injury or present any clinical impact. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Braemar began notifying customers by telephone on 2/24/2012. Braemar issued Advisory Notice letter P/N 600-066201 RevA on 3/02/2012. The letter described the product, the problem, and the action to be taken by the customer. Customers were requested to only use TESTED batteries that they have receivewd fromBraemar starting on February 23, 2012. Complete and return the attached Verification Form. If customers have any untested product in their inventory, they were instructed to segregate those battery packs and await further instruction from Braemar to coordinate the return of those untested units. Upon receipt of the attached verification form, Braemar will continue to send replacement battery packs that have passed the additional testing protocol. Customers were instructed to contact their Braemar sales representative if they have any questions regarding the notice.For questions regarding this recall call 651-286-8620. |  | 3286 | 3286 packs | Nationwide Distribution including FL, IL, LMD, MA, MN, NY, OH, PA, TN, and TX. | Terminated | July 15, 2013 | 13 | 3286 | Computer
Z-1645-2013 | 65423 | transmitters and receivers, physiological signal, radiofrequency | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-00 used with ER920 Wireless Series Arrhythmia Event Recorder. Non-rechargeable disposable battery pack. | Lithium Thionyl nonrechargeable disposable battery pack | 2 | July 02, 2013 | 2013 | BRAEMAR, INC. | Braemar Inc., became aware of a battery related incident that occurred with an Able battery pack being used in a Braemar ER920W wireless event monitor. The event did not result in any patient related injury or present any clinical impact. | PRODUCTION CONTROLS: Process Control | Braemar began notifying customers by telephone on 5/2/2011. Braemar issued Advisory Notice letter P/N 600-0662-00 RevA dated May 16, 2011. The letter was sent by UPS to all consignees.The letter described the problem and the product involved in the recall. Consignees were advised not to use the battery pack in the Braemar ER920W or Braemar Fusion device and requested consignees to complete and return the Verification Form.For questions regarding this recall call 651-286-8620. |  | 10785 | 10,785 packs | Nationwide Distribution including FL, IL, MD, MN, NY, OH, PA, SC, TN, and TX. | Terminated | July 15, 2013 | 13 | 10785 | Computer
Z-1646-2013 | 65139 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray systems Product Usage: The Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 scanners are whole body Computed Tomography X-ray systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. | Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 scanner | 2 | July 02, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 systems running software versions 2.6.1 or 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images, which may impair the diagnostic quality of affected images. | DESIGN: Software Design | Philips Healthcare sent an Urgent Medical device Correction letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that Philips Healthcare is implementing software update versions 2.6.2 or 3.5.5 to resolve the issue and a Philips Field Service Engineer will contact each site to schedule the software installation. For questions contact your local Philips representative or local Philips Healthcare office. | Contact the recalling firm for information | 2348 | 2348 total units: 2182 Brilliance CT 64 units; 95 Ingenuity Core units; and 71 Ingenuity Core128 units. | USA Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. | Open |  |  | 2348 | Computer
Z-1649-2013 | 65207 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System (product code 6802413) Product Usage: For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Products Slides, VITROS Chemistry Products MicroTip Reagents and VITROS Immunodiagnostic Products Reagents. | VITROS 5600 Integrated System | 2 | July 02, 2013 | 2013 | Ortho-Clinical Diagnostics | There is evidence that an electrical short within the systems connector has caused an accumulation of heat which has radiated back through the wiring harness causing the wires and connectors to melt. | DESIGN: Device Design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letters dated May 6, 2013 and response forms to all affected customers. Ortho Clinical Diagnostics informed their customers of the issue and advised if they detect the smell of melting plastic or observe smoke originating from your VITROSÂŽ System, to immediately power down the system and contact OCD for assistance. They were also advised to consider performing the backup procedure more frequently to help prevent the loss of data on their Systems. In the near future, an OCD Field Engineer will contact customers who are affected and perform a Modification (MOD) procedure that will replace their potentially affected Wire Connector(s). Foreign affiliates were informed by email on May 6, 2013 of the issue and instructed to notify their customers of the issue and actions. OCD also found that this wire connector is commonly used in many types of computers (i.e., medical devices, personal computers, etc.). On 29 May 2013, OCD expanded their recall to incllude ORTHO Assay Software (OAS) Servers. As a precaution, OCD sent notifications to encompass all ORTHOÂŽ Assay Software (OAS) Servers that potentially contain this type of wire connector. The letters (Ref. CL13-154) were sent via US Priority Mail to all customers who were shipped an affected OAS Server. Ortho Clinical Diagnostics has requested that customers with any questions or in need of additional information, please call Customer Technical Services at 1-800-421-3311. |  | 381 | Total 381 units ( 227 units domestically & 154 units internationally) | Worldwide Distribution - USA Nationwide including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI & WV, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open |  |  | 513 | Computer
Z-1598-2013 | 65049 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | TOSHIBA Kalare Diagnostic X-Ray System K110785 | TOSHIBA Kalare XRay System | 2 | July 10, 2013 | 2013 | Toshiba American Medical Systems Inc | Firm initiated recall because when radiographic conditions with the small focus are set and the radiography ready switch is held down, the large focus is also heated after a few seconds or more. This occurs when the Automatic Exposure Function (AEC) function is used. As a result, a radiographic tube current exceeding the set value flows and a brighter image than expected is acquired. | DESIGN: Component Design/Selection | Firm initiated recall by sending recall notification letters along with a customer reply form on April 3, 2013 via USPS return mail receipt. A follow-up letter (with customer reply form) was sent on April 11, 2013 to inform customers that they should continue to use the AEC function as usual. The updated letter along with the customer reply form dated April 11, 2013 was also sent via USPS return mail receipt. The letter titled, "URGENT: MEDICAL DEVICE CORRECTION", notified customers of the reason for recall, product description under recall, corrective action, instructions to customer under 1. A new circuit board will be installed in each involved unit by a Toshiba service representative, without any charge to the user.2. In the meantime while waiting for service, the user is to continue usingthe device with the AEC function. If the AEC function is not available, theuser should push the radiography READY and EXPOSURE switches atthe same time."Note to Customer", and contact information "(800) 421-1968 or your local Toshiba Representative at (800) 521-1968". |  | 153 | 153 units | US Nationwide Distribution | Open |  |  | 153 | Computer
Z-1730-2013 | 60480 | monitor, blood-gas, on-line, cardiopulmonary bypass | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe.Intended for use during cardiopulmonary bypass procedures when continuous monitoring of arterial and/or venous pH, PCO2, PO2, K+, oxygen saturation, hematocrit, hemoglobin, and temperature are desired. | CDI 500 Monitor | 2 | July 13, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that CDI 500 Blood Parameter Monitoring System software contains inaccuracies associated with a software algorithm, and a software update has been developed to address this issue and a premarket notification (K130425) has been submitted to the agency. Once clearance is received for this premarket notification from the US FDA, Terumo CVS will implement the | DESIGN: Software Design | On 2/27/13, all consignees will receive a recall notification letter via Federal Express. Terumo CVS will correct all affected CDl Monitors. Correction to the devices will be completed concurrently with BPM replacements related to Z-2185/2189-11 RES 57194.Terumo continues to manufacture and ship devices affected by this recall. The long term corrective actions are associated with Recall# Z-2185/2189-11 RES 57194. As committed previously, we will re-notify all customers once we are ready to implement the fix for the BPM in the field.All units produced prior to the change in manufacturing will be included in thescope of this correction. Terumo continues to manufacturer units affected by this correction.All units returned for this correction will be tagged with a service traveler to indicate reason for correction and units will be properly segregated to prevent unintended re-shipment. |  | 3972 | 3972 units since 4/26/12 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Jordan, CANADA, AUSTRALIA, Japan ,Hong Kong ,UNITED ARAB EMIRATES (UAE), BELGIUM, South Korea, Malaysia, Mexico, Philippines, Singapore, Taiwan, and Thailand. | Open |  |  | 3972 | Computer
Z-1738-2013 | 63109 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Eon - 3701 (Wall Charger) and 3711 (Portable Charger). Product Usage:Eon: Eon system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. Eon Mini: Eon Mini system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome, intractable low back and leg pain. | Eon Charging System | 2 | July 16, 2013 | 2013 | St. Jude Medical | St. Jude Medical has informed your doctor that a number of cases have been reported in which discomfort associated with heating occurred at the device site while patients were using the Charging System to charge their spinal cord stimulator. In three cases, patients received a burn to the skin (one second-degree and two first-degree burns) at the implant site. As of June 30, 2012, St. Jude Medica | DESIGN: Device Design | ST. Jude Medical sent an Important Medical Device Safety Information dated July 26, 2012 to affected customers. The letter identifed the affected product, problem and actions to be taken. For questions or concerns contact your ST. Jude Medical representative or St. Jude Medical Neuromodulation Division at 1-866-240-6741. |  | 70638 | 70,638 (combined total for Eon and Eon mini systems) | Worldwide Distribution - USA Nationwide Including Puerto Rico and countries of: Argentina, Australia, Belgium, Brazil, Canada, China, Chile, Colombia, Ireland, Israel, India, Malaysia, Mexico, Saudi Arabia, Singapore, South Africa, South Korea, Switzerland and United Kingdom. | Open |  |  | 70638 | Computer
Z-1742-2013 | 65456 | drill, surgical, ent (electric or pneumatic) including handpiece | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | RemB Electric Universal Driver REF 6400-099-000 For use with the CORE and TPS System sRx OnlyThe RemB Electric Universal Driver is intended for use with the Consolidated Operating Room Equipment (CORE TM) System. When used with a variety of accessories, the RemB Electric Universal Driver is intended for surgical procedures involving driving wire or pins into bone and hard tissue. This includes but is not limited to Dental, ENT (Ear, Nose, Throat), Neuro, and Endoscopic applications. This device can also be used with the Total Performance System (TPSTM). | REMB Universal Driver | 2 | July 16, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The safety margin values detailed in the Engineering Design were entered incorrectly into the programming software. Handpieces that have been programmed with the incorrect parameters could result in unintended activation. There is a potential for injury to the patient or user as a result of unintended forward or reverse activation. There is also the potential that the handpiece may not reach 100 | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm, Stryker Instruments, sent emails beginning May 30, 2013 "URGENT MEDICAL DEVICE CORRECTION" letter dated July 1, 2013 to its customers. The letter identified the device, the reason for the recall, possible adverse events and actions to be taken. The customers were instructed to immediately review this letter, locate the units listed, complete and return the business reply form via FAX to 866-521-2762; or email Kara.spath@stryker.com; if you have further distributed this product, please forward this letter and the attached Business Reply Form to all affected locations. After receiving this form, Stryker will send you a pre-paid shipper to send your affected device(s)back to Stryker to be reprogrammed with the correct specifications. Send back all affected devices using the pre-paid shipper provided to you by Stryker. If you require a loaner, please contact the Stryker Service Department to arrange for a loaner at 888-308-1983. For questions regarding this recall please Stryker Instruments: Monday-Friday 8am-5pm (EST) 269-389-4518 / 269-389-3808 kara.spath@stryker.com / jennifer.mars@stryker.com. |  | 9 | 9 | Worldwide distribution: US (nationwide) including states of: AZ, FL, GA, NC, NY, SC, TX and VA; and countries of: India, Netherlands and Japan. | Terminated | September 16, 2013 | 62 | 20 | Computer
Z-1750-2013 | 65620 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zee/zeego systems equipped with generator A100 PlusProduct Usage:Angiographic x-ray system | Siemens AXIOM Artis zee/zeego systems equipped with generator A100 Plus | 2 | July 17, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the course of product monitoring, Siemens became aware of a tolerance problem in the power supply of the generator control above a specific value, thus, the generator may fail. | DESIGN: Device Design | Siemens sent an Urgent Customer Information letter to their consignees. The letter identified the affected product, problem , update Instructions and actions to be taken. Siemens began to equip affected systems with a modified board assembly in January, 2013. |  | 23 | 23 | United States Nationwide Distribution - including the states of: CA, CO, FL, IA, IL, IN, MA, MD, MI, NC, NY, PA, TX, VA and WA and Puerto Rico. | Open |  |  | 23 | Computer
Z-1743-2013 | 65652 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDHistology/CytologyClassification name: Calculator/data processing module for clinical use (862.2100). Intended use: Laboratory Information System (L.I.S) | TDHisto/Cytology | 3 | July 17, 2013 | 2013 | Technidata S.A. | A recall was initiated by Technidata because use of symbols (less than or equal to or greater than or equal to or any other symbol) in Microsoft Word reports when using electronic transmissions of results by of HPRIM (including HPRIM Medecin via Kermit) or HL7 must be absolutely prohibited. Such symbols are not supported by the Microsoft Word interface used by the Histology/Cytology software and | DESIGN: Software Design | TD Technidata Medical Software sent a Medical Device Recall letter dated June 20, 2013, to one affected customer who received the TD-Synergy Histology/Cytology module. The customer was informed of the problems identified and the actions to be taken. Customers were instructed to complete the attached Reply Form and return it to the firm by fax to 514-270.2770. The customer were also instructed to contact Technidata Canada at (514) 270-7777 or at support-canada@technidata-web.com. |  | 1 | 1 site | US Distribution CA only. | Terminated | July 29, 2013 | 12 | 1 | Computer
Z-1753-2013 | 65176 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hydra Vision DR 60/80 System, Models 700539 and 700540, Liebel-Flarshiem Company. The System is used in radiologic imaging procedures. | Hydra Vision DR 60/80 System | 2 | July 18, 2013 | 2013 | Leibel-Flarsheim Company LLC | A complaint was received in which a hospital physicist indicated the dosage was too high on the HUT system. During the service call the systems Maximum Fluoro Dose was found to be 30 R/min which is three times the Maximum Fluoro Dose limit of 10 R/min. | TRAINING: Employee Error | The customers received an "Urgent Device Correction" notification letter on May 12, 2013 explaining the product issue and providing recommended actions. Customers were contacted to set up a service call to test and recalibrate equipment as necessary. Questions were directed to 800-778-7898. |  | 12 | 12 units | Nationwide Distribution (US), including the states of TN, KY, CO, KS, CT, HI, and AZ. | Terminated | July 26, 2013 | 8 | 12 | Computer
Z-1752-2013 | 65215 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix 1200 and 2400 Compounders.Used as an automated pumping system that compounds multiple sterile ingredients into a finished solution. | ExactaMix 1200 and 2400 Compounders | 2 | July 18, 2013 | 2013 | Baxa Corporation | Baxter Healthcare is recalling certain ExactaMix automated compounding devices due to the complementary metal oxide semi-conductor battery failure resulting in resetting the internal clock of the device to midnight December 31, 2010. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 6/04/13, customers were notified by 1st class mailing with mitigating instructions including reminding customers that the battery failure results in a failure of the ExactaMix system to alert the user to replace the disposable set after 24 hours of use. Users should not rely on the alert to trigger the daily replacement of disposable sets. Users should also confirm the correct time and date at the setup of each use. If the compounder displays an incorrect time and date, customers are to contact Baxter Technical Support. |  | 753 | 753 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada, Latin America, Europe/Middle East/Africa, and Taiwan. | Open |  |  | 753 | Computer
Z-1771-2013 | 65630 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs ARKON Anesthesia Workstation. Spacelabs ARKON Anesthesia Workstation may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. | Spacelabs ARKON Anesthesia Workstation | 2 | July 19, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs ARKON Anesthesia System is recalled due to software defect. The System Status Computer may incorrectly determine that a communication error has occurred. If this situation occurs, a buzzer will activate and a yellow triangle will be displayed on the System Status Computer screen that alerts the user of this error. | DESIGN: Software Design | Spacelabs Healthcare sent an Urgent Medical Device Correction" letter, dated June 20, 2013, to all affected consignees. The letter identified the product, th problem, and the action to be taken by the consignee. Consignees were instructed to advise their staff of the situation. Consignees were also advised that Spacelabs Field Service personnel would be contacting them at their facility to schedule installation of a software update to resolve the issue at not cost to the consignee. Consignees with questions were instructed to call Technical Support at 1-800-522-7025 and select option 2.For questions regarding this recall call 800-522-7025. |  | 9 | 9 units (4 in the US and 5 outside US) | Worldwide Distribution - USA including South Carolina and internationally to Canada and Great Britian. | Terminated | August 29, 2013 | 41 | 9 | Computer
Z-1774-2013 | 65688 | laser, ophthalmic | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PASCAL and PASCAL Streamline (532 nm.577 nm);Treats the patient's eye. | PASCAL and PASCAL Streamline | 2 | July 19, 2013 | 2013 | Topcon Medical Laser Systems, Inc | When using the "Favorites" function for saved settings, there is potential for a mismatch between displayed power and the actual power output. This may result in over or under treatment. | DESIGN: Software Design | Initial notification was sent on March 27, 2012 to all TMLS Distribution Partners and Sales Representative. This was followed by a follow up notification on June 4, 2012. A final notice was sent on February 11, 2013 to all those who had not sent documentation that the field action had been implemented. |  | 40 | 7 US, 33 foreign | Worldwide Distribution-USA and the countries of China, Hong Kong, Brazil, France, Japan, Netherlands, Pakistan, Malaysia, United Kingdom, South Korea, Thailand, Dubai, and South Africa. | Terminated | July 19, 2013 | 0 | 40 | Computer
Z-1773-2013 | 65565 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. This software algorithm analyzes the patient's electrocardiographic (ECG) rhythmand indicates whether or not it detects a shockable rhythm. The LIFEPAK 500 AED is intended for use by personnel who are authorized by a physician/medical director and is intended for use in the pre-hospital and workplace and community environments. | LIFEPAK 500 Automated External Defibrillator (AED) | 2 | July 19, 2013 | 2013 | Physio Control, Inc. | The LIFEPAK 500 Automated External Defibrillator (AED) could be susceptible to malfunction due to an assembly error. If affected, the malfunction would be the result of internal shorting on a printed circuit board assembly and could render the device inoperable. | PRODUCTION CONTROLS: Process Control | The firm, Physio-Control Inc., contacted their consignees by telephone or email starting on June 13, 2013 and sent and/or delivered an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 2013 to its customers. The letter described the product, problem and actions to be taken. A Physio-Control representative will be contacting the customers by direct visit, telephone, or email you to coordinate the replacement of their device(s).Customers with questions regarding this notification, please call Physio's Technical Support at 1-800-442-1142, 6:00 a.m. to 4:00. pm. (Pacific), Monday-Friday. |  | 139 | 139 total (125 units in the US and 14 units outside US) | Worldwide distribution: US (nationwide) including states of: IL, MA, NC, ND, NY, PA, RI, TX, WA, and WI; and country of: Canada. | Terminated | September 19, 2013 | 62 | 139 | Computer
Z-1780-2013 | 65622 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Radiohead Access Point 1.4 GHz Radiohead Access Point Firmware versions: C.00.04 & C.00.05 Part Number: 989803171211Product Usage:IntelliVue Smart-hopping 1.4 GHz Access Point provides wireless connectivity between IntelliVue devices (listed below) and the Philips IntelliVue Information Center. MX40 Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults and pediatrics in hospital environments and during transport inside hospitals. The MX4O is to be used primarily as a traditional telemetry medical device. It connects to the IntelliVue Information Center by way of a wireless network. | Philips Healthcare IntelliVue Smarthopping 1.4 GHz Access Point | 2 | July 20, 2013 | 2013 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of severalminutes impacting wireless connectivity | DESIGN: Software Design | Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377. |  | 20105 | 20105 units | USA Nationwide Distribution | Open |  |  | 20105 | Computer
Z-1782-2013 | 65651 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed.Angiographic x-ray system | ARTIS zee and zeego | 2 | July 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui |  | Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: 1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. | Contact the recalling firm for information | 29 | 29 | Nationwide Distribution and Puerto Rico | Open |  |  | 29 | Computer
Z-1797-2013 | 65468 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx Monitor/Defibrillator,Model numbers M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips HeartStart MRx Monitor/Defibrillator, | 2 | July 24, 2013 | 2013 | Philips Healthcare Inc. | The MRx may deliver a non-synchronized cardioversion shock resulting in the delivery of incorrect therapy, which may induce ventricular fibrillation. | DESIGN: Software Design | Philips sent an Urgent Medical Device Correction letter on June 19, 2013, to all affected customers. The letter informed customers of the issue, identified details of the units affected, gave instructions on actions to be taken by the customer and identified what action Philips planned to take to remedy the issue.Philips will perform a software upgrade and replace the label on the unit that outlines the steps for Synchronized Cardioversion with a revised label for all affected units at no charge to the customer.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving the software upgrade and device label revision, provided that you follow the instruction provided.Customers with questions were instructed to contact their local Philips representataive or call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 66150 | 66,150 units | Worldwide Distribution - USA (nationwide) and countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MONGOLIA MOROCCO NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM and YEMEN. | Open |  |  | 66150 | Computer
Z-1792-2013 | 65579 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM/CT 670 dual detector free-geometry integrated nuclear imaging system. Product Usage: The GE Discovery NM/CT 670 system is a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy. | Discovery NM/CT 670 | 1 | July 24, 2013 | 2013 | GE Healthcare, LLC | GE became aware of an incident at a VA Medical Center facility in the US in which a patient died due to injuries sustained while being scanned on an Infinia Hawkeye 4 Nuclear Medicine System. UPDATE: On July 03, 2013 GE notified hospitals that they are recalling several Nuclear Medicine Imaging Systems because serious injuries or deaths could occur due to the failure mode associated with this re | DESIGN: Device Design | GE sent an Urgent Medical Device Correction letter dated June 17, 2013 to all affected customers. The letter identified the affected product, recommended that qualified service personnel maintain the equipment and that Preventative Maintenance procedures are executed according to labeling. Also the Safety Chapter Sections should be re-reviewed with personnel to ensure proper operation of the equipment. GE notified customers again on July 03, 2013, via Urgent Medical Device Recall letter (including confirmation of delivery for US customers) and follow-up telephone call. Customers were instructed to cease using the affected system until GE Healthcare Field Engineer can complete an inspection of the system and perform any necessary repairs at no cost. A GE Healthcare representative will contact the hospitals to arrange for the inspection. | Contact the recalling firm for information | 213 | 213 | Worldwide Distribution - US Nationwide including Guam, Puerto Rico and the countries of : Afghanistan, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Vietnam. | Open |  |  | 3454 | Computer
Z-1814-2013 | 65666 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. | ED PulseCheck | 2 | July 29, 2013 | 2013 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. |  | 35 | 35 | Nationwide Distribution, including the states of: AK, AR, AZ, CA, CO, DC, DE, FL, GA, ID, IN, MA, MD, MO, NH, NJ, OH, OR, SC, TN, WA, and WV. | Open |  |  | 35 | Computer
Z-1834-2013 | 65440 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer with Automation Ready Connection (All Software Versions), Model Numbers AU5811, AU5821, AU5831 and AU5841. This system is for in vitro diagnostic use only. | AU5800 Clinical Chemistry Analyzer | 2 | July 31, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a software error on AU5800 analyzers with the Automation Ready (AR) Sampler Kit. This software error results in the failure of the concentrated wash solution pump to stop, creating an overflow or leak on each analyzer unit. | DESIGN: Software Design | An Urgent Product Correction letter dated 5/29/13 was sent on the week of June 3, 2013 to all the customers who received the AU5800 Clinical Chemistry Analyzer with Automation Ready (AR) Connection (All Software versions). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to complete and return the enclosed Response Form within 10 days and to share the information with their laboratory staff. Customers with questions are instructed to contact the 24 hours Call Center (Hotline) at (800) 854-3633 in the US and Canada. | Contact the recalling firm for information | 76 | 76 units | Worldwide Distribution, including Nationwide (US) and the countries of China, Italy, Japan, Canada, Turkey, United Kingdom, Singapore, Israel, and the Netherlands. | Open |  |  | 76 | Computer
Z-1833-2013 | 65755 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM.The device is a general purpose ultrasound system. Specific clinical applications remain the same as previously cleared: Fetal/OB; Abdominal (including GYN, pelvic and infertility monitoring/follicle development); Pediatric; Small Organ (breast, testes, thyroid etc.); Neonatal and Adult Cephalic; Cardiac (adult and pediatric); Musculo-skeletal Conventional and Superficial; Peripheral Vascular; Transvaginal; Transrectal; and Intraoperative (abdominal, PV and neurological). | GE Healthcare, VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM | 2 | July 31, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware that during use of the Voluson E6 or E8 systems, that an error message Out of Memory can be displayed which could cause the system to freeze. This is an intermittent behavior that also requires a certain workflow. This issue can be resolved by rebootingthe system. | DESIGN: Software Design | No consignee letter was sent out.Starting on July 13, 2012, Field Engineers visited consignees and carried on instructions as per Field Modification Instruction (FMI 79044A). The Engineer returned the system to the customer and ensure acceptance. Before leaving the customer site have the staff operate the system. Ensure that: all settings are customized, all peripherals are functioning and all communications are working properly. We plan to update the software on all of the affected GE VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM that were distributed.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.Further questions please call 262-513-4122. | Contact the recalling firm for information | 375 | 375 | Worldwide Distribution - US Distribution including the states AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, LA, MD, MI, MN, MO, NE, NV, NM, NY, NC, OH, OR, PA, SC, SD, TN,TX, VA, WA and WI., and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRIL, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF MOLDOVA, NETHERLANDS, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, KINGDOM OF SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA. | Open |  |  | 375 | Computer
Z-1837-2013 | 65498 | exerciser, powered | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Omnicycle Elite; US Importer and Distributor: Accelerated Care Plus, Reno NV. Manufactured by Medica Medizintechnik GmbH, Hochdorf, Germany. The product is indicated for powered exercise equipment intended for medical purposes. | Omnicycle Elite | 2 | July 31, 2013 | 2013 | Accelerated Care Plus Corporation | An anomaly in the firmware that may cause Omnicycle Elite to operate at a speed other than what was previously set when in reverse mode. |  | Letters were issued on June 12, 2013 that included a memory stick with self loading software so consumers can update the device. Letters were sent by trackable method. Customers with questions or concerns should call Customer Support at 800-350-1100. | Contact the recalling firm for information | 1001 | 1001 units | Nationwide Distribution. | Terminated | October 23, 2013 | 84 | 1001 | Computer
Z-1859-2013 | 65603 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Information System (sDIS) versions 9.5 and VA10A.The syngoÂŽ Dynamics system is intended for the acceptance, transfer, display, storage, archive, and manipulation of digital medical images, including quantification and reportgeneration.syngoÂŽ Dynamics is not intended to be used for reading mammography images. | syngo Dynamics | 2 | August 01, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens initiate this recall due to a potential issue when using the sDIS matching rules of the syngo Dynamics Information System version 9.5 or VA10A with modality device configured to end study at association close. Images or DICOM SR objects sent from a modality device to syngo Dynamics may not be saved when the modality is configured at syngo Dynamics to end study at association close, and whe | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notice dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) With sDIS implemented, use the DICOM Modality List (DMWL), without sDIS matching rules enabled OR 2) Do not implement sDIS. Update instructions were also provided. The reported issue will be resolved in a modification to syngo Dynamics version 9.5 and syngo Dynamics VA10A which will be released in the near future.We regret any inconvenience that this may cause , and we thank you in advance for your understanding.For further questions please call (610) 219-6300. |  | 9 | 9 | US Distribution only including the states of PA, OH, UT, CA, NY, MO, and TX. | Open |  |  | 9 | Computer
Z-1847-2013 | 65610 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Hospira Symbiq Infusion System is intended for the delivery of fluids, solutions, drugs, agents, nutritionals, electrolytes, blood and blood products via parenteral, enteral, intravenous, intra-arterial, subcutaneous, epidural, or irrigation routes of administration. | 1) Symbiq One Channel Infuser, 2) Symbiq Two Channel Infuser | 2 | August 01, 2013 | 2013 | Hospira Inc. | The devices secondary audio signal fails to activate. | PRODUCTION CONTROLS: Process Control | Hospira sent a Urgent Device Field Correction letter dated February 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. In addition, letters were sent by e-mail to ECRI, NRAC, Noblis and ISMP for notification of their members and posting on their websites February 5, 2013. The letter was also posted on the Hospira website February 4, 2013.Customers were requested to return a reply form indicating they have received the letter and notified their users.For further inquiries please contact Hospira Global Complaint Management at1-800-441-4100 ( 8am - 5pm CST, M-F) to report adverse events or product complaints, or Hospira Advanced Knowledge Center at 1-800-241-4002, option 4 for additional information or technical assistance. ( 24hrs a day / 7 days a week. | Contact the recalling firm for information | 2791 | 2,791 units | US Distribution including the states of AL, CA, CT, DE, FL, HI, IA, IL, IN, LA, MA, MD, ME, MI, MN, NC, NJ, NY, PA, SD, TX and WA. | Open |  |  | 2791 | Computer
Z-1849-2013 | 65759 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ e Ultrasound System.The LOGIQ i/e and Vivid e is a general purpose ultrasound system intended for use by a qualified physician for evaluation by ultrasound imaging or fluid flow analysis of the human body. Specific clinical applications and exam types include: Fetal/OB; Abdominal (GYN & Urology); Pediatric; Small Organ (breast, testes, thyroid); Neonatal and Adult Cephalic; Cardiac (adult & pediatric); Peripheral Vascular; Intra-operative (abdominal, thoracic and PV), Musculo-skeletal Conventional & Superficial, Transesophageal, Transrectal and Transvaginal, and Thoracic/Pleural for motion/sliding and fluid detection. | GE Healthcare, LOGIQ e Ultrasound System. | 2 | August 01, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue associated with the LOGIQ e system. Image data from one patient has the potential to be linked to the patient ID of another patient. | DESIGN: Software Design | Starting on November 9, 2012, GE Healthcare, will plan to update the software on all of the affected GE LOGIQ e Ultrasound systems that were distributed at no cost to the customer.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.No Customer letter was sent.Further questions please call (262) 513-4122. | Contact the recalling firm for information | 588 | 588 | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DE, FL, GA, ID, IL , IA, MD, MA, MI, MS, MO, NJ, NM, NY, NC, OH, PA, RI, SC, TN,TX, UT,VT, VA and WI., and the countries of ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, HONG KONG, INDIA, ITALY, JAPAN, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND and UNITED EMIRATES. | Open |  |  | 588 | Computer
Z-1848-2013 | 65491 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF Fluoroscopic X-Ray systemFluoroscopic X-Ray system | Siemens Luminos dRF | 2 | August 01, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using the Luminos dRF with software version VD10A. The possibility exists that a software error may occur causing the table to tilt into an upright position and may collide with the floor. Four consignees were notified of this recall by Update Instruction XP067/12/S -- all four units have been corrected. | DESIGN: Process Design | Siemens initiated their correction of this product in May, 2013 by sending an Safety Update Instruction XP067/12S to resolve the software error to all affected customers. All units have been corrected.Further questions please call (610)-219-6300. |  | 4 | 4 | US Distribution including the states of CA, IN, NJ and OK. | Terminated | December 24, 2013 | 145 | 4 | Computer
Z-1835-2013 | 65564 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit, model 8015, with version 9.12-Keyboard Processor and logic board assembly part number TC10006584.The Alaris PC Unit is the main user interface unit and power supply of the Alaris System, a modular system to be used with Alaris System modules intended for use in todays growing professional healthcare environment for facilities that utilize infusion and/or monitoring devices. | Alaris PC unit, model 8015 | 1 | August 02, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC units model 8015 (PC unit), version 9.12, because it is operating at an incorrect voltage and it could potentially experience a loss of communication between the PC Unit main Processor and Keyboard Processor. | DESIGN: Process Design | The firm, Carefusion, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 17, 2013 to all customers who purchased the Alaris PC unit (model 8015) with version 9.12-Keyboard Processor. The recall letter informs the customers of the problems identified and the actions to be taken. The customers were instructed to either discontinue use of affect PC units until remediated or continue use on patients and follow the steps provided in the letter; fully charge Alaris PC units prior to use; during operation and when not in transport, keep Alaris PC units plugged in at all times; consider having additional back up Alaris PC units available and if you experience the described error code messages on the Alaris PC unit during infusion, you may continue the infusion(s) as programmed unless a programming change is required. Obtain a back-up Alaris PC unit to replace the device. In addition, if the customer experiences a SYSTEM ERROR (111.2000 or 111.2002), or a CHANNEL DISCONNECTED error on the Alaris PC unit contact CareFusion Customer Advocacy at 888-812-3266; 24 hours a day, 7 days a week; email: customerfeedback@carefusion.com. The customers were also instructed to promptly complete and return the enclosed Customer Response Card.CareFusion will provide an update to correct the voltage on the keyboard processor for those affected Alaris PC units. Carefusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to contact CareFusion Support Center at (888) 562-6018 or at supportcenter@carefusion.com for recall related questions; hours 7:00am to 4:00pm PT, Monday-Friday. For adverse reports, customers are instructed to contact Customer Advocacy at (888) 812-3266; 24 hours a day, 7 days a week or at customerfeedback@carefusion.com. For Technical questions, customers are instructed to call (888) 812-3229; hours 6:00pm to 5:00pm PT, Monday-Friday or at dl-us-inf-techsupport@ca. |  | 26564 | 26,564 total units (26,220 units for US) | Worldwide distribution: US (nationwide) and countries of: Australia, Canada and Europe. | Open |  |  | 26564 | Computer
Z-1874-2013 | 65778 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MRidium 1145 Dose Reduction System (DERS) drug library kit. (This is an option for the MRidium 3860+MR Infusion Pump System).The MRidium 3860+MRI Infusion pump are intended for general hospital or clinical use by medical professional whenever it is required to infuse patients with intra-venous or intra-arterial fluids before, during or after magnetic Resonance imagining (MRI) scans. | MRidium 1145 Dose Reduction System (DERS) | 1 | August 02, 2013 | 2013 | Iradimed Corporation | The Dose Error Reduction System (DERS) can indicate an incorrect recommended value for the rate via a specific key sequence during the initial infusion setup. This can result in a risk of over-infusion or under-infusion if the infusion is started with this rate value. |  | The firm, Iradimed Corporation, notified by phone and sent an "URGENT MEDICAL DEVICE RECALL NOTICE" via email dated July 1, 2013 all consignees/customers. The notice described the product, problem and action to be taken. The customers were instructed to identify the MRidium 3860+ MR Infusion Pumps which have the 1145 DERS Library card (AM05) installed. If these are not installed, no further action is required at this time; however, if they are installed-remove the 1145 DERS Library Card from the 3860 Pump using the instructions provided. Product will be reconditioned by installing a software upgrade that corrects this problem. The updated software (version 3.5.1) can be seen during the 3860+ Pump's power on cycle.Should your users have any questions following these instructions, please contact Iradimed Customer Service; 866-677-8022 (toll free) or 407-677-8022 ext. 100 or via e-mail - recallsupport@iradimed.com. | Contact the recalling firm for information | 176 | 176 individuals kits | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IL, KY, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, RI, TN, TX, UT, VA, and WA; and countries of: Australia, Austria, Denmark, Epsom, France, Germany, London, Netherlands, Saudia Arabia, and United Kingdom. | Open |  |  | 176 | Computer
Z-0476-2013 | 63187 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Enterprise 8000 Bed; an AC-powered adjustable hospital bed; ArjoHuntleigh AB, Eslov, Sweden; Models 8000, 8001, 8002, 8003, 8500, 8600, 8700 and 8800General medicine, general and specialist surgery, care of the elderly, HDU, ITU, coronary care and oncology. The bed is appropriate for high dependency patients who pose a movement and handling risk and/or whose clinical condition requires that they are positioned with the minimal physical handling. | Enterprise 8000 Bed | 2 | August 06, 2013 | 2013 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received reports of unintended movement where the backrest and/or thigh sections of the bed have moved without anyone consciously pressing the control button. These incidents have occurred in High Dependency and Intensive Therapy environments where unintended movement of the patient may have serious consequences. |  | Please be aware that this is not a new recall. Posting to the Internet was delayed due to a technical error.ArjoHuntleigh sent an Urgent Field Correction Recall letters dated November 16, 2012, to all affected customers, informing the accounts of reports of unintended movement of the backrest and/or thigh sections of the bed without intentional activation of the control buttons, which could have serious consequences due to the patient's links to other life supporting devices. The accounts were requested to disseminate the information to all users of the affected beds, to attach the enclosed warning labels to the outside of both head end safety sides immediately adjacent to the controls, to include a copy of the notice in the Instruction for Use Manuals for their beds, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 94 | 94 beds in U.S., 31 beds in Canada, 26,010 beds in ROW | Worldwide Distribution-USA (nationwide, including Puerto Rico) and the countries of Algeria, Angola, Australia, Austria, Bahrain, Belgium, Bosnia and Herzegovia, Botswana, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Falkland Islands, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahira, Lithuania, Malaysia, Malta, Mauritius, Mexico, Morocco, the Netherlands, New Zealand, Nigeria, Nordic, Norway, Oman, Panama, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Kingdom, United Arab Emirates, Uruguay, Venezuela and Yemen. | Open |  |  | 232 | Computer
Z-1907-2013 | 65470 | susceptibility test cards, antimicrobial | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861.The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | August 06, 2013 | 2013 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | TRAINING: Employee Error | The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. |  | 12384 | 12,384 total kits | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Open |  |  | 12384 | Computer
Z-1879-2013 | 65486 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | Sonopet Ultrasonic Aspirator Console5450-851-000(Console 100V)Product Usage:The Sonopet Console is indicated for use in surgical procedures where fragmentation, emulsification, and aspiration of soft and hard tissue is desirable, including neurosurgery, gastrointestinal and affiliated organ surgery, urological surgery, plastic and reconstructive surgery, general surgery, orthopedic surgery, gynecological surgery, thoracic surgery, laparoscopic surgery, and thoracoscopic surgery. | Sonopet Ultrasonic Aspirator Console | 2 | August 06, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | A supplied component exhibiting corrosion was built into the power supply boards of the Sonopet Consoles listed above. Theissue is limited to 2 lots of the supplied component. The presence of the corrosion on the relay of the power supply board canpotentially lead to a conduction defect which could cause the irrigation pump to stop working. The purpose of the irrigationpump is to supply sali | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Stryker sent an URGENT Medical Device Recall Notice letter dated June 24, 2013 to all affected customers. The letter identified the affected product, product issue, risk to health and actions to be taken. Customers were instructed to locate affected product and hold for upgrade, use enclosed response form to document inventory and receipt of notice, and in case of further distribution forward the recall notice. Stryker will dispatch a technician to perform the upgrade to the affected product. For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Kara Spath, Regulatory Affairs269-389-4518kara.spath@stryker.com. |  | 215 | 215 units total | Worldwide - USA Nationwide including the states of CA, CO, FL, HI, IA, IL, KS, KY, LA, MD, MI, MN, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, ITALY, and the countries of AUSTRALIA, CANADA, UNITED KINGDOM, NETHERLANDS, INDIA, JAPAN, NEW ZEALAND, SWITZERLAND, HONG KONG, and TAIWAN. | Terminated | January 30, 2014 | 177 | 215 | Computer
Z-1931-2013 | 65786 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. MHI-TM2000 Linear Accelerator System is intended for radiation therapy of lesions, tumors and conditions anywhere in the body where radiation treatment is indicated | MHITM2000/ VERO (TM) | 2 | August 08, 2013 | 2013 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. |  | The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 14 | 14 (2 in US) | Worldwide distribution: US (nationwide) including states of TX and FL; and countries of: Belgium, Italy, Japan and Korea. | Open |  |  | 14 | Computer
Z-1943-2013 | 65133 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA INFINITY MR SYSTEM WITH EXCITE TECHNOLOGY GE 3.0T SIGNA INFINITY TWINSPEED WITH EXCITE MR SYSTEMGE SIGNA 3.0T WITH EXCITE MR SYSTEM GE SIGNA EXCITE 1.5T MR SYSTEMGE SIGNA EXCITE 3.0T MR SYSTEM Signa HDx Signa HDxt Signa HDi Signa VibrantGE 0.35T SIGNA OVATION WITH EXCITE MR SYSTEM (i.e. Signa Ovation 5 and Ovation HD) GE 1.5T SIGNA HDe MR SYSTEM GE 0.35T SIGNA OVATION WITH EXCITE MR SYSTEM GE 0.7T SIGNA OPENSPEED (4) WITH EXCITE MR SYSTEM GE 0.7T SIGNA OPENSPEED (5) WITH EXCITE MR SYSTEMThe GE Signa Infinity line of products are whole body magnetic resonance scanners designed to support high resolution, high signal-to-noise ratio and short scan times. The Signa Infinity System is indicated for use as a diagnostic imaging device to produce transverse, sagittal, coronal and oblique images, spectroscopic images, and/or spectra, dynamic images of the internal structures and organs of the entire body, including but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. The images produced by the Signa Infinity systems reflect the spatial distribution of protons (hydrogen nuclei) exhibiting magnetic resonance. These images and/or spectra when interpreted by a trained physician yield information that may assist in diagnosis. | Signa Infinity Signa Infinity Twinspeed Signa Vibrant Signa HDx Signa HDxt Signa HDxi | 2 | August 09, 2013 | 2013 | GE Healthcare, LLC | Screen save images are included in wrong exam due to duplicate UID creation. This may cause a safety issuedue to a potential patient data mismatch.While the issue has only been observed at one customer site on the Signa HDxT 1.5T system, internal investigation revealed that the issue also impacts multiple GE healthcare MR Systems. |  | GE Healthcare communicated internally to their Service Team the directions for completing the correction on the customers' systems. GE Healthcare's recall strategy was to update all affected systems that were distributed. At this time,Week of 02/06/2012, deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2608 or email paulcorrigan@ge.com. | Contact the recalling firm for information | 7016 | 7016 (2919 US; 4097 OUS) | Worldwide distribution: USA (nationwide) including DC and Puerto Rico and countries: ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIDJAN, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIAN, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU. PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF MACEDONIA, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TUR,EKY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, YUGOSLAVIA, MALAWI, MAURITANIA, NIGERIA, UKRAINE, IVORY COST, OMAN, ZAMBIA, UZBEKISTAN, ESTONIA, IRAQ, and JAMAICA. | Open |  |  | 7016 | Computer
Z-1940-2013 | 65333 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed CT750 HDThe GE LightSpeed CT750 HD Computed Tomography X-ray system is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes, including Axial, Cine, Helical (Volumetric), Cardiac, Spectral, and Gated acquisitions for all ages. | GE LightSpeed CT750 HD Computed Tomography Xray System | 2 | August 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare conducted a recall involving 2 events affecting the GE LightSpeed CT750 HD. Event 1 - There was a report regarding CT number variance in time line at posterior fossa region of head perfusion. When beam hardening and helical artifacts in this region are reconstructed into the perfusion map, these artifacts will give subtle differences in the CT numbers.Event 2 - The software does no | DESIGN: Software Design | GE Healthcare's recall strategy was to update CT750HD CT Systems with software 10MW25.6 on all systems that were distributed. Updates started on 4/5/2011. At this time, deployment of this recall is considered complete. The recall was completed 2/2/2012.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com. | Contact the recalling firm for information | 244 | 244 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT, FL, GA, HI, IL, IN, IA, KY,ME, MD,MA, MI,MN, MS, NV, NH, NJ, NY, NC, PA, TN, TX, VA, WA,WV, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANDA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, NEW ZEALAND, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUKEY, and UNITED KINGDOM. | Open |  |  | 244 | Computer
Z-1944-2013 | 65583 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 Ultrasound SystemThe device is intended for use by a qualified physician for ultrasound evaluation of Fetal; Abdominal; Pediatric; Small Organ (breast, testes, thyroid); Neonatal Cephalic; Adult Cephalic; Cardiac (adult and pediatric); Peripheral Vascular; Musculo-skeletal Conventional and Superficial; Urology (including prostate); Transrectal; Transvaginal and Intraoperative (abdominal, thoracic, vascular and neurosurgical). | LOGIQ S8 Ultrasound System | 2 | August 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue related to incorrect image measurement associated with the software of the LOGIQ S8 system. This issue was found internally. Users may notice that the system shows incorrect image measurement results on Pulsed Wave (PW) or TimeIntensity Curve (TIC) or depth measurement when using Virtual Convex mode and putting a cursor (Sample Volume, Depth measur | DESIGN: Software Design | GE initiated a field action on 8/13/12. GE plans to update the software on all of the affected GE LOGIQ S8 Ultrasound systems that were distributed, at no cost to the customer.Should you have any questions or require additional information, please contact RA Leader, US Region by telephone at 262-548-2089 or email James.Giles@med.ge.com. | Contact the recalling firm for information | 755 | 755 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IL, IN, IA, LA, MA, MI, MS, MO, MT, NE, NV, NJ, NY, NC, OH, OK, PA,RI, SC, TN, TX, VA, WA, and WI; and countries of: ALGERIA, AUSTRALIA, AUSTRIA, BRAZIL, CANADA, COLOMBIA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, LATVIA, LEBANON, MALAYSIA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, REPUBLIC OF MOLDOVIA, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and YUGOSLAVIA. | Open |  |  | 755 | Computer
Z-1942-2013 | 64992 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Optima XR200amx, and Brivo XR285amx.The series are indicated for use on adult and pediatric patients for general-purpose diagnostic radiographic examinations and procedures. The systems are indicated for taking radiographic exposures of the skull, spinal column, chest, abdomen, extremities, and other body pants with the patient sitting, standing, or lying in the prone or supine position. These devices are not intended for mammographic applications. | Optima XR220amx, Optima XR200amx, and Brivo XR285amx | 2 | August 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare conducted a recall on various mobile x-ray units (Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems) because of an Erratic Drive Performance. | DESIGN: Software Design | GE Healthcare recall strategy was to update all of the affected Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems that were distributed. At this time ( 2/14/2013), deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com. |  | 458 | 458 (306 US; 152 OUS) | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: Albania, Australia, Bahrain, Belgium, Canada, Chile, Czech Republic, Dominican Republic, France, Germany, Great Britain, Honduras, Hungary, Ireland, Israel, Italy, Japan, Korea, Kuwait, Netherlands, Nigeria, Norway, Romania, Saudi Arabia, Spain, United Arab Emirates and Venezuela. | Terminated | February 04, 2014 | 179 | 458 | Computer
Z-1945-2013 | 65721 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zeegox-ray, angiographic system | Siemens AXIOM Artis zeego | 2 | August 09, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety notice about the overlay of system information after a functional check of the brakes on the Artis Zeego system. | DESIGN: Software Design | The firm, Siemens, sent an "Important safety notice" by letter on May 13, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately notify the service department to clarify what further measures need to be taken, if after carrying out the regular brake test, the system can only be moved at a reduced speed. If this device has been sold and is therefore no longer in your possession, send this safety notice to the new owner and provide the new owner to Siemens.Update instruction AX012/13/S was issued, and a software update will be released as AX006/13/S and AX007/13/S which will modify the system software so that it will not be possible for information to be overlaid.If you have any questions call +49 (9191) 18 0. | Contact the recalling firm for information | 132 | 132 | Nationwide distribution: AL, AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MD, MN, MO,MS, NC, NE, NH, NJ, NY, OH, PA, PR, SD, TN, TX, UT, VA, WA, and WI. | Open |  |  | 132 | Computer
Z-1949-2013 | 64304 | system, immunomagnetic, circulating cancer cell, enumeration | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS AUTOPREP System (Software Version 2.4.1 & Below. Product Usage:The CELLTRACKS AUTOPREP System is a general purpose laboratory instrument used with immunomagnetic reagents that capture and enrich target cells, and labeling reagents that differentiate cells in whole blood. Cell analyzers such as the CellTracks Analyzer II, CellSpotter System, flow cytometers or microscopes may be used for cell identification and enumeration. The system is fit for in vitro diagnostic use. In the CellTracks AutoPrep System software 2.4.1 release notes Pub. No. J40135EN issued 2012-03-26 page 2, the document states that a new feature unique to software 2.4.1 allows the used to modify a batch in a new way. The user has additional optional fields in the edit sample information dialog box that allows one to enter expanded sample and patient information. The user can also edit the patient management setting for selected samples if not all samples have the same designation. | CELLTRACKS AUTOPREP System | 2 | August 10, 2013 | 2013 | Veridex, LLC | Veridex, LLC has received a report that a patient's demographics were incorrectly associated with another patient's identification (PID) number. The purpose of this notification is to inform you of the issue and provide instructions to help prevent an occurrence on your CELLTRACKS AUTOPREP System and CELLTRACKS ANALYZER II. | DESIGN: Software Design | Veridex sent an Important Product Correction Notification letter dated December 28, 2012, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 82 | US: 82 Units; Ex-US 78 Units. | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, China, India, Japan, Singapore, England, France, Germany, Italy and Spain. | Open |  |  | 152 | Computer
Z-1877-2013 | 65369 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX Evolution System, X-Ray System ---The DRX-Evolution system is a permanently installed diagnostic x-ray system for generation of x-rays for examination of various anatomical regions. | Carestream DRX Evolution System | 2 | August 12, 2013 | 2013 | Carestream Health Inc. | Due to potential for an x-ray exposure technique change. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Carestream will make the correction which includes the installation of a SW patch to the affected customer sites, and to any other affected customer sites utilizing the DRX Evolution with V5. 5 SW outside of the US. --- Once receiving FDA approval on the corrective action plan, the correction will be implemented to all affected devices at identified customer sites: 1. A notification letter was sent to all customers identifying the issue and highlighting the approved procedures of the Corrective Action Plan. The letters were sent on 8/16/13 via Certified Mail, Return Receipt Requested. 2. Carestream Health will contact the customer to schedule a service visit to install software updates to correct this issue. The installation time will be approximately one hour.3. As stated in the customer notification letter, this software update will be provided to the customer at no expense.4. All software updates are estimated to be completed within 1-3 months.5. Carestream Health will provide monthly progress reports on the effectiveness of the plan, including the number of units corrected.Further questions please call (800) 328-2910. |  | 49 | US: 49 units; Foreign: 183 units (revised) | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Canada, China, Finland, France, Germany, Italy, Korea, Kuwait, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Open |  |  | 49 | Computer
Z-1958-2013 | 64970 | tray, surgical, instrument | FSM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Case Lid and Base (Inspan, Invue, SLIFT) Caddy Lid (Inspan, Arena-C); Inspan Lid: 12-60039A, 11-60042B, 11-60043B, 11-60045B, 11-60046C, 11-60045A, 11-60046A, 11-60053A; Invue Lid: SI70095A; S-Lift Lid: SI50070A; S-Lift Base: SI50067A; Arena-C Lid: 11-72013D, 11-72014D; Generics Lid: 12-99005A.System cases are designed to hold and secure all instrumentation to be used with a surgical system. | Case/Caddy (tray, surgical, instrument) | 2 | August 13, 2013 | 2013 | SpineFrontier, Inc. | Non-toxic inks, used on the silk screen graphics of the case bases and lids, and caddy lids, were found to be either delaminating or running off from the substrate after repeated washing and sterilization cycles. | DESIGN: Labeling Design | On 2/15/2013, SpineFrontier emailed distributors informing them that they needed specific parts returned to work towards completing the conversion of product graphics. |  | 258 | 258 units (30 in distribution at the time of the recall) | Worldwide Distribution-USA including the states of FL, OH, MA, and TX, and the countries of Mexico, and Dominican Republic. | Open |  |  | 258 | Computer
Z-1956-2013 | 65743 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | StealthStation S7 Framelink Software version 5.4.Aid for locating anatomical structures and planning surgical trajectories in open and percutaneous procedures. | STEALTHSTATION TREATMENT GUIDANCE PLATFORM | 2 | August 13, 2013 | 2013 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. has initiated a field correction to their Framelink software version 5.4 which can be used on Medtronic Navigation StealthStation 57 and i7 systems or a Medtronic Planning Station due to the potential of discrepancy between the exam image and the displayed measurement or overlay under certain conditions of use. | DESIGN: Software Design | A Field Safety Notice was sent on 7/15/13 to users informing them that the Framelink software version 5.4 has been revised to correct the software anomalies described in this letter. A new version will be installed on their system(s) by a Medtronic Navigation representative. Until the software upgrade has been installed on their system, do not use high resolution examinations (greater than 256 x 256). If users are not sure of the exam resolution being used, firm recommend contacting the hospital imaging department or local Medtronic Navigation representative for assistance. Medtronic representatives will visit all identified hospitals and upgrade all systems to version 5.4.1. This will account for all consignees. |  | 572 | 572 | Worldwide Distribution. | Open |  |  | 572 | Computer
Z-1973-2013 | 65381 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂŽ Advanced Perfusion System 1 (APS1, System 1)100/120V ACProduct Usage:The TerumoÂŽ Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns TM or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | August 14, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (TCVS) has become aware that an Advanced PerfusionSystem 1 (APS1) Centrifugal Pump 'Service Pump' message may occur due to user inducedmotor movement. In this instance, the APS1 will identify the movement (as it isdesigned) and display a Service Pump message error. The user might interpret this as apump which requires service rather than a message displayed d | DESIGN: Device Design | Terumo Cardiovascular Systems (TCVS) notified their consignees of the Urgent Medical Device Recall letter dated 7/26/13 and was mailed 7/29/13. The notification provided reason for the device correction, affected product, how the device will be corrected, and the potential health hazard associated with a user's inaccurate interpretation of the service pump error message. Instructions for consignee responses to the recall notice communication may be faxed to 734-741-6149 or emailed to the address on the return form. If you have any questions or concerns, call Terumo Cardiovascular Systems (TCVS) Customer Service at 1-800-521-2818 M-F, 8am-6pm ET. Terumo Cardiovascular Systems (TCVS) will provide an addendum to the APS1 Manual to their consignees. The addendum will be placed at the beginning of Chapter 6, Centrifugal Control Unit. |  | 1657 | 1657 total | Worlwide Distribution - US Nationwide: AL,AR,AZ, CA, CO, CT, DC, DE, FL,GA, HI, IA, ID, IL, IN, KY, LA, MA , MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of: AUSTRALIA,UNITED ARAB EMIRATES (UAE),Hong Kong, onesia,Singapore,Taiwan,Thailand, Philippines,CHILE, Malaysia,USA,South Korea,BELGIUM,Japan,Mexico,CANADA | Open |  |  | 1657 | Computer
Z-1969-2013 | 65354 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂŽ Advanced Perfusion System 1: 220/240V AC, Advanced Perfusion System 1The TerumoÂŽ Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns " or similar equipment. | Terumo Advanced Perfusion System 1 | 2 | August 14, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | The operators manual does not fully describe the sequence of expected events that will occur when a specific operating condition exists. Specifically, if the user is adjusting the System 1 centrifugal pump speed via the Central Control Monitor at the same time an alarm condition occurs with configuration for a pump coast response, the labeling in the operators manual does not fully describe how th | DESIGN: Software Design | Terumo CVS, sent an URGENT MEDICAL DEVICE RECALL notification letter dated July 26, 2013, to all affected customers explaining the correction to the operators manual, potential hazards and what to do in response to the issue.If you have any questions or concern please contact Terumo CVS Customer Service 1-800-521-2818 Recall Fax 1-734-741-6149 Customer Service Hours:Monday  Friday, 8 AM - 6 PM ET. |  | 1657 | 1657 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Dubai, Hong Kong, Indonesia, Singapore, China, Mexico, British Columbia, Taiwan, Thailand, Chile, Malaysia, Japan, Belgium, United Arab Emirates (UAE), South Korea, Manila and Canada. | Open |  |  | 1657 | Computer
Z-1975-2013 | 65719 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza SW version VA20 with Mammography license.image processing, radiological system | Siemens syngo.plaza | 2 | August 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Customer Safety Notice to inform customers of a potential problem when using syngo.plaza. In some instances, it can happen that the laterality information in mammography images is not displayed and these images may not be properly aligned to mammography segments. | DESIGN: Software Design | The firm, Siemens, notified their customers on March 4, 2013 and sent a "Customer Safety Advisory Notice" letter dated April 03, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to always check whether the laterality information is displayed in the image text of mammography images; if not, the DICOM header is accessible and can be checked for the tag "Laterality" (0020, 0060). If a system is not able to provide the DICOM tag "Image Laterality" (0020, 0062), configuring the image text by adding the tag was proposed. If you have sold the device/equipment and it is no longer in your possession, forward this safety notice to the new owner of the device/equipment and inform us about the new owner of the device/equipment. This issue will be resolved in the syngo.plaza software version VA20D_HF04 which is planned to be released in 02/2013. Customers were also instructed to contact the local Siemens service office to plan the update.If you have any questions, contact the SYNGO CEO or Vice President Quality at +49 (9131) 84-4874 or email to: michael.mlb.braun@siemens.com. |  | 6 | 6 | Nationwide distribution: FL, MA, KS, NY, OH, PA, TX and WI. | Open |  |  | 6 | Computer
Z-1977-2013 | 65380 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System 1Product Usage:The TerumoÂŽ Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns TM or similar equipment. | Terumo Advanced PErfusion System 1 | 2 | August 15, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | TCVS is taking action to update the Operators Manual with the specific instructions to respond to this condition and will issue a labeling addendum for current users. In the event of a false high pressure alarm condition occurs that displays a "999" pressure reading and results in stopping a pump, this pressure display can be cleared and the pump re-started by disconnecting the pressure transduce | DESIGN: Software Design | Terumo sent an Urgent Medical Device Recall notification letter dated July 26, 2013 to all affected customers. The letter identified the affected product, reason for correction, potential hazards of the recall issue, the correction to be implemented, and customer instructions for replying to the recall notification. Terumo CVS is providing an addendum to be placed in the Operator's Manual at the beginning of Chapter 7, Modules. Questions or concerns are directed to Terumo CVS Customer Service at 1-800-521-2818, Recall Fax 1-877-275-6860. Customer Service hours are Monday to Friday 8am-6pm ET. | Contact the recalling firm for information | 1657 | 1657 | Worldwide Distribution - US Nationwide in the states of :AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MSMT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and countries of AUSTRALIA, UNITED ARAB EMIRATES (UAE),Hong Kong,Indonesia, Singapore, Taiwan, Thailand,Philippines, CHILE,Malaysia,South Korea, BELGIUM,Japan,Mexico,CANADA | Open |  |  | 1657 | Computer
Z-1967-2013 | 65797 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Discovery NM/CT 570c CT Standalone ScanThe Discovery NM/CT 570c hybrid system is a combination of a cardiac dedicated nuclear imaging subsystem and LightSpeed GE CT subsystem. This system is intended for cardiac Nuclear Medicine (NM) diagnostic imaging applications, head and body CT diagnostic imaging applications, as well as attenuation correction and anatomical localization. The Discovery product is a SPECT-CT system. It consists of two back-to-back gantries, the NM Gantry, carrying solid state Cadmium Zinc Telluride (CZT) detectors, the CT Gantry, a single patient table, power distribution unit (PDU), NMV integrated power supply (IPS) , two operator consoles with two acquisition systems (Nuclear and CT), interconnecting cables and associated accessories (ECG gating, etc.). | Discovery NM/CT 570 c | 2 | August 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident where the wrong part of the patient body wasscanned during a CT standalone exam with the Discovery NM/CT 570c | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. GE Healthcare has provided field modification instructions (FMI) to update software for the Discovery NM/CT570c hybrid system. The software update will correct the incorrect positioning of the table not prescribed by the operator. The FMI includes:A software upgrade with foreign language support to be installed in the affected models listed above,Labels (FMI label and packing label),Software upgrade workflow procedure with functional verification for all modes of operation, andCertification Report, Postponement of Medical Device Correction, Rejection of Medical Device Correction, or Refusal and/or Location Unknown Documentation.For further questions please call (262) 513-4122. |  | 10 | 10 Systems | US Distribution in the state of CT. | Terminated | January 21, 2014 | 154 | 10 | Computer
Z-2046-2013 | 65865 | camera, ophthalmic, ac-powered | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | NAVIS-EX , Image Filing Software for NIDEK AFC Fundus camera, Ver 1.1.0 to 1.3.5. NAVIS-EX is indicated for use as file management software for ophthalmic images. | NAVISEX | 2 | August 23, 2013 | 2013 | Nidek Inc | Under certain circumstances, the Left eye image may be saved as the Right eye image. |  | "Medical Device Correction Notification" letters dated August 5, 2013 were sent to consignees through their US contracted distributor - Marco Opthalmic. The letter described the product issue and provided information on the correction. Customers can contact 800-223-9044 for questions. |  | 86 | 86 | Nationwide US distribution only. | Terminated | February 25, 2014 | 186 | 86 | Computer
Z-2084-2013 | 48997 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics, Model No. V1000 | Esprit Ventilator | 2 | August 26, 2013 | 2013 | Respironics California Inc | The power to the Esprit Display's Backlight is interrupted, causing the GUI to flicker or go dark and making the display unreadable. |  | A Medical Device Field Correction letter dated 10/26/2007 was sent to all customers who received the Esprit Ventilator. The recall letter informed the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact Respironics' Customer Service Dept. at (800) 345-6443, select Option#5, then select Option #4. |  | 3255 | 3,255 units (2302 units in US) | USA Nationwide and Worldwide to Algeria, Argentina, Australia, Bangladesh, Brazil, Cambodia, Canada, Chile, Colombia, Ecuador, Egypt, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Libya, Malaysia, Mexico, Pakistan, Peru, Philippines, Puerto Rico, Russia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | August 26, 2013 | 0 | 3255 | Computer
Z-2088-2013 | 65607 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. In vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric, and integrated ion-selective multisensory technology for clinical use. | Dimension Vista Intelligent Lab Systems | 2 | August 26, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot |  | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: 1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation - | Contact the recalling firm for information | 1937 | 1937 | Worldwide distribution: US (nationwide) including PR and countries of: Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, Portugal, Republic Korea, Saudia Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, and United Kingdom. | Open |  |  | 1937 | Computer
Z-2089-2013 | 65769 | cabinet, table and tray, anesthesia | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mc Kesson Anesthesia-Rx Model Number 8.0.2 & 8.0.1 with RapiDispenseProduct Usage: Automated anesthesia cart that monitors and controls the dispensing of medications, narcotics and supplies during surgical procedures. | Mc Kesson AnesthesiaRx | 2 | August 26, 2013 | 2013 | MCKESSON TECHNOLOGIES INC | The firm issued a clinical alert after users reported an issue where the Anesthesia-Rx system was interrupted ANF became temporarily unavailable during use. | DESIGN: Software Design | McKesson sent a Clinical Alert letter dated July 22, 2013 to all affected customers. The letter identified the affected product, symptoms, and actions to be taken. McKesson committed to releasing 8.0.2 Update 1.5 to affected customers which will include enhanced code to prevent the unhandled excpetion errors. Additional resolution steps were included in the letter. For questions call 1-800-700-8737. | Contact the recalling firm for information | 106 | 106 | USA Nationwide Distribution in the states AR, MD, NJ, NY, PA, and WV. | Open |  |  | 106 | Computer
Z-2059-2013 | 65950 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Upgrade DR, Code No: 712086This system is used for making X-Ray exposures for diagnostics. | Radiographic system, digital | 2 | August 26, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | DESIGN: Software Design | Philips Healthcare sent an Urgent Field Safety Notice letter dated July 24, 2013, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed : When reading an image, a physician should not use the mirror icon for image side interpretation. 2. Clinical staff should follow the Instructions for Use and use lead letters to indicate body orientation of patient or equivalent electronic markers to provide Left/Right indications, during all X-ray image acquisitions. Philips will upgrade the system, which involves software and hardware upgrades of the system. The software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Philips will also provide updated Instructions for Use for the work spot and the system (including an addendum).If you need any further information or support concerning this, please contact your local Philips representative Technical Support Line at 1-800-722-9377.We apologize for any inconvenience this may cause and trust that this information is adequately addressing any concerns you may have. | Contact the recalling firm for information | 89 | 89 | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, GA, MO, MS, NC, NH, NY, PA and TX., and the countries of Canada, Australia, China, Czech Republic, France, Germany, Greece, Italy, Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Open |  |  | 89 | Computer
Z-2098-2013 | 65880 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Ultrasound ViewPoint 6. Product Usage: ViewPoint is intended to be used in medical practices and in clinical departments and serves the purposes of diagnostic interpretation of images, electronic documentation of examinations in the form of text and images and generation of medical reports primarily for diagnostic ultrasound. ViewPoint provides the user the ability to including images, drawings, and charts into medical reports. ViewPoint is designed to accept, transfer, display, calculate, store and process medical images and data, and enables the user to measure and annotate the images. The medical images, which ViewPoint displays to the user, can be used for diagnostic purposes. ViewPoint is intended for professional use only. ViewPoint is not intended to be used as an automated diagnosis system. | GE Ultrasound ViewPoint 6 | 2 | August 27, 2013 | 2013 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, |  | GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia). | Return product to place of purchase for a full refund | 225 | 225 | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: ANDORRA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, CZECH REPUBLIC, ESTONIA, FINLAND, GERMANY, HONG KONG, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KUWAIT, MALAYSIA, Norway, OMAN, PALESTINIAN TERRITORY, OCCUPIED, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Open |  |  | 225 | Computer
Z-2111-2013 | 65779 | processor, cervical cytology slide, automated | MKQ |  |  | Pathology |  | PrepStain Tecan US-I.For use in the screening and detection of cervical cancer. | PrepStain Tecan USI | 2 | August 28, 2013 | 2013 | Becton Dickinson & Co. | PrepStain system rack setting was incorrectly entered on one PrepStain system during a routine preventive maintenance visit by a BD Field Service Representative. | DESIGN: Software Design | Ex-US BD sites were notified via email and teleconference, and will communicate directly with customer sites in their region as required by local regulation. All US customers were initially contacted via telephone and letters were mailed. | Contact the recalling firm for information | 890 | 890 instruments | Worldwide Distribution-USA (nationwide) and Puerto Rico and the country of Canada. | Open |  |  | 890 | Computer
Z-2115-2013 | 66016 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000.Product Usage: The intended use of the system is to monitor physiologic parameter data on adult, pediatric and neonatal patients. | Dash 3000/4000/5000 Patient Monitor | 2 | August 28, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent noise on ECG and/or Respiration waveforms associated with movement of the ECG connector on your DASH, PDM, and/or Tram Module. ECG and Respiration signal noise caused by DASH, PDM, and/or Tram ECG cable connector movement may result in reduced ECG, Respiration and arrhythmia detection performance. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare "Urgent Medical Device Correction" letter dated April 3, 2013 to all affected customers. The letter was addressed to Healthcare Administrator / Risk manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Product, Safety Instructions, Affected Product Details, and Product Correction. For questions contact Technical Support at 1-800-558-7044. |  | 56240 | 28,120 *14,709 US & 13,411 OUS) | Worldwide Distribution - USA Nationwide including DC and PR except ND and the countries of: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, HONDURAS, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNSIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, BOLIVIA, HONG KONG, JAMAICA, PHILIPPINES, SOUTH KOREA, TRINIDAD TOBAGO, URUGUAY, VENEZUELA. | Open |  |  | 28120 | Computer
Z-2108-2013 | 65793 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis systems. C-arm angiographic x-ray systemx-ray, angiographic system. | Siemens Axiom Artis | 2 | August 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Axiom Artis systems installed by a third party installer that may involve incorrect channel nuts being used. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent Update Instructions dated December 2012, to all affected customers. This update was initiated on behalf of the US service organization to re-check all systems installed by MEC Group. A call center representative or designated engineer called each site on April 26, 2013, and described the work being done, amount of time needed and number of engineers that would be on site. At that time the call center or engineer scheduled the update with the customer.For further questions please call (610) 219-6300. |  | 178 | 178 | US Distribution including the states of PA, NJ, DC, VA, MD and DE. | Open |  |  | 178 | Computer
Z-2131-2013 | 65923 | device, microtiter diluting/dispensing | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | WASP (Walk Away Specimen Process). Automatic Planting and Streaking InstrumentCatalog number: W086WASP is a fully automated system used to process microbiological specimens in liquid phase for the investigation of infectious microbes. | WASP (Walk Away Specimen Process) Instrument software | 3 | August 30, 2013 | 2013 | Copan Italia | The recall of the WASP (Walk Away Specimen Process) instrument was initiated because Copan Italia SpA has upgraded the software concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. |  | Copan sent an Urgent WASP recall letter dated July 31, 2013, to all affected customers that received the WASP (Walk Away Specimen Process) instrument to notify them of the software upgrade concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. This event may happen only if damage in the door locking system happens. This could result in a minor, transient injury to the user due to the impact with the rotating table.The WASP instruments involved in this upgrade are all whose serial numbers show the configuration with the broth inoculation feature activated.The machine installed in your laboratory is one of this serial number involved in the SW upgrade. The Software upgrade will be performed on site by personnel trained by Copan and no action is required by yourself.Customers will be contacted by the WASP Technical Support in order to define a date that will fit with their planning. Customers with questions are instructed to call +39 0303666130 or write to regulatory.wasp@copanitalia.com. |  | 63 | 63 units | Worldwide Distribution - USA including Oregon and California., and the countries of France, Belgium, Italy, Denmark, Norway, Switzerland, Romania, Finland, Germany, South Africa, Canada, Ireland, UK, Sweden and Japan | Terminated | December 12, 2013 | 104 | 63 | Computer
Z-2157-2013 | 66048 | block, beam-shaping, radiation therapy | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO - Beam shaping block for Radiation Therapy. | Elekta DMLC | 2 | September 04, 2013 | 2013 | Elekta, Inc. | Clinical mistreatment may occur when there are errors in the clinical work flow. | DESIGN: Device Design | Important Field Safety Notices 200 01 406 030 (Changes to your system to improve the safe operation of DMLC) and 200 01 602 031 (Changes to your system to improve the safe operation of 3DLINE DMLC) letters were sent to customers on 8/15/2013 advising them of the issue and recommendations to resolve either by upgrading their system or to discontinue a clinical use of the device. | Contact the recalling firm for information | 31 | 31 | AL, CT, FL, GA, IA, MS, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA and Puerto Rico | Open |  |  | 31 | Computer
Z-2164-2013 | 64146 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of deliveringfluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-100ES, 621-200ES, 621-300ES, and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | Outlook ES Safety Infusion System | 2 | September 06, 2013 | 2013 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on 26 August, 2011 due to an issue in the Outlook ES Safety Infusion System where the Care Area can inadvertently be exited thereby removing DoseGuard dose limit protection. This occurs when a sequence of HOLD - HOLD key strokes are performed on the pump key panel while the user is in the Dose Dat | DESIGN: Software Design | The firm decided to conduct a field correction to correct the issue and carried it out through routine field services. The firm was notified by FDA that this is a reportable event, so they notified their consignees via letters dated 10/31/2012. |  | 1970 | 1,970 pumps added for this expansion | Nationwide Distribution. | Open |  |  | 1970 | Computer
Z-2163-2013 | 64147 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of deliveringfluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-100ES, 621-200ES, 621-300ES, and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | Outlook ES Safety Infusion System | 2 | September 06, 2013 | 2013 | B Braun Medical, Inc. | B. Braun Medical Inc. has become aware of an issue with the Outlook ES Safety Infusion System. The key panel may fail to function as intended if the underlying circuitry fails due to intermittent contact. This issue applies specifically to the HOLD button. If this occurs, the Door Processor (DP) will not receive the desired input from the user. This issue can be observed in pumps with key panels | DESIGN: Component Design/Selection | B Braun sent an URGENT: MEDICAL DEVICE FIELD CORRECTION letter dated October 24, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Support at 1-800-627-7867 for questions regarding this notice. | Contact the recalling firm for information | 28427 | 28,427 | Nationwide Distribution. | Open |  |  | 28427 | Computer
Z-2175-2013 | 64150 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-100ES, and 621-200ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | Outlook ES Safety Infusion System | 2 | September 10, 2013 | 2013 | B Braun Medical, Inc. | B. Braun Medical Inc. is informing our customers of a field correction previously performed to address the loss of the drug library in the Outlook Pump Models 620-100 and 620-200. This could have occurred after periods of lost DC (battery) power or out of tolerance low power. This could have led to a delay in therapy, or to delivery of an over or under-dose of a therapy if the user was to rely on | DESIGN: Component Design/Selection | The firm, B. Braun Medical Inc., sent an "URGENT: MEDICAL DEVICE FIELD CORRECTION" letter dated December 05, 2012 to its customers. The letter describes the product, problem and actions taken. The firm finished performing "field upgrades" on March 20, 2010 for this issue.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Infusion Systems Customer Service at 1-800-627-7867 (627-PUMP). |  | 14161 | 14,161 | Worldwide Disribution: USA (nationwide) and country of: Canada. | Open |  |  | 14161 | Computer
Z-2160-2013 | 66024 | table, radiologic | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Juno DRF System. Mobile X-ray systemThese are general, multi-function x-ray systems capable of routine radiographs and general fluoroscopy procedures | Juno DRF System | 2 | September 13, 2013 | 2013 | Villa Radiology Systems LLC | It has been discovered that the Air Kerma rate information is inconsistent in Continuous Fluoro modes 30x30 (12"x12") and 15x15 (6"x6). This could result in a deviation of the displayed values of Air Kerma Rate and Cumulated Air Kerma from the actual values greater then +/-35% and then in a failure to comply with the requirements of 21 CFR 1020.32(k)(6 | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the Villa Sistemi Medicali (U.S. Agent: Villa Radiology Systems) is responsible for the correction of all 22 noncompliant units distributed within the U.S. Should the CAP prove ineffective in bringing the noncompliant unit(s) in the field into compliance with the applicable performance standard, the CDRH can require more stringent measures to obtain compliance.The system will be modified by upgrading the generator power firmware, which will bring Juno DRF x-ray system into compliance. You provided a verification and validation test report (Juno DRF System AK fix-DAM Prom V&V; Test Report) to FDA demonstrating that these proposed upgrades addressed the problem. Field service engineers will schedule an appointment with customers to install the new software.2. The customer notification letter (URGENT: Electronic Product Radiation Warning), dated August 02, 2013, includes the statement that "the manufacturer will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you."3. The CAP will be completed within a six month timeframe (by March 1, 2014).4. The manufacturer will notify the Secretary on monthly basis providing progress with the field correction.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22.For further questions please call (203) 262-8836. | Contact the recalling firm for information | 22 | 22 units within US | US Distribution including the states of AZ, MI, IN, UT, CT, AL, NY, PA, IL, MD, CA, MA, IA, OH, MS and MO. | Open |  |  | 22 | Computer
Z-2188-2013 | 64152 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | Outlook ES Safety Infusion System | 1 | September 13, 2013 | 2013 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on August 26, 2011 due to an issue in which the Outlook ES Safety Infusion System may halt infusion but the "RUN" Light Emitting Diodes (LEDs) on the front display continue to advance as if the pump were infusing. The pump emits a backup alarm, but there are no visual indicators that the infusion | DESIGN: Software Design | The firm, B Braun Medical Inc, sent an "Urgent Expanded Medical Device Field Correction" letter dated November 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that pumps that have been upgraded (software version 151599 or higher), no additional action is required on their part at this time. For the pumps that have not yet been upgraded, within three weeks of the letter, customers should confirm the serial numbers using the enclosed response form or contact B. Braun Medical, Inc. Infusion Systems Customers Support at 1-800-627-7867 (800-627-PUMP) to schedule their upgrade.For questions regarding this recall call 972-245-2243. | Contact the recalling firm for information | 6919 | 6,919 | Nationwide distribution: USA including states of: AR, CA, FL, IL, LA, NC, NJ, NY, MD, ME, MN, SC, TN, TX, VA, WI, and WV. | Open |  |  | 6919 | Computer
Z-2201-2013 | 66156 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850.Intended for use in multiple areas within a professional healthcare facility. | GE Healthcare, CARESCAPE Monitor B850 | 2 | September 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue associated with the CARESCAPE Monitor B850 when used with the iCentral network. The iCentral must have the alarm management software license (L-NETM-00) and the alarm limit change option enabled to experience this issue. When changing the HR, SpO2 or Invasive pressure alarm limits from the iCentral over the network, the alarm priority on th | DESIGN: Software Design | Consignees were sent on 7/20/2011 a GE Healthcare "Important Product Information" letter. The letter was addressed to Healthcare Administrator / Risk Manager, chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Instructions, Product Correction, and Contact Information (Technical Support 1-800-558-7044). |  | 636 | 636 (53 US, 583 OUS) | Worldwide Distribution-USA (nationwide) including the states of CA, CO, CT, HI, IL , IA, KY, LA, MA, MI, NE, NY,OH, OK, TX, and WI, and the countries of UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF SERBIA, PORTUGAL, POLAND, PARAGUAY, NORWAY, NETHERLANDS, NEPAL, MEXICO, MALAYSIA, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BOSNIA AND HERZEGOVINA, BELGIUM, BAHRAIN, AUSTRIA, and AUSTRALIA. | Terminated | October 25, 2013 | 39 | 636 | Computer
Z-2213-2013 | 65657 | system, communication, powered | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Responder 5 is a nurse call communication system that connects patients with staff in a hospital. The nurse call system also connects team members and departments from radiology, physical therapy, transport, environmental services and more.Responder 5 is used in hospitals and long term patient care facilities to allow for remote communications between patients and patient monitoring devices and healthcare providers | Responder 5 | 2 | September 17, 2013 | 2013 | Rauland-Borg Corp | A counterfeited electronic component was identified in the products affected. The component, a data transceiver (transmitter/receiver), is associated with the intercommunication of these devices in the Responder 5 system across the L-Net data path. Failure of a data transceiver results in the loss of nurse call functionality in the associated room, and may disrupt communications on the L-Net run | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Counterfeit | Tech Bulletins sent a letter dated May 10, 2013, to all affected customers notifying them of the counterfeit parts and requesting that customers uninstall and return affected products to Rauland-Borg for replacement. A subsequent recall notification letter dated August 5, 2013 was mailed to customers. The letter included instructions to both direct accounts and end users on the process of replacing affected units with units manufactured with authentic parts.For additional information please call 847-590-5920. | Contact the recalling firm for information | 385 | Product Code 351000: 385 units; Product Codes 352000 and 352020: 14,808 units (total) | Worldwide distribution - US: (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NE, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI and WV., and the countries of : Australia, Canada, Columbia, Lebanon, Mexico and Saudi Arabia. | Open |  |  | 14808 | Computer
Z-2241-2013 | 66111 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantage Workstation, READY View Application.Image analysis software. | GE Healthcare, Advantage Workstation, READY View Application | 2 | September 18, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware that when analyzing fMRI images using the READY View application software, incorrectfMRI activation maps may be computed and displayed. | DESIGN: Software Design | GE provided Urgent Medical Device Correction letters to consignees in the following countries: Canada, Finland, France, Germany, Guadeloupe, Italy, Japan, Latvia, Martinique, Netherlands, Norway, Poland, Reunion, Romania, Saudi Arabia, Singapore, Spain,Sweden, Switzerland, Turkey and the United Kingdom.The letters were dated March 27, 2012 and were addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US 800-437-1171, Japan 0120-055-919 and for other countries contact your local GE Healthcare Service Representative).*Letters to other consignees is being planned for September 2013. |  | 318 | 318 (28 USm 290 OUS) | Worldwide Distribution-USA including the states of IL, MN, NJ, NY, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GUADELOUPE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN,KOREA, LATVIA, MARTINIQUE, NETHERLANDS, NORWAY, PHILIPPINES,S POLAND, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VIETNAM. | Open |  |  | 318 | Computer
Z-2253-2013 | 66168 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery IGS 730 Cardiovascular X-Ray Imaging System.The angiography X-ray systems are indicated for use in generating fluoroscopic and rotational images of human anatomy for cardiovascular, vascular and non-vascular, diagnostic and interventional procedures. Additionally, with the OR table, the angiographic X-ray systems are indicated for use in generating fluoroscopic and rotational images of human anatomy for image guided surgical procedures. The OR table is suitable for interventional and surgical procedures. | GE Healthcare Discovery IGS 730 Cardiovascular XRay Imaging System | 2 | September 18, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to Discovery IGS 730 Imaging Systems.A patient data management error may prevent the Discovery IGS 730 Imaging System from booting on start-up or re-booting during procedure. The boot issue may occur when a largeamount of patient data remains in the browser and is not deleted. The inability to boot the system ma | DESIGN: Software Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 19, 2013 to its Consignees/customers. The letter described the Safety Issue, Safely Instructions, Affected Product Detail, Product Correction and Contact Information. The consignees were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call : United States: 800 437 1171. For other countries, please contact your local GE Healthcare Service Representative. |  | 4 | 4 (2 US, 2 OUS) | Worldwide distribution: US (nationwide) including states of : PA and TN; countries of: FRANCE and SINGAPORE. | Terminated | October 25, 2013 | 37 | 4 | Computer
Z-2256-2013 | 66148 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650.The CARESCAPE Monitor B650 is a multi-parameter patient monitor intended for use in multiple areas and intra-hospital transport within a professional healthcare facility. The CARESCAPE Monitor B650 is intended for use on adult, pediatric, and neonatal patients and on one patient at a time. The CARESCAPE Monitor B650 is indicated for monitoring and recording of, and to generate alarms for, hemodynamic (including EGG, ST segment, arrhythmia detection, ECU diagnostic analysis and measurement, invasive pressure, non-invasive blood pressure, pulse oximetry, cardiac output, temperature and mixed venous oxygen saturation), impedance respiration, airway gases (C02, 02, N20 and anesthetic agents), spirometry, gas exchange, and neurophysiological (including electroencephalography, Entropy, Bispectral Index (HI1S) and neuromuscular transmission) status. The CARESCAPE Monitor B650 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for remote viewing and to data management software devices via a network. The CARESCAPE Monitor B650 is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a professional healthcare facility. The CARESCAPE Monitor B650 is not intended for use during MRI. | CARESCAPE Monitor 650 | 2 | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a number of potential issues associated with the CARESCAPE Monitor B650.No injuries or illnesses have been reported. | DESIGN: Software Design | GE Healthcare sent an Important Product Information letter dated July 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will provide a software and hardware correction at no charge once it is available. We will contact you to arrange for this correction.If you have any questions regarding this notification, please contact Technical Support (1-800- 558-7044) or your local Service Representative. | Contact the recalling firm for information | 9466 | 9466 (3261 US, 6205 OUS) | Worldwide Distribution - USA (nationwide) including the states of : AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV,NH, NH, NM, NY, NC, OH, OK OR, PA, PR, SC, TN, TX, VA, WA, WV, WI and WY., and the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, COSTA RICA. CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MOROCCO, NETHERLAND, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM and VENEZUELA. | Open |  |  | 9466 | Computer
Z-2257-2013 | 66126 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage, 5183685, 5183685-2, 5183685-3, 5183685-4, 5183685-5, 5183685-6, 5183685-7, 5183685-8, 5183685-9, 5342425, 5342425-2, 5342425-3, 5342425-4, 5342425-5.Seno Advantage is a medical image review station that allows easy selection, processing, filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. When interpreted by a trained physician, mammographic images displayed on the high-resolution monitors may be used as an element for diagnosis. Furthermore filmed images from all modalities may also be used as an element for diagnosis. | Seno Advantage medical image review station | 2 | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety. |  | GE Healthcare sent an Important Product Information letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety.The Contact Information (US 800-437-1171, Japan 0120-055-919, Other countries contact local GE Healthcare Service Representative). | Contact the recalling firm for information | 179 | 179 | Worldwide Distribution - USA (nationwide) including the states of : AL AZ, CA, CO, CT, FL, GA, IL, IN, MD,MI, MO, NJ, NY, NC, OK, PA, TX and WI.,and the countries of : AUSTRIA, BELGIUM, CANADA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, JAPAN, LEBANON, NORWAY, PORTUGAL, SPAIN, SWEDEN and TAIWAN. | Open |  |  | 179 | Computer
Z-2266-2013 | 65992 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE 1.5T Signa HDxt, 3.0T Signa HDxt, Signa Vibrant, HD23 Software and the following model numbers: 2395001, 2395001-2, 2395001-3, 2395001-4, 5269997-2, 5269997-4, 2377062-61, 5127452-3, 5127452, 5337894-3, 5337894-4, 5911000-4, 5911000-5, 53085735308573-2.GE Discovery MR750w 3.0T, Discovery MR750 3.0T, Optima MR450w 1.5T, Discovery MR450 1.5T.The 1.5T Signa HDx family and 3.0T Signa HDx family are a whole body magnetic resonance scanner designed to support high resolution, high signal-to-noise ratio, and short scan times. IT is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images, spectroscopic images, parametric maps, and/or spectra, dynamic images of the structures and/or functions of the entire body, including but not limited to head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. Depending on the region of interest being imaged, contrast agents may be used | Discovery MR750w /MR750, Optima MR450w, Discovery MR450, 1.5/3T Signa HDxt, Signa Vibrant. | 2 | September 23, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to non-unique series identifiers embedded in image headers of certain GE MR Products.For ClariView and Screen Save from GVTK (Gems Visualization ToolKit) Viewer, the MR Product can generate image headers with non-unique series identifiers, which when exported to some PACS systems, results in the series being associated wit | DESIGN: Software Design | Customers were sent two GE Healthcare "Urgent Medical Device Correction" letters one dated April 9, 2013, and the other May 22, 2013, each addressing one of the 2 groups of MR Products involved in the recall. The letters were addressed to Hospital Administrators / Risk managers, Radiology Department Mangers and Radiologists. The letters described the Safety Issue and provided Safety Instructions, Affected Product Details, Product Correction and Contact Information.GE Healthcare will correct all affected systems by providing a software update at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification.For further questions please call (262) 513-4122. | Contact the recalling firm for information | 725 | 1,246 (521 US and 725 OUS) | Worldwide Distribution - All states in continental USA including DC except AK, AZ, RI and WY. and the countries of : AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEANON, MARINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, LUXEMBOURG, PORTUGAL, VIETNAM and URUGUAY. | Open |  |  | 1246 | Computer
Z-2271-2013 | 66211 | cart, emergency, cardiopulmonary (excluding equipment) | BZN | Cardiopulmonary emergency cart. | Anesthesiology | Anesthesiology | 510(K) Exempt | Stryker Vision Mounting Arm, for use with other Stryker Endoscopy medical device products for human patients. Model number 0240-095-200, manufactured by Stryker Endoscopy San JoseThe intended use of the Stryker Vision Mount is to hold and position an LCD monitor. The Stryker Vision Mounting Arm can be used with any Stryker flat panel display and is It is compatible with Stryker Carts (0240099001, 0240099011 or 0240099020). It is indicated for use in a surgical operating room environment and is exposed to environmental conditions associated with the storage and shipping of the product. It is not intended to enter the sterile field. The Stryker Vision Mount does not deliver to and/or extract from the patient any energy or material. It does not process any biological materials, as it is not intended to be sterilized by the user. The Stryker Vision Mount is only used to display surgical images and video and does not sustain or support life. | Stryker Vision Mounting Arm | 2 | September 24, 2013 | 2013 | Stryker Endoscopy | When positioning the display located on top of the video cart, it is possible for the user to apply excessive force and push the display up and out of the mounting arm. This can result in the display falling upon the user/patient. | DESIGN: Device Design | Stryker Endiscopy sent an Urgent Device Correction letter dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform individuals within their organization who needed to be aware of this Medical Device Field Correction, contact Stericycle at 1-877-546-0482 or email stryker4645@stericycle.com to schedule a time for Stericycle to perform the correction, or perform the correction themselves by requesting the materials and following the correction instructions included. Customers were also instructed to report all adverse events or product quality problems to Stryker Endoscopy at 1-800-624-4422, press 3 for Customer Service. --UPDATE--On September 17, 2013, the recall was expanded to include additional customers who may have received the screw and washer included with the design change update. Notification was sent to these consignees on September 17, 2013. . | Contact the recalling firm for information | 14271 | 14271 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and internationally to Australia, Canada, Chile, China, Argentina, Guam, Taiwan, India, Japan, Korea, Latin America, Philippines, Mexico, New Zealand, Hong Kong, Singapore, and Switzerland. | Open |  |  | 14271 | Computer
Z-2275-2013 | 66051 | medical device data system | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS system Four labels of the Tegris System include:Emergency button (1703.72)MAQUET GmbH & Co. KGKehler-Strabe 31Baden-Wuerttemberg76437 Rastatt, Germany Touch screen (elo)Elo TouchSystems 301 Constitution DriveMenlo Park, CA 94025www.elotouch.comManufacturer: Tyco ElectronicsBrand Name: Elo Touch SystemsMade in ChinaTegris system (1703.01)MAQUET GmbH & Co. KGKehler Strabe 3176437 Rastatt, Germany OR Table Module (1703.71)MAQUET GmbH & Co. KGKehler Strabe 31Baden-Wuerttemberg76437 Rastatt, GermanyDesigned to be used as the central operating system in an operating room. The integration system has two main functions. -recording and distribution of images and videos -interaction with medical and non-medical devices. The system has a modular design and comprises the basic units as well as s various optional modules,which may be used or integrated in conjunction with the basic unit. TEGRIS is operated using a touchscreen. Medical devices that can be operated by TEGRIS include OR tables an OR lights. In the case of OR tables and OR lights, the main functions of these devices may be separated, e.g. adjustment of OR tables or switching the OR lights on.off. TEGRIS has no effected on the primary operating functions of the original control device. TEGRIS does not provide any other functions in addition to these primary operating functions. The purpose of TEGRIS namely the central operation of these various devices, is dependant on the individual devices. Please refer to the operating instructions of the individual devices for instructions. The integrated surgical safety checklist is a raft and must be modified to meet local requirements and approved by the User. TEGRIS is intended for short-term use pursuant to Appendix IX of the Directive 93/42/EEC of medical products in medical rooms in accordance with Group 1 and 2 VDR 0100-710, providing, among other things, protection against electric shock and connections for potential equalization lines. Intended for the following users: medical experts | Maquet | 2 | September 25, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue that might occur because of the lack of an emergency stop button on a secondary touch screen display on the TEGRIS System. | DESIGN: Device Design | Maquet sent an Urgent Medical Device Recall Letters/Field Action Response Forms dated August 19, 2013, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that a MAQUET authorized service technician would contact them to arrange an appointment to equip their system with a second emergency stop button that would be located in close proximity to the secondary touch screen. Customers were instructed to complete and return the enclosed Field Action Response Form acknowledging their receipt and understanding of the communiction by emailing a scanned copy to FieldActions@maquet.com or via fax at 973-807-1658. Customers with questions were asked to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3.For questions regarding this recall call 973-709-7361. | Contact the recalling firm for information | 21 | 21 units | Worldwide Distribution - USA including NY and PA and internationally to Australia, Brazil, Canada, Germany, Hong Kong, Japan, Sweden. | Open |  |  | 21 | Computer
Z-2276-2013 | 65996 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP SystemRadiation treatment planning | Monaco RTP System | 2 | September 25, 2013 | 2013 | Elekta, Inc. | Upgrade installation script is designed to overwrite the same pouch files in the Treatment Couch Library with the original default values. | DESIGN: Device Design | Monaco 3.3 upgrades have been temporarily halted and Important Safety Notice (LINMON0007) has been issued to effected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that couch files are checked for releases 3.20.00 and 3.30.00 to verify that the correct relative electron density values are being used in the dose calculation. If customers have edited the sample couch files, and previously upgraded to Monaco 3.20.00 or 3.30.00, their couch files have already been overwritten. Open the couch files, make the necessary updates to the couch parameters, and use the "Save As Treatment Couch" option to save the modified sample couch files to a new, unique name.Please note that the couch files can also be edited in Focal and then the patients opened in Monaco. Comparable Focal releases would be Focal4.64.00(Monaco 3.10.00), Focal4.70.00(Monaco3.20.00) and Focal4.80.00 (Monaco 3.30.00).For questions regarding this recall call 404-993-5886. | Contact the recalling firm for information | 45 | 45 | Nationwide Distribution including AK, CA, CO, CT, FL, IL, IN, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, WA and Puerto Rico. | Open |  |  | 45 | Computer
Z-2277-2013 | 65998 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System Radiation Therapy Treatment Planning System. Radiation treatment planning. |  | 2 | September 25, 2013 | 2013 | Elekta, Inc. | Patient needs to be positioned as indicated by the treatment plan. | DESIGN: Labeling Design | Computerized Medical Systems issued an Important Safety Notice dated May 22, 2013 to all affected customers. This letter idenitfies the problems the products and the actions needed to be takem by the custom. |  | 72 | 72 | Distributed in Puerto Rico and the states of AK, AZ, CA, CO, CT, FL, IL, IN, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, and WA. | Terminated | December 23, 2013 | 89 | 72 | Computer
Z-2273-2013 | 66219 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Product Name: AVEAÂŽ Ventilator.All Models. The affected devices are cleared under K073069 pm 1/25/2008 and K103211 on 05/06/2011. This device is a Class II medical Device.A continuous ventilator (respirator) is a device intended to mechanically control or assist patient breathing by delivering a predetermined percentage of oxygen in the breathing gas. Adult, pediatric, and neonatal ventilators are included in this generic type of device. | AVEA Ventilator | 1 | September 27, 2013 | 2013 | Carefusion 211 Inc dba Carefusion | CareFusion has identified a potential risk associated with AVEAÂż ventilators when used at higher altitudes. A proactive complaint review identified an error in the barometric pressure compensation.CareFusion is voluntarily initiating a field correction of the affected devices to correct the error. | DESIGN: Component Design/Selection | CareFusion sent customer notification letters on September 05, 2013 via USPS mail return receipt carrier. The customer notification letter dated September 3, 2013 was titled "URGENT PRODUCT RECALL", and was addressed to Biomedical Engineering, Respiratory Care, and Risk Management. The customer notification letter informed customers of the product under recall with product name, description, and serial numbers enclosed as list unique to each customer. The customer notification letter also states the Problem and Affected Devices: Issue, Potential Risk, Actions to be Taken by CareFusion, Action to be Taken by the Customer, and Contact Information.Care Fusion has initiated a project to develop a new version of software that will resolve this issue. Once this software is available to download you will receive a notification letter.The customers were instructed to promptly return the enclosed Response Card to expedite the correction process and acknowledge receipt of this Notification and follow the listed instructions noted in the letter.For Recall Related and Support contact CareFusion Recall Support Center at 888.562.6018 or email: SupportCenter@carefusion .com Questions. For Product Technical Support Adverse Event Reporting contact Support/Customer Advocacy at 800.231.2466, 714.283.2228 or email to support.vent.us carefusion.com. from (6:30am -5:00pm POST).UPDATED:Firm issued press release of the recall on 09/30/2013. |  | 19546 | 19,546 units | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, American Samoa, Argentina, Australia, Austria, Bahrain, Barbados, Bangladesh, Belgium, Bermuda, Belize, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile,China, Colombia, Costa Rica, Croatia, Cyprus, Czech Rep, Dominican Republic, Ecuador, Egypt, Estonia, El Salvador, Finland, France, Germany, Georgia, Greece, Guatemala, Hong Kong, Honduras, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Lybia Jamahiriya, Malaysia, Mexico, Moldova, Morocco, Netherlands, Nepal, Nicaragua, Norway, Oman, Pakistan, Palastine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Open |  |  | 19546 | Computer
Z-2301-2013 | 66195 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology |  | Selenia Full Field Digital Mammography System with software versions 5.0.x and 5.1.xThe Selenia Full Field Digital Mammography system generates digital mammography images that can be used for screening and diagnosis of breast cancer. The Selenia Full Field Digital Mammography System is intended for use in the same clinical applications as traditional screen-film mammography systems. Mammography images can be interpreted on either hard copy film or soft copy review workstations, cleared for mammography | Selenia Full Field Digital Mammography System with software versions 5.0.x and 5.1.x | 2 | September 30, 2013 | 2013 | Hologic, Inc. | Software: When the Selenia upon meeting the hard drive threshold of 4,000 images, may delete the most recently dated images | DESIGN: Software Design | Hologic sent an Urgent Medical Device Recall letter dated September 5, 2013, by Certified Mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised how to avoid loss of images until Hologic corrects the problem. A software upgrade to customers who are connected to Hologic Connect who have already been connected remotely and customers who have already been connected via a Service Call will receive notification of the problem and the date the system was remotely corrected notified by Certified Mail.A Dealer notification issued by certified mail to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00072.Selenia software version 5.2 will implement a double safeguard to both guarantee a 28 day retention period, as well as ensure images are removed in a least recently used order.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 62 | 62 systems | Worldwide Distribution - USA including the states of AL, AR, CA, CO, FL, IL, KY, LA, NE, NC, NJ, NY, OH, PA, TN, TX, UT, WA and WV., and the countries of Canada, Denmark, Iran, Italy, Mexico and Taiwan. | Terminated | December 13, 2013 | 74 | 62 | Computer
Z-2318-2013 | 66301 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. | McKesson Horizon Medical Imaging | 2 | September 30, 2013 | 2013 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. |  | 329 | 329 units | Worldwide Distribution - United States including the territories of Guam and Puerto Rico and the countries of Canada and France. | Open |  |  | 329 | Computer
Z-2295-2013 | 65821 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | MedStream Refill Kit, U.S: 91-4287Product Usage:The Medstream programmable pump is intended for the intrathecal delivery of Baclofen in the US market and for the intrathecal delivery of Morphine or Baclofen in the EMEA market. Refill Kits- are intended for use in the filling and re-filling of the Medstream pump reservoir. | Infusion Pumps & Refill Kits | 1 | September 30, 2013 | 2013 | Codman & Shurtleff, Inc. | Drug flow rates exceeding programmed flow rates. |  | Codman & Shurtleff, Inc. sent a Medical Device Field Safety Notice letter dated August 13, 2013 to US Consignees via FedEx overnight delivery and International Consignees were notified of the recall via email. The letter identifed the affected products, potential clinical impact, and updated IFU pump filling warning. The letter informed consignees that the Instructions for Use (IFU) and product training materials had been updated by adding a warning statement about potential over infusion if air gets inside the pump reservoir and additional clarifications to reinforce the proper filling technique. Consignees are asked to complete the Field Safety Notice Acknowledgement Form and return it to your local Codman Neuro Sales Representative or fax to: 508-977-6665.Consignees are asked to contact your Codman Neuro representative if they have questions, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to 1-866-491-0974, Option 2. | Contact the recalling firm for information | 5 | 5 units | Worldwide Distribution - USA Nationwide in the states of NJ, NC, TX, FL and in the countries of Austria, Belgium, Egypt, Germany, Greece, Ireland, Israel, Italy, Luxemborg, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, and UK. | Open |  |  | 4116 | Computer
Z-0009-2014 | 66381 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2.Intended for volume or pressure control ventilation. | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2 | 1 | October 15, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Avance, Avance CS2, and Aisys anesthesia devices.While operating in Pressure Control Ventilation  Volume Guarantee (PCV-VG) mode, the affected products may produce an over delivery tidal volume when given a unique sequence of user inputs and a collapsed bellows (partially filled circuit).Exposure to excessiv | DESIGN: Software Design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated September 17, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager, and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US call 1 800-345-2700, for other countries contact your local GE Healthcare Service Representative). |  | 10263 | 10,263 (3,194 US, 7,069 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico except DE, ME, NM, RI, and VT, and the countries of UNITED ARAB EMIRATES, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA , MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED KINGDOM, VIETNAM, URUGUAY, and VENEZUELA. | Open |  |  | 10263 | Computer
Z-0010-2014 | 66250 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom.Intended for volume or pressure control ventilation. | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom | 1 | October 15, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving unresponsive buttons on the display of the EngstrÂżm ventilator and the Aespire View, Aisys, and Avance anesthesia machines.As a result of a manufacturing issue, the buttons on the left, right, and bottom keypads may not always detect button presses. This may result in the inability to access certain menu functions wh | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent on 8/28/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 28, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manger and director of Biomedical / clinical engineering. The letter described the Safety Issue, Safely Instructions, Affected Product Details, Product Correction and Contact Information (1-800-345-2700 US). |  | 911 | 911 (252 US, 659 OUS) | Worldwide Distribution-USA (nationwide) including the states of US: AK, AZ, CA, CO, DE, FL, GA, IL, KY, LA, MD, MI, MN, MS, MT, NV, NY, NC, OH, OR, PA, SC, TX, WY, and WI and the countries of AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COST ARICA, CROATIA, CZECH REPUBLIC, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KAZAHHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LITHUANIA, MAURITITUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPINES, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, and VENEZUELA. | Open |  |  | 911 | Computer
Z-0023-2014 | 65958 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, Optima CT660.Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data. | GE, Optima CT660 | 2 | October 17, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the sameSeries UID. It has been seen that some PACS may combine series under a patient using only the Series UID | DESIGN: Software Design (Manufacturing Process) | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 26, 2013. The letter was addressed to Director of Radiology, Emergency Medicine Physicians and Radiologic Technologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 380 | 380 (52 US & 328 OUS) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, IS, LA, MA, MO, NH, NM, NY, OH, OK PA, SC, TX, UT, VA, WA, and WI, and the countries of UNITED KINGDOM, DENMARK, CHILE, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CZECH REPUBLIC, ECUADOR, EGYPT, ETHIOPIA, FINLAND, FRANCE, FRANCE GUIANA, GERMANY, GUADELOUPE, GUATEMALA, GUYANA, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, JORDAN, KAZAHSTAN, KOREA, KUWAIT, LATVIA, MYANMA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UZBEKISTAN, and VIETNAM. | Terminated | December 05, 2013 | 49 | 408 | Computer
Z-0013-2014 | 66296 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology and Brilliance CT 16 (Air)Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology and Brilliance CT 16 (Air) scanners are whole body Computed Tomography X-ray systems that are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient and equipment supports, components and accessories. | Philips | 2 | October 17, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There are artifacts appearing on the Tracker images during threshold-triggered bolus scans with protocols using a Rotation Time of 0.4 seconds. | DESIGN: Software Design | Philips sent an "Urgent - Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem If you need any further information or Support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377, Option 5: Enter Site ID or follow the prompts).This notice has been reported to the appropriate Regulatory Agency.Philips apologizes for any inconveniences caused by this problem. | Contact the recalling firm for information | 258 | 258 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR,TN, TX, UT, VA, VT and WV., and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Denmark, Egypt, France, Germany, Honduras, India, Indonesia, Italy, Kazakhstan, Latvia, Lebanon, Libyan Arab Jamahiriya, Mexico, Netherlands, Norway, Palestine, Philippines, Poland, Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom and Vietnam. | Open |  |  | 258 | Computer
Z-0100-2014 | 65722 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | pocH-100i; sysmex; Automated Hematology Analyzer; 7.3" x 13.8" x 18.1" / 30.8 lbs.; Sysmex Corporation; 1-5-1 Wakinohama-Kaigandori, Chuo-ku, Kobe 651-0073, Japan; Sysmex America, Inc; One Nelson C. White Pkwy, Mundeline, IL 60060, U.S.A.The Sysmex pocH-100i Automated Hematology Analyzer is an automated cell counter intended for in vitro diagnostic use in CLIA non-waived clinical laboratories (not for Point of Care use in a CLIA waived laboratory). The Complete Blood Cel (CBC) count test parameters include: WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, and NEUT%/#, LYMPH %/#, MIXED %/# (EO, BSO, MONO). | Sysmex pocH100i | 2 | October 18, 2013 | 2013 | Sysmex America Inc | The data printed on the thermal printer may omit a digit or decimal point in the results or sample identification number. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Sysmex, sent a "Product Notification-FIELD CORRECTION" letter dated May 2013 to all customers in the U.S. and Canada on May 15, 2013 by U.S. Mail. The letter describes the product, problem and actions to be taken. The customers will be contacted by a Sysmex Service Representative within the following 90 days to schedule an appointment to replace the Printed Circuit Board (PCB) on the internal printer. The repair will be performed during an on-site service visit. The customers were instructed to, until the repair can be made, use the thermal printer for reporting results; to compare the sample identification, values and decimal points on the printout with the data on the LCD screen for accuracy.If you have any questions concerning this information, please email the Sysmex Technical Assistance Center (TAC). Simply click on email TAC button on the home page of the CRC, complete the form that appears and then click the submit to send your question to TAC. Sending an email to TAC is recommended for non-urgent requests.For urgent request, call the Technical Assistance Center at 1-888-879-7639. |  | 193 | 193 units | Worldwide distribution: US (Nationwide) including Puerto Rico and states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and country of Canada. | Open |  |  | 193 | Computer
Z-0034-2014 | 64787 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Cardio Imaging 5.0 is a software only Picture Archiving and Communication System (PACS). Product Usage: It will be sold as a software only device to operate on general purpose computing hardware. Centricity Cardio Imaging receives medical images and other information from various data sources. Information can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Centricity Cardio Imaging is intended to assist trained professionals in the viewing, analysis and diagnostic interpretation of images and other information for the diagnosis and treatment of cardiac and vascular disease. These trained professionals include but are not limited to physicians, cardiologists, radiologists, nurses, medical technicians and assistants. Centricity Cardio Imaging is not intended for use in the patient vicinity. | Centricity Cardio Imaging | 2 | October 21, 2013 | 2013 | GE Healthcare It | GE Healthcare is aware of a potential safety risk where an image labeling situation exists in the Image Viewer software of the Centricity Cardio Imaging product, version 5.0 SP1. Multiple studies may be displayed side by side by selecting "Compare Study" from the Study List, or by opening prior studies in the Patient Folder. In these cases, the Viewer differentiates the studies wityh a label (Curr | DESIGN: Software Design | GE Healthcare IT mailed an IMPORTANT PRODUCT INFORMATION letter dated October 2, 2012 to the only consignee of this product. The letter notified the customer of the issue, provided temporary instructions on how to use safely use the software until a new software version can be installed, and advises that a software correction will be made available in the near future. Customers who have any questions can contact their local GEHC-IT service representative or call GE Healthcare's Remote Online Center (ROC) Customer Care Center at 1-800-437-1171. |  | 1 | 1 unit | USA Nationwide Distribution in the state of OK | Open |  |  | 1 | Computer
Z-0035-2014 | 66424 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 5.2 DB Server (running Oracle 10.1.0.4.0), medical imaging Picture Archiving and Communication System (PACS). | IMPAX | 2 | October 21, 2013 | 2013 | AGFA Corp. | System downtime and/or slow performance may result if software in use is an old version of IMPAX and is not at current supported levels. |  | An "URGENT FIELD SAFETY NOTICE" letter was sent via email to the consignees on 9/19/2013, Acknowledgement, via FAX-Back or email, that the information was received and understood has been requested from the consignees. The letter specifically describes the safety alert for the device, informs consignees of the problem and recommends the upgrades for IMPAX and Oracle. |  | 4 | 4 | Worldwide distrbution: US(nationwide) distribution including NY and OK, and country of Canada. | Terminated | October 24, 2013 | 3 | 4 | Computer
Z-0032-2014 | 66293 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | FSE Image Upgrade Kit, Catalog # 92019CThis kit is used to install and/or upgrade software used with the EVOLIS Microplate System. | Image Upgrade Kit, Catalog 92019C | 3 | October 21, 2013 | 2013 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIAQuantitative Determination in EVOLIS APF version 6.0. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites wheninstalling the EVOLIS APF V6.0b.Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option2 and then Option 3 for any questions about this product correction. |  | 23 | 23 kits | US Distribution including the states of FL, IA, ME, MO, MS, NE, NY, TN, TX, and WA | Terminated | November 13, 2013 | 23 | 23 | Computer
Z-0046-2014 | 66304 | excimer laser system | LZS |  |  | Ophthalmic |  | STAR Excimer Laser System, Model: STAR S2, Catalog Number: 0030-1479, Product is manufactured and distributed by AMO Manufacturing USA, LLC, Milpitas, CAThe STAR Excimer Laser Systems are indicated for laser assisted in situ keratomileusis (LASIK), photorefractive keratectomy (PRK), and phototherapeutic keratectomy (PTK) eye surgery in certain patient populations. | STAR Excimer Laser | 2 | October 22, 2013 | 2013 | Abbott Medical Optics, Inc. | This corrective action will address an identified failure of the lamp and illumination system within the laser unit caused by an unexpected failure of a thyristor component on the AC Power Distribution Printed Circuit Board (PCB). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Abbott Medical Optics, Inc., (AMO), sent an "ADVISORY NOTICE and a COVER LETTER FOR ADVISORY NOTICE" dated September 16, 2013 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the Star Notification Acknowledgement form via fax to 1-714-247-4530 within 3 business days or mail to Abbott Medical Optics, 510 Cottonwood Drive, MS 347118, Milpitas, CA 95035, and share this advisory notice with your staff.If you have any questions related to operation of STAR Excimer Laser System or documentation requirements associated with this letter, please call AMO Technical Support at 1-800-511-0911 and you will be directed to an AMO Representative. |  | 834 | 834 units in total | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Asia Pacific, Canada, Europe, Africa, Middle East, Japan, and Latin America. | Open |  |  | 834 | Computer
Z-0052-2014 | 65812 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson SC2000 volume imaging ultrasound systems at software version 3.0.Mfg name:Siemens Medical Solutions USA, Inc.Mountain View, CA 94043Medical Device for human use. Ultrasound imaging system used to measure anatomical structures and calculation packages that provide information to the clinician for clinical diagnosis purposes. | Acuson SC2000 | 2 | October 23, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Software issue results in inaccurate 3D image may result in ablation in wrong location. | DESIGN: Device Design | Siemens sent an Customer Safety Advisory Notification letter to all affected customers on July 10, 2013.The letter identified the product the problem and the action needed to be taken by the customerLetters provide steps to mitigate the risk and inform users that a software fix will be installed as soon as it is available. Siemens Service Engineers will contact all locations to arrange installation once fix is available. All personnel who use the ACUSON should be notified of the issue.Update September 17, 2013 : After discussion with CDRH, Siemens revised their letter to state that procedures using the ACUSON SC2000 system with BioSense Webster CARTO 3 equipped with CARTOSOUND Module and SoundSTAR Catheter should not be performed until the users system has been updated with the software upgrade that resolves the issue.For further questions please call (650) 694-5398. |  | 26 | 26 units | Worldwide Distribution - US (nationwide) including the states of IL, PA, AZ, OR, OH, NY, CO, TX and CA., and the countries of France, Germany, Japan, Portugal, Saudi Arabia, Taiwan and Turkey. | Terminated | November 25, 2013 | 33 | 26 | Computer
Z-0082-2014 | 66336 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET Getinge Group CARDIOHELP Base Unit REF 70104.7999 REF 70104.8012 Manufactured for: MAQUET Cardiopulmonary AG Made in Germany.Cardiopulmonary support system. | CARDIOHELPi | 2 | October 25, 2013 | 2013 | Maquet Medical Systems USA | It has come to the attention of MAQUET that in certain rare instances, the CARDIOHELP-i device may display an error message with audible alarm, indicating "battery defective." The alarm will occur-upon start up or when the main power supply is unplugged from the main power source. The human machine interface touch screen of the CARDIOHELP-i device may gradually become black due to the human mach |  | MAQUET issued their customers "Urgent-Medical Device Field Correction" notices/Customer Fax Response forms dated September 16, 2013 via FedEx on September 16, 2013. The notification informs the customer of the issue with the product; how to identified affected product; and the action needed to be taken by the customer. The field correction includes service on the affected product. Authorized MAQUET Service Technicians will service the affected product to correct the issue. Customers were requested to complete and fax back (or email a scanned copy) of the enclosed response form. Customers could contact their local field representative or the MAQUET Customer Service line at 1.888.627.8383 (press option 2 followed by option 2) Monday through Friday, 8am to 6pm EST.For questions regarding this recall call 973-709-7000. |  | 63 | 63 | Nationwide Distribution including AZ, FL, CO, TX, OH, VA, MN, WI, SC, IL, ND, UT, CA, OR, NY, IA, CT, SD, MI, HI, AL, MD, KS, MO | Open |  |  | 63 | Computer
Z-0067-2014 | 66205 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CTDiagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-Ray Computed Tomography (CT) systems. The CT subsystem produces cross sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other Oran imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode, dynamic, and gated acquisitions. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a "fused" (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for quantification of cardiology images and data sets including but not limited to myocardial perfusion for the display of wall motion and quantification ventricular or left-function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated. | Ingenuity TF PET/CT | 2 | October 25, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Issues with the Ingenuity TF PET/CT Software Version 4.0 could pose a risk for patients in the event that a patient rescan were to require reinjection of PET radiopharmaceutical. | DESIGN: Software Design | Philips Healthcare sent an Field Safety Notice dated August 16. 3013, via certified mail to all US customers. Market Representative will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customer.Field Safety Notice 88200460 is intended to inform you about;" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problemAdditionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning these issues, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800.722.9377, option 5: Enter Site ID or follow the prompts). |  | 13 | 13 | Worldwide Distribution - US (nationwide) including the states of NY, OH and Pa, and the countries of Australia, Austria, Egypt, France, Germany, India and Netherlands. | Open |  |  | 13 | Computer
Z-0102-2014 | 66166 | electrode measurement, blood-gases (pco2, po2) and blood ph | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDPointÂŽ 400 system blood gas analyzerProduct Usage:These systems are intended for near-patient and laboratory testing of blood gases, electrolytes, metabolites, total hemoglobin, and hemoglobin derivatives in arterial, venous, and capillary whole blood samples. The systems are capable of determining the following parameters: RAPIDPoint 400 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, Hct RAPIDPoint 405 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, tHb, F02Hb, FCOHb, FMetHb, FHHb, nBili | Siemens Healthcare | 2 | October 29, 2013 | 2013 | Siemens Healthcare Diagnostics Inc | RAPIDPointÂż 400 Series, RAPIDPointÂż 500, orRAPIDLabÂż 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample | DESIGN: Software Design (Manufacturing Process) | Siemens Healthcare sent an Urgent Field Safety Notification letter dated September 2013 via Federal Express, to affected Siemens domestic customers. Effected Siemens customers outside the US was provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures. Previously Siemens Healthcare had issued a Customer Notification on August 23, 2012. The letter identified the affected product, issue, information to mitigate the issue and actions to be taken. The firm will issue an upgrades software to users. For questions contact your local Customer Care Center - Technical Support Representative at 1-877-229-3711. |  | 828 | 828 units | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Brazil, Brunei, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Fren.Polynesia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Macedonia, Malaysia, Mayotte, Netherlands, New Caledonia, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., United Kingdom, Vatikancity, Venezuela, Vietnam, | Open |  |  | 11006 | Computer
Z-0122-2014 | 66059 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio XG Diagnostic Ultrasound System SSA-790AUsed for routine diagnostic imaging | Aplio XG Diagnostic Ultrasound System SSA790A | 2 | October 30, 2013 | 2013 | Toshiba American Medical Systems Inc | The firm initiated this recall of certain Toshiba Diagnostic Ultrasound Systems because it has been found that under certain conditions, the temperature of the transducer head may increase in 2D+PW mode or 2D+CDI (POWER, ADF) +PW mode. When the system is used under both conditions listed below, the temperature of the transducer head connected to the system may increase in temperature and possibly | DESIGN: Device Design | Toshiba sent an Urgent Medical Device Correction letter dated August 5, 2013. along with a Reply Form via USPS return receipt mail. The letter identified the product, the problem, and the action to be taken by the customer. Until the software is corrected customers were instructed to change the preset settings as described in the letter. Customers were asked to complete and return the attached form and fax it to the toll free number. The form also can be sent via email to raffairs@tams.comCustomers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968.For questions regarding this recall call 714-730-5000. | Contact the recalling firm for information | 430 | 430 (1352 total) | Nationwide Distribution including and US Territory Puerto Rico | Open |  |  | 447 | Computer
Z-0116-2014 | 65471 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. | GE Centricity PACS | 2 | October 30, 2013 | 2013 | GE Healthcare It | GE has recently become aware of potential issues due to out of context issues associated with specific workflows of the Centricity PACS RA1000 workstation. These issues are as follows: 1) In Centricity RIS-IC, and a RIS-Driven integration, the "close" button closes the exam in RIS-IC, but does not close the exam in Centricity PACS RA1000 Workstation. 2) Interrupted workflow could break the | DESIGN: Software Design | URGENT MEDICAL DEVICE CORRECTION letters dated June 11, 2013 was mailed to all customers alerting them to these potential safety issues. The software was modified to correct these defects and a GE Healthcare service representative will contact each customer to make arrangements for installation of the new software version at no charge to the customer. Customers with concerns or questions about the market action can contact their local GE Healtchcare IT service representative or GE Healthcare's Remote Online Center - Customer Care Center at 1-800-437-1171. |  | 623 | 623 units | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV) and the countries of the UK, Taiwan, Switerzland, Sweden, Spain, Singapore, Saudi Arabia, Netherlands, Malaysia, Korea, Japan, Italy, Israel, India, Hong Kong, Germany, France, Denmark, China, Canada, Bermuda, Beligum and Australia. | Open |  |  | 623 | Computer
Z-0118-2014 | 64730 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Lossy compressed mammographic images and digitized film screen images may only be interpreted using an FDA approved monitor that offers at least 5 M pixel resolution and meets other technical specifications reviewed and accepted by the FDA. Typical users of this system are trained professionals, including but not limited to physicians, radiologists, nurses, medical technicians an assistants. | GE Centricity PACSIW | 2 | October 30, 2013 | 2013 | Ge Healthcare It | Multiple issues were identified with the GE Centricity PACS-IW system that could pose a safety risk to patients. These issues include:1) With Auto-Fetch in use another patient's image may appear without user's intention. 2.a) In the case where dose administration time and image acquisition time fall on either side of midnight, the value will be under reported and close to zero for all pixe | DESIGN: Software Design | GE Healthcare sent an IMPORTANT PRODUCT INFORMATION letter dated June 15, 2012 to direct accounts informing them of the Auto-Fetch functionality issue and advising them to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers informing them of all three of the safety issues. The letter again adivsed customers to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. The letter further infomed customers that all three errors can be resolved through the installation of a new version of the software. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. |  | 847 | 847 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, District of Columbia, Puerto Rico, and the U.S. Virgin Islands.Internationally to Austria, Azerbaijan, Bahrain, Belgium, Bosnia, Brazil, Canada, Columbia, Egypt, France, Germany, Ghana, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Libya, Lithuania, Mexico, Netherlands, Nigeria, Palestine, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Yemen.Center Recommended Depth - Consumers/User** | Open |  |  | 847 | Computer
Z-0125-2014 | 66473 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. Medtronic Activa RC and Activa PC Deep Brain Stimulators are indicated for Therapy for Parkinson's Disease - Bilateral stimulation of the internal globus pallidus (GPi) or the subthalamic nucleus (STN). Indicated for adjunctive therapy in reducing some of the symptoms of advanced, levodopa-responsive Parkinson's disease that are not adequately controlled with medication. Activa PC Deep Brain Stimulator is indicated for Therapy for Dystonia. Unilateral or bilateral stimulation of the internal globus pallidus (GPi) or the subthalamic nucleus (STN). Indicated as an aid in the management of chronic, intractable (drug refractory) primary dystonia, including generalized and segmental dystonia, hemidystonia, and cervical dystonia (torticollis), for individuals 7 years of age and older. Activa PC Deep Brain Stimulator is indicated for Obsessive-Compulsive Disorder. Bilateral stimulation of the anterior limb of the internal capsule, AIC, as an adjunct to medications and as an alternative to anterior capsulotomy for treatment of chronic, severe, treatment-resistant obsessive-compulsive disorder (OCD) in adult patients who have failed at least three selective serotonin reuptake inhibitors (SSRIs) | Activa PC and Activa RC | 2 | October 31, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding over stimulation or stimulation in the wrong area related to select Medtronic neurostimulators. Upon a specific set of conditions, typically related to device recovery from an over discharge, there is a potential for over stimulation or stimulation directed to a lead electrode other than what was intended. | DESIGN: Software Design | The firm, Medtronic, sent an "Urgent: Medical Device Correction" letter dated September 2013 entitled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter described the product, problem and actions to be taken. The letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. The customers were instructed to avoid over discharge; if a rechargeable device becomes over discharged, contact Medtronic Patient Services at (1-800-510-6735) to resolve the POR and reduce the likelihood of over stimulation.If you have questions, please contact Medtronic Neurostimulation Technical Services at 1-800-707-0933 weekdays 7am-6pm CST. |  | 38294 | 38,294 | Worldwide distribution: US (Nationwide) including Puerto Rico and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Open |  |  | 102457 | Computer
Z-0090-2014 | 65805 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution Mobile X-Ray System, Model DRXR-1, Verona Street Rochester, NY 14608Catalog Numbers 1019397, 1023415, 1023423 --- Made in USA by Carestream Health, Inc. 150 The device is designed to perform radiographic x-ray examinations on all pediatric and adult patients, in all patient treatment areas. | DRX Revolution Mobile XRay System | 2 | October 31, 2013 | 2013 | Carestream Health, Inc. | During an audit, Carestream discovered a DRX-1 System that captured an image which displayed an artifact. This issue was discovered in manufacturing. It occurs when the DRX product console software is updated to v5.6b software and the detector is not calibrated before use. | DESIGN: Software Design | Carestream sent a Urgent-Medical Device Recall letter dated July 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Carestream Health Inc. recently discovered an image artifact issue due to a detector not being recalibrated with the 5.6B software upgrade. As our installation instructions did not require the user to recalibrate during an upgrade, there is a possibility your detector could do the same. When this occurs an image artifact could be evident on images captured in the Extended Exposure mode. These artifacts could potentially contribute to misdiagnosis or the necessity of patient re-exposure.We request that you take a moment to verify that your detector(s) are calibrated. At our next visit, Carestream will also verify that your detector was re-calibrated and, if not, re-calibrate the detector for you. Please note this only affects systems upgraded in the field.For questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910, 8 am through 8 pm EST in the US. |  | 23 | 23 units installed total (11 domestically & 12 internationally) | Worldwide Distribution - US (nationwide) including the states of New York, Massachusetts, Pennsylvania, Illinois, Texas, California and Louisiana., and the countries of Australia, Canada, Finland, Italy, Switzerland and United Kingdom. | Open |  |  | 23 | Computer
Z-0016-2014 | 65161 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro Systems - Luminos dRF, Ysio and Uroskop Omnia with software versions VB10D, VB10F, VB10C and VB10E. Fluroscopic X-Ray Systems | Siemens RAD Fluoro Uro Systems | 2 | November 01, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has discovered through product monitoring a potential malfunction when using system operating with software versions VB10D, VB10F, VB10C and VB10. During a RAD examination the possibility exists that radiation may continue longer than expected and may result in unnecessary dose applied during an acquisition. Additionally, it is not possible for the user to stop the radiation manually; ho | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Siemens initiated this correction on May 6, 2013 by sending Update InstructionXP005/13/S. Firm subsequently issued a letter on June 3, 2013 to all accounts.1. Siemens Field Modification of the Luminos dRF, Ysio & Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, which improves the signal transmission for the automatic exposure control, is available as of May 2013, and systems will be updated within one year.2. The revised software fix for models Luminos dRF, Ysio & Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, will be provided to all identified customers at no charge.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 workingdays of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) districtoffice coordinator noted below are to be included in the notification.2. Siemens Medical Solutions USA, Inc., will notify the FDA district office inPhiladelphia, Pennsylvania, with reports of progress regarding this fieldcorrection. |  | 23 | 23 | Nationwide Distribution | Open |  |  | 23 | Computer
Z-0138-2014 | 65782 | neurological stereotaxic instrument | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator - 3D Line stereotatic Hardware AccessoriesThe device is part of the class of medical devices that are used in radiotherapy for the treatments of head tumors. | Stereotactic Circular Collimator 3D Line stereotatic Hardware Accessories | 2 | November 04, 2013 | 2013 | Elekta, Inc. | Potential for clinical errors. |  | Elekta sent an Important Field Safety Notice dated 7/2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the issue and recommendation to resolve either by upgrading their system or to discontinue a clinical use of the device.For further questions please call ( 770 ) 300-9725. |  | 89 | 89 | Nationwide distribution: USA including the states of AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN and WV. | Open |  |  | 89 | Computer
Z-0155-2014 | 66587 | table, surgical with orthopedic accessories, ac-powered | JEA | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ORBIT PT.The ORBIT PT works in conjunction with an TRUMPF mobile operating table to automate the patient transport process. | ORBIT PT | 3 | November 05, 2013 | 2013 | Trumpf Medical Systems | An odor and smoke developed - The cause was determined to be a power supply unit sitting in the controller column. | DESIGN: Component Design/Selection | The user facility was contacted via phone in early August. They were contacted 9/17/2013 to schedule the replacement of the power pack because the new power pack became available. |  | 1 | One | US Distribution: NJ only. | Open |  |  | 1 | Computer
Z-0168-2014 | 66239 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS INTEGRA 800 Analyzer with software 9864.C2Cobas. COBAS Part Numbers 28122474001, 28122474692, 04559126001, 04589459970 corresponding to the Analyzer with or without closed tube sampling accessory A fully automated system for clinical chemistry analysis intended for the in vitro quantitative/qualitative determination of analytes in body fluids. | COBAS INTEGRA 800 Analyzer | 2 | November 05, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof | DESIGN: Software Design | Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed toRoche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336For questions regarding thisr ecall call 317-576-3911. | Contact the recalling firm for information | 527 | 527 units | Nationwide Distribution and Puerto Rico | Open |  |  | 527 | Computer
Z-0170-2014 | 66408 | ventilator, continuous, facility use | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-C2 Ventilator:Manufactured by Hamilton Medical AGCH-7402 Bonaduz, Switzerland;Imported and distributed by Hamilton Medical, Inc. Reno, NV 89502The HAMILTON C2 ventilator is intended to provide positive pressure ventilatory support to adults, pediatrics, infants and neonates depending on its configuration. Intended areas of use: In the intensive care ward or in the recovery room; During transfer of ventilated patients within the hospital. | Hamilton C2 Ventilator | 2 | November 06, 2013 | 2013 | Hamilton Medical, Inc. | Manufacturer defined limits for improperly functioning loudspeakers of the HAMILTON C2 alarm system were exceeded. | DESIGN: Device Design | Hamilton Medical sent an Medical Device Safety Alert and Corrective Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The corrective action does not require any action by the operator. In the case of an alarm situation as described in the letter, users should follow the instruction for use (Section 8 - responding to alarms in the operator manual) and have the ventilator serviced by an authorized service technician. An addendum to the Operators Manual as a result of the field correction was provided to customers with the Safety Alert Letter.Manufacture will develop and provision a test protocol for the service tech which test the proper functioning of the loudspeaker during the annual service inspection.We appreciate your support in this matter and sincerely regret any inconveniences that this action may cause you. We consider this action as necessary to ensure that our customers receive only safe and effective products with high quality.Further questions please call (775) 426-6331 Ext. 210. |  | 316 | 316 devices distributed in the US. | Nationwide distribution: USA including Puerto Rico and the states of IL, MT, IN, ND, AZ, MT, KY, NC, OH, MA, MO, VA, CA, NV, TX, WA, NY, GA, SC,MN, WY, PA, ND, UT, NE, CO and MI. | Open |  |  | 316 | Computer
Z-0169-2014 | 66659 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaCoreÂŽ software. Versions subject to the limitation:6.0, 6.1, 6.2, 6.3 (except 6.3.1 ), 6.4 (except 6.4.4 and 6.4.5), aswell as their upgrades, and 6.5. The VitreaCore (ViTALConnect) system is a medical diagnostic software system intended to process, analyze, review, and distribute multi-dimensional digital Images acquired from a variety of imaging devices including: CT, MR, CR/DR/DX, SC, US, NM, PET, XA, and RF, etc. The VitreaCore (ViTALConnect) is not meant for primary Image Interpretation In mammography. | VitreaCore software | 2 | November 06, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. has found an issue in VitreaCore software though internal testing. There is a potential for incorrect and possibly reversed image orientation for certain snapshots taken in VitreaCore software versions 6.0, 6.1, 6.2, 6.3 (except 6.3.1), 6.4 (except 6.4.4 and 6.4.5), as well as their upgrades, and 6.5. | DESIGN: Software Design | Vita Images, Inc. sent an Urgent Medical Device Recall letter on October 21, 2013, to all affected customers via USPS First class mail. The letter described the Reason for Voluntary Recall (Nature of software issue, Potential scenarios which exhibit the issue), Risk to Health (How to confirm if the Snapshots are correct), Actions to be Taken by Customer (Recommended Next Steps which includes to complete and sign the enclosed Effectiveness Check Form and send to fieldnotices@vitalimages.com or fax to 1-952-540-3717), and Actions Being Taken by Vital Images (Vital Images Contact Information: If you have questions, direct Vital Images customers and customers of Toshiba America Medical Systems (TAMS) should contact Vital Images Customer Support at +1-800-208-3005, or by e-mail to support@vitalimages.com; customers through a distributor will be contacted by that distributor).For questions regarding this recall call 952-487-9574. |  | 598 | 598 | Nationwide Distribution - all states in continental USA including DC but not OK or RI.OUS: Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Honduras, Hong Kong, Iceland, Ireland, Italy, Korea, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, r Saudi Arabia, Serbia, Singapore, Slovenia, South Africa,Spain , Sweden, Switzerland, Thailand, Turkey, Ukraine. | Open |  |  | 598 | Computer
Z-0171-2014 | 63859 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Handset component of the Handicare USA's Minivator Bath Bliss Reclining Bath Lifts (Bath Bliss), Part Number BP300-100/BP300-400RSUS (handset) - Model Numbers BBR0311USAB (blue) and BBR0311USSS (silver)Handset component devices operate the Minivator Bath Bliss Reclining Bath Lift, a non-AC-powered patient lift consisting of a motorized actuator, powered mechanical seating device, temporarily secured in a bath, and used to lift and lower a patient in the vertical or reclining position while bathing. | Handicare USA's Minivator Bath Bliss Reclining Bath Lifts (Bath Bliss), handset | 2 | November 06, 2013 | 2013 | Handicare Accessibility | Firm became aware that two handsets in the affected batches became inoperable while the user was in the tub. |  | Handicare sent an Urgent Medical Device Recall letter dated November 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately advise all of their customers affected by the recall letter to immediately remove all affected product from use. Upon receipt of affected product, replacement handsets would be sent to customers. Customers with questions were instructed to call 866-276-5438 or email Customerservice.usa@handicare.com. For questions regarding this recall call 610-266-5260. | Contact the recalling firm for information | 3301 | 3301 handsets | Worldwide distribution: USA (Nationwide) and country of: Canada. | Open |  |  | 3301 | Computer
Z-0180-2014 | 66067 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX -Mobile Retrofit Kit, Catalog number: 1019397. Made in U.S.A. by Carestream Health, Inc. 150 Verona Street Rochester, NY 14608to perform radiographic x-ray examinations on all pediatric and adult patients | DRXMobile Retrofit Kit | 2 | November 07, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. has conducted a voluntary field action which includes the installation of a software patch on the following: DRX-Revolution Mobile X-Ray, DRX-Mobile and DRX Transportable Systems with DirectView Version 5.5 and 5.6 console software when used with DRX-1 Detector version 69 firmware. Carestream recently received reports of a non-diagnostic image on exams when a Normal Exposure | DESIGN: Software Design | The firm, Carestream Health Inc., sent an "Urgent: Medical Device Recall" letter dated August 29, 2013 to its customers. The letter described the product, problem and actions taken. Carestream has developed a software patch that will allow Normal Exposure Mode and Extended Exposure Mode images to be taken in any order and will install the patch in all DRX-Revolution, DRC-Mobile and DRX-Transportable Systems with software version 5.5 or higher and Detector firmware version 69 or higher. No DRX Transportables were affected. Your Carestream service representative will install the patch at your next service call. If you have questions or concerns, or if you experience an issue prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8am through 8pm Eastern Standard Time. |  | 364 | 364 units | Nationwide distribution:AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MI, MO, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WI, and WV. | Open |  |  | 554 | Computer
Z-0197-2014 | 66120 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library.Intended to be used for the controlled administration of intravenous fluids. | SIGMA SPECTRUM Infusion Pump with Master Drug Library | 2 | November 08, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Drug Library due to reports of upstream occlusion alarms and air-in-line alarms occurring during infusions. |  | On June 10, 2013 Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Library. The notices were dated June 5, 2013 and were distributed via first class mail. Until further information is provided, users should continue to follow the SIGMA SPECTRUM Infusion Pump Operator's Manual, which identifies a number of environmental and clinical usage factors that contribute to upstream occlusion and air-in-line alarms. For high-risk patients, consider having additional devices as backup. If the biomedical engineer finds that the sensor is not functioning properly, contact Baxter Healthcare at 800-356-3454 (choose option 1). For clinical questions, please contact Medical Information Services at Baxter at 800-933-0303, Monday through Friday 8:00am to 5:00 pm Central Time. |  | 273835 | 273,835 units | Worldwide Distribution-USA including Puerto Rico and Canada. | Open |  |  | 273835 | Computer
Z-0201-2014 | 66725 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Titan 3T MRI System, MRT-3010A/5.Product Usage: Usage:MRI system. | Vantage Titan 3T MRI System, MRT3010A/5 | 2 | November 13, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is recalling due to a potential problem with Specific Absorption Rate (SAR)values when using the Toshiba Titan 3T System. When the system is used in combination with either the QD Head Coil (MJQH-142A) or the Large KneeSPEEDER (MJAJ-182A), SAR values lower than the actual SAR values may be displayed on the operatingmonitor. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Toshiba America Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION". letter dated October 28, 2013 via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. The letter listed Device, Serial Number, reason for recall, problem, corrective actions, request to customers, and contact information. |  | 18 | 18 units | USA Nationwide Distribution in the states of: CA, FL, GA, IA, NV, NY, LA, IN, PA, OR, TX, and WI. | Open |  |  | 18 | Computer
Z-0273-2014 | 66664 | controller, foot, handpiece and cord | EBW | Dental handpiece and accessories. | Dental | Dental | 510(k) | Osseocare Pro Console & Set Osseocare Pro, Model 1600870-001, 1700470-001, and 1700471-001.Device Name: Osseocare Pro console & Set Osseocare Pro.Model: 1600870-001, 1700470-001, 1700471-001.Marketing status: Class I, under K092214.Device description: Software-based, electronically controlled tabletop device.In dentistry for dental surgery, endodontics and implantology. | Osseocare Pro Console & Set Osseocare Pro | 2 | November 13, 2013 | 2013 | Nobel Biocare Usa Llc | The recall was initiated because there is a potential for malfunction of the OsseoCare Pro when used with the OseeoCare Pro iPad application software release 1.2.0.7. |  | Bien Air Dental send and Urgent Field Safety Notice dated October 24, 2013, to all affected customers. The letter notified customers of the recall, description of the problem, description of the product (OsseoCare Pro iPad app software release 1.2.0.7), action to be taken by user, an alternative solution, and contact details. The letter was also accompanied by an acknowlegement form. Customers with questions were instructed to call 1-800-433-2436.For questions regarding this recall call 714-282-4800. |  | 1082 | 1082 (391 US) | Worldwide Distribution - USA (nationwide) United States and Internationally to Austria, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Latvia, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Turkey, and the United Kingdom. | Open |  |  | 1082 | Computer
Z-0270-2014 | 66550 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AW VolumeShare 4 (ADVANTAGE WORKSTATION 4.3.), AW VolumeShare 5 (AW VOLUMESHARE 5 WITH ANGIOVIZ OPTION), AW Server.K052995: Advantage Workstation 4.3 is a review station, which allows easy selection, review, processing and filming of multi-modality DICOM images from a variety of diagnostic imaging systems. When interpreted by a trained physician, filmed or displayed images on the AW monitor may be used as a basis for diagnosis, except in the case of mammography images. K110834: AW VolumeShare 5 is a review workstation, which allows easy selection, review, processing and filming of multi-modality DICOM images from a variety of diagnostic imaging systems. When interpreted by a trained physician, filmed or displayed images on the AW monitor may be used as a basis of diagnosis, except in the case of mammography images. K081985: AW Server is a medical software system that allows multiple users to remotely access AW applications from compatible computers on a network. The system allows networking, selection, processing and filming of multimodality DICOM images. Both the client and server software are only for use with off the shelf hardware technology that meets defined minimum specifications. | AW VolumeShare 4 (ADVANTAGE WORKSTATION 4.3.), AW VolumeShare 5 (AW VOLUMESHARE 5 WITH ANGIOVIZ O | 2 | November 13, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the image orientation within GSI viewer on AW products.The GSI Viewer application running on AW/Volume Share or on the AW Server when making a batch and rotating the batch group clockwise 180 degrees (Inferior to Superior) will display images in the Superior orientation. The orientation annotation on the image is correc | DESIGN: Software Design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated September 19, 2013 to its consignees/customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (800-437-1171 for USA and Canada or contact local GE Healthcare Applications Specialist). The customers were instructed to "Please ensure that all potential users in your facility are made aware of this safety notification and the action taken". GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction. This activity will be performed at no cost to you.A second updated letter was sent on October 22, 2013 to correct the Affected Product Details communicated in the previous letter dated September 19, 2013. The new statement reads "All versions of GSI Viewer prior to version 2.00-OM on AW and prior to version 2.20- OB on AW Server are affected".If you have any questions, please contact us at 800-437-1171 (USA and Canada) or contact local GE Healthcare Applications Specialist. |  | 1458 | 1458 (336 US, 1122 OUS) | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT,DE, FL,GA, HI, IL, IN, IA, KS, MD,MA, MI,MN, MS, O. NV, NH, NJ, NY, NC, OH, OK, OR, PA, SD, TN, TX , UT, VA, WA, WI. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, BYELO RUSSIAN, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HON KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LIBYA, LITHUANIA, LUXEMBOURG, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, and UNITED KINGDOM.. | Open |  |  | 1458 | Computer
Z-0269-2014 | 66560 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips brand IntelliSpace PACS 4.4, Image Management System; The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. Product is manufactured and distributed by Philips Healthcare Informatics, Inc. Foster City, CAIntelliSpace PACS are image management systems intended to be used by trained professionals, including but not limited to physicians, nurses and medical technicians. The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. The software performs digital image processing, measurement, communication and storage. iSite supports receiving, sending, printing, storing and displaying studies received from the following modality types via DICOM: CT, MR, NM, US, XA, PET, DX, DR, RE, RT, MG, SC, VL, as well as hospital/radiology information systems. | IntelliSpace PACS | 2 | November 13, 2013 | 2013 | Philips Healthcare Informatics, Inc. | Philips Healthcare became aware of a software problem with the IntelliSpace PACS system, in that a timing issue has been identified that may result in a condition where the device does not store additional images received before processing is complete. The missing images may contain pathology which may result in a different diagnosis if they had been available at time of interpretation. |  | Philips sent an Urgent Field Safety Notice dated October 14, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Please review the following with all radiology technicians who use lntelliSpaceBY CUSTOMER/USER PACS:" Assure you and your colleagues are aware of the problem." Verify that all images sent for each study are present on the iVaultbefore the images are purged from the modality.o Use the image count from the lower right corner of thethumbnails as seen in Enterprise only. Single images (such asCR, DX and MG) show no image count, so the user shouldcount those thumbnails without image counts as one. Add all thenumbers up to come up with the total number of viewableimages, stored to lntelliSpace." Compare number of images sent from the modality and those sent fromother 3rd party source to image count in IntelliSpace PACS." If the image count on the modality or 3 party source indicates there aremore images to the study than are present on the iVault, the studyshould be resent from the modality and/or 3 party source." Resending the studies will typically correct the problem but thetechnician should again verify the image count is correct.If you need any further information or support concerning this issue, pleaseINFORMATION AND contact your Philips representative or contact Care at isitecare@philips.com. Please call (650) 293-2300 for further questions. |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) and country: Canada. | Open |  |  | 71 | Computer
Z-0271-2014 | 65792 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza image processing systemimage processing radiological system | Siemens syngo.plaza | 2 | November 13, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of an unintended behavior when using the syngo.plaza. The "Rename" functionality can cause an unintended patient merge if the patient is sent to another system. This issue affects customers who transfer patient data between two or more syngo.plaza or syngo Imaging XS systems. | DESIGN: Software Design | Siemens sent a Customer Safety Advisory Notice dated August 3, 2011, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to forward the Customer Safety Advisory Notice to their customers. If necessary, customers should have the letter translated technically into the local language and copy it themselves. A letter was sent to direct accounts to inform them of the issu and provide instructions to avoid encountering the problem.For questions regarding this recall call 610-219-6300. |  | 19 | 19 | Nationwide Distribution including MO, MA, NY, WI, OH, NY, CA, FL, NC, MN, and WA. | Open |  |  | 19 | Computer
Z-0219-2014 | 66306 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Carestream Vue PACS; Virtual Enterprise SW StandAloneThe CARESTREAM PACS is an image management system whose intended use is to provide completely scaleable local and wide area PACS solutions for hospital and related institutions/sites, which will archive, distribute, retrieve and display images and data from all hospital modalities and information systems. | Carestream Vue PACS; Virtual Enterprise SW StandAlone | 2 | November 13, 2013 | 2013 | Carestream Health Inc. | Carestream Health, Inc. has issued a recall when using the CARESTREAM Client Workstation where under very specific conditions, some images from multi-slice studies such as CT and MR may not load into the viewer application due to a software problem. | DESIGN: Software Design | Carestream sent an Urgent Medical Device Recall letter dated September 13, 2013, to all affected customers via Fed Ex on September 16, 2013, informing them that a software problem may have the potential to occur under certain conditions when using the CARESTREAM Client Workstation. When certain conditions are met, it appears that the child background process completes its tasks before the parent process and as a result the parent process considers its work done and terminates. In the event that a software problem does occur, the scan may not be fluent or the image numbers may not be sequential. A software patch is currently being deployed by Carestream Health. Customers were instructed that they would be contacted by a Carestream Service Engineer to coordinate the deployment. Customers with questions were instructed to call 1-866 737 2487 or they may call the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time.For questions regarding this recall call 1-800-328-2910. |  |  |  | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, IA, IL, MA, MD, ME, MN, MO, NC, NH, NV, NY, OH, OK, PA, RI, TN, UT, VA WA, and WI and Internationally to Canada, Japan, and Kuwait. | Open |  |  |  | Computer
Z-0281-2014 | 66600 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station Radiation Therapy Treatment Planning System, Model Numbers 3.0, 3.5 and 4.0 | RayStation | 2 | November 14, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Ray Search Laboratories found an issue with the dose calculation for Elekta motorized wedges in Ray Station 3.0, 3.5 and 4.0. The error presents when a wedge is used for a beam and the wedge angle differs from 60 degrees. | DESIGN: Software Design | Medical Device Correction Letters dated 10/23/13 were e-mailed to the user accounts on 10/23/13. |  | 195 | 195 units domestically | US Distribution in the states of: California, Colorado, Delaware, Florida, Maryland, Missouri, New Jersey, New York, Ohio, Pennsylvania, Tennessee, and Virginia. | Open |  |  | 195 | Computer
Z-0279-2014 | 66506 | stimulator, electrical, implanted, for parkinsonian tremor | MHY |  |  | Neurology |  | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Activa SC, models 37602 and 37603. Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. | Activa PC, Activa RC, Active SC | 2 | November 14, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding Loss of Stimulation of specific Implantable Neuro-Stimulators under the following specific conditions: Switching between groups with Multiple programs or creating program groups. Therapy can be restored with the Physician Recharge mode (RPM) of the Implantable Neurostimulator Recharger (INSR). | DESIGN: Software Design | In September 2013 Medtronic issued a medical device correction letter titled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. Medtronic field representatives will be updating the model 8870 software application card to version AAR/01, which is contained in the physician programmer. A Physician Reply Form was asked to be signed and returned by fax or mail for acknowledging the receipt of the Urgent Medical Device Correction letter. When the Medtronic field rep updates the software application card of the physicains' programmer, they will complete a Rep Reply For: 8870 Card Update to AAR/01. |  | 62634 | 62,634 | Worldwide distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Open |  |  | 126797 | Computer
Z-0329-2014 | 65352 | pump, air, non-manual, for endoscope | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus Endoscopic Flushing Pump, Model OPF-2.Facilitates washing of tissues to remove blood, feces and other organic matter during endoscopic procedures. | Olympus Endoscopic Flushing Pump, Model OPF2 | 2 | November 15, 2013 | 2013 | Olympus America Inc. | There is a possibility that the pump head was not shipped in the locked position. If the pump head is in an unlocked position this may, under certain circumstances, cause the pump to stop and go into power stand-by mode. | DESIGN: Component Design/Selection | Notification letters were sent to customers on May 20, 2013. |  | 263 | 263 | Worldwide distribution: USA (nationwide) and the country of Canada. | Open |  |  | 263 | Computer
Z-0333-2014 | 66635 | biomicroscope, slit-lamp, ac-powered | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | Zeiss brand IOLMaster, Version 1-5, Biomicroscope, Slit-Lamp, AC-Powered, Product is manufactured by Carl Zeiss Meditec AG (Jena),Carl Zeiss Promenade 10, Jena, Germany 07745; and distributed by Carl Zeiss Meditec, Inc., Dublin, CAThis device is intended for use primarily by physicians and health care workers and may be only used under the supervision of a physician. This device will not be sold to the general public. The IOLMaster is intended for the biometric determination of ocular measurements of axial length, anterior chamber depth, corneal radius, white-to white (WTW), and for the measurement of pupil size and deviation of the visual axis from the center of the pupil. For patients who are candidates for intraocular lens (IOL) implantation, the device also performs calculations to assist physicians in determining the appropriate IOL power and type for implantation | IOLMaster | 2 | November 15, 2013 | 2013 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value. |  | The firm, Zeiss, sent an "Urgent Medical Device Correction" letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to immediately update the Laser Interference Biometry (ULIB) lens constants, provided as an Attachment 1 and complete and return Attachment 2- Confirmation Form via mail to: Stericycle, Inc., 6026 Lakeside Blvd., Indianapolis, IN 46278; Attn: Zeiss IOLMaster.If you have any questions, call the following number for assistance: 1 (866) 891-1922. |  | 4907 | 4,907 units | Nationwide distribution. | Open |  |  | 6108 | Computer
Z-0330-2014 | 64806 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. The Centricity Laboratory System product is designed to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. It integrates with other Hospital Information Systems (HIS) through HL7 interface by order entering and also report processing. The Centricity Laboratory System is intended to interface with various lab results. Centricity Laboratory System is intended to be used by laboratory technologists, technicians, analyst and other trained/site authorized system users | Centricity Laboratory Core Lab | 2 | November 15, 2013 | 2013 | GE Healthcare It | GE Healthcare has become aware of a potential safety issue associated with the use of GE Centricity Laboratory where panel reflex does not occur if DIP.Specimen is blank. Panel reflexing fails to occur when ad-hoc re-routing functionality is used and the following are true: 1) The ordered panel is a slave panel, and; 2) The panitem specimen group field of one or more of the ordered items on the ma | DESIGN: Software Design | GE Healtchare sent an Urgent Medical Device Correction letter dated March 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers review their workflows to determine whether this series of events could happen at their sites. Customers who self-identified this series of events as a possible workflow are recommended to contact their GE service representative for a software correction. These customers may also elect to identify master panels that may be re-routed ad hoc and then to ensure that the specimen group field is updated to NULL if currently set to blank.Customers with any questions about this issue are advised to contact their GE service representative or the GE Healthcare IT Help Desk: North America at tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-277-4468. |  | 7 | 7 units | Worldwide Distribution - USA (nationwide) including GA and Internationally to Canada and the United Kingdom. | Open |  |  | 7 | Computer
Z-0343-2014 | 66069 | electrode, ion specific, sodium | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA 1200 System Software: 1) V2.00, 2) V2.01.Performs assays for general and specialty chemistries, methods for specific protein measurement, drugs-of-abuse testing and therapeutic drug monitoring. | ADVIA 1200 System Software: 1) V2.00, 2) V2.01 | 2 | November 18, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics issued a Urgent Medical Device Correction due to Ration Parameters Flagging Behavior for All Software Versions utilized on ADVIA 1200, 1650, 1800 or 2400 Chemistry Systems. | DESIGN: Software Design | On September 3, 2013, Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice for the ADVIA 1200, 1650, 1800 and 2400 Chemistry Systems due to a software issue. The software issue has caused the Chemistry Systems to generate reports without error flags. Notices were sent to domestic customers via U.S. and outside the US the distribution is determined at the country level. The customers were instructed to contact Siemens Healthcare Diagnostics Technical Solution Center if they have technical questions regarding this issue at 877-229-3711 (prompt 18). | Contact the recalling firm for information | 625 | 625 units | Worldwide Distribution-USA (nationwide) and the countries of Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Vietnam. | Open |  |  | 3759 | Computer
Z-0337-2014 | 64805 | system, monitoring, perinatal | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal and Centricity Intensive Care is intended to be used in clinical departments healthcare delivery systems such as a Labor and Delivery, Postpartum Maternal Care, newborn Nursery, critical care units including Neonatal Intensive Care Units (NICU), and may also be used in physicians offices and outpatient clinics. This product is primarily intended to serve the purpose of electronic documentation of clinical data and is designed to accept, transfer, display, calculate, store and manage clinical data. System capabilities provide the user with the ability to acquire data from the medical devices and to document, annotate, display, store, print, export and retrieve patient clinical practice information. This product is intended for professional use only. | Centricity Perinatal and Centricity Intensive Care | 2 | November 18, 2013 | 2013 | GE Healthcare It | 1) Under rare network conditions, a Centricity Perinatal process may maintain an incorrect count of Patient slots in use which may lead to an inability to access the patient roster. The user is temporarily unable to review or edit form and chart information. This could result in a potential delay of care; however active fetal strips and permanent patient data storage are not affected. 2) When | DESIGN: Software Design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 11, 2013, to all affected customers. The letters requested that direct accounts ensure that all potential users of the software were made aware of the potential problems and the recommended actions. Direct accounts were asked to return a customer response form and to return it to GE Healthcare via fax at 847-939-1522 or e-mail at QARAFMl1@ge.com.Version 6.94 of the software has been modified to correct these defects.GE Technical Support: 800-433-2009. |  | 1132 | 1,132 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Denmark, Germany, Hong Kong, Korea, South Korea, Thailand and United Kingdom. | Open |  |  | 1132 | Computer
Z-0351-2014 | 66727 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390. It is a patient monitor with arrhythmia detection or alarms.The qube is intended for use under the direct supervision of a licensed healthcare practitioner, or by personnel trained in proper use of the equipment in a hospital environment. These devices determine a) when an alarm condition is violated; b) the alarm priority (i.e. high, medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390 | 2 | November 18, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390, may fail to power ON due to incorrect programmed microcontroller component. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare qube Compact Monitor" letter, dated 01 November 2013, to the US consignees/customers on November 1, 2013. Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to advise staff of the potential for the monitor to fail to turn ON.Spacelabs Healthcare will contact their customers at the earliest possible date to schedule a convenient time for Spacelabs to correct all of your facility's affected monitors at no cost. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 54 | 54 units total (42 units in the US and 12 outside US) | Worldwide distribution: US (nationwide) including states of: Georgia, Idaho, New York, Oklahoma, Tennessee, Texas, and Wyoming and in countries of: Bahrain, Canada, France, Great Britain, Poland, and The Netherlands. | Open |  |  | 54 | Computer
Z-0368-2014 | 66662 | system, imaging, pulsed doppler, ultrasonic | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System, Model 795200, 795201, 795202.Intended for diagnostic ultrasound imaging and fluid flow analysis. | EPIQ 7 Ultrasound System | 2 | November 20, 2013 | 2013 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. Using the International System Date Format may result in calculation errors in Gestational Age Gestational Age (GA) and Established Due Date (EDD). | DESIGN: Software Design | Affected customers will be notified via the Urgent-Field Safety Notice (FSN79500286), Philips Model EPIQ 7 Ultrasound system, dated 2013 October. Customers are informed that using the International System Date Format in the Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. may result in calculation errors in Gestational Age and Established Due Date (EDD). Philips will upgrade customer systems to software version 1.1, which corrects the issue, for a free of charge. |  | 284 | 284 units total (183 international; 100 in the US). | Worldwide Distribution-USA (nationwide) including the states of CT, HI, IL, IN, KY, MA, ME, MI, NC, OR, PA, TX, and WA, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Ireland, Republic of Korea, Latvia, Netherlands, New Caledonia, New Zealand, Norway, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, and United Kingdom. | Open |  |  | 284 | Computer
Z-0367-2014 | 66631 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, & Brivo NM 615.The GE Discovery NMJCT 670 system is a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy. The GE Discovery NM/CT 670 system, combining Nuclear Medicine (NM) and Computed Tomography (CT) systems, is intended to produce: NM System: General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and imaging features designed to enhance image quality. The scanning modes include planar mode (Static, Multi-gated, Dynamic and Whole body scanning) and tomographic mode (SPECT, Gated SPECT, Whole body SPECT). The acquisition types include single and multi-isotope/multi peak frame/list mode single-photon imaging. The imaging-enhancement features include assortment of collimators, gating by physiological signals, and real-time automatic body contouring. CT System: Cross sectional images of the body by computer reconstruction of X-Ray transmission data taken at different angles and planes, including Axial, Cine, Helical, Cardiac, and Gated acquisitions. These images may be obtained with or without contrast. The CT system is indicated for head, whole body, cardiac and vascular X-Ray Computed Tomography applications NM + CT System: Combined, hybrid SPECT and CT protocols, for CT-based SPECT attenuation corrected imaging as well as functional and anatomical mapping imaging (localization, registration and fusion). The GE Discovery NM/CT 670 system may include signal analysis and display equipment, patient and equipment supports, components and accessories. The system may include data and image processing to produce images in a variety of trans-axial and reformatted planes. The images can also be post processed to produce additional images, imaging planes, and analysis results. The system may be used for patients of all ages. The GE Discovery NM 630 system is a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and in assessment of organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy. The GE Discovery NM 630 system is a Nuclear Medicine (NM) system, which is intended to yield General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and imaging features designed to enhance image quality. The scanning modes include planar mode (Static, Multi-gated, Dynamic and Whole body scanning) and tomographic mode (SPECT, Gated SPECT, Whole body SPECT). The acquisition types include single and multi isotope/multi peak frame/list mode single-photon imaging. The imaging-enhancement features include assortment of collimators, gating by physiological signals, and real-time automatic body contouring. The GE Discovery NM 630 system may include signal analysis and display equipment, patient and equipment supports, components and accessories. The system may be used for patients of all ages. - The GE Optima NM/CT 640 system is intended for General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body. It includes a general purpose Nuclear Medicine (NM) system using a variety of scann | GE Healthcare Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, & Brivo NM 615 | 2 | November 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related to a SPECT scan procedure used for pulmonary imaging.Contact between the Detector (GE Healthcare Discovery NM/CT 670, and/or Optima NM/CT 640, and/or DiscoveryNM630, and/or Brivo NM 615 all configurations) and the Patient's elbow may occur during SPECT protocols using 450 angular steps in scenarios when a Patients a | TRAINING: Employee Error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated September 25, 2013, to all affected customers. The letter identified the product the problem and the actions to be taken by the customer.We plan to correct all of the affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, & Brivo NM 615 systems that were distributedThe plan for public warning is to contact all customers with an affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, & Brivo NM 615 system by means of an Urgent Medical Device Correction letter. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to 3 titles within the affected accounts; Managers of Nuclear Medicine, Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology. A GE Healthcare Field Engineer will perform the required software update on each affected system.Contact Information (GE Healthcare Call Center: USA or Canada 800-437-1171 and Other countries:( 262-896-2890). |  | 464 | 464 (143 US, 321 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MI, ME, MD, MA, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WV and WI., and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY,. JAPAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, NETHERLANDS, NEW ZEAL NAD, NORWAY, PAKISTAN, POLAND,QUATAR, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN,K THAILAND, TURKEY and UNITED KINGDOM. | Open |  |  | 464 | Computer
Z-0373-2014 | 66459 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 ultrasound system. GE Vivid E9 is a Track 3 diagnostic ultrasound system, which is primarily intended for cardiac imaging and analysis, but which also includes vascular and general radiology applications. The Vivid E9 incorporates a variety of electronic array transducers operating in linear, curved linear, sector/phased array or matrix array format, including two dedicated CW transducers and several real time 3D transducers. It consists of a mobile console with keyboard control panel; color LCD/TFT touch panel, LCD color video display and optional image storage and printing devices. It provides high performance ultrasound imaging and analysis and has comprehensive networking and DICOM capability. | Vivid E9 ultrasound system | 2 | November 21, 2013 | 2013 | GE Healthcare, LLC | GE became aware of a potential safety issue due to a system slow down and system lock up associated with the software of Vivid E9 ultrasound system. During a critical procedure the Vivid E9 Scanner may gradually become unresponsive and eventually lock up, with the result that the vivid E9 Scanner must be rebooted. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | An Urgent Medical Device Correction letter, dated 10/11/2013, was sent to 3 departments within the affected hospitals. The letter described the safety issue, and identified affected devices. The letter stated to reboot the scanner in an unresponsive /lock up condition. Also if the system is being used for an interventional procedure, the TEE probe should remain in the patient while system is rebooting. A GE healthcare service engineer will install a correction to affected devices. |  | 710 | 710 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belarus, Belgium, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Malaysia, Netherlands, Norway, Oman, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open |  |  | 710 | Computer
Z-0385-2014 | 66605 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Combo Blood Glucose Meter (as a part of the ACCU-CHEK Aviva Combo System)\The ACCU-CHEK Combo System is indicated for the treatment of insulin-requiring diabetes and for the quantitative measurement of glucose in fresh capillary whole blood from the finger. | ACCUCHEK | 2 | November 22, 2013 | 2013 | Roche Diagnostics Operations, Inc. | There is a software synchronization issue with the ACCU-CHEK Combo system. In rare cases, when the Manual Pump option is chosen on the ACCU-CHEK Aviva Combo meter, there is a possibility of receiving an incorrect bolus advice recommendation that may cause a temporary under delivery of insulin. | DESIGN: Software Design | The firm, Roche Diagnostics, sent an "URGENT MEDICAL DEVICE CORRECTION (UMDC)" letter dated October 29, 2013, via UPS to their consignees/customers. The letter described the product, problem and actions to be taken. The UMDC informed them of the software synchronization issue with the ACCU-CHEK Combo system that may cause a temporary under delivery of insulin. The customers were instructed to ensure they properly deliver the correction bolus when utilizing the Manual Pump option, as instructed in the Owners Booklet; and complete and return the Business Reply Card via fax to 1-877-282-0223. The Distributors were instructed to provide the UMDC to their direct customers. The HCPs were instructed to remind their patients to follow the instructions in the Accu-Chek Aviva Combo Meter Advanced Owners Booklet when utilizing Manual Pump option.Questions about the information contained in this UMDC are directed to ACCU-CHEK Pump Support, 24 hours a day, seven days a week at 1-800-688-4578. |  | 6595 | 6,595 units | Nationwide Distribution. | Open |  |  | 6595 | Computer
Z-0389-2014 | 66755 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) Product Usage:The V Series Monitor is intended for intra hospital use under the direct supervision of a licensed healthcare practitioner. The indications for use for the V Series include the monitoring of the following human physiological parameters: 1) ECG waveform derived from 3, 5, 6 and 12 lead measurements 2) Heart Rate 3) Pulse Oximtery 4) ST Segment 5) Arrhythmia Detection 6) Non Invasive Blood Pressure (IBP) 7) Cardiac Output (CO) 8) Respiratory Gases 9) Respiration Rate 10) Temperature The V series monitor has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. | V Series Monitor, Mindray DS USA, Inc | 2 | November 25, 2013 | 2013 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified an issue with V Series Monitor where the V Dock power pins connecting the V Series Monitor to the V Dock may lose spring tension, potentially causing the screen to temporarily go blank, lose communication with the central station monitor (when one is in use) or shut down. No patient related events have been reported to Mindray concerning the issue. | PRODUCTION CONTROLS: Environmental Control | Mindray sent an Urgent Corrective Action letter dated October 7, 2013 to all of their customers. The letter identified the affected product, problem and actions to be taken. Customers should contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday from 830 am to 530 pm to arrange for the replacement of the affected product. |  | 1839 | 1839 units | Worldwide Distribution - USA Nationwide | Open |  |  | 1839 | Computer
Z-0374-2014 | 66231 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | (Siemens Rad Fluoro Uro Systems) Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10Image intensified fluoroscopic x-ray system | Siemens Rad Fluoro Uro Systems | 2 | November 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A potential malfunction and hazard to patients exists when using the Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10, which may result in radiation continuing longer than required and overexposed images that are not of diagnostic quality. | DESIGN: Software Design | Siemens Healthcare sent a Field Safety Notice dated September 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Field Safety Notice to any new users and provide the new owner information to Siemens Healthcare. Customers were asked to send back the attached form confirming receipt of the Field Safety Notice. Customers with questions should call 610-850-4634.For questions regarding this recall call 610-219-6300. |  | 357 | 357 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Open |  |  | 357 | Computer
Z-0140-2014 | 66314 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro system (Luminos dRF and Uroskop Omnia with software version VD10A/GImage intensified Fluoroscopic X-ray system | Siemens RAD Flouro Uro | 2 | November 27, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in radiation continuing longer than required resulting in an overexposed image that is not of diagnostic quality, prompting repetition of the examination. | DESIGN: Software Design | Siemens sent an Field Safety Notice dated September 3, 2013, to all affected customers.The letter identified the product the problem and the action to be taken by the customer. Action to be taken: 1. It is not possible for the user to prevent this malfunction. The exposure time with the use of automatic exposure control is limited to 500ms, which is the maximum duration of radiation in this mode. The exposure can be interrupted by the user at anytime by releasing the exposure button.2. There are two possibilities to solve this issue:a.) Switch off the system if the message "wait for FLC" in status bar does not disappear approximately. 30s after the OGP selection. The system gets ready after switching the system on and reselecting the Torno OGP.b.) Press Shift+F10. The pop up message has to be confirmed and you have to wait approximately.40s until! the "i" icon in the status bar stops blinking. After that a RAD OGP has to be selected and then the Torno OGP. Thereafter, the system gets ready.3. It is not possible for the user to prevent this malfunctionThe Center for Devices and Radiological Health (CDRH) understands that yourCorrective Action Plan (CAP) will involve 59 units and consist of the following elements:1. You have notified affected customers with Field Safety Notice XP032/13/S.2. You have prepared a field modification of the Luminos dRF and Uroskop Omniasystems with software version VD10A/G.3. Distribution of this field modification will began in September 2013 as UpdateInstruction XP015/13/S at no cost to the user.For further Customer Support call (610) 219-6300. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Open |  |  | 59 | Computer
Z-0426-2014 | 65985 | endoscope and/or accessories | KOG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Custom Ultrasonics System 83 Plus 2 and Plus 9 Endoscope Washer/Disinfector. | Custom Ultrasonics System 83 Plus 2 and Plus 9 Endoscope Washer/Disinfector | 2 | November 27, 2013 | 2013 | Custom Ultrasonics, Inc. | Certain System 83 Plus 2 and System 83 Plus 9 units running the Windows 7 operating system were not fully validated prior to distribution. | DESIGN: Software Design | A notification letter, dated February 26, 2013, was sent to direct accounts to inform them of the recall and that technicians would visit facilities to replace the Windows 7 operating system with an appropriately validated Windows XP operating system. |  | 118 | 118 | Worldwide Distribution: USA (nationwide) and Internationally to: Canada, Slovakia Republic, Sydney and Peru. | Open |  |  | 118 | Computer
Z-0430-2014 | 65408 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1.Diagnostic device intended to determine a variety of analytes in human body fluids. | Dimension Vista(R) Intelligent Lab Systems, Dimension Vista(R) 500 and Dimension Vista(R) 1500 Rel | 2 | December 02, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | There is the potential for the vial volume of "in use" Vista calibrators, controls, or sample diluent to incorrectly revert to full volume when removed and reloaded. Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1 | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | A notification letter was sent to consignees in May 2013. |  | 776 | 776 | Worldwide Distribution: USA (nationwide) and iinternationally to Canada. | Open |  |  | 776 | Computer
Z-0431-2014 | 66705 | system, x-ray, stationary | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Easy Diagnost ElevaUniversal Fluoroscopic applications | Easy Diagnost Eleva | 2 | December 02, 2013 | 2013 | Philips Medical Systems, Inc. | The delivered application software was an outdated version to the model of equipment. | DESIGN: Software Design | Philips Medical Systems sent an Urgent - Field Safety Notice dated October 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusromers were informed that the delivered application software was an outdated version to the model of their expipment and that a Philips Field Engineer will be implementing an upgrade of the application software. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of its contents. Customers with questions were instructed to contact their Philips representataive: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 19 | 19 devices | Nationwide Distribution including MN, NH, DE, OR, MA, TX, OH, FL, CA, WA, and SC. | Open |  |  | 19 | Computer
Z-0439-2014 | 66764 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series: Clinac, Trilogy, Trilogy Tx. , Novalis high energy linear accelerators and UNIQUE single energy linear accelerator; Versions 7, 8 and 9. The UNIQUE is not sold in the US.Product Usage: The Varian High Energy Linear Accelerator is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. The UNIQUE is a single energy medical linear accelerator. UNIQUE is indicated for precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. | CSeries: Clinac, Trilogy, Trilogy Tx. , Novalis high energy linear accelerators | 2 | December 03, 2013 | 2013 | Varian Medical Systems, Inc. | This correction is to notify users that a solution to a previous correction has been developed and Varians reps will be contacting locations to schedule installation. | DESIGN: Device Design | Varian sent an Urgent Medical Device Correction letter, P/N CP-05645 Rev D dated October 24, 2013 to their affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated all affected sites will be contacted by Varian representatives to schedule an upgrade to their devices. Questions may be directed to 888-827-4265 for USA and Canada. Europe +41 41 749 8844. |  | 3809 | 3,809 devices | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Uthuania, luxembourg, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, SriLanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Zimbabwe. | Open |  |  | 3809 | Computer
Z-0433-2014 | 65147 | software, transmission and storage, patient data | NSX |  |  | General Hospital | Enforcement Discretion | syngo(R) Lab Data Manager - Version VA11B and VA12A SystemsSecurely consolidates data from multiple lab instruments, the LIS and Siemens StreamLAB(R) Automation Solutions. It provides autoverification, integrated QC capabilities, browser-based connectivity supporting up to 25 concurrent users, remote services to help keep systems running optimally and also has audit trail capabilities. | Siemens syngo(R) Lab Data Manager Version VA11B and VA12A | 2 | December 03, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under certain conditions the system may not perform as intended, causing the release of results to the Laboratory Information System that should have been held for manual review due to auto-verification rules or the delay/omission of result transmission to the LIS. | DESIGN: Software Design | Siemens sent an Urgent Medical Device Correction letter in May 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the letter in its entirety and ensure that all operators understand the information presented. Specific actions needed to be implemented by the customer were outlined in the letter. Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 262 | 262 | Worldwide Distribution: USA (nationwide) and internationally to Canada. | Open |  |  | 262 | Computer
Z-0443-2014 | 66632 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III Data Export Software Version 1.10The Blanketrol III Data Export Software is intended for use in the electronic transfer of medical device data from the Blanketrol III to a computer, without controlling or altering the function or parameters of the Blanketrol III. | Blanketrol III Data Export Software | 2 | December 04, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a customer inquiry regarding a beta version of the Bill Data Export Software, Model USB-127, it was discovered that the wrong software version was released to a customer. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Software Change Control | Cincinnati Sub-Zero sent a Urgent Medical Device Notification letter dated July 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They instructed their customers to to do the following:1) Upon receipt of the corrected Blanketrol III Data Export Software disc, immediately collect and destroy all previous Blanketrol III Data Export Software disc from this lot # currently in possession and replace all downloaded Blanketrol III Data Export Software with the correct version provided by CSZ.2) After the Blanketrol Data Export Software Version 1.00 have been destroyed and downloaded versions replaced with version 1.10, complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that you have performed and completed the requested actions. Return the form by fax to (513) 772-9119 scan and e-mail the information to BIIIExportSoftware@cszinc.com, or mail to:Cincinnati Sub-Zero12011 Mosteller RoadCincinnati, OH 45241 U.S.A.For further questions please call 1-(800) 989-7373 or (513) 772-8810. |  | 12 | 12 Units | Worldwide Distribution - US Distribution including the following states AK, CO, FL, MA, MI, and NC. and Internationally to Canada, Germany and United Kingdom. | Open |  |  | 12 | Computer
Z-0454-2014 | 64762 | system, digital image communications, radiological | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | The GE Centricity Web; Software Version 3.0x, is intended for use under the direct supervision of a licensed healthcare practitioner. The Centricity Enterprise Web is an image and information distribution system for the clinical review of medical images and reports. The Centricity Enterprise Web is not intended for primary diagnosis. The Centricity Enterprise Web is available as a supplemental sub-system to Centricity PACS or as a stand alone Web-based image and information distribution system. The Centricity Enterprise Web provides both the server software and a client application, which utilizes off the shelf browser technology. Centricity Enterprise Web supports several interface protocols which allows the system to be integrated with any image or information systems such as a hospital information system (HIS), radiology information system (RIS), or electronic medical record (EMR), which supports one of the supported integration protocols. | GE Centricity Web | 2 | December 05, 2013 | 2013 | Ge Healthcare It | GE Healthcare has become aware of an issue associated with rejected images using the WADO protocol within GE Centricity Web that may impact patient safety. The WADO (Web Access to DICOM Objects) interface allows for embedding DICOM images as JPEGs on a web page and is available as a licensed operation on Centricity Web. When rejected images are archived by Centricity PACS to Enterprise Archive, an | DESIGN: Software Design | Urgent Medical Device Correction letters dated April 19, 2013 were sent to customers. The letter requests that customers ensure that their system is configured to disable the WADO interface using the included instructions until the software revision can be installed at their site. The letter further informs customers that updated software will be sent to them for installation.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI# 85206R. |  | 53 | 53 units | US Distribution incluidng states of: CA, FL, GA, IA, IL, LA, MO, NJ, NY, OK, TN, WA, and WI. | Open |  |  | 53 | Computer
Z-0456-2014 | 66808 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iba COMPASS SW Version 3.0, Catalog Number CS10-100 (medical linear accelerator)radiological | iba COMPASS SW Version 3.0 | 2 | December 06, 2013 | 2013 | Iba Dosimetry Gmbh | The development partner of IBA informed that there is problem in a SW component used in COMPASS when converting Region of Interest (ROI) geometries to dose grid representation. Under certain conditions this problem generates too large ROI voxel volumes causing erroneous dose statistics and DVHs. | DESIGN: Software Design | Iba sent an Field Safety Notice on August 20, 2013, by email and letter delivered by UPS, Please distribute this notice to all those who need to be aware within your area of responsibility where the affected software have been installed (e.g. colleagues, satellites).Please maintain awareness of this notice and resulting action until we have provided you with a new software version.We would kindly ask you to acknowledge the receipt of this Field Safety Notice. Please sign and send back the attached Confirmation of Receipt within four weeks after the receipt of this Field Safety Notice.We deeply apologize for the inconvenience. We are working on a new software version with corrected functionality. This should become available during next fall and we will ship it to you as soon as it is released, together with the installation and verification instructions.Should you have questions or require additional information, please contact the Customer Service Team at: Service Department IBA Dosimetry GmbH Bahnhofstrasse 5DE-90592 Schwarzenbruck, GermanyTel.:+49 9128 607 38Fax: + 49 9128 607 26Email: service@iba-group.com. |  | 33 | 33 units | Worldwide Distribution - US Distribution including the states of CT, MA, OH, CA, SC, NY and OK, and the countries of International distribution (from Germany) Poland, Germany, GB, The Netherlands, Spain, Turkey, Canada, Mexico, China, Hong Kong and Singapore. | Terminated | February 24, 2014 | 80 | 33 | Computer
Z-0469-2014 | 66596 | research use only/pathology | OTU |  |  | Pathology | Enforcement Discretion | Digital Pathology System (DPS) Software version 1.1The Omnyx" Integrated Digital Pathology system, which consists of a whole slide scanner and associated software applications, is a scanning, receiving, storing, managing, annotating and computer monitor viewing system for digital whole slide images. Omnyx" products are for in vitro diagnostic use for specific clinical applications, and are intended for research use only on other applications. Specific clinical use applications are indicated in separate clinical specific user guides supplied by Omnyx. | Digital Pathology System (DPS) Software, version 1.1 | 2 | December 09, 2013 | 2013 | Omnyx Llc | There is an issue associated with Digital Pathology System (DPS) Software, in which the annotation tools within the DPS version 1.1 Image Viewer will display incorrect measurements and area calculations for a region of interest on a scanned slide, under certain conditions. | DESIGN: Software Design | Omnyx sent a Urgent Field Safety Notice letter dated October 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Please discontinue use of the measurement and area calculation features of the annotation tools (line, rectangle/square, ellipse/circle, and freeform) in the DPS Image Viewer at this time. If any of these tools have been used for measurements or area calculations on clinical cases, verify measurements and calculations on glass slides.Please confirm receipt of these instructions by completing the attached form and scanning/emailing it to FieldAction@ omnyx.com or faxing it to 001-412-894-2137Product Correction: An updated version of the DPS SW will be released and installed on your system as soon as it becomes available.Please note that the appropriate regulatory authorities have been informed of this field action.We apologize for any inconvenience that this may have caused you and are working diligently to address this issue.If you have any additional questions or concerns, please contact you GE Healthcare account representative. Further questions please call (412) 891-2100. |  | 2 | 2 | International Distribution to Canada and Switzerland only. | Open |  |  | 2 | Computer
Z-0489-2014 | 66970 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | HGE Healthcare Patient Monitor B30, Procare Monitor B20, Procare Monitor B40, and Monitor B40K122253 (Procare Monitor B20): The PROCARE Monitor B20 is a portable multiparameter unit to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adult, pediatric, and neonatal patients in a hospital environment and during intra-hospital transport. The PROCARE Monitor B20 is intended for use under the direct supervision of a licensed health care practitioner. The PROCARE Monitor B20 is not intended for use during MRI. The PROCARE Monitor B20 monitors and displays oscillometric non-invasive blood pressure (systolic, diastolic and mean arterial pressure), invasive blood pressure, end-tidal carbon dioxide, heart/pulse rate, respiration rate, EGG (including arrhythmia and ST segment analysis), temperature with a reusable or disposable electronic thermometer for continual monitoring. Esosophageal/Nasopharyngeal/Tymnpanic/Rectal/Bladder/Axillary/Skin/Airway/Room/Myo cardial/Core/Surface temperature, and functional oxygen saturation (Sp0 2) and pulse rate via continuous monitoring , including monitoring during conditions of clinical patient motion or low perfusion. K120598 (Procare Monitor B40): The PROCARE Monitor 840 is a portable multiparameter unit to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adult, pediatric, and neonatal patients in a hospital environment and during intra-hospital transport. The PROCARE Monitor 840 is intended for use under the direct supervision of a licensed health care practitioner. The PROCARE Monitor 840 is not intended for use during MRI. The PROCARE Monitor B40 monitors and displays oscillometric non-invasive blood pressure (systolic, diastolic and mean arterial pressure), invasive blood pressure, end-tidal carbon dioxide, heart/pulse rate, respiration rate, ECG (including arrhythmia and ST segment analysis), temperature with a reusable or disposable electronic thermometer for continual monitoring. Esophageal/Nasopharyngeal/Tympaiic/Rectal/Bladder/Axillary/Skin/Airway/Room/Myocardial/Core/ Surface temperature, and functional oxygen saturation (SpO2) and pulse rate via continuous monitoring, including monitoring during conditions of clinical patient motion or low perfusion. K130584 (Monitor B40): The Monitor B40 is a portable multi-parameter unit to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adult, pediatric, and neonatal patients in a hospital environment and during intra-hospital transport. The Monitor B40 is intended for use under the direct supervision of a licensed health care practitioner. The Monitor B40 is not intended for use during MRI. The Monitor B40 can be a stand-alone monitor or interfaced to other devices via a network. The Monitor B40 monitors and displays: ECG (including ST segment, arrhythmia detection), invasive blood pressure, heart/pulse rate, oscillometric non-invasive blood pressure (systolic, diastolic and mean arterial pressure), functional oxygen saturation (SpO2) and pulse rate via continuous monitoring (including monitoring during conditions of clinical patient motion or low perfusion), temperature with a reusable or disposable electronic thermometer for continual monitoring. Esophageal/Nasopharyngeal/Tympanic/Rectal/Bladder/Axillary/Skin/Airway/Room/Myocardial/Core/ Surface temperature, impedance respiration, respiration rate, airway gases (CO2, O2, N2O, anesthetic agents, anesthetic agent identification and respiratory rate) and Entropy. Product Quantity Distributed (Int) 16,061 (343 US, 15,718 OUS) | HGE Healthcare Patient Monitor B30, Procare Monitor B20, Procare Monitor B40, and Monitor B40 | 2 | December 12, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to ECG filter settings associated with the B20/B40 V1, B30 and B40 V2 Patient Monitors.The B20, B30 and B40 Patient Monitors set the ECG filter setting incorrectly if the monitor starts with impedance respiration set in the ON position. Instead of the user selected filter, an additional 0.5 Hz high pass filter is used. The | DESIGN: Software Design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 10, 2013, to all affected consignees. The letter described the safety Issue, Safety Instructions, Affected Product, Details, Product Correction and Contact Information (Technical Support at 1-800-558-7044).Consignees were asked to retain a copy of the instructions with the user manual.For questions regarding this recall call 262-513-4122. |  | 16061 | 16,061 (343 US, 15,718 OUS) | Worldwide Distribution - USA (nationwide) to CA, CT, FL, GA, MO, NC, NJ, PA, TX, WI and Internationally to:FRANCE, PORTUGAL, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA 7 HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, NEW ZEALAND, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS,. CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, GABON, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARIA, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, MYAMMARA, NAMBIA, NAPAL, NETHERLANDS, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SENEGAL SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SIR LANKA, SWEDEN, SWITZERLAND, THAILAND, THE NETHERLANDS, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, URUGUAY, UZEBEKISTAN, VENEZUELA, and VIETNAM. | Open |  |  | 16061 | Computer
Z-0519-2014 | 66521 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Reaction Vessels.The Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access Immunoassay Systems Reaction Vessels | 2 | December 16, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for the Access Immunoassay Systems Reaction Vessels, including the Acess/Access 2 Immunoassay Systems, SYNCRON LXi 725 Clinical System, and UniCel DxC 600i SYNCHRON Access Clinical System because Beckman Coulter has observed an intermittent increase in Access System "Wash Carousel Motion Errors" with specific lots of Access SystemReaction Vessels (RVs). I |  | Beckman Coulter sent an Urgent Product Correction (PCA) letter dated October 1, 2013, along with a customer response form to all affected customers on October 2, 2013. The customer letter informed customers of the recall with the issue, impact, action (If your Access Immunoassay System experiences Wash Carousel Motion Errors that cannot be corrected with normal troubleshooting measures as outlined in the Access 2 Reference Manual section 6.3, contact Beckman Coulter Customer Technical Support.If you do not have Access System RVs with the lot numbers listed above, or if your Access Immunoassay System is not experiencing Wash Carousel Motion Errors, no additional action is required.), resolution, and contact information, website: http://www.beckmancoulter.com. Customers in the United States and Canada with questions were instructed to call 1-800-854-3633, outside the United States and Canada, customers were instructed to contact their local Beckman Coulter representataive.For questions regarding this recall call 714-961-4941. |  | 16154 | 16,154 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Japan, India, Hong Kong, Guatemala, China, Brazil, Brunei Darussalam, Costa Rica, El Salvador, JordanMalaysia, Mexico, Panama, Philippines, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Venezuela, and Viet Nam. | Open |  |  | 16154 | Computer
Z-0539-2014 | 66989 | medical computers and software | LNX |  |  | General Hospital | 510(k) | ABACUS TPN Calculation Software, Baxa: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition)ABACUS is a WindowsÂŽ - based order entry TPN software application for comprehensive calculations and label printing. The primary use of Abacus is for TPN order calculations. Other uses of Abacus include: non-TPN volumes administered to patients on TPN therapy in order calculations; CVVH calculations; Cardioplegia (Buckburg) calculations; Epidural calculations; Dilutions from concentrated ingredient; Combinations of multiple ingredients; Fractional doses from larger volumes | ABACUS TPN Calculation Software, Baxa | 2 | December 18, 2013 | 2013 | Baxter Corporation Englewood | Baxter Healthcare Corporation is issuing a voluntary correction for ABACUS Total parenteral Nutrition (TPN) Calculation Software versions 2.0 through 3.0 due to the possibility that a patient could receive another patient's TPN order as a result of data corruption. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Baxter sent an Urgent Device Correction letter dated December 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the patient order against the final bag label as part of the TPN acceptance procedure to ensure that the patient's TPN order is correct. Customers will be contacted by Baxter to arrange for installation of a software upgrade specific to their compounder's configuation. Customers were asked to acknowledge the receipt of this notification by completing the attached Customer Reply Form and return it to Baxter by fax to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. If the product was distributed to other facilities customers were instructed to notify end-user customers in accordance with their customary procedures. Customers with questions were instructed to call 1-800-678-2292.For questions regarding this recall call 800-422-9837. |  | 2211 | 2,211 | Nationwide Distribution | Open |  |  | 2211 | Computer
Z-0536-2014 | 66963 | clinical sample concentrator | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure 96 SystemVersion 1.0Software Version 2.0Robotic workstation for automated isolation of nucleic acids from different kinds of crude sample material (whole blood, serum, blood cells, culture cells, tissue, bacteria, fungi) by the use of the specially designed MagNA Pure reagent kits, automated filling of different kinds of PCR reaction vessels. | MagNA Pure 96 System | 2 | December 18, 2013 | 2013 | Roche Diagnostics Operations, Inc. | This issue is only relevant for customers importing xml files into the MagNA Pure 96 IVD instrument software for sample identification. There is a potential for sample mismatch when using the MagNA Pure 96 IVD instrument (catalog number 06541089001). Importing an order file in xml format to the MagNA Pure 96 IVD instrument software may result in a wrong sequence of samples in the graphical overvi |  | Roche Diagnostics sent an URGENT Medical Device Correction Notice dated November 25, 2013, to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Customers were advised to discontinue use of the Pipetting Report and to use the sample list file instead. The MagNA Pure 96 Instrument Operators Guide is being updated with an addendum addressing this issue. The letter included a Customer Response Form for customers to complete and return to the firm by FAX : 1-888-670-4135. Please contact the Roche Support Network Customer Support Center at 1-800-526-1247, 24 hours a day, seven days a week, if you have questions about the information contained in this UMDC. |  | 27 | 27 | US Distribution including the states of AL, AZ, CA, CT, GA, NY, NC, NJ, OH, TX, VA and WA. | Open |  |  | 27 | Computer
Z-0534-2014 | 66865 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8000 and 8015 with Alaris PCA module model 8120 and Auto ID barcode scanner model 8600.Product Usage: The Alaris PC unit is the central programming, monitoring and power supply component for the Alaris System. The Alaris PCA (Patient Controlled Analgesia) pump is one of the modules offer under the Alaris System. The Auto ID is a barcode scanner. | Alaris PC unit model 8000 and 8015 | 1 | December 20, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit (model 8000 and 8015) because the Alaris PC unit will produce an audible alarm and the Alaris PC unit will display a System Error message. | DESIGN: Software Design | CareFusion sent an Urgent Medical Device Recall Notification letter dated December 5, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to contact CareFusion Customer Advocacy at (888) 812-3266, if they experienced a SYSTEM ERROR on the Alaris PC unit. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to complete and return the enclosed Customer Response Card to expedite the corrective action process. Customers with questions were instructed to contact CareFusion Support Center at 888-562-6018, or for Technical Support call 888-812-3229. |  | 4824 | 4824 total units (4242 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, and Australia. | Open |  |  | 4824 | Computer
Z-0552-2014 | 66956 | system,planning,radiation therapy treatment | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Plan Parameters workspace withinEclipse Treatment Planning System Version 11 and 13;ARIA Radiation Oncology versions 11 and 13Acuity, versions 11 and 13.Manufactured by Varian Medical SystemsPalo Alto, CA.Used in radiation therapy simulation. | Eclipse Treatment Planning System | 2 | December 20, 2013 | 2013 | Varian Medical Systems, Inc. | An anomaly has been traced to a software deficiency within the Plan Parameters workspace. When a dose relevant parameter is changed and "do no clear MUs and Reference Point dose" is selected, the Monitor units may change. | DESIGN: Software Design | Varian Medical Systems will provide all affected consignees with an Urgent Medical Device Correction Letter commencing on November 18, 2013 by federal express. . The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the reported issue. |  | 6701 | 6701 | Worldwide Distribution. | Open |  |  | 6701 | Computer
Z-0556-2014 | 67014 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology 12.0, Picture Archiving Communication System (PACSA). | McKesson Radiology | 2 | December 20, 2013 | 2013 | Mckesson Information Solutions LLC | When merging two patient records into one patient record using McKesson Radiology 12.0, the resulting patient record is missing the Contrast Allergy information for the second patient record. | OTHER/UNDETERMINED: Pending | Consignees were notified via letter (Field Safety Notice) on 10/15/2013. |  | 6 | 6 units | MO, NY, TN, WV and France. | Open |  |  | 6 | Computer
Z-0571-2014 | 66754 | calculator/data processing module, for clinical use | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager driver : Driver sysu100i that connects to the Sysmex UF-100 analyzerSoftware all versions prior to( v7.00.0006 and v8.00.0005): v7.00.0001, v7.00.0002, v7.00.0003, v7.00.0004, v7.00.0005, v8 .00.0001, v8.00.0002, v8.00.0003, and v8 .00.0004 | Instrument Manager | 2 | December 24, 2013 | 2013 | Data Innovations, Inc. | When connected to the Sysmex UF-100, theInstrument Manager driver sysu100i driver software versions 7.00.0001, 7.00.0002, 7.00.0003, 7.00.0004, 7.00.0005, 8.00.0001, 8.00.0002, 8.00.0003,and 8.00.0004, and is configured with option "Strip Zeros from Result" set to "Yes" may result in sample error | DESIGN: Software Design | Data Innovation notified consignees by Important Medical Device Recall Notification dated 11/5/13 to all list serv recipients of the problem and posted critical alert on Customer Web Portal. A certified letter was issued to impacted customers and consignees if no response was recieved from e-mail. |  | 24 | 24 systems | NationwideCanada and Britain | Open |  |  | 24 | Computer
Z-0570-2014 | 67046 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx Monitor/Defibrillator with CO2/EtCO2 Measurement The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician. The HeartStart MRx EtCO2 Option monitors carbon dioxide (CO2) and measures end-tidal carbon dioxide (EtCO2). The option provides an EtCO2 value and a CO2 waveform on the devices display. | Philips Healthcare | 2 | December 24, 2013 | 2013 | Philips Medical Systems, Inc. | MRx defibrillator displays a -?- for EtCO2 and does not display EtCO2 values when patient CO2 level falls below 7.6 mmHg | DESIGN: Software Design | Philips Healthcare issued an Customer Information letter dated February 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer with the addendum to the Instructions for Use.If you have a ? in the parameter block and no CO2 waveform on the display, the waveform source is invalid. If invalid, check patient, confirm airway status and examine the FilterLine to make sure it is connected to the device and not kinked or pinched. For more information on the EtCO2 Option seeMonitoring Carbon Dioxide chapter in the HeartStart MRx Instructions for Use.Should you have any questions or concerns, please contact your local Philips representative at 1-800-722-9377. |  | 18529 | 18,529 units | US Nationwide and OUS. | Open |  |  | 18529 | Computer
Z-0580-2014 | 66980 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | steroEOS Workstation with software versions: 1.3, 1.4 and 1.5 used in processing 3D X-ray imagesProduct Usage: The sterEOS Workstation is a general system for acceptance , transfer, display and digital processing of two 3D x-ray images. | EOS Imaging | 2 | December 26, 2013 | 2013 | EOS Imaging | Error copying information to the clinical file and to the report when two 3D models are opened simultaneously on the sterEOS workstation | DESIGN: Software Design | EOS imaging issued a safety information letter dated November 29, 2013, to the US customer sites. The letter identifed the affected product, problem and actions to be taken. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. The letter informed customers that EOS imaging will, at it's own expense, corect all systems affected by means of a software update. An EOS imaging representative will contact affected customers to make an appointment for this update once it is available. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. For questions call +33 (0) 155 25 6060. |  | 36 | 36 units | USA Nationwide Distribution | Open |  |  | 36 | Computer
Z-0602-2014 | 67094 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Advanced Monitor with Waveform Display Model CAM01 Integra Neurosciences, 311 Enterprise Dr Plainsboro, NJ 08536 USAThe Camino Advanced Monitor is a compact, portable device for use with Camino 110-4 series of Pressure / Temperature and Pressure Transducer-Tipped Catheters. Pressure and/or Temperature are measured at the Catheter tip, eliminating the need for external transducers, fluid, pressure tubing, and flush devices. The Camino displays intracranial pressure (ICP), intracranial temperature (ICT), and calculates cerebral perfusion pressure (CPP) (CPP = mean arterial pressure - mean intracranial pressure). The Camino provides a continuous display of the ICP waveform, as well as mean ICP, CPP, ICT, or systolic and diastolic waveform values. A continuous record of mean pressure and CPP values over the most recent 24-hour period is stored in memory, and can be displayed on command as a TREND either as the most recent 12 or 24 hour period. Outputs for ICP and ICT are available for use with patient bedside monitors. An isolated analogue output provides a continuous ICP waveform for hard copy documentation or data acquisition. An input to receive mean arterial pressure from a compatible patient bedside monitor is available for use when it is desired to have a CPP reading. The mean arterial pressure cannot be acquired from a non-invasive blood pressure cuff. Although the Camino is intended to be a stand-alone system, it also conveniently connects to hospital bedside monitoring systems. | Integra | 2 | December 27, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through a complaint from a single non-US customer that if the optional SYS/DIAS mode is used on CAM01 monitor that contains Mod 4 Firmware, the displayed values for the systolic and diastolic intracranial pressure (ICP) are reversed. But, if the CAM01 is used with an auxiliary bedside monitor, the bedside monitor displays these values correctly. | DESIGN: Device Design | Integra sent an Urgent Medical Device Correction letter and Correction Acknowledgement form dated November 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you received one or more of the affected serial numbers that will require this firmware update. Integra asks that you do the following:1. Review your inventory and verify that you have Camino Advanced Monitors with a serialnumber formatted as MExxxxxx04. If so, you can continue to use the CAM01monitor(s), being aware that if you choose to use the optional SYS/DIAS mode, thevalues on the CAM01 display will be reversed.2. Complete the attached form. If you do not have the product with the affected serialnumber in service at your facility, check the box:I do not have any product with a serial number of the form MExxxxxx04.3. If you do have the affected serial number, check the box:I do have the product with a serial number of the form MExxxxxx04, and pleaseindicate the serial numbers you have.4. Complete the other information and return it by email or fax as indicated on the form.Complete the form and return it by email or fax as indicated on the form. Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 276 | 275 CAM01 Monitors | Worldwide Distribution - USA ( Nationwide ) and the countries of : UK, Saudi Arabia, South Africa, Brunei, China, France, Poland, Turkey, Greece, Sweden, United Arab Emirates, Russia, Spain, Pakistan, Finland, Israel, Italy, Slovenia, Japan, Poland, Malaysia, Germany and Panama. | Open |  |  | 276 | Computer
Z-0613-2014 | 67035 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx Monitor/DefibrillatorModels: M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips HeartStart MRx Monitor/Defibrillator | 2 | December 31, 2013 | 2013 | Philips Medical Systems, Inc. | Device Operating on Battery Power May Shutdown without Warning if exposed to elevated levels of electromagnetic interference from other radio frequency (RF) energy sources and | PRODUCTION CONTROLS: Environmental Control | Philips issued a Medical Device Correction letter April 2013 to all affected customers along with an IFU Addendum titled, HeartStart MRx Electromagnetic Compatibility. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the condition could be avoided by operating the device on external AC or external DC power. Customers were instructed to follow the device instructions when using battery power to minimize or eliminate potential interference from other RF sources. Customers with questions were instructed to contact their local Philips representative.For questions regarding this recall call 978-687-1501. |  | 64972 | 64,972 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries:AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMONGOLIAMOROCCONAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Open |  |  | 64972 | Computer
Z-0516-2013 | 61870 | upc reader (bar code reader) | RFZ |  |  |  |  | Motorola Solutions PCS3000/CS3070 LaserUtility/Peripheral Laser Products | Motorola Solutions Laser | 2 | January 30, 2013 | 2013 | Motorola Inc | It was discovered the CS3000/CS3070 was manufactured using the laser power settings for a Class 2 laser device instead of the Class 1 power levels as specified on the product label. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Motorola Solutions, Inc. agrees to provide instructions for downloading and installing a software update from the internet, which will reduce the laser power to IEC Class 1 AELs. Purchasers will be provided an internet site and instructions for notifying Motorola Solutions, Inc. that the software installation and power reduction were successful, thereby completing the corrective action.corrections to all units is anticipated to be completed by May 2013.I. Notifications to all dealers and purchasers must be made, and the FDA NY District Office Recall Coordinator is to be included in the notifications. 2.Correction of all units to be completed by May 2013.3.All products will be brought into compliance with the laser performance standard at no cost to the purchaser. For questions regarding this recall call 631-738-3474. |  | 16961 | 16,961 units domestically | Nationwide distribution | Open |  |  | 16961 | Computer
Z-0991-2013 | 64394 | susceptibility test cards, antimicrobial | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 Compact computer system when equipped with the following Hewlett-Packard PC models:(1) dc7100 8-port, Cat. #W0441; (2) dc7100 8-port (refurbished), Cat. #W0441R; (3) dc7700 8-port, Cat. #W0447; (4) dc7700 8-port (refurbished), Cat. #W0447R; (5) dc7800, Cat. #W0449; (6) dc7800 (refurbished), Cat. #W0449R(7) rp5700, Cat. #W0452;(8) rp5700 (refurbished), Cat. #W0452R;(9) rp5800 XPE, Cat. 413642;(10) rp5800WES7, Cat. #413862;(11) d530 6-port, Cat. #W0438;(12) d530 6-port (refurbished), Cat. #W0438R; and(13) d530 4-port, Cat. #W0437.VITEK 2 is an automated system designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | VITEK 2 Compact computer system | 2 | March 25, 2013 | 2013 | Biomerieux Inc | There is a potential to link an isolate result to the wrong patient and then upload the results to the LIS. | DESIGN: Software Design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" dated February 20, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Emails were issued by the recalling firm on 2/18/2013 and 2/19/2013 to their subsidiaries and to distributors. Letters began issuing to customers via e-mail on 2/20/2013. Contact bioMerieux Clinical Customer Service at (800) 682-2666 for questions regarding this notice. |  | 3 | For all 3 computers types: 16,736 originally distributed; 554 refurbished and redistributed | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaidjian, Bahrain, Belarus, Belgium, Benin, Bolivia, Bosnia, Brazil, Bulgaria, Burkina Faso, Camaroun, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Democratic Republic of Congo, Demark, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Mongolia, Morocco, Myanmar, Netherlands, Netherlands, Nicaragua, Niger Niamey, Nigeria, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Open |  |  | 3 | Computer
Z-1030-2013 | 64665 | electrode, cortical | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Integra Ojemann Cortical Stimulator; Product / Catalogue No: OCS2.Intended for intraoperative cortical stimulation mapping procedures. | Integra Life Sciences Ojemann Cortical Stimulator | 2 | April 01, 2013 | 2013 | Integra Burlington MA, Inc. | If the headphone jack is in use during a surgical procedure with the OCS2 and a non-intended voltage (such as static electricity) comes in contact with the outer case of the unit, that voltage could be transmitted to the patient and could be a potential source of injury. | DESIGN: Device Design | Integra initiated a voluntary recall on March 18, 2013 by providing written notification to consignees either by traceable courier service or traceable emails regarding the correction that Integra records indicate have an OCS2 since it was introduced to the market in 2007. The consignee notification:Advised of the nature of the issue and to immediately stop using the headphone jack while using the OCS2. Requested identification by S/N any OCS2's they have, and complete / return the Acknowledgment Return Form.Advised that Integra will contact the consignee to schedule a time / date to correct the OCS2's they have identified. Questions, contact service hotline at 1-888-772-7378. |  | 371 | 371 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Belgium, Denmark, Spain, Great Britain, Hungary, Ireland, Italy, Netherlands, Poland, Morocco, Netherlands, Poland, Portugal, Saudia Arabia, Sweden, Turkey, and South Africa. | Terminated | February 25, 2014 | 330 | 371 | Computer
Z-1082-2013 | 64427 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK CR Plus and LIFEPAK EXPRESS defibrillators. Product part's numbers: 3200731-xxx, 3202177-xxx, U3200731-xxx, and U3202177-006.Automated external defibrillators. | The LIFEPAK CR Plus and LIFEPAK EXPRESS Automated External Defibrillators (AEDs) | 2 | April 10, 2013 | 2013 | Physio Control, Inc. | LIFEPAK CR Plus or LIFEPAK Express Automated External Defibrillators (AEDs) are recalled due to the potential to experience an early internal battery depletion issue. | DESIGN: Component Design/Selection | The firm start sending the "URGENT MEDICAL DEVICE CORRECTION LIFEPAK CR Plus AND LIFEPAK EXPRESS AEDs" letter, dated February 2013, to their consignees.Consignees are advised to check the readiness of the AEDs and should contact Physio if the OK symbol is NOT visible and either the ATTENTION or WRENCH symbols are present.In the event a device is returned to Physio-Control for assessment, the firm will analyze it to determine if it is affected by the early internal battery depletion issue. If it is verified the device is affected, the consignee will be provided with a permanent replacement device.Customers with questions can visit the firm's website www.physio-control-notices.com/CHARGE-PAK or call Technical Support at 1-800-442-1142. |  | 824 | 54,197 units in US and 78,627 units internationally (total 132, 824 units) | Worldwide Distribution. | Open |  |  | 132824 | Computer
Z-1150-2013 | 64781 | accessories, cleaning, for endoscope | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessor System, P/N 50004.Designed to automatically clean and high-level disinfect flexible, submersible video or fiber-optic endoscopes. | EVOTECH | 2 | April 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because they have received reports of smoke emanating from the system due to a circuit board malfunction that may occur in the Channel Control Manifold sub-assembly. |  | An Urgent Medical Device Field Correction letter dated 3/28/13 was sent to all customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR). The letter informed the customers of the problems identified and the actions to be taken. Customers are informed that the failure can potentially cause the following outcomes for the cycles listed below:1. In the Full Reprocessing Cycle (Cleaning and HLD): there is a potential for residual rinse water and/or alcohol to remain in the endoscope channels after the cycle completes. 2. In the HLD-Only Cycle: there is potential for an adequate amount of high-level disinfectant solution to be sent through some endoscope channels.If customers observe smoke, then they are instructed to:a. Power OFF the device.b. Report your incident to ASP Professional Services at (888) 783-7723 and select option #2.c. Do not use the device until it has been serviced by ASP.Customers are instructed to return attached buyer return card acknowledging their receipt of this communication. Customers are instructed to contact Stericycle at (866) 761-9566 for any questions related to the field correction. | Contact the recalling firm for information | 761 | 761 units | Nationwide Distribution | Terminated | August 22, 2013 | 126 | 761 | Computer
Z-1170-2013 | 64837 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients.The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. | GemStar Infusion System | 1 | April 30, 2013 | 2013 | Hospira Inc. | Pump shutting off during use without warning. | DESIGN: Device Design | Hospira sent an Important Device Information letter dated March 18, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the affected product for signs of leakage, corrosion or other damage prior to each use. In addition, each time the batteries are replaced the battery compartment should be inspected for damage. If a device exhibits damage caused by leakage, it should immediately be removed from clinical service and Hospira should be contacted at 1-800-441-4100 to arrange for return of the device for repair. If the product has been further distributed they should notify their consignees and confirm they have done so by returning the reply form to Stericycle.For questions regarding this recall call 224-212-4892. | Contact the recalling firm for information | 286 | List Number 13086: 286 pumps; List Number 13087: 45,376 pumps; List Number 13088: 23,942 pumps | Worldwide Distribution - USA (nationwide) and internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open |  |  | 471 | Computer
Z-1245-2013 | 64876 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle InsuLinx Blood Glucose Monitoring System:For In Vitro Diagnostic Use only. Use only with FreeStyle InsuLinx test strips.Abbott Diabetes Care, Inc.1360 South Loop RoadAlameda, CA 94502 USAThe FreeStyle InsuLinx Blood Glucose Test Strips are for use with the FreeStyle InsuLinx Blood Glucose Meter to quantitatively measure glucose in capillary whole blood samples drawn from the fingertip. | FreeStyle InsuLinx Blood Glucose Monitoring System | 1 | May 09, 2013 | 2013 | Abbott Diabetes Care, Inc. | AT rare, extremely high glucose levels (1024 mg/dL and above), the FreeStyle InsuLinx Blood Glucose Meter may provide an inaccurate reading. | DESIGN: Software Design | Abbott Diabetes Care sent an Urgent Product Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take one of the following options immediately to address the issue. 1) Customers can access a software update to resolve the issue at www.freestyleinsulinx.com/swupdate.2) Customers can contact Abbott Diabetes Care Customers Service at 1-866-723-2697 to expedite return and replacement of the affected product at no charge. For questions customers were instructed to call Abbott Diabetes Care Customers Service at 1-866-723-2697.For questions regarding this recall call 510-239-2775. |  | 544609 | 335,273 meters; 209,336 remain at firm. | Nationwide Distribution | Open |  |  | 544609 | Computer
Z-1431-2013 | 65194 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | 8mm Precise Bipolar Forceps used in conjunction with the da Vinci Surgical System, Model IS1200.Multiple use electrosurgical endoscopic instruments. | 8mm Precise Bipolar Forceps | 2 | May 29, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of dislodging of the Cautery Plug Insert from the Cautery Plug Riser, preventing a bipolar cautery cable from being properly connected which results in loss of cautery function of Bipolar Instruments. | DESIGN: Device Design | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated May 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 697 | 697 units - all products | Worldwide Distribution-USA (nationwide) including the states of AL, CA, CO, FL, IL, KY, MI, MS, NJ, NY, PA, TX, UT, VA, WA, and WI, and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Norway, Qatar, Romania, Spain, Sweden, Switzerland, Taiwan, United Kingdom, Saudi Arabia, Singapore, and Venezuela. | Terminated | February 18, 2014 | 265 | 697 | Computer
Z-1670-2013 | 65292 | unit, phacofragmentation | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | SovereignÂŽ Compact Phacoemulsification System, v 5.3, Part Number: CMP680300. The system is used for cataract surgery. | Sovereign Compact Phacoemulsification System | 2 | July 09, 2013 | 2013 | Abbott Medical Optics, Inc. | Abbott Medical Optics Inc. (AMO) initiated this Urgent Notice of Medical Device Correction and Advisory Notice to all customers who use the AMO SovereignÂż Compact Phacoemulsification System, Version 5.3, because it has the potential to lose the vacuum tone mode during surgery. | DESIGN: Software Design | AMO issued an "Urgent Notice of Medical Device Correction and Advisory Notice" on May 23, 2013 via email, to all customers with a SovereignÂŽ Compact Phacoemulsification System v 5.3 to inform them of the issue and provide them with actions to take in the event of sound loss during use of the SovereignÂŽ Compact Phacoemulsification System, v 5.3. |  | 534 | 534 systems | Worldwide Distribution, including Nationwide (US) and countries in Latin America, EMEA, and APAC. | Open |  |  | 534 | Computer
Z-1686-2013 | 65590 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | da Vinci Si Surgeon Side Cart, model number SS3000;Manufactured by Intuitive Surgical, Sunnyvale, CA 94086.Intuitive Surgical endoscopic Instrument Control System is intended to assist in the accurate control of Intuitive Surgical Endoscopic instruments. It is intended for adult and pediatric use. | da Vinci Si Surgeon Side Cart | 2 | July 11, 2013 | 2013 | Intuitive Surgical, Inc. | Factory testing on da Vinci Si Surgical Systems may not be in compliance with UL Standards as one of the testing devices was found to be working incorrectly. | DESIGN: Software Design | Intuitive Surgical sent an Urgent Device Correction notice dated July 15, 2013, to all affected customers by Federal Express.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all affected personnel were fully informed of the notification and that the notification should be forwarded to other managers within their facility. Customers were also instructed to complete and return the attached Acknowledgment Form to acknowledge receipt of the notification. Customers with questions were instructed to call Customer Service at 800-876-1310, Option 3.For questions regarding this recall call 408-523-2244. |  | 30 | 30 total, all devices | Worldwide Distribution - USA including UT, CA, AZ, CO, FL, HI. MT, NY,TX and WA and one customer in Russia. | Terminated | January 22, 2014 | 195 | 30 | Computer
Z-2168-2013 | 65960 | table, operating-room, electrical | GDC | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | OPERON D850 Surgical TableProduct Usage: Surgical Table | OPERON D850 Surgical Table | 2 | September 09, 2013 | 2013 | Berchtold Corp. | Table was not lowering. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Berchtold sent a Customer Notification letter dated April 16, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that BERCHTOLD will upgrade the OPERON Surgical tables with a new firmware version that will eliminate the potential for experiencing problems. For questions call 843-569-6100 ext. 322. |  | 263 | 263 | USA Nationwide Distribution in the states of AL, AZ, CA, CO, CT, FL, GA, HI.IL, IA, KY, LA, MD, MA, MI, MN, MS, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WI and including Canada. | Terminated | September 10, 2013 | 1 | 263 | Computer
Z-2177-2013 | 64601 | microscope, surgical, general & plastic surgery | FSO | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Surgical microscopeThe device is used to improve the visibility of objects through magnification and illumination and can be applied for observation and documentation and for human veterinary medical treatment. | Leica M720 OH5 | 2 | September 10, 2013 | 2013 | Leica Microsystems (Schweiz) Ag | The brake in the Z axis and/or brakes in the AB and C axes failed during installation. | DESIGN: Device Design | The firm, Leica Microsystems, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to DISCONTINUE USE until your system software is updated and to complete and return the attached Acknowledgement Form via email at LMGRA@leica-microsystems.com or fax to: 847-236-3747 or +41 71 726 32 49 within 10 working days after receipt of field safety notice. Your Leica Mycrosystems will contact you to make the necessary arrangements to complete an upgrade at no charge to the customers. If you have any further questions, contact Regulatory Affairs/Quality Assurance Manager at 847-405-5413 or +41 71 726 32 16. |  | 28 | 28 units | Worldwide distribution: USA (Nationwide) including states of: AK, CA, CO, NY, and OR; and countries of: Belarus, China, Hong Kong, India, Indonesia, Italy, Norway, South Africa, Spain, Sweden, Turkey, and Vietnam. | Open |  |  | 33 | Computer
Z-2191-2013 | 66133 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Vision One Laser System Model GA-0025020, Serial No: 10050, 10051, 10053.Intended for use in the treatment of ocular pathology. | Vision One Laser System Model | 2 | September 11, 2013 | 2013 | Lumenis, Inc. | Lumenis has initiated a recall on certain models of Vision One System due to a potential for unintended laser exposure to the user. | DESIGN: Software Design | Customer's were notified via letter on 8/19/13. Service visits to consignees were scheduled with anticipated completion on 8/26/13 to replace the control board. Consignee monitoring was performed by use of return reply verification tracking cards returned by the Lumenis service engineer after completion of CPU board replacement. |  | 3 | 3 units | Distributed in the states of NJ, PA, and IL. | Terminated | December 19, 2013 | 99 | 3 | Computer
Z-0026-2014 | 66171 | implantable pulse generator, pacemaker (non-crt) | LWP |  |  | Cardiovascular |  | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models F102, F110, F111. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. Indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. | Boston Scientific TELIGEN ICD | 2 | October 17, 2013 | 2013 | Boston Scientific CRM Corp | The performance of a low voltage capacitor in a subset of COGNIS CRT-Ds and TELIGEN ICDs manufactured prior to December 2009, may be compromised over time, causing increased current drain that can lead to premature battery depletion. | DESIGN: Component Design/Selection | An Important Medical Device Information letter dated August 2013 was sent to implanting and following physicians. The letter provided a description of the issue, describes clinical implications, provides the rate of occurrence, and gives recommendations for monitoring patients. |  | 38500 | 38,500 total worldwide | Worldwide Distribution-USA (nationwide) including Guam and Puerto Ricoand the countries of Argentina, Aruba, Australia, Austria, Belgium, Bermuda, Brazil, Canada, Chile, Hong Kong, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Finland, France, Gabon, Germany, Greece, Hungary, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Peru, Poland, Portugal, Saudi Arabia, Singapore, Si Lanka, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, UK, United Arab Republic, and Venezuela. | Open |  |  | 38500 | Computer
Z-0050-2014 | 66361 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Solea 1.0 Laser Surgical instrument for use in general and plastic surgery and dermatology. | Solea 1.0 | 2 | October 22, 2013 | 2013 | Convergent Dental | Specific error condition could lead to unintended laser emission during laser activation. | DESIGN: Software Design | The following Notice of Correction letter was sent out to all customers via email on 9/12/13, and followed up with a direct phone call to confirm receipt and review the contents of the letter: Dear Dr. [Customer name],This email is to inform you of a pending notice of correction regarding your Solea system, serial #[Customer serial number].We have identified a specific error condition that could lead to unintended laser emission during laser activation. The software defect that caused the error condition has been identified and tested.Please do not use your Solea system until further notice. Resolution of this issue will require the installation of an updated release of the Solea software. A representative from Convergent Dental will be in touch with you to schedule the upgrade.Convergent Dental is committed to product quality and patient safety. If you have questions about this issue, please contact us. Software fix and testing:New software version (Release 2.0.4.0) was created by the software development group. Convergent Dental has implemented two software changes to eliminate this problem. The first is a defect fix to correct the handling of fatal errors, specifically laser communication timeout errors, so that these errors are not able to be cleared. This will prevent the user from being able to re-enter the workflow after a fatal error. The second is a defect fix to not set the laser state to disconnected when a laser timeout error condition occurs. This will ensure that the command to stop laser emission is always sent to the laser control module anytime the laser cutting state is exited due to error condition or the normal workflow event of stepping off the foot pedal.The software development team ran full Unit and Integration Tests on the new software build. The software verification activities included execution of the full verification test plan (Doc #790-00001), which has been revised to include specific defect . |  | 5 | 5 | US distribution including states of: MA, CT, NY. | Terminated | November 15, 2013 | 24 | 5 | Computer
Z-0089-2014 | 66472 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT is a gamma camera for Single Photon Emission Computed Tomography (SPECT), one unit per box.BrightView XCT is a gamma camera for Single Photon Emission Computed Tomography (SPECT) and integrates with an attenuation device consisting of flat panel x-ray imaging components. BrightView XCT produces non-attenuation corrected SPECT images and attenuation corrected SPECT images with x-ray transmission data that may also be used for scatter correction. The nuclear medicine images and the XCT images may be registered and displayed in a fused format (overlaid in the same orientation) to provide anatomical localization of the nuclear medicine data. | BrightView XCT | 2 | October 28, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips received reports from the field that the Flat Panel Detector (FPD) does not always remain securely latched in the stowed position. | DESIGN: Device Design | Philips Healthcare sent an Urgent - Medical Device Correction letter (Field Safety Notice (FSN 88200462), dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the notice with all members of their staff who need to be aware of the contents of the letter. Customers were asked to retain a copy with the equipment instruction for use. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. A Field Service Engineer will visit the site to install the corrected software. The firm will be tracking the progress.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center at 1-800-722-9377, option 5. |  | 371 | 371 units | Worldwide Distribution - USA including AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MD, MN, MO, MS, MT, NC, NJ, ND, NE, NV, NY, OH, OK, OR, PA, SD, TX, VT, WA, and WI.Internationally to Canada, Algeria, Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Finland, France, Gabon, Germany, India, Indonesia, Ireland, Italy, Japan, Korea Republic, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open |  |  | 371 | Computer
Z-0133-2014 | 66504 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | The EkoSonic Endovascular System employs high frequency (2-2.5 MHz), low power (0.585 watt average power per transducer) ultrasound to facilitate the delivery of physician-specified fluids in the peripheral vasculature and to deliver solutions into the pulmonary arteries.The EkoSonic Endovascular System is intended for the controlled and selective infusion of physician-specified fluids, including thrombolytics, into the peripheral vasculature and the infusion of solutions into the pulmonary arteries. The reusable EKOS EkoSonic Control Unit provides power to the system and the user interface for operator control. A reusable, non sterile Connector Interface Cable connects the EKOS EkoSonic Control Unit to the EkoSonic Endovascular Device. | EKOS EkoSonic Control Unit | 2 | November 01, 2013 | 2013 | EKOS Corporation | The EkoSonic Control Unit is recalled because it had the incorrect power entry module installed. It is missing an AC filter to reduce the noise and electrical interference. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | EKOS notified their customer by email on October 2, 2013, about the recall. EKOS has requested the customer to return of EkoSonic Control Unit PT3B-783, so that it can be replaced.CORRECTIVE ACTION:This single affected unit will be reworked with the correct power entry module and released into finished goods as a refurbished control unit.EKOS did not distributed any affected devices in the US.For any questions, please call EKOS Corporation at 425-415-3100. |  | 1 | one unit | Worldwide Distribution in Germany only. | Terminated | November 22, 2013 | 21 | 1 | Computer
Z-0142-2014 | 66464 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Heart-lung machine, Stockert S5 System. Catalogue No: 48-30/40/50-00.The Stockert S5 System is intended to be used during cardiopulmonary bypass for procedures lasting six(6) hours or less. | Sorin S5 Perfusionn System | 2 | November 04, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin Group is recalling certain models of S5 Perfusion Systems due to malfunction of cardioplegia control. |  | Sorin Group sent an Urgent Field Safety Notice dated September 24, 2013, to all affected customers. The letter informed users of the potential for failing of automatic cardioplegia delivery function and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to update the firmware on site. Customers were asked to complete the Response form to confirm they had received, read and understood the Field Notice. For questions regarding this recall call 303-467-6527. |  | 84 | 84 | Worldwide Distribution - USA including WA, NY, VA, TN, SC, TX, CA, CO, MA, WI, MD, AL, ND, and Puerto Rico and Internationally to Italy, Finland, Germany, Sweden, and Great Britain. | Open |  |  | 84 | Computer
Z-0348-2014 | 66773 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Candela Laser GentleLase Pro, Powered Laser Surgical Instrument for dermatological use - Model number: 9914-00-9015.and Candela Laser GentleLase Pro LE, Model number: 9914-00-9040 | Candela Laser | 2 | November 18, 2013 | 2013 | Candela Corporation | Unintended single pulse maybe emitted to handpiece prior to pressing of trigger or footswitch and cause injury. | DESIGN: Software Design | Candela notiifed Distributors on 11/8/13 through e-mail, Certified return receipt request, or FedEx package containing distributor letter, customer letter for translation and customer lists. A software update to Version 2.0 will be issued to the user by a field service representative. |  | 1258 | 1258 units | Worldwide distribution: US (nationwide) and Internationally to: Azerbaijan, Bulgaria, Chile, Czech Republic, Greece, Israel, Kuwait, Qatar, Saudi Arabia, Serbia, Turkey, and the United Arab Emirates and to Subsidiaries in Australia, France, Germany, Italy, Japan, Portugal, Spain, and the United Kingdom. | Open |  |  | 1258 | Computer
Z-0540-2014 | 66987 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MedfusionÂŽ 4000 Syringe Infusion PumpsMedfusionÂŽ 4000 Syringe Infusion Pumps are designed for use in care areas for the administration of fluids requiring precisely controlled infusion rates. These pumps are indicated for use by healthcare professionals in the delivery of IV solutions, drug solutions, blood or blood products, and enteral feedings in continuous, intermittent, and body weight delivery modes. | Medfusion 4000 Syringe Infusion Pumps | 2 | December 19, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a recall involving all MedfusionÂż 4000 Syringe Infusion Pumps.Smiths Medical has identified an issue with electrical shorting at the AC power receptacle. Shorting can occur when saline solution or other conductive fluids (e.g., potassium chloride) enter the AC power receptacle. When this happens, sparking and/ or flames, accompanied by smoke, may appear at the AC po | DESIGN: Device Design | Smiths Medical sent an "Urgent Medical Device Correction" letter dated November 22, 2013, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.All customers will be provided with instructions on inspection, cleaning, and prevention regarding this issue, for the continued safe and effective use of these pumps. Additionally, customers will be instructed to:1. Return damaged pumps for servicing or damaged AC power cords for replacement to Smiths Medical.2. Return the completed Confirmation Form to acknowledge receipt and understanding of the information provided in the Urgent Medical Device Correction letter, via fax or email, to Smiths Medical for processing."The Instructions for preventing this issue", "Mitigations", Conclusions", " Action to be taken by the User", and "Transmission of this Urgent Medical Device Correction". The consignees were also asked to complete the "Urgent Medical Device confirmation Form" and fax it to 1-800-237-8033 or by sending an electronic copy via e-mail to medfusion4000nov2013@smiths-medical.com.For questions regarding this information, consignees should contact their local Sales Representative, Smiths Medical's Customer Service Department at 1-800-258-5361, or via e-mail at medfusion4000nov2013@smiths-medical.com.UPDATE: Consignees were sent a follow-up letter dated 13 January 2014 from the one dated November 22, 2013. The letter was addressed to Risk / Safety Managers, Biomedical Professionals, Clinicians who oversee the use of Medfusion Pumps, Distributors, and other users of these devices. The "Urgent Medical Device Recall" letter included Advise on Action to be Taken by the User. This included customers to have their Pumps undergo a free correction. The correction consists of the installation of a new AC power Receptacle seal. Smith Medical provided options for completing the correction process which included: 1) Consignees were a. |  | 9560 | 9560 (8601 US, 959 OUS) | Worldwide Distribution - USA (nationwide ) including the states of : AZ, NY, TN, MS, AR, WV, VA, VT, PA, DE, DC, OH, CO, NH, MI, CA, FL, CT, WA, MD, NM, RI, MA, TX, NC, IL, KY, WI and MI., and the countries of CANADA, SAUDI ARABIA and AUSTRALIA. | Open |  |  | 9560 | Computer
Z-0616-2014 | 67002 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Distributed by STERIS Corporation, Mentor, OHThe SYSTEM 1E Liquid Chemical Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned immersible and reusable critical and semi-critical heat sensitive medical devices, including endoscopes and their accessories. | SYSTEM 1E Liquid Chemical Sterilant Processing System | 2 | December 31, 2013 | 2013 | Steris Corporation | Customers indicated fluctuations in environmental conditions sites and gradual degradation/deformation of check valve o-rings following frequent contact with S40 sterilant use solution result in nuisance aborts. | DESIGN: Device Design | STERIS sent an Urgent Voluntary Field Correction Notice dated November 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a STERIS Field Service Representative would contact their facility to schedule installation of the new hardware and updated software on their SYSTEM IE Liquid Chemical Sterilant Processing System(s). Customers were instructed that they may continue to operate the SYSTEM IE processor prior to the completion of the update. Customers with questions were instructed to contact STERIS Field Service Dispatch at 1-800-333-8828.For questions regarding this recall call 440-392-7519. |  | 6028 | 6,028 units | Nationwide Distribution - USA (nationwide) and the District of Columbia and Puerto Rico.Product was also shipped to the following countries: Afghanistan, Bahrain, Hong Kong, Italy, Saudi Arabia, Turkey & United Kingdom. | Open |  |  | 6028 | Computer
Z-0643-2013 | 63030 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips and Laerdal brands of HeartStart HS1 Defibrillator Family Model number M5066A, M5067A, and M5068A Automated External Defibrillator Manufactured by Philips Medical Systems, Seattle, WA 98121 USA.The HS1 Defibrillator is intended to treat ventricular fibrillation, the most common cause of sudden cardiac arrest. Using voice prompts, light emitting diodes (LEDs) and buttons, the user is guided through the response. The HS1 uses a SMART biphasic, impedance compensating exponential waveform to deliver a nominal 150 J to adults and nominal 50 J to infants/children. | Philips and Laerdal brands of HeartStart HS1 Defibrillator Family | 2 | January 04, 2013 | 2013 | Philips Medical Systems North America Co. Phillips | Philips determined that it is important to clarify information in the Owner's Manual and keep customers informed about the maintenance of their FRx (models 861304 and 861305), OnSite (models M5066A and M5067A), or HeartStart Home (model M5068A) automated external defibrillators (AEDs) shipped between December 2005 and July 2012. Philips had determined there is a need to emphasize the importance o |  | Philips Healthcare sent an Emergency Care and Resuscitation letter Customer Information Letter dated July 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Philip's representative to obtain the tools and information they need to ensure that their device is functioning properly.For questions regarding this recall, please call AED Recall Support at 1-800-263-3342 (Option 6). |  | 613000 | Estimate 613,000 devices shipped | Worldwide Distribution - USA (nationwide) and the countries of Japan and Malaysia. | Open |  |  | 814598 | Not_Computer
Z-0764-2013 | 63485 | pump, infusion, enteral | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Infinity Enteralite Ambulatory Feeding Pumps, List number: INFKIT2 - Addendum to Operator's Manual Product usage: provide a means for the volumetric delivery of fluids used in Enteral applications. | Infinity Enteral Feeding infusion pump | 2 | January 31, 2013 | 2013 | MOOG Medical Devices Group | The firm is adding an addendum to the operator's manual for Infinity Enteralite Ambulatory Feeding Pumps pertaining to the Infinite Dose mode and the inability of the Pump to detect air in line. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An Urgent Field Safety Notice was sent to all original consignees by certified mail beginning on Monday, October 15, 2012. Consignees were provided with an Addendum to the Operator's Manual as part of the Notice. Consignees were instructed to assure pump users are aware of the information in the addendum and to store the addendum with the Operator's Manual. Consignees were also requested to complete and return a Medical Device Recall reply form to the firm. For additional information or clarification, consignees are to contact Holly Imes, Manager, Regulatory Affairs at 801-264-1001, ext. 230 or via email at: himesstark@moog.com. |  | 92488 | 92,488 units | Worldwide Distribution and US Nationwide in the countries of (Germany, Sweeden, Croatia, Lithuania, and Dubai). | Open |  |  | 92488 | Not_Computer
Z-0749-2013 | 63761 | ventilator, emergency, manual (resuscitator) | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Giraffe and Panda Bag and Mask Resuscitation SystemThe T-piece Resuscitation System [ventilator, emergency powered resuscitator]/Panda Bag and Mask Resuscitation System provides the basic equipment required for pulmonary resuscitation of infants. | Giraffe and Panda Bag and Mask Resuscitation System | 1 | February 08, 2013 | 2013 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly. | TRAINING: Employee Error | GE Healthcare sent an Urgent Medical Device Correction letter dated October 15, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: It is recommended that a qualified biomedical technician or respiratory therapist identify the affected resuscitation system and locate the air and oxygen fittings and labels at the back of the unit. With the table provided, identify the appropriate fitting for their country and location and verify the air fitting is located on the left side and the oxygen fitting is located on the right. Verify the air label is located on the left side and the oxygen label is on the right and verify that the color of the air and oxygen labels match the picture on table provided. If the unit does not meet all of the criteria, the unit will require repair. If the unit in use with patient, disconnect the wall supply gas and switch to tank supply. Once the patient no longer requires respiratory support, transfer the patient to another bed and remove the unit from service. A Field Engineer will be deployed to your site to complete the required repair. If the unit NOT in use with Patient, take the unit ouf of service. A Field Engineer will be deployed to your site to complete the required repair. For questions or concerns regarding this notification, please call the phone number 1-800-345-2700. For questions regarding this recall call 262-513-4122. | Contact the recalling firm for information | 276 | 276 units total, 109 of those nationwide. | Worldwide Distribution - Worldwide Distribution - USA including AL, AZ, CA, CO, FL,GA, IA, ID, IL, KY, MA, MD, MI, MN, MO, MS, MT, NC, ND ,NE, NM, NV, NY, OH, OR, PA ,TX, VT, WA, WI and WV; and the countries of Australia, Belgium, Canada, China, France, New Zealand, Poland, Saudi Arabia, Switzerland and United Kingdom. | Open |  |  | 276 | Not_Computer
Z-0837-2013 | 64076 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pro Time Microcoagulation System/ Pro Time 3 Cuvette, Pro3-4, Pro3-25, and Pro3-50.International Technidyne Corporation. The Pro Time Microcoagulation System is a portable battery operated instrument with a disposable cuvette for quantitative determination of prothrombin time from fingerstick whole blood or anticoagulant free venous whole blood. | Pro Time Microcoagulation System/ Pro Time 3 Cuvette | 2 | February 19, 2013 | 2013 | International Technidyne Corp. | Some Pro Time3 Test Cuvettes within a specified lot range may recover lower than expected Prothrombin Time/International Normalized Ratio (PT/INR) results. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | ITC sent an 'Urgent Medical Device Recall' letter/Health Professional Account Tracking Form dated October 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the form whether or not they have remaining inventory of these lots and return the form promptly. All unused returned product will be replaced. Any questions can be directed to ITC Technical Support at 800-631-5945 Extension 4707 or e-mail ProTime@itcmed.com.For questions regarding this recall call 800-631-5945, ext. 4707. |  | 34710 | 11,742 boxes US (25 cuvettes per box); 22,968 boxes OUS (25 cuvettes per box) | Worldwide Distribution - USA (nationwide) and Internationally to Bermuda, Canada, Chile, Denmark, Germany, Italy, Latvia, Lithuania, Malaysia, Philippines, South Africa, Spain, Switzerland, Taiwan, and United Arab Emirates. | Terminated | January 09, 2014 | 324 | 34710 | Not_Computer
Z-0931-2013 | 64165 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI. Radiation Therapy Digital Imager. Product XVI R3.5.1, R4.2.1, and R4.5.1 Used as part of radiation therapy treatment process. | Digital Imager, Radiation Therapy | 2 | March 10, 2013 | 2013 | Elekta, Inc. | The DICOM tags do not contain sufficient information for XVI to identify the different volumes that have the same name. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated December 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The IFSN informs the user in the TPS, to give the different volumes different structure names. After importing the patient data set from the TPS, make sure that the structure contours are correct. Users are requested to return the signed acknowledgement. Contact your local Elekta office for questions regarding this notice. |  | 164 | 164 | Nationwide Distribution-DC, Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Open |  |  | 164 | Not_Computer
Z-0974-2013 | 64435 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | INTIO Inc., ClearStart-SVM Segmentation and Volumetric Measurement System, Users Manual. Part #: UM-06-20-0003. ClearStart-SVM is used for patient assessment using CT (computed tomography) images, in order to plan potential therapies or monitor tumor response to treatment. | INTIO Inc. ClearStartSVM | 2 | March 20, 2013 | 2013 | Intio Inc | INTIO Inc. distributed a SVM User's Manual and it has an editing mistake that could be interpreted as promoting an unapproved use. | MISBRANDING: Labeling False and Misleading | INTIO sent a Recall Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. INTIO has corrected this version of the SVM User's Manual. An INTIO Inc. employee will contact the customers to replace the current manual with the correct one. Consignees were notified by letter that an INTIO Inc. employee will contact you within the next 30 days to replace the current manual with the correct one. The new SVM User's Manual will be provided to you once INTIO Inc. has retrieved the incorrect version of the SVM User's Manual. Customers with questions were instructed to email support@intio.us or call 303-396-1777.For questions regarding this recall call 303-396-9291. |  | 7 | 7 | Nationwide Distribution including MN, FL, PA, WI, NY, and RI. | Terminated | April 17, 2013 | 28 | 7 | Not_Computer
Z-0988-2013 | 63494 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) Alkaline Phosphatase (ALP) CalibratorThe ALP CAL is an in vitro diagnostic product for the calibration of the Alkaline Phosphatase method on the Dimension Vista" System. | Siemens Dimension Vista(R) Alkaline Phosphatase (ALP) Calibrator | 3 | March 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Correlation slope bias and accuracy shift | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent a Urgent Medical Device Correction letter dated October 16, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to calibrate with newer lots of Dimension Vista(R) ALP Calibrator as they become available, beginning with lot 2GD011. Customers were also asked to discontinue use and discard any remaining inventory of lot 2AD052. Customers further instructed to complete the attached response form and fax to (302) 631-8467, to indicate they had received the information and to facilitate their no charge replacement.Further questions please call (302) 631-6311. |  | 1985 | 1985 cartons | Nationwide distribution: USA including Puerto Rico and the states of AL, OH, MO, GA, IL, CT, IA, SC, VA, TX, FL, NY, NC, CA, NM, AZ, CA, MI, NJ, PR, MS, NE, WV, PA, MT, MA, MD, KS, NH, TN, VT, LA, CO, ID, WA, NH, DE and OR. | Terminated | January 08, 2014 | 289 | 1985 | Not_Computer
Z-0987-2013 | 63946 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | ARCHITECT HBsAg Qualitative Assay, List Number 4P53 and ARCHITECT HBsAg Qualitative Confirmatory Assay, List Number 4P54The ARCHITECT HBsAg Qualitative assay is a chemiluminescent microparticle immunoassay (CMIA) for the qualitative detection of hepatitis B surface antigen (HBsAg) in human adult and pediatric serum and plasma and neonate serum. | ARCHITECT HBsAg Qualitative Assay, ARCHITECT HBsAg Qualitative Confirmatory Assay | 2 | March 25, 2013 | 2013 | Abbott Laboratories, Inc | Abbott has identified that assay files for the ARCHITECT HBsAg Qualitative and HBsAg Qualitative Confirmatory assays (LNs 4P53 and 4P54) on the ARCHITECT i1000SR System were mistakenly made available on abbottdiagnostics.com in October 2012. Additionally, some i1000SR customers received notification through AbbottLink in November 2012 that assay files for LNs 4P53 and 4P54 were available. The AR |  | The firm decided to correct the issue and and sent out Urgent Field Safety Notices to their customers dated 10/18/2012. | Contact the recalling firm for information | 12 | 12 | Nationwide Distribution including AZ, CA, FL, GA, HI, LA, MA, MD, MO, NJ, NY, PA, RI, TN, TX, UT, and WA. | Terminated | July 31, 2013 | 128 | 12 | Not_Computer
Z-1078-2013 | 64398 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Quasar Nuclear Medicine System, Hawkeye Option and Hawkeye 4 Option for Dual-Head Variable Angle Gamma Camera. Nuclear Medicine imaging systems. Recall includes all Infinia systems, all configurations. |  | 2 | April 10, 2013 | 2013 | GE Healthcare, LLC | Accelerated fatigue of the lateral axis motion subsystem of the Infinia Computed Emission Tomography System could result in mechanical failure causing the detector to slide; may impact patient and operators. | DESIGN: Device Design | GE Healthcare sent an " IMPORTANT PRODUCT INFORMATION" letter dated March 6, 2012. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers and Managers of Radiology/Cardiology. The letter described the Issue, instructions, Affected Product Details, and Product Correction. For questions contact your local GE Healthcare Service Representative or call the GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other Countries 262-896-2890. |  | 2180 | 2180 | Worldwide Distribution - USA (nationwide) including DC and PR and the countries of UKRAINE, TAIWAN, TURKEY, TUNISIA, THAILAND, SINGAPORE, SWEDEN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, PAKISTAN, PHILIPPINES, PANAMA, NEW ZEALAND, NORWAY NETHERLANDS, MALAYSIA, MEXICO, MALTA, MARTINIQUE, MOROCCO, LATVIA, LITHUANIA, LEBANON, KUWAIT, KOREA, JAPAN, JORDAN, ITALY, IRAN, INDIA, ISRAEL, IRELAND, INDONESIA, HUNGARY, CORATIA, HONG KONG, GREECE, GEORGIA, UNITED KINGDOM, FRANCE, FINLAND, SPAIN, EGYPT, ECUADOR, ALGERIA, DOMINICAN REPUBLIC, DENMARK, GERMANY, CZECH REPUBLIC, CYPRUS, COLUMBIA, CHILE, SWITZERLAND, CANADA, CHINA, BRAZIL, BELGIUM, AUSTRALIA, AUSTRIA, ARGENTINA, SOUTH AFRICA, VENEZUELA, URUGUAY, ALGERIA. | Terminated | January 27, 2014 | 292 | 2180 | Not_Computer
Z-1083-2013 | 64759 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | VMX, VMX Plus Mobile Radiographic Unit with the following models:45296139; 45296200; 45296350; 45296477; 45296964; 45297001; 45296966; 45297003; 2240169; 2240171; 2240172; 2240173.Product Usage: The VMX is a mobile radiography system incorporating stylish design for totally adapted use in intensive care, emergencies, pediatrics, neonatology and orthopedics. | VMX, VMX Plus | 2 | April 10, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the base casting for the articulated arm of your VMX andVMX Plus Mobile X-ray systems that may impact Operator or Patient safety. | DESIGN: Device Design | GE Healthcare sent and "Urgent Medical Device Correction" letter dated January 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions contact your local GE Healthcare Service Represenative. |  | 1296 | 1296 | Worldwide Distribution - USA (nationwide) including the states of: AR, CA, GA, IL, IN,IA,LA, MA, MN, MT, NY, OH, PA, TX, VA and the countries of BELGIUM, BURKINA FASO, BRAZIL, CHINA, CANADA, SWITZERLAND, CHILE, COLOMBIA, CYPRUS, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, EGYPT, SPAIN, FRANCE, UNITED KINGDOM, CHANA, GUADELOUPE, GREECE, GUYANA, HONG KONG, CROATIA, HUNGARY, INDONESIA, ISRAEL, ITALY, KUWAIT, LEBANON LUXEMBURG, LIBYAN, MOROCCO, MARTINIQUE, MEXICO, NETHERLANDS, NORWAY, PANAMA, PERU, POLAND, PORTUGAL , REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SWEDEN, SINGAPORE, SLOVENIA, SLOVAKIA, SIERRA LEONE, THAILAND, TUNISIA, TURKEY, VENEZUELA, MACEDONIA, PALESTINE, SOUTH AFRICA. | Terminated | February 04, 2014 | 300 | 1296 | Not_Computer
Z-1216-2013 | 60555 | light source system, diagnostic endoscopic | OAY |  |  | Gastroenterology/Urology | PMA | KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System,Instruction Manuals: PDD Telescopes, Tricam SL II, and PDD Camera Heads.Used as an optional accessory to white light cystoscopy when used in combination with the diagnostic imaging drug Cysview. | Karl Storz Photodynamic Diagnosis (PDD) DLight C System, | 2 | May 08, 2013 | 2013 | Karl Storz Endoscopy America Inc | The recall was initiated because Karl Storz has confirmed that the Photodynamic Diagnosis (PDD) D-Light C System instruction manuals have important safety-related changes. | PRODUCTION CONTROLS: Process Control | A Safety Alert letter, dated 8/16/12, was sent to all their customers to inform them that there is a labeling error in which important safety-related changes were made to the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System instruction manuals. The correct Ethylene Oxide (EtO) Sterilization set point for Relative Humidity should read, =70% RH. The correction applies to the PDD Camera Head instruction manual, PDD Rigid Telescopes instruction manual, and Tricam SL II instruction manual, the three component manuals indicated for use within the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System. The letter instructed the customers to place each of the three enclosed stickers, in the appropriate instruction manual. Customers are instructed to cover the erroneous information with the provided stickers as indicated. Customers with any questions regarding the proper use of the device, are instructed to contact Karl Storz's Technical Support Department at 800-421-0837, Ext. 5350, during office hours; or Ext. 7222, 24 hours/7 days a week for emergencies. | Contact the recalling firm for information | 15 | 15 units | Nationwide distribution: USA including states of: IL, CA, WV, TX, SC, WI, NY, PA, NE, NJ, MD, OH, IA, and NC. | Terminated | May 08, 2013 | 0 | 15 | Not_Computer
Z-1426-2013 | 65072 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Thyroidectomy Indication for the da Vinci Surgical Systems;Product Usage:da Vinci Thyroidectomy offers a minimally invasive surgery for both partial and total thyroidectomy procedures, including endoscopic and video assisted procedure that are preferable to conventional (open) procedures that utilize a cervical incision because they leave no visible neck scar. | daVinci TORS Indication of Use | 2 | May 29, 2013 | 2013 | Intuitive Surgical, Inc. | Promotional literature for use of the da Vinci system for thyroidectomy, which had not been approved with a change to the 510(k). | PREMARKET APPROVAL: No Marketing Application | Intuitive sent a notification letter on October 17, 2011 by email to all consigness whose email address were on file. Also a copy of the letter was sent by registered mail on February 8, 2012 to all surgeons whose addresses were on file. The letters informed consignees that Intuitive Surgical were in the process of reviewing labeling for the da Vinci System in regard to use for thyroidectomy procedures. During this period, Intuitive will not provide or facilitate any promotional, training, case observations, proctoring or in case procedural support for use of the da Vinci System in conjunction with thyroidectomy procedures. The letters recommend consulting the institution policy regarding use of approved or cleared medical devices for procedures that are not specified in the labeled indication for use. Questions may be directed to customer service at 1-877-408-3872. |  | 2423 | 2,423 consignee | USA Nationwide distribution | Terminated | May 30, 2013 | 1 | 2423 | Not_Computer
Z-1424-2013 | 65094 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | User Manual addendum, TORS for da Vinci Surgical System IS 1200, IS 2000 and IS 3000.Product Usage:The user manual for Transoral Surgery (TORS) provided information to update and replace the Indications for Use and Representative uses section and add the section Warnings for Transoral Otolaryngology Surgery. | daVinci | 2 | May 29, 2013 | 2013 | Intuitive Surgical, Inc. | Change to User Manual to remove one indication for use and amend warnings to emphasize restrictions for use. | MISBRANDING: Labeling False and Misleading | Intuitive sent an Important Medical Device Correction letters dated January 24, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate and discard the existing TORS addendum, place the new TORS addendum included with this letter with your System User Manual, and complete and return the attached Acknowledgement Form to Intuitive Surgical. For questions call Intuitive Surgical Customer Service at 1-800-876-1310. |  | 1374 | 1,374 | Worldwide distribution- USA Nationwide-list of countries not available. | Terminated | May 30, 2013 | 1 | 1374 | Not_Computer
Z-1481-2013 | 65140 | system, tomography, computed, emission | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView SPECT gamma camera for Emission Computed Tomography, Philips Medical Systems, Cleveland, OH 44143BrightView SPECT is a gamma camera designed for single or dual detector nuclear imaging accommodating a broad range of emission computed tomography Emission Computed Tomography (ECT) studies. | BrightView SPECT gamma camera for Emission Computed Tomography | 2 | June 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips has received one report from the field that there was an unexpected motion of the detector due to a failure of the mechanical assembly. | DESIGN: Device Design | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer were not required to take any specific action. Customers who need further information or support concerning the affected product, they should contact their local lPhilips representative or local Philips Healthcare office. For North Americ and Canada they should contact the Customer Care Solutions Center (1-800-722-9377, option 5).For questions regarding this recall call 440-483-7600. |  | 505 | 505 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV & WY.Internationally to Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, France, Germany, India, Indonesia, Iran, Italy, Japan, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Syria, Thailand, Turkey, Taiwan, Ukraine, United Arab Emirates, United Kingdom, Venezuela & Viet Nam. | Terminated | February 12, 2014 | 252 | 505 | Not_Computer
Z-1480-2013 | 64927 | bariatric bed | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | E-Z WIDER BARIATRIC CHAIR BED; Single Motion Expandable Width Technology | EZ Wider Bariatric Bed | 2 | June 05, 2013 | 2013 | Kreg Medical Inc. | An EZ Wider bed (capacity of 1000 lb. safe working load collapsed during patient use, leaving the patient in a forward tilt position. | DESIGN: Device Design | Kreg Medical sent an Urgent Voluntary Medical Device Recall letter dated May 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Kreg Medical started replacing potentially defective beds with refurbished beds in August 2012. All of the potentiall defective beds were refurbished by November 20, 2012. Customers were requested to send back the Medical Device Recall Return Response to confirm receipt of the notification letter.For questions regarding this recall call 312-275-7070. |  | 59 | 59 beds | Nationwide Distribution including IL and IN. | Open |  |  | 59 | Not_Computer
Z-1532-2013 | 65175 | ventilatory effort recorder | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | embla Embletta Gold (It functions as a Polysomnography Recorder and Ventilator Effort Recorder).The Embletta Gold is a battery operated device which records physiologic signals used in the diagnosis of sleep disorders. It is intended to be used for adult and pediatric (excluding neonatal and infant) studies. The recorder is not equiped with an alarm device and is not intended to be used as a life monitor. It may be used in a variety of configurations and the recording of physiologic signals of a sleeping patient.Item Numbers: RMIV003465RMIV017570RMRE008075UI009108UI010814UI013848RMIV019053RMIV019328RMIV019340RMIV019634RMIV020254101000555210100055851010002453101000158813515066810100039151010002669RMIV0195421010002370UI018933RMIV019638RMIV019850RMIV020166UI012625RMIV0192351667585801010007510UI007649RMIV020079RMIV0131961010002697RMIV019658UI016466RMIV0196591010002034UI016112UI0181601010007423RMRE007779UI018158RMIV02014510100013321010002431101000243210100013781010006090101000748410100047501010003658101000735110100026881010002699UI008431RMIV009745RMIV01967910100070051010002088UI00936510100036111010006891UI018570UI0225401010002262RMIV019517RMIV019572RMRE007959UI006113RMIV006948UI013856RMIV016367RMIV019395RMIV02011210100054341010006883RMRE007326RMIV0192371010006887101000199110100037321010003984UI002603101000426610100064781010006800RMIV01939710100019281010001729RMIV019166RMIV019746RMRE008072UI016332UI0076311010002320RMIV018044RMRE007777101000306710100050001010005287RMIV019076RMIV020194101000024310100035431010006931RMIV017565RMIV019454RMRE008087156096159046RMIV02002910100012561010003438RMIV019940RMRE0080741010003546UI0126391180186361010003507RMIV003685UI00521610100033811010002913101000292010100028451685741121010006576101000270210100051971010006074UI01259310100019371010003637UI0125861010002664101000453510100038171010005998101000682610100077101010003802101000096510100010011010001006101000101310100026591010002671101000269810100026701010002646101000279110100028351010002830101000279310100035591010003562101000321510100041451010004220101000558010100055761010005464101000549710100054691010005314UI019628UI007951UI011892RMRE004711RMIV013784UI015720RMIV019570RMRE008077RMRE008073101000118810100064371010004274UI006849UI009712RMIV019889UI003213UI005546UI011507UI015675RMIV019238RMIV019254RMIV019511101000125510100035651010002955UI018536RMIV019954RMIV0191771010001217UI00803810100049851010002682167617437RMIV0202431010005643UI0036791010000951UI013298RMIV020078RMIV02024710100057301010003680RMIV02005410100001991010007284UI012571UI0125751010002675UI014387RMRE007776IN000380UI014119UI016006UI016355UI017046RMIV019982UI005858RMIV0199191010039241010002837UI_000670UI_001339UI004904RMIV007125UI014813UI017060UI019513RMIV019041144913RMIV019639RMIV0197781010002678101000432510100043101596991010003947UI011502UI020772 | Embletta Gold | 2 | June 11, 2013 | 2013 | Embla System Llc | Calibration error for Embla Embletta Gold units calibrated through service center between June 15, 2012 to February 27, 2013. This calibration error could cause a clinician to over titrate the patient and prescribe a therapy pressure higher than is necessary. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Embla Systems, LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter on April 24, 2013, to their customers via courier. The letter described the product, problem and actions to be taken. The customers were instructed to return all recalled units for recalibration and to contact Technical Support at 1-888-662-7632, option #1, at your earliest convenience; Reference the EBG Field Corrective Action in order to obtain your return number. Embla will provide overnight shipping both ways and a three business day turn around for affected units. If you have transferred any of these products to another location, please forward a copy of this urgent notification to that location and notify Manager, Quality and Regulatory Affairs, at 613-254-8877 ext 1105 or via email at EBGFieldCorrectionProgram@natus.com. | Contact the recalling firm for information | 519 | 519 units were calibrated at the service center | Worldwide distribution: US (nationwide) including states of: CO, AK, WA, CA, WI, MI, MA OH, DE, FL CA, CT, TX, IL, GA, PA, AZ, DC, VA, MD, UT, IN, MO, NY, NC, OR, NV, NJ, NH, DE, and country of: Canada. | Open |  |  | 519 | Not_Computer
Z-1563-2013 | 65187 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Thermo Scientific MAS Bilirubin for Siemens and Dimension Vista Systems, , Product is manufactured and distributed by Microgenics Corporation, Fremont, CA MAS Bilirubin is intended for use as a consistent test sample of known concentration for monitoring assay conditions in many clinical laboratory determinations. Includes Bilirubin with patient serum specimens when assaying for any of the listed constituents. Assay values are provided for the specific systems listed. The user can compare observations with expected ranges as a means of assuring consistent performance of reagent and instrument. | Thermo Scientific MAS Bilirubin for Siemens and Dimension Vista Systems | 2 | June 21, 2013 | 2013 | Microgenics Corporation | It has been confirmed from in-house testing that Direct Bilirubin (DBIL), and Total Bilirubin (TBIL) claimed in Thermo Scientific MAS Bilirubin Systems are displaying reduced stability. |  | ThermoFisher Scientific sent a Product Recall Notification letter dated May 13, 2013, by Certified US mail to all affected customer. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed1. Determine if you are using or have inventory of the recalled product.2. Discontinue use of the product immediately.3. Evaluate quality controls. They may vary among laboratories and should be evaluated according to laboratory procedures.4. Retain a copy of the letter for laboratory records.5. If any of the affected lots of MAS Bilirubin had been forwarded to another laboratory, please provide a copy of the recall letter to them.6. Complete and return the response form within 5 days.Letters also provide an alternative product - MAS Bilirubin part number BC103.Questions or concerns may be directed to Technical Services at 1-800-232-3342 or 510-979-5417. |  | 188 | 188 | Nationwide Distribution: US including the states of WI, UT, FL NY, PA, IL, CO, AZ and NJ. | Terminated | August 27, 2013 | 67 | 188 | Not_Computer
Z-1519-2013 | 65372 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE Healthcare Senographe Essential System.The Senographe 2000D system generates digital mamnmographic images that can be used for screening and in the diagnosis of breast cancer. The Senographe 2000D is intended to be used in the same clinical applications as traditional film-based mamnmographic systems. | GE Healthcare Senographe Essential System | 2 | June 26, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to inaccessible Operator Instructions associated with the defective Operator Manual CD of the Senographe Essential. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | GE Healthcare will notify customers with affected systems via letter and enclose a new CD- Rom, along with notification that a GE Representative will schedule an appointment to install the User's Manual along with a GE Service Engineer will go on site to install it on the control station of the Senograph Essential system. 16- H LA 6.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter.GE Healthcare will secure and summarize documentation that the manual has been installed and is available on each affected system and provide summarize of the documentation to the district recall coordinator.For further questions please call ( 262 ) 513-4122. |  | 49 | 49 installed in the US. | Nationwide Distribution including the states of NJ, CO, LA, AL, OH, MA, NC, ME, WI, GA, PA, TX, IN, NY, MO, KY, KS, CA, MN, FL and IL. | Terminated | November 26, 2013 | 153 | 49 | Not_Computer
Z-1663-2013 | 65418 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens MEVATRON series, PRIMART (limited release)PRIMUS and/or ONCOR or ARTIST Linac Systems with BEAMVIEW capability.The intended use of the SIEMENS branded MEVATRON", ARTISTE", ONCOR" and PRIMUS" family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens MEVATRON series, PRIMART (limited release)PRIMUS and/or ONCOR or ARTIST Linac Systems | 2 | July 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The information regarding the FILM Mode calibration may have been considered difficult to interpret when translated. SIEMENS HEALTHCARE CR / RADIATION ONCOLOGY had become aware that this issue may have resulted in mistreatment of patients. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION" letter and Update instruction TH022/13/S beginning on May 29, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.July 19, 2013 - TH023/13/S was distributed is now in progress. This is the FILM Mode Physics Primer Addendum. |  | 1956 | 1956 customers | Worldwide Distribution: USA Nationwide and countries of: Algeria, Angola, Argentina, Austria, Australia, Bahamas, Belgium, Bosnia Herzeg, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Ecuador, Egypt, El Salvador, France, Germany, Greece, Hungary, India, Indonesia, Italy, Iran, Iraq, Ireland, Japan, Jordan, Kenya, DPR of Korea, Kosovo, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Pakistan, Palestinian Ter, Peru, Philippines, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sri Lanka, South Africa, Sweden, Switzerlan, Taiwan, Thailand, Trinidad Tobago,Turkey, United Kingdom, Venezula, Vietnam, Rep of Yemen, and Zambia. | Terminated | January 14, 2014 | 195 | 1956 | Not_Computer
Z-1694-2013 | 65487 | full field digital,system,x-ray,mammographic | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration - full-field digital mammographyProduct Usage: Mammography | Siemens Mammomat Inspiration | 2 | July 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Firm became aware of an unintended behavior when using the Mammomat Inspiration. The possibility exists that a malfunction may occur during Stereo or Tomo acquisitions if the tomo clutch was incorrectly adjusted. Although the possibility is rare, the system could move while a patient's breast is compressed resulting in a possible bruise. Firm released Update Instructions XP031/12/S for this iss | DESIGN: Device Design | Siemens issued an Update Instructions XP031/12/S by letter to the affected customers. For affected sites, a Siemens Service Engineer will inspect the unit and make any adjustments needed. Following the inspection and adjustments, a proper calibration will be completed. For questions call 610-448-3237 and 610-446-4634. |  | 3 | 3 | United States Nationwide Distribution in the states of ND, OH and SC. | Terminated | February 21, 2014 | 225 | 3 | Not_Computer
Z-1783-2013 | 65679 | retractor | GAD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Rultract Retractor Ratchet Cover Plate Distributed by Pemco, Inc., Cleveland, OH 44131Surgical retractor used to lift the incision site | Rultract Retractor Ratchet Cover Plate | 2 | July 22, 2013 | 2013 | Pemco Inc | The firm became aware of the problem when a ratchet cover plate on a device was returned for repair showed signs of corrosion. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Rultract Skyhook sent a Recall Notification letter dated July 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify the device and return preferably just the gearbox portion of the device to:Pemco, Inc.Attn: Cover Plate Replacement5663 Brecksville RoadIndependence, OH 44131For customers who are unable to send the gearbox for recall service, Pemco would replace the Cover Plate free of charge.Customers with questions were instructed to call 216-524-2990, ext 103.For questions regarding this recall call 216-524-2990. |  | 87 | 87 units | Worldwide Distribution - USA including AL, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NM, NV, OK, PA, TN, TX, UT, VA & WA. Internationally to Czech Republic, Italy, Japan, Korea & Singapore. | Terminated | February 11, 2014 | 204 | 87 | Not_Computer
Z-1796-2013 | 65635 | system, multipurpose for in vitro coagulation studies | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thromboelastograph Coagulation AnalyzerProduct Usage:Multipurpose System for In Vitro Coagulation Studies | Thromboelastograph Instrument | 2 | July 24, 2013 | 2013 | Haemonetics Corporation | Research Use Only (RUO) parameters were included in TEG 5000 User Manual. | MISBRANDING: Labeling False and Misleading | Haemonetics sent a Medical Device Correction letter dated June 20, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. The letter instructs customers to discard the current manuals and replace with revised manuals included with the Customer Notification Letter. A Consignee Response Form has been included with the letter, which is to be signed and returned for effectiveness purposes. Follow up will be made to customers who don't return the Response Form. For questions contact TEG Technical Support team at 1-800-438-2834. |  | 1590 | 1590 devices | US Nationwide | Open |  |  | 1590 | Not_Computer
Z-1870-2013 | 65033 | hammer, surgical | FZY | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | S-LIFT Slap Hammer, Part Number: 11-50204, Revision AThe slap hammers are used primarily to remove trial instruments that are used to estimate sizes of intervertebral body fusion devices intended to promote fusion. These devices were used in conjunction with lumbar intervertebral body fusion device intended for intervertebral body fusion of the spine of skeletally mature patients. | SLIFT Slap Hammer | 2 | August 01, 2013 | 2013 | SpineFrontier, Inc. | Malfunction of slap hammer. | DESIGN: Device Design | The firm, SpineFrontier, sent an advisory notice entitled "URGENT April 25, 2013 S-LIFT Slap Hammer Removal" dated April 25, 2013 via email to its customers. The notice described the product, problem and action to be taken. The notice requested the removal of all distributed SLIFT Slap Hammer with listed lot number. The customers were instructed to complete and return the URGENT S-LIFT Slap Hammer Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email to customerservice@spinefrontier.com.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 3 | 3 | US distributions to states of: KS and TX. | Terminated | January 09, 2014 | 161 | 3 | Not_Computer
Z-1930-2013 | 65843 | monitor, oxygen, cutaneous, for uses other than for infant not under gas anesthesia | LPP | Cutaneous oxygen (PcO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TCM4 series Base Units, and spare part Power Supply, Radiometer Medical ApS, Akandevej 21, BrÂŽnshÂŽj, Denmark, www.radiometer.com. The devices are intended for continuous monitoring of carbon dioxide and oxygen partial pressures, oxygen saturation, and pulse rate. | TCM4 series Base Unit | 2 | August 08, 2013 | 2013 | Radiometer America Inc | RADIOMETER became aware that the TCM4 series Base Unit does not comply with IEC 60601-1 Sub Clause 15B. | DESIGN: Device Design | On 11/30/2012 the firm sent an "Urgent Field Safety Notice" to their customers. The notification explained the issue and provided recommended actions. Customers may contact 1-800-736-0600 option 4 with questions. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 5874 | 5,874 units | Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, WV & WY. Product was also shipped to Canada. | Open |  |  | 5874 | Not_Computer
Z-1946-2013 | 65337 | bed, flotation therapy, powered | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | IsoGel AIR 84, Support Surface.Intended to assist in improving human patient outcomes. | IsoGel Air | 2 | August 09, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Complaints were received from the field indicating that some Air Pumps were flashing and alarming when connected to the IsoGel AIR" surface upon turning the pump on and the setting it to the IsoGel AIR Mode. Upon further investigation, it was determined that the flashing/alarm being experienced was due to the air pressure being out of the expected range; indicating a blockage within the LAL tubi | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On 7/12/13, an URGENT MEDICAL DEVICE RECALL NOTIFICATION was mailed to consignees. The notification addressed the affected product and issue, potential hazard, actions needed including faxing (269-488-8691) or emailing (ProductFieldAction@stryker.com) the attached reply form to confirm receipt of notification, and risk mitigation. Please contact the firm at 269-389-8129 if you have questions or concerns regarding this recall. Business hours are M-F 8am-5pm(ET). | Contact the recalling firm for information | 586 | 586 units | Nationwide Distribution-USA including the states of OH, ME, IL, CO , FL, TX, MI, GA, and VA. | Open |  |  | 586 | Not_Computer
Z-1964-2013 | 65839 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MobileDiagnostwDR system. Mobile Digital radiography X-Ray SystemProduct Usage: MobileDiagnost wDR is a mobile digital radiography X-Ray System. It is intended for use by a qualified/trained doctor or technologist on both adult and pediatric patients for taking diagnostic radiographic exposures of the skull, spinal column, chest, abdomen, extremities, and other body parts. Applications can be performed with patient sitting, standing, or lying in the prone or supine positions. Not intended for mammography. | MobileDiagnost wDR | 2 | August 21, 2013 | 2013 | Philips Healthcare Inc. | The Instructions for Use (IFU) for the MobileDiagnost wDR fails to comply with a Federal standard. The IFU does not provide tolerance levels for several generator values as required by 1020.30(h)(3). | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Phillips Healthcare sent an Important Notification letter dated July 10, 2013 to all affected customer. The notification informed customers that Philips will implement an update to the Instruction for Use (IFU). Philips will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. For questions contact your local Philips representative or local Philips Healthcare office. |  | 240 | 240 systems US -- 1 system within Puerto Rico | US Nationwide Distribution including Puerto Rico | Open |  |  | 240 | Not_Computer
Z-0021-2014 | 66254 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ.Intended for general purpose diagnostic fluoroscopy and radiography studies. | GE Healthcare, Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova | 2 | October 17, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving missing screws in the overhead Video Monitor Suspension of the Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Cardiovascular X-ray imaging systems. | PRODUCTION CONTROLS: Process Control | Consignees were sent on 8/12/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 12, 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology /Cardiology /Surgery Department and Radiologists /Cardiologists /Surgeons.The letter described the Safety Issue, safety Instructions, Affected Product details, Product Correction and Contact Information (US 800-437-1171, Japan 0120 055 919). |  | 2184 | 2184 (1063 US. 1121 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and except ND, SD and VT, and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY,GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN,IRELAND, ITALY, JAMAICA, JAPAN,JORDAN, KAZKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYA ARAB JAMHIRIYA, LITHUANIA, MALAYSIA MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SLOVAKIA, SPAIN, SWEDEN, SYRIA, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, and VENEZUELA. | Open |  |  | 2184 | Not_Computer
Z-0033-2014 | 66312 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR Systems: Asset 0.5T, Apollo 0.5T, Infinion 1.5T, Eclipse/Polaris 1.5T, Panorama 0.6T 1.0T, Magnetic Resonance Imaging Systems intended for clinical/radiologic diagnostic imaging applications. | Philips MR Systems: Asset 0.5T, Apollo 0.5T, Infinion 1.5T, Eclipse/Polaris 1.5T, Panorama 0.6T | 2 | October 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There is a potential for water to collect in the vent pipe elbow in the magnet venting system. Water may freeze, blocking the venting system. | DESIGN: Device Design | Philips Healthcare sent an Urgent Medical Device Correction letter via certified mail on August 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information with all of their staff who need to be aware of the contents of this communication. Philips Healthcare will perform any necessary corrective actions prior to the customer's next cryogen refill. Customers with questions were instructed to contact their Philips field service representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 511 | Total 511 (Asset 0.5T,: 35 units Apollo 0.5T: 15 units, Infinion 1.5T: 116 units, Eclipse/Polaris 1.5T: 290 units, Panorama 0.6T: 55 units) | Worldwide Distribution - US (nationwide) including FL, CO, KY, MS, PA, NC, NJ, TX and Puerto Rico and internationally to Italy. | Open |  |  | 511 | Not_Computer
Z-0093-2014 | 65897 | dilator, catheter | GCC | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | PIPELINE LS ACCESS SYSTEM, 15 Dilator, 5 mm, InsulatedThe PIPLINE LS Dilators and Holding Clip are intended for sequential dilation of soft tissue and enabling proper placement of the Pipeline LS Retractor. | PIPELINE LS ACCESS SYSTEM | 2 | October 23, 2013 | 2013 | DePuy Spine, Inc. | Dilators and Holding Clip are being used for a secondary function of neuromonitoring, for which it is not 510(k) cleared. |  | DePuy Synthes sent an Urgent Voluntary Product Recall Notification letter dated July 30, 2013, to all affected consignees via Fed Ex overnight delivery.Consignees were informed that DePuy Spine, Inc. is voluntarily recalling all lots of the PIPELINE LS ACCESS SYSTEM Dilators (1 st, 2nd, 3rd, and 4th) and Holding Clip. They were asked to review inventory located at their facility and quarantine any products listed in this recall notice. Consignees were instructed to contact their local DePuy Synthes Spine Sales Consultant to return any inventory located within their facility.Consignees with questions were instructed to call 508-828-3583. |  | 118 | 118 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, and WY. | Open |  |  | 623 | Not_Computer
Z-0132-2014 | 65586 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution, Catalog #'s 1)1019397, 2) 1023415 (w/ installation warranty, US-only) and 3) 1023423 (w/ parts warranty, US-only).The device is designed to perform radiographic x-ray examinations on all pediatric and adult patients, in all patient treatment areas | DRX Revolution Mobile Xray System | 2 | November 01, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. is conducting a recall for the Column end Cover present on the DRX Revolution Mobile X-Ray System due to the cover dislodging from the column. | DESIGN: Device Design | The firm, Carestream Health Inc., sent a "URGENT: Medical Device Recall" letter dated June 17, 2013 via courier service to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that there is no opportunity for interference between the End Cover and potential obstructions as the column is descending. Carestream will deploy a field service representative to apply a warning label to the column, and to provide an addendum to the Safety and Regulatory Information manual. Your Carestream service representative will contact you to answer any questions and schedule a visit as soon as possible. If you have any questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910 from 8am through 8pm. |  | 331 | 331 units | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, IO, KS, LA, MA, MD, MI, MN, MO, NC, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries of: Australia, Brazil, Canada, Dubai, Germany, Japan, Kuwait, Malaysia, New Zealand, Saudi Arabia, Singapore, South Africa, Taiwan, and Thailand. | Open |  |  | 331 | Not_Computer
Z-0172-2014 | 65795 | system, x-ray, angiographic | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee systemsx-ray, angiographic system | Siemens Artis zee systems | 2 | November 14, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Artis zee systems bearing model numbers 10094135, 10094137, and 10094141 equipped with a card collimator from a specific lot where a dislocation of the scatter radiation filter inside the card collimator may occur. | DESIGN: Device Design | Siemens sent an letter to all affected customers in January 2013, notifying them of the potential issue and urging customers to make appointments with Siemens to have software update AX 027/12/S realized in order to resolve the problem. The software update was made available from January 22, 2013, on. Customers were also asked to pass along this information if the device was no longer in their possession.The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 17 unit(s) will consist of the following:1. The firm sent out a customer notification letter in January 2013 notifying them of the issue and highlighting the approved procedures of the Corrective Action Plan.2. Siemens directed in the customer notification letter, for the end-user to make an appointment with the Customer Service organization to have safety update AX 027/12/S realized.3. As stated in the safety update AX 027/12/S and in the e-mail response the update to correct this issue will be provided to the customer at no expense.4. Siemens indicated in the customer notification letter that the AX 027/12/S update will be made available from January 22, 2013 and on.5. The AX 027/12/S update provides general information, work instructions and a completion certificate on how to correct the dislocated scatter radiation filter inside the collimator.You may proceed with implementation of the CAP (subject to the conditions noted above). For further questions please call (610) 219-6300. |  | 17 | 17 | US Distribution including the states of AZ, KY, PA, IA WV, FL, MO, OK, CA, AR, MT, NY, KS, and CT. | Open |  |  | 17 | Not_Computer
Z-0438-2014 | 66724 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE Linear Accelerator Linac systems with component: any photon unflat beam option (Multiple X feature) in combination with the IMRT or mARC option. A family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. The ARTISTE/ONCOR systems are manufactured by: SIEMENS AG, MEDICAL SOLUTIONS, Roentgenstrasse 19-21, Kemnath, GERMANY 95478 and distributed SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Martinez, CA | ARTISTE Linear Accelerator | 2 | December 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | SIEMENS became aware of two (2) customer complaints concerning a safety risk regarding a potentially existing dark current radiation phenomenon on the Linear Accelerator (LINAC) in combination with IMRT or mARC (rotational IMRT) treatments using unflat (flattening filter free) beams. In case there are many long periods of field shape changes in a treatment plan, the dose due to dark current radiat | DESIGN: Device Design | Siemens sent an Urgent Medical Device Correction letter dated October 23, 2013 including the release of update instruction TH024/13/S to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. |  | 26 | 26 units | Worldwide Distribution: USA (nationwide) including states of: FL, NC, NY, PA, TN, UT, and WI; and countries of: Angola, Brazil, Canada, P.R. China, Germany, India, Japan, Republic Korea, Poland, and Thailand. | Terminated | January 22, 2014 | 50 | 29 | Not_Computer
Z-0472-2014 | 66638 | absorber, carbon-dioxide | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | Multi Absorber Original, DisposableM1173310 Multi Absorber Original, Disposable, package of 6 pcs,GE Healthcare Finland OyMade in USRx OnlyThe GE Healthcare Multi Absorber Original, Disposable is a disposable product intended for use with the GE Healthcare Advanced Breathing System (later ABS), the GE Healthcare EZchange manifold, the GE Healthcare Compact Block, and the GE Healthcare Compact Block II (later Compact Block). The Multi Absorber should only be used with air, oxygen, nitrous oxide, halothane, enflurane, isoflurane, desflurane and sevoflurane. The device is intended to be used under constant attention of qualified professional healthcare personnel. | Multi Absorber Original, Disposable | 1 | December 13, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to air leakage associated with the CO2 Multi Absorber. | DESIGN: Device Design | GE issued an Urgent Medical Device Correction letter dated September 19, 2013, to all affected customers via Fed Ex overnight mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected product, isolate all affected product and return the attached form via fax to 800-535-7923. Once the information is received, customers would be contacted with an RMA and replacement information. If product was distributed further the notice should be forwarded to those customers. For questions customers should call 1-800-345-2700, option 2 followed by ooption 2. For foreign customers 1-800-932-0760, option 2.For questions regarding this recall call 262-513-4122. GE issued a press release on 12/30/2013. |  | 444115 | US: 59,721 boxes, 358,326 canisters; OUS: 3,724 boxes, 22,344 canisters | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Costa Rica, Mexico, and Venezuela. | Open |  |  | 4445 | Not_Computer
Z-0567-2014 | 66976 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision R C-Arm Interventional Fluoroscopic X-Ray System.Mobile Fluoroscopic Imaging | Ziehm Imaging GmbH | 2 | December 24, 2013 | 2013 | Ziehm Imaging Inc | Engineering change in 2006 resulted in a minor non-conformity of their devices. Specifically the emergency switch was changed from a standard component to just an option on the devices. | DESIGN: Device Design | Ziehm Imaging sent an "URGENT RECALL NOTICE" letter dated November 22, 2013, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customers. Customers were advised and instructed to:Your Model and Serial number can be found on the product label attached to the side of the mobile C-arm Stand. Our records indicate that your facility has one or more of the products listed above that are being used in your facility.Recall of the product was initiated because according to our manufacturing documentation, the referenced device installed at your facility was not equipped with an emergency stopping device for vertical travel of the c-arm imaging assembly. To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device identification and label.The issue:According to our manufacturing documentation, the referenced device installed atyour facility was not equipped with an emergency stopping device.To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device. The manufacturer Ziehm Imaging GmbH advised customers they may continue to operate the affected device with caution until the correction has been carried out on-site. During this time Ziehm Imaging GmbH advises the operator to be aware of the following safetymeasures which may resolve the possibility of potentially hazardous situations by:" Operating the power-off switch located either on the C-arm Stand top panel or on the monitor cart" Removing the power cable from the wall outlet or," Activate the push-buttons for reverse vertical movement and keeping the pushbutton pressed until patient /operator is no longer in the risk zone or the device stops Please alert all staff members working with the device to these potential hazardous situations, and familiarize them with all possibilities to resolve the hazardous situation . |  | 103 | 103 devices | USA (Nationwide) Distribution including the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MT, MA, MD, MI, MN, MN, MO,MS, NC ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Hawaii. | Open |  |  | 386 | Not_Computer
Z-0628-2013 | 63602 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Medical Surgical Bed, Model 3002SEXMedical Surgical Bed The S3" MedSurg Bed, is intended to support and transport patients within the Med/Surg and Critical Care hospital environments. The bed is typically used in pre-op, post-op and recovery areas of hospital facilities. The intended user for this product is both Health Care Providers (HCPs: nurses, nurses aides, and medical doctors) and human patients. Lockout features may limit patient accessible controls | Medical Surgical Bed, Model 3002SEX | 2 | January 03, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical Model 3002S3EX and Model 3005S3EX beds built between March 15, 2012 and May 23, 2012. The affected beds may have been built with brake rings that do not meet specifications, negatively impacting the holding force of the brakes. | PRODUCTION CONTROLS: Process Control | Stryker sent an Urgent Medical Device Correction letterdated December 18, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate the affected product and ensure that the brakes are holding properly. Consignees were directed to contact Stryker (1-800-327-0770, option 2) to make arrangements to replace all brake rings and continue to routinely inspect brake functionality. Customers were asked to Fax (269-488-8691) the enclosed acknowledgment and Stryker field service will contact customers to schedule service. Consignees were directed to forward a copy of this notice to the new users and advise us of their new location.For questions regarding this recall call 269-488-8691. | Contact the recalling firm for information | 9 | 9 | Worldwide Distribution - Nationwide and Canada | Terminated | August 30, 2013 | 239 | 1013 | Not_Computer
Z-0621-2013 | 63794 | set, tubing and support, ventilator (w harness) | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Fisher & Paykel Healthcare reusable breathing circuit.Non-heated breathing circuit, 1.8m or 6 feet in length.Oxygen therapy delivery system for adult patient. | 900MR068 Reusable Circuits | 1 | January 03, 2013 | 2013 | Fisher & Paykel Healthcare, Ltd. | FPH has become aware that one batch of tubes supplied to FPH may contain hole damage. These tubes were incorporated by FPH into certain 900MR068 reusable circuits with lot numbers 110810 and 111020. If this hole damage is not detected during the standard leak test before patient use, it could potentially result in a gas leak in the breathing system, which may lead to a loss of pressure for the int | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Fisher & Paykel sent an Urgent Medical Device Recall letter dated November 19, 2012, with a Product Recall Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and discard any affected circuits in their possession. Customers were instructed to complete and return the Product Recall Response form to the address indicated on the form, even if they have no affected product. Customers were asked to forward the recall notice to all those who need to be aware within their organization. Customers with questions should call 1-800-792-3912.For questions regarding this recall call 949-443-4095. |  | 200 | 200 units total, 26 units US | Worldwide Distribution - USA including IL, IN, AR and GA. | Terminated | July 11, 2013 | 189 | 200 | Not_Computer
Z-0631-2013 | 63719 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA (enfilcon A) Soft (Hydrophilic) Daily Wear Toric Contact Lenses | AVAIRA Soft Contact Lenses | 2 | January 03, 2013 | 2013 | Cooper Vision Caribbean Corp. | The lenses exceeded the acceptance limit for silicone oil residual. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The recall was conducted to the wholesale/distributor and Eye Care Practitioner level. The company contacted the Authorized Distributors by telephone, followed by written notification dated November 9, 2012. All written notifications to first consignees, for both revenue and Fit Sets, were sent by overnight mail with signature required as proof of receipt and the envelopes were marked "Urgent". The notifications described the issue and provided recommended actions. Customers may contact 800-341-2020 if questions. |  | 246 | 246 | US Nationwide, including Puerto Rico, and Canada. | Open |  |  | 246 | Not_Computer
Z-0627-2013 | 63878 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Maquet Yuno Extension Table DeviceMaquet GmbH & Co. KGKehler Strabe 31Baden-Wuerttemberg76437 Rastatt,Germany Used to support the lower extremities immediately before during and after surgery. The extension device is mounted to the YUNO operating table. The product may only be operated by medically trained staff within the OR environment. | Maquet Yuno Extension Table Device | 2 | January 03, 2013 | 2013 | Maquet Inc. | There is a potential issue with the screw tension of the extension device 1433.62A1 used with the YUNO table. | DESIGN: Device Design | Maquet sent Urgent Product Field Correction letters/Field Correction Confirmation Form dated November 2, 2012 on November 15, 2012, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a MAQUET authorized Service Technicians would contact them to arrange an appointment to assess their device and to replace the release lever at no cost to them at their facility. Customers were asked to completed and return the enclosed confirmation form, acknowledging their receipt and understanding of the communication, in the enclosed postage prepaid envelope, by e-mailing a scanned copy to FieldActions@maquet.com or via fax at 973-709-6508.For questions customers were instructed to call MAQUET Technical Support at 800-777-4222, Press option2, followed by option 2.For questions regarding this recall call 973-709-7660. |  | 63 | 63 devices | Worldwide Distribution - USA including AR, GA, SC, TX, TN and internationally to France, United Arab Emirates, Australia, Brazil, Canada, Switzerland, China, Germany, Spain, United Kingdom, Italy, Japan, Sweden, Thailand, Turkey, Serbia, and South Africa. | Terminated | August 21, 2013 | 230 | 63 | Not_Computer
Z-0638-2013 | 63809 | urinalysis controls (assayed and unassayed) | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Chek-Stix Positive and Negative Control Strips for Urine Chemistry(Distributed OUS Only)Catalog Number: 1360LChek-StixÂŽ Positive and Negative Control Strips for Urine Chemistry: When used with Siemens reagent strips and tablets for urinalysis, CHEK-STIXÂŽ positive and negative control strips provide defined results for glucose, bilirubin, ketone (acetoacetic acid), specific gravity, blood, pH, protein, urobilinogin, nitrite, and leukocytes | Siemens ChekStix Positive and Negative Control Strips for Urine Chemistry | 3 | January 04, 2013 | 2013 | Siemens Healthcare Diagnostics Inc | When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂżC (77ÂżF). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens issued An Urgent Field Safety Notice via e-mail to all affected SiemensHealthcare Diagnostics Regional offices on November 21, 2012 both in the United States and foreign accounts. This notice informed customers of the issue with the CLINITEK AtlasÂŽ and Chek-StixÂŽ Positive Urinalysis Control Solutions When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂŽC (77ÂŽF). You can use the reconstituted positive control solution for up to eight hours if the solution is held at 25ÂŽC (77ÂŽF) or lower. If the operating temperature is greater than 25ÂŽC and lower than 30ÂŽC (77ÂŽF and 86ÂŽF), the solution can be used for up to two hours.For questions regarding this recall call 781-269-3158. |  | 700 | 700 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Aruba, Brazil, Denmark, India, Mexico, New Zealand, Singapore,Thailand, Bahamas, South Wales, and Hong Kong, | Open |  |  | 37839 | Not_Computer
Z-0649-2013 | 63896 | sucker, cardiotomy return, cardiopulmonary bypass | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(k) | Sterile Rigid Tip Suction WandThe Suction Wands are indicated for use to remove excess fluid from the surgical field. | CalMed SU12202 | 2 | January 07, 2013 | 2013 | California Medical Laboratories Inc | California Medical Laboratories, Inc. has identified an issue with the Rigid Tip Suction Wand. Due to insufficient adhesive application, the distal tip has the potential to come loose from the wand and fall into the patient's chest cavity. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CalMed sent an Urgent Field Safety Notice dated November 2, 2012, to all affected customers by certified mail and additionally followed up via electronic mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Sorin Group USA's Customers Service Department at +1-800-650-2623 to return any unused product and order replacements. Customers were also asked to communicate the important safety information to all potential users within their organization or any other organization where the affected product may have been transferred and to maintain awareness of this communication for an appropriate period of time to ensure effectiveness of the corrective action. Customers were asked to return all unused product to CalMed Labs. Customers with questions were instructed to contact California Medical Laboratories at +1-714-556-7365.For questions regarding this recall call 714-556-7365. |  | 226367 | S112462=995 units & S111480 1430 units | Worldwide Distribution - USA including AZ, FL, CA, LA, CO, and MS.Internationally to Canada, Israel, Greece, Hong Kong, India, Ireland, Jordan, Lebanon, Malaysia, Netherlands, Russia, Saudi Arabia, South Africa and Spain. | Terminated | March 04, 2013 | 56 | 2425 | Not_Computer
Z-0648-2013 | 63852 | radioimmunoassay, thyroid-stimulating hormone | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Immunoradiometric assay.Device is an immunoradiometric assay designed for the quantitative measurement of thyroid stimulating hormone (thyrotropin, TSH) in serum. It is intended strictly for in vitro diagnostic. Used as an aid in the assessment of thyroid status. | CoatACount TSH IRMA IKTSX with calibrator TSI3X Lot 035 | 3 | January 07, 2013 | 2013 | Siemens Healthcare Diagnostics | The recall was initiated because Siemens Healthcare Diagnostics confirmed an under-recovery of dose values with some vials of Coat-A-Consta (CAC) IRMA TSH calibrators TSI4 (level B), TSI5 (level C), TSI6 (level D), TSI8 (level F), and TSI9 (level G) lot 035. The affected calibrator vials can be identified by a darker yellow color than normal. The TSI3 (level A), TSI7 (level E), and TSIX (level H) | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated November 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, complete and return the Urgent Medical Device Recall Effectiveness Check form included with the letter by fax to 302-631-7597. Customers were also instructed to determine their replacements needed, and to forward the recall letter to whomever they may have distributed the affected product. For questions customers were instructed to contact their Siemens Technical Solutions Ceter or their local technical support representative. For questions regarding this recalla call 914-524-2955. |  | 8 | Lot 742= 8 units all foreign consignees) Lot 743=7 units (all foreign consignees) | Worldwide Distribution - USA including CA, GA, IL, IN, MI, MN, NY, and Puerto Rico.Internationally to Agentina, Bolivia, Brazil, Canada, Chile, Guatemala, South Korea, Philippines, Paraguay, Taiwan, and Uraguay, | Terminated | October 28, 2013 | 294 | 15 | Not_Computer
Z-0644-2013 | 63842 | radioimmunoassay, follicle-stimulating hormone | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Laparoscope, general; plastic surgeryUsed for general, abdominal, gynecological and thoracic minimally invasive surgical procedures to establish a path of entry or to gain access through tissue planes and/or potential spaces for endoscopic instruments. | Kii Shielded Bladed Access System | 2 | January 07, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on a specific lot number of its 12x150mm Kii shielded bladed access systems. This is due to a potential for the blade to remain exposed after shield deployment. Further use of any affected product should cease immediately.The affected model number is CTB71, and the affected lot is 1179107. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Applied Medical sent an Urgent Medical Device Recall to all affected customers in November 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete the attached Recall Notification Confirmation Form to acknowledge the recall and indcated any product their facility is returning. The Recall confirmation form should be returned to Applied Medical by emailing to recall60494844#applied medical.com or fax to 949-713-8832. Return affecte4d product and a copy of the confirmation form to Applied Medical. Customers with questions were instructed to call 949-713-8662. For questions regarding this recall 949-713-8662. | Contact the recalling firm for information | 72 | 72 units | Worldwide Distribution - USA (nationwide) and France. | Terminated | January 16, 2013 | 9 | 72 | Not_Computer
Z-0653-2013 | 63226 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | UltraFill DBM labeled in part: &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot;***TISSUENET***Sterilized by T10^6 (r)***Description: UltraFill DBM***Size: 10cc***Product Code: RT53010***Donated Human Tissue Allograft***SINGLE PATIENT USE ONLY***FOR INTERNATIONAL USE ONLY***See Package Insert for Additional Information***&amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot;TissueNet's Porcine DBM product line is used as a bone void filler. | UltraFill DBM | 2 | January 07, 2013 | 2013 | Surgical Tissue Network, Inc. | Surgical Tissue Network Inc., DBA TissueNet Inc. recalled their UltraFill DBM Putty-Porcine (1 cc; 5cc; 10cc) because products may contain trace amounts of 316L Stainless Steel particulates. | PRODUCTION CONTROLS: Process Control | Surgical Tissue Network sent a Notification of Voluntary Tissue Recall dated March 22, 2012, to all affected customers. The firm issued an additional notification letter on August 30, 2012. The letter identified the product, the problem, and action to be taken by the customer. Consignees were asked to return all listed products that remain in inventory. If the product was further distributed, they were asked to forward the recall information to their customers.Customers were instructed to contact TissueNet's Customer Service Department at 800-465-8800 x283 to coordinate return/replacement of the affected product. Customers with questions were instructed to call 800-465-8800 x230.For questions regarding this recall call 407-380-2424. |  | 160 | 160 units | Worldwide Distribution including Turkey, Greece, & Mexico. | Open |  |  | 393 | Not_Computer
Z-0656-2013 | 61718 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #2.5 and Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #3.5, Howmedica Osteonics Corp. The intended use is for reconstruction of the head and neck of the femoral joint. | Accolade TMZF Plus Femoral Stem | 2 | January 08, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker has received a report that single size 3.5 stem was packaged as size 2.5 stem resulting in a possible product mix. | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker sent "Urgent Medical Device Recall Notification" letters and product accountability forms via Fed Ex with return receipt on 10/23/12 to branches/agencies and to hospitals and physicians on 10/24/12. The letter described the issue and provided recommended actions to mitigate the risk. Customers can call (201) 972-2100 if questions. |  | 18107 | 6021-2530- four units; 6021-3535- | Nationwide distribution, including the states of NY, NJ, CA and FL. | Terminated | September 11, 2013 | 246 | 4 | Not_Computer
Z-0655-2013 | 63938 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FlowGate Balloon Guide Catheter;Manufactured by Concentric Medical, Mountain View, CA.FlowGate&amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot; Balloon Guide Catheters are coaxial-lumen, braid-reinforced, variable stiffness catheters with a radiopaque marker on the distal end and a bifurcated luer hub on the proximal end. A balloon is flush mounted on the distal end. FlowGate&amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot; Balloon Guide Catheters are indicated for use in facilitating the insertion and guidance of an intravascular catheter into a selected blood vessel in the peripheral and neurovascular systems. The balloon provides temporary vascular occlusion during these and other angiographic procedures. The Balloon Guide Catheter is also indicated for use as a conduit for Retrieval devices. | FlowGate Balloon Guide Catheter; | 2 | January 08, 2013 | 2013 | Concentric Medical Inc | Firm received complaints of resistance and possible collapse of the distal tip of the FlowGate Balloon Guide Catheters during use for aspiration. | DESIGN: Device Design | Concentric Medical sent a Urgent Medical Device Recall Notification letter dated December 19, 2012, to all affected customers. The letter identified the product the problme and the action needed to be taken by the customer.Customers were instructed to follow these instructions.1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Concentric Medical (Stryker Neurovascular) if any of the subjectdevices have been distributed to other organizations.a. Please provide contact details so that Concentric Medical (StrykerNeurovascular) can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and: Fax it to: 650-237-5230 Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords. Upon receipt of the returned customer response form, a Stryker Representative will contact you to arrange for credit of returned devices.On behalf of Stryker Neurovascular, we sincerely thank you for your help and support and regret any inconvenience and appreciate your assistance in completing this recall.Should you have any questions concerning this matter please do not hesitate to contact the undersigned. |  | 13 | 13 units | US (nationwide) including the states of CA, FL, GA, OR and TN. | Terminated | February 05, 2013 | 28 | 13 | Not_Computer
Z-0660-2013 | 63850 | prosthesis, shoulder, semi-constrained, metal/polymer cemented | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PT113950 PT HYBRID Glenoid PostRegenerex porous titanium constructSterile RProsthesis, Shoulder, Semi-Constrained, Metal/Polymer Cemented | PTHYBRID | 2 | January 09, 2013 | 2013 | Biomet, Inc. | Biomet is recalling Part Number PT-113950 PT Hybrid Glenoid Post, following an investigation which identified that the male thread of the post may be oversized. This oversized condition can vary in degree and may lead to the following three events:1) If the PT Hybrid Glenoid Post is not fully seated into the Hybrid Base and the implant construct isimplanted, then a gap of 1-3 mm will be prese | PRODUCTION CONTROLS: Process Control | Biomet Inc. sent a Urgent Medical Device Recall Notices dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notice identified the reason for the recall including directions to immediately locate, discontinue use of the device. Surgeons were asked to continue monitoring of patients implanted. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with product to be returned are asked to obtain RGA # by calling 574-372-6677 . A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 574-371-1570, Monday- Friday 8 a.m. to 5 p.m. |  | 1050 | 1050 | Worldwide Distribution -- US (nationwide) including the countries of , Canada, Japan, Argentina, Australia ,Netherlands, Europe and Chile. | Terminated | November 08, 2013 | 303 | 1050 | Not_Computer
Z-0661-2013 | 63877 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassettes, P/N 10133.The STERRAD NX sterilizations systems are low-temperature general purpose sterilizers used to sterilize heat and moisture sensitive reusable medical devices. The sterilant used is hydrogen peroxide supplied in cassette form as a separate accessory. The STERRAD NX Cassettes are only used in conjunction with the STERRAD NX Sterilization system. | STERRAD NX Cassettes | 2 | January 09, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products is recalling three lots of STERRAD NX System Cassettes because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | DESIGN: Labeling Design | Advanced Sterilization Product sent an Urgent Medical Device Recall (Removal) letter dated December 4, 2012, to all affected customers. The letter informed the customers that the Sterrad NX Cassettes do not have adequate data to support the entire duration of the product's labeled shelf-life. Customers are instructed of the problems identifed and the actions to be taken. Customers are instructed to notify their customers if they have further distributed the recalled products and to have them contact Stericycle at (877) 650-7686 to arrange for return of the product. Customers with questions are instructed to call Stericycle at (877) 650-7686 Mon-Fri from 8am to 5pm, ET.For questions regarding this recall call 949-453-6400. |  | 12280 | 12,280 units | Worldwide Distribution - USA (nationwide) and internationally to the following countries:AR, CO, TW, SI, AU, MX, MY, HU, PH, BE, FR, DE, IE, IT, ES, NO, CH, and GB. | Terminated | October 17, 2013 | 281 | 12280 | Not_Computer
Z-0657-2013 | 63355 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | (l) Visum Halogen Surgical Light, (2) Visum LED Surgical Light, (3) Stryker Flat Panel &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp; Navigation ArmThe Spring Arm Circlip is part of the Stryker Spring Arm used in multiple configurations as a component of the (1)Visum Halogen Surgical Light, (2)Visum LED Surgical Light, and (3)Stryker Flat Panel &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp; Navigation Arm. The Spring Arm that is subject to this field action suspends loads, is a conduit for power/data, and facilitates multi-directional movement. The Spring Arms that have been identified within scope are replacement parts that have potentially been serviced incorrectly. | Visum Halogen Surgical Light | 2 | January 09, 2013 | 2013 | Stryker Communications, Inc. | Stryker has become aware that there is a low likelihood of incomplete engagement of the Spring Arm Circlip during service or replacementof the Spring Arm. If the Spring Arm Circlip is not fully seated, the Spring Arm could separate and fall along with the surgical lights or monitors mounted to the end of the Spring Arm. To date, there have been no serious injuries reported from the incomplete e | PRODUCTION CONTROLS: Equipment Maintenance | Stryker sent an Urgent Medical Device Correction letter dated September 25, 2012, to all affected customers. The letter identified the product, the prolem, and the action to be taken by the customer. Customers were instructed to immediately inspect their Spring Arm. If customers experienced any concerns with their Spring Arm they were instructed to contact Stryker immediately at 1-866-841-5663 to schedule a service technician to inspect and correct the Spring Arm circlips.For questions regarding this recall call 972-410-7100. |  | 199 | United States: 1,082 units and Outside United States: 199 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Japan, Taiwan, Canada, United Kingdom, France, Netherlands, Spain, Mexico, India, Korea, Singapore, New Zealand, Hong Kong, Philippines and Brazil. | Terminated | April 22, 2013 | 103 | 1281 | Not_Computer
Z-0659-2013 | 63559 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The TM Ardis Interbody System Inserter Stainless Steel, Rubber, Silicone Non Sterile Rx only.The TM Ardis Inserter is intended for delivery of the TM Ardis Implant into the cleared disc space. The implant is secured to the inserter via finger-tightening the knob of the inserter shaft allowing the lateral grasping arms to secure the implant. Once the implant is inserted and its position confirmed radiographically, the inserter is removed from the implant by turning the shaft knob counterclockwise to loosen the lateral grasping arms from the implant. | TM Ardis Interbody System Inserter | 2 | January 09, 2013 | 2013 | Zimmer, Inc. | Zimmer Spine is has received complaints of difficulty in turning the implant release knob of the TM Ardis Inserter after final positioning of the TM Ardis interbody spacer implant. There has been two reports where the inserter instrument was inadvertently removed from patient with the implant still engaged. The correction includes revised directions explaining the technique to remove the implant | DESIGN: Device Design | Zimmer sent a "MEDICAL DEVICE CORRECTION" letter dated September 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Risk Manager Confirmation Form was attached for customers to complete and return. Questions and concerns should be addressed to Zimmer Spine Customer Service at 866-774-6368.When re-design is approved, Zimmer will remove and replace all affected devices. |  | 170 | 170 | USA nationwide; Europe, Sweden | Open |  |  | 170 | Not_Computer
Z-0658-2013 | 62950 | compressor, cardiac, external | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | AutoPulse Resuscitation System Model 100,Mfg by:Zoll CirculationSunnyvale, CA.The AutoPulse Model 100 Automatic Mechanical Chest compressor is intended to be used as an adjunct to manual CPR, on adult patients only, in cases of clinical death as defined by lack of spontaneous breathing and pulse. | AutoPulse Resuscitation System | 2 | January 09, 2013 | 2013 | Zoll Circulation, Inc. | Complaints associated with an unexpected stop in AutoPulse compressions. |  | Zoll sent Customer Notification letters dated August 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Zoll Technical Support Line at 1 (800) 348-9011 for further assistance. | Contact the recalling firm for information | 3881 | 3881 in US, 1846 ROW. | Worldwide distribution. | Open |  |  | 3881 | Not_Computer
Z-0662-2013 | 62121 | system, balloon, intra-aortic and control | DSP | Intra-aortic balloon and control system | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon PumpAn electromechanical system used to inflate and deflate intra-aortic balloons. It provide temporary support to the left ventricle via the principle of counterpulsation. | CARDIOSAVE IntraAortic Balloon Pump | 2 | January 09, 2013 | 2013 | Maquet Cardiovascular, LLC | In the Cardiosave Intra-aortic Balloon Pump, there is a possibility that the coiled cord assembly, a prominent connector cord between the Cardiosave IABP housing and the dual display head which secured atop the IABP may fail as a result of mishandling. | DESIGN: Device Design | Maquet Cardiovascular sent an Urgent Product Recall Letter dated March 27, 2012, and Urgent Product Recall Acknowledgement to all affected customers . The letter identified the product, the problem, and the action to be taken by the customer. Datascope Corp determined that the best course of action was to replace the affected product. Customers were instructed that they would be contacted by a service representative to make arrangements for the removal and replacement of their affected product. Customers with questions were instructed to contact the US Call Center at 1-800-777-4222.For questions regarding this recall call 973-709-7498. | Contact the recalling firm for information | 29 | 29 units (total) | Worldwide Distribution - USA including ID and Fl and Internationally to United Arab Emirates, France, Luxembourg, Poland, Spain, Sweden, Germany, Italy, Netherlands, Republic of China, Spain, Sweden, and Thailand | Open |  |  | 29 | Not_Computer
Z-0665-2013 | 63941 | calibrators, drug mixture | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(k) | Siemens Emit(R) tox(TM) Serum Calibrators, catalog number 7B019ULThe Emit(R) tox(TM) Serum Calibrators are intended for use with the Emit(R) tox(TM) Serum Barbiturate and Benzodiazepine Assays as a reference for the semi-quantitative analis of barbiturates and benzodiazepines in human serum. | Siemens Emit(R) tox(TM) Serum Calibrators | 3 | January 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Incorrect expiration date was assigned to the D1 lot of Emit(R) tox(TM) Serum Calibrators. The correct expiration date is 2013-12-31. The incorrect assigned expiration date is 2014-12-31. | PRODUCTION CONTROLS: Error in Labeling | Siemens sent a Customer Notification letter dated December 2012 to all affected customers of Emit(R) tox(TM) Serum Calibrators lot D1, informing them that Siemens Healthcare Diagnostics, Inc. has confirmed that the expiration of 2014-12-31 is incorrect. Siemens requested customers to discontinue use and discard any remaining inventory of the D1 lots. Customers were instructed to complete the Effectiveness Check form included with the letter even if no replacement reagents were needed and fax to 302-631-8467. Siemens would replace any unused kits of the affected lots at no charge.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 302-631-6311. |  | 1998 | 1,998 cartons | Nationwide Distribution to the following states: AL, AR, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NY, OH, PA, TN, TX and WI. | Terminated | December 24, 2013 | 348 | 1998 | Not_Computer
Z-0664-2013 | 63892 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL).Homogeneous enzyme immunoassay intended for use in the quantitative analysis of Phenobarbital in human serum or plasma. | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL) | 2 | January 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed higher imprecision with the recalled lots on Beckman Coulter AU(R) Systems and the Syva(R) Emit(R) 2000 Phenobarbital (4D019UL) lots E1, E2, and E3 as an application if the AU(R) System is used, the issue may be experienced. Firm has found that within-run imprecision of the recalled lots frequency of outliers that differ from the expected value by 20%. This is an expansion of | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an "URGENT MEDICAL DEVICE RECALL" notice letter dated November 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with their laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 1979 | 1979 cartons | Distributed in the states of CA, FL, IL, KS, MA, MI, MN, MS, NM, NY, NC, OH, OR, PA, TN, TX, UT, and WA. | Open |  |  | 1979 | Not_Computer
Z-0663-2013 | 63231 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Merit Medical's Custom Procedural Tray or Kits K12T-03162B &amp;amp; K12T-02641F.Custom Procedural Tray or Kits is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. | Merit Medical's Custom Procedural Tray | 2 | January 10, 2013 | 2013 | Merit Medical Systems, Inc. | Custom Procedural Trays/Kits contain 0.9% Sodium Chloride Injection which were recalled because one confirmed customer report where four separate particulate issues were identified in four individual flexible containers. The particular matter may result acutely in local inflammation, phlebitis and/or generalized low-level allergic response to the particulate and/or embolize to other organs in the |  | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" dated September 14, 2012 to all affected customers. The letter identified the affected product, description of problem and recommended user actions to be taken. Consignees were requested to ensure that all appropriate personnel are notified of the affected Merit trays so that they can take immediate action. They were instructed to quarantine all affected kits, remove the bag of 0.9% Sodium Chloride and return the affected product to Merit for credit or replacement. Consignees were requested to complete and sign the Customer Response Form and immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031 and mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Consignees were given product return instructions. | Contact the recalling firm for information | 422 | 422 total kits/trays | Distributed in Mississippi and New Jersey. | Terminated | January 14, 2013 | 4 | 422 | Not_Computer
Z-0670-2013 | 64007 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns high-flow aortic arch cannula, 6.5 mm (20 Fr) OD with 3/8in connector, 11in (28 cm) longThe Sarns&amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot; High Flow Aortic Arch Cannula and Sarns&amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;quot; Flexible Arterial Cannulae are indicated for use in perfusion of the ascending aorta during cardiopulmonary bypass surgery | Sarns highflow aortic arch cannula | 2 | January 11, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Flexible Arterial Cannula and High-Flow Aortic Arch Cannula connectors were identified to contain flash on the exterior of the cannula connector that exceeded established specifications. Because the flash has the potential to detach in the operative field, TCVS has decided to remove all affected lots that were manufactured with the affected connector part. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | TERUMO sent an URGENT MEDICAL DEVICE RECALL notification to all consignees on January 7, 2013. Terumo CVS is advising users to discontinue use and return all affected cannulae in inventory. Terumo CVS will issue credit for returned cannulae. Terumo CVS providied a list of marketed devices that can be used as alternatives. Contact Terumo CVS Customer Service with questions or concerns: 800- 521- 2818 Monday  Friday, 8 AM - 6 PM ESTFor questions regarding this recall call 734-741-6173. | Contact the recalling firm for information | 6830 | 6830 total all products | Nationwide and worldwide: AL , AR , AZ , CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , ME , MI, MO , NC , ND , NV , NY , OH , OK , PA , SD , TN , TX , UT , VA , WA , and WI. Bahrain, BRAZIL, CANADA, CHILE, Honduras, Singapore, Taiwan, Trinidad and Tobago, and UNITED ARAB EMIRATES (UAE). | Terminated | August 26, 2013 | 227 | 6830 | Not_Computer
Z-0685-2013 | 62719 | wrench | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Spine EXPEDIUM Offset Torque WrenchProduct Code 2797-29-510The device is intended for use during the final tightening of an Expedium Offset Spinal Construct. | DePuy Spine EXPEDIUM Offset Torque Wrench | 2 | January 16, 2013 | 2013 | DePuy Spine, Inc. | A component within the handle of the Expedium Offset Torque Wrench is prematurely wearing, potentially resulting in the instrument recording torque values outside of the specification. |  | DePuy Spine issued an Urgent - Voluntary Product Recall Notification dated July 24, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue the use of the affected product and to remove the EXPEDIUMÂŽ Offset Torque Wrenches from service and return to DePuy Synthes Spine until direction can be provided by DePuy Synthes Spine regarding a replacement torque limiting device.For questions regarding this recall call 508-828-3769. |  | 140 | 140 units | Worldwide Distribution - USA (nationwide) and internationally to Belgium, Egypt, Ireland, Italy, Portugal, and the UK. | Terminated | December 30, 2013 | 348 | 140 | Not_Computer
Z-0690-2013 | 63771 | [6 empty] | Microplate, truncated V-bottomed lots included in theabove listed products Product Number: 1380 5465.DBS Microplate is a general purpose 96 well microplate used in the preparation and examination of specimens from the human body. | Perkin Elmer Microplate | 2 | January 16, 2013 | 2013 | Perkinelmer | Truncated V-Bottomed Microplate Wells used in assays contain defective wells. | PRODUCTION CONTROLS: Process Control | Perkin Elmer initiated the recall on November 23, 2012. The Recall Letter will be delivered to the end customer by email facilitated through our PerkinElmer Technical Support group located in Akron Ohio. Technical Support will be asked to make sure the customers receive the Recall Letter and implement the recommended precautions:- Inspect the wells of row A and row H of the affected microplate lots for possible holes and cracks prior to use.- If a hole or crack is visually detected in a well of the microplate, the well should not be used for testing.- If a hole or crack is not visually detected in a well of the plate, it is required during use of the well, the volume of solution in the well should be visually confirmed to be consistent before and after incubation.Please contact PerkinEimer Technical Support for further information (800) 321 -9632, select prompt 2. |  |  |  | Nationwide Distribution. | Open |  |  |  | Not_Computer
Z-0676-2013 | 63768 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius GranuFlo Dry Acid Concentrate For Bicarbonate Dialysis 2.0 K 2.5 Ca 1 Mg 16.5 GALCatalog Number: OFD2251-3BProduct Usage:GranuFlo is indicated in the treatment of acute and chronic renal failure during the hemodialysis treatment. | GranuFlo Dry Acid Concentrate | 3 | January 17, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | GranuFlo Dry Acid Concentrate may show discoloration |  | Fresenius Medical Care North America notified consignees by telephone on November 26, 2012 and follow-up letter on November 30, 2012 of the product recall. On January 10, 2013, the firm expanded there recall to an additional 32 lots (34 lots total). Customers were instructed to examine their inventory immediately, discontinue use, place all units in quarantine and return to Fresenius Medical Care North America. Customers were advised to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. For questions contact the GranuFio Recall Line at 1-888-803-0509. |  | 101108 | 101,108 units | US Nationwide Distribution | Open |  |  | 158107 | Not_Computer
Z-0691-2013 | 64039 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Q-Plus T Powered Laser Surgical InstrumentNd:YAG 1064 nm and 532 nm: Treatment of vascular lesions and pigmented lesions; Hair and tattoo removal; The incision, excision, ablation, vaporization of soft tissue for general dermatology. Ruby 694nm: Cutting, vaporization and ablation of soft tissue; Tattoo removal; Treatment of benign pigmented lesions. | QPlus T Powered Laser Surgical Instrument | 2 | January 17, 2013 | 2013 | Quanta System, S.p.A. | Quanta System S.p.A, Italy, is notifying U.S. consignees that their website contained promotional materials for products distributed in the U.S. which did not specify that some of the claims made therein were not meant to apply to the U.S. distributed devices. | MISBRANDING: Labeling False and Misleading | Quanta System sent a letter dated November 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following: quarantine and destroy all promotional literature printed from the manufacturer, website or obtained from the manufacturer. Provide any promotional literature independently produced by the initial importer or distributor containing information about Quanta System products. Provide URL of initial importers or distributors website if information about Quanta System products is contained therein. Provide a written certification that all actions have been taken. Quanta System has or will: review all promotional materials and websiteinformation provided by initial importers and distributors; advise initial importers and distributors of any actions that need to be taken to change promotional materials and/or website information.For questions regarding this recall call 303-223-4336. | Contact the recalling firm for information | 75 | 75 units | Class II Recall - Worldwide Distribution - USA including TX, AZ, OR, and CO.Only US product is affected by this recall. | Terminated | January 23, 2013 | 6 | 109 | Not_Computer
Z-0672-2013 | 63698 | cannula, ophthalmic | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Bausch and Lomb 27G sterile cannula packed in the Bausch and Lomb Amvisc 1.2% Sodium Hyaluronate ( Model 59051, 59081, 59051L, 59081L) and Amvisc Plus 1.6% Sodium Hyaluronate (Model 60081, 60051, 60051L, 60081L) ophthalmic viscosurgical device.The 27G disposable cannula is attached to the syringe standard luer fitting to deliver the Amvis or Amvis Plus during use as a surgical aid in opthalmic anterior or posterior segment procedures. | Bausch & Lomb Disposable 27G cannula | 1 | January 17, 2013 | 2013 | Bausch and Lomb, Incorporated | Some disposable cannulas (with 4-digit lot numbers) provided with the Amvisc or Amvisc Plus OVD are not properly engaging to the Luer-Lock on the Amvisc and Amvisc Plus sterile glass syringes. The cannulas may leak viscoelastic material or detach from the syringe during injection. In rare incidents, detachment has resulted in serious patient injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Bausch & Lomb, initially notified their customers via telephone November 14, 2012 and sent a follow-up letter entitled "URGENT VOLUNTARY PRODUCT RECALL AND FIELD CORRECTION" on November 15, 2012 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine the product; identify the affected cannulas; and complete and return the attached Recall and Field Correction Acknowledgment Form via Fax to (800) 822-8918 or email at: Luis.Leyva@bausch.com. Note: Based on your form selections, replacement cannulas will either be expedited directly to your facility or your Bausch & Lomb Sales Representative will contact you shortly to furnish replacements and properly dispose of the cannulas with 4-digit lot numbers in your product inventory.Please feel free to contact Customer Service with any questions or concerns at (800) 338-2020 (option 2, followed again by Option 2). |  | 2515760 | 2,515,760 | Worldwide distribution: USA including DC and PR; and countries of: Canada and Mexico accounts not available at this time. | Terminated | December 04, 2013 | 321 | 2515760 | Not_Computer
Z-0695-2013 | 63863 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Wire Bolt, Catalog number 4933-1-001 through -003. | Hoffman LRF Wire Bolt | 2 | January 17, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Two complaints have been filed where Hoffman LRF Wires have broken during load bearing application by patients. | DESIGN: Device Design | Stryker sent "Urgent Medical Device Removal" notification letters and acknowledgement forms dated 10/25/2012 via Fed Ex to Branches/Agencies with return receipt. Stryker sent notification letters/acknowledgement forms dated 10/26/202 via Fed Ex to Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letters described the device issue and provided recommended actions. Questions were directed to Stryker Customer Service. | Contact the recalling firm for information | 476 | 476 units | Nationwide Distribution, including the states of PA, OH, NY, CA and IL. | Open |  |  | 476 | Not_Computer
Z-0673-2013 | 63319 | retractor, self-retaining, for neurosurgery | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Vycor Medical ViewSite Brain Access System 17mm/ 11mm / 5cm Device***STERILE***Product Usage:The Vycor Viewsite Brain Access System (VBAS) has been designed to serve as a self-retaining retractor system for brain tissue. It is comprised of an introducer and a working channel port that are secured by a spring-loaded latch. Each device allows for the gentle retraction of tissue, visualization of the surgical site and for the smooth manipulation of introduced instruments. The range of device sizes (width and length) provides various working channel sizes. | Vycor Medical ViewSite Brain Access System | 1 | January 17, 2013 | 2013 | Vycor Medical, Inc. | Vycor Medical recalled their VBAS (Vycor Viewsite Brain Access System) because an unidentified black fiber was found on the device. | PRODUCTION CONTROLS: Equipment Maintenance | Vycor Medical, Inc. notified all customers of the product recall by telephone and follow-up letter. Customers were instructed to quarantine affected products and hold the shipment(s) in quarantine unti further notice. The units will be replaced with inspected and passed units. All distributors were instructed to notify their customers of the recall and ask to have them return their stock. For question call 561-558-2020. |  | 3000 | 3000 devices | Worldwide Distribution - USA (Nationwide) in the states of TX, IL, NJ, FL, OH, PA and (International) in the country of Japan. | Terminated | May 13, 2013 | 116 | 3000 | Not_Computer
Z-0697-2013 | 63840 | system, test, vitamin d | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 25-Hydroxy Vitamin D EIA, Catalog #AC-57F1.Product Usage: The IDS 25-Hydroxy Vitamin D EIA kit is an enzymeimmunoassay intended for the quantitative determination of 25-hydroxyvitamin D (25-OH D) and other hydroxylated metabolites in human serum or plasma. Results are to be used in conjunction with other clinical and laboratory data to assist the clinician in the assessment of vitamin D sufficiency in adult populations. | IDS | 2 | January 18, 2013 | 2013 | IDS (Immunodiagnostic Systems Ltd.) | Immunodiagnostic Systems is recalling the 25-Hydroxy Vitamin D EIA, AC-57F1 because the QC report was found to have an erroneous calibrator value assigned. The value of 24.4 ng/mL for calibrator 4 was stated on the QC report, however the correct value should have read 28.4 ng/mL. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | IDS sent a recall notification letter on November 7, 2012 to all customers that received the the 25-Hydroxy Vitamin D EIA (AC-57F1). The letter informed the customers of the problem identified and the action to be taken.Customers are instructed to complete the attached Recall and Acknowledgement Form and fax back to Immunodiagnostic Systems Inc. Customers with questions are instructed to contact their local IDS office. | Contact the recalling firm for information | 675 | 675 units | US Nationwide Distribution | Terminated | May 31, 2013 | 133 | 675 | Not_Computer
Z-0699-2013 | 63958 | kit, identification, enterobacteriaceae | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BBL Crystal Enteric/Nonfermenter ID Kit, Catalog # 24500, cartons of 20 kits, In-vitro diagnostic reagents for microbiological testing.The BBL Crystal Enteric/Nonfermentor (E/NF) identification (ID) System is for the identification of aerobic gram-negative bacteria that belong to the family Enterobacteriaceae as well as some of the more frequently isolated glucose fermenting and nonfermenting gram-negative bacilli. | BBL Crystal Enteric/Nonfermenter ID Kit | 2 | January 20, 2013 | 2013 | Becton Dickinson & Co. | In-vitro diagnostic reagents for microbiological testing may be reversed in well positions in the test kit and when utilized may lead to bacteria misidentification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Becton Dickinson initiated the recall by letter to direct accounts on 12/13/2012. Ex-US BD sites were notified via email, distributors were notified by fax and customer letters were delivered by UPS ground. The notification advised that the recalled kit contained citrate (CIT} and malonic acid (MLO} reagents in the reversed prong/well positions. Customers were advised to discontinue use of identified lot numbers and discard any remaining packages for replacements, and also to return the enclosed reply form. The recalling firm requested distributors to submit their customer list for direct BD notification to all end users. Contact the BD Customer Service Department at 1-800-638-8663 if further assistance is required. |  | 67820 | 67,820 units | Worldwide Distribution-USA including the states of MA, TN, KY, IN, GA, NC, TX, NY, FL, AZ, WA, NJ, MI, and CA, and the countries of Uruguay, Taiwan, Singapore, Australia, South Korea, Mexico, Hong Kong, Guatemala, Colombia, Chile, Canada, India, Japan, Brazil, China and Belgium. | Terminated | June 03, 2013 | 134 | 67820 | Not_Computer
Z-0701-2013 | 63949 | nucleic acid amplification assay system, group b streptococcus, direct specimen test | NJR | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BD MAX (tm) (6 channel) Instruments, Catalog number 441916, clinical diagnostic instrument. | BD MAX (tm) (6 channel) Instruments | 2 | January 22, 2013 | 2013 | Becton Dickinson & Co. | Clinical diagnostic instrument may exhibit minor defects, which could lead to erroneous test results. | DESIGN: Device Design | Becton Dickinson sent an Urgent Field Corrective Action letter September 18, 2012,by fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that BD Field Service Associates will contact them within 7-10 days of the notification to determine if their instrument is impacted. Customers were requested to return a response form indicating the serial number on their instrument. If customers needed further assistance they were instrycted ti contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4054. |  | 86 | 86 Instruments | USA including CT, FL, IL, KS, MD, MI, MO, NY, PA, RI, TX and WI. Internationally to Canada, Europe and Japan. | Terminated | February 04, 2013 | 13 | 86 | Not_Computer
Z-0698-2013 | 63912 | system, test, vitamin d | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Laboratories, ARCHITECT 25-OH Vitamin D Reagent Kit, 03L52-25 (ROW) and 03L52-27 (U.S.) (1 x 100 tests), 03L52-35 (ROW) and 03L52-37 (U.S.) (5 x 100 tests); in vitro diagnostic. | ARCHITECT 25OH Vitamin D Reagent Kit | 2 | January 23, 2013 | 2013 | Abbott Laboratories | The Architect 25-OH Vitamin D assay has the potential to generate falsely elevated or falsely depressed results when used with certain lots of Architect Reaction Vessels. |  | Abbott Laboratories sent a "Product Correction Immediate Action Required" letter dated December 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to immediately discontinue use and destroy any remaining inventory of the reaction vessel lots listed; to clear the ARCHITECT RV hopper of impacted RV lots following the instructions provided, prior to continued use of the ARCHITECT 25-OH Vitamin D assay; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 For further questions please call your local Customer Service representative at 877-422-2688. |  | 28600 | 28,600 total | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Barbados, Belgium, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Ireland, Italy, Japan, Latvia, Libya, Malaysia, Netherlands, New Zealand, Poland, Portugal, Russia, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 28600 | Not_Computer
Z-0708-2013 | 63665 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | DePuy Spine VIPER POLY SCREW 5.5 9 X70MM TlProduct Code: 1867-15-970 Spinal Implant Component. | DePuy Synthes | 2 | January 24, 2013 | 2013 | DePuy Spine, Inc. | Polyaxial screws 8x 80 and larger may disassociate when inserting in the Ilium., VIPER large diameter long screws in high-torque circumstances, the screw shank has the potential to separate from the polyaxial head | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | DePuy Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated November 7, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine the material and return remaining inventory. Customers were advised to complete the enclosed Business Reply Form and fax back to 508-828-3762 or email to lregan3@its.jnj.com. For questions call Linda Regan @508-977-6606. |  | 24 | 24 units | Worldwide Distribution - US (nationwide) including the states of: AZ, CA, FL, HI, IL, MD,MI, MO,NC, TN, TX and the country of Australia. | Terminated | December 24, 2013 | 334 | 169 | Not_Computer
Z-0717-2013 | 63873 | condom | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lifestyles Pleasure Collection 30 Premium Lubricated CondomsAssortment of Ultra Sensitive, Flavors and Colors, Skyn, Thryll, Thyn, and Ultra Thin. UPC 0-70907-02625-0. | LifeStyles Pleasure Collection Condoms | 2 | January 24, 2013 | 2013 | Ansell Healthcare Products LLC | One lot of Lifestyles Pleasure Collection Lubricated Condoms is misbranded. Ultra Thin was included rather than Skyn. | PRODUCTION CONTROLS: Packaging | Ansell Healthcare Products LLC sent an "URGENT: VOLUNTARY CONDOM RECALL" letter/return response form dated November 5, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm's Customer Service Department at 1-800-305-1054 for questions or if additional information is needed. | Contact the recalling firm for information | 445 | 445 cases (10680 consumer units) | Nationwide Distribution. | Terminated | July 31, 2013 | 188 | 445 | Not_Computer
Z-0716-2013 | 64043 | immunoassay method, troponin subunit | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care i-Stat cTnI cartridge.The i-Stat cardiac troponin I (cTnI) test is an in vitro diagnostic test for the quantitive measurement of cardiac troponin I (cTnI) in whole blood or plasma. Measurements of cardiac troponin I are used in the diagnosis and treatment of myocardial infarction and as an aid in the risk stratification of patients with acute coronary syndromes with respect to the relative risk of mortality. | Abbott Point of Care iStat cTnI cartridge | 2 | January 24, 2013 | 2013 | Abbott Point of Care Canada Limited | APOC has determined that 10 boxes of i-Stat cTnI cartridges from lot P12179 have an incorrect barcode applied to the cartridge portion pack. | PRODUCTION CONTROLS: Error in Labeling | APOC sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Abbott Point Of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point Of Care representative for questions regarding this information. | Contact the recalling firm for information | 560 | 560 boxes (25 cartridges in each box) | Worldwide Distribution. | Terminated | June 14, 2013 | 141 | 560 | Not_Computer
Z-0721-2013 | 63981 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Goldvac(TM) UltraClean(R) Accessory Electrode 6 inch Coated Blade with Extended Insulation, REF/Catalog No. 60-7523-001, Sterile, Rx ONLY -- CONMED CORPORATION, Electrosurgical, Cutting and Coagulation and Accessories. | Goldvac(TM) UltraClean(R) Accessory Electrode 6 inch Coated Blade with Extended Insulation, | 2 | January 24, 2013 | 2013 | ConMed Corporation | ConMed received complaints of some devices which had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. In no instance has it been reported to ConMed that a compromise in the sterile barrier has resulted in illness or injury. The compromised seals were discovered prior to use. | PRODUCTION CONTROLS: Equipment Maintenance | ConMed sent an Urgent Field Safety Notice dated January 7, 2013, and Business Reply Form to all affected customers via USPS Priority Mail. International notices were sent via FedEx International Priority on December 27, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II and return it with unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502Customer were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. | Contact the recalling firm for information | 313190 | 313,190 (total for all devices) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Belgium, Colombia, Czech Republic, Dubai, Guatemala, Hong Kong, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Portugal, Saudi Arabia, Thailand, and Uruguay. | Open |  |  | 313190 | Not_Computer
Z-0714-2013 | 63959 | needle, fistula | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems 15 Gauge x 3/5 inch Buttonhole Needle SetProduct Code: BH-2035PEProduct Usage:Medisystems Buttonhole Needle Sets are indicated for use as an access device for dialysis and pheresis procedures using a constant site orbuttonhole method of needle insertion. | Buttonhole Needle Set | 2 | January 24, 2013 | 2013 | Medisystems a NX Stage Company | A total of 5 (ea) single needle pouches are mislabeled as 15 Ga x1" instead of 15 Ga x 3/5". | PRODUCTION CONTROLS: Labeling Mix-Ups | Medisystems sent a notification letter dated December 20, 2012 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to immediately open and inspect inventory for the affected products. Results of the inspection are to be provided to Medisystems and the product is to be returned to Medisystems. For questions call (978) 332-5973. |  | 26 | 26 casesx 250 needles (6500needles) | US Nationwide Distribution in the states of NY, PA and TN. | Terminated | October 10, 2013 | 259 | 42 | Not_Computer
Z-0740-2013 | 63827 | device, hemostasis, vascular | MGB |  |  | Cardiovascular | PMA | Cordis, a Johnson and Johnson Company, ExoSeal(TM) 6F Vascular Closure Device, cat. No EX500, EX600, and EX700. Used for femoral artery puncture site closure, reducing times to hemostasis and ambulation in patients who have undergone diagnostic or interventional procedures. | Cordis ExoSeal Vascular Closure Device | 2 | January 25, 2013 | 2013 | Cordis Corporation | Affected lots of Cordis ExoSeal (TM) VCD have been sterilized using a process that did not achieve the sterility assurance level that is required by the standards of this class of product. Model Numbers EX500, EX600, and EX700. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Cordis Corporation sent an "Urgent Medical Device Recall (Removal ) letter dated October 29, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately identify and set aside all product listed in a manner that ensures the affected product will not be used. A product recall response form was attached for customers to review, complete, sign and return the enclosed acknowledgement form directly to Cordis.For further questions please call (786) 313-2000. |  | 98675 | 98,675 units distributed in US. | US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Terminated | November 06, 2013 | 285 | 98675 | Not_Computer
Z-0738-2013 | 64054 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT SA Blood Culture Bottle.The BacT/ALERT SA reagent bottle is intended to be recover and detect aerobic microorganisms (bacteria and fungi) from blood and other normally sterile body fluids. | BacT/ALERT SA Blood Culture Bottle | 2 | January 25, 2013 | 2013 | bioMerieux, Inc. | BacT/ALERT SA Culture Bottle lot 1031866 may contain bottles with uncured sensors. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" and Acknowledgement Form on January 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact your local bioMerieux Customer Service Representative for additional assistance. |  | 1690 | 1690 cases | Worldwide Distribution.USA including the states of AR, AZ, CA, CO, DE, FL, GA, IL, IN, KY, LA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WI, and WV, and the countries of China, Curacao and Japan. | Open |  |  | 1690 | Not_Computer
Z-0739-2013 | 64008 | culture media, general nutrient broth | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(TM) Fildes Enrichment, catalog number 211866, 10 tubes/pack, labeled in part ***Becton Dickinson and Company, Sparks, MD 21152 USA***Intended Use: Fildes Enrichment may be used to enrich a variety of media for the cultivation of various microorganisms. Fildes Enrichment is a peptic digest of sheep blood used to enhance the growth of fastidious organisms. It may be added to Tryptic/Trypticase Soy Broth or Agar in final concentration of 5% for cultivation of Haemophilus influenzae. Fildes Enrichment is intended to be a rich source of growth factors stimulatory to various microorganisms, including the X (hemin) and V (nicotinamide adenine dinucleotide, NAD) factors necessary for the growth of Haemophilus influenzae. Labeling: For Laboratory Use. | BBL(TM) Fildes Enrichment | 2 | January 25, 2013 | 2013 | Becton Dickinson & Co. | Enrichment media intended to enhance the cultivation of various microorganisms may fail to support the growth of Haemophilus species. |  | Becton Dickinson sent an Urgent Product Recall letter dated December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to discontinue use and discard any remaining product for replacement. Customers were also advised that due to an anticipated extended delay for the availability of Fildes Enrichment, BD will substitute DIFCO" Supplement B with Reconstituting Fluid as the replacement product. Customers were advised to ensure that this replacement product is appropriately validated in accordance with their laboratory's procedures. Customers with questions should contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4000. |  | 11710 | 11,710 tubes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Canada, South Korea, Taiwan, Hong Kong, Japan, Mexico, and Singapore. | Terminated | July 01, 2013 | 157 | 11710 | Not_Computer
Z-0744-2013 | 64135 | film, radiographic | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | KODAK INSIGHT DENTAL FILM in SureSoft Packets, Size 2 Single Packet Film, Model Number IP21S, REF/Catalog Number 8118226 --- COMMON/USUAL NAME: Dental Intraoral Radiographic Film -- 20 cartons of 80 packets.Product Usage:Dental intraoral radiography | KODAK | 3 | January 28, 2013 | 2013 | Carestream Health, Inc. | Missing symbols and manufacturing information on the clear carton label. The expiration dating symbol is missing but the expiration date is not missing. The Lot symbol is missing but the lot number is not missing. The Manufacturer symbol, name and address including country of origin are missing. All of the remaining symbols and information are on the case label. Missing information is only a | DESIGN: Labeling Design | Carestream sent an Urgent Customer Notification and Field Corrective Action Letters dated January 17, 2013 and response form to the domestic consignees via UPS Overnight Mail. The letter identified the affected product, problem and actions to be taken. In addition, the consignees were instructed to complete the response form and fax it back to Carestream to fax # 585-323-7643. For questions or concerns customers were instructed to contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. | Contact the recalling firm for information | 101 | 101 cartons: Domestic - 95 cartons; Foreign - 6 cartons | US Nationwide Distribution including the states of California, Florida, Texas, Illinois, Indiana, Maryland, Pennsylvania and Canada. | Open |  |  | 101 | Not_Computer
Z-0743-2013 | 63866 | phosphomolybdate (colorimetric), inorganic phosphorus | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Alfa Wassermann System Diluent 750 mLManufactured for Alfa Wassermann Diagnostic Technologies LLC 4 Henderson Drive West Caldwell, NJ 07006Alfa Wassermann B.V. Pompmolenlaan 24 3447 GK, Woerden The NetherlandsCommon name is water solution. Intended to dilute reagents and rinse the parts of a clinical chemistry instrument. | Alfa Wassermann System Diluent 750 mL | 2 | January 28, 2013 | 2013 | Alfa Wassermann, Inc. | Inorganic phosphorus assays using System Diluent Lot 155599 may not product acceptable calibration results. | DESIGN: Device Design | Alfa Wassermann sent an important advisory notice/product response form to their distributors on October 10, 2012. The letter identified the product, the problem, and the action to be taken by the customer. The letter informs customers about the affected lot and provide instruction regarding the receipt of replacement product. Any recalled product should be destroyed. If there is further distribution, those customers should provide Alfa with a distribution list so that they may be notified. Any questions can be directed to (866) 419 - 2532. |  | 4104 | 4,104 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada and Cayman Islands | Terminated | October 25, 2013 | 270 | 4104 | Not_Computer
Z-0745-2013 | 64081 | calibrator, primary | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) IRON Calibrator (DC85), Siemens material number 10445010.The Dimension(R) IRON Calibrator is an in vitro diagnostic product intended to be used to calibrate the Iron method for the Dimension(R) clinical chemistry system. | Siemens Dimension(R) IRON Calibrator (DC85) | 3 | January 29, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints of IRON calibrations outside of acceptable calibration slope (m) values (0.97 - 1.03) when using IRON calibrator lots 2DN060 and 2FN063. Lots 2GN068 and 2JN074 may also be impacted but no complaints have been received. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent a "CUSTOMER NOTIFICATION" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Siemens Technical Solutions Center at 1-800-441-9250 or your local Siemens Healthcare Diagnostics representative for assistance with this issue. |  | 5999 | 5999 | Nationwide Distribution. | Terminated | January 08, 2014 | 344 | 5999 | Not_Computer
Z-0746-2013 | 63918 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Cholestech LDX Multianalyte Control, 2 vial set, Catalog Number: 88769Alere Cholestech LDX Multianalyte Control (hereafter referred to as Controls) enables users to monitor the performance of total cholesterol (TC), high density lipoprotein cholesterol (HDL), triglycerides (TRG), glucose (GLU), alanine aminotransferase (ALT), and aspartate aminotransferase (AST) test procedures on the Alere Cholestech LDX System. The results obtained for the controls are to be compared with the assigned values given on the Expected Values Card, accompanying the package insert, to determine if the procedure is within control limits. Assayed quality control material for use with the Alere Cholestech LDX System. | Alere Cholestech LDX Multianalyte Control | 3 | January 29, 2013 | 2013 | Alere San Diego, Inc. | Out of range (increased) HDL Cholesterol results. | DESIGN: Component Design/Selection | A recall letter was sent on 10/22/12 to all customers who purchased the Alere Cholestech LDX Mutianalyte Control, 2 vial set, Catalog Number 88769, Lot #C2510 & Lot #C2510A. The letter informed the customers that these lots of controls may yield results which are out of range high for the Level 2 control when tested with the Alere Cholestech LDX HDL test. Customers are instructed of the actions to be taken.On 11/16/12, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 vial set, Catalog 88770, Lot#271097. The problem identified is that these lots of calibration verification materials may yield results which are out of range high when tested with the Alere Cholestech LDX HDL test. Customers are instruted of the actions to be taken.On 3/6/13, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 Vial set, Catalog # 88770, Lot#298148 and 301886. The recall expanded to include the Alere Cholestech LDX Multianalyte Control, 2 vial Set, Catalog # 88769, Lots#C2572, C2572A, C2593 and C2593A.Customers are instructed to complete and fax the enclosed verification form within 10 days.Customers with questions are instructed to contact Alere San Diego, 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8289, fax (858) 805-8457, and responses.ts@alere.com. | Contact the recalling firm for information | 11166 | 11,166 kits | Worldwide Distribution - USA (nationwide) Puerto Rico and South Korea. | Terminated | September 26, 2013 | 240 | 11780 | Not_Computer
Z-0751-2013 | 64129 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 45ÂŽ Angle Driver / Z-H, Greatbatch Medical 10 000 Wahrie Drive, Clarance, NY 14031.Model #: P2661, P2662, T3635, T5586,T6031, T6164,T10161, T12628,T14879, T16422,T2487,T2713,T116057, T13635, MPU260045AO01, MPU260045ZH01.The Driver is intended to act as an interface between the cutting tool and a motorized driving system (i.e. surgical drill). | 45 Angle Driver | 2 | January 30, 2013 | 2013 | Greatbatch Medical | Greatbatch Medical has initiated a global voluntary recall and market withdrawal of its entire Small Angled Driver 45Âż product line. Greatbatch has received reports of noisy and / or rough operation when using these Small Angled Drivers. | PRODUCTION CONTROLS: Error in Labeling | Greatbatch sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 15, 2012 to all affected customers. The letter described the Reason For Recall, Risk to Health, and Instructions to Customers. Advised consignees to examine their inventory and quarantine the product in question. If they have further distributed the product to notify their customers and to complete and return the "Recall Response Form". Contact the firm at 1-763-951-8235 for questions relating to this notice. |  | 913 | 913 | Worldwide Distribution-USA including the states of CA, IN, KS, MA, NJ, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CANADA, DENMARK, FRANCE, GERMANY, ITALY, JAPAN, PORTUGAL, and SWITZERLAND. | Open |  |  | 913 | Not_Computer
Z-0755-2013 | 63934 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ortho Clinical Diagnostics VITROSÂŽ Chemistry Products Vapor Adsorption Cartridge REF/CAT No. 680 0100. VITROSÂŽ Vapor Adsorption Cartridges are used to remove vapor contaminants that may potentially interfere with Signal Reagent used with MicroWell Assays on the VITROSÂŽ 3600, VITROSÂŽ 5600, and VITROSÂŽ ECi/ECiQ Systems. | VITROS Chemistry Products Vapor Adsorption Cartridge | 2 | January 30, 2013 | 2013 | Ortho-Clinical Diagnostics | Some VITROSÂż Chemistry Products Vapor Adsorption Cartridge CAT No. 6800100 cartons incorrectly contain a Particulate Cartridge Part No. J19612. In some instances, the Particulate Cartridge has been inadvertently installed on the VITROSÂż System instead of the intended VITROSÂż Vapor Adsorption Cartridge which may bias results. | PRODUCTION CONTROLS: Packaging Process Control | Ortho-Clinical Diagnostics sent URGENT PRODUCT CORRECTION Letters (dated 12/13/12) along with Q&A; and Confirmation of Receipt Forms to distributors and consignees on 12/13/12. They were advised of the issue and instructed to discard their current inventory of VITROSÂŽ Vapor Adsorption Cartridges. On 12/13/12, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 5237 | Domestic: 15,891; Foreign: 5237 | Worldwide Distribution, including Nationwide (USA) and the countries of AU, GM, FR, IT, SP, BR, BM, AR, CA, JA, MX, UK, CH, CL, CO, IN, EC, PE, PY, PA, PT, NZ, SG, VE and UY. | Open |  |  | 5237 | Not_Computer
Z-0752-2013 | 63805 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Extremity Medical, IO FiX System Accessory. Catalog Number 101-00010. This screw measuring device is an accessory to the IO FiX Screw and Washer System that is used to measure the length of the screw by placing the depth gauge over the guidewire and down to the bone. Orthopedic device. | Extremity Medical LLC | 2 | January 30, 2013 | 2013 | Extremity Medical LLC | Lag screw measurement is short (approximately 5 mm) when using current screw measuring device (p/n 101-00010) per the IO FIX Surgical Technique Rev C. which can lead to possible inappropriate implant size selection (e.g. shorter screw size selected). | DESIGN: Component Design/Selection | Extremity Medical LLC sent e-mail notification dated March 22, 2012, to all affected customers. The email identified the product the problem and the action to be taken by the customers.The email includes a buleeting which informes the customer that an upgrade to the existing instrumentation is being released. The upgrade will replace the existing gauge. Updated pdf versions of the technique guide are also included. Customers are instructed to return the recalled device after swapping out the recalled device with the replacement unit. Customers will be contacted by Extremity and Sales Reps should contact their principal.For further questions please call (973) 588-8980. | Contact the recalling firm for information | 93 | 93 units | Worldwide Distribution -- Nationwide Distribution including the states of AR, AZ, CA, FL, GA, IL, IN, MD, MT, NC, NE, NH, NJ, NY, OH, PA, TN, TX, WI and WA. | Terminated | June 13, 2013 | 134 | 93 | Not_Computer
Z-0753-2013 | 64004 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TUNNELOC ROUND HEAD INTERFERENCE SCREW8MMX25MMBiomet Sports Medicine manufactures a variety of internal fixation devices intended to aid in arthroscopic and orthopedic reconstructive procedures requiring soft tissue fixation, due to injury or degenerative disease. Implants used for this application include: screws, washers, anchors, pins, and suture. Interference Screws and Set Screws are intended for use in fixation of patellar bone-tendon-bone grafts in ACL reconstruction. | Arthrotek Interference Screw | 2 | January 30, 2013 | 2013 | Biomet, Inc. | Thru hole at the tip of the screw is undersized. | PRODUCTION CONTROLS: Reprocessing Controls | Biomet sent an URGENT MEDICAL DEVICE RECALL letter dated January 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed locate and discontinue use of the affected product and return it to Biomet. recall notification was sent to all consignees on 1/4/2013. The action requires the immediate location and discontinued use of the product and its return to Biomet. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. | Contact the recalling firm for information | 44 | 44 units distributed | Worldwide Distribution - USA including CA, FL, IL, AL, OH, KY, TX , and Internationally to Australia, Columbia, Mexico, Netherlands, and Chile. | Terminated | August 02, 2013 | 184 | 44 | Not_Computer
Z-0757-2013 | 60906 | temporary carotid catheter for embolic capture | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: RX Accunet Embolic Protection SystemCommon Name: RX Accunet EPS.Part Numbers: 101649-45, 1011649-55, 1011649-65, 1011649-75, 1011650-55, 1011650-65, 1011651-45, 1011651-55.The RX Accunet EPS is indicated for use as a guide wire and embolic protection system. | RX Accunet Embolic Protection System | 2 | January 31, 2013 | 2013 | Abbott Vascular | The recall was initiated because Abbot Vascular has discovered that the identified lots of the RX Accunet Embolic Protection System may exhibit difficult removal of the peel away sheath due to higher than normal wall thickness. | DESIGN: Device Design | Abbott Vascular sent an "URGENT DEVICE RECALL" letter dated November 30, 2011 to all affected customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to work with their local account representative to review their inventory, complete the attached Field Action Reconciliation/ Effectiveness Check Form and return any unused identified products to Abbott Vascular. | Contact the recalling firm for information | 1253 | 1253 units (962 units in the US & Puerto Rico) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, China, Germany, Spain, France, United Kingdom, Italy, Pakistan, Poland, Russia, Slovakia, Taiwan, and South Africa. | Terminated | January 31, 2013 | 0 | 1253 | Not_Computer
Z-0758-2013 | 64055 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Ethicon Endo-Surgery Inc. The ENSEALÂŽ G2 Tissue Sealers are indicated for bipolar coagulation and mechanical transection of tissue during laparoscopic and open procedures. Model Numbers: NSLG2C14NSLG2C25 NSLG2C35NSLG2C45NSLG2514NSLG2S25NSLG2S35NSLGS545 | The ENSEAL G2 Tissue Sealers | 2 | January 31, 2013 | 2013 | Ethicon Endo-Surgery Inc | Potential for self activation while the device is in use, which may result in unintended thermal damage. | PRODUCTION CONTROLS: Process Control | Ethicon Endo-Surgery Inc. sent a "URGENT MEDICAL DEVICE RECALL" letter dated November 28, 2012, to all affected customers. The letter identified the the product the problem and the action needed to be taken by the customer.The customers were instructed to:1. Examine your inventory immediately to determine if you have affected product on hand and remove the affected product(s).2. Complete the Business Reply Form and fax it to Stericycle at 1-877-787-0367 within 3 business days, even if you do not have affected product. Please reference Event 3436. If you have product to be returned, keep a copy of this form for your records. 3.To return affected product, photocopy your completed Business Reply Form, place it in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Note: Returned product will be credited.Should you have any questions, please contact your local representative or call our Customer Support Center, 1-800-873-3636, Option 6. |  | 58338 | 58,338 Instruments | Worldwide Distribution - USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, & WV. and the countries Arab Emirates, Australia, Belgium, Canada, Chile, China, Colombia, Cyprus, Hong Kong, Indonesia, Israel, India, South Korea, Kuwait, Malaysia, Panama, Philippines, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, Thailand, Turkey & Vietnam. | Open |  |  | 58338 | Not_Computer
Z-0759-2013 | 61879 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Rectal Retractors, stainless steel, item # AL07289000, a component of Varian Medical Ring and Tandem Applicator Sets; Intended for use in Brachytherapy. | Rectal Retractor | 2 | January 31, 2013 | 2013 | Varian Medical Systems, Inc. | A medical device used in brachytherapy may become contaminated and can cause infections, toxicity or allergic reactions. |  | Varian Medical Systems an an Urgent Medical Device Correction letters dated April 11, 2012 and May 8, 2012 to all affected customers. The April notification advised of an anomaly with titanium rectal retractors, and the correction was expanded in May to include stainless steel retractors. In both notifications, users were instructed to discontinue use and quarantine their cleaned and sterilized devices for collection by Varian field service engineers. Users were requested to return an enclosed reply form to respond with number of units in inventory. | Contact the recalling firm for information | 344 | 344 pieces | US Nationwide Distribution | Open |  |  | 496 | Not_Computer
Z-0766-2013 | 64099 | system, communication, powered | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | P10 Tobii battery box (product number 510410)Product Usage - Battery pack for MyTobii P10 Device. MyTobii P10 is a portable eye-controlled communication device featuring the Tobii Eye Control system. | P10 Tobii battery pack | 2 | February 01, 2013 | 2013 | Tobii Assistive Technology, Inc. | Tobii ATI customers P10 external battery pack, which contains sealed lead-acid batteries, malfunctioned while recharging and produced smoke with a chemical smell. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The recalling firm, Tobii Assistive Technology, Inc. sent a notification letter to all customers who received the recalled product via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediatley discontinue the use of the P10 battery pack and provded with instructions to receive a complimentary replacement solution. The letter instructed customers to dispose of their P10 External Battery Pack in a manner stipulated with local laws. An acknowledgment form is included to be completed by customers. For questions contact Tobii ATI at 1-800-793-9227 or P10batterypack@tobiiati.com. | Contact the recalling firm for information | 483 | 483 units | US Nationwide Distribution | Open |  |  | 483 | Not_Computer
Z-0778-2013 | 63651 | chair, with casters | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | PREMIER CARE RECLINER 5570/5574 WITH CAL TB 133 UPGRADE PURCHASED JAN 01, 2011 TO AUGUST 20, 2012. A multi position recliner with a steel frame used primarily in clinics, hospitals, and treatment centers for recovery, dialysis, oncology, examination, infusion. | PREMIER CARE RECLINER 5570/5574 | 2 | February 04, 2013 | 2013 | Winco Mfg., LLC | The performance of the chemicals in the foam and vinyl are not consistent enough to meet the TB 133 fire protection mandate required by California and 4 other states. | DESIGN: Component Design/Selection | Winco sent a Chair Field Correction Notification dated September 24, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to forward the "Field Correction Form - REPAIR" form back to Winco and the firm would send a Winco Representative to the consignees' facility to correct the chairs in the field.Customers with questions were instructed to contact Winco at 800-237-3377 or email customerservice@wincomfg.com.For questions regarding this recall call 352-854-2929, ext 110. |  | 95 | 95 units (All models) | Nationwide Distribution including CA, FL, IL, MA, MD, MN, NM, NY, OH, OK, and TX. | Terminated | May 23, 2013 | 108 | 95 | Not_Computer
Z-0785-2013 | 64044 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Solara 3G Custom Manual WheelchairProduct Usage: To provide mobility to a person restricted to a sitting position. | Solara 3G Custom Manual Wheelchair | 2 | February 04, 2013 | 2013 | Invacare Corporation | The potential exists for the wheel to rotate freely despite engagement of hub brake. | PRODUCTION CONTROLS: Process Control | Invacare sent an URGENT: RECALL INFORMATION Letter dated September 10, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1) Complete the Provider Response Card indicating your receipt and understanding of these instructions. 2) Contact your customers to inform them of the recall. It is recommended that you provide them with the enclosed Customer Letter. Set up an appointment to correct the customer's wheelchair as soon as possible. 3) Please examine the wheelchair to confirm it has one of the serial numbers listed above. The serial number can be found on the front lower crossbar located on the right while sitting in the chair. 4) Using the enclosed retrofit kit, make the necessary correction. Installation instructions can be found on page 109 of the Solara service manual, which is enclosed.5) Document the correction of the wheelchair on the enclosed tracking sheet.6) Using the enclosed pre-paid UPS mailing label and original box, return the replaced parts to the address listed on the label. 7) Fax or email the Provider Response Form and tracking sheet to 330-329-3458 or recall@invacare.com. For questions please contact Invacare Customer Service at 800-333-6900. | Contact the recalling firm for information | 3 | 3 units | US Nationwide Distribution including the states of CA & MT. | Open |  |  | 3 | Not_Computer
Z-0784-2013 | 63390 | pack, hot or cold, reuseable | IME | Cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | THERA PEARL 3-in-1 BREAST THERAPY, 2 hot and cold packs in plastic box. The Thera Pearl 3-in-1 Breast Therapy is a reusable therapeutic hot and cold pack for use with pain, swelling, engorgement, plugged ducts or mastitis associated with breast feeding or to help encourage let-down while using a breast pump. | THERA PEARL 3in1 BREAST THERAPY | 2 | February 04, 2013 | 2013 | Lansinoh Laboratories Inc | The Spanish translation on the Instructions for Use (IFU) insert and product package for hot and cold packs has incorrect heating instructions which would cause product to over heat. | PRODUCTION CONTROLS: Error in Labeling | Lansinoh Laboratories, Inc. sent an Urgent Product Recall letter dated August 14, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop distributing and quarantine the affected product. Customers were informed that a credit would be issued for the affected product and they were asked to mail the postage paid Business Reply Card even if they do not have the affected product. If necessary call 1-877-366-1182 to arrange pickup. For questions customers were instructed to call the Customer Service Center at 1-800-292-4794.For questions regarding this recall call 703-299-6477. |  | 51472 | 51,472 units | Nationwide Distribution, including the states of TX, CA, NJ, NY, SD, RI, WV, IL PA, NC and MN. | Terminated | March 08, 2013 | 32 | 51472 | Not_Computer
Z-0782-2013 | 62943 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes(R) Spine Vectra, Vectra-T and Vectra-One Anterior Cervical Plating for Spinal Fusion, Adjustable Drill Stop, Part Number 03.613.012.Synthes Vectra System, Synthes Vectra-One System and Synthes Vectra-T System is intended for anterior plate and screw fixation of the cervical spine (C2-C7). | Synthes(R) Spine Vectra, VectraT and VectraOne Adjustable Drill Stop | 2 | February 04, 2013 | 2013 | Synthes USA HQ, Inc. | Although the reported incidence is low, there is the potential for the drill stop to slip on the drill bit resulting in an adverse event. |  | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 610-719-5450 for questions related to this notice. |  | 48 | 48 | Nationwide Distribution-including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Open |  |  | 111 | Not_Computer
Z-0772-2013 | 63531 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trilogy Acetabular Shell Spiked.The Trilogy Acetabular System is indicated for either cemented or uncemented use in individuals undergoing surgery for rehabilitating hip damage. | Trilogy Acetabular Shell Spiked | 2 | February 04, 2013 | 2013 | Zimmer, Inc. | Some units of Trilogy shells manufactured since March 2009 may be below the lower range of the pore size specification (as low as 28% vs. 30% porosity specification). On 10/31/2012, the firm expanded the recall action to include lots of Uni-Hole Trilogy Acetabular System shells. | PRODUCTION CONTROLS: Process Control | Zimmer sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated September 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.The firm expanded the scope of the recall in January 2013 to include devices that use a secondary mean of fixation. New recall letters will be submitted to distributors and surgeons. |  | 275952 | 275,952 units | Worldwide Distribution-USA (nationwide) and the countries of Denmark, Germany, Greece, Hungary, Ireland, Italy, Netherlands, Spain, Sweden, UK, Poland, France, Australia, Chile, Canada, India, Japan, New Zealand, Peru, South Africa, Colombia, Ukraine, and UAE. | Open |  |  | 275952 | Not_Computer
Z-0783-2013 | 63793 | burr, orthopedic | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Effort, Inc., 98-0061 MA-15C Bearing Sleeve with Removable Bur Guard. A reusable device used with dissection tools; designed for Transphenoidal and Skull base procedures. | Bearing Sleeve with Removable Bur Guard | 2 | February 04, 2013 | 2013 | The Anspach Effort, Inc. | Contactbetween the bur and bur guard could generate metal fragments that may or may not be visible to thesurgeon and can potentially remain in the surgical site. The materials used to fabricate the bur guard are not traceable, design validation did not effectively evaluate adequate protection of adjacent tissue and inspection results for each of these bur guards were not documented. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Anspach Effort, Inc. sent an "URGENT: Medical Device Removal" letter dated November 7, 2012, to its customer. The letter identified the product, problem and actions to be taken. The customer was instructed to do the following: 1) Screen their inventory and remove and return all products immediately. 2) Complete and return the attached Customer Reply Form via fax or email to the address provided on the form. Should the customer have any questions, please contact Anspach Product Support at (800) 327-6887. |  | 8 | 8 devices | Distributed in the state of Massachusetts | Open |  |  | 8 | Not_Computer
Z-0787-2013 | 64028 | prosthesis, hip, semi-constrained, metal/ceramic/polymer, cemented or non-porous, uncemented | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech NOVATION CROWN CUP, NEUTRAL LINER, Use with Crown Cup Only, GXL UHMWPE, SN 0123456, REF 130-32-53.Exactech Hip Systems are indicated for use in surgery for hip replacement. | Novation Crown Cup Connection GXL Neutral Liners | 2 | February 05, 2013 | 2013 | Exactech, Inc. | Recall of Novation Crown Cup Connection GXL Neutral Liners Catalog No. (130-32-53) and Novation Crown Cup Connection GXL + 5mm Lateralized Liners Catalog No. (136-32-53). Exactech discovered the device labeling was inadvertently transposed. | PRODUCTION CONTROLS: Labeling Mix-Ups | Exactech, Inc. notified consignees by telephone on 11/05/12 of the affected devices and sent an "IMPORTANT PRODUCT RECALL NOTICE" letter dated 11/06/12. The letter describes the purpose of the recall, the product affected, the problem and actions to be taken by the consignees. A Product Recall Response Form was attached for the consignees to contact the Exactech inventory representative within 5 business days to confirm quantities of product in their inventory and return them to Exactech. Consignees with any questions relating to these devices should call 1-800-392-2832. |  | 26 | 26 Units | Distributed in the states of TN, FL, VA, WA and NV. | Terminated | December 24, 2013 | 322 | 26 | Not_Computer
Z-0789-2013 | 63823 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConMed Linvatec Biomaterials, PopLok bone Punch PKL-35M, 3.5mm Orthopedic manual surgical instrument - Bone Punch | PopLok Punch 3.5MM | 2 | February 05, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatech recalled one lot of their PopLok bone Punch 3.5mm because the metal shaft diameter is 4.5mm when it should be 3.5mm. | TRAINING: Employee Error | ConMed Linvatech issued a "URGENT MEDICAL DEVICE RECALL NOTIFICATION" to its customers on November 15, 2012 for a recall that was conducted at the end user/distributor level. All recall notifications were sent FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with reason for the recall health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169.For questions regarding this recall call 727-399-5205. |  | 8 | 8 | Worldwide Distribution - USA including California and Internationally to Spain and Korea | Terminated | May 22, 2013 | 106 | 8 | Not_Computer
Z-0790-2013 | 63555 | susceptibility test discs, antimicrobial | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, Aztreonam 30 ug, IVD, REF CT0264B, Oxoid Ltd., UK, oxoid@thrmofisher.com (Both Remel &amp;amp;amp;amp;amp; Oxoid, Ltd are a part of Thermo Fisher Scientific)An antimicrobial susceptibility test disc is a device that consists of antimicrobic-impregnated paper discs used to measure, by a disc-agar diffusion technique or a disc-broth elution technique, the in vitro susceptibility of most clinically important bacterial pathogens to antimicrobial agents (Aztreonam). | Oxoid Antimicrobial Susceptibility Test Discs | 2 | February 06, 2013 | 2013 | Remel Inc | The firm is recalling the product due to some cartridges may contain individual discs that are not impregnated with antibiotic. The firm notified their customers of the recall with a notification letter and requested a confirmation fax back of receipt. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ThermoFisher sent an Important Medical Device Product Recall Notice dated August 17, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review results reported using this lot and consider restesting and/or seek approproiate expert advice. Customers were requested to destroy any remaining inventory of the affected lot and contact Customer Services or their local Oxoid supplier who will issue a replacement. Customer were asked to return the attached "Acknowledgement & Receipt Form by FAX to 1-877-428-1922 or 1-913-895-4190. Customers with questions were instructed to call the Technical Support Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. | Contact the recalling firm for information | 30 | 30 | Nationwide Distribution including WI, MO, MA, CT, AL, IL, WA, CA, RI, MD, NJ, NC and TX. | Terminated | June 04, 2013 | 118 | 30 | Not_Computer
Z-0791-2013 | 64160 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Medacta Quadra Trial Broach size 0 part # 01.10.10.045, for use with the Medacta Total Hip Prosthesis System. Product Usage - used to broach the intramedullary canal of the femur in order to create space for the femoral stem during a hip replacement surgery. | Quadra Trial Broach size 0 | 2 | February 07, 2013 | 2013 | Medacta Usa Inc | The tip of the Quadra Trial Broach size 0, broke during a surgery. | DESIGN: Device Design | Medacta sent a Urgent Safety Notification letter dated December 13, 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to check inventory for affected product and contact QA@medacta.us.com for an RMA# to return items to Medacta USA. A new version of the recalled product will be provided to all affected customers. For questions call 805-437-7085 x26. |  | 26 | 26 pieces U.S.A. (226 pieces Worldwide) | US Nationwide Distribution including the states of Arizona, California, Idaho, Illinois, North Carolina, Ohio, Utah, Washington and US Virgin Islands | Terminated | February 07, 2013 | 0 | 26 | Not_Computer
Z-0792-2013 | 64115 | culture media, antimicrobial susceptibility test | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Diagnostic Hybrids, Inc., RhMK, Model #:49-0600, 49-0600A, 49-0102A, and 49-T075A. Culture media, used in the cultivation and amplification of viruses, such as respiratory viruses, enteroviruses, measles, and poliovirus from patient specimens. | Diagnostic Hybrids Inc. | 3 | February 07, 2013 | 2013 | Diagnostic Hybrids Inc | RhMK product fungal contamination. | PRODUCTION CONTROLS: Process Control | Diagnostic Hybrids sent a Urgent: Medical Device Recall letter dated December 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to immediately review your inventory and determine if you have any of the affected product.If so, please remove it from inventory. Complete the enclosed Product Recall Fax-Back form. Upon receipt of Fax-Back form, Diagnostic Hybrids will send you replacement product or issue credit.Representatives are available to assist you in this process and answer any questions you may have about this recall, and how to obtain replacement product. Please contact Technical Support at technicalsupport@quidel.com or by calling 800.874.1517, Monday through Friday, 8:00 a.m. to 5:00 p.m. EST. |  | 13005 | 13,005 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA & WI. | Terminated | December 05, 2013 | 301 | 13005 | Not_Computer
Z-0796-2013 | 64084 | tray, biopsy, without biopsy instruments | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Breast Biopsy Tray, Catalog # B 9-16261, packaged in bags, 1 tray/bag, 20 bags/case.Product Usage: A sterile medical device containing multiple components utilized by physician to perform a breast biopsy. | Breast Biopsy tray | 2 | February 07, 2013 | 2013 | Custom Healthcare Systems, Inc. | Surgical procedure trays are packed with pharmaceuticals for injection that may contain particulate matter. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Custom Healthcare Systems, Inc. sent an Urgent Recall Letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory, quarantine affected product and return affected product to Stericycle using the label provided with the letter. Customer were also requested to complete and return the enclosed recall response form. For questions contact Hospira Customer Care at 1-877-946-7747. |  | 80 | 80 trays | USA Nationwide Distribution including the states GA, OH, GA, NC and ME. | Terminated | February 08, 2013 | 1 | 80 | Not_Computer
Z-0793-2013 | 64132 | system, test, radioallergosorbent (rast) immunological | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | OPTIGEN Food 30; Catalog number: 85035. For allergen testing. Manufactured by Hitachi Chemical Diagnostics, Mountain View, CA.Distributed by Theridiag, Cedex, France | OPTIGEN Food 30 | 3 | February 07, 2013 | 2013 | Hitachi Chemical Diagnostics Inc | There is no evidence that a visual inspection was performed on one tray of coverslips. Possibility of a low fill volume which could lead to a false negative results. | PRODUCTION CONTROLS: Process Control | Hitachi sent an Urgent Product Recall Notification letter dated December 21, 2012, to the affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was instructed to discontinue distribution and destroy any remaining inventory of the affected product at their location per their local waste ordinances, determine if any of their customers are using or have inventory of the affected product, rquest that their customers review all patient results reported with the use of the affectted product. The customer was instructed to tell their customers to discontinue use and destroy any remaining inventory per their local waste ordinances. The customer was instructed to retain a copy of the notification letter for their records. The customer was also instructed to send confirmation of the destruction within 5 days Customers were instructed to complete and return the attached Certificate of Destruction. For questions customers were instructed to contact their local area Customer Service.For questions regarding this recall 650-961-5501. | Contact the recalling firm for information | 20 | 20 kits | Worldwide Distribution - France only | Terminated | May 16, 2013 | 98 | 20 | Not_Computer
Z-0794-2013 | 64069 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | SeriScaffold Surgical ScaffoldProduct Number: SCF10X25AGEN.Surgical mesh for use in open or laparoscopic procedures. | Allergan SeriScaffold Surgical Scaffold | 2 | February 07, 2013 | 2013 | Allergen Medical | Product is dual packaged in an inner and outer pouch, the outer pouch seal may be compromised and sterility may not be assured | PRODUCTION CONTROLS: Packaging Process Control | Allergan sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers are requested to quarantine and return any unused devices in inventory . In addition, customers are asked to complete and return the acknowledgement form. Customers with questions should contact Allergan at 781-658-2800, Mon- Fri, 9 a.m. to 5 p.m. EST. | Contact the recalling firm for information | 225 | 225 units (US) and 81 (Outside US) | Worldwide Distribution-USA including the states of CA, FL, GA, MI, MN, NJ, NY, OH, OR, RI, TX, VA, and WA, and the countries of Germany, Spain, and UK. | Terminated | July 29, 2013 | 172 | 225 | Not_Computer
Z-0801-2013 | 63584 | supplement, culture media | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | The product is a lyophilized growth supplement for use with dehydrated culture media to prepare BCVE Medium used for isolation of Legionellae from clinical and environmental samples. | Legionella BCYE Growth Supplement SR0110C | 2 | February 11, 2013 | 2013 | Remel Inc | The product may contain low level microbial contamination which could result in incorrect results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | ThermoFisher Scientific notified customers on October 10, 2012, with a Field Safety Notice letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised use of this lot could result in incorrect results reporting. This notice needs to be passed on to all who need to be aware within your organization.The firm request the information form be faxed back confirming receipt of notice.For further information please contact Technical Support via our main switchboard on +44 (0) 1256 841144. |  | 10 | 10 units | Distribution included the states of IL, IN and MI. | Terminated | June 04, 2013 | 113 | 10 | Not_Computer
Z-0799-2013 | 64175 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn. Ext. DL. Part Number: 050-87216.Acute and chronic peritoneal dialysis. | Fresenius Liberty Cycler Set Single Conn. Ext. DL. | 2 | February 11, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | There is a small risk of a leak in Cycler Sets and if the Cycler Set is punctured, the dialysate may become contaminated potentially resulting in peritonitis. |  | Fresenius Medical Care began contacting the affected customers by telephone and mailing an "URGENT: FMCNA LIBERTY CYCLER SET RECALL" letter and Return Receipt form dated October 31, 2012. The letters identified the specific part number and product lot and request that customers discontinue use of the lots immediately and to return unused product to FMCNA. Instructions are provided for returning the product, asking customers to isolate and discontinue use of the lot number 12BR08009 and place the product in a secure, segregated area. Consignees were instructed to contact their Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product. |  | 560 | 560 cases, totaling 5590 units | Nationwide Distribution. | Open |  |  | 560 | Not_Computer
Z-0763-2013 | 64155 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, AGA Medical corporation, AMPLATZERÂŽ TorqVueÂŽ FX Delivery System. Model #'s: 9-ITVFX06F45/60, 9-ITVFX07F45/60, 9-ITVFX007F45/80, 9-ITVFX08F45/60, 9-ITVFX08F45/80, 9-ITVFX09F45/80, 9-ITVFX10F45/80 9-ITVFX12F45/80, 9-ITVFX13F45/80. Cardiac surgical instrument, used to close atrial septal defects between the right and left atriums. | AMPLATZER TorqVue FX Delivery System | 1 | February 11, 2013 | 2013 | St Jude Medical Inc | The distal end of the core wire of the TorqVue FX Delivery System could potentially fracture when exposed to a combination of certain cardiac anatomies and usage conditions. | DESIGN: Device Design | St. Jude Medical sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were advised to discontinue use of the device and remove it from inventory. Customers should contact their SJM sales representative for questions relating to this recall. |  | 635 | 635 | Nationwide Distribution-including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, and WI. | Terminated | May 21, 2013 | 99 | 635 | Not_Computer
Z-0800-2013 | 63433 | filler, bone void, osteoinduction (w/o human growth factor) | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Bacterin International, Inc., Osteoselect Demineralized Bone Matrix DMB Putty; Graft ID# B090056-664 Donor# B090056 2.5cc Expiration Date: 12/25/2011; LBL 210-0010-1 REV: 040109; and Graft ID# B090056-675. Bone void filler for voids or gaps in bone. Indicated for treatment of surgically created osseous defects.On 07/12/2012 an additional expired product was discovered by a Bacterin sales rep. B100078-908 with expiration date 04/27/2012 | Osteoselect Demineralized Bone Matrix DMB Putty | 2 | February 11, 2013 | 2013 | Bacterin International, Inc. | Expired product was shipped to two customers. | PRODUCTION CONTROLS: Error in Labeling | Bacterin International phoned the two customers and they were instructed to return any of the unused product.For questions regarding this recall call 406-388-0480.The 3rd expired product was discovered by Bacterin Sales Rep 7/12/2012 at the consignee. | Contact the recalling firm for information | 3 | 3 packages each containing 2.5cc of putty | Nationwide Distribution in TX , GA and CA. | Terminated | May 13, 2013 | 91 | 3 | Not_Computer
Z-0808-2013 | 63556 | susceptibility test discs, antimicrobial | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Discs, Cefoxitin 30ug REF CT0119B, packaged with 50 discs per cartridge, 5 cartridges per box. The responsible firm name on the label is Oxoid Ltd., Basingstoke, Hants, UK. | Oxoid Antimicrobial Susceptibility Discs FOX30, Cefoxitin | 2 | February 12, 2013 | 2013 | Remel Inc | The firm is recalling the product because the cartridges may contain discs that have an insufficient dose of antibiotic. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | The recalling firm sent a recall notification letter dated 8/17/ 2012 and 8/21/2012. |  | 23 | 23 units | US Nationwide Distribution including the states of: CA, CT, GA, IA, KS, OK, RI, TX, WA, WI, | Terminated | June 04, 2013 | 112 | 23 | Not_Computer
Z-0805-2013 | 63244 | shunt, central nervous system and components | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap(R) - Miethke Shunt System, FT040T - Control Reservoir Set w/Distal Catheter and FT075P 0 Ventric Cathet, w/Stylet and RT-ANG GuideCentral Nervous System Shunt and Components intended to shunt cerebrospinal fluid (CSF) from the lateral ventricles of the brain into the peritoneum. | Aesculap(R) Miethke Shunt System | 2 | February 12, 2013 | 2013 | Aesculap, Inc. | The deflector may not move freely on the ventricular catheter. The deflector comes pre-positioned at the 5 cm mark on the catheter. In some cases the surgeon will want to adjust the deflector to a different location on the catheter before placing it in the ventricle. There is a small percentage of cases where aggressively attempting to adjust the deflector has a potential to cause the catheter | PRODUCTION CONTROLS: Process Control | Aesculap sent an Important Recall Notification letter dated June 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the contact their local sales representative or customer service department at 1-800-282-9000 to return the affected product. Customers were also instructed to complete the Inventory Sheet attached to the letter with the quantity being returned. If customers could not locate the affected product they should provide an explanation as to why the inventory will not be returned. Customers with questions should call 610-984-9251 or 610-984-9291.For questions regarding this recall call 1-800-234-9179, ext 5067. |  | 11 | 11 | Nationwide Distribution including MI, PA and TX. | Terminated | January 08, 2014 | 330 | 11 | Not_Computer
Z-0803-2013 | 64177 | pack, hot or cold, water circulating | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | BREG Cold Therapy Combination Units containing Sterile Polar Pads. 02356 Rev M, PAD M/U XL STER, Mfg.Date: 112012. Model #s: 10903, 10703, 09101, 09111, 09131, 09731, 09611, 09621, and 09631. Local anesthetic effect. | BREG Cold Therapy Combination Units | 2 | February 12, 2013 | 2013 | Breg Inc | The recall was initiated because Breg has determined that some Sterile Polar Pads products manufactured from January 2010 to October 2012 may have sustained damage to the product packaging which may compromise product sterility assurance. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | The firm, BREG, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product; complete the Return Response Form by email, Fax or mail; and contact Breg for a Return Authorization for the affected product and request replacement by contacting Breg Customer Care at 800-321-0607. Note: response is required even if you have no affected inventory.Should you have any questions regarding this communication or need to report an adverse event, please contact Breg Customer Care at 800-321-0607. |  | 44883 | 44,883 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Singapore, Chile, Latvia. | Terminated | July 01, 2013 | 139 | 44883 | Not_Computer
Z-0807-2013 | 63558 | reagents, clostridium difficile toxin | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Xpect Clostridium difficile Toxin A/B, IVD, 40 clostridium difficile toxin A/B tests, REF 24640.Product Usage:The test kit is a rapid in vitro immunochromatographic test for the direct, qualitative detection of Clostridium difficile Toxin A and/or B in human fecal specimens. The test is intended for use as an aid in diagnosis of Clostridium difficile-associated disease. | Remel X/pect Clostridium Difficile Toxin A/B (40 tests) | 2 | February 12, 2013 | 2013 | Remel Inc | The firm is recalling the products due to a potential for false negative test results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Thermo Fisher sent a recall notification letter with details of the recall on August 20, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were requested to follow the instructions provided and fax back the attached "Acknowledgement & Receipt Form" to 1-877-428-1924 (toll free in the US) or 1-913-895-4190 (International). For questions contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 50 | 50 units | Worldwide Distribution - USA Nationwide and the countries Canada and the United Kingdom | Terminated | June 04, 2013 | 112 | 543 | Not_Computer
Z-0810-2013 | 63991 | table, surgical with orthopedic accessories, ac-powered | JEA | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Mizuho Orthopedic Table, model 5855, sub assembly 5855-901 Orthopedic Table Spar. | Mizuho Orthopedic Systems, Inc. | 2 | February 13, 2013 | 2013 | Mizuho OSI | Possible instability of Orthopedic Table Top. If the epoxy bond fails, the aluminum plug may migrate out of the carbon fiber tube. If this occurs, the Orthopedic Table Spar may need to be replaced to ensure stability of the Orthopedic Table Top. | PRODUCTION CONTROLS: Process Control | Mizuho sent a Recall Notification letter dated December 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Orthopedic Table Spar will be replaced at no cost by Mizuho technicians. Customers were informed the potential failure and were advised inspection of tables prior to use. If tables are experiencing the described failure, contact Mizuho Resource Group at 1-800-777-4674, option 2. . | No consumer action necessary | 869 | 869 devices | USA Distribution -- Worldwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV and WY., and the countries of Australia, Canada, Egypt, Finland, France, Germany, China, Israel, Italy, Japan, Netherlands, New Zealand, P.R. China, Saudi Arabia, Singapore, Slovenia, South Korea, Spain, Taiwan R.O.C., UAE and UK. | Open |  |  | 869 | Not_Computer
Z-0809-2013 | 64110 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Chenica Gunnell Rehab Series of Wheelchairs (GRSW); Mechanical Wheelchair, Sold as Rehab Tough and Tilt (TNT), Rehab Recline and Mobility (RAM), Rehab Kidster (Kidster). Provide mobility to individuals limited to a seating position. | Rehab Tough and Tilt (TNT), Rehab Recline and Mobility (RAM), Rehab Kidster (Kidster) | 2 | February 13, 2013 | 2013 | CHENICA INC. | The wheelchairs may not meet the specifications of WC-19 (crash testing), therefore the WC-19 verbiage and corresponding logo are removed from the flyers on the Gunnell Rehab Series of Wheelchairs (GRSW). | MISBRANDING: Labeling False and Misleading | Gunnell Inc. sent a "RE: URGENT MEDICAL DEVICE CORRECTION/RECALL" letter dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm 1-800-551-0055 for questions relating to this letter. | No consumer action necessary | 53 | 53 | Nationwide Distribution-including the states of Florida, Michigan, South Carolina, California, Texas, Maryland, New York, Tennessee, Virginia, Illinois, Ohio, Massachusetts, Arizona and Washington. | Open |  |  | 53 | Not_Computer
Z-0812-2013 | 64272 | material, impression | ELW | Impression material. | Dental | Dental | 510(k) | Ultradent Products, Chromaclone PVS, Fast Set heavy body, Bubble Gum Flavor, 50 ml, Impression Material, Part No. 5013, Lot No. B7DS3, Used in prosthodontics. | Chromaclone PVS Heavy (Fast Set) | 3 | February 13, 2013 | 2013 | Ultradent Products, Inc. | Incorrect packaging of regular set and fast set cartridges. | PRODUCTION CONTROLS: Packaging Process Control | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected products. They may continue to use product or dispose unwanted cartridges and reply back to firm indicating acknowledgement of letter and if replacement is needed. Consumers with questions can call the firm at 801-553-4420.For questions regarding this recall call 801-553-4420. |  | 100 | 100 cartridges | Nationwide Distribution | Terminated | May 30, 2013 | 106 | 132 | Not_Computer
Z-0818-2013 | 63927 | electrode, cutaneous | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Foam electrodes with Silver-Carbon construction:P/N 202-9271, Cat No. 400-898, 2&amp;quot; x 2&amp;quot; Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9281, Cat No. 202-9281, 2&amp;quot; Round Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9251, Cat No. 400-893, 3&amp;quot; Round Electrode, Foam AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9291, Cat No. 400-892, 3&amp;quot; Round Electrode, Foam AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9371, Cat No. 400-888, 2&amp;quot; x 3.5&amp;quot; Electrode, Foam AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9491, Cat No. 400-869, 1.25&amp;quot; Round Electrode, Foam AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 202-9511, Cat No. 400-853, 3&amp;quot; x 5&amp;quot; Electrode, Foam AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 300-1541, Cat No. 300-154, 2&amp;quot; x 2.25&amp;quot; Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 300-1561, Cat No. 300-156, 2&amp;quot; Round Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 300-1611, Cat No. 300-161, 4&amp;quot; x 2.25&amp;quot; Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010, 22818P/N 300-1631, Cat No. 300-163, 3&amp;quot; Round Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010P/N 400-8701, Cat No. 400-870, 2&amp;quot; x 4&amp;quot; Oval Electrode, Foam, AG, All Serial Numbers distributed between 11/01/2009 and 01/31/2010.Cutaneous electrode. | Foam Electrodes with SilverCarbon construction | 2 | February 14, 2013 | 2013 | NAImco Inc dba Richmar Inc | Electrode Socket Terminal pulling away from the electrode lead wire when the electrical stim lead wire was removed. | DESIGN: Component Design/Selection | Rich-Mar sent "URGENT: MEDICAL DEVICE VOLUNTARY RECALL" letters to affected customers beginning on 11/30/2009. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 423-648-7730 ext. 214 for questions relating to t his recall. | Contact the recalling firm for information | 56980 | 56980 electrodes | Distributed in the states of GA and TN. | Terminated | September 03, 2013 | 201 | 422100 | Not_Computer
Z-0819-2013 | 64017 | antigens, cf (including cf control), influenza virus a, b, c | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Quidel, Sofia Influenza A + B FIA, Kit #20218 (25-Test Kit), in vitro diagnostic. | Sofia Influenza A B FIA, Kit 20218 | 2 | February 14, 2013 | 2013 | Quidel Corporation | Quidel Corporation has initiated a voluntary recall of specific lots of the Sofia Influenza A+B FIA kits because of false positive results which may lead to misdiagnosis. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Quidel Corporation sent a Medical Device Recall dated December 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and remove the affected product from their inventory. Customers should contact Quide Corporation for instructions on how to obtain replacement product. Customers were instructed to fax the enclosed inventory Form to Quidel at 858-552-7905.For technical support customers should call 800-874-1517 or 858-552-1100.For questions regarding this recall call 858-552-1100, ext 7922. |  | 13090 | 13,090 kits (327,250 tests) | Worldwide Distribution - USA (nationwide) and internationally to Korea, Norway, Germany, Chile, Hong Kong, Switzerland, Netherlands, Thailand, Japan, China, and the UK | Terminated | January 31, 2014 | 351 | 13090 | Not_Computer
Z-0795-2013 | 63858 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | The Lumenis VersaCut Tissue MorcellatorProduct Usage: is intended for use under direct or endoscopic visualization for the morcellation and removal of dissected tissue during pelviscopic, laparoscopic, percutaneous, and open surgical procedures whenever access to the surgical site is limited. | VersaCut Morcellator | 1 | February 14, 2013 | 2013 | Lumenis Limited | Lumenis Ltd is recalling certain models of VersaCut Morcellator devices to perform corrective labeling as a result of three reported events of air embolism secondary to incorrect reversal tubing system hook-up. | DESIGN: Labeling Design | Lumenis sent an Urgent Safety Advisory Notice letter dated January 2, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to re-label the devices with the enclosed labels provided, replace the Operation Manual with the current revised Manual, complete the information requested on the Confirmation Card and return the Confirmation Card using the enclosed pre-paid shipping label by UPS. For question contact VersaCut Morcellator field correction administrator by either Cell: +972-(54)-9599196, Tel: +972-(4)-9599296 or E-mail: VersaCutMorcellator@lumenis.com. |  | 931 | 931 units | Worldwide Distribution - US Nationwide including: CA, WI, NC, FL, PA, IN, NM, NJ, IA, WA, OH, FL, MD, MA, WY, NY, AZ, NH, MN, TN, OK, KY, TX, AK, MI, VT. Internationally: Japan, Nepal, India, UK, South Korea, Netherlands, Germany, Australia, China, Canada, Turkey, Spain, Italy, Egypt, Mexico, Bulgaria, Utd. Arab, Emir., Poland, Kuwait, Hong Kong, Saudi Arabia, Yemen, Qatar, Czech Republic, New Zealand, Malaysia, Portugal, France, Argentina, Singapore, Romania, Chile, Colombia, Switzerland, Bahrain, Brazil, Serbia/Montenero, Taiwan, Israel, Indonesia, Peru, Pakistan, Nepal, Belarus, Latvia, Sri Lanka, Russian Fed, Thailand, Azerbaijan, Belgium, and Denmark. | Terminated | May 20, 2013 | 95 | 931 | Not_Computer
Z-0654-2013 | 64029 | ir laser illuminator only/night vision system | REX |  |  |  |  | Lockheed Martin Gyrocam Systems ( LMGS)LMGS 15" camera systems Laser pointer and illuminator used for law enforcement surveillance from aircraft. | Lockheed Martin Gyrocam Systems ( LMGS | 2 | February 14, 2013 | 2013 | Lockheed Martin Gyrocam Systems, Inc. | These camera systems were manufactured by LMGS and found to fail to comply with the FDA laser performance standard during an FDA inspection. These camera systems incorporate Class IIIb surveying, leveling, and alignment lasers and were entered into US commerce without a variance. Additionally, the manufacturer did not submit a product report to FDA for the camera system. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Please note that LMGS is responsible for the correction of all noncompliant products.1. Notifications to all purchasers must be made, and the FDA district office Recall Coordinator is to be included in the notifications. 2. Correction of all units to be completed by November 8, 2013.3. All products will be brought into compliance with the laser performance standard at no cost to the purchaser.4. LMGS will submit a product report for the new camera system prior to installation of the replacement lasers as per 21 CFR 1002.10 and certify that the new camera systems comply with the general performance standard in 21 CFR 1010 and the laser performance standard found in 21 CFR 1040.10 and 1040.115. Updated operator manuals will be provided to purchasers.Further questions please call (941) 355-3206. |  | 33 | 33 units | Nationwide Distribution | Open |  |  | 33 | Not_Computer
Z-0832-2013 | 64204 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion SmartSite Gravity Blood Set Model 42082E. Used to administer blood and blood products through a needle or catheter inserted into the patient's artery or vein. | CareFusion SmartSite Gravity Blood Set | 2 | February 15, 2013 | 2013 | Carefusion 303 Inc | CareFusion is recalling the SmartSiteÂż Gravity Blood Set as a result of a potential for separation at the connection of the drip chamber and tubing . The separation may be observed at priming or observed during use. If a separation between the drip chamber and tubing occurs, this could cause a delay in treatment. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Carefusion sent an "URGENT: Medical Device Recall Notification" letter dated January 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to return all recalled product directly to the distributor from whom it was purchased immediately, complete and return the enclosed pre-addressed and postage paid mandatory Customer Recall Response Card and to contact your distributor directly, or call the Carefusion Support Center, if you have any questions.For questions Recall Related Questions contact CareFusion Support Center at 888-562-6018; 7am to 4pm (Pacific). For technical questions contact Technical Support at 888-812-3229, 7am to 5pm (Pacific). | No consumer action necessary | 5050 | 5,050 units | Nationwide Distribution including the states of AK, CA, CO, IL, MO, NV, NY, TX and UT. | Terminated | July 25, 2013 | 160 | 5050 | Not_Computer
Z-0825-2013 | 63933 | unit, cryosurgical, accessories | GEH | Cryosurgical unit and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | HealthTronics Endocare 3.8mm RenalCryo Cryoprobe, R3.8L. The Endocare CryoProbe is an accessory probe intended for use with the CryoCare Surgical System, for freezing and ablation of tissue. | HealthTronics Endocare 3.8mm RenalCryo Cryoprobe | 2 | February 15, 2013 | 2013 | Healthtronics, Inc. | The Endocare 3.8mm RenalCryo Cryoprobe is being recalled because the label on the outer box has the incorrect expiration date. | EXPIRATION DATING: Incorrect or No Expiration Date | Healthtronics sent an Urgent Device Recall letter dated December 20, 2012, to all affected customers that allows tracking. The letter included a suggested format/content for a correction status/acknowledgement letter. The customer was requested to complete the verification section of the recall letter and return all affected products to Healthtronics. A return material authorization (RMA) and pre-paid FedEx return labels was provided for all customers. All affected Healthtronics field personnell were notified of the recall via email.For questions regarding this recall call 512-439-8355. |  | 50 | 50 units | Nationwide Distribution including CA, CO, FL, GA, NC, NJ, OH, TX, and WI | Terminated | April 22, 2013 | 66 | 50 | Not_Computer
Z-0830-2013 | 63783 | prosthesis, knee, patello/femorotibial, semi-constrained, uncemented, porous, coated, polymer/metal/polymer | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Diaphyseal SleevesProduct Usage:The LPS Diaphyseal Sleeves are used with the LPS System which is an end-stage revision knee product that allows surgeons to reconstruct severe soft tissue and bony defects. | LPS Diaphyseal Sleeves | 1 | February 15, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics is initiating a voluntary recall of the LPS Diaphyseal sleeve because the product has the potential for fracture at the taper junction. The LPS Diaphyseal Sleeve to Diaphyseal Sleeve Base taper connection may not be sufficient to accommodate potential physiologic loads that may be transferred to the junction during normal gait activities by some patients. | DESIGN: Device Design | DePuy Othopaedics sent an Urgent Information Medical Device Recall Notice letter dated January 4, 2013 to all affected customers. The letter identified the affected products, reason for recall, clinical implications and actions to be taken. The letter instructed customers to check inventory, discontinue use, cease further distribution and return all inventory on hand for credit. Customers were requested to complete and return the Reconciliation Form. For product-related questions contact your local DePuy Othopaedics Sales Representative, clinical questions contact DePuy's Scientific Information Office at 1-888-554-2482 and recall notice information call 574-372-7333. | Contact the recalling firm for information | 162 | 162 total units (US) and 102 units total (OUS) | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, GA, ID, IL, MA, MD, MI, MN, NC, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV and the countries of Holland, Australia, Austria, South Africa and Canada including Veteran Administration hospitals. | Open |  |  | 162 | Not_Computer
Z-0835-2013 | 64321 | needle, hypodermic, single lumen | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | VACUETTE Safety Blood Collection Set + Luer Adapter, 23x12" Safety Blood Collect LA Product, Item number 450096. Usage: The Safety Blood Collection Set is used in routine venipuncture procedures. | VACUETTE Safety Blood Collection Set Luer Adapter | 2 | February 15, 2013 | 2013 | Greiner Bio-One North America, Inc. | The VACUETTE Safety Blood Collection Set + Luer Adapter may display leaking (small droplets) during use due to a crack on the luer connector piece. |  | Greiner Bio-One sent a notification letter dated January 31, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to stop distributing the affected product, complete the attached Product Disposition form and fax to GBO at 800-726-0052 or email to patech@gbo.com. GBO will contact customers for the return of the affected product after receiving the attached Product Disposition form. Customers were requested to provide the names of customers who may have received the product. For questions call 800-515-8112. |  | 192000 | 192,000 pieces (192 cases) | US Nationwide Distribution in the states of FL, IL, KY, NY, NC, PA, OH, TX, VA | Terminated | October 29, 2013 | 256 | 192000 | Not_Computer
Z-0826-2013 | 64248 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ELECTROLASE HYFRECATOR(R) ELECTRODES - SHARP, REF/Catalog Number 7-100-8BX, Disposable Hyfrecator(R) Tips, Sterile, Rx Only --- CONMED CORPORATION USA.Intended to be used as active RF electrodes for desiccation, fulguration and coagulation. | ELECTROLASE HYFRECATOR(R) ELECTRODES SHARP | 2 | February 15, 2013 | 2013 | ConMed Corporation | Some devices had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. | PRODUCTION CONTROLS: Equipment Maintenance | ConMed sent Urgent Device Recall Letters (dated 1/28/13) and Business Reply Forms to all domestic customers via USPS Priority Mail. International notices were sent via FedEx International Priority to the foreign consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II response form and return it with unused affected product to: ConMed Corporation 525 French Road, Utica, NY 13502. Customers were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. | No consumer action necessary | 104250 | Domestic Distribution: 97,450 units; Foreign Distribution: 6,800 units | Worldwide Distribution-USA (nationwide) and the countries of China, Dubai and Indonesia. | Open |  |  | 104250 | Not_Computer
Z-0842-2013 | 64158 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | NEUROPATH 6Fx90 GUIDING CATH MPCProduct Code: GMC60900000Intended for use in the coronary, peripheral, and neurovasculature for intravascular introduction of interventional/diagnostic devices | Codman (and Formerly MICRUS) NEUROPATH | 2 | February 20, 2013 | 2013 | Codman & Shurtleff, Inc. | Sterility of device may be compromised due to lack of packaging integrity | DESIGN: Packaging Design/Selection | Codman sent an Urgent Voluntary Product Recall Notification letter dated January 11, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check all inventory to determine if they have any affected product, use the included instructions to report inventory status, return and obtain replacement product. Customers were instructed to contact your local Codman Representative or Stericycle at 1-866-201-9077.For questions regarding this recall call 508-880-8100. | No consumer action necessary | 867 | 867 units | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Belgium, Canada, Chile, Costa, Rica, Czech Republic, Denmark, France, Germany, Great Britian/UK, Ireland, Italy, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Singapore, Slovakia, Spain, and Switzerland. | Open |  |  | 5911 | Not_Computer
Z-0838-2013 | 64246 | condom | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LifeStyles ZERO 10 uber-thin lubricated latex condoms. Mfd. for Ansell Healthcare Products LLC, Dothan, AL 96303.A sheath which completely covers the penis with a closely fitting membrane. The condom is used for contraceptive and for prophylactic purposes (preventing transmission of sexually transmitted infections). | Zero | 2 | February 20, 2013 | 2013 | Ansell Healthcare Products LLC | Out of an abundance of caution, Ansell is implementing a voluntary recall of the Zero Latex Condom (regular) product. The condom has met or exceeded test and quality standards required by the Food and Drug Administration. However, we are experiencing a number of complaints concerning the difficulty in unrolling and subsequent breakage, and so we are asking to have all products returned. | DESIGN: Device Design | Ansell HealthCare Product LLC sent an Recall letter dated January 14, 2013, via registered mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification informs customers about the problem with the affected product and provides instructions to the customers on how to handle affected product. Customers are asked to complete the Response Form and return via fax (1-800) 722-8155.If you have further questions, please contact our Customer Service Department at 1-800-952-9916. | No consumer action necessary | 469134 | 469,134 units | .Nationwide Distribution including the states of AL, AZ, CA, CO, DE, FL, GA, IA, IN, IL, KS, KY, LA, MI, MN, MO.MS, NC, NH, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WI and UT. | Open |  |  | 469134 | Not_Computer
Z-0839-2013 | 64288 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet ELS Cannula 15 FR. Used for extracorporeal membrane oxygenation (ECMO) for the treatment of respiratory insufficiency in infants. | ELS Cannula 15 FR | 2 | February 20, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | ELS 15 Fr Cannula could potentially contain an occluded arterial lumen. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Maquet sent Urgent Device Recall Letters/Recall Acknowledgement Forms to their customers on 12/26/2012. | No consumer action necessary | 257 | 41 units - US Distirbution; 216 units - OUS Distribution | Worldwide Distribution-USA including the states of AZ, CO, FL, TN, and TX and the countries of Australia, Canada, Czech Republic, Dubai/Bahrain/Kuwait/Lebanon/Saudi Arabia/Yemen, France, Germany, Hong Kong, Italy, Japan, Netherlands, Singapore, South Africa, Sweden, Switzerland, Thailand, Turkey, and UK. | Open |  |  | 257 | Not_Computer
Z-0868-2013 | 64215 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, SuperCross Microcatheter, Sterilized using ethylene oxide, Rx Only, Model numbers 5300 - 5307, 5324, 5325, 5329, 5332, 5336, 5340, and 5341.Product Usage:The SuperCross catheter is intended to be used in conjunction with steerable guidewires to access discrete regions of the coronary and/or peripheral vasculature. It may be used to facilitate placement and exchange of guidewires and other interventional devices and to subselectively infuse/deliver diagnostic and therapeutic agents | Percutaneous Catheter | 2 | February 20, 2013 | 2013 | Vascular Solutions, Inc. | Vascular Solutions, Inc has discovered a potential problem with a number of lots of Supercross microcatheters. A potential for surface contamination on aged catheters was observed during in-house testing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Vascular Solutions sent an "Urgent Medical device Recall" letter dated January 28, 29013 to all affected customers. The letter identified the affected product, described the problem and actions to be taken. The letter advised customers to remove the product from their inventory and to complete and return the VSI Account Inventory Form. For questions contact your local Sales Representative. |  | 4731 | 4731 | Worldwide Distribution - USA Nationwide including the states of AL, AZ, AR, CA, CO, CT, DE, CD, FL, GA, HI, ID, IL ,IN, IA, KY, LA, ME, MD, MA, MI,MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT VA, WA, WV, WI, WY and AUSTRIA, CANDA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, ISRAEL,. ITALY, MEXICO, NETHERLANDS, NORWAY, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Terminated | June 25, 2013 | 125 | 4731 | Not_Computer
Z-0866-2013 | 64202 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Merit Medical Systems Inc., Resolve Biliary Locking Drainage Catheter, Models RBC and RBDC, used for drainage of bile within the biliary system. | ReSolve | 2 | February 20, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems is recalling certain lots of ReSolve Biliary Locking Drainage Catheters due to a consumer complaint that a catheter was observed to be in several pieces under patient fluoroscopic visualization. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Merit initiated the recall on January 16, 2013. A voicemail script notification to users and Q&A; document was provided to those sales representative that had received and or distributed product. On January 16, 2013, Merit sales representatives began contacting each physician and/or user facility that had been provided product.Written communication distribution began on 22 January 2013. Affected sales representatives, physicians, and hospitals were provided written identification information, instruction to not distribute or continue placing these devices, and on how to return any unused product if applicable.All product on hand at Merit was inventoried and remains under quarantined status.Further questions please call (801) 208-4301,. |  | 387 | 387 units | Worldwide Distribution - USA including the states of NJ, FL, GA, MO, AZ, UT and the country of Denmark | Open |  |  | 387 | Not_Computer
Z-0840-2013 | 63692 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | STRYKER Sustainability Solutions, AC Power Adaptor for Restep Deep Vein Thrombosis (DVT) System, an external compression system. PUMPKIT (Pump Kit), and RSP159-01 (AC Power Adaptor). | Restep DVT System | 2 | February 20, 2013 | 2013 | Stryker Sustainability Solutions | Reports that the AC Power Adaptor may not operate due to wires coming loose in the unit. | PRODUCTION CONTROLS: Process Control | Stryker Sustainability Solutions sent a Urgent Medical Device Removal Notification letter dated September 5, 2012, to all affected customers. The letter identified the product, problem and the action to be taken by the customer. The firm notified that affected AC Power Adaptors would be replaced with improved Adaptors. Any affected lots would be removed and shipped back to Stryker Sustainability Solutions for final disposition. Customers were instructed to fill out the attached Recall Effectiveness Form even if no affected product was found and fax to Stryker Sustainability Solutions at (480) 809-4957. | No consumer action necessary | 1311 | 1,311 | Distributed in Arizona and New Jersey. | Open |  |  | 1311 | Not_Computer
Z-0870-2013 | 64329 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Hand Controller Sheath, Catalog number AVA 500 HCS/3010903, for cardiac catheterization procedures. The Sterile Sheath is specifically intended for use in the x-ray angiography environment to cover the Medrad Avanta Sterile Hand Controller to prevent cross contamination between patients. | MEDRAD(R) Hand Controller Sheath | 2 | February 21, 2013 | 2013 | Medrad Inc | The seal between the white polyethylene and the clear plastic may be missing on the Hand Controller Sheath package, resulting in a potential breach of the sterility of this accessory. | DESIGN: Packaging Design/Selection | Medrad sent a Recall letter dated November 5, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were instructed to review their inventory and quarantine any packages with an expiration date prior to April, 2017. Customers were also given the option of conducting a visual inspection on site to determine if the product should be returned. Alternatively, the customer has the option of returning the affected product. If the customer has affected product, they are instructed to contact Medrad Customer Support to make arrangements for the return and replacement of the product.For further questions please call our Customer Support Team at 1-877-229-3767. |  | 832175 | 832,175 | Worldwide Distribution including USA nationwide. | Open |  |  | 832175 | Not_Computer
Z-0872-2013 | 62334 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Overhead Rail System, Ceiling Mounted Rail Systems for all Liko Stationary Lifts including Likorall, Multirall and Masterlift SystemsProduct Usage - The Overhead Rail System allows overhead lifts to be mounted to the ceiling allowing versatility for use in lifting situations. It can be used with different Liko ceiling lifts. | Overhead Rail System | 2 | February 21, 2013 | 2013 | Hill-Rom, Inc. | Liko/Hill-Rom to date has received two reports from facilities alleging that the overhead rail system has failed to the extent that the overhead rails have fallen down while in use in high humidity (swimming pool) areas. Analysis of the incidents suggest that the components (hardware) used to secure the rail to the ceiling had corroded and may be unsuitable for environments with an atmosphere con | DESIGN: Device Design | Liko, a Hill-Rom company, sent an URGENT MEDICAL DEVICE RECALL notice dated October 31, 2012 to all affected customers. Letter identified the affected product, problem and actions to be taken until a new design is finalized and the field corrective action implemented. The letter instructed customers to complete and return the Response Form immediately even if you do not have any rail systems in a chlorinated environment. For questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. | No consumer action necessary | 3816 | 3816 in total | Worldwide Distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and CANADA LOCATIONS: AB, BC, MB, NB, NS, ON, and SK. | Open |  |  | 3816 | Not_Computer
Z-0869-2013 | 64185 | accessories, operating-room, table (kit) | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | MAYFIELDÂŽ Composite Series Base Unit Standard, Model (A3101), Rx ONLY, Integra LifeScience Corporation, 4900 Charlemar Drive, Building A, Cincinnati, OH 45227The MAYFIELDÂŽ Base Units are intended to be used to support a patient during diagnostic examination and/or surgical procedures where a rigid support between surgical table and headrest, or skull clamp is necessary, positional freedom is required and where X-ray imaging modalities will be used. | MAYFIELD Composite Series Base Unit Standard | 2 | February 21, 2013 | 2013 | Integra LifeSciences Corp. | Integra determined there is a possibility that the Base Units may break during use. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Integra LifeSciences Corporation sent an Urgent Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product immediately and also remove them from service to prevent any further use of the affected product. Integra would provide an alternative base unit to replace those customers could no longer use. Customers were asked to complete the Recall Acknowledgement and Return Form to ensure they received this information. For assistance customers were instructed to call 1-513-533-7944 or 1-513-533-7946. |  | 96 | 96 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA & WI. Internationally to Belgium. | Terminated | July 10, 2013 | 139 | 96 | Not_Computer
Z-0873-2013 | 64322 | lubricant, personal | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY SENSITIVE TM JELLY Personal Lubricant. The product is packaged into 3 oz tubes, which are then packed into a printed carton. Distributed by McNeil-PPC Inc. | KY SENSITIVE TM JELLY | 2 | February 22, 2013 | 2013 | Mc-NEIL-PPC, Inc. | This product is not cleared for marketing in the US. |  | Johnson and Johnson Healthcare Products sent recall notification letters dated 1/25/2013 to its customers (retailers and wholesalers). The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to return all cases of affected product in their possession. The letter included delivery confirmation and a return packing slip, return authorization label and business reply card. Contact J & J Sales & Logistics at 866-912-8414 for questions regarding the return of affected product. Contact your J & J Sales & Logistics Company Sales Representative for questions about the product. |  | 1185664 | 1,185,664 units | Nationwide Distribution. | Open |  |  | 2361232 | Not_Computer
Z-0876-2013 | 64303 | bracket, metal, orthodontic | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Avex CX Suite, Avex CX 0.22 UL4,5 HK REF #224-017, Product Usage: Avex CX ceramic brackets are formed with 99.9% pure alumina oxide. The indicated use of the Avex CX ceramic bracket is the mechanical movement of natural teeth in conjunction with other components designed to attach to the brackets. | Avex CX Suite | 3 | February 22, 2013 | 2013 | Ultradent Products, Inc. | Orthodontic brackets distributed was incorrectly labeled. | DESIGN: Packaging Design/Selection | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 8 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to return affected products for replacement. Customers with questions can call firm representative, Corey Jaseph at (801) 553-4220. |  | 124 | 124 | USA Nationwide Distribution including the states of CA, AZ, CO and WA. | Terminated | June 07, 2013 | 105 | 124 | Not_Computer
Z-0878-2013 | 64294 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, non-porous, calicum-phosphate | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Versafitcup Impacting RingModel Number: 01.26.10.0133 (46mm), 01.26.10.0134 (48mm), 01.26.10.0135 (50mm), 01.26.10.0136 (52mm), 01.26.10.0137 (54mm), 01.26.10.0138 (56mm), 01.26.10.0139 (58mm), 01.26.10.0140 (60mm), 01.26.10.0141 (62mm), 01.26.10.0142 (64mm).Designed for cementless use in total hip arthroplasty. | Versafitcup Total Hip SystemImpacting Ring | 2 | February 23, 2013 | 2013 | Medacta Usa Inc | Medacta USA is recalling Versafitcup Impacting Ring because the impaction ring may not fit properly in the implant shell in all cases. This could result in surgical delay potentially increasing possible complications under anesthesia, having to ream up to a larger shell size leaving less bone than planned and repositiong/removal of the shell after initial impaction. | PRODUCTION CONTROLS: Process Control | Medacta USA sent notification letters dated 1/15/10 to Sales Representatives for customers that purchased the Versafitcup Impacting Ring. The letter informed the customers of the problem identified and the actions to be taken. Customers are informed that Medacta USA personnel will contact them by telephone to provide them with a RMA number. Customers are informed that they will receive a new set of redesigned/reworked impacting rings and a second near the end of the January 2010. Customers with questions are instructed to contact the firm at (805) 437-7085. |  | 293 | 293 units | Nationwide Distribution including the states of ID, CO, TX, WI, NY, CT, WA, CA, UT, and AZ. | Terminated | February 23, 2013 | 0 | 293 | Not_Computer
Z-0890-2013 | 64316 | dressing, wound, drug | FRO |  |  | General & Plastic Surgery | 510(k) | Atrapro Antipruritic Hydrogel, 1.5 oz gel tube, 4 Net wt/ 1.5 oz..RX only, Professional Sample, Not for Sale.Pre-printed tube: MS31071,Shelf carton: PL51091-08Atrapro Antipruritic HydroGel is indicated to manage and relieve the burning, itching and pain experienced with various types of dermatoses, including radiation dermatitis and atopic dermatitis. | Atrapro Antipruritic Hydrogel, 1.5 oz gel tube | 2 | February 26, 2013 | 2013 | Oculus Innovative Sciences Inc | Out of specification result - stability test failure of free available chlorine (FAC) which functions as a preservative for the hydrogel. The 13 month test result obtained was 4 ppm FAC and the stability specification is | PRODUCTION CONTROLS: Packaging Process Control | Oculus sent an Urgent Medical Device Recall letter dated February 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue distribution and segregate the affected product. Immediately remove the affected lots of Recall product from the inventory of sales representative, sub contracted distribution and or fulfillment centers. Segregate the affected product in a secure location for return or destruction. If the product was further distributed customers were instructed to notify healthcare professionals of this recall. Customers were also instructed to completed and return the Business Reply Card and fax to 1-415-462-5163. Once the firm receives the affected product the customer would be contacted by an Oculus Representative and provided a Returns Goods Authorization (RGA) number.For questions customers were instructed to call 1-707-559-2445.For questions regarding this recall 707-559-7191, ext 14. |  | 56567.5 | 56,566 1.5 oz tubes | Nationwide Distribution - One distributor in PA | Terminated | November 25, 2013 | 272 | 56567.5 | Not_Computer
Z-0879-2013 | 64250 | probe, radiofrequency lesion | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | BMC RF Cannula Curved Sharp RadiOpaque, Non-Pyrogenic, Active Tip: 10mm, Gauge 16 Ga, Length: 100mm. The BMC RF Cannula is sealed in a Tyvek pouch (primary packaging) and then placed into a white shelf-box (secondary packaging) with the instructions for use. A predetermined number of white shelf-boxes are then placed into shipping cartons (tertiary packaging). The BMC RF Cannula is intended for use in radiofrequency heat lesion procedures for relief of pain. | BMC RF Cannula | 2 | February 26, 2013 | 2013 | Baylis Medical Company Inc. | A report of an event was received indicating that upon opening a box containing 10 sterile cannula by a nurse, the pouch package of 5 of the 10 individually packaged cannula had not been sealed along one edge. Thus the pouch was open and the device could fall out of the pouch and on to the floor when retrieved from the shelf box. The cannulas were not used on any patients, as the unsealed packages | PRODUCTION CONTROLS: Packaging | A recall was initiated on 10/04/2012 and a letter was sent out to the distributor with a Fax-back Acknowledgment Form. The letter described the issue and provided recommendations and a contact number. |  | 1780 | 1780 | International distribution to the country of Australia. | Terminated | May 23, 2013 | 86 | 1780 | Not_Computer
Z-0886-2013 | 64357 | coagulator-cutter, endoscopic, unipolar (and accessories) | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(R) Jarit(R) Monopolar Cable (Catalogue No. 600-290). For use in endoscopic surgery. For use when a rigid endoscopic instrument for grasping and/or dissecting if soft tissue is determined to be appropriate by the surgeon. Monopolar electrosurgical current can be used for coagulation and/or cutting. | Integra(R) Jarit(R) Monopolar Cable | 2 | February 26, 2013 | 2013 | Integra LifeSciences Corp. | Addendum to the Instructions for Use.The addendum emphasizes the warnings and precautions already contained in the device's IFU and also provides additional warnings and precautions that are applicable to the use of electrosurgical equipment in health care facility environments. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra sent an Urgent Medical Device Correction Field Safety Notice letter dated February 12, 2013 and a copy of the Instructions for Use addendum to all affected customers. The letter identified the product, nature of the issue and actions to be taken. Customers were requested to examine their inventory for the affected product and include a copy of the addendum to that inventory. Customers were instructed to complete and return the Acknowledgement Return Form and provide a copy of the IFU addendum to any of their customers that had purchased the Jarit(R0 Monopolar Cable. For questions call Integra Surgical Customer Service at 800-431-1123. |  | 22500 | 22,500 | Worldwide Distribution - USA Nationwide including Canada. | Terminated | December 24, 2013 | 301 | 22500 | Not_Computer
Z-0887-2013 | 64407 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Aesculap Bipolar Generator Foot Pedal, Catalog No. GK226, for electrosurgical cutting and coagulation. | Aesculap(R) Bipolar Coagulator/Foot Pedal | 2 | February 26, 2013 | 2013 | Aesculap, Inc. | The bipolar energy did not stop after release of the foot pedal. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Aesculap AG sent an Important Recall Notification letter dated January 21, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory and return for a replacement product or credit. The letter included an inventory sheet to be completed by the customer and returned to the recalling firm. For question call 610-984-9265 or 610-984-9291. |  | 59 | 59 | USA Nationwide Distribution including the states of: AR, CA, FL, IN, IA, KS, MD, MN, MO, OR, TN, TX, VA and WA. | Terminated | January 08, 2014 | 316 | 59 | Not_Computer
Z-0889-2013 | 64228 | test, immunity, cell mediated, mycobacterium tuberculosis | NCD |  |  | Microbiology |  | Vacuette Cellestis 1 ml QuantiFERONÂŽ TB ANTIGEN (HA) High Altitude: Shake 10x; REF 454413. HE0396. greiner bio-one; 13x75 50pcs.,1200 pcs in total. Product Usage: QuantiFERONÂŽ-TB Gold (QFT) is an in vitro diagnostic test using a peptide cocktail simulation ESAT-6, CFP-10 and TB7.7 (p4) proteins to stimulate cells in heparinized whole blood. | QuantiFERON TB Gold TB High Altitude (HA)Blood Collection Tubes | 2 | February 26, 2013 | 2013 | Cellestis Inc | Cellestis is recalling QuantiFERONÂżÂż - TB Gold TB Blood Collection Tubes and HA TB Blood Collection Tubes due to potential contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Cellestis, inc., sent a "Notification of market withdrawal of QuantiFERONÂŽ" letter dated September 27, 2012 to its domestic and international customers. The letter described the product, problem and actions to be taken. The customers were informed to discontinue use of these products and any product sequestered such that it cannot be used or distributed. Cellestis will provide replacement materials, if and when a retest is required.Please do not hesitate to contact your local QuantiFERON sales representative or customer service representative at +613 8527 3500 or 661-775-7480 should you have any questions or concerns. |  | 29690 | 29,690 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Bosnia and Herzegovina, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Ireland, Israel, Jordan, Luxembourg, Mexico, Netherlands, New Zealand, Peru, Poland, Saudi Arabia, Slovakia, Slovenia, South Africa, Switzerland, and Turkey. | Terminated | April 25, 2013 | 58 | 339090 | Not_Computer
Z-0891-2013 | 64434 | scalpel, one-piece | GDX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Technocut Plus, Sterile Single Use Safety Scalpel, Distributed by: Myco Medical. | Technocut Plus,Sterile Single Use Safety Scalpel. | 2 | February 27, 2013 | 2013 | Myco Medical Supplies Inc | Pouch labeling/printing errors and defective packaging resulting in possible compromised sterility. | PRODUCTION CONTROLS: Packaging | MYCO Medical sent an Advisory Notice dated February 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers wsere instructed to review their inventory for the affected product and immediately quarantine the affected product per their internal procedures. Customers were also instructed to complete the attached Material for Return form and return to MYCO via email QA@mycomedical.com or fax to 919-800-3919. Customers were have the returned material packaged and ready for UPS to pick up at their location. If product was shipped to other customers, customers were asked to contact a MYCO Customer Service Representative with the additional customer's contact and ship-to information. Customers with questions were instructed to call 919-460-2535.For questions regarding this recall call 919-460-2535, ext 105. |  | 110768 | 110,768 scalpels in individual pouches | Worldwide Distribution - USA including CA, FL, IL, MI, MO, NY, OH, PA, SC, TN, WI and Internationally to Canada | Terminated | July 09, 2013 | 132 | 110768 | Not_Computer
Z-0895-2013 | 62804 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthofix ISKD Limb Lengthener. ISKD - Intramedullary Skeletal Kinetic Distractor. The ISKD is intended for limb lengthening of the tibia and femur. | Orthofix ISKD Limb Lengthener Sterile Rx Only | 2 | February 28, 2013 | 2013 | Orthofix, Inc | There is a possibility that the ISKD limb lengthener may stop distracting post-operatively during treatment, which may result in premature bone consolidation (limb not achieving the desired length) leading to revision surgery to remove and/or replace the device. | DESIGN: Device Design | Orthofix sent an Urgent Product Recall Notification letter dated August 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue distribution of devices, identify the affected and remove from their inventory, and returned to the firm. Credit will be issued upon receipt of the returned product. Customers were asked to take the actions listed in the enclosed Product Return Instructions immediately. Customers with questions should contact Orthofix Customers Service by telephone at 800-266-3349 or 214-937-2047.For questions regarding this recall call 214-937-2509. | Contact the recalling firm for information | 1224 | 1224 | Worldwide Distribution - USA (nationwide) Internationally to Australia, Hong Kong, Saudi Arabia, Korea, Turkey, Greece, Finland, South Africa, Spain, Singapore, Switzerland, Brazil, Germany, France, Japan, Netherlands, New Zealand, Israel, Canada and Mexico | Terminated | May 17, 2013 | 78 | 1224 | Not_Computer
Z-0893-2013 | 64313 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker XIA 3 Iliac Screwdriver, 3-PCS. The XIA 3 Iliac Screwdriver is used to drive XIA 3 closed head Monoaxial and closed head Polyaxial screws into the Ilium. Orthopedic manual surgical instrument. | XIA 3 Iliac Screwdriver, 3PCS | 2 | February 28, 2013 | 2013 | Stryker Spine | Surgeons have experienced unthreading of the screwdrivers outer shafts during use. | DESIGN: Device Design | Stryker Spine sent a "Urgent Medical Device Removal" letter dated January 14 2013, via FedEx overnight to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine your inventory and hospital inventory locations to identify your product. Please fax a copy of the Customer Response form to (855) 632-9049 to Regulatory Compliance. For further questions please call (201) 760-8298. |  | 46 | 46 units US | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, GA, IA, IL,IN, KS, LA, MA, MD, MI, MO, MN, NC, NY. OH, OK, OR, PA, TN, TX, UT and WI. | Open |  |  | 46 | Not_Computer
Z-0894-2013 | 61554 | stimulator, electrical, implanted, for parkinsonian tremor | MHY |  |  | Neurology | PMA | Advanced Neuromodulation Systems, Brio DBS System implantable neurostimulator, IPG or R-IPG, Model 6788 Implantable Pulse Generator. | Brio DBS System | 2 | February 28, 2013 | 2013 | Advanced Neuromodulation Systems Inc. | Unintended Fluid ingress into the device header has been shown. | DESIGN: Device Design | St. Jude Medical sent an Important Product Field Corrective Action Device Information letter dated March 28, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for any unused product, isolate the affected product for return to St. Jude Medical. Their local representatiave would collect the unused inventory. Any unused product would be replaced with product once inventory becomes available. For patients with implanted product and no degradation of therapy, it is recommended dthat the devices not be unnecesssarily explanted. For patient with implanted product who experience diminished or loss of therapy customers were advised to follow the steps provided in the letter. Customers were asked to forward the notification to any institution in the event one or more patients or products could be potentially affected by this action. Customers wsere asked to maintain a record of thenotice along with the recommendations to ensure the effectiveness of the communication.For questions regarding this recall call 972-309-2154. | Contact the recalling firm for information | 388 | 388 units | Worldwide Distribution - 53 consignees out of the US. | Terminated | August 15, 2013 | 168 | 388 | Not_Computer
Z-0899-2013 | 64389 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips MX 16-Slice CT Scanner System, a whole-body computed tomography X-ray system. | Philips MX 16Slice CT Scanner System | 2 | February 28, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The four screws that secure the patient table top to the carrier became loose, which caused the table top to detach. |  | Philips Healthcare sent an Urgent - Medical Device Correction letter dated December 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to have BioMedical Engineering personnel inspect the patient table top screws that secure it to the carrier. Customers were instructed that If it is found that the table top screws are improperly seated, they should stop using the system and contact their local Philips Field Service Engineer for repair. A Philips Field Service Engineer will schedule a time to inspect and repair all the affected systems. For further information customers were instructed to contact their Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 and follow the recorded menu option to reach a Customer Solutions. For questions regarding this recall call 800-722-9377, ext 5. |  | 412 | 412 Units | Worldwide Distribution - USA including AL, AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, NJ, NV, NY, OH, PA, PR and VA. | Terminated | December 03, 2013 | 278 | 412 | Not_Computer
Z-0896-2013 | 64265 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Firebird Spinal Fixation System, Parallel Rod Connector (52-6800) and Axial Rod Connector (52-6700). Product Usage:The connector assemblies allow for rod components to be connected side-to-side, or end-to-end and eliminate the need of rebuilding the entire system with a longer rod in order to include the new vertebral levels. | Rod Connector Assembly | 2 | February 28, 2013 | 2013 | Orthofix, Inc | There is a possibility that the Set Screw (p/n 90-2923) sub-component of the Parallel Rod Connector Assembly (p/n 52-6800) & Axial Rod Connector (p/n 52-6700) may crack during final tightening. Orthofix Inc received six (6) customer complaints related to set screw cracking between November 15, 2010 and December 2012. | DESIGN: Device Design | Orthofix sent an Urgent Medical Device Recall Notification dated February 1, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to identify and remove the affected products from inventory immediately and follow the actions listed in the enclosed Product Return Instructions. For questions or concerns contact Regulatory Affairs at 214-937-2100. |  | 565 | 565 total (287 PN 52-6800 and 278 PN 52- 6700) | Worldwidewide Distribution - USA Nationwide including Puerto Rico and the countries of: Australia, Lebanon, Germany, Spain, Sweden and Israel. | Terminated | July 01, 2013 | 123 | 565 | Not_Computer
Z-0898-2013 | 64235 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | REF CMS-CT6T225, 6F Triple Lumen CT PICC with M.S.T. ComponentsProduct Usage: Power injectable infusion PICC insertion kits | Churchill Medical Systems, Inc. a Vygon Company REF CMSCT6T225, 6F Triple Lumen CT PICC with M.S.T. | 2 | February 28, 2013 | 2013 | Medical Components, Inc dba MedComp | The affected product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medcomp sent a notification letter dated December 13, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns call 215-256-4201. | Contact the recalling firm for information | 785 | 785 | USA Nationwide Distribution including the state of PA. | Terminated | July 31, 2013 | 153 | 785 | Not_Computer
Z-0902-2013 | 64383 | console, heart-lung machine, cardiopulmonary bypass | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Century Console Base, Model #043625051b. Intended for use as a platform for other ancillary equipment and disposable products in cardiopulmonary surgical procedures. | Century Console Base | 2 | March 01, 2013 | 2013 | Salyer Prn Biomedical | Concerns that the addition of the UPS Battery Backup to the Century Console Base Shelf did not comply with the FDA premarket notification regulations. | PREMARKET APPROVAL: No Marketing Application | An Urgent: Medical Device Recall letter, dated 1/31/13, was sent to all the customers that received the Century Console Base Shelf option with the Linear/1.5k 115v UPS Battery Backup for the Century Heart Lung Machine. The letter informed the customers that Salyer PRN Biomedical will have a Salyer Biomedical technician return the customers' Century Console Bases to their original configuration without the Shelf option and restore the Linear/1.5k UPS Battery Backup to its original condition as a stand-alone unit with its own wheel base for ease of use. A refund or credit will be issued for the return of the product. Customers are instructed to contact Jeff Poland at (314) 497-5232. |  | 9 | 9 units | Distributed in the states of MA,CA, and WI. | Terminated | April 15, 2013 | 45 | 9 | Not_Computer
Z-0904-2013 | 64195 | screw, fixation, intraosseous | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Synthes Multi-Vector Distractor, Multi-axial Mandible Distractor (Right), Part Number 487.966.For mandibular bone lengthening. | Synthes MultiVector Distractor, Multiaxial Mandible Distractor (Right), Part Number 487.966 | 2 | March 01, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that the sheer pin on the Multi-Vector Distractor may not be made according to product specifications. | DESIGN: Device Design | The firm, Synthes (USA) sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firmat 610-719-5450 for questions regarding this notice. |  | 24 | 24 | Nationwide Distribution-including the states of CO, LA, MD, MN, NY, PA, RI, TN, TX and UT. | Open |  |  | 44 | Not_Computer
Z-0892-2013 | 64136 | electrode, cortical | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Ad-Tech Medical Instrument Corporation Macro Micro Subdural Electrode. For temporary ( | Subdural Electrode | 1 | March 02, 2013 | 2013 | Ad-Tech Medical Instrument Corporation | The microelectrodes are not flush with the silastic substrate, may contain sharp edges capable of abrading the cortical surface, and may have remaining pieces which may fracture and remain in the cortical tissue with the electrode is explanted. Recall expanded on 5/2/2013 to include additional Macro Micro Subdural Electrodes. | DESIGN: Device Design | AD-TECH Medical Instrument Corporation sent an "Urgent Medical Device Recall" letter dated December 18, 2012 to all affected consignees. The letter described the Reason for the Voluntary Recall, Risk to Health, Actions to be taken by the Customer, Product and Distribution Information, Type of Action by Ad-Tech Medical Instrument Corporation and Other Information. Consignees were instructed to quarantine the affected product and contact the firm for a Return Material Authorization number. If the product have been further distributed consignees were advised to notify their customers of the recall. The letter requested consignees to return the Acknowledgement and Receipt Form.For questions contact your Ad-Tech Clinical Specialist.Expanded recall letters dated May 2, 2013 were sent out April 29, 2013. The letters were formatted as the initial ones sent on 12/18/2013. |  | 144 | 144 (103 US; 41 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, GA, LA, MA, MD, MI, MN, NH, PA, RI, TX, UT, WI, NY IA, WI and the countries of: Australia, CANADA, ENGLAND, JAPAN, KOREA, SPAIN, FRANCE, and SWITZERLAND. | Open |  |  | 144 | Not_Computer
Z-0908-2013 | 64417 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DPL Single Stage Venous Cannula with Metal Tip, used in cardiopulmonary bypass procedures. Model numbers: 67312; 67314; CB67314; 67316; 67318; CB67318; 67320; 69312; 69314; 69316; 69318; 69320; 69322; 69324; 69328; 69331. | Medtronic DPL | 2 | March 05, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Potential for uncharacteristically rough metal edges to occur along the inner diameter of the metal tip side ports of the cannula. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent Medical Device Recall" letter dated February 19, 2013. The letter described the problem and the product involved in the recall. Advised consignees to quarantine affected product, and return the product to Medtronic. Requested consignees the complete and return the attached Customer confirmation certificate. If product has been forwarded to another facility, requested consignees to notify the facility of the issue. |  | 6980 | 13,691 (6711 USA; 6980 OUS) | Worldwide distribution, including USA nationwide, Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Kazakhstan, Korea, Libya, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Viet Nam. | Terminated | November 21, 2013 | 261 | 13691 | Not_Computer
Z-0912-2013 | 64404 | surgical device, for ablation of cardiac tissue | OCL | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | Cardiovascular | 510(k) | Atricure Inc., Max1 Isolator Transpolar Pen. Model A000163. Used to ablate cardiac tissues during cardiac surgery using radiofrequency energy when connected to the AtriCure Ablation and Sensing Unit or for temporary cardiac pacing, sensing, recording, and stimulation during the evaluation of cardiac arrhythmias during surgery when connected to a temporary external cardiac pacemaker or recording device. | Isolator Linear Pen, Electrosurgical cutting and coagulation device and accessories | 2 | March 06, 2013 | 2013 | Atricure Inc | Products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. | TRAINING: Employee Error | AtriCure sent a Urgent Medical Device Recall/Correction letter dated January 14, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.AtriCure has identified that products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. If the MAX1 is used as intended there is no risk to the patient as a result of this issue.The correct instructions for Use for the MAX1 Pen is included with this notification. AtriCure will replace any affected product if requested by the facility.Please review and fill out the attached " Acknowledgement Form" and either fax to AtriCure at (513) 644-4753. If you need to contact customer service, their number is 1-866-349-2342. | Contact the recalling firm for information | 5 | 5 devices | Nationwide Distribution including the states of MD, PA and TX. | Terminated | June 11, 2013 | 97 | 5 | Not_Computer
Z-0914-2013 | 64464 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Lifeguard Infusion Set with Smartsite, 20Ga, 1, CLGY-2010Huber type needle set used for the infusion of medical fluids into implantable ports. | Lifeguard Infusion Set with Smartsite, 20Ga, 1 | 2 | March 06, 2013 | 2013 | Churchill Medical Systems, Inc. | The label on the device states that the item number is CLGY-2210 when the item number on the device label should read CLGY-2010. This is a 20Ga Huber needle set; however, the product label states the pouch contains a 22Ga set. |  | Vycon sent an Urgent Medical Device Recall letter dated January 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check their stock for the affected product and immediately cease use and distribution and quarantine all affected product immediately. Count and document their affected inventory on the Recall Acknowledgement and Inventory Return Form. Customers were instructed to fax or email the form to Vygon including customer contact information. All affected product would be credited or replaced at no cost to the customer. Customers with questions were instructed to call 1-800-473-5414, or by e-mail at customerservice@vygonus.com.For questions regarding this recall call 603-743-5988. |  | 2340 | 2340 | Nationwide Distribution including MO, IL, FL, PA, VA, MO, NY, NE, NH, and OH. | Terminated | September 10, 2013 | 188 | 2340 | Not_Computer
Z-0917-2013 | 64367 | motor, surgical instrument, ac-powered | GEY | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | D3000I Advantage Drive System, 230v, surgical instrument motor. Product Usage: The Advantage Drive Systems function as a powered instrument system consisting of handpieces and accessories to perform cutting of soft tissue and bone | Advantage Drive System | 3 | March 06, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The Membrane Switch Panel (referred to as "front panel" of the component part number A34-003-207, used on both the D3000 and D3000I Advantage Drive System, contained the incorrect symbology per third edition IEC60606-01 Medical Electrical Equipment: General requirements for basic safety and essential performance. |  | Linvatec Corp sent an "URGENT: Medical Device Recall Notification" letter dated January 24, 2013 to its customers after initial phone notifications. The letter describes the affected product, problem and actions to be taken. Customers were instructed to immediate check facility's inventory for the affected product, use the enclosed FedEx air-bill when returning the product, and complete and return to the Recall Reply Form. For questions contact your Customer Service department at 800-535-8536 or fax at 727-319-5701 or email Custserv1@linvatec.com. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) including the states of: CA, FL, GA, IL, KS, MA, MD, MO, NE, NY, and TX and the countries of: Canada, Germany, Romania, Spain, Palestine, South Africa, France, Sweden and Singapore | Terminated | July 05, 2013 | 121 | 54 | Not_Computer
Z-0909-2013 | 64425 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Hickman 9.0 French Dual Lumen Catheter with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer system (pediatric product code 0600580; adult product code 0600600).Product Usage: Hickman and Leonard Catheters are designed for long-term vascular access and for use in patients that lack adequate peripheral venous access. | Hickman and Leonard Catheters | 2 | March 06, 2013 | 2013 | Bard Access Systems | Bard Access Systems is recalling two lots of Hickman 9.0 French and one lot of Leonard 10.0 French Dual Lumen Catheters with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer Systems due to the wrong size introducer being included with the kit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Bard sent an Urgent Recall Notification letters dated February 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete the enclosed Reply Form and Inventory Reconciliation Form. For questions contact Bard Access Systems Customer Service at 1-800-290-1689. |  | 195 | 195 units. | USA Nationwide Distribution and the countries Europe, Belgium, Australia, and Latin America. | Terminated | January 03, 2014 | 303 | 268 | Not_Computer
Z-0911-2013 | 64013 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PST Gel and Lithium Heparin (LH); 83 Units; Plus Blood Collection Tubes; Sterile; IVD; REF 367962; 4.5 mL, 13 X 100 mm; BD, Franklin Lakes, NJ USA; Made in USA.Used in hospitals, outpatient clinics and physicians' offices to collect blood from patients for specialized chemistry testing. | BD Vacutainer Plastic Plasma Separator Tube (PST) Plus | 3 | March 06, 2013 | 2013 | Becton Dickinson & Company | BD is conducting a recall of the BD Vacutainer Plus Plastic Plasma Separator Tube (PST) with Light Green/BD Homogard Closure due to a deformed top of the tube reservoir reported in two customer complaints. | PRODUCTION CONTROLS: Equipment Maintenance | BD sent advisory letters via UPS dated 3/5/2012 to customers. The letter identified the affected product and the reason for recall. BD has taken steps to modify the existing manufacturing and assembly processes to help reduce the potential for this condition. Questions regarding this issue should be directed to BD Global Technical Services at 1-800-631-0174 or email at www.bd.com/Vascutainer/contact. | Contact the recalling firm for information | 225 | 225 million units | Nationwide Distribution | Open |  |  | 225000000 | Not_Computer
Z-0918-2013 | 64343 | stapler, surgical | GAG | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | PSX PROXIMATE Skin Staple Extractor,ETHICON ENDO-SURGERY, LLC.For routine skin closure in a wide variety of surgical procedures. | PROXIMATE Skin Staple Extractor | 2 | March 07, 2013 | 2013 | Ethicon Endo-Surgery Inc | Damage to the packaging may have compromised the sterility of the device. | DESIGN: Packaging Design/Selection | On 10/28/2012, the firm sent Urgent: Device Recall, Event 5928 notification letters to their customers. The letters identified the affected product, the reason for the recall, the action to be taken by customers, as well as the actions taken by the firm. Customers are to examine their inventory immediately and remove the affected product. The Business Reply Form should be completed and faxed to Stericycle at 1-877-650-0360 within 3 business days, even if there is no affected product on hand. Customers must complete the "Replacement Purchase Order Number Field" on the Business Reply Form to ensure adequate identification of the return product in their facility. Customers should follow the instructions provided in order to return the product. The information should be shared with all appropriate staff at the customer's facility. EES planned to replace any unopened an unexpired recall product that was returned by December 31, 2012. Questions regarding the issue should be directed to Stericycle at 1-877-650-5412. | Contact the recalling firm for information | 5894975 | 5,894,975 pieces | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, & WY; and, the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, NICOSIA, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PROVINCE OF CHINA, PUERTO RICO, REPUBLIC OF KOREA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA &, YEMEN. | Terminated | June 11, 2013 | 96 | 5894975 | Not_Computer
Z-0922-2013 | 63551 | accessories,arthroscopic | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Endotrig Trigger Release Sterile Micro Hook BladeProduct Usage:The blade is intended to release A-1 pulley in endoscopic trigger release procedures. | Trigger Finger Blades | 2 | March 08, 2013 | 2013 | Instratek, Incorporated | Instratek, Inc.has received multiple complaints of the 1052 Trigger Finger Release Blades which could not be advanced into the cannula through which it is used during trigger finger release procedures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm decided to recall the product and sent notification letters to consignees dated 09/18/2012. |  | 235 | 235 | USA Nationwide Distribution including the states of: AZ, MI, NE, NV, and TX. | Terminated | April 10, 2013 | 33 | 235 | Not_Computer
Z-0920-2013 | 64318 | device, neurosurgical fragmentation and aspiration | LBK |  |  | Neurology | 510(k) | Integra CUSA excel. CEM Nosecone for CUSA EXcel System. The CEM Nosecone is intended for use CUSA Excel/CUSA Excel + Ultrasonic Surgical Aspiration System, CUSA handpiece, and Covidien Force FX Electrosurgical Generator. Used in surgical procedures where combined ultrasonic dissection and electrosurgical coagulation is desired, either simultaneously or independently. | Integra CEM Nosecones | 2 | March 08, 2013 | 2013 | Integra LifeSciences Corp. | There is a potential for erosion of the CUSA EXcel tip used with the Nosecone. Specifically, the potential for tip erosion occurs when using a CUSA CEM Nosecone if the side of the tip, rather than the annulus of the tip, is in contact with the tissue. This erosion can potentially lead to tip breakage. There were no reports of a patient injury associated with these complaints. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Integra sent a Urgent Medical Device Recall Field Safety Notices dated January 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the nature of the issue and how if affects CUSA top erosion/breakage; advises customers and provided a copy of the CEM Nosecone IFU additional Warning; and requess customers to complete and return the Acknowledgement Return Form.Please feel free to contact our service hotline at ( 888 ) 772-7378. | Contact the recalling firm for information | 29450 | USA 6,190, EU 22,972 and ROW 288 units | Worldwide Distribution - USA and the countries of Austria, Australia, Bangladesh, Belgium, Bulgaria, Brazil, Switzerland, Canada, China, Costa Rica, Czech Republic, Germany, Dominican Republic, Egypt, Emirates, Spain, Finland, France, Great Britian, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Japan, Korea (South), Kuwait, Lebanon, Sri Lanka, Latvia, Macedonia, Mexico, Malaysia, Netherlands, Philippines, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Singapore, Slovenia, Thailand, Tunisia, Taiwan, United Arab, Viet Nam and South Africa. | Terminated | August 06, 2013 | 151 | 29450 | Not_Computer
Z-0921-2013 | 64421 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | 3.75mm x 12mm Fixed Cervical Screws. Components of the Genesys Spine Anterior Cervical Plate SystemProduct Usage: is intended for anterior screw fixation to the cervical spine. | Fixed Cervical Screws | 2 | March 08, 2013 | 2013 | Genesys Orthopedic Systems, LLC | Manufacturing residue may be present on the cervicalscrews. Use of this product may result in patient infection, delayed fusion or non-fusion (pseudarthrosis), inflammatory response, or revision surgery. | TRAINING: Employee Error | Genesys Orthopedic Systems, LLC sent a Recall Notification letter dated February 20, 2013 to all affected customers. The letter identified the affected product, problem and necessary actions to be taken. Customers were instructed to immediately return the affected product to the Genesys Spine office and complete and return the enclosed Response Form as soon as possible. For questions call 512-381-7070. |  | 9 | 9 cervical screws | USA Nationwide Distribution including the states of: CA and TX. | Terminated | September 09, 2013 | 185 | 9 | Not_Computer
Z-0919-2013 | 64306 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Command Instrument Broach Handle. Used in Total Hip Arthroplasty. Catalog Number: 6266-0-100 | Stryker Howmedica Osteonics Corp. | 2 | March 08, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Reports of handle fracture and subsequent disassociation of cam assembly subcomponents during use. | PRODUCTION CONTROLS: Process Control | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (201) 972-2100 for questions regarding this issue. | Contact the recalling firm for information | 48 | 48 units | Nationwide Distribution. | Open |  |  | 48 | Not_Computer
Z-0923-2013 | 64477 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Genesis Straight Esthetic 4.0 mm 1mm Cuff AbutmentPart Number: G71200Genesis implant for single stage or two stage surgical procedures all types of of bone in partially or fully edentulous mandibles and maxillae. | Straight Esthetic 4.0 mm 1mm Cuff Abutment | 2 | March 08, 2013 | 2013 | Keystone Dental Inc | Mislabeled: Genesis Straight Esthetic 5.0mm Contour ,3mm cuff labeled as Genesis Straight Esthetic 4.0 mm, 1 mm cuff abutment, | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Keystone issued Market Withdrawal notice via fax on July 5, 2011, to affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the fax-back form to facilitate whether they hae any affected product and indicate whether they would like to receive a replacement of a full refund. Customers ewre asked to contact Keyston'es Customer Relations Department to coordinate shipment of the returned product. For questions regarding this recall call 781-328-3390. |  | 35 | 35 units | Nationwide Distribution including CA, IN, MA, MD, MN, MI, , NH, NY, NC, PA, and WI. | Terminated | March 11, 2013 | 3 | 57 | Not_Computer
Z-0925-2013 | 64354 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF S313152 Straight Acetabular Inserter Handle Ringlocnon-sterileOrthopedic manual surgical instrument | Stright Ringloc Button Latch Inserter Handle | 2 | March 08, 2013 | 2013 | Biomet, Inc. | The Straight Ringloc Button Latch Inserter Handle may have been manufactured at the higher end of the tolerance causing interference with the Ringloc cup. The inserter handle may lock onto the cup during impaction. If alternative inserters are not available this could cause a surgical delay greater than 30 minutes. | PRODUCTION CONTROLS: Process Control | BIOMET sent an Urgent Recall Notice beginning February 13, 2013, to all affected customers. The notice included possible adverse events and directions to locate and remove identified devices by returning a Fax Back Response Form to 574-372-1683 and obtaining an RGA # prior to return of product to Biomet, Inc. 56 East Bell Drive Warsaw, IN 46580. Questions should be directed to 574 372-1570 M-F, 8 a.m.-5.p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 58 | 58 | Worldwide Distribution - No USA distribution -Internationally to Canada, Japan, and Chile | Terminated | September 30, 2013 | 206 | 58 | Not_Computer
Z-0927-2013 | 64348 | reagent, borrelia serological reagent | LSR | Treponema pallidumtreponemal test reagents. | Microbiology | Microbiology | 510(k) | ANTI-BORRELIA (LYME) MICROPLATE EIA Kit, Catalog # 32507 Kit Lot Number 120960. Package Insert Supplement: Calibrator-1206 U/mL, Positive Control-228 - 532 U/mL, and Negative Control- | BIORAD ANTIBORRELIA (LYME) MICROPLATE EIA Kit | 2 | March 08, 2013 | 2013 | Innominata Dba Genbio | Multiple incidents of low-reacting assay plates. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Innominata dba GenBio, sent an "Urgent Product Correction" letter dated February 11, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the listed questions to determine sites risk and take the appropriate actions. For questions regarding this issue, customers may contact Bio-Rad Technical Support at : 1-800-224-6723, option #2, and then option #3 for assistance. | Contact the recalling firm for information | 195 | 195 | Nationwide distribution: CA, IL, MA, MD, ME, NC, ND, NJ, NY, OR, VA and WI. | Terminated | August 27, 2013 | 172 | 195 | Not_Computer
Z-0928-2013 | 64344 | defibrillator, implantable, dual-chamber | MRM |  |  | Cardiovascular |  | Isoline Implantable Defibrillation Lead, Model Isoline 2CR5, 2CR6 and 2CT6.These lead models are indicated for ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life threatening arrhythmias | Implantable Defibrillation Lead | 2 | March 08, 2013 | 2013 | Sorin Group USA, Inc. | Isoline defibrillation leads distributed may have internal insulation breach. | DESIGN: Device Design | Sorin C.R.M., srl sent an Urgent Medical Device Field Safety Notice dated January28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return any unused products by certified mail with signed acknowledgment form, patients should continue to be monitored by their physician field reps during notification visit, and customers should make sure all personnel involved in the management of patients implanted with affected product are aware of the information described in the Field Safety Notice. Customers with questions should contact their local Sorin representative at 1-610-304-9728.For questions regarding this recall call 303-467-6222. |  | 1376 | 1376 Isoline leads | Worldlwide Distribution - USA (Nationwide) and worldwide to Argentina, Australia, Austria, Belgium, Brazil, Canada, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Morocco, The Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Ukraine, Uruguay and the United Kingdom. | Open |  |  | 1376 | Not_Computer
Z-0930-2013 | 64231 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 Automated Breast Volume Scanner (ABVS); Model number 10042541Distributed by Siemens Medical Solutions USA Inc.business Unit: Ultrasound. The S2000 Automated Breast Volume Scanner (ABVS) Ultrasound System for imaging of the breast allows volumetric acquisition of the breast image. | ACUSON S2000 Automated Breast Volume Scanner (ABVS); | 2 | March 09, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | On some ACUSON S2000 ABVS systems, the inside of the transducer pod may contain sharp edges. | DESIGN: Process Design | Siemens sent a Customer Safety Advisory notification letter to Siemens Service Personnel on January 18, 2013 and which would be provided to all affected sites during their visit to correct the problem. The notification identified the product, the problem and the action to be taken. The local Customer Service Engineer will perform the test at their site. If the system passes the test, no further action is required. If the ABVS pod fails the inspection, the Service Engineer will correct the sharp edge.For questions regarding this recall call 650-694-5398. | Contact the recalling firm for information | 214 | 214 affected devices | Worldwide Distribution - USA including WV and OH; Internationally to Australia, Austria, Bahrain, Belgium, Brazil, China (P.R.) Estonia, France, Germany, Greece, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Macedonia, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Portugal, Qatar, Rep of Korea, Romania, Sweden, Taiwan, Thailand, Turkey, Unite Arab Emirates, United Kingdom and Venezuela. | Terminated | July 16, 2013 | 129 | 214 | Not_Computer
Z-0929-2013 | 64015 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic |  | AMO Sensar Intraocular Lenses (IOLs), Model AR40e, Sensar Foldable IOL with OptiEdge Product Usage: These are ultraviolet-absorbing posterior chamber intraocular lenses for use after cataract extraction. | Sensar Foldable IOL with OptiEdge | 2 | March 09, 2013 | 2013 | Abbott Medical Optics Inc (AMO) | Two production orders of these IOLs that contain mislabeled IOLs. | PRODUCTION CONTROLS: Labeling Mix-Ups | AMO sent an Urgent Field Safety Notice letter to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to stop using and remove affected product from inventory. The letter advice customers to complete and return the enclosed Facsimile Form. For questions contact AMO Customer Service Representative. |  | 155 | 155 units | Worldwide Distribution - USA Nationwide including the state of OH and the countries of: Germany, Switzerland, Sweden, Italy, China, France, Singapore, UK, Croatia, Hungary, Turkey and Brazil. | Terminated | January 08, 2014 | 305 | 155 | Not_Computer
Z-0932-2013 | 62993 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Cynch Lordotic T-PLlF Implant. Intervertebral fusion device with bone graft, for use with supplemental fixation. | Cynch Lordotic TPLlF Implant | 2 | March 10, 2013 | 2013 | Spine Smith Partners LP | Implants were incorrectly laser marked as 23mm instead of 28mm. | PRODUCTION CONTROLS: Labeling Mix-Ups | Spine Smith sent a Field Action Notice dated August 24, 2012 to its affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 512-637-2097 for questions concerning this notice. |  | 3 | 3 units | Distributed in Texas. | Terminated | May 17, 2013 | 68 | 3 | Not_Computer
Z-0941-2013 | 64159 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Large Combination Clamp/Large External Fixation System, Part Number 390.005Product Usage: Synthes Large External Fixation is intended for use to provide treatment for long bone and pelvic fractures that require external fixation. | Large Combination Clamp component of Synthes External Fixation System | 2 | March 11, 2013 | 2013 | Synthes USA HQ, Inc. | The L1 hex nuts, used to tighten the Large Combination Clamp, were oversized and the 11 mm sockets and wrenches will not fit on the hex nuts. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated January 11, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product(s), remove them from inventory and return them to Synthes. For questions call 610-719-5450 or email FieldAction@synthes.com. |  | 136 | 136 | USA Nationwide Distribution including the states of AL, AR, AK, CA, FL, ID, KS, LA, MD, MA, MO, NJ, NY, OH, PA, TN, TX, and VA. | Open |  |  | 136 | Not_Computer
Z-0934-2013 | 64274 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ. Cardiovascular X-ray imagingsystems. For use in generating fluoroscopic images of human anatomy. | GE Healthcare, Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ | 2 | March 11, 2013 | 2013 | GE Healthcare, LLC | Exposed high voltage wires inside the cabinet . On a cable located in the system cabinet, two of the shrink tubings covering the 400V live cable leave an area potentially unprotected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | No consignee letter sent out. GE Healthcare's recall strategy was to replace the impacted cable on all affected Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ systems that were distributed. At this time, deployment of this recall is considered complete as of 9/6/2011. |  | 48 | 48 | Worldwide Distribution-USA including the states of FL, IL, IN, NJ, NC, OK, TX, and WI, and the countries of UNITED KINGDOM, TUNISIA, SERBIA, RUSSIAN FEDERATION, NORWAY, NIGERIA, LITHUANIA, LEBANON, KOREA, REPUBLIC OF KOREA, ITALY, INDONESIA, INDIA, FRANCE, FINLAND, COLOMBIA, CHINA, and BRAZIL. | Terminated | January 27, 2014 | 322 | 48 | Not_Computer
Z-0936-2013 | 64326 | drainage catheter with antibiotic | OEI | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Bacterin Elutia, Hubless Coated Silicone Surgical Wound Drain 24 Fr Round Hubless, Full-Fluted Coated with Silver Sulfadiazine STERILE, Box of 10 Part #7118. Indicated for use in surgical wound drainage. | Elutia Hubless Coated Silicone Wound Drain | 2 | March 11, 2013 | 2013 | Bacterin International, Inc. | A portion of a silicone wound drain was found to be cracked. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were informed of the recall via e-mail beginning 1/28/13 with a request for a return response. This was followed by a "Voluntary Product Recall Information" letter dated 2/13/13 that described the product issue and provided recommended actions. The firm asked for all product to be returned to them. | Contact the recalling firm for information | 15 | 15 boxes (10/box) to 4 consignees in US and 81 boxes (10/box) to 1 foreign consignee | 13 US consignees located in CA, CO, MT, TX, UT, and VA. This represents 4 customers and 9 other individual contacts in USA that received marketing samples. Plus 4 Bacterin representives nationwide received samples.Foreign Consingee | Terminated | July 08, 2013 | 119 | 15 | Not_Computer
Z-0937-2013 | 64458 | catheter,intravascular,therapeutic,long-term greater than 30 days | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems, Inc. A Vygon Company - All sizes of CMS-CT6T225 6F Triple Lumen CT PICC with M.S.T. Components, Peripherally Inserted Central Catheter. | Churchill Medical Systems, Inc. A Vygon Company All sizes of CMSCT6T225 6F Triple Lumen CT PICC | 2 | March 11, 2013 | 2013 | Vygon Corporation | Product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Churchill Medical Systems sent a letter to their consignees on December 19, 2012. |  | 210 | 210 | Product was distributed to 4 hospitals (under one hospital network) within the State of Pennsylvania | Terminated | January 14, 2014 | 309 | 210 | Not_Computer
Z-0933-2013 | 64063 | ferrozine (colorimetric) iron binding capacity | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Siemens Total Iron Binding Capacity (IBCT) Flex Reagent Cartridge. An in vitro diagnostic test intended to quantitatively measure Total Iron Binding Capacity in human serum. | Siemens Total Iron Binding Capacity (IBCT) Flex Reagent Cartridge | 3 | March 11, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The analytical sensitivity claim of 6 ug/dL (1.1 umol/L) in the Instructions for Use (IFU) is no longer valid. Siemens is conducting a voluntary corrective action for Dimension Total Iron Binding Capacity (IBCT) to modify the analytical sensitivity claim to 36 ug/dL (6.44 umol/L). |  | Siemens Healthcare sent a "Urgent Medical Device Correction" letter dated Dec 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were notified that theanalytical sensitivity of 6 ug/dL (1.1 umol/L) in the Dimension(R) IBCT FLex(R) Reagent Cartridge instructions for us (IFU) is no longer valid. Please make adjustments to middle wave products or Laboratory Information Systems ( LIS ) as appropriate for your Laboratory.Please complete the attached form and fax it to (302) 631-8467, to indicate you have received this information.If you have any questions please contact your Siemens Technical Solutions Center or your local Siemens technicial support representative at (302) 631-6311. |  | 522 | 19, 503 cartons | Nationwide Distribution | Open |  |  | 19503 | Not_Computer
Z-0938-2013 | 64173 | orthosis, spondylolisthesis spinal fixation | MNH | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Synthes Spine 5.2mm Titanium Click ' X Pedicle Screw Dual Core 40 mm Thread Length, Part Number 487.072.Intended to provide precise and segmental stabilization of the spine in skeletally mature patients. | Synthes Spine Click X Pedicle Screw | 2 | March 11, 2013 | 2013 | Synthes USA HQ, Inc. | Lot number 7542279, of the 45 mm Click 'X Pedicle Screw (Part number 487.072), was incorrectly etched as "Click 'X Pedicle Screw 05.2 L40 TAN violet" although the actual length of the screw is 45 mm. | DESIGN: Process Design | The firm initiated a recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter, dated January 11, 2013, to all consignees. The letter identified the affected product and the issue, as well as requested that consignees examine their inventory for the product(s) with the identified Part and Lot Numbers, remove them from use, and return them to Synthes. Customers were to follow the steps provided in the letter regarding if they had or did not have any of the identified affected product. Questions were to be directed to 610-719-5450 or Fieldaction@synthes.com. |  | 18 | 18 | Distributed in the states of AL, CA, MS, MO, NJ, OR, and TN. | Open |  |  | 18 | Not_Computer
Z-0942-2013 | 64287 | plate, bone | JEY | Bone plate. | Dental | Dental | 510(k) | Synthes Titanium Midface Distractor, MR Conditional, Part number 487.982. For temporary stabilization and gradual lengthening of the cranial or midfacial bones. | Synthes Titanium Midface Distractor | 2 | March 12, 2013 | 2013 | Synthes USA HQ, Inc. | There is the potential for the anterior footplates to not fit onto the Titanium Midface Distractor Assembly as the tab of some distractors may be oversized. | TRAINING: Employee Error | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to all consignees. The letter identified the affected product and the problem. Also, the firm requested consignees to examine their inventory for products(s) with the specified part number, remove them from use, and return them to Synthes. Customers were to follow the instructions provided on what to do if they have or do not have any identified affected product. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 596 | 596 | Worldwide Distribution -- USA, including the states of CA, LA, MA, NY, and PA, and the countries of Canada and Switzerland. | Open |  |  | 596 | Not_Computer
Z-0946-2013 | 64479 | amplifier, physiological signal | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Nicolet EEG Wireless Amplifier 32/64 Channel. Catalog/Part Numbers : 515-015400 for the 32 channel amplifier and 515-015500 for the 64 channel amplifier.Product Usage:The Nicolet Wireless EEG Amplifier is intended to be used as a front end amplifier to acquire, store, and transmit electrophysiological signals in a wired or wireless mode for the Nicolet Neurodiagnostic system. | Nicolet EEG Wireless Amplifier | 2 | March 13, 2013 | 2013 | Natus Neurology Incorporated | Natus Neurology Incorporated is recalling the Nicolet EEG Wireless Amplifier which may overheat and become uncomfortable to touch around the area of the on/off button. The overheating may result in discomfort ranging from too hot up to the potential for a burn. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Natus sent an "Urgent Medical Device Correction and Removal Notification' letter dated February 27, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter advised customers that a Natus Field Service technician will remove and replace the affected product at no charge to the users. The letter also advised if any of the affected products have been transferred to another location to please forward a copy of the letter and to notify Natus. For questions call 800-356-0007 ext. 5129. |  | 114 | 114 (31 USA; 83 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, FL, LA, MD, MS, MO, NJ, NY, PA, TX and the country of AUSTRALIA.BAHRAINBRAZILCANADACountry DENMARKEGYPTGERMANYHONG KONGINDIAINDONESIAIRELANDITALYJAPANKOREAKUWAITMALAYSIAMEXICONETHERLANDSROMANIARUSSIASINGAPORESPAINSWITZERLANDTAIWANTHAILANDTURKEYUNITED KINGDOMURUGUAYVIET NAM | Open |  |  | 114 | Not_Computer
Z-0940-2013 | 64423 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | GOPump Elastometric infusion PumpKit with/ and without accesssories all containing Dual GOPump, part number 510076: One infusion pump comprising two, 150 mL elastomeric reservoirs, each of which connects to its own catheter. Liquid flow rate is labeled at 2 mL/hr per elastomer.The Symbios Disposable Infusion Pump Kit is a disposable, self-contained infusion system utilizing an inflatable elastomeric reservoir to mechanically provide percutaneous infusion of prescribed solutions at a pre-set rate for post-operative pain management | GOPump Elastomeric infusion PumpKit | 1 | March 13, 2013 | 2013 | Symbios Medical Products, LLC | Received 2 complaints out of 458 distributed pumps in this lot, where flow restrictor bead became displaced, which were reported to permit fast flow of contents. |  | Symbios sent an Urgent Medical Device Recall letter dated February 25, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to follow the enclosed 'Steps for Voluntary Field Action' in their entirety. These steps contain information about: I) identifying the product to be returned, 2) obtaining a Returned Goods Authorization number, and 3) returning the recalled product. If product(s) is (are) located, please call 317-450-0145 from 7:00 a.In '0 5:00 p.m. EDT to receive a Returned Goods Authorization umber ("RGA"). Please indicate the RGA number on the outside of the package. Attached is a Verification Form. Customers were instructed to complete and return the form even if they do not have any product to return. Their local sales representative can assist them in completing this form. Distributors of the affected product were instructed to contact Symbios at the number provided and forward the recall notice to all their affected accounts.For questions regarding this recall call 317-225-4447. |  | 458 | 458 | Nationwide Distribution including PA, CA, OH, AL, FL, WV, TN, NJ, NC, TX, CO, NY, TX, MO, MS, SC, and GA. | Terminated | February 13, 2014 | 337 | 458 | Not_Computer
Z-0943-2013 | 64342 | nail, fixation, bone | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162380, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 80 MM, TITANIUM, Smith & Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. | INTERNAL HEX RECON SCREW | 2 | March 13, 2013 | 2013 | Smith & Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label. | PRODUCTION CONTROLS: Labeling Mix-Ups | Smith & Nephew sales representatives and Smith & Nephew international distributors were notified of problem and the recall via email on February 7, 2013. They were instructed to cease use and distribution of the product and quarantine for return. Non responding consignees were notified again on February 28, 2013 by email. Hospitals were notified by letter beginning on February 7, 2013. Smith & Nephew also contacted the surgeons who implanted the screws by letter the week of February 11th.For questions regarding this recall call 978-749-1330. | Contact the recalling firm for information | 69 | 69 units | Worldwide Distribution - USA including NM, TN, TX, NC, CA, MI, KY, MA, NJ, OH, and HI. Internationally to Canada, United Arab Emirates, Mexico, Argentina, Japan, and Chile. | Open |  |  | 159 | Not_Computer
Z-0947-2013 | 64178 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Advantx, Innova2000, Innova3100, Innova4100, Innova2100IQ,Innova3100IQ, and Innova4100IQ equipped with Video MonitorSuspension Cardiovascular X-ray imaging systems. Model numbers 2236709, 2353620 ,2223039-2, and 2270677-2. | GE Healthcare Advantx, Innova2000, Innova3100, Innova4100, Innova2100IQ, | 2 | March 13, 2013 | 2013 | GE Healthcare, LLC | A locking nut might loosen out of the overhead Video Monitor Suspension yoke affecting Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, and Innova 4100IQ Cardiovascular X-ray imaging systems. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | GE Healthcare sent "URGENT MEDICAL DEVICE CORRECTION" letters dated January 18, 2013 and February 11, 2013 to all affected customers. The letters were addressed to Hospital Administrators / Risk Managers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information.The second letter dated February 11, 2013 was sent out to consignees which incorporated an additional model number (2270677-2) which was initially missed. Contact the firm 800-437-1171 for questions regarding these letters. |  | 297 | 297 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, ME, MD, MA, MN, MS,MO, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT , VA, WA, WV, WI and the countries of VENEZUELA, UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SPAIN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, NORWAY, NIGERIA, NICARAGUA, MOROCCO, MALTA, LEBANON, KAZAKHSTAN, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, FRANCE, EGYPT, CHINA, CANADA, BRAZIL, BELGIUM, and NEW ZEALAND. | Terminated | January 16, 2014 | 309 | 297 | Not_Computer
Z-0958-2013 | 64162 | alkaline picrate, colorimetry, creatinine | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Creatinine, Part Numbers OSR6178, OSR6678Product Usage: System reagent for the quantitative determination of Creatinine in human serum and urine on Beckman Coulter AU analyzers. | Creatinine | 2 | March 15, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. | DESIGN: Device Design | Beckman Coulter sent an Urgent Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed the customers that Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. Customers are instructed to complete and return the enclosed response form. Customers with questions about the notification letter are instructed to contact the call center (hotline) at (800) 854-3633 in the US and Canada. |  | 57962 | 57,962 units (50,452 units in the US) | Worldwide Distribution - USA Nationwide and the countries of: Canada, Brazil, Colombia, Mexico, Puerto Rico, Venezuela, Guam, Honduras, Peru, Ecuador, Costa Rica. | Terminated | October 17, 2013 | 216 | 57962 | Not_Computer
Z-0956-2013 | 64194 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Compression Plate/VA-LCP Dorsal Distal Radius Plate, Part Number 02.115.431.Indicated for: Dorsally displaced fractures, Open joint reconstruction, and Corrective osteotomies after distal radius malunion. | Synthes Variable Angle LCP Dorsal Distal Radius System | 2 | March 15, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that one of the VA-LCP Dorsal Distal Radius Plates was incorrectly etched as Part Number 02.115.431 (110 degree) instead of the correct Part Number 02.115.231 (90 degree). | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to consignees. The letter identified the affected product and the issue. Synthes requested customers to examine their inventory for the affected product, remove them from use, and return them to Synthes. Additionally, customers were to follow the steps provided in the event that they did or did not have any affected product on hand. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 4 | 4 | Distributed in the states of IN, OH, PA, and TX. | Open |  |  | 4 | Not_Computer
Z-0954-2013 | 64452 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Spine Polaris 5.5 Button Lock Screw Inserter. Catalog Number 14-500178. The Polaris Spinal System is a non-cervical, pedicle screw spinal fixation device. | Biomet Spine Polaris 5.5 Button Lock Screw Inserter. | 2 | March 15, 2013 | 2013 | Ebi, Llc | Increased risk of failure at the driver tip during screw insertion or removal. |  | Biomet Spine sent a "Urgent Medical Device Recall" notices dated February 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately locate and remove affected items from circulation and return affected product to EBI, LLC.Please follow the instructions on the enclosed "FAX Back Response Form, "and fax a copy of the Response Form to (973) 257-0232, prior to return of the items.Questions related to this notice should be directed to (973) 299-9300 ex. 2322.Monday through Friday, 8am to 5pm. | Contact the recalling firm for information | 59 | 59 units. | Nationwide Distribution including the states of AL, IN, MO, NE TN and TX. | Open |  |  | 59 | Not_Computer
Z-0957-2013 | 63767 | enzyme immunoassay, tracrolimus | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Siemens Dimension(R) Tacrolimus Flex(R) reagent cartridge (DF107). The TACR method is used as an in vitro diagnostic test. | Siemens Dimension(R) Tacrolimus Flex(R) reagent cartridge (DF107) | 2 | March 15, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints of low patient sample recovery with Dimension(R) TACR Flex(R) lot FA3085. Quality Control (QC) materials have not exhibited low recovery of the same magnitude.Internal method regression testing versus LC/MS has shown an average 25% low bias compared to LC/MS across the assay range. Individual patients, particularly samples with lower TACR values | DESIGN: Component Design/Selection | The firm initiated their recall of this product on November 27, 2012 by sending a letter to all consignees. The letter included a recall effectiveness check form to be completed by the consignee and returned to the firm. The firm expanded their recall on January 29, 2013 by sending another letter to consignees. This letter also included a recall effectiveness check form to be completed by the consignee and returned to the firm. | Contact the recalling firm for information | 2976 | 2,976 | Product was distributed within the United States to AL, AZ, CA, DC, FL, LA, MD, MI, MO, NJ, NM, NC, OH, OR, PA, TN, TX, UT, VA, and WA. | Open |  |  | 2976 | Not_Computer
Z-0953-2013 | 64465 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Logic Comprehensive Knee Prosthesis System, Proximal Tibial Spacer and screws, size 3.5, 8mm, Catalog # 02-012-42-3508. | Proximal Tibial Spacer | 2 | March 15, 2013 | 2013 | Exactech, Inc. | A dimensional mismatch was identified which can potentially allow for interference between devices. | DESIGN: Device Design | Exactech Inc. sent a *** Important *** Product Market Recall Notice dated January 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Customers were instructed to immediately cease distribution or use of these products. 2. Extend this information to your accounts that may have this product in their possession.3. Verify if you have any of the subject Optetrak Logic Proxima Tibial Spacer (PTS) in the specified lots.Please complete the attached fax back form and return it to Exactech within the next 5 working days. Thank you for your prompt attention to this matter. For further questions please call (352) 377-1140. |  | 123 | 123 devices | Nationwide Distribution including the states of FL, AL, LA, PA, VA, IL, OH, NY, NC, KS, TN, NJ, CO, GA, CA, MD, MN, IN, & Puerto Rico. | Open |  |  | 123 | Not_Computer
Z-0952-2013 | 64330 | passer, wire, orthopedic | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Extremity Medical - Disposable Orthopedic Surgical Instrument - 1.6 mm Guidewire - Reference Number 101-00006 - Non-Sterile. Contained in Extremity Medical Surgical Kits including CompressX, Hallu.X, IO Fix and XMCP. Used in the selection of screw size when used with appropriate depth gauge. | Extremity Medical Disposable Instrument 1.6 mm Guidewire | 2 | March 15, 2013 | 2013 | Extremity Medical LLC | The 1.6 mm Guidewire is undersized up to 2 mm in length. When used with the depth gauge, the measurement reading will not correspond with the desired screw length. The selected screw size based on the depth gauge reading will be about 2 mm longer than intended. | DESIGN: Device Design | Extremity Medical sent an "URGENT: FIELD SAFETY NOTICE" dated January 10, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return. Contact Customer Service at 1 (973) 588-8980 for questions regarding this notice.On April 10, 2013, Extremity sent notifications to their customers at the surgeon level. They were asked to complete the attached product field action response form. |  | 238 | 238 units | Worldwide Distribution-USA including the states of Florida, Pennsylvania, New Hampshire, Oklahoma, California, Maryland, Arizona, Ohio, North Carolina, Texas, Virginia, Illinois, New York and Tennessee, and the countries of Australia, Belgium, Israel, Netherlands and United Kingdom. | Terminated | July 18, 2013 | 125 | 238 | Not_Computer
Z-0967-2013 | 55155 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ILLICO Straight CP TI Rod, -5.5MM x 110M; REF#73601-110; Qty 1; LBL-005; REV.B; RX ONLY; Alphatec Spine, Inc. The ILLICOÂŽ MIS Posterior Spinal Fixation System is intended to facilitate the surgical correction of non-cervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. | IILICO MIS Posterior Spinal Fixation System | 2 | March 18, 2013 | 2013 | Alphatec Spine, Inc. | The rods were incorrectly laser etched as ILLICOÂż Straight CP Tl Rods, 110mm, when they were in fact manufactured to 100mm in length. | PRODUCTION CONTROLS: Process Control | Alphatec Spine sent an Urgent Medical Device Recall letter dated July 6, 2011, to all affected customers. This notification described the product, problem and actions to be taken. The consignees were instructed to review their inventory to verify the affected device; abstain from use; immediately contact Alphatec Spine Sales Administration (salesadministration@alphatecspine.com) for instructions on how to return this affected product; and complete, sign and return confirmation form to Alphatec Spine, Inc., 5818 El Camino Real, Carlsbad, CA 92008.If you have any questions, call 760-431-9286. |  | 7 | 7 | Worldwide Distribution - USA including CO and NC; Internationally to Belgium. | Terminated | March 20, 2013 | 2 | 7 | Not_Computer
Z-0959-2013 | 64509 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Orthopedic | 510(K) Exempt | Allen Traction Boot, Modular, OT 1000 Traction Device Boot Assembly, Fracture Table Orthopedic Leg Holders, non-invasive medical devices designed to attach to orthopedic fracture tables to aid clinical staff in applying traction to the leg of a patient during surgical procedures. | Ankle Traction Boot | 2 | March 18, 2013 | 2013 | Allen Medical Systems, Inc. | Some units may have been assembled incorrectly which could result in the boot breaking free from the mounting base and allowing the patient's leg to fall. | PRODUCTION CONTROLS: Process Control | Allen Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken. Customers were instructed to: 1. Immediately remove from service your Allen Traction Boot with the serial number(s) XXXXX 2. Contact Allen Medical Customer Service at (800) 433-5774 for an RMA number and instructions to return your Allen Traction Boot for inspection and repair if necessary. There will be no cost to you. 3. Return your Allen Traction Boot to Allen Medical as soon as possible. In order to minimize the inconvenience, we will repair your Allen Traction Boot and return the device to you within 48 hours of receipt. |  | 24 | 24 devices | Worldwide Distribution- USA including the states of TX, NV, KY, GA, OH, PA, NY, AL, WA, IL and the countries of New Zealand and Japan. | Terminated | August 07, 2013 | 142 | 24 | Not_Computer
Z-0969-2013 | 64471 | accessories, blood circuit, hemodialysis | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sterile Access & Return Line Set, 36" Length, Part # UM-872-36. Intended for use with medically prescribed hemodialysis blood tubing. | Sterile Access & Return Line Set | 2 | March 18, 2013 | 2013 | Molded Products Inc | The recalled product leaks at the connection site. | PRODUCTION CONTROLS: Process Control | The firm's one customer was notified of the recall by phone call on 5/8/2012. The customer returned the affected product once the replacement product was received. | Contact the recalling firm for information | 1200 | 1200 sets | Distributed to one customer in Michigan. | Terminated | August 14, 2013 | 149 | 1200 | Not_Computer
Z-0964-2013 | 64456 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp. Stryker Orthopaedics Specialty Triathlon Left and Right A//P Sizer Adjustment Housing PER File K2401. For use in knee prosthesis implantation. NON-STERILE | Triathlon | 2 | March 18, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | The standard Triathlon Femoral Stylus does not fit into the Specialty Sizers and the "R" and "L" markings on the Right Sizer are reversed. | DESIGN: Device Design | A Notification Letter/Product Accountability Form dated 01/28/2013 was sent via FedEx on 01/28/2013 to the sole US branch receiving the affected devices. The branch was instructed to also contact the hospitals in its territory that have the affected product to arrange return of the product. Affected products to be returned to Stryker Orthopaedics, 325 Corporate Drive, Mahway, NJ 07430. |  | 4 | 4 units (total of 2 sets) | 1 Branch (NJ) with one consignee (DE). | Open |  |  | 4 | Not_Computer
Z-0962-2013 | 64072 | cement, bone, vertebroplasty | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | DePuy CONFIDENCE SPINAL CEMENT SYSTEM 11C Plus KITCatalog Number: 283910000.Delivery of CONFIDENCE Spinal Cement, which is indicated for fixation of pathological fractures. | CONFIDENCE SPINAL CEMENT SYSTEM 11C Plus KIT | 2 | March 18, 2013 | 2013 | DePuy Spine, Inc. | Cement mixer used to prepare cement for surgery may not turn properly and may jam or stuck | PRODUCTION CONTROLS: Process Control | DePuy Synthes Spine issued notification letter dated 1/11/13 issued via UPS Next Day Air Saver on 17 January 2013. Accounts requested to review inventory, quarantine any products listed in this recall notice and to complete and return the enclosed Business Reply Form. Contact the firm at 508-977-6606 for questions relating to this notice. |  | 7773 | 7773 | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Armenia, Kazakhstan, Russia, Ukraine, Australia, Austria Belgium, Luxembourg , Brazil, Canada, China, Colombia, Cyprus, Jordan Lebanon, Saudi Arabia, UAE, Czech Republic, Finland ,France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Latvia, Malaysia Mexico, Netherlands, New Zealand, Norway, Sweden, Poland, Portugal Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey, and UK. | Open |  |  | 16013 | Not_Computer
Z-0966-2013 | 64410 | lens, condensing, diagnostic | HJL | Diagnostic condensing lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | 20 Diopter Binocular Indirect Ophthalmoscope (BIO) LensFor visualization during diagnosis and laser therapy of the human retina (fundus) using a binocular indirect ophthalmoscope. | 20 Diopter Binocular Indirect Ophthalmoscope (BIO) Lens | 3 | March 18, 2013 | 2013 | Volk Optical Inc | The defect was with mis-engraving of the lot # directly on the product and its packaging. Mixed up device engraving of V20LC (lot # BB09213) with that from V60C (lot # BB09205). | TRAINING: Employee Error | Volk sent a Voluntary Medical Device Field Correction letter dated October 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were asked to correct their records with the correct lot number and to notify their customers of the voluntary field correction and request their compliance in updating their records. For questions customers were instructed to call 440-510-0800 or 440-942-6161.For questions regarding this recall call 440-942-6161. | Contact the recalling firm for information | 40 | 40 devices | Worldwide Distribution - USA including the states of AZ, CA, GA, MN, NJ, and NM. Internationally to Japan. | Terminated | June 11, 2013 | 85 | 40 | Not_Computer
Z-0970-2013 | 64529 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CR MD1.0 General Set, x-ray imager image plate. CR 10-X Image Plate 35CM X 43CM. An image plate, contained in an X-Ray cassette, used to capture images from X-Ray exposures. The cassette with the image plate, once exposed, is inserted in the Agfa Digitizer CR 10-X where the image plate is taken out of the cassette and then scanned so that the scanned image can be displayed for viewing. | CR 10X Image Plate 35CM X 43CM | 2 | March 19, 2013 | 2013 | AGFA Corp. | The cassettes were shipped with the wrong IP (image plate) size bar code. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On February 7, 2013, AGFA called dealers involved to alert them of the safety issue. Also, an 'URGENT FIELD SAFETY NOTICE" letter dated February 26, 2013 was sent to the customers. The letter described the affected product, problem and actions to be taken. Customers were requested to check their inventories for the affected product and block any identified product within their inventory. Customers were instructed to complete the attached Feedback form and return via FAX or email. For questions call 877-777-2432 Prompt 5. |  | 17 | 17 | USA Nationwide Distribution including the states of: AL, CA, FL, IN, MN, NE, NV, OH, SC and UT. | Terminated | June 06, 2013 | 79 | 17 | Not_Computer
Z-0971-2013 | 63973 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 IM Rod, Non-Sterile and associated Surgical Techniques:Sigma HP Instruments Classic Surgical Technique,0612-89-510Sigma HP Instruments Balanced Surgical Technique, 0612-88-510Sigma HP Instruments Fixed Reference Surgical Technique, 0612-87-510PFC Sigma Knee Systems Primary Cruciate-Retaining & Cruciate-Substituting Procedure Surgical Technique,SP2-007Product Usage:The SP2 IM Rod is used in both primary and revision Sigma knee procedures to align the femoral locating device and distal femoral cutting block. It is also used with the IM tibial resection. This rod is included in both the SP2 and HP instrument kits. | Specialist 2 | 2 | March 19, 2013 | 2013 | DePuy Orthopaedics, Inc. | Potential for the IM rod to break, leaving fragments in the patient. Surgical Techniques guidance has been updated until design changes can be made to reduce likelihood of potential failure. | DESIGN: Device Design | DePuy Orthopaedics sent an Urgent Medical Device Correction Notice letter dated February 5, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete and return the FAX back form to FAX 574-371-4939 or email kseppa@its.jnj.com Questions about device correction information call Post Market Surveillance, 574-372-7333, Sales force questions call Product Specialist, 574-372-7303 and Surgeon questions DePuy's Scientific Information Office at 1-888-554-2482. |  | 2600 | 4000; 2600 in USA | Worldwide Distribution - USA Nationwide. | Terminated | January 06, 2014 | 293 | 2600 | Not_Computer
Z-0955-2013 | 64511 | adaptor, stopcock, manifold, fitting, cardiopulmonary bypass | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, Inc., Guardian II and Guardian II NC Hemostasis Valve, used in cardiopulmonary bypass. Model numbers: FH101, FH101-T, FH101-25, FH101-50, 8210, 8211, FH102, FH102-T, FH102-25, FH102-50, 8215, 8216, 8241, 8242, 8244, 8272 and 8274.Vascular Solutions Zerusa Limited.Intended to maintain hemostasis during the use of diagnostic/interventional devices during diagnostic/interventional procedures. | Guardian II and Guardian II NC Hemostasis Valve | 1 | March 20, 2013 | 2013 | Vascular Solutions, Inc. | Risk that air may be introduced into the device which may lead to an air embolism. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees were sent a Vascular Solutions "Urgent Medical Device Recall" letter dated February 28, 2013. The letter described the problem and the product involved in the recall. In addition, consignees were advised to identify the location of and immediately remove the recalled product from their inventory and secure it. Consignees were requested to complete and return the VSI Account Inventory Form. The letter also stated that VSI Customer Service would contact customers within 1-2 business days after receiving the completed VSI Account Inventory Form. All returned affected product will be replaced upon receipt. Customers are to contact their local Sales Representative if they have any questions or concerns regarding this recall. |  | 79869 | 79,869 | Worldwide Distribution -- USA, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT , VT, VA, WA, WV, and WI, and countries of GREECE, DENMARK, AUSTRIA, UNITED KINGDOM, SWITZERLAND, IRELAND, INDIA, TURKEY, POLAND, HUNGARY, ITALY, EGYPT, KOREA, PORTUGAL HONG KONG, GERMANY, SPAIN, MEXICO, SWEDEN, NORWAY SLOVENIA, NETHERLANDS, and BELGIUM. | Terminated | August 05, 2013 | 138 | 79869 | Not_Computer
Z-0972-2013 | 64285 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ReClaim Reamer Extension; Product code 2975-00-500, used with Reclaim Modular Revision Hip Prosthesis System. The ReClaim Reamer Extension is as an attachment to the distal reamers to allow the surgeon to clear soft tissue while reaming the distal femoral canal. | ReClaim Reamer Extension | 2 | March 20, 2013 | 2013 | DePuy Orthopaedics, Inc. | The tabs on the reamer extension device have broken. DePuy is notifying surgeons of the issue to increase the awareness and to ask surgeons and users to look for tabs before and after surgery to reduce any potential that the tabs are left in the patient. Once the root cause is determined and a design change is implemented, DePuy will conduct a formal swap-out of the affected devices. | DESIGN: Device Design | DePuy Synthes sent a Urgent Information Medical Device Correction Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed upon visual inspection, if the reamer extension is found to have one or more tabs broken, please do not use the reamer extension. The reamer extension will still engage with one tab intact and no clinical implications have been identified with the use of the reamer extension with one tab, but DePuy recommends it be returned and the alternate reamer extension device be used.If both tabs are broken, the reamer extension will not engage and DePuy recommends it be returned and the alternate reamer extension device be used. A reamer extractor is available in the instrument kit for removal of the reamer if necessary.In order to reduce the risk of breaking the tabs during surgery, the Reamer Extension should only be used to drive the Distal Reamer axially within the patients femoral canal. The instrument is NOT intended to be used for side-cutting or lateralizing the femoral canal, therefore, no off-axis force should be applied during distal reaming.The purpose of this communication is to inform you of this device correction and request acknowledgement of receipt of this letter by signing and returning the Surgeon Response Card within 14 days of receiving this notice. Response cards should be faxed to Fax Number: 574-371-4939.For product-related questions, please contact your local DePuy Orthopaedics Sales Representative. For clinical-related questions from surgeons, please have them contact DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am to 5 pm EST.) For questions about device correction information provided, 574-372-7333 (M-F: 8 am to 5 pm EST.). |  | 750 | 750 | Worldwide Distribution - USA including the states of ME, MA, PA, MD, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, TN, NY, AR, NC and VA and the countries of Canada, Chile, Germany, Czech Republic, Israel , Norway, Poland, Slovenia, Australia, New Zealand, South Africa, Belgium, Austria, Switzerland, France, England and County Cork, Ireland. | Terminated | December 30, 2013 | 285 | 750 | Not_Computer
Z-0973-2013 | 64633 | accessories, implant, dental, endosseous | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | PrimaConnex Tapered Implant TC, SD Internal Connection 3.5 x 10mmCatalog Number: l5613K.Implants in support for fixed bridgework. | Keystone PrimaConnex Tapered Implant TC | 2 | March 20, 2013 | 2013 | Keystone Dental Inc | Mislabeled: Cover screw that was in the package was a WD cover screw and not a SD cover screw as labeled | PRODUCTION CONTROLS: Packaging Process Control | Keystone Dental initially called customers on 3/11/13 and forms were completed to document these calls. In addition, a letter sent via Federal Express/or hand delivered to all distributors and customers. Product requested to be returned. | Contact the recalling firm for information | 120 | 120 units | Worldwide Distribution-USA including DC and the states of CA, FL, IN, MA, MN, NJ, NY, NH, PA, WI, and NY, and the countries of Republic of China, Germany, and France. | Terminated | September 26, 2013 | 190 | 120 | Not_Computer
Z-0975-2013 | 64485 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Intraosseous Fixation System (IFS) 1.7 mm, 2.0 mm & 2.5 mm Cannulated Drill Bit.Intended to stabilize and aid in fixation of fractures. | Intraosseous Fixation System | 2 | March 22, 2013 | 2013 | Orthohelix Surgical Designs Inc | The firm was notified by a customer that there was a burr on the inside end of the drill bit resulting in the drill bit not fitting over the associated guide wire. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | OrthoHelix sent a "FIELD REMOVAL NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 330-247-1444 for questions concerning this notice. |  | 9596567 | Lot # 2399041, 97 pcs; Lot # 2399051, 143 pcs; Lot # 2399051R, 4 pcs & Lot # 2399061, 119 pcs | Nationwide Distribution including the states of CA, CO, IL, MI, MD, NC, NM, OH, OK, PA, SD, TX, UT, VA, WA and WI. | Terminated | December 05, 2013 | 258 | 363 | Not_Computer
Z-0976-2013 | 64310 | nucleic acid amplification assay system, group b streptococcus, direct specimen test | NJR | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX PCR Cartridges Catalog #437519, box of 24 cartridges labeled in part:***BD MAX PCR Cartridges, GeneOhm Sciences Canada, Inc. 2555 boul du Parc-Technologique, Quebec, Canada G1P 4S5***BD Max Group B Streptococcus (GBS) Assay is an in vitro diagnostic (IVD) test for detection of GBS DNA in Lim Broth cultures.The BD MAX" PCR Cartridge is a consumable required for the BD MAX" System. | BD Max (tm) GBS Assay | 2 | March 22, 2013 | 2013 | Becton Dickinson & Co. | Cartridges are not sealing properly during PCR testing and may cause false results, indeterminates and/or invalid runs. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS delivery on December 11, 2012. The letter identified the product, the problems, and the action to be taken by the customer. Customers were instructed to take the following actions: Discontinue and discard any affected product. BD would send replacements for the discarded material. Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification. Fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 or Email to RegulatoryComplianceFax@BD.com.. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, contact BD Technical Services Department at 1-800-638-8663. The BD site in Belgium was notified via email.For questions regarding this recall call 410-316-4054. | Contact the recalling firm for information | 27 | 27 boxes/648 cartridges | Worldwide Distribution - USA including CA, FL, IN, IL, KS, NY, OH, OR, PA, TX, and WI and the country of Belgium | Terminated | July 22, 2013 | 122 | 27 | Not_Computer
Z-1001-2013 | 64585 | skin prep tray (kit) | OJU | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Sepp, Compound Benzoin Tincture Packaged and Tincture 3,000 Bulk. Product Usage: Preoperative skin product | Sepp | 3 | March 26, 2013 | 2013 | CareFusion 213, LLC | An internal review of CareFusion's labeling for preoperative skin prep product codes was recently conducted. The labeling of the product codes identified above state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterilized. | MISBRANDING: Labeling False and Misleading | CareFusion sent an Urgent Recall Notice dated March 4, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, quarantine and return the affected product for replacement. For questions call (913) 345-3504. |  | 13749000 | 13,749,000 | USA Nationwide Distribution | Open |  |  | 13749000 | Not_Computer
Z-0995-2013 | 64689 | wrap, sterilization | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | KIMGUARD Container Filters, 7.5in (19cm), Round and 9in. X 9in. (23cm x 23cm), Square.Intended to be used to enclose another medical device that is to be sterilized by a health care provider. | KimberlyClark KIMGUARD Container Filters | 2 | March 26, 2013 | 2013 | Kimberly-Clark Corporation | Some individual container filter units might contain thin areas that may potentially compromise the ability of the filter to maintain a sterile barrier. | DESIGN: Component Design/Selection | All US distributors and clinics were sent an Urgent: Voluntary Product Recall letter via Federal Express, and the Canadian importer/distributor was also notified with the same letter via Federal Express and e-mail. Distributors were instructed to contact their customers by sending a copy of the letter. The letter identified the affected product and stated the reason for the recall. The recall notification includes a Response Sheet that allows the customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The customer is instructed to immediately cease further distribution of the product, physically quarantine and destroy the product, and fax the completed Response Sheet to the Kimberly-Clark Recall Coordinator. |  | 144717 | Code 68507 - 6420 cases; Code 68509 - 1281 cases | Worldwide Distribution -- USA, including the states of AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NV, NJ, NM, NY, NC, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI, and the countries of Canada, Australia, Belgium, Hong Kong, Japan, Portugal, and Singapore. | Terminated | September 16, 2013 | 174 | 7701 | Not_Computer
Z-0998-2013 | 64609 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 System Carriage, P/N 10208.Used as a tool to expedite loading and unloading processed items from a STERRAD 200 sterilizer. | STERRAD 200 Carriage | 2 | March 26, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 System Carriage because it can potentially become dislodged when an excessive load is applied. | DESIGN: Device Design | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter informed the customers that ASP recently determined that the push bar located on the STERRAD 200 System Carriage can potentially become dislodged when an excessive load is applied. Customers with questions are instructed to contact Stericycle at (877) 931-0642. |  | 2457 | 2,457 units | Worldwide Distribution-USA (nationwide) and the countries of Egypt, Germany, Italy, France, Saudi Arabia, UAE, Japan, Singapore, China, Thailand, Hong Kong, Hungary, Israel, Slovania, Poland, Czech Republic, Taiwan, Sweden, Greece, Korea, Ireland, and United Kingdom. | Open |  |  | 2457 | Not_Computer
Z-0996-2013 | 63621 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Model 3500 Syringe Pump, Model #3500-415 and 3500-500.Intended for use in critical care areas for the administration of fluids requiring precisely controlled infusion rates. | Medfusion Model 3500 Syringe Pump | 2 | March 26, 2013 | 2013 | Smiths Medical ASD, Inc. | An issue with the plunger rod seal was identified where the seal had the potential to migrate inside the pump case top during slow rate/ high back pressure infusions. | DESIGN: Device Design | On May 14th, 2010, Smiths Medical contacted the facility that received the pumps via e-mail and requested that they immediately return the pumps. The pumps were returned per Returned Goods Authorization (RGA) and exchanged for replacement pumps manufactured with plunger rod seal G6000077. | Contact the recalling firm for information | 6 | 6 | Distributed in the state of NH. | Terminated | April 11, 2013 | 16 | 6 | Not_Computer
Z-1013-2013 | 64327 | susceptibility test cards, antimicrobial | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux VITEK 2 AST-ST01 Streptococcus Susceptibility Card, REF 410 028.For use with VITEK 2 systems to determine the susceptibility of Streptococcus to antimicrobial agents. | VITEK 2 ASTST01 Streptococcus Susceptibility Card | 2 | March 27, 2013 | 2013 | Biomerieux Inc | The product package did not contain the correct version of the package insert listing the limitations of use and performance requirements. | PRODUCTION CONTROLS: Error in Labeling | bioMerieux sent an "PRODUCT SAFETY CORRECTION NOTICE" dated February 13, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at (800) 682-2666 for questions regarding this notice. |  | 24 | 4/20-card cartons | Distributed only in HI. | Terminated | October 30, 2013 | 217 | 4 | Not_Computer
Z-1006-2013 | 64531 | pessary, vaginal | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Incontinence Dish w/ Support Pessary #0, 55mm. MilexPart Number: MXKPCONDS00A pessary is a medical device that is inserted into the vagina to function as a supportive structure of the uterus and/or bladder and rectum. | Cooper Surgical Incontinence Dish w/ Support Pessary 0, 55mm | 3 | March 27, 2013 | 2013 | Cooper Surgical, Inc. | Pessary mislabeled outer package may not reflect the size of the product contained in the package | PRODUCTION CONTROLS: Packaging Process Control | Cooper Surgical initiateda recall by letter on February 26, 2013, advising accounts to examine inventory and report quantity back to the firm for replacement. For questions customers should call Customer Service at 1-800-243-2974.For questions regarding this recall call 203-601-9825. |  | 45 | 45 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, Colombia, Demarkt, France, Holland, New ZeaJand, andThailand. | Open |  |  | 750 | Not_Computer
Z-1015-2013 | 53991 | cabinet, table and tray, anesthesia | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System Model 2000 (PAS 2000)Product Usage: The PAS 2000 is an anesthesia medication dispensing device that provides medications and inventory management for the anesthesiologist at the point of care within the operating room. | Pyxis | 2 | March 28, 2013 | 2013 | CareFusion Corporation | The recall was initiated because Carefusion has confirmed that when the PAS 2000 drawers due to inaccessability of drawers caused by system down time or lock-up. | DESIGN: Device Design | Carefusion sent a Safety Alert Notification letter dated January 13, 2010 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to ensure back-up procedures were in place in case of system down-time or lock-up, including review of the back-up procedures with staff. The letter included an update to the PAS 2000 User Guide. The letterstates that a CareFusion field representative will be contacting customers to make arrangements to apply warning labels to the PAS 2000 electronics doors. Customers were instructed to complete and return the Response Card included. For questions contact your local field support representative ot the Advocacy department at customer.advocacy@carefusion.com. |  | 2637 | 2637 | Worldwide Distribution - USA (Nationwide) including CANADA. | Terminated | March 28, 2013 | 0 | 2637 | Not_Computer
Z-1016-2013 | 64271 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Dens Instrument Set (Dens Graphic Case), Part Number 687.030 (also known as the Anterior Column Retractor Set)The Synthes Dens instrument Set has been designed to facilitate internal fixation of the odontoid process for the reduction of Type II and shallow Type III fractures, as proposed by Anderson and d'Alonzo. | Synthes Dens Instrument Set (Dens Graphic Case), Part Number 687.030 (also known as the Anterior Col | 2 | March 28, 2013 | 2013 | Synthes USA HQ, Inc. | The Dens Instrument Set has been designed to facilitate the anterior internal fixation of the Dens (odontoid process). There are currently no screws specifically indicated to treat Dens fractures via anterior screw fixation therefore the affected product will be removed. | PREMARKET APPROVAL: No Marketing Application | Synthes sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory, remove the recalled product from use and return it to Synthes. Customers were asked to completed the attached verification Section at the end of the letter by checking the appropriate box indicating that no affected product had been located and fax to 610-251-9005.For questions customers were instructed to call 610-719-5450 or email at FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 148 | 148 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, KS, KY, LA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, TN, TX, VA, VT, WI, WV, and WY. | Open |  |  | 148 | Not_Computer
Z-1022-2013 | 64501 | platelet factor 4 radioimmunoassay | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | Gen-Probe LIFECODES PF4 Enhanced assay, Catalog No. X-HAT45GTI PF4 Enhanced is designed as a solid phase enzyme-linked immunosorbent assay (ELISA). | LIFECODES PF4 Enhanced assay | 2 | March 28, 2013 | 2013 | Genetic Testing Institute,inc | Lifecodes PF4 Enhanced assay, lot number 3000389, is being recalled because of observed inconsistent replicate values when testing the kit Positive Control and patient samples. The variation in reactivity can lead to invalid assays. If an invalid assay is obtained, those results should not be reported and the test should be re-run. An invalid assay would not result in significant harm to a pati | PRODUCTION CONTROLS: Process Control | Hologic/Gen-Probe sent an Urgent Medical Device Recall letter dated February 4, 2013, via FedEx overnight deliver yto all affected customers. The letter identified affected product, stated the issue, and asked that product from the affected lot stop being used. Reported results for all samples tested using the affected lot should be reviewed. The referring physician should be notified that any negative patient result may be erroneous. Remaining inventory should be returned to Gen-Probe. A response form was also asked to be returned. For questions customers were instructed to contact Technical Support at 262-754-1000 or waukeshatechsupport@gen-probe.com.For questions regarding this recall call 262-754-1026. |  | 386 | 386 kits (370 US, 16 OUS) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, China, Slovakia, and Spain. | Open |  |  | 386 | Not_Computer
Z-1018-2013 | 64510 | diaphragm, contraceptive (and accessories) | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Omniflex Diaphragm - Size 65 (Milex Wide Seal Silicone Vaginal Diaphragm) Model MXWF65Intended for use in contraception. | Diaphragm (contraceptive device) | 2 | March 28, 2013 | 2013 | Cooper Surgical, Inc. | Incorrect size was printed on the outer product packaging. | PRODUCTION CONTROLS: Packaging | CSI notified consignees affected by this recall by a Recall Notification letter, dated February 26, 2013, sent via Federal Express with confirmed delivery receipt. The letter identified the affected product, and asked customers to contact the firm's Customer Service Department at 1-800-243-2974, if they had any of the affected lot in their inventory. Customers also have the option of completing the form attached to the letter and faxing it back to 800-262-0105. At CooperSurgical's expense, arrangements will made to replace the affected products that customers have on hand. The firm will provide further instruction and disposition of the product at the time of replacement. |  | 80 | 80 individual diaphragms | Worldwide Distribution -- USA, including the states of NE, MA, FL, TN, CA, ME, SC, CA, NH, OK, TX, NM, MO, NY, PA, and AZ, and countries of Germany, Italy, and Canada. | Open |  |  | 80 | Not_Computer
Z-0982-2013 | 63951 | tray, surgical, instrument | FSM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Needle Counters with Blade Disarmer; 12.2 x 10.5 x 1.7 cm; 30 Count/60 Capacity; Sterile; Foam Strip; Magnetic Area (Cat. No. 4FS60SSA)Needle Counters with Blade Disarmer; 12.2 x 10.5 x 1.7 cm; 30 Count/60 Capacity; Sterile; Double Foam Strips (Cat. No. 4FS61SSA)Needle Counters with Blade Disarmer; 12.2 x 10.5 x 1.7 cm; 30 Count/60 Capacity; Sterile; Double Numbered Magnets (Cat. No. 4MM60SSA)Product Usage:Disposal system for counting of used surgical blades and needles. | Needle Counters | 2 | March 28, 2013 | 2013 | Cardinal Health | Small holes (approximately 1 mm or less) found in device unit level packaging potentially compromising product sterility. | DESIGN: Packaging Design/Selection | Cardinal Health sent an Urgent Product Recall letter dated December 18, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immedaitely stop using affected product and discard any product that may be on hand, return the enclosed acknowledgement form, notify any customers to whom you may have distributed product affected by this recall, and contact the appropriate Customer Service group to arrange for credit of any affected product discarded. For questions contact Cardinal Health Professional Services at 800-292-9332. |  | 20713 | Cat. No. 4FS60SSA: 576; Cat. No. 4FS61SSA: 14,760; Cat. No. 4MM60SSA: 5,184 | Worldwide Distribution - USA (Nationwide) and the countried of Australia, Canada and Japan. | Open |  |  | 403891 | Not_Computer
Z-1019-2013 | 64623 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W. Sterile using ethylene oxide. Product Usage:The Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewires are intended to facilitate catheter placement and exchange during diagnostic or interventional procedures in the aorta, where increased support, distal flexibility, and low surface friction of the guidewire is needed. | Archer 0.035 inch (0.89mm) Super Stiff Guidewire | 2 | March 28, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue involving specific lot numbers of the Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W, where a change in the manufacturing process has been identified as the cause for potential kinking and breaking at the proximal weld end of the outer spring coil during in-vivo use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Medtronic sent an Urgent Medical Device Recall letter dated March 2013 to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to immediately quarantine all affected products and returned to Medtronic. Customers were requested to complete the attached Customer Confirmation Certificate and fax it to Medtronic to the attention of Customer Focused Quality. For question contact your local Medtronic representative or Medtronic Customer Service at 1-888-283-7868. |  | 27682 | 27,682 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Argentina, Armenia, Australia, Austria, Belgium, Bolivia, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Israel, Italy, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Viet Nam. | Terminated | February 06, 2014 | 315 | 27682 | Not_Computer
Z-1017-2013 | 64620 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer Super Stiff Guidewire 0.035 inch (0.89mm), Model ARCSJ200W.Intended to facilitate catheter placement and exchange during procedures in the aorta. | Archer Super Stiff Guidewire | 2 | March 28, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue with a specific lot of the Archer 0.035 inch (0.89mm) Super Stiff Guidewire, model ARCSJ200W, where the PTFE (Teflon) coating has the potential to separate from the wire during in-vivo use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Medtronic sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Medtronic requested any unused guidewires from the lot number to be quarantined, and returned by contacting the customer service phone number at 1-888-283-7868. A Customer Confirmation Certificate was requested to be faxed back to Medtronic at 651-367-0612. |  | 225 | 225 units | Worldwide Distribution-USA including the states of FL, OH, and MN and the countries of Belgium, France Germany, Netherlands, Turkey, and United Kingdom. | Terminated | November 21, 2013 | 238 | 225 | Not_Computer
Z-1023-2013 | 64579 | computer, diagnostic, programmable | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole Eco Cable, Catalog #EM-5050-060Product Usage:The 20-Pole Eco Cable is intended to be used as an interface cable between the LASSO NAV Eco Catheter and CARTO 3 System. | 20Pole Eco Cable | 2 | March 28, 2013 | 2013 | Biosense Webster, Inc. | Biosense is recalling the 20-Pole Eco Cable because it was incorrectly calibrated. When a catheter is connected to an affected to an affected 20-Pole Eco Cable, the catheter will not be recognized by the impacted 54 cable serial numbers within the field. | TRAINING: Employee Error | Biosense Webster sent an Urgent Field Notice Medical Device Letter dated March 5, 2013 to all affected customers. The letter identified the affected product, problem and the actions to be taken. Customers with questions are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Mon-Friday from 7amto 8pm EST. |  | 54 | 54 units | Worldwide Distribution - USA (Nationwide) including the states of: DE, NL, CH, FR, GB, DK. | Terminated | October 17, 2013 | 203 | 54 | Not_Computer
Z-1021-2013 | 64469 | quality control slides | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics KWIK-QC SLIDES ( KWIK-QC Acid Fast Stain Slides and KWIK-QC Mycobacterium Slides).Slide preparations serving as quality control challenges to demonstrate satisfactory performance of staining reagents. | Microbiologics KWIKQC SLIDES | 2 | March 28, 2013 | 2013 | Microbiologics Inc | Microbiologics Inc is initiating a voluntary recall on products KWIK-QC Acid Fast Stain Slide, SL42-10 and KWIK-QC Mycobacterium, SL43-10. These products are being recalled because the control organisms are not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could pote | DESIGN: Device Design | Microbiologists an "Urgent Medical Device Recall" letter dated February 11, 2013. The Letter described the product and the reason for the recall. Advised consignees to "Do not Distribute" and "Return" the product. Urged consignees to "Notify" end users and to complete the attached "Recall Response Form". Contact the form at 320-229-7057 for questions regarding this notice. | Contact the recalling firm for information | 396 | 396 boxes of 10 slides each. | Worldwide Distribution-USA including the states of CA, IN, MS, NY, OH, PA, and SD, and the countries of AUSTRALIA, AUSTRIA, UNITED ARAB EMIRATES, MALAYSIA, TAIWAN, JAPAN, TURKEY, BELGIUM, COLOMBIA, CANADA, SOUTH AFRICA, and VENEZUELA. | Open |  |  | 396 | Not_Computer
Z-1025-2013 | 64075 | marker, skin | FZZ | Skin marker. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Dual Tip Marker, SterileProduct Usage:A writing instrument intended to mark paper and plastic items around the sterile field during surgery. It has two tips. One is black for miscellaneous use. The other is for surgical skin marking. | Dual Tip Marker | 3 | March 29, 2013 | 2013 | Viscot Medical LLC | Dual Tip Marker has inconsistent ink flow on the black permanent ink side. | DESIGN: Device Design | Viscot Medical LLC sent a Customer Notification Letter dated August 22, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product for a replacement. To return the product call 1-800-221-0658 or e-mail jeffl@viscot .com for the RGA. | Contact the recalling firm for information | 458589 | 458,589 markers | USA Nationwide Distribution | Open |  |  | 458589 | Not_Computer
Z-1026-2013 | 64692 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Sterilization System, P/N 10104Product Usage:The STERRAD 100NX Sterilization System is a general purpose, low temperature sterilizer which uses the STERRAD 100NX process to inactivate microorganisms on a board range of medical devices and surgical instruments. | STERRAD 100NX | 2 | March 30, 2013 | 2013 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) has determined that the STERRAD 100NX and STERRAD 200 Sterilization Systems may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. |  | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notification letter dated March 22, 2013 to all customers. The letter identified the affected products, problem and actions to be taken. Customers are instructed to report any complaints or suspected problems with their STERRAD 100NX and/or STERRAD 200 System to ASP Professional Services directly at (888) 783-7723 and select option #2. | Contact the recalling firm for information | 1500 | 1500 units | US Nationwide Distribution | Terminated | September 30, 2013 | 184 | 2538 | Not_Computer
Z-1028-2013 | 64453 | real time nucleic acid amplification system | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Molecular Systems cobas x 480 System 50 mL reagent reservoir, Catalog number 05232732001. Used with the cobasÂŽ CT/NG 4800 System, for in vitro diagnostics. | Roche Molecular Systems cobas 4800 System | 2 | April 01, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Certain lots of Reagent reservoirs, may not have a separation in the lower part of the reservoir. As a consequence, pipetting errors can occur due to uneven distribution of reagent. These pipetting errors may lead to delayed results and wasted reagent. | PRODUCTION CONTROLS: Process Control | Roche sent a Urgent Medical Device Correction letter via UPS to all affected customers ( receipt required) The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to visually inspect affected product for the molding defect and to verify equal distribution of reagent in both chambers of the reservoir after adding reagent. All customers are requested to fill out the FAX Back form to schedule shipment of defective part. Non-responding accounts will be monitored on an ongoing basis and follow-up attempts will be made. Questions and concerns are to be addressed to Roche Molecular Diagnostics Technical Support 24 hours/day 1-800-526-1247. |  | 189 | 189 units | Nationwide Distribution including Puerto Rico and the states of MO, NE, CO, MT, NJ, FL, WA, TX, GA, NY, KS, NH, VA, HI, PA, MI, IA, CA, NC, AZ, UT, LA, IL, TN, ME, MN, KY, CT, OR, NV, AL and NM. | Open |  |  | 189 | Not_Computer
Z-1045-2013 | 64569 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Zimmer NexPosure, retractor portals of the NexPosure MIS Access System, Part Numbers or REF numbers: N1830R80-100, N1830R80-090ÂŽ, N1830R80-080ÂŽ, N1830R80-070ÂŽ, N1830R80-060, N1830R50-100, N1830R50-090, N1830R50-080, ÂŽ N1830R50-070, N1830R50-060, N1830L80-100, N1830L80-090, N1830L80-080, N1830L80-070, N1830L80-060, N1830L50-100, N1830L50-090, N1830L50-080, N1830L50-070, N1830L50-060, ÂŽ N182450-085, N182450-075, N182450-065, N182450-055, N182450-045.ÂŽ Retractor access system that allows placement of posterior spinal fixation implants. | NexPosure Portal System | 2 | April 01, 2013 | 2013 | Zimmer, Inc. | Possibility that the inner pouch may be compromised. The product is packaged in two pouches - a sterile inner pouch within an external pouch. NexPosure Retractors contained in a breached inner pouch have an increased risk of contamination (i.e.; loss of product sterility) if proper aseptic technique is not followed when transferring the inner package into the sterile field. Patient infection may o | DESIGN: Packaging Design/Selection | Consignees were sent on 2/8/13 a Zimmer Spine "Urgent Medical Device Recall" letter dated February 08, 2013. The letter was sent to Surgeons and Facilities using the NexPosure MIS Access System. The letter described the product involved in the recall and the problem. Advised consignees to discontinue the use of the product and to contact their sales representative for removal.A consignee letter dated February 8, 2013 was also sent to Distributors, Sales Representatives, and Distribution Operation Managers Distributing the NexPosureÂŽ MIS Access System. The letter contained instructions if they had the affected product or not and the steps to follow. |  | 58 | 58 | Nationwide Distribution including the states of AZ, CO, DE, FL, ID, LMA, MI, NH, NJ, NM, PA, RI, TX, and VA. | Open |  |  | 58 | Not_Computer
Z-1035-2013 | 64332 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 600i Synchron Access Clinical System; P/N 107405SProduct Usage: The Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access Immunoassay Systems | 2 | April 01, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall due to a defective newly released aluminum incubator belt pulley that may rub against the light shield, which causes the formation of debris. Debris that is caused by the incubator belt pulley rubbing against the light shield can fall into the reaction vessels (RVs) and present a risk of erroneous test results. | DESIGN: Device Design | Beckman Coulter sent an "Urgent Product Correction" letter dated February 12, 2013 via mail and email to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to place your Access system in the Ready or Not Ready mode. Lift the cover (as you do to access the probes) and look at the left side of the incubator unit under the peripump by the aspirate probe, if you see a silver pulley, immediately discontinue use and contact Technical Support also check for QC failures and system check failures; If you see a white pulley, this is the correct part and you can resume normal activity. Beckman will replace the metal pulley with the plastic pulley if you verified the your instrument has the silver pulley. The customers were also instructed to share this information with staff; and complete and return enclosed Response Form within 10 days.If you have any questions regarding this notice, please contact Technical support from the website: http://www.beckmancoulter.com or by phone at: 1-800-854-3633 in the United States and Canada. |  | 956 | 956 units | Worldwide distribution: USA (nationwide) and countries of: Aurstria, Belgium, Bulgaria, Canada, Croatia, Czech Republic, Denmark, Djibouti, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | December 23, 2013 | 266 | 5118 | Not_Computer
Z-1049-2013 | 64463 | phosphomolybdate (colorimetric), inorganic phosphorus | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Siemens Dimension(R) Phosphorus (PHOS) Flex(R) reagent cartridge (DF61)The Phosphorus (PHOS) method used on the Dimension(R) clinical chemistry system is an in vitro diagnostic test intended for the quantitative measurement of phosphorus in human serum, plasma, and urine. | Siemens Dimension(R) Phosphorus (PHOS) Flex(R) reagent cartridge (DF61) | 2 | April 02, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has initiated a voluntary field action for Dimension(R) PHOS Flex(R) Reagent due to reagent interaction causing falsely depressed Creatinine results when EZCR is processed from open wells that are in close proximity to open wells of PHOS reagent. | DESIGN: Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Customer:1. Customers with two or more Dimension systems should perform testing of PHOS and EZCR on separate Dimension systems.2. Customers with only one Dimension system should discontinue use of either PHOS or EZCR on the Dimension system. Please contact your local sales representative to discuss alternate solutions that will best meet the needs of your laboratory.Please complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at (302) 631-8467, to indicate you have received the information. If you have any questions, please contact the Siemens Technical Solutions enter or your local Siemens technical support representative. We apologize for the disruption this situation will cause your laboratory. We are working diligently to resolve this issue with a permanent solution.. |  | 60674 | 60,674 | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Open |  |  | 60674 | Not_Computer
Z-1053-2013 | 64528 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | LEEP REDIKIT Product No. 6060 In a Loop Electrosurgical Excision Procedure, a thin, low-voltage electrified wire loop is used to cut out abnormal cervical tissue. The cartridges of Polocaine and Xylocaine with Epinephrine contained in the LEEP RediKit are local anaesthetic which the physician may choose from during the procedure. | LEEP | 2 | April 03, 2013 | 2013 | Cooper Surgical, Inc. | LEEP RediKit Products (for use in Loop Electrosurgical Excision Procedures) have Xylocaine (Lidocaine) HCL (2%) listed on the tray lid label but this is not included in the contents of the kit, which contained Polocaine (Mepivacaine) HCL (3%) instead (which was not listed on the tray label). | TRAINING: Employee Error | CooperSurgical sent a Recall Notification dated February 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CSI intends to track and quarantine returned product in a segregated location at its manufacturing facility to ensure its separation from acceptable product and prevent its misuse. CSI intends to render used product useless by physical destruction and to rework unused product by correcting the labeling for redistribution.If you have any one of these affected lot numbers in your current inventory, please contact our Customer Service Department at 800-243-2974, or simply complete the attached form and fax to 800-262-0105. Arrangements will be made to replace product in your possession at CooperSurgical's expense. Further instruction and disposition of product will be communicated to you at that time.CooperSurgical deeply regrets the inconvenience that this problem has caused. | Contact the recalling firm for information | 4395 | 4395 kits | Nationwide Distribution including Puerto Rico | Open |  |  | 4395 | Not_Computer
Z-0993-2013 | 64370 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020 Insulin Infusion PumpProduct Usage:This product is indicated for continuous subcutaneous infusion of insulin for the treatment of insulin-requiring diabetes. | Animas(R) 2020 Insulin Infusion Pump | 1 | April 03, 2013 | 2013 | Animas Corporation | Animas has identified a component issue affecting a small supply of the Animas(R) 2020 insulin pumps. The component issue may trigger the pumps to sound a false alarm or warning related to one of the following: - Loss of prime; -Occlusion; or -No cartridge detected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Animas sent an Urgent Recall letter dated December 17, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that Animas will send a replacement pump free of charge, along with instructions for returning their pump. For questions or concerns contact at 866-796-6373. |  | 7157 | 7,157 (494 US pumps affected) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and Puerto Rico and the countries of: Australia, Austria, Bermuda, Canada, Czech Republic, Finland, France, Germany, Ireland, Israel, Italy, Mexico, New Zealand, Norway, Spain and the United Kingdom. | Open |  |  | 7157 | Not_Computer
Z-1056-2013 | 64494 | container, specimen, sterile | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | BD Vacutainer C&S; Transfer Straw Kit C&S; Preservative Plus Urine Tube. 4.0 ml, 13 X 75 mm STERILE REF 364953. Becton Dickinson & Company, Franklin Lakes, NJ USA,The kit consists of a transfer straw and urinalysis tube intended for UA chemistry urine testing. | BD Vacutainer C&S; Transfer Straw Kit C&S; Preservative Plus Urine Tube | 3 | April 04, 2013 | 2013 | Becton Dickinson & Company | The BD Vacutainer Urinalysis Transfer Straw Kit #364991, Lot 1335484 has preprinted polybags which are incorrectly labeled as Catalog #364953. BD advises any Customer who purchased this Kit that the Tube in this Kit is not a C&S; tube. If a Customer relies solely on the polybag labeling and uses the Tube inside the Kit for the purpose of collecting urine for culture and sensitivity, it is possible | TRAINING: Employee Error | BD Diagnostics sent a "Product Advisory Notice" dated February 14, 2013 via UPS Second Day Air to all end user customers. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product.For questions regarding this recall call 201-847-5033. |  | 236400 | 236,400 units | Nationwide Distribution including Arizona, California, Colorado, Georgia, Florida, Illinois, Iowa, Kentucky, Louisiana, Maryland, Massachussets, Michigan, Minnesota, Missouri, New Jersey, New Mexico, New York, North Carolina, Ohio, Oregon, Pennsylvania, Tennessee, Texas, Utah, Virginia, Washington DC, Washington, West Virginia, and Wisconsin. | Terminated | August 06, 2013 | 124 | 236400 | Not_Computer
Z-1054-2013 | 64704 | assembly, thigh/knee/shank/ankle/foot, external | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | LeMaitre Stent Guide: 1102-00 (100 strips); 1102-50 (50 strips); 1102-20 (20 strips).Intended to be placed on the skin to assist during imaging procedures. | LeMaitre Stent Guide | 2 | April 04, 2013 | 2013 | LeMaitre Vascular, Inc. | Devices were not sealed correctly during the manufacturing process, and the sterility of these products have been compromised. | PRODUCTION CONTROLS: Packaging Process Control | LeMaitre sent an "URGENT:VASCUTAPE RADIOPAQUE TAPE DEVICE FIELD SAFETY NOTICE" dated March 20, 2013 to affected customers. The recall letter contains a form that requests customers to return to LeMaitre Vascular, Inc. as a record of notification and reconciliation. The letter provides instructions on how the customer can inspect the defective tape seals and/or return the affected products for replacement/credit. |  | 6970 | 6970 | Worldwide Distribution-USA (nationwide) and the countries of Canada, EU, and Asia. | Open |  |  | 20530 | Not_Computer
Z-1057-2013 | 64593 | catheter, flow directed | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow/Teleflex Medical Berman Angiographic Balloon Catheter and Reverse Berman Angiographic Balloon Catheter.The Arrow Berman(TM) Angiographic Catheter is indicated for use in right heart catheterization for cardiac angiography, and measurement of the right heart and pulmonary artery pressures. It can be sued to inject radiopaque media and measure pressures in any chamber or vessel that can be entered from the right heart. The Arrow Berman(TM) Angiographic Catheter is also indicated for use in Balloon Occlusion Femoral Angiography (BOFA) of the ipsilateral leg for adult patients. The Arrow Reverse Berman(TM) Angiographic Catheter is indicated for balloon occlusion pulmonary angiography. | Teleflex Medical Berman Angiographic Balloon Catheter and Reverse Berman Angiographic Balloon Cathet | 2 | April 04, 2013 | 2013 | Arrow International Inc | The product specifications sheet, part of the Instructions for Use (IFU), for the specified product numbers, reference Renografin-76 contrast medium (viscosity of 8.4 centipoise), which is no longer available. Firm is in the process of updating the labeling for these products. In the meantime, when choosing an alternative contrast medium, health care providers are instructed to select a contrast | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Arrow sent an Urgent Medical Device Customer Notification letter dated March 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Arrow is in the process of updating the labeling; in the meantime they wanted to instruct all customers to select a contrast medium with a viscosity of 8.4 centipoise or lower for use with the pressure and flow settings specified in the current IFU. In addition, they are placing an insert card in kits for the specified product codes alerting them that Renografin-76 is no longer available and to select a contract medium with a viscosity of 8/4 centipose or lower for use with the pressure and flow setting specified in the current IFU.Arrow remains committed to providing high quality, safe and effective products. If you require additional information or clarification regarding this matter, please contact your Sales Representative or Arrow's Customer Service Department at (866) 396-2111. |  | 99627 | 99627ea | Worldwide Distribution - USA including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. and the countries of Australia, Dominican Republic, Uruguay, Ecuador, Chile, Taiwan, Pakistan, Colombia, Argentina, Korea Republic, New Zealand, Guatemala, Thailand, Peru, Bahamas, Philippines, Trinidad and Tobago, Japan, Ireland, India, Singapore, China, Canada and Mexico. | Terminated | December 21, 2013 | 261 | 99627 | Not_Computer
Z-1059-2013 | 64537 | reagent, occult blood | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Test Strips packaged in the OC-Light Manual iFOBT Kit, Catalog No. FOB50.For the rapid, qualitative detection of human hemoglobin in feces. | OC Light Manual iFOBT strips | 2 | April 04, 2013 | 2013 | Polymedco, Inc | The test strips in the affected product lots of the OC Light immunological fecal occult blood test are being recalled because they may not be stable through their labeled shelf-life. There may be a possibility of obtaining false negative results when using the affected lots. | PRODUCTION CONTROLS: Process Control | Polymedco initiated phone calls to clients on 2/12/13. When the contact was made with the client, an URGENT DEVICE RECALL letter and Fax Form were sent by e-mail/fax. Mass mailing of the URGENT DEVICE RECALL letter dated February 12, 2013 and Fax Form was done by First Class USPS on 2/14/13-2/18/13. The recall affects the test strips only. For questions, please contact Polymedco Technical Support at 1-800-431-2123 ext 285. |  | 355350 | 355,350 test strips | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada and Guam. | Open |  |  | 584950 | Not_Computer
Z-1063-2013 | 64420 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Covidien Endo GIA AutoSuture Universal Loading Unit, 30mm-2.0mm Product Code: 030450Applications in abdominal, gynecologic, pediatric anti thoracic surgery for resection, transection and creation of anastomosis. | Covidien Endo GIA | 2 | April 05, 2013 | 2013 | Covidien LP | Single Use Loading Unit contained two staples loaded in each cartridge pocket and may result in difficulty firing and removing the device from the application site, which may require medical intervention | PRODUCTION CONTROLS: Process Control | Covidien notified customers with an Urgent Medical Device Recall notification on February 11, 2013. The letter identified the affected product and the reason for the recall. Accounts were requested to quarantine and discontinue use of the affected devices, and notify Universal Loading Unit 30mm-2.0mm users of this recall. Customers should follow the instructions provided in order to return the affected product. Customers are to complete and return the Product Return Form regardless if they have any affected product or not. For questions regarding the RGA I return process, please contact Covidien Customer Service, M-F, 8 am - 6:30 pm ET at (800) 962-9888, option 1, and then option 2. |  | 6751 | 6,751 units | Worldwide Distribution -- USA, Asia Pacific, Canada, Europe, and Japan. | Open |  |  | 6751 | Not_Computer
Z-1062-2013 | 64200 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flexible Grip which is part of the Synthes Universal Nail System.Indicated for the conventional (non-locking) technique of treating stable diaphyseal fractures. | Synthes Flexible Grip | 2 | April 05, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary recall of the Flexible Grip (Part number 355.28) which is part of the Universal Nail System, due to the potential lack of the ability to thoroughly clean the instrument. | DESIGN: Device Design | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter requested consignees examine inventory, remove the recalled products from use and return them to Synthes. If you have any questions, contact the firm at 610-719-5450. |  | 328 | 328 | Nationwide Distribution including the states of AZ, CA, CO, FL, IA, ID, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, TN, TX, UT, VT, WA, WI, and WV. | Open |  |  | 328 | Not_Computer
Z-1061-2013 | 64655 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Baseplate Impactor Extractor TRIATHLON Instruments. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430 USA A subsidiary of Stryker Corporation. Made in USA. Stryker France ZAC Satolas Green Pusignan Av de Satolas Green 69881 MEYZIEU Cedex FranceThe Baseplate Impactor Extractor is intended to securely attach to the tibial baseplate or baseplate trial for impaction and extraction during surgery. | Stryker Triathlon | 2 | April 05, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received reports from the field indicating that the handle detached from the cam of the Baseplate Impactor/Extractor Assembly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker sent Notification Letters/Product Accountablility Forms dated March 1, 2013 via FedEx to their Branches/Agencies on March 1, 2013.. Notification Letters/Product Recall Acknowledgement Forms dated March 4, 2013, were sent via FedEx to Hospital Risk Management on March 4, 2013. The Notification letters identified the product, the problem, and instructed customers on how to identify affected product; informs customers of the issues and potential hazards with the affected product and instructs them to contact their representative to arrange for return/replacement of any affected product. Customers are also asked to fax back the provided "Urgent Medical Device Removal Notification Acknowledgement Form" to Stryker within 5 days of receipt.Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5523. |  | 871 | 871 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, Japan, Australia, China, New Zealand, France, Italy, Netherlands, Poland, Spain, Sweden, Switzerland and the United Kingdom. | Open |  |  | 871 | Not_Computer
Z-1065-2013 | 64527 | system, balloon, intra-aortic and control | DSP | Intra-aortic balloon and control system | Cardiovascular | Cardiovascular | 510(k) | Arrow International FiberOptix Intra-Aortic Balloon Catheter Kit (IAB) 8 Fr 40 cc;Product Number: IAB-05840-LWS.The IAB is utilized for intra-aortic balloon counter-pulsation therapy. | Arrow Internationa IABP | 2 | April 06, 2013 | 2013 | Arrow International Inc | Potential to have an internal leak at the fill port of the stainless steel tip, which may necessitate removal and replacement ofthe catheter. | PRODUCTION CONTROLS: Process Control | Arrow International issued Urgent Medical Device Recall notification on 11/27/12. The letter identified the affected product and the reason for the recall. Accounts are requested to check their stock for the products included within the scope of this recall. Also, customers should cease use and distribution, and quarantine all affected product immediately. Contact Arrow's Customer Service Department at 800-523-8446 for a Return AuthorizationNumber. Customers are to complete and fax the Recall Acknowledgement & Stock Status Form to the number provided, and return the affected product. If customers have questions or need clarification regarding this issue, they should contact Arrow's Customer Service Department at 1-800-523-8446. |  | 637 | 637 units | Worldwide Distribution -- USA, Belgium, Great Britain, Italy, Kazakhstan, Latvia, Russia, Slovenia, Turkey, and South Africa. | Open |  |  | 825 | Not_Computer
Z-1066-2013 | 64505 | hip prosthesis, semi-constrained, cemented, metal/polymer, + additive, porous, uncemented | OQG | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Cross-over Acetabular Shell and Liner Hip System, Part Numbers SC3458-52, SC3474-52, SC3459-52, and SC3460-52.Intended for use in reconstruction of the articulating surface of the acetabular portion of the hip. | Stelkast Crossover Acetabular Shell and Liner Hip System | 2 | April 07, 2013 | 2013 | Stelkast Co | Firm was made aware of incidents relating to the difficulty of properly engaging and securing the Cross-Over 36mm liner into the Cross-Over shell. | DESIGN: Device Design | The firm initiated their recall of this product on January 24, 2013 by notifying their consignees by telephone informing them of the recall. |  | 254 | 254 | Nationwide Distribution including the states of NV, OK and PA. | Open |  |  | 254 | Not_Computer
Z-1071-2013 | 63622 | restraint, protective | FMQ | Protective restraint. | General Hospital | General Hospital | 510(k) | Posey Biothane Restraint Cuffs, Catalog Number 2900.The device is intended to limit a patient's movement to the extent necessary for treatment or protection of the patient or others. | Wrist restraint, wrist cuff | 2 | April 09, 2013 | 2013 | J T Posey Company | The firm voluntarily recalled Biothane Restraint Cuffs, Catalog No. 2900, that did not meet the required specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. | DESIGN: Process Design | J. T. Posey sent an Medical Device Recall (Removal) letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customer that Posey Company is voluntarily recalling all units of Posey BiothaneQD Cuffs, Catalog Number 2900. All lots of this product are being recalled. Reason for recall was listed:"Posey has identified rivets on the Biothane~ Restraint Cuffs, Catalog No. 2900, that did not meet therequired specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. No patient injuries or deaths have reported."Please immediately discontinue further distribution or use of Posey Biothane~ Cuffs, Catalog Number 2900.Please Complete and return the Acknowledgement and Receipt Form to Posey within 7 days using the provided self-addressed, stamped envelope, by fax to 626-443-5418 or via e-mail to RegulatoryAffairs@Posey.com. Return this form even if you do not have any devices in your facility.Please return directly to Posey ill! units you have of Biothane~ Cuffs, Catalog Number 2900, using the enclosed pre-paid shipping label containing an RMA number. Posey will provide replacement product or credit upon receipt of the product.If you have any further questions please call (626) 443-3143. |  | 1673 | 1673 units | Nationwide Distribution | Open |  |  | 1673 | Not_Computer
Z-1073-2013 | 63804 | test, time, partial thromboplastin | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent.For use in the determination of the activated partial thromboplastin time (APTT). | Dade(R) Actin(R) FSL Activated PTT Reagent | 2 | April 09, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has observed that the listed lot numbers show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated February 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The safety notice instructed consignees to discontinue use of the affected lots and discard any remaining inventory. Contact the firm at 1-800-242-3233, option 1 for questions about this notice. |  | 13083 | 13,083 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Puerto Rico and the country of Canada. | Terminated | December 24, 2013 | 259 | 13083 | Not_Computer
Z-1072-2013 | 64664 | couch, radiation therapy, powered | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table.To be used as part of radiation therapy treatment process. | Precise Treatment Table | 2 | April 09, 2013 | 2013 | Elekta, Inc. | The Column rotation can accidentally be moved to a different position which in the locked position. | DESIGN: Device Design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Elekta office for questions about this notice. |  | 443 | 443 | Nationwide Distribution and Puerto Rico. | Open |  |  | 443 | Not_Computer
Z-1058-2013 | 64751 | mask, gas, anesthetic | BSJ | Anesthetic gas mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Respironics SealFlex Single Port Ribbed Mask (Infant).For general use in respiratory and anesthesia applications as a conduit for various gases to resuscitate or to enhance breathing abilities of the patient. . | SealFlex | 1 | April 10, 2013 | 2013 | Engineered Medical Systems,Inc | The mask inside the package has been identified as "Neonatal" in size and not "Infant" as the labeling indicates. | PRODUCTION CONTROLS: Process Control | Consignees were notified 9/12/2011 with a Medical Device Voluntary Recall letter identifying the product and the problem. Consignees were directed to quarantine, cease further distribution, and destroy the affected product, returning a proof of confirmation/disposal letter, to be submitted for credit. Correspondence should be sent to Tamara Lefevers, Quality Manager, at tlefevers@ engmedsys.com or fax 317-246-5501. | Contact the recalling firm for information | 1595 | 1500pcs, 75cs {20pcs per cases) | Distributed in the states of PA, IN, NY, CO, and CA. | Terminated | August 02, 2013 | 114 | 1575 | Not_Computer
Z-1079-2013 | 64418 | bath, incubators/water, all | JTQ | Microbiological incubator. | Microbiology | Microbiology | 510(K) Exempt | CO2 Incubators (General Purpose Incubator), model numbers 370, 371, 380, 381, 3307, 3308, 3310, 3311, 3540, 3541, 3542, 3543, 3584, 3585, 3598, and 3599.Product Usage:Incubators are designed to maintain an optimal environment for the incubation of tissue and cell samples. These models are designed to maintain temperature, carbon dioxide, and humidity levels. | CO2 Incubators (General Purpose Incubator), | 2 | April 10, 2013 | 2013 | Thermo Fisher Scientific (Asheville) LLC | Incubators manufactured prior to mid 2009 have coated metal hinges that have reportedly failed in the field. | DESIGN: Component Design/Selection | Thermo Fisher sent an Urgent Field Correction Notice letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate all affected products and determine if it is within the recall. If so, they are to return the Acknowledgement Form in order to get replacement hinges and instructions. Upon receipt of the new hinges they are to replace them according to the instructions. |  | 15567 | 15,567 incubators | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SN, TN, TX, UT, VA, VT, WA, WI, WV & WY. and the countries: AFGHANISTAN, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, COSTA RICA, DENMARK, EGYPT, ENGLAND, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, LEBANON, MALAYSIA, MEXICO, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PALESTINE, PEOPLES REPUBLIC OF CHINA, PERU, PHILIPPINES, PORTUGAL, QATAR, REPUBLIC OF PANAMA, ROMANIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, VENEZUELA, VIETNAM & WEST INDIES. | Open |  |  | 15567 | Not_Computer
Z-1088-2013 | 64776 | first aid kit with drug | LRR |  |  | General & Plastic Surgery | Enforcement Discretion | LHB, First Aid Type I Individual Kit, NSN: 6545-00-656-1092, antiseptics, Kit number: 2114002. First Aid Kit. | LHB, First Aid Type I | 3 | April 10, 2013 | 2013 | Lighthouse For The Blind | The product contains an iodine 10% Ampule (10% Povidone Iodine swab) which is being recalled by the firm's distributor. The labeling of the iodine 10% Ampule labels state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterili | MISBRANDING: Labeling False and Misleading | The firm sent "URGENT DEVICE RECALL" letters dated March 22, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-542-3697 for questions regarding this notice. |  | 56133 | 56,133 kits | Nationwide Distribution including DC and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, ID, IL, IN, KS, KY, MI, MD, MN, MO, MT, NC, NJ, NM, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA and WY. | Terminated | November 25, 2013 | 229 | 108939 | Not_Computer
Z-1085-2013 | 64754 | pump, infusion, insulin, to be used with invasive glucose sensor | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion PumpsModel Numbers: MMT-511, MMT-512, MMT-712, MMT-712E, MMT-515, MMT-715, MMT-522, MMT-522K, MMT-722, MMT-722K, MMT-523, MMT-523K, MMT-723, MMT-723K, MMT-554, and MMT-754.Product Usage:Paradigm insulin infusion pumps are intended for use to deliver basal and bolus insulin to persons with diabetes. | Medtronic MiniMed Paradigm Insulin Infusion Pumps | 2 | April 10, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic Paradigm Insulin Pump because the pump's drive support cap may become detached from the pump case and protrude from the lower right side of the pump. When the pump is expose to water it may result in damage to the pump's internal electronics. This moisture damage can prevent the pump's buttons from working properly or can cause the pump to alarm. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Safety Notification letters to all affected customers. The recall letter informs the customers of the problem identified and the actions to be taken. The letter informs the healthcare professionals that Medtronic will inform their patients about the potential safety issues. Customers are instructed to find online information at www.medtronicdiabetes.com/support/product-updates. Customers are instructed to call Medtronic at 1.888.204.7616 (option 1 for loose drive support cap related questions and option 2 for water damage related questions). Distributors are instructed to provide Medtronic with an Excel spreadsheet list of all end users to whom they have shipped any of the listed serialized products. |  | 428000 | 428,000 units | Worldwide Distribution - USA (nationwide) the countries of: AE, AR, AT, AU, BA, BH, BM, BR, BS, CA, CH, CL, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HR, HU, IE, IL, IN, IQ, IT, JO, JP, KR, KW, LB, LC, LV, MK, MX, MY, NL, NO, NZ, OM, PA, PE, PL, PR, PY, QA, RO, SA, SE, SG, SI, SK, SV, SY, TF, TH, TR, TT, TW, UA, UY, VE, ZA. | Terminated | November 04, 2013 | 208 | 428000 | Not_Computer
Z-1076-2013 | 64559 | dental cement w/out zinc-oxide eugenol as an ulcer covering for pain relief | MZW | Dental cement. | Dental | Dental | 510(k) | Kerr Maxcem Elite, self-etch/self-adhesive resin cement; Intro Kit Part No. 34417; DMC 0446 Rev 0.Used for cementation of all indirect restorations. | Kerr Maxcem Elite | 2 | April 10, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is initiating this recall because raw material used to produce the Maxcem Elite is causing the material to prematurely polymerize in the syringe. The product could be difficult to extrude or demonstrate a shorter work-time as the material may prematurely set prior to placing restorations on the tooth or core build-up. As a result, the restoration may not be seated properly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Kerr Corporation/Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 22, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers are instructed to determine if you have any of the affected products in your inventory; If you have the affected lots, please contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number for quick return; complete and return the Maxcem Elite Acknowledgement and Recall Return Form via fax to 714-288-4609 regardless of whether you have any product in your inventory. If you are an authorized Kerr Corporation distributor, identify and contact those customers that may have been shipped the affected product within 48 hours of receipt of this field correction notification.If you have any question, contact Kerr Customer Care at 1-800-537-7123. |  | 40819 | 40,819 kit and syringe units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Belgium, Czech Republic, France, Great Britain, Poland, Albania, Austria, Bulgaria, Switzerland, Cyprus, Germany, Denmark, Estonia, Spain, Finland, Croatia, Hungary, Ireland, ITtaly Luxembourg, Latvia, Netherlands, Norway, Portugal, Romania, Serbia (Republic of Serbia), Sweden, Slovenia, Turkey, Ukraine, United Arab Emirates, Congo, The Democratic Republic of the (formerly Zaire), Kazakhstan, Lebanon, Russian Federation, South Africa, Mexico, Thailand, Brazil, China, Hong Kong, Israel, India, Japan, Malaysia, Panama, Singapore, Viet Nam. | Terminated | January 22, 2014 | 287 | 40819 | Not_Computer
Z-1075-2013 | 56584 | hemoglobin a2 quantitation | JPD | Hemoglobin A2assay. | Hematology | Hematology | 510(k) | G7 HPLC.Used as an aid in the detection and presumptive identification of abnormal hemoglobins in whole blood. | Tosoh Automated Glycohemoglobin Analyzer HLC723G7 Betathalassemia Analysis Mode | 3 | April 10, 2013 | 2013 | Tosoh Bioscience Inc | The Beta-thalessemia Mode of the G7 Analyzer-Beta Thalassemia Mode is an automated High Performance Liquid Chromatography (HPLC) system that separates and reports HbF and HbA2 quantitative percentages in whole blood. The firm received complaints that the A2 peak could not be detected with the G7 Beta-thalessemia columns. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Please be aware that this is not a new recall. The firm has taken action. The firm stated that they have notified consignees via mail, with the list of notified customers maintained by the Regulatory Affairs/Quality Assurance department. The firm stated that they have notified distributors in the same way that they have notified other consignees, with a required confirmation to the firm of corrective actions. | Contact the recalling firm for information | 633 | 633 | Worldwide Distribution-USA (nationwide) including DC, Puerto Rico, Virgin Islands and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV, and the countries of Canada, Columbia, Chile, Brazil, and Jamaica. | Open |  |  | 633 | Not_Computer
Z-1084-2013 | 64749 | tube, orthodontic | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Bracket Buccal Tube U7L DB/UL22 Roth MT HK 0/-14/0 Product Usage: Bracket Buccal Tubes are intended to hold in orthodontic wires used to move teeth during orthodontic treatment | Bracket Buccal Tube | 2 | April 10, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one (1) lot of Bracket Buccal Tubes as a result of a packaging error. More specifically, the BracketBuccal Tubes product package correctly identifies the product as gingival (0Âż) offset brackets, butactually contains distal (10Âż) offset brackets. Use of the affected product may cause unintended rotationof the molars. Although reversible, a delay in | PRODUCTION CONTROLS: Packaging | Ormco sent an Urgent Medical Device Recall letter dated March 8, 2013 to all affected customers, via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees were instructed to complete the Return Form and return any affected product identified in stock. For questions contact Ormco Customer Care. | Contact the recalling firm for information | 8210 | 8210 | Worldwide Distribution - USA (nationwide) including the country of EU. | Terminated | September 04, 2013 | 147 | 8210 | Not_Computer
Z-1097-2013 | 64678 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select.Angiographic x-ray system. | Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fu | 2 | April 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential malfunction and possible hazard to patients and staff when using the Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select. This malfunction is possible in cases where the table vertical lift movement is obstructed by a solid object which could potentially result in collision. This can re | DESIGN: Process Design | Siemens sent a Customer Safety Advisory letter dated January 15th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The SIEMENS service representative will investigate if parts need to be exchanged in your product, an if so, which parts need to be exchanged to restore the system safety and function. Please kindly confirm receipt of the Field Safety Notice by sending back the attached form.For further questions please call (610) 448-3237. |  | 51 | 51 | Nationwide Distribution including the states of CA, FL, GA, IA, IN, KS, LA, ME, MI, MO, NC, NH, NJ, NY, NV, OH, PA, TN, TX, VA and WA. | Terminated | April 23, 2013 | 12 | 51 | Not_Computer
Z-1099-2013 | 64853 | endoscopic irrigation/suction system | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Torrent" irrigation tubing. Product Usage: The Torrent irrigation system (tubing and accessories to accommodate various endoscopes and irrigation pumps) is intended to provide irrigation via irrigation fluids, such as sterile water, during gastrointestinal endoscopic procedures when used in conjunction with an irrigation pump or electrosurgical unit. | Torrent" irrigation tubing | 2 | April 11, 2013 | 2013 | US Endoscopy Group Inc | The firm was notified by their customers that there was a water leakage from the white tubing connector. | DESIGN: Device Design | US Endoscopy Group sent an Urgent Recall Notification letters to their customers. The letter identified the affected product, problem and actions to be taken. They instructed their customers to remove any of the affected products from their inventory and quarantine them. They also requested each customer complete the attached Customer Response Card and return it to US Endoscopy via fax at 1-440-639-4495 or by email at recall@endoscopy.com as soon as possible. For questions contact Customer Service at 1-800-769-8229. | No consumer action necessary | 21189 | 21,189 units | USA Nationwide Distribution - including the states of: AK, AL, AZ, CA, CO, CT, DE, FL, GA, ID, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI & WV. | Open |  |  | 21189 | Not_Computer
Z-1052-2013 | 64308 | prosthesis, knee, patello/femorotibial, semi-constrained, uncemented, porous, coated, polymer/metal/polymer | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ShapeMatch Cutting GuidesProduct Usage: The ShapeMatch Cutting Guides are intended to be used as patient-specific surgical instrumentation to assist in the positioning of total knee arthroplasty components intraoperatively and in guiding the marking of bone before cutting. | ShapeMatch Cutting Guides | 1 | April 12, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | With regards to ShapeMatch Cutting Guides using MRI Imaging, Stryker Orthopaedicshas become aware that there are potential issues associated with internal processesfor planning cases and producing ShapeMatch Cutting Guides. | DESIGN: Device Design | Stryker sent Product Correction Bulletins on 11/8/2012 dated 11/9/2012 and 11/19/2012 to all affected customers. Stryker sent an updated Urgent Medical Device Product Notification letters and Acknowledgement letters dated 1/24/2013 through 1/28/2013 to all affected customers via Fed Ex with return receipt: The letter identified the affect product, problem, potential hazards, risk mitigating factors and actions to be taken. For questions call 201-831-6878.Stryker issued press release on 4/10/2013. |  | 7868 | 7868 units | USA Nationwide Distribution | Open |  |  | 7868 | Not_Computer
Z-1106-2013 | 64577 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Product Code 6801375 and VITROS 5,1 FS Chemistry System Refurbished, Product Code 6801890.Intended for use in the measurement of a variety of chemistry analytes. | VITROS 5,1 FS Chemistry System and VITROS 5,1 FS Chemistry System Refurbished | 2 | April 13, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics, Inc. (OCD) determined that extension springs on some VITROSÂż Systems may beout of its manufacturing specifications. Extension springs are part of the slide alignment guide assembly. During an internal evaluation in its manufacturing facility, it was discovered that an extension spring had insufficient force. If the extension spring does not provide sufficient force, | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | URGENT MEDICAL DEVICE CORRECTION letters were sent to the consignees on 2/07/13. On/about 2/07/13, OCD trained service personnel started to perform an assessment of all the VITROSÂŽ System(s) in End-User facilities worldwide to determine if the extension springs are out of manufacturing specifications. In addition, the slide alignment guides will also be inspected. If the extension springs are out of specification they will be replaced. Consignees will also be provided with a communication that explains the issue. | Contact the recalling firm for information | 670 | VITROS 5,1 FS: USA: 670 units, Foreign: 1249 units; VITROS 5,1 FS Refurbished: USA: 52 units, Foreign: 2 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Republic of Panama, Singapore, Spain, and Venezuela. | Open |  |  | 1503 | Not_Computer
Z-1110-2013 | 64755 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | IS3000 USER MANUAL for Intuitive Surgical da Vinci Si Surgical System and da Vinci Si-e Surgical System, Model SI3000.Arm drapes protect equipment from contamination. | IS3000 USER MANUAL for Intuitive Surgical da Vinci Si Surgical System | 2 | April 13, 2013 | 2013 | Intuitive Surgical, Inc. | Revised instructions for users of the da Vinci S, Si and Si-e systems with additional information regarding sterile adapter engagement and instrument installation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated April 1, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-876-1310 for questions regarding this letter. | Contact the recalling firm for information | 1642 | 1642 customers | Worldwide Distribution.-USA (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Open |  |  | 1642 | Not_Computer
Z-1111-2013 | 64782 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Duraflo coated Femoral Cannula. Product Usage:The cannulae are intended for use in situations which require rapid femoral venous and arterial access for short-term ( | Edwards Lifesciences Duraflo coated FemFlex II Femoral Arterial Cannula | 2 | April 14, 2013 | 2013 | Edwards Lifesciences, LLC | Two unsealed pouches were found at an Edwards' distribution location in Japan. Incorrectly sealed pouches may result in a sterility barrier breach. | PRODUCTION CONTROLS: Packaging Process Control | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of any affected product remaining in inventory and to return any unused devices to Edwards. A completed Acknowledgement Form should be faxed to Edwards at 800-422-9329 within three days of receiving the notice. Questions regarding the recall should be directed to Edwards Customer Services at 800-424-3278, 6:00AM - 4:30PM PST. |  | 379 | 379 units | USA Nationwide Distribution including the states of: FL, MA, MI, NY, PA, TX. | Open |  |  | 379 | Not_Computer
Z-1115-2013 | 64810 | blade, saw, general & plastic surgery, surgical | GFA | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Product is distributed by Zimmer (Warsaw, Indiana)/Synvasive Technology, Inc. (El Dorado Hills, CA) & manufactured under Synvasive Technology, Inc. El Dorado Hills, CA Reciprocating saw blade: A sterile surgical saw blade used to resect damaged cartilage and/or bond in an orthopaedic reconstructive procedure such as a total knee replacement. | Synvasive | 2 | April 15, 2013 | 2013 | Synvasive Technology Inc | Various reciprocating and oscillating saw blades were incorrectly labeled. The labeling errors include: (1) The pouch says oscillating saw blade when it actually contains a reciprocating saw blade; (2) The box bar code date does not match the human readable expiration date; and (3) The label incorrectly states use with Zimmer/Synthes handset and not the Zimmer Universal Power System. | PRODUCTION CONTROLS: Packaging Process Control | Zimmer sent a Urgent Device Removal letter dated April 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers that received affected product from both Synvasive and from Zimmer Surgical will receive two customer letters, as each letter sets out specific item and lot numbers dependent upon distribution entity. Zimmer Distributors will also receive an e-mail notification. All notifications include a response form for both Synvasive and Zimmer distributors as well as customer consignees to account for product disposition (disposal in the field per facilities' sharps protocol).(1) Notifications will be mailed via Fed Ex to Risk Managers and Zimmer Distributors with inventory.(2) Notifications will be emailed to international geographies.(3) Notifications will be emailed all Zimmer Distributors.(4) A second Fed Ex mailing will be sent to domestic non-responding consignees.(5) Effectiveness phone calls will be made to any remaining domestic non-responding consignees. (6) If a customer reports they have further distributed the product, the secondary consignee identified will receive a notification from Zimmer.The scope of the recall is limited to certain confirmed lots which were shipped to the affected consignees..E-mails, certified mail receipts, and telephone records will be the firm's course of action in monitoring effectiveness checks. All returns will be stored in a separate area away from regular inventory. Final disposition of returned product is yet to be determined.The firm's recall strategy appears to be adequate.For further questions please call (916) 939-3913. |  | 3442 | 3442 units per the part numbers listed | Worldwide Distribution - USA (nationwide) and countries of: Austria, Australia, Belgium, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, Panama, Spain, Sweden, Switzerland, Tunisia, United Kingdom and Yemen. | Open |  |  | 3442 | Not_Computer
Z-1117-2013 | 64841 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. | GemStar Infusion System | 2 | April 15, 2013 | 2013 | Hospira Inc. | During infusions of 2.0 mL/hr the motor assembly may rotate backwards capturing additional medication resulting in over-infusion. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated March 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refrain from using the affected produc and contact Hospira at 1-800-441-4100 to arrange for return of their device for repair. Customers were directed to notify their customers if they further distributed the product and confirm they have done so by return of the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 286 | 1) List Number 13086: 286 pumps; 3) List Number 13087: 45,376 pumps; 4) List Number 13088: 23,492 pumps | Worldwide Distribution - USA (nationwide) and Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open |  |  | 69154 | Not_Computer
Z-1124-2013 | 64679 | prosthesis, knee, femorotibial, unicompartmental, semi-constrained, metal/polymer, mobile bearing | NRA |  |  | Orthopedic | PMA | REF 32-422760 Oxford Uni Knee System,Toffee Hammer, non-sterile, found in REF 32-422763 Oxford Mlcroplasty Knee SystemTibial instrument Tray small surgical hammer supplied with the Oxford Partial Knee Microplasty Instrument set, used for impacting implants. | Oxford Unicompartmental Toffee Hammer | 2 | April 15, 2013 | 2013 | Biomet U.K., Ltd. | Investigation of complaints found excessive use of the instrument could result in cracking on the weld between the head and the handle, which could lead to part of the head becoming detached from the handle. | DESIGN: Device Design | BIOMET sent an Urgent Medical Device Recall Notice on March 20, 2013, to all affected customers with directions to forward the notice if product was further distributed to Hospital Operating Room managers. The notice identified the recalled device, reason for the recall, and possible adverse events directing consignees to locate and discontinue use of the device. A FAX back Response form was to be returned to 574-372-1683 within 3 days of receipt prior to return of the product to Biomet, USA. Questionsregarding this recall should be addressed to 574 372-1570 M-F, 8a.m. -5 p.m. |  | 430 | 430 | Worldwide Distribution - USA (nationwide), Europe, Japan, Australia, Canada, Singapore, South Africa, and Costa Rica. | Open |  |  | 430 | Not_Computer
Z-1114-2013 | 64460 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek XS Pro meters (catalog number 05530199160)bar code scanner.Used with the CoagCHEK XS Pro Meters to quantitatively determines prothrombin time ("PT"). | CoaguChek XS Pro meters | 2 | April 15, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche investigation into complaints found Accu-Chek Inform II and CoaguChek XS Pro System barcode reading meters had the potential of intermittent erroneous decoding of patient identification if barcodes being read fail to meet size and/or quality standards. Errors have been confirmed with the EAN 13 barcodes which should not be used for patient or operator identification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Consignees of both devices were notified by Urgent Medical Device Correction letters sent via UPS Ground (receipt required) beginning 3/06/2013. The notice contained recommendation not to use EAN-13 barcode system and ensure barcodes used meet quality standards; use operator and patient lists (downloaded from a data manager) in order to check any decoded identification against the respective list. The notice directed forwarding of the notice if the product was further distributed. Customers were asked to complete the attached fax form. Contact the Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 for Accu check questions and 1-800-428-4674 for CoaguChek questions about the information contained in this UMDC. |  | 4467 | 4467 both devices | Nationwide Distribution. | Open |  |  | 4467 | Not_Computer
Z-1123-2013 | 64700 | accessories, cleaning, for endoscope | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner & Reprocessor System, P/N 50004Designed to automatically clean and high-level disinfect flexible, submersible video or fiber-optic endoscopes. | EVOTECH Endoscope Cleaner & Reprocessor System | 2 | April 15, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has determined that a limited number of EVOTECH ECRs manufactured between 2011 and 2012 did not receive sufficient High Potential (HIPOT) DC voltage testing. |  | An Urgent Medical Device Field Correction letter, dated 3/22/13, was sent to customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR) HIPOT Testing. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed on what to do if electrical shock occurs. Customers are instructed that they will be contacted in the next 2 weeks on behalf of ASP to schedule HIPOT testing. Customers with questions are instructed to contact Stericycle at (888) 965-5816. Customers are instructed to report complaints or suspected problems with their EVOTECH ECR system by contacting ASP Professional Services directly at (888) 783-7723 and select option #2. | Contact the recalling firm for information | 273 | 273 units | Nationwide Distribution-USA including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 18, 2013 | 217 | 273 | Not_Computer
Z-1118-2013 | 63945 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Clinical Chemistry Phenobarbital, list number 1E08 | Clinical Chemistry Phenobarbital | 2 | April 15, 2013 | 2013 | Abbott Laboratories, Inc | Abbott Diagnostics is recalling Clinical Chemistry Phenobarbital reagent lots 62299UN12 and 85773UN12 due to imprecision issues which increase with the age of the reagent lot. The imprecision of the Phenobarbital assay is d 7% Total CV%; however, internal testing confirmed CVs up to approximately 12% for in-date reagent lots over five months old. Although lot 52803UN12 is now expired, it exhibit | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Laboratories sent an Immediate Action Required Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures, order replacement lot 27426UN12 which will expire 18 DEC 2012, and is the only lot currently available (Additional lots will be available in December and January), and complete and return the enclosed Customer Reply Form. If lots were forwarded to another laboratory, a copy of the letter should be forwarded to them. For questions regarding this recall call 1-877-4ABBOTT. For customers outside of the US they should contact their local area Customers Service. | Contact the recalling firm for information | 1289 | 1289 | Worldwide Distribution - USA (nationwide) and Internationally to Internationally distributed to Argentina, Australia, Austria, Bahamas, Brunei, Canada, Chile, Curacao, Denmark, France, Germany, Greece, Ireland, Italy, Lebanon, Mexico, Netherlands, Norway, Oman, Paraguay, Portugal, Puerto Rico, Russia, Saudi Arabia, Serbia, South Korea, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | June 18, 2013 | 64 | 1289 | Not_Computer
Z-1121-2013 | 64610 | calibrator, primary | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Vitamin D Calibrator 2-pack; Reference Number:10493589.For use in the quantitative determination of Vit D. | Siemens ADVIA Centaur Vitamin D Calibrator 2pack | 2 | April 15, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Negative bias in patient samples for Vitamin D |  | Siemens Healthcare issued Urgent Medical Device Recall was sent to all affected Siemens Healthcare Diagnostics customers in the United States on March 13th, 2013. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Medical Device Recall instructs customers to discontinue the use of ADVIA Centaur VitD Assay Kit Lots (REF 10491994 and 10631021) and ADVIA Centaur Systems VitD Cal (REF 10493589 and 10630911) released prior to the recalibration of the assay and request replacement lots. This notice also directs customers to review the contents of the recall notice with their laboratory or medical director. A confirmation fax-back form is included to ensure that all customers in the United States have been notified and have complied with the mandatory action. |  | 372 | 372 pk Cals | Nationwide Distribution. | Open |  |  | 8489 | Not_Computer
Z-1125-2013 | 64568 | ventilator, emergency, powered (resuscitator) | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Instrumentation Industries BE183-SU and BE 183 SU-12 Manual Jet Ventilators; BE 183-SUR and BE 183 SUR-12 Manual Jet Ventilators with Regulator and Gauge; BE 183-2 4 ft. Small Bore Tubing Assembly w/Inline Filter; and BE 183-R RegulatorThe BE 183-SUR and BE 183-SU Manual Jet Ventilators are designed to provide transtracheal ventilation in specific emergency situations of upper airway obstructions. | Manual Jet Ventilators, Manual Jet Ventilators w/Regulator and Guage, Small Bore Tubing Assembly | 2 | April 16, 2013 | 2013 | Instrumentation Industries Inc | Firm initiated an update to Directions for Use necessitating replacement of previously distributed Directions for Use. | PRODUCTION CONTROLS: Error in Labeling | Instrumentation Industries, Inc. initiated a recall of this product on September 7, 2012 by sending a letter via priority mail to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to replace the package insert of any affected product with the enclosed corrected package insert (now at Rev. H). The Directions for Use can alsbe downloaded from their website if they should need additional copies. If the product was further distributed consignees were asked to notify their consignees about this Field Correction.For questions regarding this recall call 412-854-1133. |  | 2998 | 2998 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. Shipped internationally to Canada, Ireland, Hong Kong, the Netherlands, New Zealand, and Thailand. | Terminated | July 05, 2013 | 80 | 2998 | Not_Computer
Z-1020-2013 | 64734 | unit, x-ray, intraoral | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | DeGotzen XRay unit XGenus Dental X-ray unit. | Generator, HighVoltage , X Ray Diagnostic | 2 | April 17, 2013 | 2013 | DeGotzen | It was discovered at the WEAC Labs method sample did not meet requirements for labeling and certification | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | DeGotzen will add new labeling or change existing labeling to correct for the 5 mentioned deficiencies. This correction of the noncompliant products must be made at no cost to the customers.For further questions please call 800-289-6367 ext. 390. |  | 159 | 159 units | US (nationwide) | Open |  |  | 159 | Not_Computer
Z-1138-2013 | 64486 | filler, bone void, osteoinduction (w/o human growth factor) | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | OsteoSelect DBM PuttyProduct Usage:Orthopedic bone filler | OsteoSelect | 2 | April 17, 2013 | 2013 | Bacterin International, Inc. | The firm's retesting procedure was not validated. | PRODUCTION CONTROLS: Process Control | Bacterin sent a Voluntary Product Recall Notification letter via e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are being asked to return any remaining product by obtaining a Return Authorization number from them. Bacterin is also asking customers to return the verification form mailed with the recall letter regardless of whether or not they had any product on hand. For questions contact Bacterin sales representative or Bacterin Customer Service. |  | 165 | 165 | USA Nationwide Distribution | Terminated | July 01, 2013 | 75 | 165 | Not_Computer
Z-1136-2013 | 64395 | endoscopic suture/plication system, gastroesophageal reflux disease (gerd) | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Plus Device with Serosa Fastener and Accessories Part number C02033-01. Label reads in part "EsophyX2 2-LINK Device Qty: (1) Fastener Delivery System ***"Designed for the treatment of Gastroesophageal Reflux Disease (GERD), with erosive and non-erosive esophagitis | Endoscope Clip applier, Implantable Fastener and Accessories | 2 | April 17, 2013 | 2013 | Endogastric Solutions Inc | EsophyX2Plus device was not validated under actual or simulated use conditions prior to distributing the product to physicians. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | EndoGastric Solutions, Inc. sent an Urgent Medical Device Recall letter dated January 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate and record lot numbers for any affected product. If no product remain in inventory customers were instructed to report such to customer service. Upon receipt of the product a credit will be issued for each device returned. For any questions customers were instructed to call 425-307-9226.For questions regarding this recall call 425-307-9200. |  | 190 | 190 | Nationwide Distribution including NC, VA, WI, TN, KY, IN, FL, MI, IA, IL, TX, LA, AR, HI, CA, NV, OR, and WA, | Terminated | July 02, 2013 | 76 | 190 | Not_Computer
Z-1132-2013 | 64669 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | TRIGEN (TM) L-P SCREW, 4.5 MM X 37.5 MM, TI-6AL-4V, QTY: (1), STERILE R, REF 71645437, Smith & Nephew, Inc. Memphis, TN 38116 USAorthopedic surgery | TRIGEN (TM) LP SCREW, 4.5 MM X 37.5 MM | 2 | April 17, 2013 | 2013 | Smith & Nephew Inc | One batch of 4.5 mm diameter Trigen Low-Profile Bone Screws were incorrectly anodized with a gold coating (indicating a 5 mm diameter) instead of a gray coating (indicating a 4.5 mm diameter). | PRODUCTION CONTROLS: Process Control | Smith & Nephew sent an Urgent Product Recall Notification letter dated March 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the affected product and quarantine for return. Non-responding consignees were notified again on March 15, 2013, by email. Surgeons were to be notified by letter beginning on March 20, 2013.For questions regarding this recall call 978-749-1330. |  | 23 | 23 units | Nationwide Distribution including CA, FL, GA, MA, MI, NC, SC, OH, and MO. | Open |  |  | 23 | Not_Computer
Z-1134-2013 | 64747 | aortic valve, prosthesis, percutaneously delivered | NPT |  |  | Cardiovascular |  | Edwards Lifesciences Crimpers, Model 9100CR23. The Crimpers are used to crimp the transcatheter heart valve onto the delivery system prior to use. | Edwards Lifesciences Crimpers | 2 | April 17, 2013 | 2013 | Edwards Lifesciences, LLC | Holes in the sterile barrier pouch may be caused by handling of the device during manufacturing or once pouches have been removed from the unit box and stored on the shelf by customers. |  | Edwards Lifesciences notified customers via letter and instructed them to maintain the device in its protective unit carton until ready for use and to inspect the pouch for holes or tears prior to use as instructed in the product's labeling. If any holes or tears are found, customers are instructed to contact Edwards Customer Service to arrange for return and replacement of damaged product. | Contact the recalling firm for information | 5472 | 5,472 units | Nationwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | September 16, 2013 | 152 | 11589 | Not_Computer
Z-1128-2013 | 64698 | scale, patient | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | SR Instruments FG3105 Patient Scales Used on various Joerns Healthcare Patient Lifters. Model# FG3105 Patient Scales. Product Usage:The FG3105 Scale is designed for Joerns Healthcare for use in portable patient weighing systems. The scale is designed to be connected between the patient hoist mechanism and patient sling apparatus. The scale is not intended to be used for patient transport. | Patient Scales | 2 | April 17, 2013 | 2013 | Sr Instruments Inc | SR Instruments Inc. is implementing a field correction for the Joerns Healthcare Patient Lifts after they received two incident reports regarding two field failures of a tensile type load cell involving the FG3105 Patient Scale. The failures resulted in patient falls. |  | SR Instruments, Inc. sent a Notification letter dated December 13, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 716-693-5977. |  | 806 | 806 patient scales | USA Nationwide Distribution in the state of Texas. | Open |  |  | 806 | Not_Computer
Z-1129-2013 | 64478 | single (specified) analyte controls (assayed and unassayed) | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan brand OneTouchÂŽ Ultra Control Solution, Part Numbers/Model Numbers: 01045808, 02141603.Testing of Blood Glucose Levels. | LifeScan OneTouch Ultra Control Solution | 2 | April 17, 2013 | 2013 | Lifescan Inc | When OneTouchÂż Ultra and OneTouchÂż Select test strips are tested with Ultra and Select control solution the meter may display a result outside the control solution test range printed on the strip vial label even though the system is working properly. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | LifeScan sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers may contact OneTouch Healthcare Professional Line at 866-461-0089 for questions about this notice. | Contact the recalling firm for information | 132301 | 132,301 units total | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, 10, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NO, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SO, TN, TX, UT, VA, VI, VT, WA, WI,WV, and WY, and the countries of Russia and Ukraine. | Terminated | October 23, 2013 | 189 | 132301 | Not_Computer
Z-1139-2013 | 64671 | hydrazone colorimetry, alt/sgpt | CKD | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX ALT AST Test Cassette, Model #12-788.For the in vitro quantitative determination of alanine aminotransferase (ALT) and aspartate aminotransferase (AST). | Alere Cholestech LDX ALT AST Test Cassette | 2 | April 17, 2013 | 2013 | Alere San Diego, Inc. | The recall was initiated because Alere San Diego has an update for the Alere Cholestech LDX ALT AST cassette, and it has the potential for humidity changes to impact results for certain analytes on the Alere Cholestech LDX System. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Alere sent an "URGENT MEDICAL DEVICE RECALL UPDATE" letter dated March 6, 2013 to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers with any questions about the recall notification, are instructed to contact Alere San Diego 9975 Summers Ridge Rd, San Diego, CA 92121; (877) 308-8289; Fax: (858) 805-8457. Customers are instructed to complete the customer response form even if they do not have any involved product. |  | 38571 | 38,571 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada, South Africa, Netherlands, Italy, Brazil, Singapore, Russia, Hong Kong, Argentina, South Korea, Philippines, El Salvador, Indonesia, and Trinidad. | Terminated | August 13, 2013 | 118 | 38571 | Not_Computer
Z-1127-2013 | 64646 | prosthesis, shoulder, hemi-, humeral, metallic uncemented | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Oxford Partial Knee; Oxford Knee Resection Procedure 3 Pack CEMENTEDFor use with Stryker System 5, System 4, System 2000 and EDH. Product Usage: For use with Stryker System 5, System 4, System 2000 and EDH. A set of three surgical saw blades is used to resect damaged cartilage and/or bone in an orthopedic reconstructive procedure such as a partial knee replacement. The set contains both oscillating blades for tissue resection along a small arc in a left/right or up/down direction, and reciprocating blades for resection back and forth in a single plane. | Surgical Saw Blade Procedure Pack | 2 | April 17, 2013 | 2013 | Synvasive Technology Inc | Biomet part # 506076, lot 928182 was received from Synvasive Technology containing the incorrect blade. |  | An Urgent Medical Device Recall Notice letter dated March 21, 2013 was sent to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use; remove the products from inventory and complete and return the attached Response Form. For questions call 574-372-1570. |  | 100 | 100 kits | USA Nationwide Distributor including the state of IN | Terminated | January 13, 2014 | 271 | 100 | Not_Computer
Z-1145-2013 | 64840 | injector and syringe, angiographic | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Medical Systems Inc.The ACIST Angiographic Injection System is intended to be used for the controlled infusion of radiopaque contrast media for angiographic procedures. | ACIST | 2 | April 18, 2013 | 2013 | Acist Medical Systems | ACIST Medical Systems is issuing a voluntary recall notice for three specific lots of BT2000 Automated Manifold Kits. For these specific lots, ACIST has received field reports related to ingress of air into the manifold body. This has been detected by the Air Column Detect Sensor halting the injection, and normally replacement of the BT2000 Automated Manifold Kit is necessary in order to proceed w | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | ACIST Medical Systems sent an Urgent Medical Device Recall letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter advised customers that if they had inventory of any of the lots under the scope of the recall, to discontinue use or further distribution of the product. If you possess inventory containing any of the lot numbers listed above, please discontinue use or further distribution of these products. Please return this inventory to ACIST. Upon receipt of the returned BT2000 Automated Manifold Kits, ACIST will arrange for a shipment of replacement kits at no charge. To order additional or immediate replacement product, please contact BRACCO customer service at 1-877-BRACCO9 (1-877-272-2269), option 1.Please also complete and return the enclosed Medical Device Recall Form (see enclosure). Even if you do not possess any of the affected product in your inventory, the form needs to be completed and returned, indicating that you have no remaining units containing these lot numbers in your inventory. |  | 6640 | 6640 units | Nationwide Distribution including the states of AL, AK, AZ, AR, CA, CO, CT, DC, GA, HI, ID, IL, IN, IA, KS, KY, MD,MI, MN, MS, MO, NV, NH, NJ, NM, NY,NC, OR, OK, OH, PA, TN, TX, VA and WA. | Terminated | August 07, 2013 | 111 | 6640 | Not_Computer
Z-1143-2013 | 64538 | catheter, peripheral, atherectomy | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Elana Catheter Sizer 2.0***The Elana Catheter Sizer 2.0 is designed to aid the surgeon to assess the bypass graft of the ELANA technique. | Elana Catheter Sizer 2.0 | 2 | April 18, 2013 | 2013 | Elana, Inc. | Distribution of an unapproved device into interstate commerce. | PREMARKET APPROVAL: No Marketing Application | Elana Inc. issued an Urgent: Medical Device Recall letter on February 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the Elana Catheter Sizer 2.0 as soon as possible. The Sizer 2.0 was intended as an optional accessory to the Elana Surgical Kit. The Elana Surgical Kit may be used properly without the Elana Catheter Sizer 2.0. Please acknowledge receipt of these instructions immediately by Email: Claartje@elana.com or Fax: 301-215-7038. The return shipping address for the Sizer 2.0 is Elana, Inc., 10480 Little Patuxent Parkway, Suite 400, Columbia, MD 21044 USA. Contact information for questions: Phone: 301-215-7010 or 443-538-4923 or +1 31 30 2537276 www.elana.com. |  | 3 | 3 devices | Distributed in the states of Pennsylvania, Texas, and New York. | Terminated | May 09, 2013 | 21 | 3 | Not_Computer
Z-1147-2013 | 64770 | computer, diagnostic, pre-programmed, single-function | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | HCS3011 AV Loop Kit For use with Transonic CO status System Only, -HCS3021 AV Loop Kit For use with Transonic CO status System Only. Product Usage: is indicated for use in patients greater than or equal to 2 years old with arterial and venous lines in the diagnostic assessment of cardiovascular status including cardiac output and associated hemodynamic parameters. | AV Loop Kit | 2 | April 18, 2013 | 2013 | Transonic Systems Inc | Transonic Systems Inc. have received complaints of blood leakage involving the AV Loop Kit. | DESIGN: Component Design/Selection | Transonic Systems Inc. sent an Urgent: Medical Device Recall dated March 27, 2013 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine, discontinue use and return any recalled product via FedEx. Transonic Systems Inc. has offered to reimburse the firm for the return of these recalled products by check or credit memo for the returned goods and postage. Transonic Systems Inc. has requested that if a firm has further distributed any of the above lots, that they immediately contact their accounts, inform them of the recall, and have them return their outstanding recalled stocks to you. Transonic Systems Inc. have included copies of the product labeling and a Recall Return Response Form in their packet. Completed Recall Return Response Forms should be faxed to 607-257-5310. |  | 195 | 195 units total (140 units Domestic, 55 units Foreign) | Worldwide Distribution - USA Nationwide including the states of AZ, CA, FL, GA, NM, NY, TX, VA and WA and countries of Netherlands, Canada and Russia. | Open |  |  | 195 | Not_Computer
Z-1151-2013 | 64596 | enzyme immunoassay, digoxin | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Siemens Dimension Vista(R) System Vista Drug 4 Calibrator. The device is used as an in vitro diagnostic product. | Dimension Vista(R) System Vista Drug 4 Calibrator | 2 | April 18, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed complaints of low QC and patient recovery with the Dimension Vista(R) LOCI(R) Digoxin(DIGXN) Flex(R) reagent cartridge when calibrating DIGXN with either Drug 4 Cal (KC460) lot 2KD052 or Drug 4 Cal (KC460A) lot 2KD053. Internal testing has shown an average 15% low bias in the therapeutic range of 0.9 to 2.0 ng/mL. |  | The firm initiated their recall of this product by sending an Urgent Medical Device Recall letter dated February 2013 to all Dimension Vista Drug 4 Calibrator customers who received the affected lots. The letter notified them of the issue, the potential risk to health, and actions to be taken by customers. |  | 1638 | 1,638 | Distributed within the US to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | January 07, 2014 | 264 | 1638 | Not_Computer
Z-1131-2013 | 64848 | system, blood collection, vacuum-assisted, manual | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | Heritage Labs Path Study Extra Supplies Kit, Product # A 2376, kit contains 3 zippered plastic bags which hold support supplies for health care providers when treating patients. Bag #3 contains 5 different blood collection tubes (5 each) and 2 different gauge needles--all for drawing blood. One of the tubes is the recalled BD blue top citrate tube (5 tubes). These tubes were manufactured by BD, Franklin Lakes, NJ. The kit is used by health care providers when treating patients. | A2376 Extra Supplies Kit | 2 | April 18, 2013 | 2013 | Heritage Labs International LLC | The product has a Lithium Heparin contaminant which could adversely affect coagulation results potentially causing erroneous lab results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The recalling firm contacted 25 of their 26 customers by phone on February 5, 2013, to notify them of the recall and sent a follow up memo on February 6, 2013, During the call, the firm reminded their customers how many recalled kits they recieved and instructed them to go through the kits and pull the recalled lots. The recalled kits are to be returned to the recalling firm and the firm sent a follow-up memo on February 6, 2013, with the shipping label their customers to place on the returning kit shipping box. This memo was sent to the customers by overnight Fed-ex mail and the ship label for the returned product was an overnight UPS label. On February 7, 2013, the final customer was notified by phone and instructed to return the recalled kits.For further questions please call (913) 764-1045. |  | 25 | 25 kits | Nationwide Distribution including the states of AL, CA, GA, IL, IN, KY, MA, MD, NC, NY, OK and VA. | Terminated | November 25, 2013 | 221 | 25 | Not_Computer
Z-1142-2013 | 64884 | interventional fluoroscopic x-ray system | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100IQ, Innova 4100IQ, Cardiovascular and Interventional Imaging System. | GE Healthcare, Innova 2100IQ, Innova 3100IQ, Innova 4100IQ | 2 | April 18, 2013 | 2013 | GE Healthcare, LLC | As part of GE Innova IQ table introduction, a label was designed to be affixed on the front Cover of theGantry Pivot section. This safety label is missing on systems manufactured between March 2010 and July 2011. In the situation a staff member (doctor, nurse, operator, anesthetist, etc&) stands in the region between the C-arm gantry pivot and the head section of the table, there is a risk of co | DESIGN: Labeling Design | A field action was initiated on October 21, 2011, and completed March 9, 2012, . The issue was corrected to all affected systems by applying the missing label onto the gantry. No recall letter was issued.For further questions please call (262) 513-4122. |  | 48 | 48 | Worldwide Distribution - Nationwide Distribution including the states of FL, IL, IN, LA, MI, MN, NJ, NM, NY, OH, PR, SC, TX and VA and the countries of Vietnam, Latvia, Italy, Israel, Indonesia, Hong Kong, United Kingdom, Finland, Spain, Czech Republic, Canada, Bahrain and Australia. | Terminated | February 04, 2014 | 292 | 48 | Not_Computer
Z-1148-2013 | 64790 | system, test, infectious mononucleosis | KTN | Infectious mononucleosis immunological test system. | Immunology | Immunology | 510(k) | IMO-402 Mononucleosis Rapid Test Device (Whole Blood/Serum/Plasma). Catalog #IMO-402/94201The MONO test (Whole Blood) is a rapid chromatographic immunoassay for the qualitative detection of Infectious Mononucleosis heterophile antibodies in whole blood to aid in the diagnosis of infectious Mononucleosis. | IMO402 Mononucleosis Rapid Test Device | 3 | April 18, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego is recalling the Clearview Mononucleosis Cassette because the positive control included in the rapid test may produce false negative results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Alere sent an Urgent Medical Device Recall letter dated February 26, 2013, to all affected customers. The letter informs the customers of the problems identified and the actions to be taken. Customers with unused product remaining are instructed to immediately discard it. Investigations of customer complaints indicate that the positive control included with this lot may produce false negative results. Customers are instructed to return the attached Verification form. Customers that have consumed all product are instructed to indicate zero under "Quantity Discarded" and complete and fax back the form. Replacement product will be provided for any unused portion of the affected lot based on completion and return of the attached Verification Form. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at later.flow.support@alere.com.For questions regarding this recall call 858-805-2000, ext 3015. |  | 78 | 78 kits | Nationwide Distribution including TN, WA, UT, KS, KY, MN, GA, MA, PA, NY, VT, and the Dist. of Columbia | Terminated | August 29, 2013 | 133 | 78 | Not_Computer
Z-1140-2013 | 64499 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Duane Reade TRUEtrack monitor kit, UPC# 6 39194 01679 7 Part # A4006-81The TRUEtrack Blood Glucose System is intended for the quantitative determination of glucose in human whole blood taken from finger or forearm. The System is intended to be used to assist the patient and healthcare professional in the management of diabetes. | Duane Reade TRUEtrack monitor kit | 2 | April 18, 2013 | 2013 | Nipro Diagnostics, Inc. | Meter kit outer boxes containing 10 ct. test strips state incorrect storage temperature range. | PRODUCTION CONTROLS: Process Control | NIPRO Diagnostics sent a Voluntary Product Field Action Notice dated August 12, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check with all Duane Reade retail locations and Duane Reade Disgtribution Center to determine if they have any of the affected product ath their location. If product is located customers were instructed to call 1-800-803-6049 for a Return Authorization Number and instructions for returning the affected product.For questions regarding this recall call 954-677-9201. | Contact the recalling firm for information | 1627 | 1,627 Kits | Nationwide Distribution only in NY. | Terminated | July 08, 2013 | 81 | 1627 | Not_Computer
Z-1146-2013 | 64229 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Pressure Monitoring sets with VAMP Plus closed blood sampling system model # T001691A, PXVP0566, PXVP2260, PXVP0826, PXVP0827, PXVP2284AT3, PXVP23X3AT3, T398807D, T100605B, T001762A, T391T00A, VP2.Blood sampling system used on patients requiring periodic withdrawal of blood samples. | Pressure Monitoring sets | 2 | April 18, 2013 | 2013 | Edwards Lifesciences, LLC | The device is being recalled because the firm has received multiple complaints related to incorrect assembly that may result in dilution of blood samples. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm initiated a recall in Europe via customer letters starting January 30, 2013 to notify the recall of model number T001691A lot number 59326333. The recall was expanded and a worldwide recall via customer letters was initiated on February 11, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 30467 | 30,467 | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, HI, ID, IN, KS, KY, LA, MA, MI, MN, MO, NC, ND, NM, NV, NY, OH, OK, PA, SD, TN, TX, VA, WA, and WI) and Europe. | Open |  |  | 30467 | Not_Computer
Z-1153-2013 | 64638 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip Xpress Glu-Test Strips (Glucose test strips)Catalog # 42214 | Nova StatStrip Glucose Test Strips | 2 | April 19, 2013 | 2013 | Nova Biomedical Corporation | Glucose test strips reports no result message when tested | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Nova Biomedical, sent an "Urgent Field Safety Notice 01-13SS" dated March 18, 2013 and US customers notification by phone contact and to the International customers including a Faxback Form to be communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call). The customers were instructed to INVALIDATE the Nova StatStrip Glucose Test Strip lots indicated in the list on all StatStrip meters within your institute; REPLACE all Nova StatStrip Glucose Test Strips from the indicated lots with replacement test strips provided by Nova; REMOVE all unused boxes of the Nova StatStrip Glucose Test Strip lots indicated from all locations within your institute; RETURN all unused boxes of the Nova StatStrip Glucose Test Strips lots indicated to Nova Biomedical U.K. or your local Nova distributor following the country specific instructions provided by your authorized Nova representative, and complete and return the Fax-Back Form via email or FAX to your Regional Service Manager at 781-8914806.On April 08, 2013, the firm expanded the recall notification to include one additional test strip lot #0313021309 which was shipped under two catalog numbers(Catalog # 42214 and Catalog# 51493). US Notification by phone contact with key customer contact facilitated by Senior Manger of Technical Support Services. International - Field Safety Notice and Faxback Form is being communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call).Questions regarding this notice should be directed to the following Nova Biomedical EU Authorized Representative at: Nova Biomedical Innovation House, Aston Lane South Runcorn,WA7 3FY Telephone: + 44 1244 287087Email: Europe@novabiomedical.co.uk. or US at (781) 894-0800. |  | 7433 | 7433 cases | Worldwide distribution: USA (nationwide) including states of: CA, IL, KY,NY, WV and countries of: United Kingdom and Italy. | Open |  |  | 9633 | Not_Computer
Z-1152-2013 | 64707 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHBR (Phenobarbital) Slides, REF 822 1384.For the quantitative measure of phenobarbital (PHBR) concentration in serum and plasma. | VITROS Chemistry Products PHBR Slides | 2 | April 19, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in accuracy and control fluid values outside of expected intervals when using VITROS Chemistry Products PHBR Slides. OCD's investigation confirmed that positively biased values for patient samples could be obtained that may not be detected by quality control fluids. |  | On 3/14/13, URGENT PRODUCT CORRECTION NOTIFICATION letters (dated 3/14/13), Questions & Answers, and Response Forms were sent to consignees via FedEx courier or US Postal Service Priority Mail. Consignees were informed of the issue and advised to immediately discontinue using VITROS PHBR Slides. On 3/14/13, distributor letters were sent via US Postal Service Priority Mail to the distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS PHBR Slides. On 3/14/13, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 594 | Domestic: 11,164 units, Foreign: 594 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open |  |  | 11758 | Not_Computer
Z-1156-2013 | 64584 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Orthopedic Alliance Spine System (alias "Blue & Gold")Part Numbers:BGC3010 Screw CapBGR0040 Straight Rod 5.5 x 40mmBGR0050 Straight Rod 5.5 x 50mmBGR0060 Straight Rod 5.5 x 60mmBGR0070 Straight Rod 5.5 x 70mmBGR0080 Straight Rod 5.5 x 80mmBGR0090 Straight Rod 5.5 x 90mmBGR0100 Straight Rod 5.5 x 100mmBGR0110 Straight Rod 5.5 x 110mmBGR0120 Straight Rod 5.5 x 120mmBGR0130 Straight Rod 5.5 x 130mmBGR0140 Straight Rod 5.5 x 140mmBGR0150 Straight Rod 5.5 x 150mmBGR0200 Straight Rod 5.5 x 200mmBGR0300 Straight Rod 5.5 x 300mmBGR0400 Straight Rod 5.5 x 400mmBGR0500 Straight Rod 5.5 x 500mmBGS5535 5.5 x 35mm Polyaxial ScrewBGS5540 5.5 x 40mm Polyaxial ScrewBGS5545 5.5 x 45mm Polyaxial ScrewBGS6035 6.0 x 35mm Polyaxial ScrewBGS6040 6.0 x 40mm Polyaxial ScrewBGS6045 6.0 x 45mm Polyaxial ScrewBGS6050 6.0 x 50mm Polyaxial ScrewBGS6535 6.5 x 35mm Polyaxial ScrewBGS6540 6.5 x 40mm Polyaxial ScrewBGS6545 6.5 x 45mm Polyaxial ScrewBGS6550 6.5 x 50mm Polyaxial ScrewBGS7035 7.0 x 35mm Polyaxial ScrewBGS7040 7.0 x 40mm Polyaxial ScrewBGS7045 7.0 x 45mm Polyaxial ScrewBGS7050 7.0 x 50mm Polyaxial ScrewBGS7055 7.0 x 55mm Polyaxial ScrewBGS7535 7.5 x 35mm Polyaxial ScrewBGS7540 7.5 x 40mm Polyaxial ScrewBGS7545 7.5 x 45mm Polyaxial ScrewA pedicle screw system intended to provide immobilization and stabilization of spinal segments in skeletally mature patients. | Orthopedic Alliance Spine System (alias "Blue & Gold") | 2 | April 22, 2013 | 2013 | Orthopedic Alliance LLC | Orthopedic Alliance is recalling Blue & Gold Implants and Instruments because the Blue & Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product peformance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. | DESIGN: Device Design | A recall letter dated March 7, 2013 was sent to one customer who purchased the Blue & Gold Implants and instruments. The letter informed the customer that the Blue & Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. Customer is informed of the actions to be taken. Customer is instructed to complete the customer response form and return it by FAX to (858) 764-9739. |  | 3300 | 3,300 units | Distributed in Texas. | Open |  |  | 3300 | Not_Computer
Z-1157-2013 | 64578 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | APLIF Implants and Instruments, Part Numbers: 184008-10 Anterior Psoas Lumbar 18 x 40 x 08mm - 10degree184010-10 Anterior Psoas Lumbar 18 x 40 x 10mm - 10degree184012-10 Anterior Psoas Lumbar 18 x 40 x 12mm - 10degree184014-10 Anterior Psoas Lumbar 18 x 40 x 14mm - 10degree184016-10 Anterior Psoas Lumbar 18 x 40 x 16mm - 10degree184008-00 Anterior Psoas Lumbar 18 x 40 x 08mm - 00degree184010-00 Anterior Psoas Lumbar 18 x 40 x 10mm - 00degree184012-00 Anterior Psoas Lumbar 18 x 40 x 12mm - 00degree184014-00 Anterior Psoas Lumbar 18 x 40 x 14mm - 00degree184016-00 Anterior Psoas Lumbar 18 x 40 x 16mm - 00degree184510-08 Anterior Psoas Lumbar 18 x 45 x 08mm - 10degree184510-10 Anterior Psoas Lumbar 18 x 45 x 10mm - 10degree184512-10 Anterior Psoas Lumbar 18 x 45 x 12mm - 10degree184514-10 Anterior Psoas Lumbar 18 x 45 x 14mm - 10degree184516-10 Anterior Psoas Lumbar 18 x 45 x 16mm - 10degree184510-00 Anterior Psoas Lumbar 18 x 45 x 10mm - 00degree184512-00 Anterior Psoas Lumbar 18 x 45 x 12mm - 00degree184514-00 Anterior Psoas Lumbar 18 x 45 x 14mm - 00degree184516-00 Anterior Psoas Lumbar 18 x 45 x 16mm - 00degree185008-10 Anterior Psoas Lumbar 18 x 50 x 08mm - 10degree185010-10 Anterior Psoas Lumbar 18 x 50 x 10mm - 10degree185012-10 Anterior Psoas Lumbar 18 x 50 x 12mm - 10degree185014-10 Anterior Psoas Lumbar 18 x 50 x 14mm - 10degree185016-10 Anterior Psoas Lumbar 18 x 50 x 16mm - 10degree185008-00 Anterior Psoas Lumbar 18 x 50 x 08mm - 00degree185010-00 Anterior Psoas Lumbar 18 x 50 x 10mm - 00degree185012-00 Anterior Psoas Lumbar 18 x 50 x 12mm - 00degree185014-00 Anterior Psoas Lumbar 18 x 50 x 14mm - 00degree185016-00 Anterior Psoas Lumbar 18 x 50 x 16mm - 00degree185508-10 Anterior Psoas Lumbar 18 x 55 x 08mm - 10degree185510-10 Anterior Psoas Lumbar 18 x 55 x 10mm - 10degree185512-10 Anterior Psoas Lumbar 18 x 55 x 12mm - 10degree185514-10 Anterior Psoas Lumbar 18 x 55 x 14mm - 10degree185516-10 Anterior Psoas Lumbar 18 x 55 x 16mm - 10degree185508-00 Anterior Psoas Lumbar 18 x 55 x 08mm - 00degree185510-00 Anterior Psoas Lumbar 18 x 55 x 10mm - 00degree185512-00 Anterior Psoas Lumbar 18 x 55 x 12mm - 00degree185514-00 Anterior Psoas Lumbar 18 x 55 x 14mm - 00degree185516-00 Anterior Psoas Lumbar 18 x 55 x 16mm - 00degreeThe Eminent Spine Interbody Fusion System is indicated for intervertebral body fusion of the lumbar spine, from L2 to S1, in skeletally mature patients who have had six months of non-operative treatment. | APLIF Implants and Instruments | 2 | April 22, 2013 | 2013 | Spinal Solutions, LLC | Spinal Solutions is recalling the APLIF system because it is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, movement, or inadequate sterilization. | PREMARKET APPROVAL: No Marketing Application | Spinal Solutions LLC sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to discontinue use of the affected product and return to the firm. Customers were also instructed to complete the bottom of the form and to return it by fax to (858) 764-9739.For questions regarding this recall call 951-304-9001. |  | 220 | 220 units | Nationwide Distribution including NV, WI, MD,and CA. | Open |  |  | 220 | Not_Computer
Z-1158-2013 | 64715 | collector, urine, (and accessories) for indwelling catheter | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Belly Bag Urine Collection Bag with Hip Belt.The product is a sterile urine collection device that collects urine by mechanical means when attached to an indwelling Foley or suprapubic catheter. | Belly Bag Urine Collection Bag with Hip Belt. | 2 | April 22, 2013 | 2013 | Teleflex Medical | Sterile packaging may be compromised. |  | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed if you have any of the affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed.If you and your customers have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document your receipt of this letter.To return product complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document the amount of product you have on hand for return.If you have any other questions feel free to contact your local sales representative or Customer Service at 1 (866) 246-6990. |  | 190103 | 190,103 | Worldwide Distribution - USA and the countries of Singapore, Canada, New Zealand, Germany, Ireland, Japan and Australia. | Open |  |  | 190103 | Not_Computer
Z-1162-2013 | 64508 | microspheres radionuclide | NAW |  |  | Radiology |  | Sirtex SIR-yttrium 90 microspheres (Radionuclide); SIR-Y001.Indicated for the treatment of unresectable metastatic liver tumors. | Sirtex SIRyttrium 90 microspheres (Radionuclide) | 3 | April 23, 2013 | 2013 | Sirtex Medical Limited | Incorrect expiration date | MISBRANDING: Labeling False and Misleading | Sirtex notified medical end users on 2/27/13 via email correspondence with a copy of the recall letter attached. This was done to ensure advice was received at the same time that the goods were arriving. Hard copies of the recall letter, which were stamped "copy", were also sent to the individual facilities via UPS, processed for delivery Thursday, 28 February, 2013. Medical end users were advised in the recall letter to contact the US Customer Service staff at (888) 474-7839 x 702 to arrange for the pick-up if they decide to return the unused product. |  | 21 | 21 doses | Nationwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CT, GA, ID, IL,KY, LA, MD, ME, MO, MT, NJ, NV, WA, and WV. | Open |  |  | 21 | Not_Computer
Z-1166-2013 | 64783 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series CT Scanner System, Part number (PN): 989605651321. This is a computed X-ray, Tomography system.NeuViz Dual Multi-slice CT Scanner System is intended to produce cross-section images of head and body by computer reconstruction of X-ray transmission data taken at different angles. | NeuViz Dual series CT Scanner System | 2 | April 23, 2013 | 2013 | Philips And Neusoft Medical Systems Co., Ltd. | User error of the NeuViz 16 or NeuViz Dual may cause personnel injuries. There was a report where a finger was pinched when the table was moved backward while the tabletop was controlled by another person from the gantry left control panel. |  | The firm, Neusoft Medical Systems Co., Ltd., sent an "URGENT DEVICE CORRECTION" NeuViz Dual and NeuViz 16 letter, dated Mar. 4, 2013, to their US consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review this information with all members of their staff and retain a copy of the letter with the equipment Instruction For Use.Customers with questions should contact the Service Support Department by email to nms-service@neusoft.com. |  | 25 | 25 units | Nationwide distribution: USA including states of: California, Connecticut, North Carolina, Nebraska, Ohio, South Carolina, Texas, and Puerto Rico. | Terminated | November 22, 2013 | 213 | 48 | Not_Computer
Z-1161-2013 | 64739 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | VITROSÂŽ Immunodiagnostic Products Anti-HBs (total antibody to hepatitis B surface antigen (anti-HBs)) Reagent PacksFor the quantitative in vitro determination of total antibody to hepatitis B surface antigen (anti-HBs) in human serum. | VITROS AntiHBS Reagent Pack | 3 | April 23, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling certain lots of the Vitros Anti-HBs Reagent packs an inability to obtain a valid calibration. |  | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated November 1, 2012, to all affected customers via USPS. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue and discard any remaining recalled lots of VITROS Anti-HBS Reagent Packs. Foreign affiliates were informed by email on November 1, 2012, of the issue and instructed to notify their consignees of the issue and actions . On March 22, 2013, OCD sent customer letters via FedEX overnight and/or via USPS to US consignees for two additional lots of product being recalled. The letter provided similar recall instructions as the initial notification. Ortho Clinical Diagnostics has requested that their clients discontinue and discard any remaining inventory of VITROS Anti-HBS Reagent Packs subject to this recall. OCD has also requested that recall product recipients also complete a Confirmation of Receipt and return that document to OCD, fax to: 1-888-557-3759 or 585-453-4110.We apologize for any inconvenience this may cause your laboratory. If you have technical questions please contact Customer Technical Services at 1-800-421-3311. If you have questions about your replacement order, please call our Customer Service Center at 1-800-828-6316. | Contact the recalling firm for information | 1464 | 1464 units | Worldwide Distribution- USA (nationwided) and countries of: Australia, Brazil, Canada, Chile, Chao Yang District, Beijing, Columbia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Open |  |  | 1464 | Not_Computer
Z-1178-2013 | 64701 | c-reactive protein, antigen, antiserum, and control | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | The Alere Cholestech LDXÂŽ high sensitivity C-Reactive Protein (hsCRP) Test Cassette, Model Number 12-807. Product Usage: Immunoassay for the determination of C-Reactive Protein (CRP). The Alere Cholestech LDXÂŽ high sensitivity C-Reactive Protein (hsCRP) is an in vitro diagnostic test for the quantitative determination of C-reactive protein in whole blood or serum. | Alere Cholestech LDX hsCRP Test Cassette, | 2 | April 25, 2013 | 2013 | Alere San Diego, Inc. | The Alere Cholestech LDXÂż hsCRPcassette (PN 12-807) may have increased imprecision relative to the performance data in the package insert,which may cause differences in quantitative results for hsCRP compared to a reference method orrepeat testing on the same patient sample. |  | Alere sent an Urgent Medical Device Recall letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to block all stock of the affected lot in their warehouses; discontinue use and discard all stock on hand, complete the provided Customer Verification form to document disposal, contact their primary consignees and advise them of the required discontinuation and disposal of the affected lot. For question call 877-308-8289. | Contact the recalling firm for information | 2333 | 2333 kits | USA Nationwide Distribution | Open |  |  | 2333 | Not_Computer
Z-1176-2013 | 64706 | suture, surgical, absorbable, polydioxanone | NEW | Absorbable polydioxanone surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ORTHOCORD One Violet and One Blue Braided Composite Suture with OS-6 Reverse Cutting Needles, Product Code: 223116.For use in general soft tissue approximation, and/or ligation, including orthopedic procedures. | ORTHOCORD One Violet and One Blue Braided Composite Suture with OS6 Reverse Cutting Needles | 2 | April 25, 2013 | 2013 | DePuy Mitek, Inc., a Johnson & Johnson Co. | The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six | MISBRANDING: Labeling False and Misleading | On March 22, 2013, firm notified consignees of recall via letter. Consignees were informed of the following: The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six weeks after implantation. As Blue and Violet ORTHOCORD free strand sutures share the same IFU, IFU-108147 will be revised. Consignees were requested to confirm the receipt of the letter by March 29, 2013. |  | 46501 | 46501 | Nationwide Distribution. | Terminated | February 24, 2014 | 305 | 449009 | Not_Computer
Z-1180-2013 | 64972 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI)The STERRAD CYCLESURE 24 Biological Indicator (BI), is intended to be used as a standard method for frequent monitoring of the STERRAD Sterilization System cycles. | STERRAD Cyclesure | 2 | April 26, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products is recalling certain lots of STERRAD CYCLESURE 24 Biological Indicator product because it may not have adequate data to support the entire duration of the labeled shelf-life. |  | Advanced Sterilization Products sent an Urgent Medical Device Recall letter dated April 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers that have installed and validated STERRAD system between Feb 2012 and March 2013, are instructed to contact Stericycle at (888) 943-4901 to determine if their STERRAD System is eligible for an optional revalidation. If customers would like to have their system validation repeated, ASP will revalidate their system without charge. Customers are instructed that the revalidation can occur as part of a routine Planned Maintenance (PM) within 12 months. Customers are also notified that the optional procedure will require that their Sterrad System will be unavailable for use for approximately 31 hours (depending on their system) while revalidation is performed. Customers are instructed to contact Stericycle at the number above with their STERRAD System serial number by May 31 2013. Customers are instructed to contact Stericycle at (888) 943-4898 or contact their ASP representative. Stericycle's hours of operations are Mon-Fri , 8am-5pm, ET. | Contact the recalling firm for information | 69994 | 69,994 cases | Worldwide Distribution - Nationwide Distribution including the states of AR, AL, AK, AZ, CA, CT, CO, DE, FL, GA, HI, IA, IL, IN, ID, KY, LA, MA, ME, MD, MI, MS, MN, MO, MT, NC, ND, NE, NJ, NY, NV, NH, NM, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV., and the countries of AR, AU, BE, BR, CA, CL, CN, CO, DO, EC, GT, HK, ID, IN, JP, KR, MV, MX, PA, PE, PH, PR, RU, SG, TH, TW, UY, VE and VN. | Terminated | August 01, 2013 | 97 | 69994 | Not_Computer
Z-1183-2013 | 64504 | system, dialysate delivery, central multiple patient | FKQ | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Diacap(R) Ultra Dialysis Fluid Filter, Catalog Number 7107367.Intended to filter bacteria and endotoxins from dialysis fluid. | Diacap(R) Ultra Dialysis Fluid Filter | 2 | April 26, 2013 | 2013 | B. Braun Medical, Inc. | There is the potential of the residual moisture of the Diacap Ultra membrane being reduced over time, which results in the reduction of the membrane permeability (ultrafiltration coefficient). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | B. Braun Medical Inc. sent and "URGENT RECALL NOTICE" dated February 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Support Department at (800) 848-2066 for questions regarding this issue. |  | 1580 | 1580 | Nationwide Distribution including the states of AZ, CA, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MN, MO, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, WA and WV. | Open |  |  | 1580 | Not_Computer
Z-1186-2013 | 64621 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | SL3 SOFT-TISSUE Laser (Model LR2002); Manuf: Discus Dental, LLC, 8550 Higuera Street, Culver City CA 90232.The SL3 is a soft-tissue diode laser intended to be used for dental procedures | SL3 SOFTTISSUE Laser | 2 | April 29, 2013 | 2013 | Discus Dental LLC | The firm is initiating this recall of the SL3 Soft Tissue Laser equipped with a fiber optic Unifiber system which contains a fragile glass filament because the fiber broke near the hand piece. | DESIGN: Device Design | DISCUS Dental, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 21, 2011. The letter described the product, problem and action to be taken. Customers were instructed tocontact Discus Dental at 1-888-427-9279 for a no-cost replacement.For questions regarding this recall call 310-845-8200. | Contact the recalling firm for information | 150 | 150 units | Nationwide Distribution including AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MO, MT NH, NJ, NM, NV, NY, OH, PA, SD, TX, UT, VA, and WA. | Terminated | April 29, 2013 | 0 | 150 | Not_Computer
Z-1188-2013 | 64889 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspeed VCT, Discovery CT750HD.Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission. | GE, VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspeed VCT, Discovery CT7 | 2 | April 29, 2013 | 2013 | GE Healthcare, LLC | Fastener bolts on the heat exchanger were not properly torqued and part of the heat exchanger can detach from the gantry and hit the inside of the covers and other components of the gantry possibly resulting in the oil conduit breaking, spilling oil within the gantry and outside the system. This issue affects GE Healthcare VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspe | PRODUCTION CONTROLS: Process Control | No consignee letter was sent out. At customer site Field Engineer replaced heat exchanger, the tube and the pump. Field Engineer also verified torque on all bolts. Field action was initiated on 7/25/2011 and has been completed as of 3/29/2013. |  | 166 | 166 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CO, CT, FL,GA, IL, IN, KY, MA, MI, MS,MO, NH, NJ, NM, NY, NC, OK, OR, PA, SC, SD, TX, UT, VA, WA, and WI and the countries of ALBANIA ARGENTINA AUSTRALIA BELGIUM BRAZIL CANADA CHINA COLOMBIA COSTA RICA CZECH REPUBLIC DENMARK ECUADOR EGYPT FINLAND FRANCE GREECE HONG KONG HUNGARY INDIA IRELAND ITALY JAPAN KOREA, REPUBLIC OF MEXICO MOROCCO NIGERIA NORWAY POLAND RUSSIAN FEDERATION SINGAPORE SPAIN SWEDEN SWITZERLAND TURKEY UKRAINE UNITED KINGDOM. | Terminated | February 04, 2014 | 281 | 166 | Not_Computer
Z-1192-2013 | 64919 | applicator, absorbent tipped, non-sterile | KXF | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | MediChoiceÂŽ Rayon Tipped OB/GYN Applicator, 8 inch, Plastic Shaft, Non-Sterile, Item Numbers WOD5001 and WOD5002.Product Usage: Usage:An absorbent tipped applicator is a device intended for medical purposes that consists of an absorbent swab on a wooden, paper, or plastic stick. The device is used to apply medications to, or to take specimens from, a patient. | MediChoice Rayon Tipped OB/GYN Applicator | 2 | May 01, 2013 | 2013 | Owens & Minor Distribution, Inc. | Owens & Minor Distribution, Inc. is conducting a market withdrawal of MediChoice Applicator OB-GYN 8 inch product that does not meet our quality standards because the rayon may loosen or unravel from the applicator during use. | PRODUCTION CONTROLS: Process Control | Owens&Minor; sent an Urgent: Medical Device Recall letter to the first set of customers on March 26, 2013 and to their second set of customers on April 9, 2013. The letter identified the affect product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all affected products, contact service representative to discuss whether affected products should be returned or destroyed, and complete and return the Recall Reply Form. For questions call 804-723-7417. |  | 4298.5 | 4298.5 cases | US Nationwide Distribution. | Open |  |  | 4298.5 | Not_Computer
Z-1191-2013 | 64439 | implanted subcortical electrical stimulator (motor disorders) | MRU |  |  | Neurology | HDE - Humanitarian Device Exemption | Medtronic DBS Therapy for Dystonia Kit, models 3317, 3319, 3337 and 3339. Sterile and Non-Pyrogenic. Product Usage:Dystonia Therapy Kit is indicated for unilateral or bilateral stimulation of the internal globus pallidus (GPi) or the subthalamic nucleus (STN) to aid in the management of chronic, intractable (drug refractory) primary dystonia, including generalized and/or segmental dystonia, hemidystonia, and cervical dystonia (torticollis). | Activa Dystonia Therapy Kit | 1 | May 01, 2013 | 2013 | Medtronic Neuromodulation | There is a potential for lead damage due to the use of the lead cap provided in DBS and Dystonia kits. Medtronic has received reports of DBS leads being damaged at the connector end of the lead when the lead cap is used. The connector end of the lead is the end of the lead connected to the lead extension. Tightening or loosening of the setscrew may twist the setscrew connector block and may damag | PRODUCTION CONTROLS: Process Control | Medtronic sent an Urgent Medical Device Correction letter dated February 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to follow the modified instructions if the lead cap is used during the implant procedure. For questions call Medtronic Neuromodulation Technical Services 1-800-707-0933. UPDATE: Medtronic mailed an Urgent: Medical Device Removal notification to customers on June 07, 2013, who may have unused DBS lead kits that were manufactured prior to the process change to correct the issue. On June 10, 2013 Medtronic representatives began visiting the identified accounts to facilitate removal of any unused product. The Medtronic Rep documented the hospital visits on an Account Specific Customer Confirmation Form and left a copy with the Hospital. |  | 30000 | 30,000 | Worldwide Distribution - USA Nationwide and the countries of: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam. | Open |  |  | 30000 | Not_Computer
Z-1210-2013 | 64942 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics Corp. AxSOS System Instrument.The AxSOS Locking Plate System is designed to treat periarticular or intraarticular fractures of the Distal Femur, Proximal Humerus, Proximal Tibia, and the Distal Tibia. Per the operative technique, " A 4.3 mm Drill (REF 702743) is then used to drill the core hole for the locking screw. Using a fluoroscopy, check the correct depth of the drill, and measure the length of the screw. The Drill Sleeve should now be removed, and the correct length 5.0 mm Locking Screw is inserted using the Screwdriver T20 and Screw Holding Sleeve." | Howmedica Osteonics Corp. | 2 | May 01, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received a report indicating that the scale of the AxSOS Calibrated Drill Bit 4.3,,X263mm, flat , manufactured by Stryker Osteosynthesis, is incorrect. The incorrect scale could result in the surgeon selecting a screw that is 10mm longer than intended. | DESIGN: Device Design | Stryker Orthopaedics sent a Urgent Medical Device Removal letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Please inform users of this Medical Device Removal and pass thisnotice to all those individuals who need to be aware within yourorganization.2. Return all affected products available at your location toStryker Osteosynthesis or Contact Stryker Customer Service.3. Complete and sign the enclosed Business Reply Form and fax a copyto: 1-865-252-3635 .4. Keep a copy of the completed and executed Business Reply Form foryour records.Report any adverse events or product quality problems to StrykerOrthopaedics: 1-866-0R-ASSIST. (1-866-672-7747).As we strive for products that meet your expectations for quality and reliability, please do not hesitate to contact us, in case you have any further questions at (201) 831-5970.. | Contact the recalling firm for information | 107 | 107 units | Nationwide Distribution including the states of AL, AZ, IL, KY, LA, NJ, NC, FL, CA, NM, MO, TN, VA, CO, MA, TX, IN and MT. | Open |  |  | 107 | Not_Computer
Z-1212-2013 | 64885 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR430s MRI Scanner; 5000-0002is indicated to use as a magnetic resonance imaging device of the leg (excluding the thigh), knee, ankle, foot, elbow, forearm, wrist, and hand. | Optima MR430s MRI Scanner | 2 | May 01, 2013 | 2013 | GE Healthcare, LLC | Coolant supply lines may leak liquid onto the floor resulting in a slippery surface and fall hazard. | PRODUCTION CONTROLS: Process Control | GE Healthcare sent an Urgent Medical Device Correction letter dated April 23, 2013, to all affected customers. The letter informed customers of a potential safety issue due to leakage from gradient coolant lines. Consignees were instructed to monitor the floor for leakage and continue to use if leakage is not observed. If leakage is observed consignees are instructed to stop using the system and notify their local service representative. GE Healthcare will correct all affected scanners by providing an update at no cost.Customers with questions were instructed to contact their local service representataive.For questions regarding this recall call 262-513-4122. |  | 81 | 81 | Worldwide Distribution - USA including AL, CA, CO, FL, ID, IL, MA, NH, NJ, NY, OH, PA, SC, SD, TN, TX, UT, VA, WA, WI; Worldwide: Argentina, Australia, Austria, Brazil, Canada, Chile, Denmark, Egypt, Finland, Germany, India, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Peru, Poland, Russia, Singapore, South Africa, Spain, Sweeden, Switzerland, UK, Ukraine, and UAE. | Terminated | October 25, 2013 | 177 | 151 | Not_Computer
Z-1214-2013 | 64820 | dialysate concentrate for hemodialysis (liquid or powder) | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution, Model/Catalog Number: RFP-401, packaged in a case The NxStage PureFLow-B Solution is indicated for use with renal replacement therapy systems that utilize sterile premixed dialysate during hemodialysis. | NxStage PureFlowB Solution | 2 | May 02, 2013 | 2013 | NxStage Medical, Inc. | Product may be mislabeled. | PRODUCTION CONTROLS: Labeling Mix-Ups | NxStage sent a letter dated March 18, 2013, via email to all consignees instructing them to check their entire inventory and to segregate and quarantine lot F023444 pending further instruction from NxStage.On 4/2/13 NxStage sent a letter via email to all consignees requesting a return of all RFP-400 and RFP-401 labeled with lot F023444.. Consignees were instructed to complete the recall reply form and provide a list of customers (name address, phone number) and fax them to NXStage Medical, Inc.Customers with questions were instructed to call 1-866-697-8243.For questions regarding this recall call 1-866-697-8243. |  | 1480 | 1480 cases (2 bags per case) | Nationwide Distribution including IL, SC, TX, VA, FL, AR, and NY. | Open |  |  | 1480 | Not_Computer
Z-1213-2013 | 64859 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer + additive/metal/polymer + additive | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert. The insert is used along with the other components that comprise the 3DKnee System. | Tibial Insert | 2 | May 02, 2013 | 2013 | Encore Medical, Lp | U.S. Agent contacted Director Commercial Logistics, notifying him of incorrect color coded labels on two packages of e+tibial inserts. There were 8 e+ tibial inserts of the same lot number in Finished Goods as the two from the agency. These devices were inspected and all were incorrectly labeled. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm sent their US customer a "Urgent Field Safety Notice" letter on April 1, 2013. The letter explained the product problem and provided recommended actions as well as customer contact numbers. |  | 5 | 5 | US Distribution to the state of California and International Distribution to the country of Venezuela. | Terminated | July 01, 2013 | 60 | 5 | Not_Computer
Z-1215-2013 | 64444 | culture media, for isolation of pathogenic neisseria | JTY | Culture medium for pathogenicNeisseriaspp. | Microbiology | Microbiology | 510(k) | BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack, REF 441124, contains 12 each Microwells labeled in part ***BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack REF 441124 ***For use with the BD Viper System *** Becton, Dickinson and Company 7 Loveton Circle, Sparks, MD 21152 USA www.bd.com/ds EC REP Benex Limited, Pottery Road, Dun Laoghaire, Co. Dublin, Ireland. Assay is indicated for use with asymptomatic and symptomatic individuals to aid in the diagnosis of gonococcal urogenital disease. | BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack | 2 | May 03, 2013 | 2013 | Becton Dickinson & Co. | Increased levels of false positives in BD Probe Tech Neisseria gonorrhoeae Q Amplified DNA assay kits. | PRODUCTION CONTROLS: Environmental Control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter via UPS overnight delivery on February 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to verify that these kits have been consumed, returned or destroyed by completing the enclosed form. The firm will issue a no charge replacement for customers' product. Any product remaining in customers' inventory should be discarded or destroyed per their facility procedures. Complete the attached form regardless if any inventory remains, so that the firm may acknowledge customers' receipt of this notification. It is not necessary to contact BD by telephone to receive replacements. Complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 OR E-mail to RegulatoryComplianceFax@BD.com. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. |  | 77 | 77 kits | Distributed in the states of CA, FL, GA, NJ, and TX. | Terminated | May 13, 2013 | 10 | 77 | Not_Computer
Z-1218-2013 | 64682 | apparatus, suction, operating-room, wall vacuum powered | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Stryker Irrigator Reusable Tips:5mm X 32cm Regular Tip, Vented, Metal Finish.Suction irrigator tips aid in cleaning and clearing the operative site during laparoscopic surgery. | Stryker Irrigator Reusable Tips | 2 | May 03, 2013 | 2013 | Stryker Endoscopy | The pre-vacuum steam sterilization (wrapped method) temperature in the Instructions for Use incorrectly listed 121-123ÂżC , instead of the correct 132-133ÂżC. | PRODUCTION CONTROLS: Error in Labeling | Stryker sent "Urgent Device Correction" letters dated March 27, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. | Contact the recalling firm for information | 147921 | 147,921 - all sizes | Worldwide Distribution: USA (nationwide) and countries of: Mexico, Portugal, Latin America, India, Greece, Taiwan, Singapore, New Zealand, Phillipines, Netherlands, France, Australia, Canada, Japan, Scandinavia, Iberia, Brasil, Romania, South Africa, Chile, Korea, China. | Terminated | September 24, 2013 | 144 | 147921 | Not_Computer
Z-1238-2013 | 64771 | prosthesis, elbow, hemi-, radial, polymer | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Synthes Radial Head Prosthesis: Trial Radial HeadElbow joint prosthesis: Synthes Radial Head Prosthesis System is a two-piece modular system comprised of titanium alloy stem and cobalt chrome components with an integral screw and side-loading application to allow in-situ assembly. | Synthes Radial Head Prosthesis: Trial Radial Head | 2 | May 03, 2013 | 2013 | Synthes USA HQ, Inc. | The Trial Head may come loose from the implant stem during manipulation of the arm during surgery. To date, no adverse events have been reported related to this issue and this recall is not being initiated as a result of adverse events. | DESIGN: Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 22, 2013, to all affected consignees. The letter requested consignees examine their inventory, remove them from use and return them to Synthes. Customers were asked to complete the attached Verification Section at the end of the letter indicating whether or not they have any of the affected product. For questions customers were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. | Contact the recalling firm for information | 417 | 417 | Nationwide Distribution: USA including states of: IL, IN, MA, MI, MO, NJ, NY, OH, OR, PA, SC, TX, VA, WA and WI. | Open |  |  | 417 | Not_Computer
Z-1206-2013 | 64736 | anesthesia breathing circuit kit (adult & pediatric) | OFP | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Presource PBDS, Knee Arthroscopy, Kit, Circulator | Presource PBDS, Knee Arthroscopy, Kit, Circulator | 1 | May 03, 2013 | 2013 | Cardinal Health, Medical Products & Services | Various Presource Kits contain a pre-assembled Filter and Anesthesia Circuit that may contain outer plastic packaging on one or more components. If the packaging is removed without disassembling the components, remnants of the plastic from the packaging material may become lodged in the filter potentially causing an obstruction of airflow. | PRODUCTION CONTROLS: Packaging Process Control | Affected consignees were notified with a letter, dated March 26, 2013, sent via FedEx Priority, of this action. Cardinal Sales Representatives will also be advised and will verbally communicate the action to affected customers. The letter identified the affected product and the problem. Customers are to locate any affected product and quarantine them. In addition, clinicians should be informed of the issue; and, warning labels should be affixed to the front of each kit instructing users to remove and discard the Anesthesia Circuit and Filter Assembly. In order to confirm receipt of this action and completion of labeling the product, the enclosed Acknowledgement Form should be completed and returned to the number provided. If customers do now wish to accept overlabeled product or to overlabel the product in their possession, they are to contact Presource Sales Operations at 800-766-0706, or their sales representative for further instruction. Customers are to contact Cardinal Health Quality Systems at 800-292-9332, if they have any questions or need special assistance. |  | 4571 | 4,571 Kits total | Distributed in the states of AR, MS, SC, and TN. | Open |  |  | 4571 | Not_Computer
Z-1244-2013 | 64705 | catheter, assisted reproduction | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Sure-ProÂŽ Single Stage Embryo Replacement Catheter Product Code: PP623Sure-ProÂŽ Embryo Replacement Catheters are used for the introduction of embryos into the uterine cavity following in vitro fertilization. Sure-Pro UltraÂŽ Embryo Replacement Catheters have the added benefit of being visible under ultrasound. These products can be used in either a one- stage or two-stage procedure. In the one-stage procedure, the clinician pushes the inner catheter through the outer sheath and loads the embryo prior to performing the embryo transfer into the uterus. In the two-stage procedure, the clinician places the outer sheath through the cervical canal into the uterus, then loads the embryo in the inner catheter, and advances the inner catheter through the outer sheath to perform the embryo transfer into the uterus. | Wallace SurePro Single Stage Embryo Replacement Catheter | 2 | May 03, 2013 | 2013 | Smiths Medical ASD, Inc. | Unable to pass catheter through outer sheath during embryo transfer | PRODUCTION CONTROLS: Process Control | The firm, Smiths Medical, sent an "URGENT FIELD SAFETY NOTICE" dated March 28, 2013 to its customers. The notification describes the product, problem and actions to be taken. The customers were provided with two options for proceeding: 1) continue to use product if the product passes a pre-test; or 2) return affected product to Smiths Medical for replacement or credit. In addition, the customers were instructed to inspect their inventory and quarantine all affected products; perform a count of affected products currently in inventory; complete and return the attached Confirmation Form via fax to +1-866-258-6068 or email to pamela.harrell@smiths-medical.com; send a copy of this notice to your customers identified having received affected product; and provide email confirmation of your customers notifications to Smiths medical at pamela.harrell@smiths-medical.com. Customers should report any issues with these products to Smiths Medicals Global Complaint Department at +1-866-216-8806 or globalcomplaints@smiths-medical.com. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at +1-800-258-5361. | Contact the recalling firm for information | 800 | 800 units | Worldwide distribution: USA (Nationwide) and countries of: Arab Emirates (AE) Ireland (IE), Albania, Israel (IL), Argentina (AR),Italy (IT) Armenia (AM) , Japan (JP), Austria (AT) Jordan (JO), Bahrain (BH) Kuwait (KW) Belgium (BE), Lebanon (LB), Belarus , Netherlands (NL), Brazil (BR),Nigeria (NG), Bulgaria (BG) , Pakistan (PK), Canada (CA) Poland (PL), China (CN) Romania (RO), Columbia (CO), Russia (RU), Cyprus(CY) , Saudi Arabia (SA) Czech (CZ) , Singapore (SG), Denmark (OK), South Korea (KR), Egypt (EG), Spain (ES), Finland (FI), Sri Lanka (LK),Germany (DE) Turkey (TR) Ghana (GH) , Uganda (UG), Great Britain (GB) Ukraine (UA), Greece (GR) Uruguay (UY), and Hong Kong (HK). | Open |  |  | 136349 | Not_Computer
Z-1267-2013 | 64962 | instruments, surgical, cardiovascular | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Codman(R) Collins Radioparent Sternal Blade Nylon,4 3/4" (121 mm) long, 3" (76 mm) wide, 1 1/2" (38 mm) deep, REF 50-8081sternum retractor | Codman(R) Collins Radioparent Sternal Blade Nylon | 2 | May 07, 2013 | 2013 | Symmetry Medical/SSI | There is a possibility that the Collins Radioparent Sternal Blades Nylon may break during use. |  | Symmetry Surgical sent an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. If so, customers were instructed to discontinue use of the product immediately and promplty call 800-251-3000 for a return material authorization (RMA) number and to identify if they would like credit or an alternative product. Customers were asked to fill out the attached Recall Response Form and fax the completed form to 615-964-5567. Customers with questions were instructed to call 615-964-5515.For questions regarding this recall call 615-964-5515. |  | 524 | 524 boxes of 10 | Worldwide Distribution - USA including CA, MS, FL, PA, TN, CO, IL, TX, OH, MI, MA, IL and internationally toBelgium, Canada, Germany, UK | Open |  |  | 524 | Not_Computer
Z-1248-2013 | 53487 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Mercury Spinal System DIA 5.5 x 50 MM Lordosed Rod,Part Number: 60100-050.Intended to provide immobilization and stabilization of the spine in skeletally mature patients. | Mercury Spinal System DIA 5.5 x 50 MM Lordosed Rod | 3 | May 07, 2013 | 2013 | Spinal Elements, Inc | Spinal Elements has initiated a recall for the Mercury Spinal System due to a raw material issue. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spinal Elements sent Urgent Medical Device Recall notification letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Chief Technology Officer at Spinal Elements for questions about the recall letter. |  | 68 | 68 units total | Distributed in OH and PA. | Terminated | May 07, 2013 | 0 | 68 | Not_Computer
Z-1269-2013 | 64696 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | A) Product Code 2N3371: Non-DEHP Y-Type Catheter Extension Set, 5.7", Vol. 0.8 mL, 2 Injection Sites, Male Luer Lock Adapter, Single use only Product Usage:This device is indicated for use in blood sampling and the administration of solutions. The InterLink Injection Site feature is intended to eliminate accidental needle sticks when used in conjunction with an InterLink Cannula as part of an I.V. needless access system. | NonDEHP Catheter Extension Set | 2 | May 07, 2013 | 2013 | Baxter Healthcare Corp. | Customers have reported that when separating an individual package from its attached grouping, the adjacent package has opened, compromising its sterile barrier properties. | PRODUCTION CONTROLS: Packaging Process Control | Baxter sent an "Urgent Product Notification" dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taked. Customers are advised to locate and remove all of the affected products from their inventory. If the products were further distributed, consigneees are requested to notify their customers of the recall. Products can be returned for credit by contacting Baxter. An acknowledgement form is required to be compelted by customers. |  | 231573 | A) Product Code 2N3371: 228,200 units | Worldwide Distribution - USA Nationwide and the countries of Puerto Rico, Brazil, Canada and Columbia. | Open |  |  | 1035695 | Not_Computer
Z-1276-2013 | 64988 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ExoShape Soft Tissue Fastener 12 mm x 30 mm REF 1101-00-1230.For use in the fixation of ligaments, tendons or soft tissue grafts to bone. | ExoShape Soft Tissue Fastener 12 mm x 30 mm | 2 | May 08, 2013 | 2013 | Medshape Solutions, INC. | Expiration date on the patient label and date on the product carton were not the same. | PRODUCTION CONTROLS: Error in Labeling | Medshape sent an "URGENT: MEDICAL DEVICE RECALL" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 26 | 26 | Nationwide Distribution including the states of CO, FL, GA, IA, MD, MO, OR, TX, VT, VA, and WI. | Terminated | August 07, 2013 | 91 | 26 | Not_Computer
Z-1272-2013 | 64699 | panels, test, susceptibility, antimicrobial | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens Healthcare Diagnostics Inc. MIcroScan MICroSTREP plus Type 1 B1027-201 panels..MICroSTREP panels are designed for use in determining antimicrobial agent susceptibility of aerobic nonenterococcal streptococci (including Streptococcus pneumoniae) and Haemophilus spp. | Siemens Healthcare Diagnostices Inc. MIcroScan MICroSTREP plus Type 1 B1027201 panels. | 2 | May 08, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Customer complaint investigation confirmed the issue of false skipped wells and false susceptible misreads with A. agalactiae affecting multiple antimicrobial agents on MicroScan Microbiology Systems. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Siemens Healthcare sent a Urgent Field Safety Notification letter dated March 6, 2013, to all affected customers of MicroScan MICroSTREP plus type 1, who process panels in a Walk Away System. The letter identified the product the problem and the action needed to be taken by the customer.As stated in our product Instructions for Use, test results should be interpreted in conjunction with the patients medical history, clinical presentation and other findings. We recommend discussing the content of this letter with your laboratory director regarding the need to review previous test results, conduct patient follow-up, and/or repeat testing for testing conducted in the last twelve months.Confirmation of receipt of this letter is required. Attached you will find a form indicating you have received and understood the information. We would greatly appreciate your assistance in notifying us that you have received the information by taking a moment to complete and return the form below.If you require further information or assistance, please contact your local Siemens Healthcare Diagnostics Representative. Please forward this notification to anyone to whom you may have distributed this product.We apologize for the inconvenience that this situation has caused. For further questions please call 1-800-677-7226 Option 1 (USA/Canada)Thank you for your patience and continued support. |  | 453768 | 21,608 boxes = 432,160 panels | US Nationwide Distribution including Puerto Rico. | Terminated | October 28, 2013 | 173 | 21608 | Not_Computer
Z-1273-2013 | 64898 | clamp, umbilical | HFW | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cypress Medical Products LLC, MediChoice Umbilical cord clamp, sterile; single use; Distributed by Owens & Minor, Mechanicsville, VA. The product is used to clamp the umbilical cord after childbirth. | MediChoice Umbilical cord clamp | 2 | May 08, 2013 | 2013 | Cypress Medical Products LLC | Cypress Medical Products LLC is voluntarily recalling one lot of the Owens and Minor (O&M;) Umbilical Cord Clamps because recent complaints indicate the clamps are breaking upon use. The product may not perform as intended. |  | Cypress Medical Products sent an "Urgent Voluntary Medical Device Recall" notification and an associated response form dated April 15, 2013 to their only customer for this device - Owens & Minor. The product problem was identified and recommended actions provided. |  | 150 | 150 Cases (75,000 Units) | US Distribution to the state of VA only. | Open |  |  | 150 | Not_Computer
Z-1278-2013 | 64913 | accessories,arthroscopic | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek Tissue Liberator -Blade UpCatalog Number: 214623The Reusable Arthroscopy Instruments are reusable, non-powered surgical instruments used during arthroscopic surgery of any joint except the spine. | DePuy Mitek Tissue Liberator Blade Up | 2 | May 09, 2013 | 2013 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Labeled incorrectly as a Blade Up configuration instead of a Blade Down configuration. | PRODUCTION CONTROLS: Process Control | DePuy Synthes sent an Urgent Voluntary Product Recall letter dated April 15, 2013, to all affected customers. The letter iidentified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check all inventories to locate and return the affected product. For returns customers were instructed call 1-877-379-4871 to obtain a goods authorization (RGA) number. Affected product should be returned to:Mitek Sports MedicineATTN: Mitek Complaints325 Paramouont DriveRayham, MA 02067For questions customers were instructed to call 1-877-379-4871. |  | 27 | 27 units | Worldwide Distribution - USA including KY, NJ and Internationally to Argnetina, Australia, Belgium,Columbia, India, France, Japan, Jordan, Latvija, Israel, Korea, and Great Britain.For questions regarding this recall call 508-977-3813. | Open |  |  | 27 | Not_Computer
Z-1282-2013 | 64778 | dura substitute | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Dural Graft Matrix DuraGenÂŽ Dural Graft Matrix is an absorbable implant for repair of dural defects. DuraGen matrix is an easy to handle, soft, white, pliable, nonfriable, porous collagen matrix. DuraGen matrix is supplied sterile, nonpyrogenic, for single use in double peel packages in a variety of sizes.. | DuraGen Graft Matrix | 2 | May 10, 2013 | 2013 | Integra LifeSciences Corporation | Due to the process deviation, product lots in question may have been released with higher levels of pyrogens than permitted by the specifications for the products. | PRODUCTION CONTROLS: Process Control | The firm, Integra, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 9, 2013, to all US, Canada, and non-EU Integra consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify if they have any lots of the listed product, if so, STOP using that product; remove the product from service; and promptly complete and return the PRODUCT RECALL RETURN ACKNOWLEDGMENT FORM via Email: FCA@integralife.com or FAX to : 1-609-275-9445. Customer Service department will contact you to return and replace the product."However, if you have already implanted or used collagen sponge products affected by this recall, we recommend you monitor the patient for a fever in the immediate postoperative period according to the standard hospital or clinician protocol."For assistance or any other questions that you may have, please contact Customer Service at 1-855-532-1723. |  | 15607 | 15,607 boxes | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Arab Emirates, Argentina, Belgium, Brasil, Canada, Chile, China, Costa Rica, Denmark, Egypt, Spain, France, Great Britain, Italy, Japan, Peru, Singapore, Thailand, and South Africa. | Open |  |  | 552525 | Not_Computer
Z-1281-2013 | 64887 | cannula, ophthalmic | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Insight Instruments MIS ( Multi-Port Illumination System) Illuminated pilot tubes work like a ceiling light while simultaneously protecting the entry sites. Illumination automatically follows the movement of the two surgical instruments inserted into the M.I.S. cannulas. The large distance of more than 16mm between cannula ends and posterior pole guarantees sufficient light in the working area and a large observation field especially when used in conjunction with modern wide field panoramic observation systems. | Insight Instruments MIS ( MultiPort Illumination System) | 2 | May 10, 2013 | 2013 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling M.I.S. (Multi-Port Illumination System), Ophthalmic Cannula due to a misclassification of the medical device. |  | Insight Instruments, Inc. sent a recall letter dated March 15, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to do the following: 1)Perform a physical count of their inventory immediately for any of the recalled lots and complete and return the Return Response Form. 2) If they have any products listed, contact the firm at (800) 255-8354, to arrange for return and prompt replacement. 3) Discontinue using and dispensing all M.I.S. devices of the recalled lots. |  | 776 | 776 | US Distribution IN and IL | Open |  |  | 776 | Not_Computer
Z-1280-2013 | 64875 | device, incontinence, mechanical/hydraulic | EZY | Implanted mechanical/hydraulic urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology |  | AMS 800ÂŽ Urinary Control System. The AMS 800ÂŽ Urinary Control System is used to treat urinary incontinence due to reduced outlet resistance (intrinsic sphincter deficiency) following prostate surgery. The AMS 800ÂŽ Urinary Control System is an implantable, fluid filled, solid silicone elastomer device used to treat stress urinary incontinence. It is designed to restore the natural process of urinary control. The device simulates normal sphincter function by opening and closing the urethra at the control of the patient. The AMS 800 system consists of three inter connected components: a cuff, a pump, and a pressure-regulating balloon. The three components are connected with kink-resistant tubing. The AMS 800 can be implanted at either the bulbous urethra or the bladder neck. | AMS 800 Urinary Control System | 2 | May 10, 2013 | 2013 | American Medical Systems, Inc. | American Medical Systems, Inc. is initiating a recall on two components of the AMS 800Âż Urinary Control System, Control Pump & Pressure-Regulating Balloon. These products were released despite not meeting all final product release criteria. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | American Medical Systems sent two written communications distributed within the United States. The first is a letter that is being delivered to customers who purchased the affected lots using their ship to address from the original sales order. The second is a letter being delivered to implanting physician customers who are known by AMS to be regular users of the devices. Similar letters will be sent to international customers, including both end users and distributors (where applicable). American Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letters identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to return the recalled product and to complete the Product Recall Acknowledgement Form. The letters addressed to "Physicians" advised consignees against prophylactic removal of implanted control pumps and pressure-regulating balloons contained within the scope of the product recall.For further questions please call AMS Customer Service at 1 (800) 328-3881. |  | 77 | 77 | Worldwide Distribution -- Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY,NC, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VT, VA, WA and WI., and the countries of AUSTRIA, AUSTRALIA, CANADA, BELGIUM, SWITZERLAND, CZECH REPUBLIC, GERMANY, DENMARK, ESTONIA, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, IRELAND, ITALY, LEBANON, NETHERLANDS, NORWAY, SWEDEN, TURKEY, ARGENTINA, BRAZIL, IVORY COST, COLOMBIA, ECUADOR, INDIA, JAPAN, MEXICO, NEW ZEALAND, PANAMA and VENEZUELA. | Open |  |  | 682 | Not_Computer
Z-1305-2013 | 55358 | resectoscope | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Electrode, cutting, 24FR (27050G/6).Resectoscope accessories (ie. electrode, cutting loops and cold knives). | Karl Storz Electrode, cutting | 2 | May 11, 2013 | 2013 | Karl Storz Endoscopy America Inc | Possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products, which may result in airborne contamination. | PRODUCTION CONTROLS: Packaging | A Safety Alert dated May 29, 2009 was sent to affected customers. The letter informed them of possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products which may result in airborne contamination. Karl Storz requested consignees to inspect their inventory of the subject products thoroughly and look for small holes in the blister trays. Any identified units should be returned to Karl Storz, and a replacement package (6 units per package) will be provided to at no charge. The Karl Storz Customer Support number 1-800-421-0837 ext. 7577 was provided to obtain a Return Authorization Number (RMA) and arrange replacement. Questions and comments were forwarded to Technical Support at 1-800-421-0837 ext. 5350. |  | 9432 | 9432 units | Distributed Nationwide and in Canada. | Terminated | May 13, 2013 | 2 | 9432 | Not_Computer
Z-1304-2013 | 64908 | kit, surgical instrument, disposable | KDD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Curity Staple Remover Kit; Product Code: 66701.Staple Remover Kit. | Staple Remover Kit | 3 | May 11, 2013 | 2013 | Covidien LLC | On March 26, 2013 Covidien received a recall notice from one of their suppliers CareFusion. CareFusion is recalling a variety of OTC drug products including the 10% Povidone Iodine SEPP Applicator which is utilized in specific convenience kits. | PRODUCTION CONTROLS: Error in Labeling | US Customers were notified of the recall by letter sent by Federal Express on April 12, 2013. The letter requests that they immediately stop using any affected product, quarantine it, and return it to Covidien. Consignees were asked to return a verification form and return it to Covidien via fax. |  | 52363 | 52363 distributed | Worldwide distribution: USA (nationwide) and in the country of: Panama. | Open |  |  | 70103 | Not_Computer
Z-1317-2013 | 65112 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | PadPro(R) Adult Radiotransparent Multifunction Electrodes, REF/Catalog No. 2001M-C. The PadPro(R) Adult Radiotransparent Multifunction Electrodes with Carbon Leadwires and Medtronic(R) Connector are intended to be used in external pacing, defibrillation and monitoring applications | PadPro(R) Adult Radiotransparent Multifunction Electrodes | 2 | May 13, 2013 | 2013 | ConMed Corporation | ConMed was notified by the contract manufacturer that their wire/connector supplier reported damage to the connector tooling which could pinch or compress the lead wire filaments and potentially render the device incapable of delivering the appropriate energy or shock to the patient and/or could cause potential injury to user or patient. | DESIGN: Component Design/Selection | ConMed sent a URGENT: Field Safety Notice letter dated May 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review your inventory for any of the devices listed on Attachment I.Business Reply Forms were sent to the domestic and foreign consignees on May 3, 2013. The domestic notices were mailed by USP Priority Mail pouches. The recall is being conducted to the end user level.We ask that you contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please call 1 315-624-3237 orfax to +1 315-624-3225. | Contact the recalling firm for information | 600 | Domestic: 49,010; Foreign: 600 | Worldwide distribution: USA (nationwide) and the country of: Saudi Arabia. | Open |  |  | 49610 | Not_Computer
Z-1319-2013 | 65034 | knife, ear | JYO | Ear, nose, and throat manual surgical instrument. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Summit Medical, BL-2006, Myringotomy Lance (Juvenile) Blade, Sterile P, Summit Medical, Inc., 815 Northwest Parkway, Suite 100, St. Paul, MN 55121Use to make an incision in the tympanic membrane. | Summit Medical, BL2006, Myringotomy Lance (Juvenile) Blade | 2 | May 14, 2013 | 2013 | Summit Medical, Inc. | Summit Medical has initiated a recall of Myringotomy Lance Blade (Juvenile) because the box contains the wrong blade. The box contains Myringotomy Spear Blade (Juvenile). | PRODUCTION CONTROLS: Process Control | Summit Medical sent a Medical Device Recall letter dated March 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately examine your inventory, and quarantine the product subject to this recall. Please complete and return the enclosed "Recall Response Form" as soon as possible. The fax number is (651) 789-3961.For further questions please call (651) 789-3921. |  | 9 | 9 boxes | Distributed in MS, NY, PA and TX. | Terminated | August 06, 2013 | 84 | 9 | Not_Computer
Z-1320-2013 | 64650 | antibody igm,if, cytomegalovirus virus | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | QIAGEN artus CMV RG PCR ASR (96) (Catalog number 4503225) Product Usage: Test is intended for use in the detection and quantitation of human cytomegalovirus (CMV) specific DNA, as part of an in house validated real-time PCR assay. One service package containing the handbook. | Qiagen artus CMV RG PCR ASR | 2 | May 14, 2013 | 2013 | QIAGEN Gaithersburg, Inc. | CMV RG PCR kits may contain a mixture of incorrect vials among the correct vials. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Qiagen Gaithersburg, Inc. issued a PRODUCT FIELD ACTION letter dated February 4, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the vials prior to use to ensure that the vials are labeled "CMV RG Master" Vials labeled "EBV TM Master" should NOT be used and count the vials and discard per internal laboratory procedures. Qiagen will replace free of charge or credit any kits discovered with incorrect master mix vials. Please complete page 2 of this notice as soon as possible and email or fax it back to us at: (661) 702-3854, techservice-na@qiagen.com. If you have additional questions or concerns polease do not hestiate to call QIAGEN Technical Services at 1-800-344-3631 option 2 or by e-mail techservice-na@qiagen.com. |  | 90 | 90 kits in total | USA Nationwide Distribution including the states of: CA, DC, FL, GA, MA, MD, MI, and IL. | Terminated | July 01, 2013 | 48 | 90 | Not_Computer
Z-1322-2013 | 64221 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Proplege Coronary Sinus Catheter model PR9, 11 French Introducer, kitted with PR9. Product Usage: The catheter introducer sheath kit provides access to the central venous circulation to facilitate catheter and guidewire insertion. The Proplege coronary sinus catheter provides retrograde cardioplegia during cardiopulmonary by-pass surgery. | Proplege Coronary Sinus Catheter | 2 | May 14, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain Proplege Coronary Sinus Catheter, model PR9 due to confirmed reports of blood leakage through the hemostasis valve in the introducer that is kitted with the catheter. | DESIGN: Device Design | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated February 1, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review entire inventory for the affected product, quarantine and return to Edwards. Contact Customer Services at 800-424-3278 for any question, to obtain a Returned Goods Authorization number and replacement product. |  | 807 | 807 | Worldwide Distribution - USA Nationwide including the states of: MD, NY, PA, CA, FL, AZ, AL, TX, VA, OH, MO, NC, GA, TN, CO, OR, SC, ME, IL, HI, NJ, WA, WI, CT, KY, LA and Canada. | Open |  |  | 807 | Not_Computer
Z-1323-2013 | 64390 | ventricular (assisst) bypass | DSQ |  |  | Cardiovascular | PMA | REF 1103***HVAD Pump Implant Kit.Used as a bridge-to-cardiac transplantation in patients who are at risk of death from refractory end-stage left ventricular heart failure. | HVAD Pump Implant Kit | 2 | May 14, 2013 | 2013 | Heartware Inc | HeartWare, Inc. of Miami Lakes, FL is recalling their Ventricular Assist System due to the loosening of the HVAD Pump driveline connector. | DESIGN: Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their customers on February 7, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. They were asked to sign and return the acknowledgement form to HeartWare. For questions contact your HeartWare representative or Customer Service at 1-877-367-4823. |  | 3544 | 3,544 | Worldwide Distribution-USA (nationwide) including the states of MN, IL, MI, FL, WI, TX, MO, CA, MA, TN, OH, NC, GA, VA, UT, KY, MD, IN, PA, NY, NJ, LA, OR, WA, AL and CO, and the countries of Poland, Australia, Austria, Italy, Germany, New Zealand, South Africa, Czech Republic, France, Brazil, Switzerland, Singapore, United Kingdom, Saudi Arabia, Netherlands, China, Chile, Belgium, Canada, Norway, Lebanon, Israel, Malaysia, Turkey, India, Argentina, Greece, Japan and Lithuania. | Terminated | July 09, 2013 | 56 | 3544 | Not_Computer
Z-1274-2013 | 64975 | stent, superficial femoral artery, drug-eluting | NIU |  |  | Cardiovascular |  | ZilverÂŽ PTXÂŽ Drug-Eluting Peripheral Stent - 6 mm X 80mm; 125 cm, Rx, SterileProduct Usage:The Zilver PTX Drug -Eluting Peripheral Stent is a self expanding, small metal, mesh tube (stent) with the outer surface coated with the drug Paclitaxel that can be Implanted in an artery in the thigh (femoropopliteal artery). The Paclitaxel coating helps Preven the artery from narrowing again (restenosis). The Zilver; PTX stent acts as a scaffold to: Hold open a narrowed artery in the thigh (femoropopliteal artery) caused by PAD . improve blood flow to the extremity. | Zilver PTX DrugEluting Peripheral Stent | 1 | May 15, 2013 | 2013 | Cook, Inc. | Cook Medical has received a small number of complaints related to the delivery system for the Zilver PTX Drug-ElutingPeripheral Stent involving fractures of the delivery system inner catheter after stent deployment, and separation of theinner catheter tip section (13 complaints with an occurrence rate of 0.043%) Two adverse events, including one death,occurred in cases where a tip separation | PRODUCTION CONTROLS: Process Control | Cook Medical sent an Urgent Medical Device Recall letter dated April 24, 2013 to all affect customers. The letter identified the affected products, description of the problem, and actions to be taken. The letter instructed customers to stop use, quarantine, and return all affected products. Customers were asked to complete product reply form and return affected products for credit to Cook Medical, ATTN: Retrun Goods/RGA # 2013C0003. For questions contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1057 | 1,057 units in USA | Worldwide Distribution - USA Nationwide; in the countries of Japan, Europe, Brazil, and Australia | Terminated | December 05, 2013 | 204 | 1592 | Not_Computer
Z-1328-2013 | 64910 | urinalysis controls (assayed and unassayed) | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Iris Diagnostics Division iQ Control/Focus Set, Part Number 800-3104, Lot Number 153-12.Product Usage: For in vitro diagnostic use with the iQ200 series. | Iris Diagnostics Division iQ Control/Focus Set | 2 | May 15, 2013 | 2013 | Iris Diagnostics | The firm initiated this recall because complaints from customers reported that the recovery compared to assigned assay values is significantly below specification. This results in the instrument not passing controls and the customer not being able to run patient samples. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | IRIS sent a Medical Device Recall letter dated September 18, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to dispose of all unopened or partially used inventory, complete the attached worksheet certifying the disposal of the affected product and email or fax to Iris Diagnostics. |  | 942 | 942 boxes | Worldwide Distribution - USA Nationwide and the countries of: Canada, Malaysia, Spain, Switzerland, Taiwan, Peru, France, Russia, Japan, United Kingdom, Germany, and Paraguay. | Terminated | January 27, 2014 | 257 | 942 | Not_Computer
Z-1330-2013 | 64959 | multi-analyte controls, all kinds (assayed) | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | Liquichek" ToRCH Plus IgM Control - PositiveLiquichek ToRCH Plus IgM Control is intended for use as an unassayed quality control serum to monitor precision of IgM laboratory testing procedures for the analytes listed in this package insert. This product is not intended for use in blood donor screening assays. Analytes claimed with this lot of Liquichek ToRCH Plus IgM Control: ÂŽ Cytomegalovirus (CMV) IgM ÂŽ Herpes Simplex Virus Type 1/2 (HSV-1/2) IgM ÂŽ Rubella Virus IgM ÂŽ Toxoplasma gondii IgM | Liquichek ToRCH Plus IgM Control | 3 | May 16, 2013 | 2013 | Bio-rad Laboratories, Inc. | The Liquichek" ToRCH Plus IgM Control - Positive, contains detectable levels of HCV RNA which originatedfrom the non-specific human IgM component used in the mousemonoclonal antibody conjugation process for this product. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Bio-rad Laboratories initiated this recall on April 18, 2013, via telephone call and certified mailing to customers. The notification letter titled "Important Product Information Notice of Correction" was accompanied by a "Field Correction Notification Acknowledgement" form that notified customers of the product for recall (description and codes), issue (This product contains detectable levels of HCV RNA which originated from the non-specific human IgM component used in the mouse monoclonal antibody conjugation process for this product), Customers were instructed to immediately discontinue use this lot of Liquichek ToRCH Plus IgM Control. Discard all remaining inventory of this lot of Liquichek ToRCH Plus IgMControl that you have at your location in accordance with therequirements of your local waste management authorities. Handle this product with the same precautions used with patientspecimens. In accordance with good laboratory practice, all humansource material should be considered potentially infectious. Have the appropriate laboratory personnel complete theacknowledgment form and email or fax it to us. It is very important to usto know that you have received the information provided in this Notice.Contact information (1-800-2-BIORAD). | Contact the recalling firm for information | 211 | 211 boxes and 5 MiniPaks | Worldwide Distribution - Distributed in San Juan Puerto Rico and the states of NM, NY, MN, OR and WI., and the countries of Australia, Brazil, France, Italy, Korea, and New Zealand. | Terminated | August 01, 2013 | 77 | 211 | Not_Computer
Z-1329-2013 | 64800 | enzyme immunoassay, cocaine and cocaine metabolites | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | PROFILE-V MEDTOXScan Drugs of Abuse Test System.For the rapid, qualitative detection of drugs of abuse in human urine, | PROFILEV MEDTOXScan Drugs of Abuse Test System | 2 | May 16, 2013 | 2013 | Medtox Diagnostics Inc | Review of validation information and customer complaints which indicated preliminary screen positive results for cocaine that were not positive on confirmatory testing. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | MEDTOX issued "URGENT MEDICAL DEVICE RECALL" letters dated April 3, 2013 to consignees informing them of the recall. The letter requested that they examine their inventory and immediately discontinue all use and distribution of the affected lots. The letter also requested that they immediately inform clinicians of the recall at the testing sites that have utilized the recall product. The letter also included a Customer Verification Form which was to be returned. |  | 221625 | 221,625 | Worldwide distribution: USA (nationwide) and in the country of: Canada. | Terminated | September 12, 2013 | 119 | 221625 | Not_Computer
Z-1331-2013 | 65078 | dress, surgical | FYE | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | DePuy LifeLiner Stick & Cut Resistant Gloves Rt, LgPart No. 5200-68-000The gloves are designed to resist sticks and cuts and are not cut or puncture proof. These gloves are not required by any surgical technique. | LifeLiner | 3 | May 16, 2013 | 2013 | DePuy Orthopaedics, Inc. | LifeLiner Stick and Cut Resistant Gloves are being recalled because cases and individual glove packages were mispackaged, with right handed, large sized glove packaging containing left handed, large sized gloves. | PRODUCTION CONTROLS: Labeling Mix-Ups | DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.,sent a " Urgent Information Medical Device Recall Notice" dated May 8, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please cease further distribution or use of the affected gloves immediately. Our records indicate your hospital purchased the affected lot. The purpose of this communication is to inform you of this recall and requestacknowledgement of receipt of this letter by signing and returning the Hospital/User Facility Reconciliation Form within 10 business days of receiving this notice.Response cards should be faxed to (574) 371-4964. Please contact your DePuy Synthes Joint Reconstruction Sales Representative for assistance returning any unused affected gloves.For product related questions, please contact your local DePuy Synthes Joint Reconstruction Sales Consultant. For clinical questions from surgeons, please contact DePuy Synthes Joint Reconstructions Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.- 5 p.m. EDT). For questions about device recall information provided, please call (574)371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 180 | 180 gloves (120 gloves distributed in the US) | Worldwide Distribution - US (nationwide) including the states of IL, OH, NC, SC, MN, NV, TN and WA., and the countries of UK, Netherlands and Canada. | Terminated | January 14, 2014 | 243 | 180 | Not_Computer
Z-1334-2013 | 64883 | culture media, non-selective and non-differential | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(tm) CDC Anaerobe 5% Sheep Blood Agar.Used for culturing microorganisms. | BBL CDC Anaerobe 5 Sheep Blood Agar. | 2 | May 16, 2013 | 2013 | Becton Dickinson & Co. | Microbiological media may be contaminated with bacteria. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | BD Diagnostics notified consignees with a "Field Correction Notification" sent on 3/11/13 to distributors, direct accounts, and distributors' customers as available, via fax and email, and followed by a letter sent via UPS. The notice identified the product, the problem, and the action to be taken by customers. BD Diagnostic Systems informed customers of the bacterial contamination of some microbiological media, including BBL Trypticase Soy Agar with 5% Sheep's blood, part number 221239, lot 2348199; part number 221261, lots 2339019, 2348108, 2348208, 2348074, 2348157, 2355145, 2348202, and 2361011; BBL Trypticase Soy Agar with 5% Sheep's blood plus MacConkey II-I plate, part number 221291, lots 2355138 and 2355164; and BBL CDC Anaerobe 5% Sheep blood agar, lot 2361003. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. The notice included a response form to indicate receipt of the notice. BD Diagnostic Systems contacted distributors and end-users via fax, email, and letter on 4/25/13 to notify them of a recall expansion to include lot 3038371 from part/catalog number 221261. The recall notice included an acknowledgement form to document receipt. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. | Contact the recalling firm for information | 521160 | 521,160 total pieces (all three recalled products) | Worldwide Distribution-USA (nationwide) including Puerto Rico, Virgin Islands, and Guam and the countries of Columbia, Japan, Singapore, Canada, Turks, Calicos Islands. | Terminated | June 24, 2013 | 39 | 521160 | Not_Computer
Z-1336-2013 | 64468 | lenses, soft contact, daily wear | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR, Soft, disposable contact lenses.The Vistakon (galyfilcon A) Contact Lens is indicated for daily wear for the correction of refractive ametropia (myopia and hyperopia) in aphakic or non-aphakic persons with non-diseased eyes who may have 1.00D of astigmatism or less. | ACUVUE ADVANCE Brand Contact Lens | 2 | May 17, 2013 | 2013 | Johnson & Johnson Vision Care, Inc. | Limited number of individual contact lens packages may not have been completely sealed. | PRODUCTION CONTROLS: Packaging | The firm, Johnson & Johnson, Vision Care, Inc. sent a letter dated January 9, 2013 to all their consignees stating the following actions:" Please contact your patients that may have received any of the affected product and instruct them to discontinue use immediately and return product to you for replacement " Please return the enclosed Business Reply Card (BRCl confirming for our records that you have received this notice, and indicating whether or not you have in stock, or have distributed to patients,any product from the lots noted above. " Please return any remaining affected product in your possession using the enclosed pre-paid mailer. " Pass this notice on to anyone in your practice that needs to be informed " Customer service will arrange for the return and replacement of any affected product. If you have any questions and/or if replacement lenses are needed, please contact customer Service (1-800-87 4-5278). | Contact the recalling firm for information | 3546 | 3546 lenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico; and countries of: Austria, Belgium, Canada, Costa Rica, Denmark, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea, Japan, Luxembourg, Netherlands, Portugal, Norway, South Africa, Slovenia, Spain, Sweden, and United Kingdom. | Terminated | July 05, 2013 | 49 | 39876 | Not_Computer
Z-1337-2013 | 64901 | calibrator, multi-analyte mixture | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products CK-MB Calibrators, REF 128 6293, IVD, Ortho-Clinical Diagnostics High Wycombe HP12 4DP, UK For use in the calibration of the VITROS ECi/ECiQ Immunodiagnostic Systems, the VITROS 3600 Immunodiagnostic System and the VITROS 5600 Integrated System for the quantitative measurement of CK-MB in human serum and plasma (EDTA or heparin). | VITROS CKMB Calibrators | 3 | May 17, 2013 | 2013 | Ortho-Clinical Diagnostics | Internal testing demonstrated that opened-reconstituted VITROS CK-MB Calibrators stored at -20 degrees Celsius (-4 degrees Farenheit) generated invalid calibrations that were detected by positively biased quality control results. | PRODUCTION CONTROLS: Process Control | Ortho Clinical Diagnostic sent a " Important Product Correction Notification" letter dated April 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed of the issue and advised they must discontinue storage of opened-reconstituted vials of VITROS CK-MB Calibrators frozen at -20 degrees Celsius (-4 degrees Farenheit). The same day, the Distributor letters were sent via US Postal Service Priority Mail or FedEx overnight courier to US Distributor consignees who were informed of the issue and instructed to notify their customers of the issue and actions. Also on April 3, 2013, the foreign affiliates were informed of the issue by e-mail and instructed to notify their consignees of the issue and actions. Confirmation of Receipt - Important Response Required so that we can complete our records, please return this form to us no later than April 19, 2013.FAX: 1-888-557-3759 or 1-585-453-4110We apologize for any inconvenience this may have caused your laboratory. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. | Contact the recalling firm for information | 3019 | Domestic: 3019 units, Foreign: 905 units | Worldwide Distribution - Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open |  |  | 3969 | Not_Computer
Z-1341-2013 | 64943 | mesh, surgical, synthetic, urogynecologic, for apical vaginal and uterine prolapse, transabdominally placed | OTO | Surgical mesh. | General & Plastic Surgery | Obstetrics/Gynecology | 510(k) | Rostorelle L, Polypropylene Mesh, Catalog number 501440, Manufacturer Coloplast A/S 3050 Denmark. | Rostorelle L, Polypropylene Mesh | 2 | May 20, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is recalling one lot of Restorelle L prosthesis because it is mislabeled. The incorrect Instructions For Use (IFU) pamphlet is included in the packaging. | PRODUCTION CONTROLS: Labeling Mix-Ups | Coloplast sent a letter to dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included an attachment with the correct IFU to replace the incorrect IFU. Customers with questions were instructed to call Customer Service at 800-258-3476.For questions regarding this recall call 612-302-4992. |  | 111 | 111 | Nationwide Distribution including AL, CA, CO, DC, IL, IN, KY, MD, MA, MO, NH, OH, RI, and VA. | Terminated | January 21, 2014 | 246 | 111 | Not_Computer
Z-1344-2013 | 49026 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cardinal Health hCG Cassette Rapid Test, Part Number: B1077-22 (Mfg PN: 4582105026)The SP hCG Cassette Rapid Test is a rapid chromatographic immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine, as an aid in the early detection of pregnancy. This test is for professional use. | SP hCG Cassette Rapid Test | 3 | May 21, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego Inc. is recalling the Cardinal Health hCG Cassette Rapid Test because some of them were incorrectly packaged into Cardinal Health hCG Combo Rapid test kit boxes. The kit box incorrectly indicates the product may be used for both urine and serum sample types. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 21, 2013 to all their customers. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of the product; immediately discard all units of the the product in accordance with your local regulations; share this information with your laboratory staff and retain this letter for your records; if you have forwarded the product to another laboratory, provide a copy of this letter to them and include the name, address and a contact for the laboratory on the enclosed Customer Verification Form; complete and return the Customer Verification Form within 10 days via fax to: Technical Service at 858-805-8457 or e-mail to responses.ts@alere.com. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at lateral.flow.support@alere.com. |  | 2167 | 2,167 kits | Nationwide distribution: USA including states of: AL, AK, AR, AZ, CA, CO, CT, DE, HI, IA, ID, IL, IN, FL, GA, KY, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, UT, TN, TX, VA, WA, WI, WV AND WY. | Terminated | August 28, 2013 | 99 | 2167 | Not_Computer
Z-1342-2013 | 64512 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Smart SWL150 kg Product Usage:To lift and transport a patient in the horizontal or other required position from one place to another. | Molift Smart SWL | 2 | May 21, 2013 | 2013 | Moller Vital | The Molift Smart (now known as the Molift Smart 150) is a foldable, mobile hoist with a safe working load ("SWL") of 150 kg (approximately 330 lbs.) and is intended for lifting and transferring a disabled person. The hoist is made out of lightweight materials for ease of operation. It has castors and is battery-powered. The Molift Smart includes a leg spreading mechanism to give optimal functiona | TRAINING: Employee Error | Molift Group AS issued a Field Safety Notice letter to all affected customers. The letter identified the affected products, problem and actions to be taken to the Smart hoists that are already in use. Customers were instructed to make sure that all affected customers receive this information and encorage to increase attention until the affected product has been upgraded. For questions contact your local Molift representative or Customer Service at: groupsales@molift.com. |  | 15110 | 7,555 units subject to review; US: 1,280 units - Foreign: 6,275 units | US Nationwide Distribution: CA, FL, GA, ID, IL, KY, MA, ME, MI, MN, NC, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, VT, WA and Hawaii.Foreign Distribution:: Australia, Austria; Belgium, Denmark, Finland, France, Germany, Iceland, Israel, Italy, Japan, Netherlands, New Zealand, Norway, Spain, Sverige, Switzerland, United Kingdom. | Open |  |  | 7555 | Not_Computer
Z-1179-2013 | 65018 | unit, x-ray, intraoral | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | Fona Srl IntraOs 70 intra-oral x-ray systemIntraOS 70 is an extra oral source X-ray system for dental radiographic examination and diagnosis of diseases of the teeth. | FONA Blue X Imaging Srl | 2 | May 21, 2013 | 2013 | Chicago X-Ray Systems, Inc. | The IntraOs 70 system was found to have missing or incomplete system labels as required by 21 CFR 1010.2 & 1010.3. Also, the user's manual was found to have missing or improperly labeled information as required by 21 CFR 1020.30(h). | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | FONA will have upgraded copies of the manuals that will be made available on the firms website for free download and may also be mailed to each customer or end user should they request it. They will also provide users and distributers with corrected system labels to be placed on the x-ray control, beam limiting device, and tube housing assembly along with instructions for the proper application of these labels. They will provide users and distributers with an amended "Technical Data" page for insertion int the user's manual.Further questions please call (847) 459-3889. | Contact the recalling firm for information | 1960 | 1960 units | Worldwide Distribution- USA (nationwide) and Latin America. | Open |  |  | 1960 | Not_Computer
Z-1347-2013 | 64618 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | PROGRESSIVE MEDICAL Laparoscopic Electrode PTFE Coated L Hook 33cm.Intended for the cutting of soft tissue and coagulation. | PROGRESSIVE MEDICAL Laparoscopic Electrode | 2 | May 22, 2013 | 2013 | Unimed Surgical Products, Inc. | Firm is recalling due to concerns regarding changes that were made to the PTFE coating materials. |  | Unimed Surgical Products, Inc. issued a Medical Device Correction and Removal letter dated February 15, 2013 to its customers. Customers were instructed to quarantine any undistributed product, and notify its customers of the recall to retrieve any unused product. For further information, contact Unimed Surgical Products Inc. at 1-800-531-2627. |  | 120416 | 120,416 units | Distributed in the states of MO, TN, and PA. | Terminated | August 07, 2013 | 77 | 120416 | Not_Computer
Z-1355-2013 | 64931 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Progressive Medical PTFE Coated Disposable Electrosurgical Electrodes -Laparoscopic Electrode PTFE Coated Curved Blade 33cm Length, Product number P0019 Rx Only. Sterilized by Radiation. Manufactured for Progressive Medical, Inc., St. Louis, MO 63128. Electrodes used in Electrosurgeries | Progressive Medical Laparoscopic Electrode PTFE Coated Curved Blade | 2 | May 22, 2013 | 2013 | Progressive Medical Inc | Product coatings require 510K approval. | PRODUCTION CONTROLS: Manufacturing Material Removal | Progressive Medical Inc. sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and segregate the affected units prior to return to Progressive Medical, Inc. If product was further distributed customers should notify their accounts of ther ecall and return the products. Customers were asked to complete and return the attached form by fax to 314-961-4535, even if they had no inventory remaining. Customers with questions were instructed to call 1-800-969-6331.For questions regarding this recall call 314-961-5786. |  | 38 | 38 boxes with 6 units each. | Nationwide Distribution including AR, AZ, CA, CO, FL, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, NE, OH, OK, OR, TN, TX, WA, and WI. | Open |  |  | 717 | Not_Computer
Z-1365-2013 | 65158 | medical device data system | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MEDITECH Bi-Directional interactive Communication Analyzer Interface; Integrates with MEDITECH's Laboratory Information System (LIS) with Other Vendor systems to manage the transfer of patient results and demographic information. | MEDITECH BiDirectional LAB Analyzer Interface | 2 | May 22, 2013 | 2013 | Medical Information Technology, Inc. | Incorrect transmission of laboratory results. | DESIGN: Device Design | Beginning on May 7, 2013, MEDITECH distributed a notification to MAGIC and Client Server customers who have the MEDITECH Bi-Directional Interactive Communication Analyzer Interface in use with Vitros 3600, 4600, or 5600 Analyzer in TEST and/or LIVE. This notification was made via remailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.This notification informs customers of the potential patient safety issue that could result in reporting inaccurate qualitative results. Customers will receive a corrective change, in which the previous test's results will no longer be associated with the test not performed. Until customers receive the corrective change, they are instructed to monitor all No Result values from their analyzer and confirm that the result in MEDITECH matches. For further questions please call 781-821-3000, ext. 5742. |  | 56 | 56 devices | Worldwide Distribution - USA, including Puerto Rico and the country of Canada. | Open |  |  | 56 | Not_Computer
Z-1366-2013 | 61315 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Hip Solutions ALTRX Polyethylene Acetabular Liner, Ref 1221-36-452Product Usage:An acetabular liner is used in combination with other implants for total hip arthroplasty is intended to provide increased patient mobility and reduce pain by replacing the damaged hip joint articulation in patients where there is evidence of sufficient sound bone to seat and support the components. | Pinnacle | 2 | May 22, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of one lot of the AltrX" liner because the product was incorrectly labeled on the box. The box was labeled for a 36IDX52OD +4 Neutral Liner, but the product enclosed was a 36IDX56OD Neutral Liner. Three ( 3 ) complaints were received in January 2012. | PRODUCTION CONTROLS: Packaging Process Control | DePuy Orthopaedics sent an URGENT INFORMATION- MEDICAL DEVICE RECALL NOTICE dated March 7, 2012 via email to all affected distributors and customers. The letter identified the affected product, problem and actions to be taken. Distributors and customers are instructed to cease further distribution or use or of recalled product and to return them to DePuy. The sales representatives are expected to aid customer in returning of recalled product. Distributors and sales representatives are to return the recalled product through the normal DePuy Returns process. Reconciliation form should be completed and returned to your DePuy Sales Representative or faxed to 574-372-7567. For questions about recall information contact the Manager of Customer Quality 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA Nationwide including the state of VA and the country of Ireland. | Terminated | May 24, 2013 | 2 | 20 | Not_Computer
Z-1217-2013 | 65036 | unit, x-ray, extraoral with timer | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | MyRay RXDC dental unit x-ray unit | MY Ray RXDC xray unit | 2 | May 23, 2013 | 2013 | CEFLA DENTAL GROUP | The RX DC dental x-ray unit was found to be noncompliant with 21CFR 1010,2, 1020.30(e) & 1020.30(h). Specifically, 1. The tube housing did not contain a Certification label.2. The collimator did not contain an identification label.3. The user's manual did not contain leakage technique factors, tube housing cooling curves and rating charts, and technique factors which lead to operation at ma | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | The firm, CEFLA Dental Group will update the user's manual to include the information required by 21 CFR 1020.30(h). Future production will ensure that system labels required by 21 CFR 1010.2, and 1020.30(e) are properly applied including the 53 units manufactured awaiting shipment. The updated user's manual will be made available online for free download. Alternatively, system purchasers, subsequent transferees, dealers and distributors may elect to receive a hard copy of the updated user's manual at no charge. All corrections will be made at no cost to the purchaser, subsequent transferee, dealer, or distributor.The above actions will be completed within 4 weeks. For further questions please call 1-800 323-2690. | Contact the recalling firm for information | 173 | 173 shipped to US | GA, TX, IL, CA, FL, NY, MA, AL, NJ, SC, WA, LA, MD, WY, CT, NM, WI, CO, NC, VA, PA, MS, and MICanada | Open |  |  | 173 | Not_Computer
Z-1370-2013 | 64941 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon Femoral Distal Augment- Left; 1 single unit per packageHowmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430 USA Stryker France, ZAC Satolas Green Pusignan, Av de Satoles Green 69881Meyzieu, Cedex, FranceThis device is used when there is a deficiency in the bone present during femoral preparation these devices are used to "fill the gap" between the implant and the bone | Triathlon Femoral Distal Augment Left | 2 | May 23, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that a label on the external packaging (box) of a Triathalon Femoral Distal Augment 10mm- Size 5 Left depicted a glyph of an acetabular shell rather than a femoral distal augment. In addition the label was missing a reference to the "10mm" augment thickness which per the label print is required to appear below the size reference. | DESIGN: Software Design (Manufacturing Process) | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 25, 2013 and Product Accountability Forms via Fed Ex to its Consignees/ customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the attached Notification Acknowledgment Form within 5 days via fax to 855-251-3635; contact the hospitals in your territory that have the affected product to arrange return of product; and return affected product to the attention of Regulatory Compliance Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, attach flourescent orange PRODUCT REMEDIATION sticker to your return indicating the PR #RA2013-011 and mark outer box with the words "Product Recall".If you have any questions, feel free to contact Manager, Divisional Regulatory Compliance at (201) 972-2100. | Contact the recalling firm for information | 11 | 11 units | Nationwide distribution: USA including states of: HI, IN, MS, NC, NJ, OK, TX, and WI. | Open |  |  | 11 | Not_Computer
Z-1367-2013 | 64718 | cytocentrifuge | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | SLIDERPREP PLUS SLIDE PREPARATION CENTRIFUGE.Produce a monolayer of cells onto a glass slide from any fluid suspension. | SLIDERPREP PLUS SLIDE PREPARATION CENTRIFUGE | 2 | May 23, 2013 | 2013 | Separation Technology, Inc. | Separation Technology, Inc. recalled SlidePrep Plus Cytology Centrifuges manufactured prior to January 1, 2013 for having the potential to start up unexpectedly when encountering an electrostatic discharge. | DESIGN: Device Design | On March 13, 2013, Separation Technology, Inc. sent their customers letters instructing them to complete and return the attached Acknowledgement & Receipt form via fax to (407) 788-3677 and requesting the return of the SlidePrep Plus Cytology Centrifuges to STI for correction. For any questions regarding this recall, call STI Customer Service at (800) 777-6668. |  | 102 | 102 | Worldwide Distribution-USA (nationwide) including the states of PA, NC, IL, KS, and FL, and the countries Canada and Turkey. | Terminated | December 24, 2013 | 215 | 102 | Not_Computer
Z-1378-2013 | 64588 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Femoral component, CR, cemented, #1, left2103-1310Femoral component, CR, cemented, #2, left2103-1320Femoral component, CR, cemented, #3, left2103-1330Femoral component, CR, cemented, #4, left2103-1340Femoral component, CR, cemented, #5, left2103-1350Femoral component, CR, cemented, #6, left2103-1360Femoral component, CR, cemented, #1, right2103-1410Femoral component, CR, cemented, #2, right2103-1420Femoral component, CR, cemented, #3, right2103-1430Femoral component, CR, cemented, #4, right2103-1440Femoral component, CR, cemented, #5, right2103-1450Femoral component, CR, cemented, #6, right2103-1460Femoral component, PS, #1, left2103-3110Femoral component, PS, #2, left2103-3120Femoral component, PS, #3, left2103-3130Femoral component, PS, #4, left2103-3140Femoral component, PS, #5, left2103-3150Femoral component, PS, #6, left2103-3160Femoral component,PS, #1, right2103-3210Femoral component,PS, #2, right2103-3220Femoral component,PS, #3, right2103-3230Femoral component,PS, #4, right2103-3240Femoral component,PS, #5, right2103-3250Femoral component,PS, #6, right2103-3260Variety of hip and knee implants and instruments, multiple uses. | Femoral component | 2 | May 24, 2013 | 2013 | Orthopedic Alliance LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Orthopedic Alliance because the directions for use and surgical techniques were not being provided to every physicianand for every surgery. | DESIGN: Labeling Design | A recall letter dated 3/8/13 was sent to all customers who purchased loaner bins and sterilization trays for implants and instruments from Orthopedic Alliance. The letter informed the customers of the problem identifiedand the actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 136 | 136 units | Nationwide Distribution including the states of CA, MD, TX, NV, and WI. | Open |  |  | 821 | Not_Computer
Z-1377-2013 | 64886 | drills, burrs, trephines & accessories (simple, powered) | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ***REF S-3D*** 3mm Diamond Ball.Used for cutting and shaping bone. | 3mm Diamond Ball | 2 | May 24, 2013 | 2013 | The Anspach Effort, Inc. | Due to three batches of product may have diamond grit size larger than intended applied to the head of the cutting burr. | PRODUCTION CONTROLS: Labeling Mix-Ups | Anspach sent an URGENT: Medical Device Correction letter dated March 4, 2011 to its customers identifying the affected product and actions to be taken by the customers. Customers questions or concerns should be directed to Anspach Customer Support at (800) 327-6887. |  | 1827 | 1,827 burrs | Worldwide Distribution-USA (nationwide) including DC and the states of OH, TX, IL, CA, GA, KY, FL, WI, NC, VA, WI, MN, MD, KS, OK, PA, OR ,HI, AZ, WA, TN, AR, IA, MI, CO, LA, NJ, IN, WV, MO, AL, NY, VT, and SC and the countries of France, Canada, Great Britain, Mexico, Australia, Sweden and Netherlands. | Terminated | December 17, 2013 | 207 | 1827 | Not_Computer
Z-1376-2013 | 65042 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NEXGEN Complete Knee Solution Monoblock Tibial Drill with stop 10.7 mm. Used to prepare cylindrical holes in the native tibia. | Trabecular Metal" Tibia Stop Drill | 2 | May 24, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Tibia Stop Drill due to the potential of the outside diameter being oversized. As a result, there is a potential for the drill to not pass through the applicable drill guide. | PRODUCTION CONTROLS: Process Control | Zimmer sent an "URGENT MEDICAL DEVICE RECALL LETTER" dated April 19, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 69 | 69 units | Nationwide Distribution. | Open |  |  | 69 | Not_Computer
Z-1398-2013 | 63204 | injector, vertebroplasty (does not contain cement) | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthobiologics Aliquot Delivery System Manual Surgical Instrument - Aliquot Beveled Needle, 6", Part Number 2110-0524Product Usage:Facilitate placement of bone cement in weakened or diminished bone | Stryker Orthobiologics Aliquot Beveled Needle, 6", Part Number 21100524 | 2 | May 24, 2013 | 2013 | Orthovita, Inc., dBA Stryker Orthobiologics. | There have been a few reported cases of the catheter breaking when removing the Aliquot catheter from the needle during vertebroplasty procedures. This has been attributed to bending of the catheter during the retraction from the needle. It was determined to be use errors. The correction is limited to a revision of the Directions for Use. | DESIGN: Labeling Design | Stryker sent an Urgent Medical Device Correction letter dated November 10, 2012 to all affected customers. The letter identified the affected product, problem issue, potential hazards and actions to be taken. Customers were instructed to immediately review this notification and the attached IFU. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to 1-484-323-8803. For questions contact Stryker Orthobiologics. |  | 12 | 12 | USA Nationwide Distribution in the states of AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, ND, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Open |  |  | 1413 | Not_Computer
Z-1400-2013 | 63367 | acid, lactic, enzymatic method | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX System(s) and UniCel DxC 600/800 System(s) Lactate Reagent, Part Number: A95550Product Usage: Lactate (LACT) reagent, when used in conjunction with Synchron LX System(s), UniCel DxC 600/800 System(s) and Synchron Systems Multi Calibrator, is intended for the quantitative determination of lactate concentration in human plasma and cerebrospinal fluid (CSF). | Synchron LX system and UniCel DxC 600/800 System(s) Lactate Reagent | 2 | May 24, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter initiated the recall for the Lactate Assay (LACT) Lot M206209 as they received a customer complaint of error message (RXN RATE HI), i.e. no numerical values reported for samples containing Lactate concentration greater than 6.5 mmol/L. |  | Beckman Coulter sent an Urgent Product Correction letter dated April 17, 2013 to all affected customers. The letter identified the affected product, problem actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center (Call Center) on the Beckman Website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. |  | 8774 | 8,774 units | Worldwide Distribution - USA Nationwide and the countries of Canada, Australia, Belgium, Czech Republic, France, French Polynesia, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Jordon, Kuwait, Lebanon, Malawi, Morocco, Namibia, Netherlands, oman, Panama, Paraguay, Qatar, Reunion, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom | Open |  |  | 8774 | Not_Computer
Z-1408-2013 | 65150 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical PerfOrma Anglographic Catheter 5F MultipackProduct Usage:Angiographic catheters are designed to be used for delivering radiopaque media to selected sites in the vascular system in conjunction with routine diagnostic procedures. | Performa Cardiac Multipack | 2 | May 25, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems is recalling various lots of Performa Cardiac Multipacks containing a JR4.0 catheter due to a slight variation in tip shape. | PRODUCTION CONTROLS: Process Control | Merit sent an Urgent Product Recall Notice letter dated May 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to identify, quarantine, discontinue using products, and contact their sales representative to arrange product return and replacement or credit. For questions call 1-801-208-4468. |  | 1465 | 1,465 | USA Nationwide Distribution. | Terminated | July 22, 2013 | 58 | 1465 | Not_Computer
Z-1403-2013 | 65067 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS PANELS, Glucose Test Strips for use with CardioChek Brand AnalyzersProduct Usage: used by healthcare professionals and individuals with diabetes to measure glucose in whole blood | PTS Panels(R) Glucose Test Strips | 2 | May 25, 2013 | 2013 | Polymer Technology Systems, Inc. | Internal investigation demonstrated under-recovery for glucose when tested against a reference method. This could result in a low to moderate health risk. |  | Polymer Technology Systems sent a Field Correction Action letters dated April 4, 2013 to distributors and customers. The letter identified the affected product, problem and actions to be taken. The letter directed customers to discontinue use and dispose of remaining inventory of recalled product. Customers were asked to acknowledge and confirm recall directions were follwed by fax to 1-317-870-5608. |  | 7048 | 7048 total, 3499 in USA | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: UK, CHL, ENG, BEL, ITA, TKY, BEL, LIT, LTV, EGY, SPG, HOK, SWE, MEX, UAE, COR, ISR, MAL, AUS, POR, POL, SLN and STQ. | Open |  |  | 7048 | Not_Computer
Z-1405-2013 | 65119 | cannula, catheter | DQR | Catheter cannula. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Micro Introducer Kits under the following labeling: 1) AngioDynamics 4F Standard Micro-Introducer Kit, Sterile, Catalog Number: 06597005, 2) AngioDynamics 5F Standard Micro-Introducer Kit, Sterile, Catalog number: 06597012; 3) AngioDynamics 4F Standard Micro-Introducer Kit, Sterile, Catalog Number: 06597013; 4) AngioDynamics 5F Standard Micro- Introducer Kit, Sterile, Catalog Number: 06597018; and 5) AngioDynamics 4F Standard Micro-Introducer Kit, Sterile, Catalog Number: 06597019. Product kits are individually wrapped in plastic, 10 units included per box.Product Usage:This device is intended to be used to measure the refractive power and the radius of corneal curvature of the human eye. | AngioDynamics Micro Introducer Kits | 2 | May 25, 2013 | 2013 | Angiodynamics, Inc. | AngioDynamics Inc., is recalling Micro Introducer Kits and VenaCure EVLT NeverTouch Procedure Kits, because the potential exists that the micro Introducer (sheath/dilator) packaged within the kit could be the incorrect French size. | PRODUCTION CONTROLS: Labeling Mix-Ups | AngioDynamics sent an Urgent Medical Device Recall letter dated April 29, 2013, to all customers by Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to lmmediately remove the affected product from inventory and call AngioDynamics, Inc. Customer Service at 1-800-772-6446 to obtain a replacement or credit for your returned product and to obtain UPS Account Number to return the package via second day delivery. Customers are instructed to segregate and return all affected devices to: AngioDynamics 603 Queensbury Avenue Queensbury, NY 12804 Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinator. Write the RMA number on the shipping box (Provied on the Recall Verification Tracking Form). Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on the notification form and the Reply Verification Form. Send the Reply Verification form to AngioDynamics either via Fax or Emai. Fax Reply Verification Tracking Form to: Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinatory Fax number 1-518-798-1360 or Email Reply Verification Tracking Form to: rdenino@angiodynamics.com. |  | 1973 | Domestic(1,760 units): 1)260 units/26 boxes, 2)310 units/31 boxes,3)260 units/26 boxes,4)880 units/88 boxes,5)60 units/6 boxes,Foreign:10 units/1 box | USA Nationwide Distribution including the states of: IL, GA, CA, NY OH, AZ, TX, NE, FL, PA, NC, WI, AL, IN, CO, MO, MT, NE, MA, VA, AK, IN, MI, NJ, MD, TN, CT, WA, MN | Open |  |  | 1770 | Not_Computer
Z-1402-2013 | 65060 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 3.0Product Usage:The Agility multileaf collimator is indicated for use when additional flexibility is required in conforming the radiation beam to the anatomy to be exposed. | Multileaf Collimator | 2 | May 25, 2013 | 2013 | Elekta, Inc. | Unexpected movement during beam delivery may not be detected if there are incorrect static tolerances for Linac Geometric Axes loaded in Integrity R3.0. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Elekta sent an Important Field Safety Notice (IFSN) letter dated April 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the FCO Action Notification Report to local Elekta Office or Representative. |  | 31 | 31 | USA Nationwide including the states of CA, FL, IL, MD, MA, NJ, NY, NC, ND, OH, OR, PA, TN, TX, VA, WA, DC and Puerto Rico | Open |  |  | 31 | Not_Computer
Z-1401-2013 | 65048 | visual, pregnancy hcg, prescription use | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere" hCG Combo Cassette (20/10 mIU/mL) Kit.Product Usage: The AlereTM hCG Combo (20/10 mIU/mL) test is a rapid chromatographic immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in serum or urine, as an aid in the early detection of pregnancy. It is for healthcare professionals only. | Alere hCG Combo Cassette (20/10 mIU/mL) KIt | 2 | May 25, 2013 | 2013 | Alere San Diego, Inc. | The firm initiated recall of Alere" hCG Combo Cassette (20/10 mIU/mL) kit PN 92215, lots hCG1110133 and hCG1110135, because an investigation has indicated a portion of these Alere" hCG Combo Cassette (20/10 mIU/mL) PN 92215 kit lots were incorrectly packaged with the Alere" hCG Cassette (20 mIU/mL) PN 92210 Product Insert. The Alere" hCG Cassette (20 mIU/mL) Product Insert contains information for | PRODUCTION CONTROLS: Packaging | Alere initiated this recall by sending recall notification letters, titled "URGENT MEDICAL DEVICE RECALL", dated April 16, 2013 via via fax, email or direct mail. The letter informed customers of the product recalled, reason for recall, customer required action, and contact information, "Alere Technical Services at 866-216-0094 or by e-mail at lateral.flow.support@alere.com.". A verification form was attached to the notification letter. |  | 3975 | 3975 kits | United States Nationwide Distribution | Terminated | October 18, 2013 | 146 | 3975 | Not_Computer
Z-1409-2013 | 64586 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | 45 mm Screw, Item #BG7045Product Usage:Spinal Solutions distributes a variety of spinal fixation devices in loaner bins or on consignment. The company has decided to provide directions for use and surgical techniques electronically on the companys website. The company has created a label intended to inform the user of the existence of the websites and the information it contains. Indications for use: multiple. | PEDICLE SCREW SYSTEM | 2 | May 27, 2013 | 2013 | Spinal Solutions, LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Spinal Solutions because the directions for use and surgical techniques were not being provided to every physician and for every surgery. | DESIGN: Labeling Design | Spinal Solutions sent a recall letter dated March 8, 2013 to all customers who purchased loaner bins and sterilization trays for implants and instruments from Spinal Solutions. The letter identified the affected product, problem identified and actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 24 | 24 units | USA Nationwide Distribution in the states of MS: TX, NV, CA, MD, and WI. | Open |  |  | 503 | Not_Computer
Z-1421-2013 | 65165 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Leucadia Autolok" Polyaxial Pedicle Screw System.Intended to help provide correction, immobilization and stabilization of spinal segments. | Leucadia Autolok" Polyaxial Pedicle Screw System | 2 | May 28, 2013 | 2013 | Alphatec Spine, Inc. | Alphatec Spine initiated this medical device recall for the following reason: due to tolerancing, the flange of the Phygen Leucadia" Autolok set screw may interfere with the top of the screw head prior to fully locking the rod. If this condition occurs, the polyaxialscrew head and rod may fail to provide a rigid construct under loading. To date there have beenno reports of patient injury, adve | DESIGN: Device Design | Alphatec Spine initiated recall by contacting customers via telephone and/or email on 05/06/2013. An Urgent Medical Device Recall Notification letter along with a response form were sent via FedEx to all affected accounts. The letter identified the product, problem, and actions to be taken by the customers. | Contact the recalling firm for information | 1814 | 1,814 | Nationwide Distribution. | Open |  |  | 1814 | Not_Computer
Z-1372-2013 | 65070 | nebulizer, medicinal, non-ventilatory (atomizer) | CCQ | Medicinal nonventilatory nebulizer (atomizer). | Anesthesiology | Anesthesiology | 510(K) Exempt | Asthmanefrin Starter Kit contains ten (10) Asthmanefrin vials and one (1) EZ Breathe Atomizer.For the delivery of liquid medications for respiratory usage. | Asthmanefrin Starter Kit | 1 | May 29, 2013 | 2013 | Nephron Pharmaceuticals Corp. | Possible dislodgement of the "Plate A", if this instance occurs, it presents a potential choking hazard. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Nephron Pharmaceuticals Corporation, notified its consignees on April 30, 2013, and May 3, 2013, of the problem via a recall notice and a Press Release, respectively. The firm directed consignees to discontinue use of any EZ Breathe atomizer units contained in the Asthmanefrin Starter Kits with the affected lot numbers and also those that were sold individually. Customers were directed to return all affected product. | Contact the recalling firm for information | 241538 | 241,538 kits | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AZ, CA, CO, CT, FL, IL, IN, IO, GA, KS, KY, LA, MD, MN, MS, MO, NK, NV, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN , TX, VT, VA, WA, and WV, and the countries of Guatemala, Poland, Panama, and United Kingdom. | Open |  |  | 242892 | Not_Computer
Z-1423-2013 | 65089 | discs, strips and reagents, microorganism differentiation | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL DrySlide Oxidase Kit, catalog number 231746Product Usage:Presumptive diagnostic aid for gram-negative bacteria. | BBL DrySlide Oxidase kit | 2 | May 29, 2013 | 2013 | Becton Dickinson & Co. | In-vitro diagnostic agents were placed into incorrect packaging for distribution to microbiological laboratories. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | BD Diagnostic Systems sent an Urgent Product Recall" letter dated April 2013 to all affected consignees and end users. The letter identified the affected product, problem and actions to be taken. The notice instructed consignees and end users to discontinue use of the products, discard remaining products, and return a completed response form. Consignees were requested to provide contact information for end users who received recalled product in distribution, to allow for recall notice. The letter included instructions to report injuries and/or illnesses to Medwatch, and address recall inquiries to BD Technical Services at 1-800-638-8663. |  | 610 | 610 total test kits (PYR and Oxidase) | Worldwide Distribution - USA Nationwide and tthe countries of Australia, Canada, Europe, Hong Kong, Mexico, New Zealand, Singapore, and Taiwan. | Terminated | September 28, 2013 | 122 | 610 | Not_Computer
Z-1425-2013 | 65073 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Tip Cover Accessory, 8 MM, MCS, box of 10Product Usage:The disposable Tip cover accessory is used in conjunction with the 8mm Monopolar Curved Scissors Instruments to cut and coagulate tissue. The tip cover is intended to provide electrical insulation over a section of the Monopolar Curved Scissors instrument so that FR energy is only available at the scissors blades. | EndoWrist MCS Tip Cover Accessory | 2 | May 29, 2013 | 2013 | Intuitive Surgical, Inc. | Clarification of existing labeling and related cautions for the tip cover accessory of the da Vinci Surgical System. | DESIGN: Device Design | Intuitive Surgical sent an Important Product Notification letters dated October 10, 2011 to all affected customers. A second letter was sent to those that had not responded on December 4, 2011. The letters identified the affected product, problem, and actions to be taken, The letter included recommendations and suggestions for proper use and prevention of damage to the Tip cover during surgery. Also attachments include recommendations and further information for the following areas: avoid instrument collisions; carefully install tip cover; straighten wrist prior to removal; inspect cannula prior to use and do not exceed maximum monopolar cautery settings.Attached is a list of validated ESU settings and suggestions and instruction on being aware of critical anatomy in contact with the instrument and to change the tip cover in cases with extended cautery use.Questions may be referred to Customer service at 1-800-876-1310 in the US and at + 41 21 821 2020 or +800-0821-2010. |  | 10 | 50,000 boxes of 10 tip covers | Worldwide distribution - USA Nationwide and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom. | Terminated | May 30, 2013 | 1 | 50000 | Not_Computer
Z-1428-2013 | 65083 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Flush Port (used in almost all da Vinci instruments)Assists in cleaning the instrument. | Flush Port | 2 | May 29, 2013 | 2013 | Intuitive Surgical, Inc. | Clarification of existing labeling information and operating procedures. | MISBRANDING: Labeling False and Misleading | Intuitive Surgical sent an "IMPORTANT PRODUCT NOTIFICATION" letter dated October 17, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 1-800-876-1310 for questions regarding this notice. |  |  |  | Worldwide Distribution- USA Nationwide and countries of: Argentina, Australia,Canada, Chile, China, Czech Republic, Egypt, India, Israel, Italy,Korea-Republic of, Poland, Qatar, Romania, Saudi Arabia, Singapore, Spain, Thailand, Turkey, and Venezuela. | Terminated | May 30, 2013 | 1 |  | Not_Computer
Z-1430-2013 | 65118 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU2700 Clinical Chemistry AnalyzerAU5400 Clinical Chemistry AnalyzerPerform automated analysis of serum, urine, and other body fluids, including whole blood. | AU2700 and AU5400 Clinical Chemistry Analyzers | 2 | May 29, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling AU2700/AU2700 Plus & AU5400 Clinical Chemistry Analyzers becuase it was determined that the tubing/joint connections associated with these MODs may leak concentrated detergent. Beckman Coulter has received complaints that the joint connecting the tubing appears to disintegrate over time, causing a leak. | DESIGN: Component Design/Selection | Beckman Coulter sent an Urgent Product Correction letter on April 26, 2013 along with a response form. The letter notified customers of the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-854-3633 for questions regarding this notice. |  | 108 | 108 units | Worldwide Distribution-USA Nationwide and countries of: AU, CA, DE, ES, GB, IT, PT, and JP. | Open |  |  | 108 | Not_Computer
Z-1437-2013 | 49033 | ventilator, non-continuous (respirator) | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Fisher & Paykel Healthcare ICON CPAP (continuous positive airway pressure).Models: ICON Auto: ICONAAN, ICONAAN-HT, ICONAAJ.ICON Novo: ICONNAN, ICONNAN-HT, ICONNAJ.ICON Premo: ICONPBN, ICONPBN-HT, ICONPBN-HTC.The ICON CPAP unit is for use on adult patients for the treatment of Obstructive Sleep Apnea (OSA). The device is for use in the home or sleep laboratory. The device is not intended to be used as a life supporting device. | Fisher & Paykel Healthcare ICON CPAP | 2 | May 30, 2013 | 2013 | Fisher & Paykel Healthcare, Ltd. | Fisher and Paykel Healthcare is recalling the ICON CPAP units because there is a potential power issue on certain ICON CPAP. These ICON CPAP units could potentially become inoperable, which would result in a lack of treatment to the patient's obstructive sleep apnea symptoms until a replacement unit is obtained. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Fisher & Paykel Healthcare Limited, sent a "Medical Device Recall" letter dated April 22, 2013 to all their customers who purchased the Fisher & Paykel Healthcare ICON CPAP Replacement. The recall letter described the product, notified the customers of the problem identified and the actions to be taken. Customers are instructed to locate and verify if any of the affected products are in stock; check the part and lot numbers of your device; complete and return the enclosed Product Recall Response form, even if you have not received or distributed any affected products via Email: FA-2013-001@fphcare.com; Fax: (949) 453-4001; or Post: Fisher & Paykel Healthcare,15365 Barranca Parkway, Irvine, CA 92618; transfer this notice immediately to all those who need to be aware and if affected product has been distributed to any other organizations or distributors, notify them regarding the recall within 10 business days upon receipt of this letter; prepare any products to be returned to Fisher & Paykel Healthcare and contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at 1-800-792-3912 ext. 1426 to arrange for a return shipping label to be sent to you. If you have any questions related to this recall, contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at (800) 792-3912 ext. 1426. |  | 1407 | 1,407 units | Worldwide distribution: USA (nationwide) and countries of: Canada, Japan, and Taiwan. | Terminated | October 18, 2013 | 141 | 1407 | Not_Computer
Z-1438-2013 | 65120 | pacemaker, cardiac, external transcutaneous (non-invasive) | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | adult defibrillation, cardioversion, pacing, monitoring Disposable Electrodes, Model P-211-M1 and P-214-M1.Product Usage:Used by trained professionals in hospitals, doctors offices, and Emergency Medical Services for adult low-energy defibrillation, transcutaneous pacing, cardioversion and monitoring. | Medtronic/Physio Control Transparent Defibrillation Electrodes | 2 | May 30, 2013 | 2013 | TZ Medical Inc. | TZ Medical Inc is recalling its P-211-M1 and P-214-M1 Medtronic / Physio Control Compatible transparent defibrillation electrodes because it may contain a defective white wire on the connector assembly.If a failure were to occur, the end user would be able to identify the failure by inspecting the white wire 1/4 inch from the Gray Medtronic / Physio adapter. The consignees will notice discolor | PRODUCTION CONTROLS: Process Control | TZ Medical Inc. started notifying their customers by telephone on April 30, 2013. The firm followed up with an Urgent: Medical Device Recall notification letter dated May 14, 2013, to their consignees. TZ Medical Inc., Representatives have will contact all affected customers and arrange for the replacement of the effected product in stock with inspected product. For any questions about this recall, please fax to 503-639-0239. |  | 100 | US: 10,410 pieces; Outside US: 100 pieces | Worldwide distribution - USA (nationwide) and Chile. | Terminated | September 05, 2013 | 98 | 100 | Not_Computer
Z-1435-2013 | 65085 | electrode, pacemaker, temporary | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Vermed, Inc Radiotransparent Defibrillation ElectrodesMonophasic or biphasic compatible defibrillation/cardioversion/pacing/ECG Monitoring electrodes.Used in connection with a defibrillator to deliver an electrical current (shock) to the patient's heart during cardiac emergencies. | Vermed, Inc Radiotransparent Defibrillation Electrodes | 2 | May 30, 2013 | 2013 | Heart Sync, Inc | On 4/18/13 HEARTSYNC became aware of a recall from their wire/connector sub-assembly supplier which is used in the C100-Physio Electrodes. There were reports of equipment damage due to misalignment from tooling installed in August 2012. The result of this damage could render the device incapable of delivering the appropriate energy or shock to the patient. | DESIGN: Process Design | HeartSync sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 26, 2013, to all consignees of Heart Sync for C100-Physio Radiotransparent Defrillation Electrodes. On April 29, 2013, an URGENT MEDICAL DEVICE RECALL notification was sent to Vermed Inc for the D304 Defrillation Electrodes . The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the affected product to the firm for full credit. Customers were asked to complete the form and send back via fax to 734-213-5640 or call 734-213-5530.For questions regarding this recall call 734-213-5530.. |  | 17180 | 17180 units for all part numbers | Worldwide Distribution - USA including AL, GA, MI, MN, NC, ND, OH, PA, SD, VA, VT, OH, TX, MA, and Internationally to Korea. | Open |  |  | 17180 | Not_Computer
Z-1440-2013 | 65109 | microscope, surgical | EPT | Surgical microscope and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Zeiss brand OPMI Pentero C, Surgical microscope and accessories, Model: OPMI Pentero C, Product Numbers: 302583-9001-000 (PENTERO C), and 305961-0000-00 (Ceiling Mount for PENTERO C),Product Usage:A surgical microscope and accessories is an AC-powered device intended for use during surgery to provide a magnified view of the surgical field. | OPMI Pentero C | 2 | May 31, 2013 | 2013 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec is conducting a field corrective action for its OPMI Pentero C surgical microscope system. During a routine inspection, one microscope installation was found to have 3 (three) broken screws out of 6 (six) in the lower section of the ceiling mount. |  | Carl Zeiss Meditec, Inc sent a Customer Notification letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that firm's sales representative will be scheduling an on-site visit, informing them of the affected product and providing instructions on the recall. For question call 925-580-5371. | Contact the recalling firm for information | 6 | 6 units | US Nationwide Distribution. | Terminated | June 26, 2013 | 26 | 6 | Not_Computer
Z-1442-2013 | 65216 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | 8 mm Monopolar Curved Scissors (a.k.a. Hot Shears)Manufactured and Distributed by:Intuitive SurgicalSunny vale, CA The EndoWrist Monopolar Curved Scissors Instrument is a multiple-use endoscopic instrument utilizing a single use, tip cover accessory, to be used in conjunction with the Intuitive Surgical Endoscopic Instrument Control System. | 8 mm Monopolar Curved Scissors | 2 | May 31, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a potential for some units of the Monopolar curved Scissors version 09 and 10 instruments to develop very small cracks near the distal end. This may create a pathway for electrosurgical energy to leak to tissue and potentially cause thermal injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Recall letter dated May16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that they did not need to quarantine or return the affected device at that time. a Intuitive Surgical representative would contact them to arrange for replacement of the affected product once they became available. In the interim customers were instructed to follow the precautions and warnings outlined in Attachment A in the letter as well as the Instructions for Use. Customers were also asked to ensure all affected personnel were fully informed of the contents of the Recall Notification. The letter should be forwarded to other surgeons and/or managers within their facility. Customers were also asked to complete and return the attached Acknowledgement Form. For further questions customers were instructed to contact Customer Service:North and South America: 800-876-1310 Option 3Japan: 0120-56-5635 or 003-5575-1362Korea: 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 or +41 21 821 2020For questions regarding this recall call 408-523-2244. |  | 132441 | ~132,441 instruments | Worldwide distribution - USA (nationwide) Foreign Distribution to Argentina, Austria, Australia, Belgium, Brazil, Bulgaria, Canada, China, Chile, Czech Republic, Denmark, Finland, France, Germany, Cyprus- Greece, Egypt, Guadeloupe, Israel, Ireland, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Portugal, Poland, Qatar, Russia, Romania, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Taiwan, Turkey, United Kingdom, Saudi Arabia, Singapore, South Korea, and Venezuela. | Open |  |  | 132441 | Not_Computer
Z-1439-2013 | 65255 | instrument, ent manual surgical | LRC | Ear, nose, and throat manual surgical instrument. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | GYRUS ACMI, MYRINGOTOMY KNIFE DISPOSABLE, SICKLE BLADE FLAT HANDLE, REF 130715, Rx only, STERILE EO, GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104Disposable myringotomy knife intended to perforate the tympanic membrane to allow the placement of a bent tube or to directly provide ventilation or drainage of the middle ear. | MYRINGOTOMY KNIFE DISPOSABLE, SICKLE BLADE FLAT HANDLE | 2 | May 31, 2013 | 2013 | Gyrus Acmi, Incorporated | lack of sterility assurance |  | Cyrus ACMI notified their direct accounts of the recall by Priority Mail with delivery confirmation on May 24, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and immediately cease any further use of any affected product they may have, remove it from inventory, and quarantine it until they ship it back to the firm. For questions customers were instructed to contact their Olympus customer service representative at 1-800-848-9024 to obtain a Returned Goods Authorization so that they may return the product with no charge to them. Customers were asked to note on the enclosed questionnaire that they have received this information and fax it to 484-896-7128.For questions regarding this recall call 508-804-2739. |  | 100 | 100 boxes of 6 units | Worldwide Distribution - USA (nationwide) and countries of: Canada and United Kingdom. | Open |  |  | 100 | Not_Computer
Z-1369-2013 | 64803 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Corporation - Mobile DaRt EvolutionThis device is a mobile X-ray system in order to take general radiography of patients in the hospital who cannot move and/or of out patients in emergency . | Mobile DaRt Evolution | 2 | June 03, 2013 | 2013 | Shimadzu Medical Systems | The recall was initiated because Shimadzu Corporation has identified a potential issue in certain manufacturing lots of the Shimadzu MobileDaRt Evolution in combination with CXDI-70C, CXDI-55C, and CXDI-55G. As a result of Shimadzu's investigation, there is a potential risk for screws becoming loose inside the Shimadzu Mobile X-ray systems. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | Shimadzu sent a recall letter dated March 18, 2013, to all customers who purchased the Shimadzu MobileDart Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to follow the instructions of the recall notice step by step. Customers with questions are instructed to contact their local Shimadzu representatives or contact Shimadzu National Technical Support at (800) 228-1429.A Field Safety Notice dated March 18, 2013, was sent to Service Managers in regards to the Shimadzu Mobile DaRt Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The Safety Notice informs the Service managers of the problems identified and the actions to be taken. |  | 200 | 200 units | US Nationwide Distribution | Terminated | February 13, 2014 | 255 | 200 | Not_Computer
Z-1443-2013 | 65164 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Canister: CO2 canister with GE part numbers 1407-3200-000, 1407-7004-000, or M1084850. Intended to provide general inhalation anesthesia and ventilatory support. | Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems | 2 | June 03, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the reusable CO2 absorbent canister accessory used with Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems. Reusable CO2 absorbent canisters accessories may not seal properly. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter April 30, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Fax back form was included with the letter requesting the consignees to complete and return. |  | 800 | 800 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, CT, ID, MK, MI, NJ, NY, OH, TN, TX, and VA, and the countries of INDIA, AUSTRALIA, BAHRAIN, BANGLADESH, BOLIVIA, BOSNIA & HERZEGOWINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDONESIA, IRAQ,ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA,LEBANON, MALAYSIA, MEXICO, NETHERLAND,NEW ZEALAND, OMAN, PARAGUAY, PERU., PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, THAILAND, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, and VIET NAM. | Open |  |  | 800 | Not_Computer
Z-1446-2013 | 64968 | cassettes, tissue | IDZ | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | LPC Fine Mesh Cassette for the Leica LPC Printer. Part No: 14060546848; 14060546846; 14060546845; 14060546933; 14060546932; 14060546852; 14060546850; 14060546847. The intended use of this device is for storage of specimens/tissues during processing. | LPC Fine Mesh Cassettes | 2 | June 03, 2013 | 2013 | Leica Microsystems, Inc. | An LPC Fine Mesh Cassette failed to stay closed during processing operations. | DESIGN: Device Design | A Medical Device Correction letter dated April 17, 2013 was sent to customers via USPS 1st class mail. The product issue was described and recommended actions were provided. Customers were requested to acknowledge receipt of the letter by fax-back. Customer questions were directed to (847) 405-5413. |  | 20160 | 20,160 units | Worldwide distribution: US (nationwide) to states of AZ and KY; and the country of New Zealand. | Open |  |  | 20160 | Not_Computer
Z-1444-2013 | 64826 | catheter, conduction, anesthetic | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | I-Flow ON-Q Pain Relief System, T-Block Needle/Catheter Tray, Model #TBTB1089TThe intended use of each component in the finished final trays is consistent with the intended use of the component stated in the respective 510(k) submission. | IFlow ONQ Pain Relief System | 2 | June 03, 2013 | 2013 | I-Flow LLC | I-Flow initiated two voluntary product removals of the ON-Q Pain Relief System: T-Block Needle/Catheter Tray and T-Block Needle/Catheter Kit because they may have potentially malfunctioned. |  | IFlow sent the customer notification letter on May 29, 2013, to 28 customers via FedEx along with a copy of the Service Guide to inform customers of the issue that can cause the Sting Ray to disconnect from the Catheter and how to prevent this problem from occuring. Customers are informed that in certain circumstances, ultrasound gel on the catheters significantly reduces the amount of force required to pull the catheter out of the Stingray connector. The In Service Guide clarifies that the method of catheter securement. It also instructs that wiping off any excess ultrasound gel from the catheter prior to inserting it into the Stingray connector will help with connecting the catheter to the Stingray connector. Customers are instructed that there is no action required for the I Flow ON Q Pain Relief System and ON Q T Bloc. The recall letter is only to provide further instructions on successful set up of the SPIROL catheter with the Stingray Connector included in the ON-Q T-bloc trays. Customers with any questions are instructed to call IFlow Customer service at (800) 448-3569 or IFlow 24 hr Tech support at (800) 444-2728.On June 1, 2012, an I-Flow sales representative removed the catheters and connectors from the one consignee/customers.For questions regarding this recall call 949-206-2700. | Contact the recalling firm for information | 450 | 450 units | Nationwide distribution: USA including states of: CA, CT, FL, IL, IN, KY, NY, OH, OR, TN, TX, VA and WI. | Terminated | June 25, 2013 | 22 | 525 | Not_Computer
Z-1420-2013 | 64631 | stimulator, nerve | ETN | Surgical nerve stimulator/locator. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | REF 8229705 x1 NIM TriVantage EMG Endotracheal Tube 5.0 mm I.D. x 6.5 mm O.D. Medtronic Xomed, Inc., 6743 Southpoint Drive North, Jacksonville, FL 32216-0980 USA STERILE Rx OnlyThe electrodes are designed to make contact with the patient's vocal cords to facilitate electromyographic (EMG) monitoring of the laryngeal musculature during surgery when connected to a multi-channel EMG neuromonitoring device. | NIM TriVantage EMG Endotracheal Tube | 1 | June 04, 2013 | 2013 | Medtronic Xomed, Inc. | In March 2013, Medtronic issued a recall of the NIM TriVantage EMG Endotracheal Tube due to cuff leakage in the intraoperative cuff deflation which then requires re-inflation or replacement of the deflated tube in order to ensure the continued ventilation of the patient. |  | The firm, Medtronic Xomed, Inc. sent an "URGENT Product Recall Notification" letter dated March 14, 2013, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to return any affected devices still in their possession per the following procedure: 1. Check inventories for the affected products listed on the enclosed checklist. 2. Fill-in the "quantity on-hand" column on the checklist. 3. Fax checklist to Medtronic ENT at 1-904-296-2386. 4. If you have any of the subject product, contact Medtronic ENT Customer Service at 1-800-874-5797 to arrange for returns and replacement or credit. 5.When returning products, clearly mark the outside of the container Returned Goods Authorization (RGA) number. If you have any questions regarding the subject action or the content of the letter, contact Gabriela Anchondo at 904-279-7550. |  | 2591 | 2,591 units | Worldwide distribution: AL, AR, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MT, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, and WA; and countries of : Australia, Canada, Germany, Italy, Norway, Poland and United Kingdom. | Terminated | January 30, 2014 | 240 | 2591 | Not_Computer
Z-1451-2013 | 65160 | mask, surgical | FXX | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Comfort brand Ear-Loop Face Mask, Item #: 3100, Model Numbers: # 3100B, # 3100P, # 3100G, # 3100Y, # 3100W; Product is manufactured and distributed by Maytex Corporation, CAEar-loop face mask for basic medical/dental procedures. | Comfort EarLoop Face Mask | 2 | June 04, 2013 | 2013 | Maytex Corp | Recent Lab testing conducted in March 2013 indicates that the product does not meet the standard of "Differential Pressure of only 1.6 Delta P (mm H2O/cm2)" as indicated on the box/package of the product; and also product manufactured between 2008 - 2012 cannot be guaranteed as Fluid Resistant. | MISBRANDING: Labeling False and Misleading | The firm, Maytex Corporation, sent a "MEDICAL DEVICE SAFETY NOTICE" dated May 15, 2013, to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately contact your medical professional for treatment and contact Maytex Corporation to file a complaint should any illness or injury occur; discontinue use of the product if you detect any potential health hazards; if this product has been further sold or distributed, notify your customers to the retail level and send copies of safety notice to customer/users; and complete and promptly return the enclosed Medical Device Safety Notice Return Response Acknowledgement and Receipt form via mail to: Maytex Corporation 23521 Foley Street, Hayward CA 94545; fax at; 510-786-0209; email: www.maytex18@aol.com or by delivery to Maytex Corporation.Maytex Corporation has temporarily taken down their Website at www.maytexcorp.com to correct the wrong product specifications as described in the notice.If you have any questions, contact Maytex Corporation at 1-800-462-9839 or email: maytex18@aol.com Monday through Friday, 8:30 AM to 5:00 PM Pacific Time. |  | 43915 | 43,915 Cases | Worldwide distribution: USA (nationwide) including states of: AL, AZ, CA, CO, FL, GA, HI, ID, IL, IN, MA, MD, ME, MI, MN, MO, MT, NE, NH, NJ, NM, NY, OK, OR, PA, PR, TN, TX, UT, WA, and WI; and country of: Mexico. | Terminated | October 29, 2013 | 147 | 43915 | Not_Computer
Z-1477-2013 | 64164 | antisera, cf, herpesvirus hominis 1,2 | GQO | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Plastic bottle containing 10.5 mL of IgG ConjugateProduct Usage:Product is a used as a counterstain for various indirect fluorescent assay kits | IgG Conjugate | 3 | June 05, 2013 | 2013 | Bion Enterprises Ltd | Affected products have a labeled expiration date of one year too long. | PRODUCTION CONTROLS: Error in Labeling | MBL Bion sent an Urgent Product Recall letter dated January 28, 2013 to all affected customers via email. The letter identified the affected product, problem and actions to be taken. Customers were requested to destroy the affected product and to notify Bion of the number of units consumed and destroyed. For questions call 847-544-5044. |  | 207 | 207 units | Worldwide Distribution - United States (nationwide) in the states of IN, SC and the countries of Canada (Toronto); South Africa (Randburg) and Brazil (Rio de Janeiro) | Open |  |  | 912 | Not_Computer
Z-1476-2013 | 64550 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | SlingBar Standard.Designed to meet the needs for lifting humans. | SlingBar Standard | 2 | June 05, 2013 | 2013 | Hill-Rom, Inc. | Retrospective review found 4 sling bars that may present a potential for impalement if a patient, caregiver or bystander were to fall onto the hooks based on a simulated test method. In effort to eliminate any risk of injury, Hill-Romis requesting products be removed from field use andreplaced with an updated design. | DESIGN: Device Design | On 4/30/13 an URGENT MEDICAL DEVICE RECALL CORRECTION was issued to all consignees detailing the hazards and actions that are needed. Included with this notification is a Response Form/Receipt and a Replacement Guide, complete and return the form to Hill-Rom as soon as possible. Hill-Rom will send you the applicable number of sling bars to you at no cost. After you receive the new sling bars, we request that you discard the old ones. |  | 726 | 726 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, IL, IN, KS, LA, MA, MI, MN, MO, NC, NH, NJ, NV, OH, OK, PA, SC, TN, UT, WA, and WI, and the countries of Canada, Australia, Austria, Belgium, Switzerland, Germany Denmark, Spain, Finland, France, United Kingdom, Greece, Hong Kong, Ireland, Israel, Italy, Japan, Netherlands, Norway, New Zealand, Portugal, Sweden, and Turkey. | Open |  |  | 4710 | Not_Computer
Z-1485-2013 | 65258 | dermatome | GFD | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Zimmer Air Dermatome II, REF 00-8851-001-01handpiece, w/o hose, Rx, non-sterilegrafting instruments intended to provide variable graft thickness and width capabilities. | Air Dermatome II | 2 | June 05, 2013 | 2013 | Zimmer Surgical Inc | This removal is in response to complaints alleging that the Zimmer Air Dermatomes and Zimmer Air Dermatome II products would either not operate oroperate intermittently. Zimmer's investigation determined that the planetary gear teeth were broken. This may cause a delay or postponement of the surgery or result in a suboptimal graft retrieval with potential of requiring an additional donor site. | PRODUCTION CONTROLS: Process Control | The firm, Zimmer, sent an "URGENT: DEVICE REMOVAL" notice dated May 7, 2013, to all consignees/customers via FED Ex and Distributors via E-mail. Consignees who had devices repaired received telephone calls as well as the mailed Recall Notice. The Notice identified the devices being recalled; possible adverse events and directions to immediately stop use of and remove affected product, and obtain an RGA (Return Goods Authorization) number by calling (800)348-2759 or email information to dover.repairs@zimmer.com . to return the devices for repair to:Zimmer Surgical, Attn: QA/RA Dept.- Recall, 200 West Ohio Avenue Dover, Ohio 44622 USA; and to complete and return the Response Form-Immediate Response Required via email to CorporateQualitv.PostMarket@zimmer.com. If you have further distributed affected product(s), please provide the customer's information on the Business Response Form to Zimmer.If you have any questions regarding this action, please email us at CorporateQuality. PostMarket@zimmer.com or call 1-330-364-0989. |  | 13 | 13 units | Worldwide distribution: USA Nationwide including Puerto Rico; UAE, TAIWAN, SWEDEN, THAILAND, KOREA, SINGAPORE, HONG KONG CHINA, CHILE, CANADA, AUSTRALIA, JAPAN, ENGLAND, GERMANY, and EGYPT. | Open |  |  | 223 | Not_Computer
Z-1466-2013 | 65257 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 1000 Bed, Product No. P1160. The Hill-Rom 1000 Bed is intended for low to moderate acuity patient in the medical surgical area of the hospital. | HillRom 1000 Bed | 2 | June 05, 2013 | 2013 | Hill-Rom, Inc. | Complaints have been reported for brakes disengaging when the bed is subjected to a hard jolt or force to the frame. | PRODUCTION CONTROLS: Process Control | On 4/16/2013 an URGENT MEDICAL DEVICE CORRECTION notification letter was sent to all consignees. The firm plans a visit to all facilities to correct each bed. Call Hill-Rom Technical Support at 800-445-3720 if you have any questions. |  | 8737 | 8,737 devices | Worldwide distribution: US (Nationwide) and countries of: Argentina , Bahamas, Bahrain, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Ecuador, Egypt, Ethiopia, France, Ghana, Guam, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, N. Mariana Islands, Pakistan, panama, Paraguay, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, United States, US Virgin Islands, Venezuela and Vietnam. | Open |  |  | 8737 | Not_Computer
Z-1472-2013 | 65093 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Versaport Bladeless Optical Fixation Cannula - 5mmProduct Code: ONBFCA5STThe Versaport" Bladeless Optical 5mm trocars are intended for use in a variety of gynecologic, general, thoracic and urologic endoscopic procedures to create and maintain a port of entry. | Versaport Bladeless Optical Fixation Cannula 5mm | 2 | June 05, 2013 | 2013 | Covidien LP | Seals may disengage from the cannula which may result in a component inadvertently disengaging into the patients cavity | DESIGN: Device Design | Covidien sent an Urgent Medical Device Recall letter dated May 6, 2013,via federal express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the product and return all inventory to Covidien Surgical Solutions, 195 McDermott Road, North Haven, CT 06473. Customers were requested to return completed Response Form.Customers with questions were instructed to contact their Covidien representative or Covidien Customer Service, Monday through Friday, Sam - 6:30pm ET, at (800) 962-9888, option 1, and then option 2.For questions regarding this recall call 203-492-8165. | Contact the recalling firm for information | 57461 | 57461 untis | Worldwide Distribution - USA (nationwide) and internationally to the following countries:CanadaAustriaBahrainBenoniCenturionChatsworthCzech RepublicDenmarkEgyptFinlandFranceGermanyIrelandIsraelItalyJapanKuwaitNetherlandsNorwayPortugalQATARSaudi ArabiaSerbiaSingaporeSouth AfricaSpainOMANSwedenSwitzerlandUAEUnited KingdomVenezuela | Open |  |  | 218221 | Not_Computer
Z-1465-2013 | 65209 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Active Articulation E1.Intended for use with either primary or revision hip arthroplasty. | E1 Advantage Head | 2 | June 05, 2013 | 2013 | Biomet, Inc. | Biomet is recalling Active Articulation E1 Hip Bearing implants. Biomet has initiated this action following an investigation which identified that outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size.The implant itself is etched with the correct O.D. size. | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet sent an Urgent Medical Device Recall letter dated April 26, 2013 to consignees detailing that the outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size. This action requires the immediate location and discontinued use of the product and its return to Biomet. Ciustomers were advised to:Immediately locate and remove the identified device(s) listed from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. | Contact the recalling firm for information | 8 | 8 units | Distributed in New York. | Terminated | August 29, 2013 | 85 | 8 | Not_Computer
Z-1482-2013 | 64991 | impactor | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Straight Impactor - Dorado IBC, Part Number: SI02140. The device is used to position implants after they have been inserted into the disc space during implantation of intervertebral body fusion devices. | Dorado Straight Impactor | 2 | June 05, 2013 | 2013 | SpineFrontier, Inc. | During two separate occasions an impactor plate on the handle broke off while being struck by a mallet. | DESIGN: Process Design | Firm states they notified consignees of recall between 5/10/08 and 5/22/2008. Zero product is currently in distribution. |  | 10 | 10 | Worldwide distribution: US (nationwide) to states of: CO, MA, and TX; and country of: Jamaica. | Open |  |  | 10 | Not_Computer
Z-1486-2013 | 65314 | suture, absorbable, synthetic, polyglycolic acid | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Digitex Absorbable Polydiaxanone Suture, Size 2-0. Catalog Number / Code of product: 52031, 5203101400.The Digitex Delivery Device is a sterile, disposable device intended to deliver a suture to the operative site. The device assists in suturing by passing a needle through the tissue to capture the suture, and suture ligation and know-typing by holding the suture. | Digitex 20 Polydioxanone (PDO) suture | 2 | June 05, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall on three lots of Digitex Absorbable Polydiaxanone Suture, Size 2-0 because it is mislabeled with a three year expiration date. | EXPIRATION DATING: Incorrect or No Expiration Date | The firm, Coloplast, sent a letter dated May 8, 2013 to its customers. The letter described the problem, the product and actions to be taken. The customers will be contacted by Coloplast Customer Service with instructions for the handling and return of any unused product.If you have any additional questions, the dedicated Customer Service number regarding this effort is 800-258-3476. |  | 772 | 772 units | US Distribution to: KY, MI, OH, PR, TX, and VA. | Open |  |  | 772 | Not_Computer
Z-1463-2013 | 65168 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | da Vinci 5 mm Instrument Cannulae for use with the da Vinci S and Si Surgical Systems.Used to establish a port of entry for Intuitive Surgical EndoWrist instruments, | da Vinci 5 mm Instrument Cannulae | 2 | June 05, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical is implementing a voluntary Field Action to implement a modified inspection process for the 5 mm Cannula (PN 420011) and 5 mm Flared Cannula (PN 420262). |  | Intuitive Surgical sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 14, 2013 to all affected customers. The letter identified the product,problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 2101 | 2101 consignees | Worldwide distribution: US (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Open |  |  | 2101 | Not_Computer
Z-1516-2013 | 65113 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | American Optisurgical's TX1 Tissue Removal System Console. The console is a software unit that is indicated for use in surgical procedures. | TX1 Tissue Removal System | 2 | June 06, 2013 | 2013 | American Optisurgical Inc | The recall has been initiated due to an overheating protection voltage suppressor on the redundant safety protection circuit board. Secondary to the recall, American Optisurgical, Inc. will be making a correction to add spacers between the valve and cassette to correct a small irrigation leak that can occur after the priming cycle. Use of this product may cause the device to lose power and stop wo | DESIGN: Device Design | "Urgent Medical Device Recall" notifications letters were sent to customers on May 20, 2013 via USPS registered mail. Included in the Recall Notification Letter were the recommended actions and instructions for customers. Questions were referred to (949) 580-1266. |  | 194 | 194 units | Worldwide distribution, US, including the states of United States: Alabama, Arkansas, Arizona, California, Colorado, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington Wisconsin, West Virginia and the countries of Singapore and the Philippines. | Open |  |  | 194 | Not_Computer
Z-1490-2013 | 65321 | gown, surgical | FYA | Surgical apparel. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Stryker Instruments0400-720-000 T4 Pullover Toga (L/XL) Sterile, EO; single use; Product Usage:Protection System and are intended to protect the patient, healthcare personnel and operating room personnel against contamination, exposure of infectious bodily fluids, and the transfer of microorganismsand particulate material. | Hytrel Toga | 2 | June 06, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Based on an accelerated aging test for this product packaging, there is an indicator of potential for a breach in the sterile barrier in the header bag and potentially the blue sterilization wrap for togas within the second and third year of shelf life. The test method used involved high heat to accelerate the aging of the packaging. Stryker is conducting additional testing and assessing this pr | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker sent an Urgent Urgent Recall noitifications letter dated May 7, 2013 via certified US mail and E-mail l to all affected customers, distributors and direct accounts. The letter identified the affected products, problem, risk to health, and actions to be taken. Customers were instructed to check inventory, immediately quarantine any affected product, and open each package before discarding to prevent use. Customers were directed to fill out and return a Business Reply Form to Stericycle, Phone: 888-943-2396 Fax: 888-943-4192, Email: stryker5893@stericycle.com verifying directions and confirming recalled product was destroyed. For questions call 269-389-4354 or 269-389-4518. |  | 4804 | 48,040 units, ( 4804 boxes) | USA Nationwide Distribution | Open |  |  | 540081 | Not_Computer
Z-1505-2013 | 65098 | skin prep tray (kit) | OJU | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | FAST Procedure Kit.Intended for use as a Minor Surgical Procedures Tray. | FAST Procedure Kit | 2 | June 06, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc initiated a recall of ChloraPrep Hi-Lite Orange 3mL Applicator, in response to a sub-recall letter from Medline Industries, Inc due to lack of sterility assurance caused by an incomplete seal on the individualunit packaging. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | American Optisurgical Inc. sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL AND PRODUCT EXCHANGE" letter dated May 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-580-1266 for questions regarding this notice. | Contact the recalling firm for information | 63 | 63 individual applicators | Nationwide Distribution-USA (nationwide) including the states of AZ, CA, FL, GA, IL, KS, MA, MI, NC, NJ, OH, PA, TN, TX, VA, and WI. | Open |  |  | 63 | Not_Computer
Z-1506-2013 | 65006 | instrument, vitreous aspiration and cutting, ac-powered | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | The Constellation Vision System Pneumatic Handpiece.Indicated for both anterior and posterior segment opthalmic surgery. | Constellation Vision System Pneumatic Handpiece | 2 | June 06, 2013 | 2013 | Alcon Grieshaber AG | Alcon is conducting a Medical Device Removal for the Constellation Pneumatic Handpiece, DSP, product number 725.01. This action is being taken because the forceps or scissors tips (sold separately) may dislodge from the hand piece, if not attached by the user according to the Directions for Use. |  | Alcon sent an "URGENT: MEDICAL DEVICE REMOVAL" letter on April 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Alcon Sales Representative for questions regarding this notice. |  | 8455 | 8,455 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Canada, Chile, China, Egypt, France, Germany, Great Britain, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Kyrgyzstan, Lebanon, Korea, Malaysia, Mexico, Pakistan, Peru, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Turkistan, Venezuela, and Veit Nam. | Open |  |  | 8455 | Not_Computer
Z-1488-2013 | 65062 | table, operating-room, electrical | GDC | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Hand Controls for Operon D750, D752, D760, D820 and D850The hand controls are used to operate OPERON surgical tables | OPERON Hand Controls for Operon D750, D752, D760, D820 and D850 | 2 | June 06, 2013 | 2013 | Berchtold Corp. | The firm received complaints for unintended movements. | DESIGN: Device Design | Berchtold sent an Urgent Medical Device Recall letter in March 2013, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were recommended to clean the affected product and use the instructions to reduce the possibility of unintended movements until the firm could replace the affected product. Customers with questions were advised to contact the Berchtold Operon Pendanct Recall Coordinator at 800-243-5135, or 843-569-6100, ext 306.For questions regarding this recall call 843-569-6133. |  | 3763 | 3763 | Nationwide Distribution | Open |  |  | 3763 | Not_Computer
Z-1487-2013 | 65223 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray System DX-D 100Agfa's DX-D 100 is indicated for use in providing diagnostic quality images to aid the physician with diagnosis. Systems can be used with MUSICA2 image processing to create radiographic images of the skeleton including skull, spinal column and extremities: chest, abdomen, and other body parts. | DXD 100 MOBILE DR | 2 | June 06, 2013 | 2013 | AGFA Corp. | Customers could potentially experience intermittent, unintended and illogical movement when using the product. |  | Agfa sent an "URGENT SAFETY NOTICE" letter dated May 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed if you experience unintended movement, the unit has to be taken out of operation. Please notify you Agfa service contact at once.Please distribute this information within your facility to all those who need to be aware of this. Please complete the feedback form as soon as possible and return to us.For further questions please call 1-877-777-2432. |  | 104 | 104 | Worldwide Distribution - US Nationwide including the states of AL, FL, IL, IN, IA, KY, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OR, SC, TN, VA, WV, WY and the country of Canada | Open |  |  | 104 | Not_Computer
Z-1458-2013 | 65192 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Part 510450 Disposable Infusion Pump, ...Elastomeric Pump 600 mL( 10mL/hr ),... found in the following kitsGoBlock Kit 510458,Disposable Single Infusion Pump | GOBlock | 1 | June 06, 2013 | 2013 | Symbios Medical Products, LLC | Symbios Medical Products, Indianapolis, Indiana, is performing a voluntary recall of all GoPumps found in the GoPump Rapid Recovery System kits and all GoBlock kits manufactured with flow control components assembled prior to July 2012. The recall is being expanded after receiving additional reports relating to excessively high flow rates on the 150 mL reservoir, 2 mL/hr GOPump (PN510031), and the |  | Symbios sent UrgentMedical Device Recall notices to all consignees beginning May 10, 2013. A second corrected/amended notice was sent on May 14, 2013 to consignees and a National Press release was issued. May 30,2013, a third amended Urgent Recall Notice was sent to all consignees correcting the pump description and lot number range table on page 5 of the notice. The notices and press release provided identification of the recalled devices including devices in the previous February 2013 recall, described possible adverse events, and included directions to to quarantine product in a secure location, in preparation for return. Consignees were asked to call 317 225-4447 ext. 25, 8:00 a.m to 4:00 p.m. EDT to receive a Returned Goods Authorization Number prior to returning product. Consignees were asked to complete and return the Attached Verification Form. Question and Concerns should be addressed to 317-225-4447 ext 25. |  | 576 | 33, 543 for all kits | US Nationwide distribution. | Open |  |  | 33543 | Not_Computer
Z-1507-2013 | 65182 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXSTENT Biliary Self-Expanding Stent System. Product Usage:The FSS FLEXSTENT Biliary SES product is indicated for use in the palliation of malignant strictures in the biliary tree. | FLEXSTENT Biliary SelfExpanding Stent System (US Commercial 80 cm Length) | 2 | June 06, 2013 | 2013 | Cordis Corporation | Holes and/or tears can be introduced in the FLEXSTENTÂż inner pouch under some conditions, and to a lesser degree, in the outer pouch. | DESIGN: Packaging Design/Selection | Cordis sent an Urgent Medical Device Correction and Potential Removal letter dated April 24, 2013 e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected products to prevent further use of the affected product, until inspected. ÂŽ Use only product identified with a blue dot next to the lot number on the carton face label, carton end label, and outer pouch label, or product with upgraded sterile barrier system (not yet available, as of the date of this letter), review, complete, sign and return the enclosed Acknowledgement Form and maintain a copy of this notice with the affected product. For questions related to the issue, please contact our FSS/Cordisrepresentative, or please contact Cordis Direct at (800) 781-0282, Mondaythrough Friday from 9:00 AM to 5:00 PM EDT. | Contact the recalling firm for information | 23 | 23 units | Worldwide Distribution - USA (nationwide) and the states of: AZ, CA, CT, FL, GA, IA, IL, LA, MD, MI, MS, NC, NJ, NY, OH, PA, RI, SC, TX, WA, WI, WV, and the countries of: Czech Republic, Italy, Netherlands, and New Zealand | Open |  |  | 2724 | Not_Computer
Z-1515-2013 | 65224 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CentaurÂŽ iPTH Calibrator 6 pk; REF: 10492387, 10630628.For in vitro diagnostic use in calibrating Intact PTH assays. | Siemens ADVIA Centaur iPTH Calibrator | 2 | June 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | iPTH Assay- reagent and calibrator lots may have demonstrated higher than anticipated variability | PREMARKET APPROVAL: No Marketing Application | Siemens sent an Urgent Medical Device Correction letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions, contact your Siemens Technical Solutions Center or your local Siemens technical support. |  | 198 | 198 (6 pk) | Nationwide Distribution. | Open |  |  | 3908 | Not_Computer
Z-1517-2013 | 65289 | catheter, embolectomy | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | lnvatec SpA brand Diver C.E., Model/Catalog Numbers: DVR010014145, DVR011014145, Product is manufactured by lnvatec SpA, Roncadelle, Italy, and distributed by Invatec, Inc., Bethlehem, PA and by Medtronic Endovascular Therapies, Santa Rosa, CAThe Diver C.E. product is an aspiration catheter indicated for the removal of fresh, soft emboli and thrombi from vessels in the arterial system. | lnvatec SpA brand Diver C.E | 2 | June 07, 2013 | 2013 | Invatec Llc | Devices labeled for an intended use was not included in the existing 510(k) or PMA, in that, the Indications for use was modified. | PRODUCTION CONTROLS: Labeling Mix-Ups | Invatec sent an Urgent - Device Field Action letter dated March 4, 2009, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were required to return a confirmation form, acknowledging the action taken, and identifying affected units in their inventory for return of the product for replacement. The firm indicated that this recall was fully executed in 2009, and this 806 captures the events that took place at that time.For questions regarding this recall call 707-566-1143. | Contact the recalling firm for information | 177 | 177 units | Nationwide Distribution | Terminated | June 10, 2013 | 3 | 177 | Not_Computer
Z-1436-2013 | 65248 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: AFX" Introducer System, Model Number S17-45.Lot Numbers: 1079840, 1079843, 1079844, 1079845.Product Usage:The AFX Introducer System is intended for use to facilitate the introduction of catheters and other medical devices into the vasculature and to minimize blood loss associated with such introduction. | Catheter, Introducer | 1 | June 07, 2013 | 2013 | Endologix Inc | Endologix, Inc. initiated this voluntary recall of certain lots of Endologix, Inc. due to reports of dilator breakage during clinical procedures. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Endologix Inc sent an "URGENT-MEDICAL DEVICE RECALL NOTICE" letters dated May 13, 2013 to all affected customers. The letter identified the affected product, reason for recall, and actions to be taken. The letter instructed customers to do not use or further distribute any affected product. The letter instructed customers to share this information with physicians that perform these procedures at the facility. The firm will call to confirm that this device is located at the facility. After confirmation the firm will retrieve the device. For questions call (800) 983-2284. |  | 53 | 53 | Worldwide Distribution - USA Nationwide in the states of: NY, NH, IN, MI, FL and NJ and in the countries of: Europe and Latin America. | Terminated | August 19, 2013 | 73 | 53 | Not_Computer
Z-1522-2013 | 65283 | bur, diamond coated, reprocessed | NME | Dental bur. | Dental | Dental | 510(k) | StarDental Classique(TM) Diamond Dental Bur. Used for specific applications in dental procedures. | StarDental Classique(TM) Diamond Dental Bur | 2 | June 07, 2013 | 2013 | Dental EZ Stardental Division | Recall was initiated because a manufacturing issue was found that prevented the diamond coating from adhering properly to the shank of the diamond bur for these lots. | PRODUCTION CONTROLS: Equipment Maintenance | Dental EZ Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter on May 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 866-383-4636 ext.4350 for questions regarding this recall. |  | 3560 | 3,560 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IA, IN, NJ, NV, NY, PA, TN, TX, and WI, and the country of Japan. | Open |  |  | 3560 | Not_Computer
Z-1520-2013 | 65081 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Glenosphere orientation guide Product Usage: The glenosphere orientation guide is an instrument used in reverse shoulder arthroplasty. The orientation guide is used to properly align the glenosphere. | Glenosphere Orientation Guide for the Delta XTEND Reverse Shoulder System | 2 | June 07, 2013 | 2013 | DePuy Orthopaedics, Inc. | The Glenosphere Orientation Guide Instrument Used with the Delta XTEND Reverse Shoulder is being recalled because an arrow that is etched on the instrument was incorrectly placed on the opposite side of the instrument. Using the improperly etched arrow as a guide could potentially lead to incorrect orientation of the glenosphere implant. | PRODUCTION CONTROLS: Process Control | DePuy Synthes sent and Urgent Information - Medical Device Instrument Recall Notice letter dated May 3, 2013 via email to affected customers. The letter identified the affected product, problem, clinical implications, and potential adverse events. Arrangements will be made for all affected Glenosphere Orientation Guides (part #230795000/lot #5120443) to be returned. Customers were asked to return the response form via fax (574-371-4964). Product related questions should be directed to DePuy Synthes Joint Reconstruction (1-888-554-2482). Recall inquiries should be directed to Kim Earle, Recall Coordinator, 574-371-4917 (M-F; 8a.m.-5 p.m. EDT). |  | 20 | 20 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Canada, Czech Republic, England, Germany, Switzerland and Turkey. | Terminated | August 28, 2013 | 82 | 20 | Not_Computer
Z-1518-2013 | 65108 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | TX1 Tissue Removal System disposable handpieces.Product Usage: The TX1 Tissue Removal System is indicated for use in surgical procedures where fragmentation, emulsification and aspiration of soft tissue are desirable, including General Surgery, Orthopedic Surgery, Laparoscopic Surgery and Plastic and Reconstructive Surgery. | TX1 Tissue Removal System | 2 | June 07, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc. is recalling certain lots of TX1 Tissue Removal System disposable handpiece due to inventory being released into distribution without a cleared FDA 510(k) to cover recent product changes. Use of this product may lead to an underperformance of the device. | PREMARKET APPROVAL: No Marketing Application | American Optisurgical Inc. (AOI) sent an Urgent Medical Device Recall letter along with a response sheet via USPS registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine your inventory and quarantine product subject to recall. Once the product has been identified, return all products by FedEx to AOI using the AOI FedEx account 30655659-2, and referencing the Returned Goods Authorization (RGA) number:RGA# (to be assigned). For questions contact American Optisurgical Inc. Customer Service 949-580-1266. | Contact the recalling firm for information | 1436 | 1,436 units | Worldwide Disitribution - USA Nationwide in the states of Alabama, Arkansas, Arizona, California, Colorado, Delaware, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, Wyoming and the countries of Quezon City, Philippines, Singapore, and Panama City, Panama. | Open |  |  | 1436 | Not_Computer
Z-1521-2013 | 65217 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Applied Medicals EpixÂŽ and Direct Drive Laparoscopic Graspers, Model Numbers C4130 and C4140.Used for grasping and manipulating tissue during general or laparoscopic surgery. | Epix and Direct Drive Laparoscopic Graspers | 2 | June 07, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on specific lot numbers of its EpixÂż and Direct Drive LaparoscopicGraspers. When the ratchet trigger is activated, if a great enough force is applied, there is a potential for the triggerto fracture and the jaws to remain in the closed position. The likelihood of this situation to occur and result inpermanent patient injury is highly unlikely; | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Applied Medical notified customers of the recall beginning on May 13, 2013. The "URGENT: MEDICAL DEVICE RECALL" letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact the firm at (949) 713-8041. |  | 127305 | 127,305 units | Worldwide Distribution: USA Nationwide and countries of: Argentina, Austria, Australia, Bahrain, Belgium, Canada, Chile, Columbia, Denmark, Ecuador, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Jordan, Kuwait, Malaysia, Lebanon, Lybia, Mexico, Netherlands, New Zealand, Pakistan, Peru, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Spain, South Africa, Switzerland, Taiwan-Province of China,Turkey, United Arabs Emirates United Kingdom, Uruguay, and Venezula-Bolivarian Republic. | Terminated | February 06, 2014 | 244 | 127305 | Not_Computer
Z-1525-2013 | 65188 | panels, test, susceptibility, antimicrobial | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens brand MicroScan Dried Negative Breakpoint Combo 47 panels (B1017-417), SMN #: 10483099), For use in determining quantitative and/or qualitative antimicrobial agent susceptibility. | MicroScan Dried Negative Breakpoint Combo | 2 | June 08, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The product has the potential for false negative or delayed reaction for Arginine (ARG) results. In addition, 42 hour identification panel holds may occur due to a weak ARG reaction. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Siemens Healthcare Diagnostics Representative for information concerning this notice. | Contact the recalling firm for information | 683 | 683 units | Nationwide Distribution. | Terminated | September 17, 2013 | 101 | 683 | Not_Computer
Z-1523-2013 | 65296 | separator, automated, blood cell and plasma, therapeutic | LKN |  |  | Gastroenterology/Urology | 510(k) | AMICUS Exchange Kit; Product Usage:This kit is designed for use with the AMICUS separator for the Therapeutic Plasma Exchange (TPE) procedure. Sterilized by irradiation. Sterile fluid path. Non-pyrogenic fluid path. | AMICUS Exchange Kit | 2 | June 08, 2013 | 2013 | Fenwal Inc | Fenwal has initiated a voluntary Urgent Product Recall for lot FA12C07070 of product code R4R2339, AmicusExchange Kits. Fenwal identified a labeling issue with this batch of Product Code R4R2339 (which was cleared for European use only) in which this batch was distributed within the United States market. The problem is identified as labeling that did not include the written description for the s | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Fenwal sent an Urgent Product Recall letter dated May 24, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, discontinue the use of this prod.uct, return the affected product to Fenwal, and complete and return the attached reply form. For questions call 1-800-333-6925. |  | 19 | 19 Cases (114 Units) | USA Nationwide Distribution | Open |  |  | 19 | Not_Computer
Z-1524-2013 | 65046 | clip, implantable | FZP | Implantable clip. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BARD Biopsy Systems UltraClip Dual Trigger breast tissue marker. Intended for use to attach to soft breast tissue at the surgical site during an open surgical breast biopsy. | UltraClip Dual Trigger breast tissue marker | 2 | June 08, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some devices with the recalled product code / lot numbercombination may have the metal tissue marker separated from the needle prior to use. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Bard Peripheral Vascular, Inc. initiated a recall of certain product code and lot number combinations of Bard UltraClip Dual Trigger Tissue Marker by contacting customers on 04/16/2013. On 04/18/2013, BPV sent a Customer Recall notification Letter via FedEx with proof of delivery notice. The letter included reason for recall, a table of affected product code and lot numbers, instructions (Do not use or further distribute any affected product), a response form (titled Recall and Effectiveness Check Form) enclosed, and contact information. |  | 3085 | 3,085 | Worldwide Distribution: USA Nationwide and Internationally to Europe. | Terminated | January 30, 2014 | 236 | 3085 | Not_Computer
Z-1526-2013 | 64276 | marker, radiographic, implantable | NEU | Implantable clip. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | BARD Biopsy Systems (A Business Unit of Bard Peripheral Vascular, Inc.), UltraClipÂŽ Dual Trigger Breast Tissue Marker 17g X 10cm needle, Ultrasound Enhanced Coil Marker 5 Pack; Sterilize, Non-pyrogenic, Rx only, Single Use; Bard Peripheral Vascular, Inc., 1625 West 3rd Street, Tempe, AZ 85281 USA.PK4320100 Rev. 0 03/10The UltraClip Dual Trigger breast tissue marker is a sterile, single use device comprised of a disposable introducer, and a metal implantable tissue marker with polyvinyl alcohol (PVA). | BARD Biopsy Systems | 2 | June 10, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is initiating this recall because a combination of BardÂż UltraClipÂż Dual Trigger Tissue Marker's and support catheters (Microsheath and Usher) were inadvertently distributed to customers without completing the sterilization process (non sterile). | PRODUCTION CONTROLS: Process Control | The firm, BARD Peripheral Vascular, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter on January 23, 2013, to its customers. Bard EXPANDED its recall to include (MicroSheath and Usher support Catheters) and sent another "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 5, 2013 describing the additional products, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Do not use or further distribute any affected product; check all inventory locations within your institution; remove any identified product from your shelves and return to Bard Peripheral Vascular, Inc. 1415 W. 3rd Street, Tempe, AZ 85281; if you have further distributed any of the product, immediately contact that location and advise them of the recall and have them return to BPV; If you have used the affected product, consider notifying, educating and monitoring those affected; and complete and return the Recall and Effectiveness Check form via fax to: ATTN; Recall Coordinator in Customer Service at 1-800-994-6772, even if you no longer have possession of the product; if unable to FAX call BPV at 1-800-321-4254 Option #2 Ex 2727 and report verbally.Please call our Recall Coordinator at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at HUsilvia.carrillo@crbard.com with any questions. | Contact the recalling firm for information | 815 | 815 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KY, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, TX, VA, WA, WV and WY. | Terminated | June 10, 2013 | 0 | 1370 | Not_Computer
Z-1531-2013 | 65341 | pump, infusion, analytical sampling | LZF | Infusion pump. | General Hospital | General Hospital | 510(k) | VIA Medical Collection Bag; The collection bag is an accessory waste bag used in conjunction with the sterile LVM blood gas sensor with the LVM Monitor. The collection bag is packaged in a tyvek pouch which is included in either replacement kit (777-3103) or LVM sensor kit (777-3102). | VIA Medical Collection Bag | 2 | June 11, 2013 | 2013 | International Biomedical, Ltd. | International Biomedical has received reports of interference between the collection bag and the extension set connectors. Evaluation of the affected components indicates that the collection bag connector can become occluded due to an over-tightening of the luer lock by the end user. To alleviate the occurrence of this issue in the fixture, the connector on the collection bag has been improved to | DESIGN: Device Design | International Biomedical decided to recall and sent "Product Bulletin" notification letters to their consignees on 05/24/2013 or 05/30/2013. The letter described the product issue and provided customers with recommended actions. Questions were directed to (512) 873-0033. | Contact the recalling firm for information | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 29, 2013 | 48 | 7000 | Not_Computer
Z-1530-2013 | 65330 | staple, implantable | GDW | Implantable staple. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ECHELON 60mm Endoscopic Linear Cutter Reloads Black in Flexible Blister pack with Tyvek lid.The ECHELON and ECHELON FLEX families of Endoscopic Linear Cutters (articulating and straight) are sterile, single patient use instruments that simultaneously cut and staple tissue. There are six staggered rows of staples, three on either side of the cut line. The ECHELON 60 instruments have a staple line that is approximately 60 mm long and a cut line that is approximately 57 mm long. The shaft can rotate freely in both directions and an articulation mechanism on articulating instruments enables bending the distal portion of the shaft to facilitate lateral access of the operative site. The instruments are shipped without a reload and must be loaded prior to use. A staple retaining cap on the reload protects the staple leg points during shipping and transportation. The instruments'ÂŽ" lock-out feature is designed to prevent a used reload from being refired. | ECHELON 60mm Endoscopic Linear Cutter Reloads Black | 2 | June 11, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary worldwide recall for ECHELON 60mm Black Reload (ECR60T) due to the potential for incomplete staple line formation from reload damage during the firing sequence. | DESIGN: Device Design | Ethicon sent a Urgent: Medical Device Recall letter dated May 29, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were instructed1.To examine your inventory immediately to determine if you have affected product manufactured by Ethicon and remove the affected product(s).See Attachment A for assistance in identifying affected product.2. Complete the Business Reply Form (BRF) Attachment C and fax or email it to Stericycle at 1-888-943-4190. Please reference Event 4706. If you have product to be returned, keep a copy of this form for your records.Your JJHCS number has been pre-populated on the BRF. Replacement product will be issued for all returned recalled product.3. All affected product must be returned by August 30, 2013, to receive replacement product. To return affected product, photocopy your completed Business Reply Form, placeit in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Returned ECHELON 60mm Black Reload (ECR60T) will be replaced with corrected ECHELON 60mm Black Reload (ECR60T) as soon as supply becomes available.4. If you do not have affected product in your inventory, you should still complete and return the Business Reply Form, indicating you have no affected product.5. Please share this information with all of the appropriate staff at your facility.6. If you need assistance identifying an alternative product code, please speak directly with your Ethicon Sales Representative. If you need clinical or product support, please contact your local sales representative or call ourCustomer Support Center, 1-877-ETHICON (877-384-4266 OPTION 6 to speak with a registered nurse). If you need additional shipping labels or a communications package, contact the Stericycle support center at 1-888-943-2397 and reference Event 4706. |  | 57540 | 57,540 Units | Worldwide Distribution - US Nationwide all fifty states and including the countries of Canada, Belgium, India, Malaysia, Saudi Arabia, Brazil, Israel, Mexico, Singapore, Jordan, New Zealand, Taiwan, Chile, Korea, Puerto Rico, Turkey, Columbia, Kuwait, Qatar, United Arab Emirates, Egypt, Lebanon and Russia. | Terminated | February 11, 2014 | 245 | 57540 | Not_Computer
Z-1533-2013 | 65395 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD PERIPHERAL VASCULAR VACCESSÂŽ PTA Balloon Dilatation Catheters, 8mm x 4cm x 80cm.Product Usage: VaccessÂŽ PTA Balloon Dilatation Catheters are recommended for use in Percutaneous Transluminal Angioplasty of the femoral, iliac and renal arteries and for the treatment of obstructive lesions of native or synthetic arteriovenous dialysis fistulae. This device is also recommended for post-dilatation of balloon expandable stents, self-expanding stents and stent grafts in the peripheral vasculature. This catheter is not for use in coronary arteries. VaccessÂŽ PTA Balloon Dilatation Catheters are supplied sterile and intended for single use. | PTA Balloon Dilatation Catheters | 2 | June 12, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some product code/lot number combination of devices may have the incorrect balloon diameter labeled on the inflation hub. Specifically, the balloon size printed on the hub may read 7mm x 4cm, when the actual balloon size is 8mm x 4cm. All other product labeling is consistent with the actual 8mm x 4cm balloon size. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | BARD sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated May 30, 2013 via FedEx overnight with proof of delivery notice to all affected customers. The letter identified the affected product, reason for recall, actions to be taken, and contact information. Customers were instructed to complete the enclosed Recall and Effectivenss Check Form and fax to the attention of Recall Coordinator in Customer Service at 1-800-994-6772. | Contact the recalling firm for information | 415 | 415 | US Nationwide Distribution | Terminated | October 23, 2013 | 133 | 415 | Not_Computer
Z-1536-2013 | 65091 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 90 Degree Cannulated Infant Blade Plate 35mm x 5mm x 3 HoleProduct Usage:The OrthoPediatrics Locking Cannulated Blade Plate System is intended for fixation of long bone fractures and osteotomies in all pediatric subgroups (except neonates) and in small stature adults. Specific indications include: intertrochanteric derotation and varus osteotomies, femoral neck and pertrochanteric fractures, intertrochanteric valgus osteotomies, proximal and distal tibial osteotomies and humeral fractures and osteotomies. | OrthoPediatrics PediLoc Locking Cannulated Blade Plate System | 2 | June 12, 2013 | 2013 | OrthoPediatrics Corp | Post Market Surveillance of the 90 degree Cannulated Infant Blade Plate indicated that there have been five reports concerning bending of the plate intra-operatively. | DESIGN: Device Design | OrthoPediatrics sent an Urgent Voluntary Recall notifiation via telephone and email to all affected consignees. The letter instructed consignees to take possession of any affected devices and quarantine immediately. A Mandatory Reply Form is included with the notification letter to be completed and returned. Please contact Greg Teghtmeyer if you have any questions (574-268-6379). | Contact the recalling firm for information | 130 | 130 devices were shipped | USA Nationwide Distribution | Terminated | September 19, 2013 | 99 | 130 | Not_Computer
Z-1543-2013 | 65374 | transducer, blood-pressure, extravascular | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Deltran Disposable Pressure Transducer Kit (DPT Kit) Extracorporeal disposable pressure transducer and tubing set with integral stopcock and flush device, Product Usage: Deltran is a disposable pressure transducer that provides consistent and accurate readings during invasive blood pressure monitoring of humans and animals. | Deltran Kit | 2 | June 13, 2013 | 2013 | Utah Medical Products, Inc | Utah Medical Products, Inc. is recalling certain Deltran Disposable Pressure Transducer Kits due to complaints related to packaging integrity of the sterile packaged devices. | PRODUCTION CONTROLS: Packaging | Utah Medical Products (UTMD) sent a Recall Notice letter dated June 3, 2013 via fax and email to all affected customers and will be completing follow up telephone calls by June 7, 2013. The letter identified the affected product, problem, health risk, and actions to be taken. UTMD instructed customers to locate and return affected products for exchange. For questions call Customer Service at 800-533-4984. | Contact the recalling firm for information | 1800 | 1,800 kits | Worldwide Distribution - United States Nationwide in the states of (MO, NE, KY, FL, OH, NY), and countries of Columbia, Brazil, Venezuela. West Indies, Latin America, Ecuador are distributed through a US direct consignee exporter. | Terminated | September 19, 2013 | 98 | 1800 | Not_Computer
Z-1539-2013 | 65285 | excimer laser system | LZS |  |  | Ophthalmic |  | STAR S3 ActiveTrack Excimer System (STAR S3)Mfg by AMO Manufacturing USA, LLC510 cottonwood DriveMilpitas, CA 95035The STAR Excimer Laser System, a class Ill device, is indicated for laser assisted in situ keratomileusis (LASIK) eye surgery in certain patient populations. | STAR S3 ActiveTrack Excimer System (STAR S3) | 2 | June 13, 2013 | 2013 | Abbott Medical Optics, Inc. | Under specific conditions, there is an unlikely potential to result in inadvertent laser firing of the laser. |  | AMO Manufacturing sent an Advisory Notice dated May 28, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the attached advisory notice to all affected customers in their assigned region. Customers were instructed to notify each surgeion at their site and complete and return the attached form to AMO via fax as an an acknowledgement of receipt of this advisory notice and that they understand the information contained within the letter. Customers with questions were instructed to contact their regional contact.For questions regarding this recall call 408-273-4016. |  | 1882 | 1,882 total all systems | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Great Britian, Greece, India, Ireland, Italy, Netherlands, New Zealand, Norway, Poland, Puerto Rico, Singapore, Sweden, Turkey, Japan, Spain, Chile, Argentina, Colombia, Brazil, Taiwan, Mexico, South Korea, Venezuela, Paraquay, South Korea, Russian Federation, Costa Rico, Portugal, Uruquay, Martinique, Israel, Bolivia, Ecuador, Lebanon, Thailand, Saudi Arabia, Egypt, Dominican Republic, United Arab Emirates, Czech Republic, Indonesia, Kuwait, Ukraine, Guatemala, Phillipines, Trinidad, Tobago, Mongolia, Jordan, Malaysia, Hong Kong, Bulgaria, Oman, Kazakhstan, Bulgaria, Peru, Bolivia, Croatia, Romania, Yemen, Vietnam, Uzbekistan, Libyia, Slovakia, Algeria, El Salvador, and Cyprus.. | Open |  |  | 1882 | Not_Computer
Z-1542-2013 | 65173 | instruments, surgical, cardiovascular | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | QuickDraw Venous Cannula with Insertion Kit. ThruPort Systmes. Models QD22 (22 Fr) and QD25 (25 Fr). The product is packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Two pouches (cannula and accessories) go into one shelf carton. One shelf carton goes in one shipper carton.Use of the QuickDraw venous cannula is indicated for patients undergoing cardiopulmonary bypass. The QuickDraw venous cannula serves to drain non-oxygenated blood from the venae cavae or right atrium during cardiopulmonary bypass. The product is packaged sterile and nonaÂŽ pyrogenic in a sealed, peel-type pouch. | QuickDraw Venous Cannula | 2 | June 13, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling the QuickDraw Venous Cannula due to complaints of cannula separation during withdrawal of the device when the Percutaneous insertion method is used. | DESIGN: Process Design | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL - ACTION REQUIRED" notice dated May 14, 2013 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to review our entire inventory for any QuickDraw Venus Cannula that has not expired; quarantine affected product from your inventory; discontinue use and return unused devices; complete and return the acknowledgment form via fax to Edwards Customer Service on 800-422-9329 within three days of receipt of this Field Safey Notice; contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number and replacement product; return product to Edwards Lifesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020 Attention: RECALL, RGA#xxxxxx.If you have any questions that have not been answered by this notice, please call Edwards Customer Service at 800-424-3278 from the hours of 6:00AM - 4:30PM PST. |  | 22616 | 22,616 | Worldwide distribution: US Nationwide and country of Canada. | Open |  |  | 22616 | Not_Computer
Z-1544-2013 | 64216 | plugger, root canal, endodontic | EKR | Dental hand instrument. | Dental | Dental | 510(K) Exempt | SybronEndo Elements Obturation Unit; REF 973-0322; SN 91-1234; 2005-10; SybronEndo, 1332 S. Lone Hill Ave. Glendora, CA 91740; EU Rep: Kerr Italia, S.r.l, Via Passanti, 332, 1-84018 Scafati (SA), ItalyThe intended use of the Elements Obturation Unit and Replacement A/C Power Cord is to provide continuous heat at the tip of the dental instrument to test tooth response to thermal stimulus, for tissue cauterization and coagulation and to backfill and down pack gutta percha during endodontic root canal treatment. | SybronEndo Elements Obturation Unit; | 2 | June 14, 2013 | 2013 | Ormco/Sybronendo | A recall has been issued by SybronEndo because a Canadian Standards Association (CSA)-directed cautionary statement label was inadvertently omitted from the Elements Obturation Unit and its Replacement A/C Power Cord. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm SybronEndo sent a "FIELD CORRECTION ACTION" letter dated January 30, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were requested to review and determine if they have any of the affected equipment in their inventory; Note: the equipment itself has not been modified in any way, and therefore there is no change in the performance; customers should continue to use the equipment in the same manner they have in the past:; Note: SynbronEndo has enclosed the cautionary label along with instructions for proper placement of the affected equipment; the customers should apply the enclosed cautionary statement label to Elements Obturation Unit and/or Replacement AC Power Cord as per instructions and complete and return the FIELD CORRECTION ACKNOWLEDGEMENT FORM ,whether or not they have any affected equipment in their inventory, via Fax to: 909-962-5605.For questions regarding this recall call 714-516-7602.Note: If you are an authorized SybronEndo international distributor, we request that you identify those customers that may have been shipped the affected lots of equipment and contact these customers to inform them of the issue within forty-eight (48) hours of receipt of this notification.If you have any questions, contact SybronEndo Customer Care at 1-800-346-3636. | Contact the recalling firm for information | 3615 | 3615 units total | Worldwide Distribution - USA (nationwide) and the countries of Australia, Azerbaijan, Belgium, Canada, China, Cyprus, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Italy, Japan, Kuwait, Latvia, Lebanon, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Romania, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | November 07, 2013 | 146 | 3615 | Not_Computer
Z-1546-2013 | 64424 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Patient Helper Bed Frame Adapter AssemblyThe bed frame adapter assembly is intended to be used by qualified medical professionals to adapt a patient helper to a hospital bed. The patient helper is used to aid and assist the patient's mobility within a hospital bed and during bed ingress and egress. | Zimmer Patient Helper Bed Frame Adapter Assembly | 2 | June 14, 2013 | 2013 | Zimmer, Inc. | The bed frame adapter assemblies may have a weak weld where the patient helper adapter mounting bars attach to the tube. | DESIGN: Device Design | Zimmer sent an Urgent Device Removal letter dated February 14, 2013, to all affected consignees. The letter identified the product, the problem,and the action to be taken by the consignee. Consignees were advised to review the notification letter, follow the instructions contained in the letter, remove from use and quarantine the affected product. Consignees were also instructed to complete and return the fax back from to 877-787-0375 or zimmer3632@stericycle.com. Consignees were instructed to return the affected product back using the shipping label provided. If affected product was further distribution, consignees were asked to provide the customer's information to the firm on the back of the form. For questions regarding this recall call 866-201-9081. |  | 344 | 344 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, FL, HI, IL, KY, MA, MD, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, WA, and WI; and country of Canada. | Open |  |  | 344 | Not_Computer
Z-1547-2013 | 65210 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Trauma Phoenix Retrograde Femoral Connecting Bolt Trauma Fixation Systems.Product Usage:The Phoenix Retrograde Femoral Nail is indicated for alignment, stabilization, and fixation of fractures caused by trauma or disease and the fixation of femurs that have been surgically prepared (osteotomy) for correction of deformity. | Phoenix Retrograde Femoral Nail | 2 | June 15, 2013 | 2013 | Biomet, Inc. | Biomet Trauma ("Biomet") has initiated a recall of Retrograde Femoral Connecting Bolt, which involves Part Number 14-442021. The connecting bolt has an undersized diameter specification that may cause an interference fit with the 4mm hex driver (Part Number: 41024) near the edge of the tolerance. If the 4mm hex driver becomes stuck in the connecting bolt it cannot be taken apart and a delay in su | DESIGN: Device Design | Biomet sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 29, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to: immediately locate and remove the identified affected device(s), carefully follow the instructions on the enclosed "FAX Back Response Form", fax a copy of the Response Form to 574-372-1683 prior to return of product, use priority carrier for your shipment, if you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. For questions call 574-372-1570. |  | 75 | 75 | Worldwide Distribution - USA Nationwide: CA, TX, NY, PA, NJ, FL, OH, MI, CO, SD, IN, MT, GA, and WI. and the countries of Netherlands, Costa Rica, Japan, and San Juan. | Open |  |  | 75 | Not_Computer
Z-1549-2013 | 33728 | microscope, phase contrast | IBM | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Aperio ScanScope System. Aperio ScanScope Models, XT, AT, AT Turbo, CS and CSO.Automated digital slide creation and viewing device. It is intended to create and view electronic image files representative of tissue specimens fixed to glass slides. | Aperio ScanScope System | 2 | June 17, 2013 | 2013 | Aperio Technologies Inc | Aperio initiated this recall because the User Guides incorrectly state the devices had been cleared for specific applications. The User Guides should state the devices are for Research Use Only and are not for use in diagnostic procedures. | PREMARKET APPROVAL: No Marketing Application | Aperio sent an Urgent Medical Device Recall letter dated April 22, 2013 to all affected consignees via certified USPS mail. The letter included the reason for the recall, recall product information, instructions, acknowledgement form, contact information ( Aperio Technical Support at 1-866-478-3999), and product labeling. The consignees were requested to acknowledge receipt of the Medical Device Correction letter. In addition, an Acknowledgement Form is required to be signed and e-mailed back to the Manufacturer asevidence of letter receipt.For questions regarding this recall call 760-304-6211, ext 6211. |  | 364 | 364 | Nationwide Distribution - including all states, none internationally. | Terminated | September 25, 2013 | 100 | 364 | Not_Computer
Z-1548-2013 | 65212 | prosthesis, shoulder, semi-constrained, metal/polymer + additive, cemented | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Superior Approach Resection Guide Assembly.Indicated for primary, fracture, or revision total shoulder replacement for the relief of pain. | Superior Approach Resection Guide Assembly | 2 | June 17, 2013 | 2013 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the slot on the resection guide is offset in the wrong direction.When the slot on the resection guide is offset in the wrong direction, the Steinman pin that is intended to hold the guide into place may not come into contact with the bone. | DESIGN: Device Design | On 5/15/2013 an URGENT MEDICAL DEVICE RECALL notification was sent to all consignees detailing to immediately locate and discontinue use of the product and return to Biomet.The following actions are REQUIRED:-Immediately locate and remove the identified device(s) listed below from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .-Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product. you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. |  | 17 | 17 | Worldwide Distribution-USA (nationwide) including the states of UT, TX, AL, KY, and WI, and the countries of Australia, The Netherlands, Korea, and Canada. | Open |  |  | 17 | Not_Computer
Z-1550-2013 | 65245 | catheter, biliary, diagnostic | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | IDEV Technologies SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. Indicated for palliative treatment of biliary strictures produced by malignant neoplasms | SUPERA VERITAS | 2 | June 18, 2013 | 2013 | IDEV Technologies, Inc. | There was an error on the outer labeling of this device. The inner label was correct. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm decided to rework the products that were in their control at the time and the one that was not in their control was corrected by an IDEV employee. That hospital was called and notified of the situation in May 2012. |  | 3 | 3 | Nationwide Distribution to the states of FL, NY, and PA. | Terminated | July 29, 2013 | 41 | 3 | Not_Computer
Z-1556-2013 | 64874 | urinary homocystine (nonquantitative) test system | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products-HCY Reagent packaged in a kit carton. Catalog# 6802002. Used to quantitatively measure total homocysteine (HCY) concentration in human serum and plasma. | VITROS Clinical Chemistry Products Homocysteine (HCY) Reagent | 2 | June 19, 2013 | 2013 | Ortho-Clinical Diagnostics | The firm is recalling the specific lot due to lower than expected results. | PRODUCTION CONTROLS: Process Control | Ortho Clinical Diagnostics sent an "Important Product Correction Notification" letter dated April 1, 2013 to all affected customers. Ortho Clinical Diagnostics recommends that their customers immediately discontinue using and discard all VITROS HCY Reagent remaining in your inventory for the affected lot. They have requested that their customers complete and return the attached Confirmation of Receipt form no later than April 12, 2013. Ortho Clinical Diagnostics requested that these notifications be forwarded to all laboratories and posted beside each VITROS System that utilizes VITROS HCY Reagent within their customers facilities. To place a product order for an alternate lot, please contact Ortho Clinical Diagnostics customer service representatives at 1-800-828-6316. The VITROS Chemistry Products HCY Reagent is used on the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System to quantitatively measure total homocysteine concentration in human serum and plasma. | Contact the recalling firm for information | 478 | 478 total (339- Domestic) (139- International) | Worldwide Distribution-USA (nationwide) including the states of AR, AZ, CA, CO, CT, FL, HI, ID, IL, KY, LA, ME, NJ, OK, TN, TX, WA, and WV, and the countries of Colombia, India, Japan, Mexico, Puerto Rico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, and Spain. | Open |  |  | 478 | Not_Computer
Z-1553-2013 | 65263 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BC Hammock Mesh Sling (Large), product number, 537030.Product Usage:The fabric sling is used for the transfer of a patient, for example from a bed to a wheelchair. | BC Hammock Mesh Sling | 2 | June 19, 2013 | 2013 | Prism Medical Services USA | The product does not meet the firm's size standards, the product was produced in the incorrect size. | TRAINING: Employee Error | Prism Medical contacted their US consignees on May 28, 2013 with an e-mail notification of the recall. The notification explained the Field Safety Notice Letter dated May 22, 2013. The notice identified the affected product, problem and actions to be taken. Customers were instructed to contact Customer Service at Prism Medical to receive a return goods authorization to replace the affected product at no charge. For questions call 905-850-9003 Ext. 249. |  | 5 | 5 slings | Worldwide distribution: USA (nationwide) including states of: CA, IL, OH, and TX. and country of: Canada. | Terminated | November 25, 2013 | 159 | 21 | Not_Computer
Z-1557-2013 | 64830 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 11-Hole, 1/3 Tubular Plate DCL Plate with Collar, Part Number 241.000.011 | Synthes 11Hole, 1/3 Tubular Plate DCL Plate with Collar | 2 | June 19, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes One-Third Tubular DCL Plate (11-Hole, 1/3 Tubular Plate w/Collar), was released to the United States market in error. This action is not being initiated as a result of adverse events. |  | The firm initiated their recall of the product on February 22, 2013 by sending an "Urgent Notice: Medical Device Recall" letter to the firm's sales representatives. The letter requested they examine their inventory, and/or customers inventories for the Part Number and if found, return it to Synthes for appropriate disposition. |  | 10 | 10 | Nationwide distribution. | Open |  |  | 10 | Not_Computer
Z-1558-2013 | 65129 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope (GVL) 3, 4, 5 and GlideScope (AVL) 2, 3, 4 and 5 reusable bladesProduct Usage: The GlideScope Video Laryngoscopes are intended for use by qualified medical professionals to obtain a clear, unobstructed view of the vocal cords for medical procedures. The GlideScope Video Laryngoscopes system should only be used by individuals who have been trained and authorized by a physician or the institution providing patient care. The device is not life sustaining or life supporting. The GlideScope AVL Video Monitor system has the same intended use as the GVL System. The difference between systems is that the AVL has a digital platform and the GVL an analog platform. | The GlideScope Video Laryngoscope GVL 3, 4, 5 and AVL 2, 3, 4 and 5 | 1 | June 20, 2013 | 2013 | Verathon, Inc. | Verathon GlideScope Reusable GVL and AVL Blades are being recalled due to the potential risk of premature failure/breakage of the blade tip that may not be readily visible during routine inspection before or after intubation. | DESIGN: Device Design | Verathon issued an Urgent - Safety Alert and Medical Device Recall letter, dated May 10, 2013, to all affected customers The letter identified the affected product, problem and actions to be taken. The letter advises customers to discontinue the use of the specific serial numbers listed in the letter and asked that the affected product be returned to Verathon for replacement. The customers are also directed to fill out a return response form (included in the letter) regarding this recall.Consignees with questions are directed to contact their Verathon representative or Verathon Customer Care at 800-331-2313 (US and Canada) and 1-425-867-1348 for those consignees outside of US and Canada. the support team will be available Monday through Friday from 6 AM - 5 PM PST for the US and from 6 AM to 4:30 pm PST outside the US. Alternately they can e-mail Verathon at customerservice@verathon.com. |  | 4784 | 4784 units worldwide (3,381 in the US and 1403 international) | Worldwide distribution: USA (nationwide) including states of : AK, AL, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Washington DC, Guam, and countries of : Argentina, Australia, Brazil, Canada, China, Ecuador, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Japan, Malaysia, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, The Netherlands, United Kingdom, and Venezuela. | Terminated | December 17, 2013 | 180 | 4784 | Not_Computer
Z-1561-2013 | 64856 | regulator, pressure, gas cylinder | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | MEDICYCL-E- Lite Portable Oxygen SystemAn integrated portable oxygen delivery system intended to provide supplemental oxygen to adults. The device is MRI safe, MRI-compatible, and intended for use during MR Imaging for MRI systems up to 3.0 T. For emergency use only when administered by properly trained personnel for oxygen deficiency and resuscitation. For all other medical applications, Rx only. | MEDICYCLE Lite Portable Oxygen System | 2 | June 20, 2013 | 2013 | Linde Gas North America Llc | A number of MEDICYL-e Portable Oxygen Systems may have misapplied valve flow knobs. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, The Linde Group, sent a "Voluntary Medical Device Recall" letter dated June 21, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to view the enlarged photo on the reverse side of the notice for guidance. If you or your customers have a Unit with lot number 25/07, then please return it with the Product to your local LifeGas fill center as soon as possible. No other lot numbers are subject to this removal action.Customers with questions should contact Linde's Customer Care Center at 1-866-LIFEGAS (543-3427). |  | 52264 | 52,264 units (1106 units) | Nationwide distribution. | Open |  |  | 52264 | Not_Computer
Z-1572-2013 | 65417 | cannula, injection | FGY | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | MAQUET GETINGE GROUP Arterial Cannula OD Used to introduce blood and sterile solutions into the systemic circulation during pediatric extracorporeal circulation procedures. | Maquet | 2 | June 21, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The specified part numbers and lot numbers of Maquet Pediatric Arterial Cannula Packs may contain an incorrectly sized connector. | PRODUCTION CONTROLS: Reprocessing Controls | Maquet sent an "Urgent Product Recall"/"Recall - Device Corrective Action Response" forms dated May 14, 2013, via Federal Express to all affected customers. The notification identifies affected product; informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; and advises customers to follow instructions on the "Recall-Device Corrective Action Response" forms included as part of the notification.If you have any further questions please contact your local MAQUET representative or call Customer Service at (888) 627-8383. |  | 100 | 100 units (60 Domestic; 40 International) | Worldwide Distribution - US Distribution including the states of Delaware, Illinois, Minnesota , New York., Ohio and Pennsylvania and the countries of Germany, Turkey and France. | Open |  |  | 100 | Not_Computer
Z-1569-2013 | 64530 | controller, temperature, cardiopulmonary bypass | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sarns " TCM II Temperature Control ModuleThe Sarns" TCM II Temperature Control Module is a source of temperature-controlled water for blood heat exchangers used in an extracorporeal circuit and is indicated for controlling and monitoring patient temperature | Sarns " TCM II Temperature Control Module | 2 | June 21, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Internal testing found certain Sarns" TCM II systems were released with incorrectly assembled heatsink assemblies which could cause the units to fail. | TRAINING: Employee Error | Terumo sent beginning May 30, 2013, an URGENT MEDICAL DEVICE RECALL NOTIFICATION were sent to all customers with instructions to notify all customers of the recall requesting all affected units to be quarantined until the heatsink units are replaced Confirm receipt by filling out the response form containing communication by faxing or emailing the attached Affiliate Response Form to the fax number or email address indicated on the form.Questions should be directed to Terumo CVS customer service in the U.S.:Telephone: +1-734-663-4145; Fax: +1-734-741-6149. | Contact the recalling firm for information | 7 | 7 units | Worldwide Distribution including the countries of UAE (United Arab Emirates) HONG KONG and THAILAND. Not distributed in USA | Terminated | September 17, 2013 | 88 | 7 | Not_Computer
Z-1577-2013 | 65243 | syringe, antistick | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD SafetyGlide Allergy. REF 305950 Sterile Allergy Treatment Syringe Tray. The device is used for aspiration and injection of medications. | BD SafetyGlide Allergy Syringe Tray | 2 | June 21, 2013 | 2013 | Becton Dickinson & Company | There has been an increase in complaints regarding the needle/safety mechanism disengaging from the syringe. |  | Direct consignees were contacted by phone and email on 05/01/2013 and instructed to return the affected product. End users were mailed "Urgent: Product Recall" letters/Response Forms dated 05/01/2013. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product. | Contact the recalling firm for information | 1326000 | 1,326,000 units | Worldwide Distribution including US Nationwide and the countries of Canada and Suriname. | Terminated | September 25, 2013 | 96 | 1326000 | Not_Computer
Z-1568-2013 | 65362 | prosthesis, wrist, carpal trapezium | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | CMC SaddleProduct Usage: The Ascension CMC is intended to resurface the proximal end of the first metacarpal in cases of rheumatoid arthritis, traumatic arthritis, osteo-arthritis, or post-fracture deformation or bone loss which present as either a painful, unstable thumb, or a thumb with limited range of motion. This is an uncemented prosthesis designed for press fit use only. | Ascension Orthopedics Carpal/Metacarpal (CMC) Saddle | 3 | June 21, 2013 | 2013 | Integra LifeSciences Corp. | As a result of a quality review, labelling content discrepancies were identified between the Instruction for Use (IFU) distributed in the US with the Ascension CMC and the content that is cleared by FDA. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Integra sent an Urgent: Medical Device Recall letter to all affected product. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, discontinue use, and complete and return the attached form. When the form is received, Customer Service Returns and Repairs will send an RMA number to return the affected inventory. For quwstions contact your regional managers or call 512-520-7697. | Contact the recalling firm for information | 467 | 467 devices | USA Nationwide Distribution | Open |  |  | 467 | Not_Computer
Z-1564-2013 | 65458 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Launcher Guiding Catheter, Model # LA8EBU45SH, Lot # 0006722866, 8F and Model # LA6IMAD, Lot # 0006724224, 6FProduct Usage: The Medtronic Guiding Catheter is designed to provide a pathway through which therapeutic devices are introduced. The guiding catheter is intended to be used in the coronary or peripheral vascular system. | Launcher Coronary Guiding Catheter | 2 | June 21, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified a labeling issue for 24 units of the Launcher coronary Guide Catheter. The two affected catheter lots are mislabeled and do not accurately reflect the catheters contained within each package. | PRODUCTION CONTROLS: Labeling Mix-Ups | Medtronic "Urgent Field Safety Notice" letter dated June 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter instructed customers to remove and quarantine all affected product within their inventory and return the affected product to Medtronic. For questions contact your local Medtronic representative. |  | 15 | 15 | Internationally distributed: to countries of Germany and Poland. | Terminated | November 19, 2013 | 151 | 15 | Not_Computer
Z-1565-2013 | 65254 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P, Baxter Healthcare Corporation, Deerfield, IL 30015 USA.Intended to isolate the Easy-Lock connector of the solution transfer set. | MiniCap with ProvidoneIodine solution | 2 | June 21, 2013 | 2013 | Baxter Healthcare Corp. | leaking pouches | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | The firm, Baxter Healthcare Corporation sent an "URGENT PRODUCT RECALL" letter dated May 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For general questions regarding this communication, contact The Center for One Baxter at 1-800-422-9837. | Contact the recalling firm for information | 629280 | 629,280 units | Nationwide Distribution. | Open |  |  | 629280 | Not_Computer
Z-1566-2013 | 65343 | balloon aortic valvuloplasty | OZT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | TRUE Dilatation Balloon Valvuloplasty Catheter, REF Numbers: 02045-11, 02245-12, 02445-12, and 02645-13; Sizes: 20mm x 4.5cm, 22mm x 4.5cm, 24mm x 4.5cm, and 26mm x 4.5cm; Intended to be used for balloon aortic valvuloplasty. | TRUE Dilatation Balloon Valvuloplasty Catheter | 2 | June 21, 2013 | 2013 | Loma Vista Medical | Slow deflation time: Slower than normal balloon deflation time of greater than 10 seconds. The balloon normally deflates in less than 5 seconds. The device performance specification for deflation time is 10 seconds maximum. | DESIGN: Process Design | The firm sent an "URGENT VOLUNTARY DEVICE REMOVAL" letter dated March 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 2379 | 2379 units | Worldwide Distribution-US Distribution in New Jersey and the countries of Germany, Switzerland, Italy, Norway, The Netherlands, Spain, Portugal, Ireland, and Austria. | Terminated | July 17, 2013 | 26 | 2379 | Not_Computer
Z-1578-2013 | 65193 | lift, patient, ac-powered | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-Series Carry Bar, Class I, model number 360755. The intended use is a carry bar attached to a C-450 or C-625 ceiling lift used to lift, transfer or position clients. | CSeries Carry Bar | 2 | June 21, 2013 | 2013 | Prism Medical Services USA | The firm replaced the recalled carry bar with a plastic insert with a carry bar with a metal insert due to improper installation of the recalled carry bars. |  | The firm, Prism Medical, sent a letter dated January 27, 2013 to is customers.The letter described the product, problem and actions to be taken. The customers were instructed to perform an immediate safety check checking that both the pin and the plug have been inserted into the carry bar as outlined in the attached User Instruction guide. Prism offered to replace the recalled product with a different product with greater functionality.If you have any questions, contact VP-Global Product Development at 416-726-3107. |  | 1131 | 1131 carry bars/New England | Worldwide distribution: US (nationwide) and country of: Canada. | Open |  |  | 1131 | Not_Computer
Z-1574-2013 | 65342 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | Medtronic Intrathecal Catheter Pump Segment Revision Kit, model 8596SC. Contents: 60-cm pump segment with attached sutureless pump connector, Spinal Segment Strain-relief sleeves, Pump segment strain-relief sleeves, Connector pin. Contents of inner package are STERILE.The Medtronic Model 8596SC Pump Segment Revision Kit is used when a revision to the pump segment of the Model 8731 or Model 8731SC catheter is required. The catheter is part of an infusion system that stores and delivers parenteral drugs to the intrathecal space. The implanted infusion system components consist of a Medtronic pump and a Model 8731 or Model 8731SC catheter. The catheter connects to the pump at the catheter port. | Medtronic Intrathecal Catheter Pump Segment Revision Kit | 1 | June 24, 2013 | 2013 | Medtronic Neuromodulation | The Sutureless Connector (SC) Intrathecal Catheter connector has been redesigned to reduce the potential for occlusion at the catheter to pump interface. Medtronic is removing the unused products from the market that were manufactured with the previous design, and recommend the previous design no longer be used due to greater potential for misalignment and subsequent occlusion. | DESIGN: Device Design | The firm, Medtronic, sent an "Urgent: Medical Device Removal" letter dated May 2013 to its customers. The letter described the product, problem and actions to be taken. Representatives (Rep) are visiting all locations, beginning June 3, 2013, to retrieve devices with a Use By date of 2014 08 14 (August 14, 2014) or sooner. The Rep is leaving a letter with the hospitals to tell them that the Sutureless Connector Intrathecal Catheter connector has been redesigned and that they are removing unused devices. They also do not recommend using any of the old design. An Account Specific Customer Confirmation Form will be completed by the Rep. A copy of the completed form will be left with the Hospital along with the Urgent Medical Device Removal letter.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 115722 | 115,722 total | Worldwide distribution: US (nationwide) and countries of: Aruba, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Open |  |  | 115722 | Not_Computer
Z-1587-2013 | 65371 | automated breast ultrasound | PAA |  |  | Radiology |  | Soma v Automated Breast Ultrasound System (ABUS), Automated Breast Ultrasound.Indicated as an adjunct to mammography for breast cancer screening. | Soma v Automated Breast Ultrasound System | 2 | June 24, 2013 | 2013 | U-systems Inc | Devices labeled for an intended use not included in the existing 510(k) or PMA, in that, the user manual contains the Indication for Use of the soma v Platinum ABUS [a 510(k) product]. | DESIGN: Software Design | On 6/5/13, an "URGENT MEDICAL DEVICE CORRECTION" letter was distributed to all affected users, with a description of the problem and user corrective action steps. One consignee received an Urgent Notice to Customer - Medical Device Correction letter on 5/16/13. Contact the firm at 408-328-7925 for questions regarding this notice. |  | 3 | 3 units | Distributed in California and Illinois. | Terminated | November 04, 2013 | 133 | 3 | Not_Computer
Z-1586-2013 | 65411 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DC power cable (p/n 3010556) supplied with the MEDRAD(R) Veris MR Vital Signs Monitor: Product Usage:Monitor, Physiological, Patient | DC power cable (p/n 3010556) supplied with the MEDRAD(R) Veris MR Vital Signs Monitor | 2 | June 25, 2013 | 2013 | Medrad Inc | The firm is recalling these power cables due to a latent design reliability issue and the potential for shorting which can result in heating/melting of the cable jacket. | DESIGN: Device Design | Medrad Inc sent an Urgent Medical Device Component Recall letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement products will be available in June and shipped in the coming months along with additional instructions for returning the affected product. Customers were instructed to complete the enclosed recall response form and fax it back to Bayer at 412-406-0942. For questions contact customer support team at 1-877-229-3767 and select option 2. |  | 1066 | 1,066 | Worldwide Distribution - US Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: AE, AU, AZ, BN, BR, CA, CH, CL, CO, DE, DK, DO, DZ, EG, ES, FR, GB, HK, HN, HR, HU, IN, IT, JO, JP, KR, LB, MU, MX, MY, NL, PH, PS, SA, SG, TH, TR, TW, UY, VE, and ZA. | Open |  |  | 1066 | Not_Computer
Z-1593-2013 | 65297 | filler, recombinant human bone morphogenetic protein, collagen scaffold with metal prosthesis, osteoinduction | NEK |  |  | Orthopedic |  | INFUSE(R) Bone Graft, SMALL KIT, REF 7510200AUS, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132orthopaedic | INFUSE(R) Bone Graft | 2 | June 25, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | Kits contain a collagen sponge recalled by Integra due to possible elevated endotoxin levels resulting from a process deviation. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic on behalf of Integra LifeSciences Corporation sent an Urgent Voluntary Removal (Recall) letter on May 24, 2013, directly to all Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. Foreign consignees were notified by email on May 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product and discontinue use. The letter should be forwarded to surgeons or attending physicians who have used the affected products. Customers were instructed to use the attached questionnaire to account for product in their possession, document which physicians they notified and report any patient injuries that may have occurred related to the use of the affected product by calling the phone number in the questionnaire. Customers were asked to return the questionnaire using the enclosed self-addressed stamped envelope, via their Medtronic Sales Representative or via fax to 763-367-8134For questions regarding this recall 901-399-3182.Customers were instructed to contact their Medtronic Sales Representative to coordinate the return and replenishment of affected product . Customers with questions should call 800-876-3133, ext 3197. |  | 1319 | 1319 units | Worldwide distribution: USA (nationwide including Puerto Rico) and Internationally to Australia, Brazil, Canada, Chile, Hong Kong, India, Mexico, Netherlands, , New South Wales, Panama, and Paraguay. | Open |  |  | 57508 | Not_Computer
Z-1596-2013 | 65138 | sterilant, medical devices | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connect QLC1677E, QRC1699E, QPC1721E and QFC1683E. STERIS Corporation, Mentor, OH 44060QLC1677E, QRC1699E, QPC1721E and QFC1683E contain processing instructions, a Quick Reference Wall Chart and a flow unit. The flow unit consists of a Quick Reference Card, tubing, tethers and adapters for connecting specified devices to Flexible Processing Trays/Containers. | Quick Connect | 2 | June 26, 2013 | 2013 | Steris Corporation | During a routine label review, the firm identified four SYSTEM 1E Quick Connects that will be updated to reflect the removal of obsolete device models. This labeling review process also highlighted the opportunity to combine two existing quick connects into a single product while eliminating the obsolete models. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | STERIS sent a Customer Notification letter dated February 19, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to replace their existing labels with labels they provided in their customer notification package.We apologize for any inconvenience this may have caused you and as always STERIS Corporation is dedicated to support our products and valued Customers. If you have any questions regarding this matter, please contact your STERIS Capital Account Manager. For further questions please call (440)392-7348. | Contact the recalling firm for information | 128 | QLC1677E: 38; QRC1699E: 128 & QPC1721E/QFC1683E: 0 | Nationwide Distribution including the states of AK, AZ, CA, CO, FL, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, NC, NE, NH, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI and WY. | Terminated | July 02, 2013 | 6 | 166 | Not_Computer
Z-1614-2013 | 65247 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products K+ (Potassium) Slides (Reagent), REF/Catalog Number 815 7596, Product Usage:For in vitro diagnostic use only. VITROS Chemistry Products K+ Slides quantitatively measure potassium (K+) concentration in serum, plasma, and urine using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. | VITROS Chemistry Products | 3 | June 27, 2013 | 2013 | Ortho-Clinical Diagnostics | While performing interval testing, Ortho Clinical Diagnostics (OCD) determined that a surfactant was unintentionally added to VITROS Chemistry Products K+ Slides during the manufacturing process. | TRAINING: Employee Error | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated May 9, 2013 via FedEx or US Priority Mail to all US customers. Foreign Affiliates were notified by e-mail on May 9, 2013 informing them of the issue. Distributors were instructed to notify their customers of the recall. Customers were instructed to discontinue use, discard the affected product, comp[lete and return the attached Confirmation of Recipt Form, and post this notification by each VITROS System. For questions call Customer Service Center at 1-800-828-6316. |  | 244703 | Domestic: 236,364; Foreign: 8,339 | Worldwide Disribution - USA Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Open |  |  | 245626 | Not_Computer
Z-1608-2013 | 65464 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis a Johnson & Johnson Company, RadialSource Spring Wire Transradial Access Kit. Manufacturer Greatbatch Medical 2300 Berkshire Lane North, Minneapolis, MN 55441, Distributor Cordis Corporation 14201 North West 60th Avenue, Miami Lakes, Florida 33014.Baxter's lnfusors SV Elastomeric Infusion Device is a single-use, disposable elastomeric infusion pump indicated tor patients requiring slow, continuous intravenous, intra-arterial, subcutaneous, or epidural administration of medications at a constant flow rate. The device is also indicated for the administration of bolus doses of medication upon patient demand when used in conjunction with the Patient Control Module. lt is suitable for use in the hospital or home setting. | Cordis Radial Source Spring Wire Transradial Access Kit Product | 2 | June 27, 2013 | 2013 | Greatbatch Medical | Cordis Corporation, Inc and Greatbatch Medical are recalling one lot of the cordis RADIALSOURCE Spring Wire Transradial Access Kit product. Some units were manufactured with an undersized vessel dilator tip, too small to slide over the guidewire. | PRODUCTION CONTROLS: Process Control | Greatbatch Medical and Cordis sent an "Urgent Field Safety Notice" dated May 31, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer. Customers were instructed:1.Read the Description of the problem section.2. Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used.3. Review, complete, sign and return the enclosed Acknowledgement Form inaccordance with the directions on the form.4. Either return any affected product per the attached instruction, or contact yourlocal sales representative to facilitate return of the affected product. Credit willbe provided.5. Pass on this notice to anyone in your facility that needs to be informed.6. If any product listed above has been forwarded to another facility, contact thatfacility to arrange return.7. Maintain awareness of this notice until all product listed above has beenremoved from the facility and returned to Cordis.8. Maintain a copy of this notice with affected product until product is returned.We apologize for any inconvenience this communication may cause. We knowthat you place high value in our products and we appreciate your cooperation inthis matter. Cordis and Greatbatch are committed to maintaining your confidence in the safety and quality of the products that Greatbatch manufactures and Cordis supplies. Further questions please call (763) 951-8181. |  | 550 | 550 kits (OUS) | Worldwide Distribution - AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, SPAIN, GREECE, HUNGARY, IRELAND, ISRAEL and ITALY. | Open |  |  | 550 | Not_Computer
Z-1612-2013 | 65397 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Terumo Glidesheath (tm), 6 French, Nitinol wire, product code 15-1060.Used to facilitate placing a catheter through the skin into a vein or artery. | Terumo Glidesheath | 2 | June 27, 2013 | 2013 | Terumo Medical Corporation | Medical devices were incorrectly labeled with extended expiration dates. | PRODUCTION CONTROLS: Error in Labeling | The recalling firm notified consignees by letter on 5/30/2013. The letter was flagged as an urgent product bulletin addressed to risk managers. The notification advised customers to immediately discontinue use and quarantine any available stock for subsequent instructions to be provided for pick-up and replacement. | Contact the recalling firm for information | 550 | 550 units | Nationwide Distribution. | Terminated | October 30, 2013 | 125 | 570 | Not_Computer
Z-1600-2013 | 65365 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Infusor SV and LV Elastomeric Infusion Devices.Indicated for the intravenous administration of medications. | Infusion Pump Infusor SV and LV Elastomeric Infusion Devices | 2 | June 27, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation (Baxter) is issuing an Urgent Product Recall for the products listed above due to an increase in rupture complaints. Baxter has implemented product and process improvements which have substantially improved product quality. Complaint rates for ruptures have since decreased. Baxter wants to ensure that no affected product remains in the field. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An Urgent Product Recall notification was sent to all affected consignees on 06/13/2013. The letter states the affected products, problem description, hazard involved and actions to be taken by customer/user, with contact information. The letter was accompanied by a Customer Reply Form and a table of Affected Product List. The recall notification instructs customers to segregate affected product and contact Baxter for instructions for return of product. | Contact the recalling firm for information | 3255096 | 3,255,096 total | Distributed Nationwide and in Puerto Rico. | Open |  |  | 3255096 | Not_Computer
Z-1609-2013 | 65364 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Coiled Tube Infusors (Portable Elastomeric Infusion Systems). Baxter's Infusor SV Elastomeric Infusion Device Device Listing: D020871Baxter's lnfusors SV Elastomeric Infusion Device is a single-use, disposable elastomeric infusion pump indicated tor patients requiring slow, continuous intravenous, intra-arterial, subcutaneous, or epidural administration of medications at a constant flow rate. | Infusion Pump | 2 | June 27, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation (Baxter) is issuing an Urgent Product Recall for the products listed above due to an increase in complaints for leaks at the distal male Luer and Luer cap. Baxter implemented enhancements to the blue winged cap to improve the seal between the cap and the Luer. Complaint rates for leaks have since decreased. Baxter wants to ensure that no affected product remains in t | DESIGN: Device Design | Baxter sent an Urgent Produd Recall notification letter dated June 13, 2013, to all affected consignees via USPS.The letter was accompanied by a Customer Reply Form and Affected Product List (table). The letter stated affected products, problem description, risk to health (or health risks), actions to be taken by customers/users and contact information.The recall notification instructs customers to segregate affected product from current inventory and contact Baxter for instructions to return affected product. Customers with questions were instructed to call Baxter at 1-800-437-5176,For questions regarding this recall call 847-270-5556. | Contact the recalling firm for information | 601230 | 601,230 total | Worldwide Distribution - USA (nationwide), Puerto Rico, and Internationally to Brazil, Colombia, Panama, Mexico, Singapore Japan, Australia, New Zealand, Hong Kong, India, Thailand, Taiwan, and Korea. | Open |  |  | 601230 | Not_Computer
Z-1613-2013 | 65377 | acid, lactic, enzymatic method | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON System(s) Lactate (LACT) Reagent.Product Usage: Lactate (LACT) reagent, when used in conjunction with Synchron LXÂŽ System(s), UniCelÂŽ DxC 600/800 System(s) and SynchronÂŽ Systems Multi Calibrator, is intended for the quantitative determination of lactate concentration in human plasma and cerebrospinal fluid (CSF). | Synchron System(s) Lactate Reagent | 2 | June 27, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has confirmed customer reports of LACT Lot M210077 failing calibration witherror condition OCR LOW. The calibration failures appear to be related to improper shipping orstorage conditions (frozen reagent). Failed calibration of LACT may cause a delay in reporting results. |  | Beckman Coulter Inc. sent an Urgent Product Correction letter dated May 28, 2013 to affected customers. The letter identified the affected product, the issue, impact, actions to be taken, and resolution. The letter instructed customers to discontinue use of the affected product and request a replacement product. For questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call 800-854-3633 in the United States and Canada. Outside the United States or Canada, contact your local Beckman Coulter Representative. |  | 3353 | 3,353 units | Worldwide Distribution - USA Nationwide and country of Canada. | Open |  |  | 3353 | Not_Computer
Z-1607-2013 | 65433 | pulse generator, pacemaker, implantable, with cardiac resynchronization (crt-p) | NKE |  |  | Cardiovascular |  | Medtronic, ConsultaÂŽ CRT-P, Model C4TR01, C3TR01 (not approved for US), and SyncraÂŽ CRT-P, Model C2TR01, Sterile EO. Dual chamber and atrial tracking modes are indicated for patients who may benefit from maintenance of AV synchrony. | Medtronic, Consulta CRTP, Syncra CRTP | 2 | June 27, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic recently identified an issue with a subset of ConsultaÂż CRT-P and SyncraÂż CRT-P devices during production. As of May 30, 2013, there have been no reported or confirmed device failures. However, because of the potential for malfunction, Medtronic is requiring the return of non-implanted devices manufactured between April 1 and May 13, 2013 for re-inspection. The issue is unique to specifi |  | Medtronic sent an Urgent Medical Device Recall letter dated June 7, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately return non-implanted devices to Medtronic and to register the implanted ones immediately by calling the Medtronic Patient Registration Department at 1-800-848-9300. Customers with questions were instructed to contact their local Medtronic representative or Medtronic Customer Service at 1-800-848-9300.For questions regarding this recall call 763-526-6000. UPDATE: An Urgent Medical Device Correction letter, dated June 2013, was sent to affected implanting and following physicians beginning June 28, 2013. In the US these letters were hand delivered by a Medtronic representative. Outside the US, the notice will be mailed or hand delivered as determined by the geographic center. The letter informed physicians that they have implanted or are following a patient implanted with a device potentially affected by the connector bracket weld issue. The letter described how a loss of hermeticity may affect the device functionality and what symptoms a patient may experience. The letter also offers patient management recommendations. | Contact the recalling firm for information | 1912 | 1912 (159 USA & 1137 OUS) UPDATED: 1128 affected | Worldwide Distribution - USA including CA, FL, GA, KY, LA, LMN, MO, NE, NY, OK, PA, TN, TX, WI, and PR and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, Ireland, Italy, Japan, Latvia, Lebanon, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia ,United Arab Emirates, United Kingdom, Uruguay , and Yemen. UPDATED Physicians location in USA: AL, AR, FL, GA, IL, IN, KS. LA, ME, MI, MN, MO, NY, OH, PA, SC, TN, TX and VA. OUS: Luxembourg | Open |  |  | 1912 | Not_Computer
Z-1617-2013 | 65506 | catheter, electrode recording, or probe, electrode recording | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole LASSO Nav Catheter Eco, Catalog # D-1349-02S.Designed to facilitate electrophysiological mapping of the atria of the heart. | 20Pole LASSO Nav Catheter Eco | 2 | June 28, 2013 | 2013 | Biosense Webster, Inc. | Biosense Webster is recalling the 20-Pole LASSO NAV Catheter Eco because it has been incorrectly calibrated causing them to be recognized and displayed as a 10-Pole LASSO NAV Catheter Eco by the CARTO 3 System. |  | An Urgent Field Notice letter was sent to all their customers who purchased the 20-Pole LASSO NAV Catheter Eco. The letter informed the customers of the problem identified and the actions to be taken. Customers are instructed to complete and return the attached Voluntary Field Removal Certification Form in accordance with the instructions listed on the form. Customers are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday-Friday from 7:00AM to 8:00 PM EST if they have any questions.For questions related to the Voluntary Field Removal Certification Form and its return, customers are instructed to contact the Field Action Coordinator, at (909) 839-8468. |  | 10 | 10 units | Distributed in the states of MN, FL, and UT. | Terminated | February 07, 2014 | 224 | 10 | Not_Computer
Z-1618-2013 | 65202 | device, intravascular catheter securement | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StatLock CV Plus w/Pigtail, Catheter Stabilization Device Sterile EO Single Use Only, C.R. Bard, Inc. The StatLockÂŽ device provides stabilization for compatible medical tubes and catheters. | StatLock | 2 | June 28, 2013 | 2013 | Bard Access Systems | BARD Access Systems is recalling lot JUWJF283 of StatLock CV Plus Stabilization Butterfly Devices due to incorrect pad assembly used during manufacture. The absence of the hydrocolloid strip on the bottom may pose an incremental risk including localized skin irritation. Missing directional arrow may lead to mis-orientation of the StatLock stabilization which may impact the securement of the cathet | PRODUCTION CONTROLS: Packaging | On 6/18/13, an "Urgent Kit Recall" notification letter was mailed to all customers informing them of the recall and instructing the consignee to remove and return any affected product lot numbers to Bard Access Systems through return goods authorization. Customers may contact 800-290-1689 if questions. | Contact the recalling firm for information | 3875 | 3,875 units. | National and Global distribution (Europe, European Economic Area, Argentina, China, and Brazil). | Open |  |  | 3875 | Not_Computer
Z-1622-2013 | 65159 | table, examination, medical, powered | LGX | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General Hospital | 510(K) Exempt | MAQUET Getinge Group Mobile Operating Table YUNO OTN GA 1433.01 EN 05.The YUNO OTN is used to support and position patients for surgical procedures immediately before, during and after the surgical phase. The table top is radio translucent and enables intraoperative use of X-ray equipment. | MAQUET YUNO OTN Mobile Operating Table and Extension Device | 2 | July 01, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue with the extension device handle (part number 1433.62A1) and will make enhancements to the YUNO table (part number 1433.01F0).Identification of the Extension device issue: By fully actuating the lever of the extension device handle, the operator's skin can enter a gap between the handle lever and a recess in the handle bar. When the lever is released, | DESIGN: Device Design | MAQUET Getinge Group issued an Urgent Medical Device Recall letter/"Field Action Response Form dated April 26, 2013, via Federal Express to all affected customers. The recall notification identifies the product, issue with the product, and measures to be taken by the user. Customers were instructed that the firm would contact them to arrange for a loaner YUNO table and extension device while the table and extension device are upgraded. Upon completion of the upgrades, MAQUET would reinstall their upgraded YUNO table and extension device and remove the loaner device. Customers were instructed to complete and return the enclosed Field Action Response Form, acknowledging their receipt and understanding of the communication by e-mailing a scanned copy to FieldAction@maquet.com or via fax at 973-807-1658. Customers with questions were instructed to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3, followed by option 1.For questions regarding this recall call 973-709-7691. |  | 37 | 13 Tables; 24 Extension devices | Worldwide Distribution - USA including Arizona, Georgia, North Dakota, South Carolina, Texas and Tennessee, Internationally to United Arab Emirates, Austria, Brazil, Canada, Switzerland, China, Germany, Spain, France, United Kingdom, Italy, Japan, Singapore, Thailand, Turkey, Serbia, and South Africa. | Open |  |  | 37 | Not_Computer
Z-1623-2013 | 65205 | knife, surgical | EMF | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | -Instratek Carpal Tunnel Release Blades, (1) Triangle Blade, (1) Hook Blade. REF R2056 - LOT 122279-Instratek Carpal Tunnel Release Hook Blade. REF 2055 Lot 122669 -Linvatec CTS Relief Kit. REF 9971, lot 122504-Biomet Indiana Tome Blade for Carpal Tunnel Release (White). REF 200060. Lot 122504 and Lot 123190.-KMI SafeGuard Knife, For SafeGuard Mini Carpal Tunnel Release System, (White). REF 08-0003. Lot # 122609Surgical Knife | Instratek Carpal Tunnel Release Blades | 2 | July 01, 2013 | 2013 | Surgical Instrument Service And Savings, Inc. | The seal which maintains a sterile barrier for reprocessed medical devices became compromised to the point where product may fall out of the pouch. | PRODUCTION CONTROLS: Packaging Process Control | MEDISISS sent a Correction Immediate Action Request letter dated May 2, 2013, to all affected customes. The letter identified the product, the problem, and the action to be taken by the customer. The letter directed the customer to inspect the products prior to use for seal integrity as specified on the product labeling: "Sterile unless package has been damaged or opened. Do not use if package is damaged."This will be followed up by a visit from the customer's sales representatives who are tasked with inspecting all the lots that may be remaining in the customer's inventory. The sales representative is directed to complete a "Verification of Seal Inspection" form and return it to MEDISISS Customer Relations. Devices with unacceptable seals are to be returned to MEDISISS. The sales representative is responsible for obtaining a shipping label from customer service to enable the devices to be returned to the recalling firm via Fed Ex.Customers with questions were instructed to call 866-866-7477. |  | 56 | 56 units | Nationwide Distribution. | Terminated | July 11, 2013 | 10 | 2270 | Not_Computer
Z-1619-2013 | 65166 | fastener, fixation, biodegradable, soft tissue | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | GENESYS Cross FT Suture Anchor with Two #2 (5 metric) Hi-FI Registered Trade Mark Sutures Rx Only STERILE EO CONMED LINVATEC REF CFBC-4502. Reattachment of soft tissue to bone in orthopedic surgical procedures | GENESYS Cross FT Suture Anchor with Two 2 (5 metric) HIFI Sutures | 2 | July 01, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Incorrect blue/white suture was used. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery on March 6, 2013, to notify all direct consignees of the recall. If the account no longer had the product in possess, they were asked to respond via a reply form or to contact ConMed Linvatec Customer Service to obtain a Service Request number for return and credit of unused unopened product. All returned product for this action would be identified, segregated and quarantined per the firm's procedures. Any returned product not part of the recall will be dispositioned in accordance with the firm's procedures.Please accept our sincere apology for any inconvenience this action may have caused . Should you have any questions please contact our Customer Service department at (800) 237-0169. |  | 60 | 60 each of both units | US Distribution including the states of FL, KY, LA, MI, NC, and NY. | Terminated | July 22, 2013 | 21 | 60 | Not_Computer
Z-1620-2013 | 65301 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Bigliani/FlatowÂŽ The Complete Solution Fukuda Retractors Large Nonsterile 4305-19-40. Used in retracting the humeral shaft posteriorly and exposing the entire glenoid surface. | Bigliani/Flatow | 2 | July 01, 2013 | 2013 | Zimmer, Inc. | This is a recall expansion from the 2010 recall which now includes seven additional lots that were not previously included. Zimmer, Inc. is initiating a recall of small and large Fukuda-type retractors due to fractures occurring at the grooves around the perimeter of the blades while being used during surgery. | DESIGN: Device Design | In 2010, Zimmer initiated a recall of the Bigliani/FlatowÂŽ Fukuda Retractors due to fractures occurring at the grooves around the perimeter of the blades. Zimmer is expanding this recall to include seven lots that were not previously included. On 5/23/13, an Urgent Medical Device Recall notification letter went to all consignees explaining the recall expansion and risk associated with the recall. All consignees are to locate and quarantine the device immediately. Please return the product along with the completed inventory certification form to Zimmer. If you have any questions regarding this recall, call Zimmer at 877-946-2761. |  | 722 | 722 units (includes 2010 market withdrawal) | Worldwide Distribution-USA (nationwide) including the states of NY, MI, MN, OH, IN, IL, VA, MD, PA, WA, KY, CT, RI, NH, TX, MT, and CA and the countries of Chile, China, South Korea, and Australia. | Open |  |  | 722 | Not_Computer
Z-1654-2013 | 64888 | punch, surgical | LRY | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | HealthLink***BIOPSY PUNCHES***Disposable manual instrument for skin biopsy. | HealthLinkBIOPSY PUNCHES | 2 | July 02, 2013 | 2013 | Clorox Healthcare Holdings, LLC (dba HealthLink) | Due to lack of sterilization. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | On March 19, 2013, HealthLink sent their customers letters instructing them to do the following: 1) Screen their inventory and quarantine the product. 2) Notify all customers that received the affected products and ask them to return the lots to their firm. 3) Complete the attached form and return the lots to HealthLink Distribution center. For further assistance, customers were directed to call customer service at 1-800-638-2625. |  | 1044600 | 1,044,600 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AZ, CA, CO, FL, GA, IA, IL, IN, MI, MO, MS, MT, NC, NJ, NY, OH, SC,TX ,VA ,WA, WI, and HI, and the countries of Canada, Mexico, and New Zealand. | Terminated | February 25, 2014 | 238 | 2849900 | Not_Computer
Z-1653-2013 | 65286 | vaporizer, anesthesia, non-heated | CAD | Anesthetic vaporizer. | Anesthesiology | Anesthesiology | 510(k) | Delta VaporizerAn anesthetic vaporizer is a device used to vaporize liquid anesthetic and deliver a controlled amount of the vapor to the patient. | Penlon Sigma Delta Vaporizer | 2 | July 02, 2013 | 2013 | Penlon, Ltd. | It has come to Penlon Ltd's attention of an instance where the installation and pre-use checks listed in the User Manual may not have been followed prior to the clinical use of the vaporizer. Penlon Ltd has updated the Installation and daily Pre-Use Check sections of the User Manual to ensure that the instructions are explicit and clear. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Penlon sent an Urgent Field Safety Notice dated May 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded and requested to follow the instructions for daily pre-use checks, initial checks, and servicing. Customers were instructed to distribute the notice to all personnel who need to be aware within their organization and to any organization where the devices are used.For questions regarding this recall call 01235 547000. | Contact the recalling firm for information | 9208 | 9208 (of which 576 were within the US) | Worldwide Distribution - USA including FL, KY, PA, ND, MN, IL, NJ, TX, TN, and VA. Internationally to Belgium, China, Czech Republic, France, Germany, Switzerland, Australia, Finland, Holland, Lithunia, Malta, Poland, and Slovakia. | Open |  |  | 9208 | Not_Computer
Z-1658-2013 | 64973 | graft,bypass,coronary artery | NAV |  |  | Cardiovascular | HDE - Humanitarian Device Exemption | ***REF SC2101***eMax 2 Plus Console, Single Port, with Irrigation***SN G17307901802***Product Usage:The RPM of the Motor Handpiece may be less than 80,000 RPM (approximately 72,000 RPM) in the free running state when using the optional Hand Control. The low Motor Handpiece RPM would only be experienced in the Free running state (Hand Control fully depressed and no load on the cutting bur). This condition would not affect the power output of the Handpiece in the loaded working range (typically 50,000 to 70,000 RPM. | Emax Consoles | 2 | July 02, 2013 | 2013 | The Anspach Effort, Inc. | The RPM of the Motor Handpiece may be less than 80,000 RPM. | PRODUCTION CONTROLS: Process Control | Anspach Effort sent an "URGENT: Medical Device Correction" letter dated June 24, 2011 to their customers via mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete a reply form promptly confirming receipt of their letter and return it by fax to the number provided on the form. If any of the affected products were distributed to other services or facilities, please forward this information as appropriate. For questions contact Anspach Customer Support at 800-327-6887 or email customer.support@synthes.com. |  | 61 | 61 units | Worldwide Distribution - USA Nationwide including the states of: AZ, CA, FL, KY, ME, MI, MO, OH, PA, TN, TX , VA, and the countries of: Australia, Egypt, Italy, Japan, Switzerland, Taiwan, and Turkey; | Terminated | December 24, 2013 | 175 | 61 | Not_Computer
Z-1660-2013 | 64078 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Angio KitsA custom procedural tray is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. Contained within the tray are hand controller sheaths which are intended to enable increased flow control (when compared to hand injection) but with the feel of a hand syringe. It provides more control over the total volume delivered by allowing the physician to adjust contrast delivery during the injection and monitor total volume delivered to the patient. | Custom Procedural Trays | 2 | July 03, 2013 | 2013 | Merit Medical Systems, Inc. | Custom Procedural trays contain Medrad Hand Controller Sheaths which were recalled because the seal between the white polyethylene and the clear plastic may be missing on the package resulting in a potential breach of the sterility of this accessory. | PRODUCTION CONTROLS: Packaging | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" letter dated November 21, 2012, to affected customer. The letter identified the product, the problem, and the action to be taken by the customer: Ensure all personnel within their institution to which the affected Merit trays were distributed are notified, quarantine all affected trays and return the affected product to Merit - Maastricht. Complete and sign the Customer Response Form attached to this letter and perform the following steps:(1) Immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031. (2) Mail the completed original Customer Response Form from (1) above using the provided postage-paid, self-addressed envelope to Merit. Product Return Instructions: Please return the affected products by shipping them back to Merit via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department, RGA #NC-54948 Merit Medical Maastricht, Amerikalaan 42, 6199 AE Maastricht - Airport, The Netherlands. NOTE: A copy of the Customer Response Form must accompany the shipment. If you have any questions, please call 1-804-416-1048. | Contact the recalling firm for information | 134 | 134 trays | U.S. Military Hospital on base in Landstuhl, Germany | Terminated | July 12, 2013 | 9 | 134 | Not_Computer
Z-1661-2013 | 65493 | sleeve, limb, compressible | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | VenaFlow Elite 30B-B and 30BI-BProduct Usage: The VenaFlow Elite System is a compression pump used for the prevention of Deep Vein Thrombosis. | AIRCAST VENAFLOW ELITE SYSTEM | 2 | July 03, 2013 | 2013 | DJO, LLC | DJO, LLC has recently identified a product safety issue with the battery assembly for the VenaFlow 30B-B and 30 BI-B products. There is a potential for battery leakage which could result in thermal meltdown (fire). | DESIGN: Device Design | DJO sent a Customer Notification letter dated May 28, 2013 to Domestic customers via USPS certified return receipt and to International customers via UPS on June 14, 2013. The Customer Notification informed customers of the product safety issue, product description, and instructions to follow. DJO plans to install replacement batteries at no charge. For questions contact your local DJO Customer Service Representative. |  | 2491 | 2491 | Worldwide Distribution - USA Nationwide and countries of: Australia, Belgium, China, Guatemala CentroAmeria, France, Hong Kong, India,Italy, Korea, Lebanon, Malaysia, Morocco, Pakistan, Paraguay, The Netherlands, Romania, Singapore, Spain, Sweden, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezula. | Terminated | February 13, 2014 | 225 | 2491 | Not_Computer
Z-1662-2013 | 65431 | nucleic acid amplification assay system, group b streptococcus, direct specimen test | NJR | Streptococcusspp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX Instrument, catalog #441916 Product Usage:The BD MAX System, used in conjunction with the appropriate extraction kits and reagents is capable of automated amplification and detection of target nucleic acid sequences for use in clinical, industrial, and research settings. | BD MAX Instrument | 3 | July 03, 2013 | 2013 | Becton Dickinson & Co. | Electrical power module installed on some in vitro diagnostic instruments does not meet industry standards as labeled. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Becton Dickinson (BD) notified their customers in the US by phone on May 31, 2013 and by an Urgent Field Correction Action letter dated May 2013. Customers were advised of the affected product, problem and actions to be taken. The letter informed customers that an onsite visit by a BD Service representative is planned to perform a system check so to determine if the instrument is affected. BD's field service team plans to immediately replace the module on site if the affected part is found. For questions contact BD Technical Services Department at 1-800-638-8663. |  | 71 | 71 instruments | Worldwide Distribution - USA Nationwide and in the countries of: Japan, Singapore, Hong Kong, China, Australia, Europe, and Chile. | Terminated | August 19, 2013 | 47 | 71 | Not_Computer
Z-1665-2013 | 65429 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Inserter Outer Sleeve Product Usage:The Inserter Sleeve is an optional instrument to add further compression of the Inserter into the Implant during implantation of S-LIF Intervertebral Body Fusion Devices using bone graft to facilitate fusion of the spine of skeletally mature patients. Model: SI50003 RevA | SLIFT | 2 | July 05, 2013 | 2013 | SpineFrontier, Inc. | The outer sleeve would not fit over the S-LIFT Inserter assembly because the hole at the top of the sleeve handle was too small. | DESIGN: Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed consignees that a modification has been implemented to the Outer Sleeve component of the S-LIFT Inserter and that SpineFrontier has initiated removal of the affected product from the field.. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 devices | USA Nationwide Distribution including the states of: FL, TX, OK, MO, NJ, MA | Open |  |  | 12 | Not_Computer
Z-1664-2013 | 65432 | tray, surgical, instrument | FSM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | BASE, S-LIFT Instrument Case Product Usage:System cases are designed to hold and secure all instrumentation to be used with a surgical system. Case and caddy components are used to contain implants intended for intervertebral body fusion and fixation of the spine. | SLIFT | 2 | July 05, 2013 | 2013 | SpineFrontier, Inc. | The S-LIFT Instrument Case contained a bracket with peeling nylon coating. | DESIGN: Component Design/Selection | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a modification has been implemented for the improved durability of S-LIFT Case Brackets and that SpineFrontier has initiated removal of these cases from the field. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 | USA Nationwide including the states of: FL, CO, Wash DC, VA, KS, TX, OH, TN, OK, CA, IL, MO, and MN. | Open |  |  | 12 | Not_Computer
Z-1666-2013 | 65442 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InSpan Inserter, Right Assembly; InSpan Inserter, Left AssemblyProduct Usage:Inserters are stainless steel instruments that have an inner shaft. The left and right inserters are used in combination with each other to place the implants into the body. Inserters are used to install the InSpan Spinous Plate System for the purpose of achieving supplemental fixation. | InSpan Inserter | 2 | July 05, 2013 | 2013 | SpineFrontier, Inc. | Potential for set screw to be come deformed. | DESIGN: Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated May 30, 2013 to all affected consignees. The letter identified the affected product. problem and actions to be taken. The letter informed consignees that a correction was implemented for InSpan Inserters. Consignees were instructed to remove parts from the field immediately for modification and that replacement Inserters will be shipped upon receipt of completed acknowledgement form and RMA request for part return. For questions call 978-232-3990. | Contact the recalling firm for information | 61 | 61 individual inserter assemblies | Worldwide Distribution - USA Nationwide: in the states of: FL, CO, VA, KS, TX, OH, TN, OK, CA, IL, MO, MN, Wash DC, including PR and the countries of Dominican Republic, Medico, Panama | Open |  |  | 61 | Not_Computer
Z-1669-2013 | 65462 | dc-defibrillator, low-energy, (including paddles) | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Cardio Medical Products, Inc. Defibrillation/Cardioversion/Monophasic or Biphasic Pacing/ECG Electrode Rx Only Single Patient Use OnlyManufactured for Cardio Medical Products 385 Franklin Avenue, Suite L Rockaway, New Jersey 07866www.cardiomedicalproducts.com(800) 227-3633 (973) 586-9624 FAXDefibrillator | Cardio Medical Products Inc. | 2 | July 08, 2013 | 2013 | Cardio Medical Products | Covidien received customer reports of Arcing/Sparking on the defibrillation electrode leadwire which may affect Cardio Medical Products Inc. defibrillator electrodes Model Number O101. | PRODUCTION CONTROLS: Process Control | Cardio Medical Products, Inc. issued an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and segregate any product with the affected lot numbers and return affected product according to the instructions for product credit by contacting the Customer Service group at 1-800-962-9888, option 1, then option 2, and reference the recall letter to request a Return Goods Authorizing (RGA). Customers were instructed to complete the enclosed Verification Form and return it to Covidien by fax to 800-895-6140, even if they did not have any of the affected lots. Customers were also instructed to report any issues to the Quality Assurance Group to ensure proper device reporting procedures are followed. Customers with questions were instructed to call 1-800-962-9888, option 8, then ext 2500.For questions regarding this recall call 973-586-9500. | Contact the recalling firm for information | 657 | 657 units | Nationwide Distribution | Open |  |  | 657 | Not_Computer
Z-1667-2013 | 65449 | tube, tracheostomy (w/wo connector) | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Tracoe Mini 3.0mm Tube, for pediatric use.Pediatric tracheostomy tube used to provide safe airway access. | Tracoe Mini 3.0mm tracheostomy tube. | 1 | July 08, 2013 | 2013 | Bryan Medical Inc | The pediatric tracheostomy tubes were mislabeled as the smaller infant tracheostomy tubes. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The consignees were contacted by telephone and asked to return the recalled product. |  | 5 | 5 units | Distributed in KY and AL. | Terminated | October 29, 2013 | 113 | 5 | Not_Computer
Z-1668-2013 | 65430 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion SetsThe Paradigm infusion sets are intended for use with Paradigm insulin infusion pumps. | Medtronic MiniMed Paradigm Insulin Infusion Sets | 1 | July 09, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic MiniMed Paradigm Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion sets because of a potential safety issue that can occur if insulin or other fluids come in contact with the inside of Medtronic Paradigm infusion sets. Under certain conditions, the infusion set may malfunction and deliver an incorrect volume of insulin. | DESIGN: Device Design | Medtronic, Inc. sent an Urgent Medical Device Safety Notification letter dated June 10, 2013, to all affected consignees. The letter identified the product and notified healthcare professionals of the potential for over or under delivery of insulin if insulin or other fluids contact the inside of Medtronic Paradigm Tubing Connectors. The letter informed the consignees that it applies to all Medtronic infusion sets designed for use with Medtronic Paradigm family infusion pumps. Paradigm Insulin Pump users, distributors, and healthcare professionals were all instructed to contact Medtronic's 24 hour helpline at (888) 204-7616.For questions regarding this recall call 818-567-4700. |  | 11000000 | 11,000,000 units | Worldwide distribution: US (nationwide) distribution: Algeria, Argentina, Australia, Bahamas, Bahrain, Bermuda, Boliva, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Southern Territories, Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Macedonia, Malaysia, Mexico, Morocco, Netherlands, Norway, New Zealand, Oman, Panama, Paraguay, Peru, Poland, Province of China, Taiwan, Qatar, Republic of Korea, Romania, Russia Federation, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, Uruguay, andÂŽVenezuela. | Open |  |  | 11000000 | Not_Computer
Z-1672-2013 | 65317 | catheter, nephrostomy | LJE |  |  | Gastroenterology/Urology | 510(k) | NephroMax high Pressure Nephrostomy Balloon Catheter. Material Number/Catalog Number: M0062101400/210-140, M062101440/210144Dilation of the nephrostomy tract. | NephroMax high Pressure Nephrostomy Balloon Catheter. | 2 | July 09, 2013 | 2013 | Boston Scientific Corporation | Incorrect sized renal sheath was packaged with the NephroMax high pressure nephrostomy Balloon Catheter. | PRODUCTION CONTROLS: Process Control | The firm, Boston Scientific, sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated May 28, 2013, consignees via Federal Express Priority Mail to its consignees/customers.OUS - Customer notification letters has also been sent to consignees/customers in the EU. The letter instructs the customers to Immediately discontinue use of and segregate recalled product; check their inventory; immediately remove all affected product from their inventory, and return it to Boston Scientific at Boston Scientific Corporation, US Distribution Center, Customer Fulfillment Center, 500 Commander Shea Blvd., Quincy, Massachusetts 02171 include original Verification Tracking Form and RGA #. The letter also requests the completion and return of the Account Reply Verification Tracking Form via email to: MapleGroveFieldActionCenter@bsci.com or Fax to: Field Action Center 1-866-213-1806. Follow up communications will be made to the customers that do not respond to this inquiry. Distributors are instructed to forward the recall notification to their customers.If you have any questions call 763-494-1133 or email to MapleGroveFieldActionCenter@bsci.com. |  | 27 | 27 | Worldwide distribution: USA (nationwide) including states of: NJ, NY, OH, IL, MI, MD, MA and countries of: Denmark, Great Britain, and Spain. | Terminated | January 23, 2014 | 198 | 42 | Not_Computer
Z-1674-2013 | 65332 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn./Ext DLProduct Number: 050-87216Automated Peritoneal Dialysis System with integrated stay safe patient connector and extended drain line. | Fresenius | 2 | July 10, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Cassette may leak during set up or during peritoneal dialysis treatment |  | Fresenius Medical Care sent a Urgent Recall letter dated June 11, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their stock immediately and discontinue use and return product. Please contact your Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions to return the product. |  | 5480 | 5480 units | Nationwide Distribution | Open |  |  | 5480 | Not_Computer
Z-1681-2013 | 65320 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***REF C7122***QTY 1***APEX Arthroscopy Tubing Set.Intended for use for joint irrigation and dissention during arthroscopic procedures. | APEX Arthroscopy Tubing Set | 2 | July 10, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Conmed Linvatec received complaints that their products C7120 APEX Arthroscopy Tubing Set, One Connection and C7122 APEX Arthroscopy Tubing Set, Two Connection, contained a scratch/grooved area on the inside diameter of the pressure fitting that may cause the APEX Irrigation (pump) Console to sense a pressure drop and increase fluid flow to the surgical site. | PRODUCTION CONTROLS: Equipment Maintenance | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery or equivalent on 5/30/2013 to all domestic direct consignees. ConMed Linvatec is translating customer notifications letters into the necessary languages and will send International notification letters no later than 6/4/2013. |  | 4691 | 4,691 units, total of each | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, GA, IL, IN, KS, MO, NC, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, and WI, and the countries of Austria; Belgium, Canada, China, Czech Republic, Finland, France, Great Britain, Italy, and Japan. | Terminated | September 04, 2013 | 56 | 4691 | Not_Computer
Z-1671-2013 | 65347 | dilator, esophageal | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RigiflexTM 11 Single-Use Achalasia Balloon Dilator, Material Number (UPN) M00554510Product Usage:The Boston Scientific Corporation RigiflexTM II Single- Use Achalasia Balloon Dilator is indicated for dilatation of the cardia in patients with achalasia. | RigiflexTM 11 SingleUse Achalasia Balloon Dilator | 3 | July 10, 2013 | 2013 | Boston Scientific Corporation | Product labeled with incorrect expiration date. | EXPIRATION DATING: Incorrect or No Expiration Date | Boston Scientific sent an Urgent: Medical Device Recall letter to all affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The letter also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Follow-up communications will be made to the accounts that do not respond. For questions contact your local Sales Representative. |  | 91 | 91 | Worldwide Distribution - USA Nationwide including the state of NY, NE, MD, OH, TX, PA, VA, PR, CA, NJ, NY, CA and the countries of: AT, FR, DE, GB, IL, IT, LB, NL, ES, SE, TN, Argentina, and Japan. | Terminated | February 10, 2014 | 215 | 91 | Not_Computer
Z-1678-2013 | 65577 | pump, infusion, elastomeric | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter's elastomeric infusion pumps are non-electronic, ambulatory, disposable devices that deliver continuous medication for periods of 30 minutes up to 7 days. They are designed for medication therapies requiring continuous infusion.The infusion pumps can be utilized for slow, continuous delivery through clinically acceptable routes of administration such as intravenous (IV), intra-arterial (IA), and subcutaneous or epidural infusion of medications directly into an intraoperative site or subcutaneously for post-operative pain management. | INFUSOR SV Elastomeric Infusion Devices | 2 | July 10, 2013 | 2013 | Baxter Healthcare Corp. | There are reported incidents of leakage on the INFUSOR and FOLFUSOR pumps at the connection of the blue winged cap and the distal male Luer. Investigation of these reports indicates the incidence of leakage has only occurred after pumps are filled, primed, re-capped (using the blue winged Luer cap) and stored for a period of time, typically overnight. In addition, the reported instances of leakage | PRODUCTION CONTROLS: Equipment Maintenance | Baxter sent an "Important Product Information" notification dated July 31, 2009, with a follow-up "Important Product Information" notification mailed on March 5, 2010. The notifications informed customers of the potential issue, and included instructions to ensure that the winged Luer cap is tightly secured after filling and priming. Customers were also instructed to follow normal clinical practices to reduce contact with any leaking medication, and to store pumps with pre-dispensed medication in a sealed pouch. If customers noticed a leak they were instructed to place the pump in a plastic bag, quarantine the device and call Baxter's Corporate Product Surveillance at 1-800-437-5176 for further instructions. Customers were asked to return the attached reply form via FAX to 1-847-270-5457. If customers further distributed the affected product they were instructed to forward the notification to other facilities. Customers with questions were instructed to call 1-800-933-0303.For questions regarding this recall call 224-270-4667. |  | 116266 | Code 2C1700KP: 7,128 units; Code 2C1702KP: 105,732 units | Worldwide Distribution - USA (nationwide), Puerto Rico and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, Columbia, Czech Republic, Denmark (including Faroe Islands), Ecuador, Finland, France (including Martinique and Reunion Islands), Germany, Greece, Iceland, India, Ireland, Isreal, Italy, Luxemburg, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweded, Switzerland, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open |  |  | 573199 | Not_Computer
Z-1683-2013 | 65183 | system, irrigation, urological | LJH |  |  | Gastroenterology/Urology | 510(k) | Two Lead Arthroscopic Irrigation Set, for use with ARTHROMATIC Plastic Containers, Product Code 2C4030 - Approximate Length 98" (2.5. m)Usage:Baxter irrigation sets are intended for the delivery of irrigating solutions from the fluid container to the irrigation site during bladder irrigation or endoscopic procedures including cytoscopy, transurethral resection and arthroscopic procedures. the sets are intended for use with Baxter irrigating solutions in Uromatic or Arthromatic plastic containers. | Two Lead TUR Irrigation Set | 2 | July 10, 2013 | 2013 | Baxter Healthcare Corp. | A firm's internal investigation led to the discovery of pinholes in the packaging of the products. The pinholes are not likely to be detected by the end user prior to use and could compromise the sterile barrier properties of the packaging. | PRODUCTION CONTROLS: Packaging | Baxter sent an Urgent Product Recall letter dated May 15, 2013 to all affected custorners. The letter identified the affected product, problem and actions to be taken. Customers were requested to remove all affected products and return for credit by contacting Baxter Healthcare Center for Service at 888-229-0001. Customers were instructed to complete the attached Customer Reply Form and return it to Baxter by either faxing to 224-270-5457 or scanning and e-mailing the reply form to fca@baxter.com. |  | 530488 | 530,488 units | US Nationwide Distribution | Open |  |  | 1136792 | Not_Computer
Z-1676-2013 | 65290 | prosthesis, shoulder, hemi-, humeral, metallic uncemented | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/FlatowÂŽ The Complete Shoulder Solution, Humeral Stem Provisional 12mm x 170mmThis reusable trial device is used in orthopaedic joint replacement and is specific to the replaced joint (e.g. hip, knee, shoulder, ankle, elbow). It can be made of metal or plastic. | Bigliani/Flatow The Complete Shoulder Solution | 2 | July 10, 2013 | 2013 | Zimmer, Inc. | The firm is initiating a removal of one lot of the Bigliani/Flatow Humeral Provisional Stem (00-4301-012-17, lot 62283991) as the stems manufactured under the lot are 14mm x 170mm devices incorrectly etched and packaged as 12mm x 170mm devices. | PRODUCTION CONTROLS: Packaging Process Control | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification dated June 11, 2013, to all consignees via email and courier describing the affected device and problem and the risks associated with using the affected device. The notifications provided instructions for returning the affected device to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. Any questions or concerns concerning the recall, please call the customer call center at 1-877-946-2761. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer by email at zimmer.per@zimmer.com.For questions regarding this recall call 800-613-6131. |  | 1 | 1 | Nationwide Distribution including IL and MO. | Terminated | January 30, 2014 | 204 | 1 | Not_Computer
Z-1675-2013 | 65439 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | AngioDynamics RITA Main Cable, 9 Pin to 14 Pin, REF 700-101892. RITA Main Cables are indicated for use with the RITA RF Generators and RITA electrosurgical devices. | RITA Main Cable | 2 | July 10, 2013 | 2013 | Angiodynamics, INC | The RITA Main Cable product contains the incorrect Instructions for Use. | PRODUCTION CONTROLS: Packaging Process Control | Consignees were notified via certified mail on 06/11/2013 with an "Urgent Medical Device Field Safety Correction" notification. This included the correct Instructions for Use and the recommended actions. |  | 13 | 13 units | Worldwide Distribution, including the US, Austria, China, Italy, Russia and the UK. | Open |  |  | 13 | Not_Computer
Z-1692-2013 | 65554 | general surgery tray (kit) | LRO | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | Enforcement Discretion | GEO-MED ORTHO TOTAL JOINT TRACECART(R), REF 53-1831.07, Rx Only, NON-STERILEProduct Usage: general surgical | ORTHO TOTAL JOINT TRACECART(R) | 2 | July 11, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal recalled kits that contained Stryker togas which were recalled for a potential problem with sterilization. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | DeRoyal sent an Urgent: Voluntary Product Recall letter dated May 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, remove any affected product from their inventory, open its package to prevent use, and dispose in a standard trash receptacle. Customers were asked to complete the enclosed notice of destruction form and return it to DeRoyal by faxing it to 865-362-3716 or emailing it to recalls@deroyal.com. For question call 865-362-1034 or contact your DeRoyal representative. | Contact the recalling firm for information | 87 | 87 units | USA Nationwide Distribution in the state of PA | Terminated | July 12, 2013 | 1 | 87 | Not_Computer
Z-1695-2013 | 65459 | prosthesis, knee, femorotibial, semi-constrained, cemented, metal/polymer | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 154200 Oxford Fixed Bearing Partial KneeCemented Lateral TibiaSize A3 ARCOM UHMWPE / CO-CR-MO AlloyOnly for use in Lateral compartmentSterile RThe Oxford Fixed Bearing Knee is a Partial knee replacement system consisting of a femoral component and a fixed bearing tibial component for use in the lateral compartment of the knee. The Oxford Fixed Lateral bearings are indicated for use in the lateral compartment and intended to be implemented with bone cement. | Oxford Fixed Bearing Partial Knee Replacement | 2 | July 11, 2013 | 2013 | Biomet, Inc. | Biomet investigation found the design of the tibial bearing creates the potential for anterior impingement between the bone and the poly bearing component in terminal extension. | DESIGN: Device Design | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated June 7, 2013 to its consignees/customers. The letter identified the device being recalled, the reason for the recall, potential adverse events associated with the recall and actions to be taken. The consignees/customers were instructed immediately locate and remove the identified device(s) listed in the letter; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the FAX Back Response Form to 574-372-1683 prior to return of product within three (3) business days; return Goods to Biomet, Inc 56 East Bell Drive Warsaw, IN 46580. To confirm receipt of this notice by call 574-372- 1570. Obtain an RGA # by phone (574)372-6677 email rqarequest@biomet.com or through FAST; use priority carrier for your shipment; and if you have further distributed this product you MUST notify personnel of the letter. Questions should be addressed to 574-372-1570 M-F 8 a.m. to 5 p.m. |  | 26 | 26 | Worldwide distribution: USA including NC, IN, PA, and OH; and countries of: Canada and Netherlands. | Open |  |  | 794 | Not_Computer
Z-1693-2013 | 65090 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Service Replacement 0-Ring. Used in as a service part in 4 sub-assemblies manufactured and/or refurbished Fresenius hemodialysis machines, models 2008K, 2008K2, 2008K@Home and 2008T | Fresenius | 2 | July 11, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Incorrect rubber nitrile O-rings distributed instead of the correct material EPDM. (Ethylene Propylene Dine Monomer (M-class) rubber) | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Fresenius notified customers on May 24, 2013, by telephone and by certified mail with signature confirmation and fax-back form. Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they will be instructed to immediately discontinue use and place all 0-rings in a secure area for return to Fresenius Medical Care North America.Please contact your Fresenius Medical Care Technical Service Team at 1-800-227-2572, (choose option 4, then option 2) for instructions how to return the recalled product. |  | 725 | 725 units | Nationwide Distribution including the states of AR, CA, FL, GA, KY, LA, MA, MD, NC, NJ, NY, PA, PR, SC, TX and VA. | Open |  |  | 725 | Not_Computer
Z-1685-2013 | 65503 | screwdriver, surgical | LRZ |  |  | General & Plastic Surgery | 510(k) | Integra Hallu-Fix cannulated screw driver (model number P/N 129733ND) and cannulated screwdriver tip (model number P/N 129735ND).Indicated for use in fixation of fractures, osteotomies and arthrodesis of the first metatarso-phalangeal joint. | Integra HalluFix cannulated screw driver and cannulated screwdriver tip | 2 | July 11, 2013 | 2013 | Integra LifeSciences Corp. | The firm has received numerous complaints of the cannulated screw driver tip breaking during surgery. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm sent recall notification letters to all consignees via traceable email or traceable fax on 02/01/13. The firm requested that the consignees fill out and return the Recall Acknowledgement and Return Form. | No consumer action necessary | 221 | 154 screwdrivers were distributed; 67 screwdriver tips were distributed. | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WA and the countries of Canada, Austria, Belgium, Switzerland, Czech Republic, Germany, Denmark, Spain, Finland, France, Ireland, Israel, Italy, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, United Arab Republic, United Kingdom, and South Africa. | Terminated | July 25, 2013 | 14 | 221 | Not_Computer
Z-1725-2013 | 65148 | ventilatory effort recorder | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Finger Sensor for Remmers Recorder (Formerly SnoreSat Sleep recorder)To be used to gather heart rate and blood oxygen data during a sleep study. | Remmers (formerly SnoreSat) | 2 | July 11, 2013 | 2013 | SagaTech Electronics, Inc. | The outer jacket is receding away from the black sensor prematurely. | DESIGN: Process Design | SagaTech Electronics Inc, sent a letter to their Customers dated May 2008. The letter identified the product the problem and the action needed to be taken by the customer.Please take particular attention when cleaning these sensors by holding the cable somewhere in the middle and clean towards each end. This should become standard practice with all sensors to prolong their lifespan.If you have further questions please call (403) 228-4214 ext. 234. |  | 27381 | 152 units 77 devices from lot 13367 and 75 devices from lot 13710 | Worldwide Distribution: US (nationwide)ÂŽand countries of: ÂŽCanada and Mexico. | Terminated | July 22, 2013 | 11 | 152 | Not_Computer
Z-1726-2013 | 65670 | external mandibular fixator and/or distractor | MQN | Bone plate. | Dental | Dental | 510(k) | Synthes CMF Distraction SystemThe product is intended for use as a bone stabilizer and lengthening (and/or transport) device for correction of congenital deficiencies or post-traumatic defects of the mandibular body and ramus, where gradual bone distraction is required. It is intended for single use only. | Synthes Craniomaxillofacial (CMF) Distraction System | 2 | July 12, 2013 | 2013 | Synthes USA HQ, Inc. | The firm initiated a voluntary recall of the BC Distractor Body which is part of the Craniomaxillofacial (CMF) Distractor System, due to a mis-alignment issue. There have been instances reported within the impacted lots in which the slot in the barrel of the BC distractor was rotated such that it was not possible to attach the B-type and C-type foot plates. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated June 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to examine their inventory for affected product, remove it from use, and return it to Synthes. If they have affected product they were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter and indicate the number of devices found and note the Return Authorization Number and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers did not have the affected product they were asked to complete the Verification Section and faxt to 610-251-9005. Customers with questions were instructed to call 610-719-5450 or e-mail FieldAction@synthes.com.For questions regarding this recall call 610-719-5450. | Contact the recalling firm for information | 10 | 10 | Nationwide Distribution including PA, CT, DE, and NY. | Open |  |  | 10 | Not_Computer
Z-1727-2013 | 65163 | column chromatography & color development, hydroxyproline | JMM | Hydroxyproline test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | B-CrossLaps/serum (B-CTx in serum)REF 11972308 160 100 testsImmunoassay for the in vitro quantitative determination of degradation products of type I collagen in human serum and plasma as an aid in assessing bone resorption. The test may be used as an aid in monitoring antiresorptive therapies (e.g. bisphosphonates, hormone replacement therapy - HRT) in postmenopausal women and individuals diagnosed with osteopenia. The electrochemiluminescence immunoassay ECLIA is intended for use on Elecsys and cobas e immunoassay analyzers. | ELECSYS CrossLaps/Serum Immunoassay | 2 | July 12, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche Diagnostics internal investigations in R&D; have shown that the claims for biotin interference and the sample stability given in the Elecsys Âż-CrossLaps (Âż-CL) assay package insert version 5, could not be reproduced. The Âż-CrossLaps test results contribute to the monitoring of antiresorptive therapy. False low or false normal results may lead to delayed diagnosis and treatment of osteoporosi | DESIGN: Device Design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" (UMDC)letter dated May 2, 2013, to its consignees/customers via UPS Ground. The letter describes the product, problem and actions to be taken. The consignees/customers were instructed to apply the correct sample storage conditions as stated in the UMDC and to inform all relevant parties about the correct biotin interference claim. The consignees/customers were instructed to keep the UMDC on file for reference until the revised package insert is available. The consignees/customers were asked to complete and return a form via fax (1-888-943-4205). Consignees/customers were instructed to call Roche Diagnostics Technical Support, 24 hours a day, seven days a week, at 1-800-428-2336 if they had any questions about the recall. Non-responding accounts are monitored on an ongoing basis. |  | 3056 | 3056 units | Nationwide distribution: USA including states of: AL, CA, GA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, PA, TX, VA, VT, and WA. | Terminated | September 27, 2013 | 77 | 3056 | Not_Computer
Z-1729-2013 | 65664 | lancet, blood | FMK | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Safety Lancets - Odyssey, Button Activated;Tri- Bevel Tip for Virtually Painless Sampling; Needle automatically retracts, No lancing device required.26G - Single Use Only, Made In KoreaDistributed by Hummingbird Medical, Savannah, GA100 count Boxes; 40 boxes per caseBlood collection by penetrating skin with needle. | Odyssey | 2 | July 13, 2013 | 2013 | KooJoo Trading Company | Lancets were not sterilized within the dose range established by sterilization validation. | PRODUCTION CONTROLS: Process Control | The firm, KooJoo, sent a "MEDICAL DEVICE VOLUNTARY T RECALL" letter dated July 3, 2013 via first class mail to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to identify and hold all safety lancets you have in inventory that was received before June 2013; notify your retail customers of this product recall and complete and return the attached Customer Reply Form via email to heegonkim@hotmail.com and also return form along with any unused/undistributed product for replacement.If you have further questions, please email heegonkim@hotmail.com or call 209-403-9191 during 9:00am to 5:00pm M-F PST. |  |  | 7, 712, 000 pieces (1928 cartons) | US Distribution GA only. | Open |  |  | 1928 | Not_Computer
Z-1728-2013 | 65505 | catheter, balloon type | GBA | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | 27023WU Balloon Catheter, 3 Fr., 2/pkg, sterile, contains latex. The RUSCH ureter occlusion catheter is indicated for the routine occlusion of the ureter to prevent the movement of stone fragments during treatment. The coaxial catheters will also allow simultaneous drainage of urine or infusion of other solutions. | 27023WU Balloon Catheter | 2 | July 13, 2013 | 2013 | Karl Storz Endoscopy America Inc | Karl Storz Endoscopy-America has initiated the recall of the 27023WU Balloon Catheter because the packaging may be compromised, and therefore the sterility of the product could be affected. | PRODUCTION CONTROLS: Packaging | The firm, Karl Storz, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated May 31, 2013 to all customers who purchased the 27023WU Balloon Catheter. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and to return the recalled product back to Karl Storz.. The customers were also instructed to complete and fax back the Recall Response Form attached with the safety alert letter at Fax #424-218-8559. If you have any questions, please feel free to call 424-218-8289. |  | 40 | 40 units | US distribution to: AZ, FL, IL, IN and TX. | Terminated | July 15, 2013 | 2 | 40 | Not_Computer
Z-1735-2013 | 65088 | sera, animal and human | KIS | Animal and human sera. | Hematology | Pathology | 510(K) Exempt | Various types of Fetal Bovine SerumBovine sourced Animal Sera for in-vitro diagnostic device and as a growth promoting nutrient in cell culture systems. | GE Healthcare/PAA Healthcare | 2 | July 15, 2013 | 2013 | Paa Laboratories Inc | Current product labeling band data sheets state that FBS Pharma Grade and FBS/FBS Nativeare pure Fetal Bovine Serum (FBS). These products may contain added adult Bovine SerumAlbumin (BSA)of United States origin, water, and/or cell growth promoting additives. All otherFBS products, as listed below, including FBS Gold and FBS Standard, described as specialtysera', may contain adult BSA of Un |  | GE Healthcare notified customers via USPS on April 12, 2013 and April 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers being contacted by phone as well.Customers were asked to contact their local customer service representative if the product does not meet their needs. Customers were instructed to call (732) 457-8455 or 301-660-0256. For questions regarding this recall call 609-514-6888. |  | 280000 | approximately 280,000 liters | Worldwide Distribution - USA (nationwide), Novia Scotia, and Canada | Open |  |  | 280000 | Not_Computer
Z-1691-2013 | 65420 | pump, infusion, implanted, programmable | LKK |  |  | General Hospital |  | MedStream Programmable Infusion Pump , Implantable for the intrathecal delivery of Baclofen or MorphineProduct Code:91-4201US (40ml pump)91-4201 (40ml pump)Intended for the intrathecal delivery of Baclofen or Morphine. | MedStream Programmable Infusion Pump | 1 | July 16, 2013 | 2013 | Codman & Shurtleff, Inc. | Miscalibrated Fill Level Sensor may affect dosing level | DESIGN: Device Design | Codman issued a Medical Device Correction Notification to Medical Facilities dated June 7, 2013 via FedX to inform them of the potential problem and a Worksheet (instructions) to Identify Pumps with Miscalibrated FLS during each patient's next scheduled refill session or sooner if patient is symptomatic. If a pump is identified as having a miscalibrated FLS, notify Codman Neuro using the same worksheet. In addition a sample letter was provided to Physi cans to provide to patients on the system.Follow the recommendations for managing patients who are identified as having a pump with a miscalibrated FLS.For additional information,contact your Codman Neuro representative, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to Codman Neuro at 1-866-491-0974, Option 2. |  | 365 | 5 US and and 360 OUS | Worldwide Distribution-USA (nationwide) including the states of FL, NC, NJ, and TX, and the countries of Austria, Belgium, Spain, Great Britain, Germany, Italy, Ireland, Luxemburg, Russia, Saudi Arabia, Portugal, Greece, Egypt, Switzerland, and Turkey. | Open |  |  | 1102 | Not_Computer
Z-1737-2013 | 64989 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | KRD1 PedFuse Rods, Model: IM80100-13 Rev BProduct Usage: Pedicle rods are cylindrical rods which act as a linking mechanism between pedicle screws. The rod sits with the pedicle screw tulip and they are locked into place with a set screw or locking cap. The rod may be placed along several pedicle screws. Component rods are used during implantation of pedicle screw system devices intended to provide immobilization and stabilization of spinal segments of skeletally mature patients. | Rods Pedicle screw spinal system | 2 | July 16, 2013 | 2013 | SpineFrontier, Inc. | One of the rods in each of two test constructs did not meet fatigue requirement as outlined in ASTM F1717-04 dynamic compression testing 5 million cycles. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | SpineFrontier issued a Recall Notification to the affected distributor via telephone. The distributor was instructed to return the affected product for a replacement product. |  | 2 | 2 rods | USA Nationwide Distribution including the state of TX | Terminated | January 08, 2014 | 176 | 2 | Not_Computer
Z-1736-2013 | 65406 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Integra PANTA Arthrodesis Nail System.Intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. | PANTA Arthrodesis Nail Support Device | 2 | July 16, 2013 | 2013 | Integra LifeSciences Corporation | As a result of the investigation into an adverse trend of complaints for reports of misalignment of the calcaneal or tibial screws during use of the Panta Nail Support Device, Integra has decided to issue this voluntary medical device correction. The misalignment can result in the inability to properly insert the screws into the Panta Nail and create the potential need for revision surgery. | DESIGN: Device Design | Integra initiated a voluntary correction on May 9, 2013 by providing a written notification and copy of the revised Surgical Technique Manual as well as the specific highlighted sections of the manual addressing calcaneal and tibial screw alignment to all consignees that Integra records indicate have been shipped a PANTA Nail. The written notifications and Surgical Technique Manual have been delivered to US Consignees by traceable courier service and by traceable email or facsimile for Canada and Non-US Consignees. The notification informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; advises customers to review highlighted revisions of the Surgical Technique Manual to minimize the potential for misalignment of the screws and requests customers to complete and return the attached Acknowledgement Return Form. | Contact the recalling firm for information | 692 | 692 | Worldwide Distribution-USA (nationwide) and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Israel, Italy, Netherlands, Portugal, Russia, Saudia Arabia, South Africa, Spain and United Arab Emirates. | Terminated | November 07, 2013 | 114 | 692 | Not_Computer
Z-1748-2013 | 65303 | bronchial thermoplasty system | OOY |  |  | Anesthesiology |  | Alair Bronchial Thermoplasty Catheter, Bronchial Thermoplasty System; Material Number: M005ATS25010, Catalog Number: ATS 2-5. Product is manufactured and distributed by Boston Scientific Corporation (BSC), Sunnyvale, CA The Alair Bronchial Thermoplasty System (including the Alair Bronchial Thermoplasty Catheter) is indicated for the treatment of severe persistent asthma in patients 18 years and older whose asthma is not well controlled with inhaled corticosteroids and long acting beta agonists. | Alair Bronchial Thermoplasty Catheter | 2 | July 17, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific has discovered an inconsistent "Use By" expiration date on a single lot/batch of Alair Bronchial Thermoplasty Catheters. The expiration date indicated on the inner tray, 2012-04, is incorrect. The correct expiration date is 2014-04 as indicated on the outer box label. | PRODUCTION CONTROLS: Error in Labeling | The firm, Boston Scientific, sent a letter dated May 28, 2013, via express mail to all its direct consignees. The letter described the product, problem and actions to be taken. The customers were informed that you can continue to use the device until the expiration date on the outer box lablel, 2014-04; No product is being recalled and the you are not required to return product to Boston Scientific. The customers were instructed to read the enclosed information carefully and post the information in a conspicuous location near the affected product and/or take such other steps, consistent with your internal policies and procedures, to ensure that the information is easily accessible to all users of this device; and to complete and return the Customer Acknowledgment Form even if you no longer have any units from the affected lot/batch via email: MapleGroveFieldAc@bsci.com or Fax to: Field Action Center - 1866-213-1806. If you have any questions, call 763-494-1133 or email: MapleGroveFieldAc@bsci.com . |  | 35 | 35 units | Worldwide distribution: US (nationwide) including states of: CA, FL, IA, ID, MI, OK, TA and WA; and countries of: Canada and Thailand. | Terminated | August 28, 2013 | 42 | 35 | Not_Computer
Z-1749-2013 | 65492 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand, models S099 and S0998BSuction wands are intended for hand held use and pericardial sumps are intended for stationary use in removing excess fluid from the surgical field. | Edwards Lifesciences Suction Wand | 2 | July 17, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling certain lots of Rigid Suction Wands because they discovered plastic particulate inside the sterile pouch and within the internal diam ether of the wand. | DESIGN: Component Design/Selection | Edwards sent an Urgent Field Safety Notice Product Recall letter, dated June 19, 2013, via FedEx to all affected customers. Two additional attempts will be made to notify customers to return products. The letter idenditifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for any affected device in their inventory, quarantine all affected product from their inventory and return all devices to Edwards. Customers were instructed to acknowledge that they have reviewed the Recall Notice and confirm that they have taken appropriate action by completing, sign and dating the enclosed Recall Response Form and returning the form to Edwards by fax to Edwards Customer Service at 800-422-9329. Customers should contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number to return their current product to and receive replacement product. Customers with questions were instructed to call Edwards Customer Service at 800-424-3278. |  | 676892 | S099: 171,587 units and S099B: 505,107 units | Worldwide Distribution - USA (nationwide) and Canada. | Open |  |  | 676892 | Not_Computer
Z-1745-2013 | 65561 | screw, fixation, intraosseous | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Multi Vector Distractor Pin Holding Clamp, limited bone stockThe holding clamp is part of the Multi Vector Distractor which cis used for mandibular bone lengthening, including conditions such as congenital mandibular deficiencies or post-traumatic defects. The Multi Vector Distractor can be utilized to perform bone transport procedures as an alternative to bone grafts and free flaps. | Synthes | 2 | July 17, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary Medical Device removal for specified part numbers and lot numbers of the Mutis Vector Distractor Pin Holding clamp, limited bone stock. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however there | DESIGN: Device Design | Synthes sent a Urgent Notice: Medical Device Removal letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however, there is a remote potential that additional intervention would be required due to a non-functioning clamp.Synthes requested the product to be removed from inventory and returned to Synthes by calling 1-800-479-6329 to obtain a Return Authorization Number. Direct Accounts were asked to complete and return the Verification Section at the end of the letter indicating the quantity of affected product was found. Negative responses were also requested.If you have any questions please call (610) 719-5450. |  | 121 | 121 | US Nationwide Distribution including the states of MA, TX, MN, SC, UT, VA, and NY. | Open |  |  | 121 | Not_Computer
Z-1747-2013 | 65558 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic |  | Afinity CQ2015A Intraocular LensThe visual correction of aphakia in persons 60 years of age or older in whom a cataractous lens has been removed by cataract extraction. | STAAR Surgical Afinity IOL | 2 | July 17, 2013 | 2013 | Staar Surgical Co. | STAAR Surgical Company is recalling a limited number of nanoFLEX CC4204A and Afinity CQ2015A intraocular lenses due to the potential presence of glass particles in the storage vial. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, STAAR Surgical, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated June 10, 2013 to all the customers that received a limited number of nanoFLEX CC4204A and Afinity CQ2015 Intraocular lenses. The letter described the product, problem and actions to be taken. The customers were instructed to check your inventory; remove any lenses from the attached list; contact STAAR Surgical Customer Service for Returns Goods Authorization Number via phone: 1-800-352-7842; email: customerservice@staar.com; and/or fax: 1-800-952-4923; return listed lenses to STAAR Surgical; complete and return the attached response form by fax to Customer Service at (800) 952-4923, and pass this notice on to all those who need to be aware within your organization or any organization where the devices have been transferred.If you have any questions or concerns regarding the product return, contact STAAR Surgical Customer Service at (800) 352-7842. |  | 3361 | 3,361 units total | Nationwide distribution: USA including states of: AL,AZ, CA, CO, CT, FL, GA, HI, KY, IA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV. | Terminated | October 17, 2013 | 92 | 3361 | Not_Computer
Z-1744-2013 | 65355 | light, microscope | IBJ | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Polarizer D Fixed w/Rotating Red 1 lambda plate.Utilized in upright Zeiss microscopes to aid in diagnosis by means of color variation. | Polarizer D Fixed w/Rotating Red 1 lambda plate | 2 | July 17, 2013 | 2013 | Zeiss, Carl Inc | Carl Zeiss Microscopy, LLC is recalling the Polarizer D component used in All upright ZEISS microscopes (Axio Imager, Axio Scope.A1, Axio Lab.A1) due to Misaligned Full-Wave (Lambda) Plate. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On May 28, 2013, Carl Zeiss Microscopy, LLC recalled a Polarizer Component due to Misaligned Full-Wave (Lambda) Plate. Carl Ziess Microscopy distributed notices to their customers via e-mail. Carl Zeiss Microscopy, LLC requests that all customers immediately discontinue the use of the component for any applications which rely on accurate color-change sequence results. Your service provider will contact you to schedule a visit in which a factory trained service technician will confirm the alignment of your Polarizer D component;s lambda plate. If it is defective, a new Polarizer D component will be installed. There will be no charge for this action. Should you have any questions related to this issue, please do not hesitate to the Carl Zeiss facility in New York by telephone 914-681-7815 or via email at ed.mancini@zeiss.com. |  | 12 | 12 (1 in quarantine) | Distributed in the states of MI, TX, SD, FL, GA, NY, and NY. | Open |  |  | 12 | Not_Computer
Z-1751-2013 | 65647 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield; Sterile; Rx only; List No. 19666-28; SYMBIQ PUMP SET, CONVERTIBLE PIN, PIGGYBACK, 2 CLAVES PORTS AND OPTION LOK, NON-DEHP; NOMINAL LENGTH: 97 IN (249 CM); PRIMING VOLUME: 14 ML; For the administration of I.V. fluids; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USAThe Hospira Symbiq Infusion System is a general infusion system which provides IV infusion across the general spectrum of clinical care. The Symbiq pumps are standard sized pumps which are typically facility-based infusion pumps which may be used to deliver continuous and/or intermittent infusions to a variety of patient populations. | LifeShield Symbiq Set | 2 | July 18, 2013 | 2013 | Hospira Inc. | Hospira has received reports of customers experiencing air-in-line alarms related to administration sets that have a drip chamber with a ball valve. Preliminary investigation findings found that there is a potential for the ball valve to not seal the drip chamber as intended, leading to air in the line of the administration set. |  | Hospira sent a Urgent Device Recall letter dated June 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were instructed to check their inventories, quarantine any affected products, complete an attached reply form, and return the product to Stericycle for credit using a label included in the mailing. Customers who further distributed the products are requested to forward the recall notification information to their customers.A 100% effectiveness check will be performed for direct Hospira customers who do not return the reply forms and/or the product. Any direct or indirect accounts that return a reply form indicating they have product but do not return it will also be contacted as part of the effectiveness check. In addition, customers who further distributed the product will be requested to provide confirmation that they have notified their customers who purchased the product.Please contact Hospira Customer Care at 1-877-946-7747 ( M-F 7am - 6pm CT) or your Hospira representative regarding replacement product availability and for questions regarding this recall. | Contact the recalling firm for information | 14400 | 14,400 units | Worldwide Distribution - US (nationwide) including the states of IA, IL, MN, SD and WA., and Costa Rica. | Open |  |  | 14400 | Not_Computer
Z-1778-2013 | 65488 | light, ultraviolet, dermatological | FTC | Ultraviolet lamp for dermatologic disorders. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-080, one unit per box. The Dermalight 80 is a Phototherapy light unit that provides ultraviolet radiation therapy for diagnosed skin disorders such as Psoriasis and Vitiligo. | Dermalight 80. UVB080 | 3 | July 19, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-080 devices may have included the manual for a UVA-083 device. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 124 | 124 units | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, CT, FL, GA, IA, ID, IL, LA, MA, MN, MO, MS, NE, NJ, NY, OH, PA, SC, TN, TX, VA, WA, and WI; and countries of: Canada, Brazil, Malaysia, and Mexico. | Terminated | February 24, 2014 | 220 | 124 | Not_Computer
Z-1777-2013 | 65141 | applier, surgical, clip | GDO | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | LIGACLIPÂŽ Endoscopic Rotating Multiple Clip Applier (ER320), 10mm M/L, bulk non-sterileProduct Usage:The LIGACLIP Endoscopic Rotating Multiple Clip Applier is a single patient use instrument designed to provide a means of ligation through ENDOPATHÂŽ surgical trocars. The instrument delivers 20 titanium clips that individually advance after each firing. The shaft is made of a low glare material that minimizes reflective distortion. It is designed to rotate 360ÂŽ in either direction. The rotating knob is located to allow for a one-handed technique. | LIGACLIP Endoscopic Rotating Multiple Clip Applier | 2 | July 19, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary recall for LIGACLIPÂż 10mm M/L Endoscopic Rotating Multiple Clip Applier (ER320) due to potential clip formation and feeding issues which may result in improper clip formation, vessel damage and/or insufficient occlusion of the vessel or other structure |  | Ethicon sent an Urgent : Medical Device Recall letter dated April 30, 2013 to all primary and secondary consignees via overnight UPS to allow tracking of the receipt. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately and return affected products for credit. This credit is intended to facilitate replacement purchase. For questions contact your local representative or call Ethicon Customer Support Center at 1-877-384-4266 option 6. |  | 10488 | 10,488 | Worldwide Distribution - USA Nationwide to all fifty (50) states, and internationally including the following countries: ARGENTINA, ARGENTINA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, ECUADOR, EGYPT, EL SALVADOR, GREECE, HONDURAS, INDIA, INDONESIA, ISRAEL, JAMAICA, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA and YEMEN. | Open |  |  | 798201 | Not_Computer
Z-1779-2013 | 65494 | light, ultraviolet, dermatological | FTC | Ultraviolet lamp for dermatologic disorders. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-083, one unit per box.The Dermalight 80 is a Phototherapy light unit that provides ultraviolet radiation therapy for diagnosed skin disorders such as Psoriasis and Vitiligo. | Dermalight 80, UVB083 | 3 | July 19, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-083 devices may have included themanual for a UVA-080 device. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 6 | 6 units | US distribution including the states of: AL, CA, MN, and NC. | Terminated | February 24, 2014 | 220 | 6 | Not_Computer
Z-1772-2013 | 65504 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Endopath Electrosurgery Probe Plus II, twelve (12) different devices, including hooks, spatulas, right angles, curved dissectors, suction/irrigation devices, accessory ports, pool/sump devices and stone retrieval cannula, packaged and sterilized individually in a transparent flexible blister sealed to a Tyvek lid. 6 flexible packages are placed into a sales unit carton, with 4 cartons per corrugated shipper. The Endopath Electrosurgery Probe Plus II system has application in minimally invasive procedures to facilitate tissue dissection, coagulation, irrigation, and fluid evacuation through a common trocar sleeve. | Endopath Electrosurgery Probe Plus II | 2 | July 19, 2013 | 2013 | Ethicon Endo-Surgery Inc | The firm is recalling the ENDOPATH Probe Plus II Shafts because the possibility that a tear in the Tyvek packaging may affect the sterility of the device. | PRODUCTION CONTROLS: Packaging | The firm, Ethicon Endo-Surgery, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 22, 2013 to all consignees/customers via overnight UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to do the following: EFFECTIVE IMMEDIATELY-DO NOT USE ANY OF THE FOLLOWING TWELVE (12) PRODUCT CODES within the noted Expiration Dates provide in the letter; examine your inventory immediately to determine if you have affected product on hand; remove the affected product; complete and return the Business Reply Form (BRF) and fax to: 1-888-912-7089 or email to: ees3118@stericycle.com within 3 business days, even if you do not have affected product; return affected product with a photocopy of the completed BRF; and share this information with all of the appropriate staff at your facility. If you need additional shipping labels or a communications package, contact the Stericycle's support center at 1-888-943-4896 and reference Event 3118."Returned ENDOPATH Probe Plus II Shafts will be replaced with corrected ENDOPATH Probe Plus II Shafts once supply becomes available. If replacement product is not available and/or has been discontinued, a credit will be issue."If you need clinical or product support, please contact your local representative or call our Customer Support Center, 1-877-ETHICON (877-384-4266). |  | 327648 | 327,648 instruments | Worldwide distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, lA, ID, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, & WY and countries of: Canada, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Ecuador, Egypt, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Libya, Malaysia, Mexico, Pakistan, Panama, Philippines, Poland, Puerto Rico, Russia, Singapore, Slovenia, South Africa, Spain, Saudi Arabia, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, & Venezuela. | Terminated | December 24, 2013 | 158 | 327648 | Not_Computer
Z-1781-2013 | 65542 | device, surgical, cryogenic | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | Myoscience brand iovera Smart Tip, Myoscience Cryo-Touch IV, a Cryogenic Surgical device, REF Catalog Number: STA0311-10.Product Usage:The Myoscience Cryo-Touch IV (a.k.a. iovera) is used to destroy tissue during surgical procedures by applying freezing cold. It can also be used to produce lesions in peripheral nervous tissue by the application of cold to the selected site for the blocking of pain. The Cryo-Touch IV (a.k.a. iovera) is not indicated for treatment of central nervous system tissue. | iovera | 2 | July 20, 2013 | 2013 | Myoscience Inc | The outer box label of the product was labeled with a down revision label which did not include the US labeling requirements, specifically the symbol descriptors and also the "Caution: Federal Law (USA) restricts this device to sale by or on the order of a physician." This product was previously labeled for EU distribution only. | PRODUCTION CONTROLS: Labeling Mix-Ups | Myoscience representatives visited both affected sites on June 14, 2013 and relabeled the product in person. Neither site had used the product and was unaffected by the relabeling. No formal communication was issued. |  | 3 | 3 units | US distribution including CA and ID. | Open |  |  | 3 | Not_Computer
Z-1784-2013 | 65399 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | HemosIL RecombiPlasTin 2G, Part Number 0020002950 (8 ml size). The product is used for the evaluation of the extrinsic coagulation pathway and the monitoring of Oral Anticoagulant Therapy. | HemosIL RecombiPlasTin 2G, | 2 | July 22, 2013 | 2013 | Instrumentation Laboratory Co. | Instrumentation Laboratory is recalling HemosIL RecombiPlasTin 2G (RPT2G), Part No. 0020009580 (8 mL Size) due to being Out of Specification. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Between April 9, 2013 and April 11, 2013, Instrumentation Laboratory distributed "Important Product Correction" notification letters to their U.S. and foreign customers informing them that the on-board instrument stability claim has been reduced to 3 days for all current and future product lots of HemosIL RecombiPlasTin 2G. A second notification to non-respondents was initiated and completed on May 7, 2013. Technical questions were referred to 1-800-678-0710, Option # 2. |  | 2443507 | Kits per product:1)11080, 2)11603, 3)11487, 4)11483, 5)11290, 6)11562, 7)11535, 8)11577, 9)11505, 10)11708,11)11579,12)11580, 13&14) not distributed | Worldwide Distribution, including the US states of MA, ME, CT, NJ, NY, PA, MD, VA, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, MN, SD, MT, IL, MO, KS, NE, LA, AR, OK, TX, CO, ID, AZ, CA, OR, WA, AK, RI, VT and the countries of Angola, Argentina, Australia, Austria, Belgium, Bermuda, Brazil, Bulgaria, Canada, Caribbean, Chile, Colombia, Costa Rica, Czech, Denmark, Dubai, Ecuador, Egypt, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Iran, Israel, Italy, Jamaica, Jordan, Korea, Kossovo, Kuwait, Lebanon, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, Turkey, Ukraine, Uruguay, Vietnam, and Japan. | Open |  |  | 2443507 | Not_Computer
Z-1786-2013 | 65573 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Orthopedic | 510(K) Exempt | STERIS Fine Traction Device, one unit per boxProduct Usage:The Fine Traction Device is used to provide traction and enable fine adjustments to traction and orientation of the leg during orthopedic procedures. | STERIS Fine Traction Device | 2 | July 22, 2013 | 2013 | Steris Corporation | A bearing used in the assembly of the Fine Traction Device did not perform successfully in pre-release reliability testing; a replacement bushing passed reliability testing and was substituted per a change order. The change order authorized use of the substitute bushing in assembly of the Fine Traction Accessory; however, inadvertently the new part was not implemented. During complaint evaluati | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | STERIS Account Managers notified affected customers via onsite visits. All consignees were notified as of 5/29/2013. Customers were instructed to return affected product for a replacement unit. |  | 16 | 16 units | USA Nationwide Distribution in the states of CO, KY, LA, VA, and WI. The affected product was also distributed to Canada. | Terminated | August 28, 2013 | 37 | 16 | Not_Computer
Z-1785-2013 | 63947 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Eon and Eon Mini charging system, model numbers Eon -3701 (Wall Charger), 3711 (Portable Charger) and Eon Mini -3721 (Charger); This is a continuation of the firm's July 2012 recall.Product Usage:The Eon Wall Charger (3701) provides the capability to recharge the IPG Battery. The charger transmits RF energy through the antenna to the IPG battery to recharge it. The Eon Portable Charger (3711) provides the capability to recharge the IPG battery. The charger transmits RF energy through the antenna to the IPG battery to recharge it. The patient can be mobile during the recharging process with a portable charger. The Eon Mini Charger (3721) provides the capability to recharge the IPG battery while stimulation is either on or off. The charger transmits RF energy through the antenna to the IPG battery to recharge it. | Eon and Eon Mini | 2 | July 22, 2013 | 2013 | St. Jude Medical | There is issues of warmth or heating at the implant site during charging for the Eon Mini implantable pulse generators (IPGs). In the July 2012 letter St. Jude Medical informed customers that they would be implementing design improvements to the charger to address possible increased energy dissipation when the charger is misaligned with the IPG or the IPG is implanted too near the surface of the | DESIGN: Device Design | St Jude Medical sent an Important Medical Device Safety Information letter dated November 30, 2012 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a replacement charger for the Eon Mini IPGs is now available and will be provided at no charge. For questions or concerns call St. Jude Medical at 877-650-0367. |  | 75663 | 75,663 | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Luxembourg, Martinique, Mexico, Netherlands, Poland, Russian Federal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 75663 | Not_Computer
Z-1757-2013 | 65553 | graft, vascular, synthetic/biologic composite | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | AlboGraft Knitted with Collagen, Bifurcate 18x09Catalog Number: AMC1809AlboGraft Vascular Grafts are made of synthetic material and designed to replace sections of damaged or malfunctioning arteries. | LeMaitre Albograft | 1 | July 22, 2013 | 2013 | LeMaitre Vascular, Inc. | Blood blushing/leaking from the surface of the graft after implantation | PRODUCTION CONTROLS: Process Control | LeMaitre Vascular, Inc.sent an Urgent Field Safety Notice daed June 19, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and return the affected device to MeMaitre Vascular which they will replace with a new unit. Customers with questions were instructed to call 781-221-2266, ext. 183. For questions regarding this recall call 781-425-1670, ext 108. |  | 3 | 3 units | Worldwide distribution: USA including PA and countries of: BRAZIL, BULGARIA, FRANCE, GEORGIA, GERMANY, GREECE, ITALY, LEBANON, NETHERLAND, POLAND, SPAIN, and TURKEY. | Open |  |  | 39 | Not_Computer
Z-1794-2013 | 65259 | bed, flotation therapy, powered | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare Bed used with AccuMax Quantum VPC mattressAccuMax Quantum Complete mattress, and AccuMax Quantum Convertible mattressThe TotalCare Bed System is intended to provide a patient support ideally suited to be used in health care environments. The TotalCare Bed System may be used in a variety of settings including, but not limited to, acute care, including critical care, step down/progressive care, medical/surgical, high acuity sub-acute care, post anesthesia care unit (PACU), and sections of the emergency department (ED). The TotalCare Bed System is capable of being used with a broad patient population as determined appropriate by the caregiver or institution. | TotalCare Bed (1900) | 2 | July 23, 2013 | 2013 | Hill-Rom, Inc. | Hill-Rom is initiating this recall due to a potential issue with TotalCareÂż Beds equipped with AccuMax Quantum"mattresses. If the bed is put into the chair egress position, the magnets on the mattress may not holdthe bottom of the mattress flat to the bed frame. Ifthis happens, the foot section of the mattress willextend away from the bed at about a 30Âż angle (asshown in the photo). | DESIGN: Device Design | The firm, Hill-Rom, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 29, 2013 was sent to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Until the correction is implemented, after placing the bed into chair position, regularly check the foot section of the mattress. If necessary, push it back against the bed frame until the magnets hold it in place before allowing the patient to get back into or out of bed. In addition, forward a copy of this letter to any other facility personnel you deem appropriate.If you have any questions concerning this notification, please contact your distributor, your Hill-Rom representative, or Hill-Rom Technical Support at 800-445-3720 (Mon-Thurs, 8 am-6:30 pm; Friday, 8 am-6 pm). |  | 4121 | 4,121 Units | Worldwide distribution: USA (nationwide) and countries of: Canada, Columbia, Nigeria, Kuwait and Saudi Arabia. | Open |  |  | 4121 | Not_Computer
Z-1795-2013 | 65604 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos MX Solid State Flat Panel Digital Imager, Model number 5895003. The device is used for diagnostic X-ray imaging. | Axiom Aristos MX | 2 | July 24, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Notice about the potential hazard to patients or operators when using the function "table top float" of the AXIOM Aristos MX. When repositioning a patient lying on the table top and releasing the table top, it is possible that the fingers of the patient or operator are between the table top and the detector tray and if the table top is moved, the fingers can be squee | DESIGN: Device Design | On 6/5/13, consignees were notified by letter of a "Field Safety Notice" and received instructions about how to avoid the potential risk of this issue while using the device. Siemens also rolled out a modification of AXIOM Aristos MX that will avoid these potential injuries in the future. | Contact the recalling firm for information | 401 | 401 | Nationwide Distribution. | Open |  |  | 401 | Not_Computer
Z-1805-2013 | 65614 | culture media, selective and differential | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL Campylobacter CSM Agar (Charcoal-Based Selective Medium), Catalog # 299614, packaged in cartons of 20 plates, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA 800-638-8663 www.bd.com/ds***BCYE Selective with PAV is used in qualitative procedures for isolation of Legionella species from clinical specimen and nonclinical (environmental) samples. It is similar to the Edelstein formula, except that the concentration of polymyxin B is reduced by half, and vancomycin is substituted for cefamandole. | BBL Campylobacter CSM Agar (CharcoalBased Selective Medium) | 2 | July 24, 2013 | 2013 | Becton Dickinson & Co. | Microbiological identification media may exhibit reduced levels of Vancomycin. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated June 2013 to Distributors by fax, US customers via letters by UPS and OUS BD sites via email and teleconference. The letter described the product, problem and actions to be taken. All customers were instructed to discontinue use of the listed product; discard any remaining packages, and complete and return the attached form, whether or not you have any inventory remaining, via Fax to: Attention: Regulatory Compliance, 410-316-4258. BD will issue replacements for the discarded material. It is not necessary to contact BD by telephone to receive replacements.If you have any questions, call 410-316-4000. | Contact the recalling firm for information | 293 | 293 cartons | Worldwide distribution: US (nationwide) and countries of:Belgium, Canada, Columbia, Mexico, Singapore, and Taiwan. | Terminated | September 03, 2013 | 41 | 1744 | Not_Computer
Z-1810-2013 | 65600 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | Medtronic Preimplantation Test Kit, Catalog No. 21047. The kit is indicated for use with Medtronic CSF-Flow Control Valves and Delta Valves, when a simple preimplantation test is desired to verify that the valve conforms to labeled product specifications. | Medtronic Preimplantation Test Kit, | 2 | July 25, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Preimplantation Test Kit because the outer carton had expiration dates that were incorrect. The date on the outer carton is later than the expiration date for one or more of the individually packaged sterile components in the kit. | EXPIRATION DATING: Incorrect or No Expiration Date | An "Urgent Field Safety Notice" recall letter dated 6/28/13 was sent to customers who purchased the Medtronic Preimplantation Test Kit, catalog number 21047, to inform them of the recall due to the incorrect outer carton expiration dates. The recall letter informed the customers of the problems and actions identified. Customers were instructed to contact the Quality Department at (805) 571-8725 for questions. | Contact the recalling firm for information | 915 | 915 units | Nationwide Distribution in the US. | Terminated | September 23, 2013 | 60 | 915 | Not_Computer
Z-1808-2013 | 65654 | intracranial neurovascular stent | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Stryker brand Neuroform 3 Microdelivery Stent System, 4.5 mm X 20 mm, REF Catalog Number: E345020, Model/Product Number: M003450200, Product is manufactured by Boston Scientific Corporation, Fremont, CA and distributed by Stryker Neurovascular, Fremont, CA The Neuroform Microdelivery Stent System is intended for use with embolic coils for the treatment of wide neck, intracranial, saccular aneurysms arising from a parent vessel with a diameter of greater than equal to 2 mm and greater than equal to 4.5 mm that are not amenable to treatment with surgical clipping. | Stryker brand Neuroform 3 Microdelivery Stent System | 2 | July 25, 2013 | 2013 | Stryker Neurovascular | The device is labeled incorrectly, in that, the device Model Number on the outer carton and inner pouch labels did not match, although both referenced the same lot number. The Model Number on the outer box states UPN M003E345020 (4.5mm x 20 mm Neuroform 3 Stent) while the pouch label states UPN M003E3450300 (4.5mm x 30 mm Neuroform 3 Stent). | PRODUCTION CONTROLS: Packaging Process Control | Stryker Neurovascular sent an Urgent Field Safety Notice on June 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their internal inventory and quarantine all subject devices pending return to Stryker. Circulate the Field Safety Notice internally to all interested/affected parties. Maintain awareness of this notice internally until all requred actions have been completed within their facility. Inform Stryker if any of the subject devices were distributed to other organization. Complete the attached customer response form and return to their Stryker Representataive.For questions regarding this recall call 510-413-2500. |  | 23 | 23 units | International Distribution including Austria, China, and Germany. | Terminated | September 30, 2013 | 67 | 23 | Not_Computer
Z-1809-2013 | 65391 | manipulator, plunger-like joint | LXM |  |  | Physical Medicine | 510(k) | Advanced Orthogonal Percussion Adjusting Instrument Indicated for use by a licensed Chiropractor to assist in making Chiropractic adjustments to the Atlas vertebra. | Advanced Orthogonal Percussion Adjusting Instrument | 2 | July 25, 2013 | 2013 | Advanced Orthogonal Equipment, Incorporated | Unapproved medical device | PREMARKET APPROVAL: No Marketing Application | The firm, Advanced Orthogonal Equipment, sent a "CORRECTION NOTICE" to its consignees/customers. The notice described the product, problem and actions to be taken. The customers were instructed that " none of these devices be used pending final clearance by FDA".Questions and concerns regarding this notice should be directed to the President of Advanced Orthogonal Equipment by phone at: 727-528-8700 or email: Atlasortho@aol.com. | No consumer action necessary | 54 | 54 | Nationwide distribution: US including states of: AL, CA, CO, FL, GA, ID, KY, MD, NC, NJ, NY, OH, PA, TN, VA, and WA. | Open |  |  | 54 | Not_Computer
Z-1806-2013 | 65703 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | ProSpace Peek Implant 5 degree x 8.5 x 22 mm The device is intended for vertebral body replacement or intervertebral body fusion to aid in the surgical correction or stabilization of the spine. The Aesculap PEEK Spinal Implant System is indicated for use in the thoracolumbar spine (T1 to L5) for partial or total replacement of a collapsed, damaged, or unstable vertebral body due to tumor or trauma to achieve anterior decompression of the spinal cord and neural tissues, and to restore the height of a collapsed vertebral body. | Aesculap PEEK Intervertebral Body Fusion System | 2 | July 25, 2013 | 2013 | Aesculap, Inc. | Aesculap Implant Systems, LLC initiated a recall of the Aesculap Prospace PEEK 5 degree Implant - 8 x 8.5 x 22 mm (SN038P) because the device is incorrectly etched with an 8 degree lordosis instead of a 5 degree lordosis. The device is labeled correctly, only the etching of the implant may be incorrect. | TRAINING: Employee Error | Aesculap notified customers via email and phone on June 18, 2013, and were instructed to check their inventory and return affected product to Aesculap.Aesculap believes that the risk to patient is low.For further questions please call 1-800-234-9179. |  | 9 | 9 | US Distribution including the states of NV and MI | Terminated | January 13, 2014 | 172 | 9 | Not_Computer
Z-1807-2013 | 65401 | chair, adjustable, mechanical | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Swift Mobil shower chair, Swift Mobil 160 shower chair, and Swift Mobil Tilt shower chair.Height adjustable four wheeled shower chairs for persons of various weights. | Swift Mobil Shower Chair | 2 | July 25, 2013 | 2013 | Etac Supply Center Ab | Etac is voluntarily conducting a field correction of the Etac Swift Mobil shower chair. If the Swift Mobil shower chair loses a caster and thus becomes instantly unstable, there is potiential risk for harm or injury. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Etac sent an "URGENT FIELD CORRECTION" letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 7200 | 7200 internationally | International distribution: Austria, Australia, Belgium, Brazil, Canada, Czech Republic, Denmark, Esonia, Finland, France, Great Britain, Germany, Greece, Hong Kong, Ireland, Isreal, Iceland, Italy, Kingdom of Bahrain, Kuwait, Lithuania, Netherlands, Norge, New Zealand, Spain, Switzerland, Sweden, and United Arab Emirates. | Open |  |  | 7200 | Not_Computer
Z-1811-2013 | 65686 | [6 empty] | Sodium Bicarbonate, USP Provided bulk as raw material - label not applicable according to firmDialysis Grade 1.5Dialysis Grade 2As an API or excipient in the manufacture of Medical Devices and Drugs | Church and Dwight Inc. Sodium Bicarbonate, USP | 2 | July 25, 2013 | 2013 | Church & Dwight Inc | Church and Dwight Co., Inc. has initiated a recall for Sodium Bicarbonate, USP produced by the Old Fort, OH Sodium Bicarbonate Production facility due to metal and other contamination. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Church and Dwight Inc., issued an Urgent Recall letter dated June 11, 2013, and return response form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine and return the affected product. For questions regarding this recall call 609-279-7748. |  | 234 | 234 sacks at 2500 lbs. each of Dialysis Grade 1.5 ; 880 bags at 50 lbs. each of Dialysis Grade 2 | Nationwide Distribution including IL, PA, IN, OH, MI, TN, TX, MN, MO, NJ, KY,and MD. | Open |  |  | 234 | Not_Computer
Z-1812-2013 | 65619 | dna probe, yeast | MLA | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD Affirm VPIII Microbial Identification Tests, packaged in kits, 120 test\kit, Catalog # 446257 and 24 test\kit, Catalog # 446252; kits are labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle Sparks, MD 21152 USA, 800-638-8663, www.bd.com/ds***The Affirm VPIII Microbial Identification Test is a DNA probe test intended for use in the detection and identification of Candida species, | BD Affirm VPIII Microbial Identification Tests | 2 | July 26, 2013 | 2013 | Becton Dickinson & Co. | In vitro diagnostic test kit may exhibit invalid internal negative control results. |  | Becton Dickinson Diagnostic Systems notified distributors by fax on June 13, 2013, and advised consignees to discontinue use and discard affected product for credit. Distributors were requested to return their distribution list to BD for direct notification by the recaller. End users were contacted by letter on June 13, 2013 and the notification addressed to Microbiology Lab Director/Risk Managers advised that a small number of tests associated with the referenced lots of product may exhibit invalid internal negative control results. If encountered, the issue would result in an invalid test, as described in the package insert. End users were requested to discontinue use of these lot numbers, discard any remaining packages for replacement and return the enclosed response form.Customers with questions were instructed to contact BD Technical Services Department at 1-800-638-8663. For assistance regarding replacements customers were instructed to contact BD Customer Service at 1-800-675-0908.For questions regarding this recall call 410-316-4054. |  | 177 | 177 kits | Nationwide Distribution | Terminated | September 28, 2013 | 64 | 177 | Not_Computer
Z-1813-2013 | 65328 | ventricular (assisst) bypass | DSQ |  |  | Cardiovascular |  | HeartWare Ventricular Assist SystemFor use as a bridge-to-cardiac transplantation in patients who are at risk of death from refractory end-stage left ventricular heart failure. | Ventricular Assist Device | 2 | July 29, 2013 | 2013 | HeartWare Inc | Potential to electrostatic discharged (ESD) | DESIGN: Device Design | HeartWare sent an Urgent Medical Device Correction Safety Notification letter on May 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to educate their patients about the HeartWare Controller and instructions for equipment handling stated in the IFU and recognized practices to avoid electrostatic discharge. The direct accounts were also told to scan and return the signed Acknowledgment Form via fax to 305-364-2665 or e-mail to quality@heartwareinc.com.Customers with questions were instructed to contact their HeartWare Representative or HeartWare Customers Service.For questions regarding this recall call 305-364-1575. |  | 76 | 76 (US) and 82 (International) | Worldwide Distribution - USA including AL, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. Internationally to:Australia, Austria, Belguim, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zeland,Norway, PolandSingapore, South Africa, Switzerland, Turkey, and United Kingdom. | Open |  |  | 76 | Not_Computer
Z-1823-2013 | 65473 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Cardinal Steam Biological Indicator, SteriTec Product Code PL 218, Distributor Product Code T40360EZEZTest Steam Biological Indicator: Mesa EZTest Steam Biological Indicator is a self contained Biological Indicator intended for monitoring the efficacy of 121ÂŽC, 132ÂŽC, 134ÂŽC and 135ÂŽC steam sterilization cycles. Mesa EZTest has a recommended incubation time of 24 hours. | Cardinal Steam Biological Indicator | 2 | July 29, 2013 | 2013 | Steritec Products Mfc. Co., Inc. | Mesa Labs, based out of Bozeman, MT, informed SteriTec Products, Inc. of a product recall involving the EZTest Steam and Smart-Read Biological Indicators (BI) because the readout time is not in specification with the label claims. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, SteriTec, sent all customers recall notifications via e-mail. The e-mail included a .pdf file of the recall notification letter entitled "URGENT: MEDICAL DEVICE RECALL" dated June 11, 2013. The letter described the product, problem and actions to be taken. The letter also informed customers of MesaLabs, "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to its customers. The customers were instructed to notify SteriTec products of receipt of this notification via email at: Jon@SteriTecproducts.com or fax to: 303-660-4213 Attn: Director Regulatory Affairs; examine the lot number printed on the packaging components; place affected items in hold and do not sell or distribute these lots; record inventory of all items remaining in your warehouse and all items that have been distributed; advise SteriTec of the status of all inventory you received used or unused, if you have distributed the affected product to other locations within your healthcare system, notify them immediately of this recall, ask them to provide inventory status and proceed with destruction or return via instructions provided in letter-If you have shipping questions, please contact SteriTec directly at 303-660-4201, or email us at: Info@SteriTecproducts.com, and also provide this notice to anyone in your facility that needs to be informed.Should you have any questions or concerns, please do not hesitate to contact Director of Regulatory Affairs and Quality Assurance or our offices at 303-660-4201. SteriTec's hours of operation are 7:30am Mountain Standard Time (MST) through 5:00pm MST. |  | 200 | 200 | Worldwide distribution: US (Nationwide) and countries of: Argentina, Australia, Canada, British Virgin Island, Czech Republic, Italy, and Turkey. | Terminated | December 19, 2013 | 143 | 57665 | Not_Computer
Z-1816-2013 | 65669 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Large External Fixation, MR Conditional, Tube to Tube Clamp, Part Number 390.007, Lot Number 6800599Synthes Large External Fixation is intended for use to provide treatment for long bone and pelvic fractures that require external fixation. | Synthes Large External Fixation, MR Conditional | 2 | July 29, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered internally that there was a discrepancy between the insert provided with the Tube to Tube Clamp and what was called for in the Bill of Materials. Following a detailed investigation, it was determined that the incorrect insert was used which indicates that the device is MR Safe, when it is actually MR Conditional. | PRODUCTION CONTROLS: Packaging | Synthes sent a Urgent Notice Medical Device Recall letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is requesting that you please examine your inventory for products with the above Part and Lot number and remove them. Call Synthes at 1-800-479-6329.If you have any further questions please call 610-719-5450. |  | 8 | 8 | US Distribution including the states of CA, MN, NC and TN. | Open |  |  | 8 | Not_Computer
Z-1821-2013 | 65668 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Synapse SD Drill Guide, Part Number SD389.477, Lot Number 6913900Indicated for Hooks, Plate/Rods, Rods and Screws when intended to promote fusion of the cervical spine fusion of the cervical spine and occipitocervical junction (occiput-T3), the plate/rod, rod, hook and screw (3.2 mm cortex) components of the Synthes Cervifix; Axon and Synapse Systems are indicated for the following: -Degenerative Disk Disease (DDD) (defined as neck pain of discogenic origin with degeneration of the disc as confirmed by patient history and radiographic studies); -Spondylolisthesis; -Spinal Stenosis; -Fracture/dislocation; -Atlantoaxial fracture with instability; -Occipitocervical dislocation; -Revision of previous cervical spine surgery; and -Tumo | Synapse SD Drill Guide | 2 | July 29, 2013 | 2013 | Synthes USA HQ, Inc. | A complaint was received from the field, that the drill guide was etched incorrectly, 12mm etch instead of 18mm drill guide. The firm identified that this constituted a field action and retrospectively reported this action. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, Synthes USA, sent an email to its customer. The email identified the product, problem and the actions to be taken. The customer was instructed to do the following: do not use this product; do not remove any product from your account or return any product without further instructions. If you have identified that you have the product, please have the product located in an area that ensures it will not be used. Further instructions will be sent out with regards to any action that needs to be taken next.Please call Synthes if you have any questions at 610-719-5364. |  | 2 | 2 | US distribution: PA only.. | Open |  |  | 2 | Not_Computer
Z-1818-2013 | 65718 | monitor,physiological,patient(without arrhythmia detection or alarms) | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | ViSi Mobile Chest Sensor (3 lead-wire ECG, AAMI), Catalog No. 93-10014The ViSi Mobile Monitoring System is intended for use by clinicians and medically qualified personnel for single or multi-parameter vital signs monitoring of adult patients. | ViSi Mobile Chest Sensor (3 leadwire ECG, AAMI), | 2 | July 29, 2013 | 2013 | Sotera Wireless, Inc. | Sotera Wireless, Inc. is recalling the ViSi Mobile Monitoring System due to the potential for the Chest Sensor to cause a blister wound. | DESIGN: Device Design | Sotera sent an Urgent Product Recall letter dated May 23. 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to report any adverse reactions experienced with the use of this product and or quality problems to the FDA's MedWatch Program by: www.fda.gov/medwatch, (800) FDA-1008, fax (800) FDA-0178. Customers were also instructed to contact Sotera Wireless Technical Support at (866) 794-5526 for questions. |  | 75 | 75 units | US distribution including CA and UT. | Terminated | September 09, 2013 | 42 | 165 | Not_Computer
Z-1819-2013 | 65068 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Custom Spine ISSYS LP Inverted Screw System. 8.5MM X 45 MM ISSYS LP Screw Rx Non-Sterile. Custom Spine, Inc. 9 Campus Dr. Parsippany, NJ 07054 USA. TEL: 973-808-0019The ISSYS LP Spinal Fixation System is intended to help provide immobilization and stabilization of spinal segments as an adjunct to fusion of the lumbar and/or sacral spine for pedicular and non-pedicular fixation. | Custom Spine ISSYS LP Inverted Screw System | 2 | July 29, 2013 | 2013 | Custom Spine, Inc. | On March 28, 2013, Custom Spine, Inc. issued a minor removal of a lot of 18 screws that were found to be out of specification. The screw length was 5 mm shorter in length than the required specification for the screw size. The product is laser marked 8.5mm X 45mm, but its length measured 40mm. |  | Custom Spine initiated a voluntary removal On 3/28/2013, Custom Spine initiated a voluntary removal of all 10 pieces of the ISSYS LP (8.5mm X 45mm screws) Lot 363. Distributors were contacted via phone and/or email and were asked to immediately return these parts to Custom Spine. |  | 18 | 18 units | Nationwide Distribution including Nebraska and Virginia. | Terminated | October 24, 2013 | 87 | 18 | Not_Computer
Z-1815-2013 | 63935 | processor, tissue, automated | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Leica ASP6025 Vacuum Tissue ProcessorThe Leica ASP6025 is a modular tissue processor for the following laboratory applications: fixation, dehydration, infiltration with intermedium, and the paraffin infiltration of histological tissue specimens. | Leica ASP6025 Advanced Smart Processor Vacuum Tissue Processor | 2 | July 29, 2013 | 2013 | Leica Microsystems, Inc. | The Instruction for Use 1v9 Ref F and all previous versions specify incorrect dimensions for the Ready To Use (RTU) bottles from other suppliers than Leica. In case of a pressure failure of the device in combination with a bottle with incorrect dimensions, overflow can cause contamination of reagents. Additionally, the current instructions for use does not highlight the importance to check the fil | DESIGN: Device Design | Leica Biosystems sent an Urgent Field Safety Notice dated December 12, 2012, to all affected customers. The notification included a description of the problem and a recommendation that the personnel who use the equipment only use Leica Ready To Use (RTU) bottles that are already supplied as a standard delivery item with the ASP6025 machine. The personnel are further reminded that reagent bottle levels need to be checked and corrected when necessary. Customers with questions or questions regarding this recall call 847-405-5413. | Contact the recalling firm for information | 128 | 128 machines | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Brazil, Chile, France, Germany, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands, Philipines, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, and United Kingdom. | Open |  |  | 128 | Not_Computer
Z-1830-2013 | 65576 | fluorometer, for clinical use | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA Analyzer Pipette Tips used on Tosoh AIA-2000, AIA-1800, AIA-900 and AIA-600II Analyzers, 1000 tips/bagAIA-Analyzer Tips are disposable pipette tips packaged in bulk quantities and used on Tosoh Bioscience AIA Analyzers for lab analysis of various immunoassays. | AIA Analyzer Pipette Tips | 2 | July 30, 2013 | 2013 | Tosoh Smd Inc | Visual inspection of the pipette tips indicated that some of the tips were not uniform in size/length which was causing error messages on the analyzers. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Tosoh Bioscience sent an Urgent Recall Notification letter dated May 15, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to respond on the amount of product that was destroyed by completing the attached form and faxing it to 614-317-1941.Customers with questions were instructed to call Tosoh Technical Support at 1-800-248-6764.For questions regarding this recall call 614-317-1909. |  | 11309 | 11,309 bags | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.Internationally to Canada, Japan, Mexico, Chile, and Uruguay. | Terminated | November 27, 2013 | 120 | 11309 | Not_Computer
Z-1831-2013 | 65084 | albumin, antigen, antiserum, control | DCF | Albumin immunological test system. | Immunology | Immunology | 510(K) Exempt | uTASWako DCP, Lot TG340 2013.11, REF 995-60701, Wako Pure Chemical Industries, Ltd. 1-2 Doshomachi 3-chome, Chuo-Ku Osaka 540-8605 Japan Distributed by Wako Diagnostics Wako Life Sciences, Inc. 1025 Terra Bella Ave., Suto A. Mountain View, CA 94043 USA. The product is a single reagent cartridge run on an automated system and each assay performed on asingle disposable chip using microfluidic electrophoretic separation. The reagent cartridge contains various buffers and antibodies. | uTASWako DCP | 2 | July 30, 2013 | 2013 | Wako Life Sciences, Inc. | Product fails to meet the accuracy at the low end of the measuring range. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Wako Life Sciences, Inc. sent an email and issued a Recall Notice letter via FED EX on April 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue the use of these kits, fill out the attached RMA form, and return unused kits using Wako's FedEx account (provided on RMA form). Partially used reagent cartridges should be disposed of according to your local and state requirements. Wako Life Sciences will relplace your existing stock of uTASWako DCP kits with kits from a different log number that has been tested and has shown to not be affected. If you have any questions regarding the Recall, please submit your inquiries to liver@wakousa.com or call 443-535-0702. |  | 36 | 36 cartridges -100 tests per cartridge | Distribution to CA, GA, MN, NC, UT and the District of Columbia. | Terminated | August 19, 2013 | 20 | 36 | Not_Computer
Z-1846-2013 | 65646 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAMMA3 System Long Nail Kit, R2.0, Ti, LeftDistributed in USA by Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 USAA single use device intended for the temporary fixation, correction of stabilization of the right femur. The kit consists of intramedullary nail and set screw. The kit is used in combination with lag screw and (distal) locking screws and end cap as well. | Stryker | 2 | July 31, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Osteosynthesis has become aware that LEFT GAMMA3 nails are marked as right nails. | PRODUCTION CONTROLS: Process Control | Stryker sent notification letters and product accountability forms dated May 21, 2013. via Fed Ex with return receipt to all branches/agencies and on May 22, 2013, via Fed Ex with return receipt to all hospital risk management, chief of orthoapedics, and surgeons.Our records indicate that you may have received and/or used the above referenced product(s). It is Stryker'sÂŽÂŽ responsibility as the manufacturer to ensure that customers who may have received and/or used these affected products also receive this important communication.Please contact your StrykerÂŽÂŽ Sales Representative to arrange for return of the product if you have any of the above listed items. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form as soon as possible to 855- 251-3635.Please note that your signature on the following form only confirms that you received this notification and does not obligate you to take any additional action beyond what is called for in this notification letter.We regret any inconvenience this action may cause you and if you have any questions, please call (201) 972-2100. | Contact the recalling firm for information | 5 | 5 units | US Distribution including the states of IL, KY and TN. | Open |  |  | 5 | Not_Computer
Z-1836-2013 | 63108 | stimulator, spinal-cord, totally implanted for pain relief | LGW |  |  | Neurology |  | Eon Mini - 3788 (IPG), Eon Mini system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back, surgery syndrome, intractable low back and leg pain. | Eon Mini | 2 | July 31, 2013 | 2013 | St. Jude Medical | As part of St. Jude Medical's routine tracking and product monitoring, they have received a total of 214 reports, out of 34,617 Eon Mini IPGs, that lost the ability to communicate or recharge due to an inner battery weld issue resulting in loss of pain relief and subsequent explant. Explant surgery, as with any surgery, presents a risk to patient health. Adverse events associated with an unplanne | DESIGN: Device Design | St. Jude Medical sent a Important Medical Device Recall Advisory dated July 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To further assist in your patient care we are providing you with a list of all serial numbers we show have been distributed to you. ( See Attachment A ).Further questions please call (972) 309-2154. | Contact the recalling firm for information | 34617 | 34,617 | Worldwide Distribution--USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Belgium, Brazil, canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Saudi, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open |  |  | 34617 | Not_Computer
Z-1832-2013 | 65568 | driver, surgical, pin | GFC | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Exactech Articulated Driver A nonpowered hand-held device intended for medical purposes to manipulate tissue, or for use with other devices in orthopedic surgery. | Articulated Driver | 2 | July 31, 2013 | 2013 | Exactech, Inc. | When torque is applied to the Articulated Driver at the maximum angle, the cross pin may shift out of position, potentially leading to disassembly of the instrument. |  | Exactech sent a" Important Product Recall Notice" dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to Immediately cease distribution or use of the products. Extend this information to your accounts that may have this product in their possession. Verify if you have any of the subject ARTICULATE DRIVER (catalog number 181-00-04) in the specified lots. FAx back the attached form.Further questions please call 1-800-392-2832. | Contact the recalling firm for information | 57 | 57 | US Distribution including the states of AR, CA, FL, NJ, NY, OK, and TX | Terminated | September 19, 2013 | 50 | 57 | Not_Computer
Z-1820-2013 | 65591 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Lower Extremity Dovetail Intercalary component.Intended for use in replacement of the mid-shaft portion of the femur. | Dovetail Intercalary component | 1 | July 31, 2013 | 2013 | DePuy Orthopaedics, Inc. | The LPS Lower Extremity Dovetail Intercalary component has the potential for fracture of the female component when exposed to certain physiological loads. | DESIGN: Device Design | On 7/11/13, URGENT MEDICAL DEVICE RECALL NOTIFICATIONS were sent to consignees: Depuy Distributors, hospitals, and surgeons. End-user Hospital Customers will submit reconciliation forms to DePuy. Possible clinical implications of device failure are stated in these notifications, including poor mechanics and loss of function, pain, component dislocation, adverse tissue reaction, bone damage, potential need for revision surgery and the risks associated with this type of surgery. A patient letter will be included with the surgeon letter to assist in surgeon's notification and discussion with patients. The devices will be returned to DePuy Warsaw through the normal DePuy Returns process, to attention of Returns and marking H13-13 on the outside of the box.If a patient presents with a fractured LPS" Lower Extremity Dovetail Intercalary component with well-fixed proximal and distal stems and the surgeon determines LPS" Lower Extremity Dovetail Intercalary component is the best treatment option, DePuy plans to make the LPS" Lower Extremity Dovetail Intercalary component available through authorized approvals. Directions on how to purchase the component are provided in the Recall Communications.Any questions or concerns about the recall, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8am-5pm EDT). | Contact the recalling firm for information | 58 | 58 U.S. / 39 OUS | Worldwide Distribution-USA including the states of FL, GA, ID, KS, KY, MI, MO, NC, NV, NY, OH, OK, PA, TN, TX, VA, and WA and the countries of AUSTRALIA, CANADA, ENGLAND, GERMANY, INDIA, COUNTY CORK, IRELAND, and TURKEY. | Open |  |  | 58 | Not_Computer
Z-1869-2013 | 64987 | tray, surgical, instrument | FSM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Caddy - PedFuse, Screw Caddy, 8MM and 9MM, Part Number: 11-80009 Rev. BThe caddy holds, in an organized fashion within the case, 8mm and 9mm pedicle screws used during implantation of pedicle screw system devices intended to provide immobilization and stabilization of spinal segments of skeletally mature patients | Screw Caddy, 8MM adn 9MMPedFuse | 2 | August 01, 2013 | 2013 | SpineFrontier, Inc. | Two sections that are designated for the 8mm and 9mm screws were switched; the holes that hold the 8mm screws are designed for 9mm screws, and the holes that hold the 9mm screws are designated for 8mm screws. | DESIGN: Device Design | The firm, SpineFrontier, sent an Advisory Notice entitled "URGENT May 1, 2013, PedFuse Screw Caddy Removal" dated May 1, 2013 to their consignees/customers. The notice described the product, problem and action to be taken. Consignees/customers were informed that a removal has been initiated to remove the parts from the field to ensure that all sets contain correctly marked screw caddies. Consignees/customers wre instructed to complete and return the URGENT PedFuse Screw Caddy Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email: customerservice@spinefrontier.com. Replacement caddies will be shipped upon receipt of completed acknowledgement form and RMA request for part return.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. | Contact the recalling firm for information | 11 | 11 caddies | US distribution in states of: CO, DC, MA and TX. | Open |  |  | 11 | Not_Computer
Z-1864-2013 | 64625 | bronchoscope accessory | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Inspira AIR Balloon Dilation system.Size 12 x 40 mm, Syringe volume 12 cc.Dilation of airway tree. | Inspira AIR Balloon Dilation system | 2 | August 01, 2013 | 2013 | Acclarent, Inc. | Labeling correction for all sizes of the Inspira AIR Balloon Dilation System to include additional language and warnings as well as to clarify some instructional steps. | DESIGN: Device Design | Urgent Medical Device Labeling correction letters will be sent to all consignees through a delivery system that can track delivery. After CDRH review and corrections, Acclarent began issuing customer notification on August 7, 2013. |  | 1217 | 1217 | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, and Saudi Arabia. | Terminated | October 25, 2013 | 85 | 14330 | Not_Computer
Z-1867-2013 | 65544 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Mesa EZtest BIOLOGICAL INDICATOR MONITORING SYSTEM. Intended for use in determining the efficacy of steam sterilization processes. | EZTest, Ensure, ETIGAM BrightCheq, ParaSure, ConFirm 24 & 10, BT Sure, Assure Accufast, Valisafe | 2 | August 01, 2013 | 2013 | Mesa Laboratories, Inc., Bozeman Manufacturing Facility | Mesa EZTest , Self-contained steam biological indicator, are recalled because Mesa recently determined that the read out time is not in specification with the label claims. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Mesa Labs sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions contact Mesa Labs at BI-Support@mesalabs.com. | Contact the recalling firm for information | 2211 | 865/25 ct boxes, 414/100 ct boxes, 307/500 ct boxes | Worldwide Distribution: US (nationwide) and countries of: Australia, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Egypt, Greece, India,Israel, Japan, Jordan, Malaysia, Mexico, Milanese, Netherlands, Peru, Portugal, Pozan, Romania, Spain, Sweden, Taiwan, Thailand, United Kingdom, and Uruguay. | Terminated | November 22, 2013 | 113 | 4758 | Not_Computer
Z-1872-2013 | 65724 | marker, skin | FZZ | Skin marker. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Non-sterile Viscot Surgical Skin MarkersPart Number 1437-100 Non-sterile Viscot Part Number 143-100 Surgical Skin Markers are pen-like devices intended to be used to write on the patient's skin, e.g., to outline surgical incision sites or mark anatomical sites for accurate blood pressure measurement; available in nonsterile configuration. | Nonsterile Viscot Part Number 1437100 Surgical Skin Markers | 2 | August 02, 2013 | 2013 | Precision Dynamics Corporation | PDC Healthcare initiated this recall of Non-Sterile Viscot Part Number 1437-100 Surgical Skin Markers, becuase they were inadvertently mislabeled as PN STER-VIS Sterile Skin Marker Pen with Labels and Flex Ruter. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Precision Dynamics Corp. (PDC) sent a Urgent: Skin Marker Pen Recall letter dated July 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All affected customers have also been contacted by phone and email as of July 12, 2013We want to take every precaution to ensure patient safety.1. Immediately examine your inventory2. Quarantine the product subject to recall3. Complete and return the enclosed response form as soon as possibleYour PDC Healthcare customer service representative will contact you directly to arrange the return of the product. In addition, if you may have distributed this product, please identify your customers and notify them at once of this product recall. Your customer notification may be enhanced by including a copy of this recall notification letter.If you have any questions please call the PDC Healthcare Returns Department at 800-435-4242 ext. 5003.Thank you for your understanding and assistance related to this matter. |  | 1200 | 1,200 pens | US Distribution including the states of SC, TX, TN, WA, OK. | Terminated | January 06, 2014 | 157 | 1200 | Not_Computer
Z-1871-2013 | 64985 | guide, surgical, instrument | FZX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Drill Guide - Chameleon FacetFuse Fixation System, Model: 11-00028 Rev. AFacetFuse Drill Guides were designed to provide a concentric pathway for the drills and taps to access the bone while protecting the soft tissue from the cutting geometry of the drills/taps. In addition the drill guide is used as a reference to measure depth that the drills and taps are protruding into bone, which is measured by depth markings at the proximal end of the drills/taps relative to the proximal end of the drill guide. These devices were used during implantation of spinal intervertebral body fusion devices intended for intervertebral body fusion of the spine of skeletally mature patients. | Drill Guide (guide, surgical, instrument) | 2 | August 02, 2013 | 2013 | SpineFrontier, Inc. | Prior to use in surgery, the Drill Guides were presenting noticeable degrees of denting and/or bending. As a result, drills and taps could not be directed down guides during surgery due to mechanical interference. | DESIGN: Device Design | Spine Frontier sent an Advisory Notice dated November 9, 2012, via email to all affected customers. Customers were advised that their FacetFuse Drill Guides need to be replaced. Customers were instructed to write their number ( included in the email) on the outside of the package and provide us with the tracking information. A replacement part has been sent out to you.If you have any further questions please call (978) 232-3990. | Contact the recalling firm for information | 13 | 13 drill guides | US Distribution including Puerto Rico and the states of CO, KS, MA and TX. | Open |  |  | 13 | Not_Computer
Z-1873-2013 | 63605 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer, All Serial Numbers.An automated chemistry analyzer that measures analytes in samples, in combination with appropriate reagents, calibrators, quality control (QC), material and other accessories. | Beckman Coulter AU5800 Clinical Chemistry Analyzer | 2 | August 02, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the AU8500 Clinical Chemistry Analyzer because there is no labeling on the diluted detergent tank to show the proper disposition of the on/off valve. | DESIGN: Software Design | A recall notification letter was forwarded to all their customers on 11/1/12 who purchased the AU5800 Clinical Chemistry Analyzer. The customers were informed of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Reponse form within 10 days. Customers with questions are instructed to call (800) 854-3633 or contact their local BCI Representative. |  | 338 | 338 units (11 units in the US) | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Germany, China, Austria, Brazil, Italy, Argentina, Korea, Vietnam, Hungary, Ireland, France, Spain, Turkey, Switzerland, Israel, Australia, India, United Kingdom, Netherlands, Malaysia, Russian Federation, Japan, Taiwan, and Canada | Terminated | October 18, 2013 | 77 | 338 | Not_Computer
Z-1875-2013 | 65605 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobiletta Mira Mobile X-ray System | Mobilett Mira Mobile Xray | 2 | August 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Alert about the potential risk to users when operating the Mobilett Mira. When a Mobilett Mira is regularly exposed to extensive mechanical strain (driven over high thresholds for doors or elevators), it is possible that one of its main wheel shafts could break because of material fatigue, resulting in the chassis falling on the ground on the affected wheel side. This | DESIGN: Device Design | Siemens sent a Field Safety Notice dated May 2. 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to :1) Never step or sit on the unit or otherwise put substantial weight on it.2) Avoid moving the unit over high thresholds for doors or elevators where possible, or do so carefully and slowly. 3) Make sure that no person had his/her extremities under the chassis, also when the unit is not moving. Update instructions were also released to alert customers of the potential risk and to provide a hardware change which is intended to improve corrosion resistance.An acknowledgement of receipt form was also sent out for the field safety notice and requested to be returned to Siemens.For further questions please call (888) 826-9702. |  | 39 | 39 | US Distribution including the states of IA, SD, NC, WI, OH, NE, TX, KS, ID, MI, GA, AR and IL. | Open |  |  | 39 | Not_Computer
Z-1851-2013 | 65689 | set, blood transfusion | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | List No. 14203-28; BLOOD SET; 200 Micron Filter, 80 Inch, Non-DEHP; For Blood Product Administration; Made in Costa Rica; Product Inquiries should be directed to Hospira., Inc., Lake Forrest, IL 60045 USAHospira Blood sets are designed for Blood Product Administration | Blood Set | 1 | August 05, 2013 | 2013 | Hospira Inc. | It was discovered that there was a potential for the piercing pin on certain Hospira blood sets to puncture the outer wall of blood bags during insertion of the pin into the blood bag. | DESIGN: Device Design | Hospira sent an Important Safety Information letter dated April 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exercise EXTREME CAUTION when using Hospira blood bags and to make sure that all instructions for use are included with the blood bag and their facility's protocol for spiking blood bags are completely followed in order to minimize the possibility of puncturing the outer wall of the blood bag. Customers with issues were instructed to call Hospira at 1-800-441-4100 or email ProductComplaintsPP@hospira.com. Customers were asked to complete the attached replay formand return it to the fax number on the form. If customers distributed the product further they were instructed to notify their accounts that received the affected product and ask them to contact Stericycle at 1-866-608-3936 to receive a reply form. Customers with questions were instructed to contact Hospira at 1-800-441-4100.For questions regarding this recall call 224-212-2000. |  | 1152 | 1,152 units | Worldwide Distribution - USA (nationwide) and Puerto Rico and Internationally to Costa Rica. | Open |  |  | 701870 | Not_Computer
Z-1876-2013 | 65757 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Mullins Transseptal Adult 7F, REF 008551.Intended for the introduction of cardiovascular catheters into the left side of the heart. | Mullins Transseptal Catheter Introducer Sheath | 2 | August 05, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic identified a specific labeling issue for specific lots of Mullins Transseptal Catheter Introducer Sheath, model 008551, in which packages labeled as Mullins 7F (French) may actually contain an 8F(French) dilator. Medtronic is requesting the return of potentially mis-labeled Mullins Transseptal Catheter Introducer 7 French Sheaths. Due to the differences in size, an 8F dilator will not | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | Consignees were sent on 7/8/2013 a Medtronic, Inc " Urgent Medical Device Recall" letter dated 8 July 2013. The letter was addressed to Healthcare Professionals. The letter identified the problem and the product involved in the recall. Advised consignees to remove and quarantine the product from their inventory and return the product. Requested consignees to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. If the product has been further distributed advised consignees to please notify them of the recall. |  | 269 | 269 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, AK, CA, DE, GA, IL, KY, MA, MI, MN, MS, NV, NY, NC, OH, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CHILE, INDIA, and SINGAPORE. | Open |  |  | 269 | Not_Computer
Z-1881-2013 | 65162 | radioimmunoassay, thyroid-stimulating hormone | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra ReadyPack ((REF 06491072 - 100 Tests, REF 06491080 - 500 Tests, REF 04862625 - 2500T).Intended to quantitatively measure human thyroid stimulating hormone (TSH, thyrotropin). | ADVIA Centaur Systems TSH3 Ultra ReadyPack | 2 | August 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | A rare variant of TSH, identified in a small cluster of patients, is not detected. | DESIGN: Device Design | On May 10, 2013, an Important Safety Information notice was sent to all affected Siemens Healthcare Diagnostic customers in the United States. On May 10, 2013, an Urgent Field Safety Notice was sent to all customers outside of the United States. These notices inform the customer that there is a rare variant of TSH, identified in a small cluster of patients that is not detected by some Siemens assays. This notice informs customers that they should continue to use the affected TSH assays as the observed rate of occurrence during a 30-month time period was 0.6 X 10 -7 and there are strong mitigating factors. These mitigating factors include co-interpretation of TSH results in conjunction with thyroxine and TS, patient history and clinical signs and symptoms. Although Siemens believes a look back is not required as the expected frequency of the variant is extremely rare, these notices state that this decision should be made by the medical director. | Contact the recalling firm for information | 172913 | 172913 | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AR, AT, AU, BA, BD, BE, BG, BO, BR, BY, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HN, HR, HU, ID, IE, IL, IN, IR, IS, IT, JO, JP, KA, KR, KW, KZ, LB, LK, LT, LV, LY, MA, MK, MT, MX, MY, NI, NL, NO, NP, OM, PA, PE, PH, PK, PL, PT, PY, QA, RO, RU, SA, SE, SG, SK, SV, SY, TH, TM, TN, TR, TW, UA, UY, UZ, VE, VN, XS, and ZA. | Open |  |  | 330020 | Not_Computer
Z-0834-2013 | 64244 | slide stainer, automated | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VENTANA (a member of the Roche Group) DISCOVERY ULTRA; REF/MU 750-601; Research Use OnlyProduct Usage:Research Use Only-The DISCOVERY ULTRA Advanced Staining System is intended to automatically stain histological or cytological specimens on microscope slides with specific inununohistochemistry or in situ hybridization reagents. | Discovery Ultra Advanced Staining System Research Use Only (RUO) | 2 | August 06, 2013 | 2013 | Ventana Medical Systems Inc | Potential for a waste fluid overflow condition to accur while running decontamination cycles on the BenchMark ULTRA automated staining platform or the DISCOVERY ULTRA automated staining platform.The waste fluid drainage system may become restricted by build-up of debris in the tubing and in-line filter, thereby inhibiting optimal flow rates. This type of overflow event can lead to migration of flu | DESIGN: Device Design | The firm, Roche sent a "Technical Bulletin" and an letter dated October 3, 2012 to its customers. The letter and bulletin described the product, problem and actions to be taken. The customers were informed that Ventana is aggressively working on solutions to this issue. Until mitigations to the overflow condition are available for implementation; inform all personnel in your organization that work with the BenchMark ULTRA or DISCOVERY ULTRA instrument about the potential for fluid overflow to occur during decontamination cycles; follow the instructions provided in the letter to avoid the documented failure mode and follow the instructions provided in the technical bulletin.Contact your local support center with any questions or concerns or call 520-877-7191. | Contact the recalling firm for information | 54 | 54 instruments | Worldwide distribution: USA (nationwide) and country of: Germany. | Open |  |  | 1789 | Not_Computer
Z-1922-2013 | 65781 | tube, feeding | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nestle Nutrition, Compat, Enteral Feeding Tube, sterile, 8 French (REF 08310400), 10 French (REF 08310600), and 12 French (REF 08310500). Product Usage:The Compat Nasogastric Feeding Tube is indicated for the administration of nutrition, fluids and medications into the stomach r small bowel by the naso-enteric, route for patients who have an intact gastrointestinal tract but are physically unable to ingest enough nutrients orally to meet their current needs. | Compat Nasogastric Feeding Tube | 2 | August 06, 2013 | 2013 | Nestle HealthCare Nutrition | Nestle HealthCare Nutrition, Inc. is initiating a recall on some of Compat NG Tubes because they have been delivered in an insufficiently sealed pouch, thus compromising the sterility claimed on the label. | PRODUCTION CONTROLS: Packaging Process Control | Nestle Health Science sent a "Medical Device Recall" letter dated July 24, 2013. The letter was addressed to Risk Manger. The letter described the product involved, problem and actions to be taken. Customers were advised that products currently in use should not be returned. Customers were requested to examine their inventory, quarantine the product, carry the recall to the user facility level and to complete and return the Acknowledgment / Response Form.For questions call Nestle HealthCare Nutrition Customer Service at 1-877-463-7853. | Contact the recalling firm for information | 37310 | 37,310 | US Nationwide Distribution - including the states of: AK, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA. | Open |  |  | 37310 | Not_Computer
Z-1896-2013 | 65819 | closure, wound, adhesive | MGO |  |  | General & Plastic Surgery | 510(k) | Walgreens ("W" and "Well at Walgreens") Wound Closure Adhesive Surgical Tape Strips 30 strips 1/4 in. x 4 in. Sterile latex freeUPC # ("W") 311917111728 item # 479241UPC # ("Well at Walgreens") 311917143354 item # 156744Used to secure, close and support small cuts and wounds. | Walgreens ("W" and "Well at Walgreens") Wound Closure Adhesive Surgical Tape Strips | 2 | August 06, 2013 | 2013 | ASO, LLC | This recall has been initiated by Aspen Surgical Products, Inc., due to the inability to confirm sterility. | PRODUCTION CONTROLS: Process Control | ASO sent a letter on July 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine your inventory and stop using/distributing the product. Place in quarantine to be returned or destroyed. ASO customers will receive credit for returned product or a returned completed Certificate of Destruction. Please use Return Authorization Number 1675 when returning the product.Please complete and return the enclosed response form as soon as possible. Return the form to:Address:ASO LLCRegulatory Affairs300 Sarasota Center Blvd.Sarasota, FL 34240Fax:941-378-9040Attn: Regulatory AffairsFor questions regarding this recall call 888-364-7004, ext 146. |  | 2689560 | Total 2,689,560 pouches (2 cards/pouch) | Nationwide Distribution including FL, IA, IL, MI, NC, NY, OH, OR, PA, RI, TX, WA and WI. | Open |  |  | 2689560 | Not_Computer
Z-1925-2013 | 65152 | calibrator, secondary | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Mass Creatine Kinase MB Isozyme Calibrator (RC420)The Mass Creatine Kinase MB isozyme Calibrator is an in vitro diagnostic product intended to be used to calibrate the Creatine Kinase MB Isozyme method for the Dimension(R) clinical chemistry system with heterogeneous immunoassay module. | Dimension Mass Creatine Kinase MB Isozyme Calibrator | 2 | August 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints that the lyophilized cake is not fully dissolved after the stated time in the MMB calibrator Instructions For Use (IFU). The frequency of this occurrence is low but if it occurs, gel-like clumps may be observed. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated May 2013, to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Until a permanent solution is identified, please follow the instructions listed in the letter in order to optimize calibration preparation; complete and return the attached Correction Effectiveness Check form via fax to the Siemens Technical Solutions Center at (302) 631-8467 and forward this notification to anyone to whom you may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center or your local Siemens technical representative or call 800-441-9250. | Contact the recalling firm for information | 7470 | 7,470 | Nationwide distribution: US including states of : AK, AL, AR, AZ, CA, CO, CT, DC, GA, GU, FL, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Open |  |  | 7470 | Not_Computer
Z-1924-2013 | 65706 | enzyme immunoassay, tracrolimus | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex reagent cartridge (DF107) and Dimension TACR CAL (DC107)Product Usage: The TACR method is an in vitro diagnostic test intended to quantitatively measure Tacrolimus in human whole blood on the Dimension clinical chemistry system. Measurements of Tacrolimus are used as an aid in the management of tacrolimus therapy in kidney and liver transplant patients. | DimensionTACR (DF107) and Dimension Tacrolimus Calibrator (DC107) | 2 | August 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens initiated a recall due to confirmed complaints of low patient sample recovery with Dimension TACR Flex Reagent cartridge (DF107) lot BB4087 and its linked calibrator (DC 107) lot 3BD029. Falsely low Tacrolimus results across the assay range may lead to alterations in dosing. |  | Siemens sent an Urgent Medical Device Recall letter dated July 2013 to affected customers The letter identified the affected product, problem, potential risk, and actions to be taken by customers. Customers were requested to discuss the letter with thier laboratory director to determine if additional follow up is appropriate for patient tests conducted with the affected device lots. Additionally, customers were asked to discard any remaining inventory and to complete and return the field correction effectiveness check for replacement product. They were also asked to forward the letter to anyone wha had received this affected product.For questions contact your Siemens Technical Solutions Center or your local Siemens Technical support representative. |  | 5423 | 5423 | USA Nationwide Distribution including the states of: TX, NC, GA, TN, UT, NJ, MI, VA, AR, LA, MO, FL, AZ, CA, DC, PA, NH, WI, SD, SC, NM, MD, and IN. | Open |  |  | 5423 | Not_Computer
Z-1887-2013 | 65482 | filter, intravascular, cardiovascular | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Retrieval Inferior Vena Cava (IVC) Filter***REF Cat. No.***466F220A and 466F220B***Vena Cava Filter and Introduction Kit***(BRITE TIP Catheter Sheath Introducer and Angiographic Vessel Dilator)***Cordis***a Johnson and Johnson Company***ENDOVASCULAR***Label No. SB 10247519.2/***LB466F220AO.5***omoraoro/2012.050794***Do not resterilize***. Product Usage:The OPTEASE Vena Cava Filter is indicated for use in the prevention of recurrent pulmonary embolism (PE) via percutaneous placement in the vena cava in several situations: Pulmonary thromboembolism when anticoagulants are contraindicated; Failure of anticoagulant therapy for thromboembolic disease; Emergency treatment following massive pulmonary embolism where anticipated benefits of conventional therapy are reduced and Chronic, recurrent pulmonary embolism where anticoagulant therapy has failed, or is contraindicated. | OptEase Vena Cava Filter | 1 | August 07, 2013 | 2013 | Cordis Corporation | Correction to provide clarification and additional labeling to minimize likelihood on implanting filter backwards. | DESIGN: Labeling Design | Cordis Corporation sent an "URGENT Medical Device Correction" notification to their consignees in the US by letter dated April 3, 2013 and an e-mail was sent to the Canadian consignees. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to read the description and recommendations sections carefully, then sign and return the enclosed Acknowledgement Form directly to Cordis to their fax number(908) 429-8287 or Scan and e-mail to MiamiQA@crdus.jnj.com. The letter states to pass on the notice to anyone in the consignees' facilities that need to be informed. Additionally, the notification request consignees to maintain awareness of this notification until the information has been incorporated in the Cordis OPTEASE Vena Cava Filter labeling. For questions related to the issue contact Cordis Endovascular Medical Information Line 877-338-4235. | Contact the recalling firm for information | 33000 | 33,000 units | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, Ny, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Hawaii and Puerto Rico and the countries of Canada & Panama | Open |  |  | 33000 | Not_Computer
Z-1923-2013 | 65322 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Multiple Sample and Direct Draw Luer AdapterBD Franklin Lakes, NJ 07417 USAA sterile non-invasive device used to connect venous access devices such as needles, blood collection sets, ad infusion sets to blood collection tubes. | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter | 2 | August 07, 2013 | 2013 | Becton Dickinson & Company | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter may have blood leakage and low draw volume. |  | BD sent a Product Advisory Notice dated May 3, 2013, via UPS to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard filled tubes since it is possible they will contain less blood volume (low draws), replace the Luer Adaptor unit being used with a fresh unit, and continue to draw blood samples using the fresh unit. Customers with questions were instructed to contact BD Technical Service at 1-800-631-0174, Option 4, or contact them through website: www.bd.com/vacutainer/contact.For questions regarding this recall call 201-847-6800. | Contact the recalling firm for information | 37906400 | 37,906,400 units | Nationwide Distribution | Open |  |  | 37906400 | Not_Computer
Z-1932-2013 | 64993 | instrument, manual, surgical, general use | MDM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | T-Handle, Strike Plate, Push Connection, Part Number: SI50023 Rev AT-handles are silicone handled instruments that have a quick release end. The silicone T-handle has an impactor cap on the proximal end of the handle which is designed to be struck with a mallet. T-handles are used primarily to attach to, and operate a wide range of instruments used during implantation of spinal intervertebral body fusion devices intended for intervertebral body fusion of the spine of skeletally mature patients. | THandle | 2 | August 08, 2013 | 2013 | SpineFrontier, Inc. | During an operation the strike plate of a T-Handle loosened and separated from the T-Handle. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Vendor Change Control | On August 29, 2011, The Supply Chain Team was notified of the recall via email. They were informed that all handles are being pulled from the field. This resulted in the return of all T-Handles to SpineFrontier.Further questions please call (978) 232-3990. |  | 11 | 11 | US Distribution including the states of MA, NV and TX. | Terminated | October 18, 2013 | 71 | 11 | Not_Computer
Z-1926-2013 | 64984 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Rev E PedFuse Return Screw Inserter, Part Number: 11-80030Screw inserters are instruments which capture, hold and drive pedicle screw assemblies. There is a capturing feature that engages and holds the screw in place. There is also a feature used to lock the screw to the inserter before insertion of the screw. They can also be used to remove screws or back screws out. Screw inserters are used during the implantation of spinal pedicle screws intended to provide immobilization and stabilization of spinal segments in skeletally mature patients. Used in conjunction with Pedicle Screw Spinal system marketed as KRD1 PedFuse Pedicle Screw System . | Screw inserter (screwdriver orthopedic manual surgical instrument) | 2 | August 08, 2013 | 2013 | SpineFrontier, Inc. | Inserters have a set screw that fixes a collar to the inserter. The screw set can loosen and fall out of the inserter which allows the collar to slide off the proximal end of the inserter. When the collar is not in place the locking sleeve can also slide off the inserter. | DESIGN: Device Design | Spine Frontier sent an email dated December 10, 2012, to the affected customers that a Pedfuse Inserter with part number PM3979 needed to be returned.Further questions please call (978) 232-3990. |  | 8 | 8 | Distribution only in Texas. | Open |  |  | 8 | Not_Computer
Z-1933-2013 | 65653 | tray, surgical, instrument | FSM | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Indus Invue Screw Caddy, Model: SI70097The caddy holds, in an organized fashion within the case, a variety of shapes and sizes of plates and screws used during fixation of anterior cervical plates to provide stabilization during the development of spine fusions. There is an array of 108 holes that store 3 types of screws. O4.2mm Self-Drilling, O4.2mm Self-Tapping, and O4.5mm Self-Tapping screws which are separated by type and screw length. The caddy has features to hold the screws in place. Recall expanded on 8/21/2013 to include additional lots of this product. | Indus Invue Screw Caddy | 2 | August 08, 2013 | 2013 | SpineFrontier, Inc. | Difficulty removing screws from the Invue caddies. | DESIGN: Device Design | SpineFrontier sent a Urgent Invue Caddy Advisory letter dated July 10, 2013, to all affected customers. Customers were advised the removal of the Invue Screw Caddy, SI70097 from the field for modification. Customers were asked to return the parts to SpineFrontier and to complete an Acknowledgement Form indicating that they have read the notice and will initiate the immediate return of the parts noted above.Please feel free to contact us at (978) 232-3990.On 8/21/2013 the recall was expanded. |  | 13 | 13 | Worldwide Distribution - USA, including the states of CT, VA, MA, FL, TX, CO and MS. and the country of Jamaica. | Open |  |  | 13 | Not_Computer
Z-1928-2013 | 65370 | prosthesis, hip, semi-constrained, uncemented, metal/polymer, porous | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) XLPE ACETABULAR LINER UHMWPE, 20o OVERHANG, 36 MM ID, 62 MM OD, QTY: (1), STERILE H2O2, REF 71335762.Orthopaedic. | R3 (TM) XLPE ACETABULAR LINER UHMWPE, 20o OVERHANG, 36 MM ID, 62 MM OD | 2 | August 08, 2013 | 2013 | Smith & Nephew Inc | One batch of 60 mm 20o R3 XLPE liner may be laser etched and packaged as a 62 mm 20o liner and one batch of 62 mm 20o R3 XLPE liner may be laser etched and packaged as a 60 mm 20o liner. | PRODUCTION CONTROLS: Packaging Process Control | All affected Smith & Nephew Sales Representatives were notified of problem and the recall via email on 5/29. They were instructed to cease use and distribution of the product and quarantine for return. Surgeons were notified by letter beginning on 6/12. International consignees were notified on 05/29 via email. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 53 | 53 units | Worldwide Distribution-USA including the states of PA, WA, TX, FL, IA, MI, WI, TN, KY, IN, NY, MO and OR, and the countries of India, Canada, Chile, Venezuela, Dubai, South Africa, Australia, Poland, Romania, Belgium, Germany and United Kingdom. | Open |  |  | 53 | Not_Computer
Z-1927-2013 | 65444 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Split Driver, Assembly with Handle, Model SI70006 Rev G. Used in conjunction with the Invue Anterior Cervical Plate System under 501(k) clearance number K121060. The Invue Driver is an instrument consisting of a Radel handle, a shaft and hexalobe-like tip with a 4.0mm split, which allows the driver to collapse to fit into a screw and expand to retain the screw. | Invue Split Driver, Assembly with Handle | 2 | August 08, 2013 | 2013 | SpineFrontier, Inc. | Device design. | DESIGN: Device Design | SpineFrontier sent an URGENT Advisory Notice dated June 27, 2013, to all affected consignees regarding the use of Invue Drivers. The notice advises users that the Split Tip Stylet Driver Assembly (11-70099) and Sold Driver, AO Connection (11-70070 or 11-70077) are being shipped in place of the Split Driver Assembly w/Handle (SI70006). The uses of the replacement parts are explained. Consignees are instructed to complete the Acknowledgement Form and return it to SpineFrontier.On June 28, 2013, customers were sent a second Advisory Notice requesting the removal of the Split Driver Assy w/ Handle Part Number SI70006 Lot Numbers 027104 and 027105. Consignees were instructed to complete the Acknowledgement Form and return it to SpineFrontier.Consignees with questions were instructed to call 978-232-3990. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 12 | 12 devices | Nationwide Distribution including Texas, Ohio, and Missouri. | Open |  |  | 12 | Not_Computer
Z-1817-2012 | 61253 | instrument, ultrasonic surgical | LFL |  |  | General & Plastic Surgery | 510(k) | TX1 Tissue Removal System Console. Model: TX1TX1 Tissue Removal System is an ultrasonic aspirator that emulsifies and removes soft tissue. The TX1 Console provides control over each mode of operation, as well as the user interface via LCD touchscreen and audible tones for confirmation of user selections.Indiciated for use in surgical procedures where fragmentation, emulsification, and aspiration of soft tissue are desirable, including General Surgery, Orthopedic Surgery, Laparoscopic Surgery and Plastic and Reconstructive Surgery. | TX1 Tissue Removal System Console | 2 | August 08, 2013 | 2013 | American Optisurgical Inc | Two reasons for recall.1. Incorrect pressure relief valve installed inside console, leading to a remote chance of a risk to health. 2. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect. | PRODUCTION CONTROLS: Process Control | The firm, American Optisurgical Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter via email on February 23, 2012. An updated "URGENT: MEDICAL DEVICE RECALL" recall letter dated March 1, 2012 was sent via email on March 1, 2012 to its customer. This updated recall letter described the product, problem and actions to be taken. The letter notified the two reasons for recall: " An incorrect pressure relief valve has been installed inside the console. Should the regulator that controls pressure delivered to the inflation cuff fail, the pressure relief valve would not open until a pressure of 10psi was achieved, which may present a health risk". "The Operator's Manual included with the system, Rev. 1, has a typographical error. Specifically, in Section E. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect". The updated recall included the above second reason for recall. The firm instructed the customer, for replacement of the pressure relief valve, and to return the revision 1 operator manual to the company's address using a pre-paid FedEx account. For questions call 949.580.1266 M-F, 7:00am to 5:00pm PST. | Problem is exhibited on insertion. Device, after successful implantation, should perform as intended. | 5 | 5units | Nationwide distribution: MN only. | Terminated | June 15, 2012 | -419 | 5 | Not_Computer
Z-1938-2013 | 65698 | mesh, surgical, polymeric | FTL | Surgical mesh. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | C-QUR V-Patch Mesh (All sizes and shapes).Intended for use in soft tissue deficiencies. | CQUR VPatch Mesh | 2 | August 09, 2013 | 2013 | Atrium Medical Corporation | Additional Instructions for Use and Storage Conditions as Coated mesh can adhere to the inner packaging liner due to exposure to high humidity conditions |  | Atrium Medical issued Recall Letter via UPS and e-mail on 7/19/13 to the accounts and field representatives. The notification identifies the problem, product, and risk factors. If the Product is exposed to excessive humidity for an extended period of time, then the increased humidity occurring inside the pouch can potentially cause the coating on the mesh to strongly adhere to the inner handling sleeve. A reply for is requested to be completed to acknowledge receipt to the notification.Additional language to the instructions for use (IFU) to include:Prolonged exposure to high humidity may result in increased rate of adherence of the C-QUR mesh to its handling sleeve. Store in a Controlled Room Temperature (25 C / 77 F ) or less. Brief exposure to up to 40 C (104 F ) is acceptable.Questions please contact Atrium Medical Customer Service at 1- 800-528-7486 Monday through Friday 9:00 am to 5:00 pm EDT. |  | 95286 | 95,286 units | USA (nationwide) including Puerto Rico and the countries of Australia Austria Bahrain Brazil Canada Chile Colombia Dominican Republic Ecuador El Salvador France Germany Great Britain Greece Honduras Hong Kong India Ireland Israel Italy Japan Jordan Korea Malaysia Mexico Netherlands New Zealand Nicaragua Norway Panama Peru Portugal Romania Saudi Arabia Singapore South Africa Spain Sri Lanka Switzerland Taiwan Thailand Turkey and Venezuela. | Open |  |  | 145251 | Not_Computer
Z-1935-2013 | 65563 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | 5.0mm Unit steel rod, pedicle screw spinal system.Intended to provide immobilization and stabilization of spinal segments in skeletally mature patients. | USS Small Stature | 2 | August 09, 2013 | 2013 | Synthes USA HQ, Inc. | DePuy Synthes has issued a Medical Device Correction notice for the 5.0mm Stainless Steel rods. This product was produced using a finishing process not identified as part of the manufacturing specification. No adverse event reports associated with the use of this product were received, and an extensive engineering and medical safety review concluded that the product poses no incremental risk to t | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On June 5, 2013, De Puy Synthes notified by letter the direct acc acccounts which were identified to have received 5.0mm stainless steel rods and requested return of the product by calling 1-800-479-6329 to obtain a return authorization number. Completion of the verification section of the letter, including negative responses was also requested. | Contact the recalling firm for information | 218 | 218 | Nationwide Distribution. | Open |  |  | 218 | Not_Computer
Z-1903-2013 | 65617 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova Max Plus Monitor System (Kit w/10 Count Vials)Catalog Number: 43435Product usage: The Nova Max Glucose Test Strips are for use with the Nova Max Family of Glucose Monitors (Nova Max, Nova Max Plus, Nova Max Link, BD Logic and Paradigm Link). These monitors are intended to be used for the quantitative measurement of glucose in fresh capillary whole blood. They are intended for use by people with diabetes mellitus in the home and by healthcare professionals in clinical settings as an aid to monitor the effectiveness of diabetes control | Nova Max Glucose Test Strips | 1 | August 09, 2013 | 2013 | Nova Biomedical Corporation | Reports of false abnormally high glucose readings from test strips | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Nova Biomedical U.S.A issued an Urgent Medical Device Voluntary Recall Letter dated July 19, 2013 via registered/certified letters to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use of the affected product; contact Nova Diabetes Care Customer Service for replacement strips; do not stop testing blood glucose while waiting for replacement strips; if affected strips must be used follow precautions provided in notification letter. For questions call Nova Diabetes Care Customer Service at 1-800-681-7390.. | Contact the recalling firm for information | 1316420 | 1,316,420 kits | Worldwide Distribution - USA Nationwide and the countries of Chile, England , Peru, Nicaragua | Open |  |  | 58911847 | Not_Computer
Z-1941-2013 | 65828 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.2mm Drill Bit. This is a component within the PediLoc Locking Plate System utilized to pre-drill a hole in bone prior to placement of a bone screw. | OrthoPediatrics PediLoc Locking Plate System | 2 | August 09, 2013 | 2013 | OrthoPediatrics Corp | OrthoPediatrics has initiated a voluntary recall of Part # 01-1050-002, 2.5mm Drill Bit, Lot #91300335. Only drill bits that are in their original packaging are being recalled. The description on the package of this drill bit has a wrong size in the description (3.2mm). The recall affects only packaged devices still in original packaging due to misbranding of the product. The drill bit itself has | DESIGN: Process Design | On 7/24/13, OrthoPediatrics notified their consignees of this URGENT MEDICAL DEVICE RECALL via phone call. A follow-up email message was then sent with the recall notification, and then a hard copy of the notice was sent via certified letter. The notification addresses the affected product, recall issue, reply form instructions, and directs consignees to return affected product to OrthoPediatrics via a Fed-Ex return prepaid pouch. The firm directs consignees to forward the notification and reply form if the product was further distributed. | Contact the recalling firm for information | 68 | 68 | Nationwide distribution, including the states of MO, MI, WA, FL, NY, TN, CO, AZ, VA, TX, CA, MN, LA, NC, and UT. | Terminated | August 26, 2013 | 17 | 68 | Not_Computer
Z-1934-2013 | 65837 | tray, surgical, ent | MMO | Ear, nose, and throat manual surgical instrument. | Ear Nose & Throat | Ear Nose & Throat | 510(K) Exempt | Tray Surgical, ENT (various instrument trays)Used for stot age and transportation of reusable surgical instruments utilized in various ear, nose and throat (ENT) and head and neck surgical procedures. | Tra, Surgical, ENT | 2 | August 09, 2013 | 2013 | Medtronic Xomed, Inc. | The firm will recall and replace the identified instrument trays in the field, with trays that have improved labeling, including lids that contain the statement indicating 'PREVACUUM ONLY" |  | Medtronic sent a Customer Commincation Recall Letter to all affected customers with the following information: Subject: URGENT PRODUCT Recall  Immediate Attention RequiredMEDTRONIC XOMED INSTRUMENT TRAYS- model numbers provided on Attachment One- all lots distributed prior to July 2013Dear Purchasing Manager, Risk Manager and Central Supply Manager/Supervisor:This letter is to inform you that Medtronic Surgical Technologies (ENT Division) is initiating a voluntary recall to remove certain instrument trays from the field and replace them with trays that have modified labeling. Medtronic Xomed Instrument Trays are intended for the storage and transportation of reusable surgical instruments. These trays are not intended and not designed to maintain sterility. We are aware, however, that some users may also choose to use the trays for holding instruments during steam sterilization. If the user chooses to sterilize instruments within these instrument trays, they must validate the sterilization process. These users should be aware of the following issue:Issue Description:During the validation of a new instrument, testing revealed that use of a GRAVITY steam cycle with one of the listed instrument trays did not produce the Sterility Assurance Level (SAL) of 10-6 recommended by industry standard (ANSI/AAMI ST77:2006). A subsequent investigation of all Medtronic Xomed instrument trays identified additional trays as having certain locations within thetray that would not meet the recommended sterilization level of SAL10-6 during GRAVITY steam sterilization cycles. The trays identified through the testing are the subject trays involved in this field action. Instruments that are sterilized in the worst-case locations of the identified trays have a higher potential to be inadequately sterilized. An instrument that is inadequately sterilized can become the source of cross-contamination or cross-infection when used in multiple surgical procedures. At this time. | Contact the recalling firm for information | 3562 | 3,562 trays | Worldwide Distribution - USA (nationwide) Nationwide, Hawaii, Puerto Rico and Guam.Internationally to: Canada, Austria, Belgium, Cyprus, France, Germany, Greece, India, Iran, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Morocco, Netherlands, Nigeria, Portugal, Qatar, Reunion, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Australia, New Zealand, Singapore, Thailand, Vietnam, China, Hong Kong,Taiwan, Korea, Japan, Brazil, Chile, Colombia, Costa Rica, Ecuador, Guatemala, Mexico, Nicaragua, and Venezuela. | Open |  |  | 3562 | Not_Computer
Z-1951-2013 | 65655 | electrode, ion specific, potassium | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ISE Electrolyte Reference reagent (for Synchron LX Systems and UniCel DxC 600/800 Systems), Part No. A28937, A28945, 467935, 467915.Product Usage: ISE Electrolyte Buffer reagent and ISE Electrolyte Reference reagent, when used in conjunction with SYNCHRON LX System(s), UniCel DxC 600/800 System(s) and SYNCHRON Systems AQUA CAL 1, 2 and 3, are intended for the quantitative determination of potassium concentration in human serum, plasma or urine. | Synchron Systems LX20/ UniCel DxC ISE Reference Reagent | 2 | August 10, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the Synchron Systems LX20/UniCel DxC ISE Reference Reagent because it has a potassium (K) response near the instrument calibration flag limits and may exceed the limits on some instruments. | PRODUCTION CONTROLS: Process Control | Beckman Coulter Inc sent an Urgent Product Corrective Action (PCA) letter dated September 22, 2010 to all affected customers. Customers were notified of the the problems identified and the actions to be completed. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and in Canada. |  | 718 | 718 units total (132 units in US) | Worldwide Distribution - USA Nationwide and the countries of Belgium, France, Germany, Netherlands, Poland, Spain, Sweden, Switzerland, and United Kingdom | Terminated | August 10, 2013 | 0 | 718 | Not_Computer
Z-1954-2013 | 65502 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 00-4309-071-40 Trabecular Metal Reverse Shoulder System Instrumentation, Glenosphere Helmet, 40 mm Diameter Nonsterile, Qty -1Instrumentation used in shoulder replacement surgery. The glenosphere helmet is used to insert the glenosphere onto the base plate prior to impaction with the glenospere impactor. | Trabecular Metal | 2 | August 12, 2013 | 2013 | Zimmer, Inc. | There is potential for one or both tabsof theTrabecular Metal Reverse Glenosphere Helmet to fracture during use and become disassociated from the device. There have been 47 reported complaints of tab fracture. | DESIGN: Device Design | Zimmer Inc sent Urgent Medical Device Recall notifications on June 28, 2013, to all affected customers. The initial letter identified the issue, possible adverse consequences and informed customers that a replacement would be available in the coming months. In addition, the letter included references to both the TM Reverse Surgical Technique and to the video available on Zimmer TV and Zimmer.com. Both resources explain in detail the proper technique for utilization of the Glenosphere Helmet. Proper utilization of the Glenosphere Helmet dramatically reduces the likelihood of fracture. Distributors will receive a copy of the surgeon letter for their reference along with the list of affected surgeons. The affected surgeons will be identified using invoice data and in cases where the surgeon is unknown, the distributor will be instructed to assist with the surgeon list. In addition, information regarding the return process for the damaged devices will be included. Questions and concerns are directed to Zimmer's customer call center (1-877-946-2761).Phase 2: As replacement devices with the new design became available for distribution, a 2nd letter was sent to the distributors. Estimated replacement device availability will begin in July 2013. The distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices were included.For questions regarding this recall call 800-613-6131. |  | 3163 | 3,163 both sizes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Korea, Lebanon, Malta, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom, Utd.Arab.Emir. | Open |  |  | 3163 | Not_Computer
Z-1955-2013 | 65822 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Fisherbrand Sterile Swab, Calcium Alginate Fiber Tipped, Wood Applicator Swab 100 pouches/boxes, 10 boxes/case.Used in applying medications to or take specimens from a patient. | Fisherbrand Sterile Swabs | 2 | August 12, 2013 | 2013 | Fisher Scientific Co | The firm is recalling Fisherbrand Sterile Swabs due to 4 customer complaints reporting the presence of non-pathogenic organisms in two lots. The finding was confirmed in lot 61513 and lot 8710 is also being recalled as a precautionary measure. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Customers were notified by letter on June 18, 2013 and requested to discontinue use and return all affected product. The firm also asked that any additional affected customer base be contacted and notified of the recall if the lots had been further distributed. If you have any questions, contact Stericycle (reference "Event 4070") at 1-877-410-5558. | Contact the recalling firm for information | 1230 | 1230 | Nationwide and the country of Canada. | Open |  |  | 1230 | Not_Computer
Z-1952-2013 | 65560 | reamer | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Medullary Reamer.Intended to be used to facilitate the preparation of the intramedullary cavity. | Synthes Flexible Medullary Reamer | 2 | August 12, 2013 | 2013 | Synthes USA HQ, Inc. | Due to the coiled design for this product, the product is difficult to clean and the potential for corrosion on the device exists. The potential for harm exists if there is an introduction of foreign material at the operative site originating from the presence of corrosion and/or the retained undefined material that could not be adequately cleaned. | DESIGN: Device Design | Sythes notified direct accounts by letter on 6/11/13 and requested them to check their inventory and remove affected lots from stock. Synthes requested that affected lots be returned and to call Synthes at 1-800-479-6329 for a Return Authorization Number. Completion and return of the Verification section at the end of the letter was requested, including negative responses. | Contact the recalling firm for information | 13050 | 13,050 | Nationwide Distribution. | Open |  |  | 13050 | Not_Computer
Z-1957-2013 | 64994 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Solid Driver, Tipped, AO. Part Number: 11-70083.Allows the user to retain the screw when changing axis of insertion during fixation of spinal intervertebral. | Tipped Tapered Solid Driver (Screwdriver) | 2 | August 13, 2013 | 2013 | SpineFrontier, Inc. | The nipple of the tipped driver broke off when surgeons were attempting to insert the screw. | DESIGN: Device Design | On May 9, 2012 an email was sent to consignees requesting that the product be returned for servicing and that a replacement part would be sent. |  | 9 | 9 | Distributed in DC and the states of MA, and OK. | Terminated | January 10, 2014 | 150 | 9 | Not_Computer
Z-1959-2013 | 65857 | thermometer, electronic, clinical | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Disposable Temperature Probes, (sterile), Sensor diameter: 3mm/0.12 in; 25 units per box. Vital Signs Inc. a GE Healthcare Company, 20 Campus Road, Totowa, NJ 07512 USA, www.gehealthcare.com/vitalsigns.Continuous temperature monitoring. | General Purpose Temperature Probe | 2 | August 13, 2013 | 2013 | Vital Signs Devices, a GE Healthcare Company | GE Healthcare has become aware of a potential safety issue associated with the use of GE General Purpose 9 French Temperature Probe, Disposable, 400 Series [M1024229]. GEs Disposable General Purpose 9FR Temperature Probe M1024229 is intended for oro-esophageal and rectal use. As it pertains to esophageal placement, it has been brought to our attention that some users are inserting the probe |  | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 23, 2013 via Fed Ex beginning on July 29, 2013 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE will update the Instructions for Use (IFU) and pouch label with new content related to insertion locations. GE has will also place the probe (M1024229) on quality hold until the new labeling changes are implemented.If you have any questions or concerns regarding this notification, please contact GEMSIT Customer Service at 1-800-588-7044 (Domestic) or Vital Signs Customer Service at 1-800-932-0760 (international). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 12239 | 12,239 units (5,002 units US, 7,237 units OUS). | Worldwide distribution: US (nationwide) and countries of: Australia, Belgium, Canada, Czech Republic, Finland, France, Germany, Italy, Lebanon, Netherlands, Norway, Poland, Saudi Arabi, Singapore, South Africa, Spain, Sweden and United Kingdom. | Terminated | January 02, 2014 | 142 | 12239 | Not_Computer
Z-1962-2013 | 64711 | antigens, cf (including cf control), influenza virus a, b, c | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Remel X/pect Flu A&B; Control Swabs, For In Vitro Diagnostic Use, Ref R246003, each package contains 20 Flu A Control Swabs and 20 Flu B Control Swabs, each swab is individually packaged in a foil pouch. The firm name on the label is Remel Europe, Ltd., Dartford, Kent, United Kingdom.Product is a test kit containing characterized positive and negative controls intended for use in qualitative procedures as external quality control material for the Remel Xpect Flu A&B; test kit. | Remel Xpect Flu A&B; Control Swabs | 3 | August 13, 2013 | 2013 | Remel Inc | Flu A+ Control Swabs incorrectly give Flu A- results when used for quality control testing with the Xpect Flu A&B; test kit. | TRAINING: Employee Error | The firm, ThermoFisher Scientific, sent a "MEDICAL DEVICE RECALL" letter dated February 20, 2013 to its customers via regular mail. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory for affected product lots; discard the affected product; notify any personnel who needs to be notified of the potential for a false negative indication when using the swab, and complete and return the attached Medical Device Recall Return Response Acknowledgment & Receipt Form via Fax to: Attn: Technical Service & Regulatory Affairs at 1-877-428-1924.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 96 | 96 kits of 20 swab sets | Nationwide distribution: AL, AR, CA, CO, CT, FL, GA, IA, IL, KS, MA, MD, MI, MO, NM, NV, NY, OH, PA, TX, WA, and WI. | Terminated | August 14, 2013 | 1 | 96 | Not_Computer
Z-1960-2013 | 65005 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3i Dental Implant, Parallel Walled 6mm (D) x 10mm (L) | T3 Dental Implant | 2 | August 13, 2013 | 2013 | Biomet 3i, LLC | A small number of outer implant boxes may be mislabeled. The inner tray, containing the implant, is label correctly. This condition, if present and not recognized could potentially result in the dental implant not being able to be placed into the osteotomy or seated to the desired depth if placement is attempted. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm issued "Urgent Medical Device Recall" notification letters to their customers dated May 10, 2013. The notification described the issue and provided recommendations actions regarding affected product. Customers with questions may contact 1 800-342-5454. |  | 51 | 51 | Distribution to the state of Massachusetts, and the foreign countries of Australia, Belgium, Spain, Denmark, France, Luxembourg, Switzerland, Germany, Bonaire, Sweden, and Italy. | Open |  |  | 84 | Not_Computer
Z-1947-2013 | 64971 | ventilator, emergency, manual (resuscitator) | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | 1) Adult Manual Resuscitator, 2K8004, Oxygen Reservoir Bag, Adult Mask; 2) Adult Manual Resuscitator, 2K8005, 40" (1.0 m) Oxygen Reservoir Tubing, Adult Mask; 3) Adult Manual Resuscitator, 2K8017, Variable Volume Oxygen Reservoir Tubing, Adult Mask; 4) Adult Manual Resuscitator, 2K8034, 40" (1.0 m) Oxygen Reservoir Tubing, PEEP Valve, without Mask; 5) Adult Manual Resuscitator, 2K8035, Oxygen Reservoir Bag, PEEP Valve, Adult Mask; 6) Pediatric Manual Resuscitator, 2K8037, 40" (1.0 m) Oxygen Reservoir Tubing, PEEP Valve, Pediatric Mask, Pressure Limiting Valve with Lock; 7) Infant Manual Resuscitator, 2K8040, Oxygen Reservoir Bag, PEEP Valve, Infant Mask, Pressure Limiting Valve with LockProduct Usage: Pulmonary resuscitation | AirLife | 2 | August 13, 2013 | 2013 | Carefusion 2200 Inc | CareFusion has received four (4) customer reports of a component (oxygen connector) within the resuscitation bag being occluded which disables the flow of supplemental oxygen to the resuscitation bag and thus may pose a serious patient safety risk. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | CareFusion sent an URGENT RECALL NOTICE letter dated April 22, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to complete and return an attached Distributor Acknowledgement Form within fifteen (15) days of receipt indicating their acknowledgement of the URGENT RECALL NOTICE and indicate if they have any affected product on hand. They are further requested to perform a 100% physical inventory to verify if any of the affected product codes/lot numbers are in stock within their distribution network/facility. In addition, customers are requested to forward the Customer Acknowledgement Form, the Customer Notification Letter, the Certificate of Destruction Form, the Pictures of Location of Lot Numbers on Case/Product, and the Frequently Asked Questions (FAQs) to all end users which they have distributed the affected product. For questions call 847-473-7097. |  | 14112 | 14,112 units | US Nationwide in the states of AZ, CA, CO, DE, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NM, NY, OH, PA, TX, UT, WA | Open |  |  | 14112 | Not_Computer
Z-1970-2013 | 65626 | calibrator, primary | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Calibrator A. Cat. #04800735, SMN 10285904 (6 pack); Cat. # 04800840, SMN 10285905 (6 pack - REF).For calibrating the following assays using the ADVIA Centaur or ACS:180 systems: FT3, T3, T4, Tup, FT4, THEO2. | ADVIA Centaur Calibrator A | 3 | August 14, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibrator A lot CA77 is under-recovering to the assigned target for the ADVIA Centaur Total T4 Assay. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An Urgent Medical Device Correction (UDMC) was sent to all affected Siemens Healthcare Diagnostic customers in the United States on June 28, 2013 via FedEx. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 27, 2013. FedEx Return Receipts will be used to confirm that the customers received the UMDC. If confirmation of receipt by any customer cannot be verified through a return receipt, a Siemens Technical Solutions Center Representative will contact those customers by telephone to ensure that they received the Urgent Medical Device Correction and completed the required actions. These notices inform the customer that they should discontinue the use of Calibrator A, lots ending in 77, for use with the ADVIA Centaur Systems T4 assay. Customers are informed that they may continue to use Calibrator A lot CA77 with the other assays FT3, FT4, Total T3, T-Uptake. For questions or additional information contact your local Technical Support Provider or Distributor. |  | 265 | 265 Cal A 6 packs (US and OUS) | Worldwide Distribution-Nationwide including PR and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NV, NY, OH, PA, TN, TX, VA, VT, and WV and the countries of Slovakia, Germany, Turkey, TW, HK, CH, TH, BR, CN, SG, and ZA. | Open |  |  | 265 | Not_Computer
Z-1966-2013 | 65732 | filter, blood, cardiopulmonary bypass, arterial line | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Affinity Pediatric Arterial Filter, model number S4014, sterilized using ethylene oxide. Recall Update 9/6/13: Medtronic, Red cross Infant Pack 1/B, South-Africa, REF M272403C, Sterile EO.The Affinity Pediatric Arterial Filters are single-use, sterile, nonpyrogenic devices designed to filter microemboli greater than the specified micron size from the circuit for periods up to 6 hours during cardiopulmonary bypass surgery. These devices are available both in an uncoated and a Carmeda coated option. The Carmeda coating is a BioActive surface that is non-leaching and provides a thromboresistant blood contact surface. | Affinity Pediatric Arterial Filter | 2 | August 14, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is recalling 148 Affinity Pediatric Arterial Filters from 2 manufacturing lots because a limited quantity of filters may have a small breach in the filter media that may allow unwanted particulate or gaseous emboli to pass through the filter and the outlet of the device, which could result in serious injury to a pediatric patient including neurological damage, or potential patient death. | PRODUCTION CONTROLS: Process Control | Customer communication was initiated verbally on June 14, 2013, to the 4 affected customers to quarantine any un-used units. An Urgent Medical Device Recall letter, dated June 26, 2013. was mailed to customers on June 28, 2013. The Urgent Medical Device letter referenced that the letter is in follow-up to the previous phone call. It identified the affected product, stated the issue with the 2 affected lot numbers and health risks, and asked that the affected devices be quarantined for return to Medtronic.If you have any questions please contact your Medtronic Sales representative or Lifeline Technical Services at 1-877-526-7890. 1-763-526-7890.Update: Customer communication to the newly identified customer was initiated verbally on August 20,2013, with the customer letter delivery beginning AUG 27, 2013. A Medtronic representative is available to facilitate the return of affected product. |  | 153 | 153 devices. | Worldwide Distribution - US Distribution only to MS and IL, and the countries of Argentina, South Africa and Singapore. | Open |  |  | 153 | Not_Computer
Z-1972-2013 | 65744 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards EndoReturn Arterial Cannula, model ER21B and ER23BProduct Usage:The ThruPort EndoReturn arterial cannula kit (with or without hemostasis valve), an introducer, a guidewire,and a connector hub. The ThruPort catheter introducer sheath kit includes a wire reinforced cannula with hemostasis valve, an introducer, and a guidewire. The cannulae have a wire reinforced section to provide kink resistance and flexibility, tapered tips to aid in insertion and advancement into the femoral artery, a non-reinforced clamp site, a 3/8 in. (9.5 mm) barbed connector, and a hemostasis valve (available on some versions) that allows passage of catheters, such as the EndoClamp aortic catheter. The introducers accept a .038 in. (0.97 mm) guidewire and are marked to simplify assembly and indicate alignment. On some versions, a connector hub secures and immobilizes the introducer within the cannula for easier, one person insertion of the cannula/introducer assembly. | Edwards EndoReturn Arterial Cannula, model ER21B and ER23B | 2 | August 14, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of theEndoReturn Arterial Cannula that were built with an incorrect component. This component can create aperformance problem with the IntraClude IntraAortic Occlusion device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Edwards sent an Urgent Field Safety Notice letter dated July 16, 2013 via Fed-Ex. The letter identified the affected product, problem and action to be taken. The letter included a Customer Acknowledgement Form to be completed and faxed to 800-422-9329 Attn: Recall Coordinator. If the customer does not send back the acknowledgement form, two additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. Customer were requested to assess their inventory to assure that all affected product is identified and returned. Returned products will be destroyed. For questions call Edwards Customer Service at 800-424-3278. |  | 855 | 855 | US Nationwide including Canada and VA sites. | Open |  |  | 855 | Not_Computer
Z-1971-2013 | 65126 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc BGEM Test Card (Catalogue number CT-1004-00-00).quantitative testing of samples for pH, pCO2, pO2, Sodium Potassium , Ionized Calcium, Glucose, Lactate and Hematocrit. | epoc BGEM Test Card | 2 | August 14, 2013 | 2013 | Epocal | This recall is being initiated because the affected lots of this product may yield results that are not consistent with stated precision performance criteria at concentration values at the high-end of the Glucose test sensor ( | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | On March 13th 2013, Epocal placed all CT-1004-00-00 test card on shipment hold. Affected product in Epocal control has been segregated and quarantined. Beginning on March 26 2013, Epocal Inc. plans to begin distribution of the Recall Notification Letter to the 4 direct consignees (customers/distributors) via certified mail. Included with the Recall Notification Letter is a faxback verification form. Notification will include all distribution channels to the end-user level. Direct consignees are instructed to discontinue use and discard all stock on hand, complete the provided faxback verification form to document disposal and contact their consignees to advise them of the notification and disposal of affected product. |  | 828750 | 828,750 Test Card | Worldwide Distribution-USA (nationwide) including the District of Columbia, US Virgin Islands and Puerto Rico, and the states of CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, MI, MN, MS, NC, ND, NE, NH, NJ, NM, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, VE, WA, WI, and WV and the coutries of India, Argentina, Australia, Azerbaijan, Belarus, Canada, China, Colombia, Egypt, Germany, Guatemala, Israel, Japan, Malaysia, Mexico, Nepal, Netherlands, Peru, Philippines, Russia Fed., Singapore, South Africa, South korea, Thailand, and Vietnam. | Open |  |  | 828750 | Not_Computer
Z-1965-2013 | 65829 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. Two-Lumen CVC Kit with Blue Flex Tip Catheter. The multiple lumen catheter permits venous access to the central circulation of patients. | Arrow TwoLumen CVC Kit with Blue Flex Tip Catheter | 2 | August 14, 2013 | 2013 | Arrow International Inc | Arrow is recalling the product due to a labeling inconsistency. One portion of the lidstock states the kit contains a 3-Lumen catheter while the rest of the lidstock correctly identifies the catheter as 2-Lumen. | MISBRANDING: Labeling False and Misleading | An "Urgent Medical Device Recall Notification" letter dated July 31, 2013 was sent to direct accounts requesting immediate discontinuation of use and quarantine of all affected product on hand. Negative responses were also requested and the immediate return of product to Arrow was authorized. Questions were directed to Customer Service at 800 343-2935. |  | 156 | 156 | AL, PA, CT, NC, VA, KS, RI, AZ, NJ, and IL. | Open |  |  | 156 | Not_Computer
Z-1963-2013 | 65838 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Anastaflo Intravascular Shunt, Model code IVS12512, IVS1512, IVS17512, IVS2012, IVS2512, IVS3012, IVS1530, IVS2030 and IVS25Anastaflo shunts are intended for use in preventing ischemia by shunting and/ or perfusing blood or cardioplegic solution distal to the anastomosis site during the construction of coronary artery bypass grafts30. | Anastaflo Intravascular Shunt | 2 | August 14, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of the Anastaflo intravascular Shunt due to concerns of excessive adhesive on the shunt body that may interfere with suturing of by-pass grafts. | PRODUCTION CONTROLS: Process Control | Edwards Lifesciences sent a Urgent Field Safety Notice letter dated July 30, 2013, via Fed-Ex to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters are tracked through customer service and as part of the customer acknowledgement letter. Customers are asked to return products.If the customer does not send back the acknowledgement letter, three additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. We reconcile the shipments made to customers against the returned product to determine impact to distribution. Once you have verified your inventory, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329 within three days of receipt of this Field Safety Notice. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 to obtain a Returned Goods Authorization number and replacement product.We sincerely regret the inconvenience caused by this action and appreciate your immediate attention to this matter. The Customer Service organization can answer questions about when replacement IVS Shunts will be available. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST or contact your Edwards sales representative concerning the recall. |  | 29618 | 29,618 units | Nationwide and the country of Canada | Open |  |  | 29618 | Not_Computer
Z-1979-2013 | 65798 | implantable cardioverter defibrillator (non-crt) | LWS |  |  | Cardiovascular |  | Boston Scientific, ENERGEN DR ICD, Model E143. The device is an Implantable Cardioverter Defibrillator. | ENERGEN DR ICD | 2 | August 16, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM manufacturing quality system recently discovered test artifacts (related to shock charge time) in a small number of implantable defibrillators (ENERGEN DR ICD model E143) which require further investigation. While these devices successfully passed all manufacturing tests and met all requirements, they showed a slightly longer charge time than other devices during a manufactu |  | Consignee was sent on 7/15/13 a Boston Scientific "Medical Device Retrieval" letter dated July 12. 2013. The letter described the product problem and informed the consignee that they will replace the product removed from their inventory. For further questions, contact Boston Scientific Technical Services at 1 800 227-3422. |  | 1 | 1 (4 devices were implanted and are not part of this action) | US distribution in the state of Tennessee. | Terminated | November 12, 2013 | 88 | 1 | Not_Computer
Z-1981-2013 | 65833 | prosthesis, hip, femoral component, cemented, metal | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem, several head Diameter, 203 mm Stem length.Used to seat the prosthesis with or without bone cement. | Moore Hip Prosthesis Long Fenestrated Stem | 2 | August 17, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a single packaging tray configuration, which | PRODUCTION CONTROLS: Packaging Process Control | Zimmer is removing the devices from the field including distributor and hospital inventories. A notification was issued to known users of the product on 07/16/2013. |  | 1710 | 1,710 units are in distribution | Worldwide Distribution-USA (nationwide) including the states of CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VI, and WI, and the countries of Singapore, Canada, Mexico, Arab Emirates, Saudi Arabia, and Hong Kong. | Open |  |  | 1710 | Not_Computer
Z-1982-2013 | 65855 | endoscopic suture/plication system, gastroesophageal reflux disease (gerd) | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Device with Serosa Fastener and Accessories. Models R2001 and R2002.Indicated for the treatment of symptomatic chronic gastroesophageal reflux disease. | EndoGastric Solutions EsophyX2 Device | 2 | August 17, 2013 | 2013 | Endogastric Solutions Inc | Endogastric Solutions, Inc. has received a limited number of reports relating to the loss of tissue mold control when operating the R2001 or R2002 EsophyX2 device. In one case, surgical intervention was necessary to remove the device. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Endogastric sent the Safety Alert: EndoGastric Solutions (EGS), EsophyX2 Device letter, dated June 5, 2013, to their consignees.Endogastric sent the second letter URGENT: MEDICAL DEVICE RECALL EsophyX2, dated July 26, 2013. This letter advised customers that the firm is voluntarily recalling EsophyX2 Devices with SerosaFuse Implantable Fastener and Accessories (Models R2001 and R2002) manufactured before February 2012.Customers who have UN-EXPIRED devices in their inventory are advised to discontinue use and fill out the Medical Device Recall Return Response form and return it to Endogastric. The Customer Service will contactcustomers with instructions on how to return the product to the company.Customers who may have EXPIRED devices are advised to with their central supply departments to ensure all identified products have been removed from inventory and destroyed. They should fill out the Medical Device Recall Return Response form with the lot numbers and quantity destroyed and return thecompleted form to the firm.Customers can call the Customer Service and Support at 425-307-9269, Monday through Friday, 8:00AM to 5:00 PM, Pacific Time for questions. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. |  | 5192 | 5192 units in the US and 41 units outside the US | Distributed nationwide and Italy. | Open |  |  | 5192 | Not_Computer
Z-1980-2013 | 65552 | chair, positioning, electric | INO | Electric positioning chair. | Physical Medicine | Physical Medicine | 510(k) | CHAMPION Medical Recliner, AC Powered with heat and massage, one chair per carton, or one chair wrapped in a blanket if delivered in person. Recliner provides heat and massage after surgery. | CHAMPION Medical Recliner, AC Powered with heat and massage | 2 | August 17, 2013 | 2013 | Invacare Corporation | There is a possibility of elevated temperatures in the seat pad. In addition, the hospital cord set does not agree with the labeling (information in manual). | PRODUCTION CONTROLS: Process Control | Champion Manufacturing plans to initiate this recall within approximately 10 business days after the submission of this report through a phone campaign and notification letters. |  | 285 | 285 units | Nationwide distribution: USA including states of: CA, FL, IN, MN, NJ, NY, OH, RI, SC, SD, TX, and VA. | Terminated | January 09, 2014 | 145 | 285 | Not_Computer
Z-1983-2013 | 65785 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | GYRUS ACMI, REF 3740, Everest BiCOAG Bipolar Dissecting Forceps 5mm/45 cm, Sterile EO.GYRUS ACMI, REF 3740PK, or REF 94201PK, PKS LYONS, Dissecting Forceps 5mm/45cm, Sterile EO.The Dissectors are sterile, single use electrosurgical devices that are indicated for the electrosurgical coagulation, mechanical grasping, and dissection of tissue during the performance of laparoscopic and general surgical procedures. | GYRUS ACMI, REF 3740, Everest BiCOAG Bipolar Dissecting Forceps 5mm/45 cm, Sterile EO | 2 | August 19, 2013 | 2013 | Gyrus Medical, Inc | On April 29, 2013 Olympus Surgical Technologies America R&D; received results from an accelerated aging study conducted as part of a new market submission. The study found that one of 60 Dissector units that underwent 3 year accelerated aging failed the dye penetration test indicating that the seal on sterile packaging was compromised. The dissector products are labeled with a 3 year shelf life. | PRODUCTION CONTROLS: Packaging Process Control | Olympus sent a "Medical Device Recall" letter dated July 24, 2013, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.. Customers were advised to immediately cease any further use of the product and remove it from inventory and quarantine it until is shipped back to the company. To return the product, call 1-888-524-7266 Olympus customer service representative. Customers were requested to complete and return by fax the enclosed questionnaire to 484-896-7128. If the product was further distributed then they need to identify their customers and notify them at once of the product recall. |  | 1140 | 1140 | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, IN, KS,KY, MD, MA, MI, MN, MS, MO, NY, NC, OH, PA, SD, TN, TX, VA and WI., and the countries of BRAZIL, CHINA, AUSTRALIA, CHINA, GERMANY and INDIA. | Open |  |  | 1140 | Not_Computer
Z-1984-2013 | 65908 | intraocular lens | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Product Name: nanoPOINT Injector System, item number 37-0000-08.Intended to fold and insert a STAAR Surgical CollamerÂŽ nanoFLEX or Afinity IOLs, through surgical procedure in a human eye. | nanoPOINT Injector System | 2 | August 20, 2013 | 2013 | Medicel Ag | The firm initiated this recall of one lot (121402) of nanoPOINT Injectors after they noticed that Tyvek lids might have been damaged and, therefore, the stability of the sterile barrier might be negatively affected. |  | Medicel AG sent recall notification to its one direct consignee via email on 08/05/2013. The recall notification letter was titled "URGENT-Medical Device Correction Notice for nanoPOINT Injectors LOT 121402". The recall notification letter informed the customer of the product device description and lot number recalled, reason for recall, instructions to "not use this affected lot...and return them", contact information and an area on the bottom of the letter to acknowledge receipt of the correction notice.The firm's direct consignee who further distributed this product to the end users will send out notifications. |  | 748 | 68 boxes or 680 units | Nationwide Distribution. | Open |  |  | 748 | Not_Computer
Z-1991-2013 | 65879 | tubes, gastrointestinal (and accessories) | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ConvaTec Flexi-Seal SIGNAL +, Rx OnlyProduct Usage:For use to manage fecal incontinence through the collection of liquid to semi-liquid stool and to provide access to administer medication. | ConvaTec | 3 | August 20, 2013 | 2013 | Convatec Inc. | ConvaTec notified their customers on 6/25/2013 regarding the manufacturing irregularity that occurred related to the packaging of certain lots of Flexi-Seal Signal FMS. In certain lots, the lot number convention was duplicated on the packaging. | DESIGN: Process Design | ConvaTec issued a notification letters to their customers dated June 25, 2013 via UPS. The letter informed the customer of the issue with the related product and the action to be taken. If the customer has any question regarding the notification they can contact customer service at 1-800-422-8811. |  | 22896 | 22,896 units | US Nationwide Distribution | Terminated | January 08, 2014 | 141 | 22896 | Not_Computer
Z-1985-2013 | 65869 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand Model S099B, packaged in Sorin Group Heart/Lung Perfusion Packs. | Edwards Lifesciences Suction Wand Model S099B | 2 | August 20, 2013 | 2013 | Sorin Group USA, Inc. | Sorin Group is recalling certain lots of Edwards Lifesciences Suction Wand Model S099B due to presence of plastic particles inside the pouch. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | On 7/8/13, customers and distributors were sent a "Field Safety Notice Product Recall" letter and requested to return unused Edwards Suction Wands found in the Perfusion Packs to the Sorin Group. Customers are to contact Sorin Group Customer Service at 1-800-650-6502 to receive instructions for returning the affected Edwards Suction Wand to Sorin. Customers were also advised to not discard affected product. |  | 18927 | 18,927 packs | Nationwide Distribution | Terminated | January 28, 2014 | 161 | 18927 | Not_Computer
Z-1986-2013 | 65644 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Reservoirs, Model No: MMT-326A (1.8mL) and MMT-332A (3.0mL)The model MMT-326A and MMT-332A are syringe type insulin reservoirs intended for use with Medtronic Paradigm series insulin infusion pumps. The two reservoirs are identical except for barrel length and as a result, maximum insulin capacity. The MMT-326A holds a maximum of 1.8ml of insulin whereas the longer MMT-332A reservoir can hold up to 3.0ml of insulin. Both reservoirs are supplied sterile and are labeled for single use. Typically, the reservoir is replaced every tow or three days at the time of infusion set replacement. The Paradigm reservoirs have a shelf life of three years post sterilization. Product Quantity Distributed (Int) 669,672 units | Medtronic MiniMed Paradigm Reservoirs | 2 | August 20, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling certain lots of Medtronic MiniMed Paradigm Reservoirs MMT-326A (1.8mL) and MMT-332A (3.0mL) used with Medtronic Paradigm insulin pumps because they may have increased risk for leaking. A leak in the reservoir may result in delivery of less insulin than intended. In addition, if there is a leaky reservoir and an insulin blockage occurs in the infusion set, the pump may not |  | An Urgent Medical Device Recall notice was sent to customers on July 3, 2013, who purchased the Medtronic Paradigm Reservoirs MMT-326A and MMT-332A. The recall letter informed the customers of the problems and the actions to be taken. Medtronic instructed customers that they will replace all reservoirs from the affected lots at no cost. Customers are instructed to refer to www.medtronicdiabetes.com/support/product-updates for a list of frequently asked questions. Customers are instructed to report adverse reactions or quality problems experienced with the use of Medtronic Diabetes insulin pumps, infusion sets or reservoirs to FDA's MedWatch Adverse Event Reporting Program. Customers are instructed to call for replacements at (866) 222-2584 Monday through Friday (8am-6pm) central time. Customers with an urgent issue may use the 24 hour helpline. |  | 669672 | 669,672 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Bermuda, Brazil, Canada, Chile, Columbia, Ecuador, El Salvador, Germany, Japan, Mexico, Netherland, Paraguay, Puerto Rico, Singapore, Uruguay and Venezuela. | Terminated | January 16, 2014 | 149 | 669672 | Not_Computer
Z-1987-2013 | 65678 | bottle, collection, vacuum | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | PleuraGuide Disposable Chest Tube Kit; Product codes: 17100, 17150, and 17200.Indicated to assist with the insertion of thoracic catheters. | PleuraGuide Kit | 2 | August 20, 2013 | 2013 | Atrium Medical Corporation | Individual component's IFUs are not included with the kit. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Notice of correction was sent on July 18, 2013 via UPS and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. It provides instructions how to handle the product that may be affected. Additional labeling will be added to the kit to include instructions for use for the additional components. Atrium Medical will confirm through tracking of the UPS letters that each consignee has received the letter of correction. This shall be 100%. Where a UPS delivery cannot be confirmed, Atrium Medical will make good faith effort to contact the consignee to ensure that the letter was received. |  | 6516 | 6,516 units | Nationwide Distribution. | Open |  |  | 6516 | Not_Computer
Z-1989-2013 | 65862 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500DProduct Usage:The Precision 500D is designed to perform radiographic and fluoroscopic x-ray examinations. | Precision 500D | 2 | August 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the overhead Video Monitor Suspension involving Precision 500D Systems. There is a potential for the locking nut to loosen out of the Video Monitor Suspension yoke which could result in impact. This will not lead to an immediate fall of the suspended monitor since it will be held in place by the remaining bushing and | DESIGN: Device Design | GE Healthcare "Urgent Medical Device Correction" letter dated May 1, 2013 to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter informed customers that GE Healthcare will correct all affected systems at no cost. A GE Healthcare service representative will contact you to arrange for this correction. For questions or concerns regarding this notification call: US 800-437-1171, Japan: 0120-055-919. |  | 445 | 445 (421 US, 24 OUS) | Worldwide Distribution - USA Nationwide in the states of AL, AK, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, NMI, MN, MS, MO, NE, NV,NH,NJ, NM, NY, NC,ND, OH, OK OR, PA,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY and the countries of CANADA, FRANCE GERMANY, GREAT BRITAIN ITALY, PANAMA, SAUDI ARABIA. | Terminated | November 26, 2013 | 98 | 445 | Not_Computer
Z-1988-2013 | 65745 | strip, adhesive, closure, skin | FPX |  |  | General & Plastic Surgery | 510(k) | Distributed in USA: Product Number: 40-2894: Reinforced Wound Closure, Latex Free Sterile EO Rx, 1/4" X 4 ", 6 mm X100 mm, (2 Cards/pouch, 1,000 Pouches/box); Distributed in Europe: "Product Number: 39-0058: Medicare Reinforced Wound Closure Strips 1/4" X 4 ", Sterile, Box 1000 pouches 5 strips /pouch, Pouch , 5 strips 6 mm X100 mm Sterile EOProduct Number: 40-2893: Wound Closure Strips , Sterile (1,000pouches/box)Rx; Pouch Label: Wound Closure Strips 6 mm X100 mm, 1/4" X 4 ", RxProduct Usage: intended for skin closure and closure support for sutures and staples | Wound Closure Strips | 2 | August 20, 2013 | 2013 | Aspen Surgical Products, Inc. | Aspen Surgical is recalling Wound Closures because firm could not confirm sterility, after FDA sampling indicated the product may not be sterile. Using the recalled product could lead to infection which may require further treatment including the use of antibiotics. There have been no reports of injuries or illnesses as of 7/18/2013. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Aspen sent an Urgent Medical Device Recall letter dated July 18, 2013 via certified mail to all customers with directions to forward the notification and provide Aspen with sub-account information. Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. Questions or concens should be addresses to Kelli Jona Directo of QA / RA at 888-364-7004 ext 177 jonask@aspensurgical.com or Michelle Pulaski, Quality System and Compliance Manager 888-364-7004 ext 146 pulaskim@aspensurgical.com 8-5 EST M-F. |  | 3286 | 3286 boxes ( 6,353,000) | Worldwide Distribution - US Nationwide in the state of Florida and in the countries of , Netherlands, and Reditch Worcestershire England | Open |  |  | 3286 | Not_Computer
Z-1994-2013 | 65866 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Getinge Assure Accufast Biological Indicators, Catalog Number 61301606636 (50 Pack) and 61301606637 (100 Pack), STEAM --- Manufactured for: Getinge USA, Inc. 1777 E. Henrietta Rd., Rochester, NY 14623. Getinge Assure AccuFast Biological Indicators (BIs) are designed for biological testing of 121ÂŽC, 132ÂŽC, 134ÂŽC and 135ÂŽC steam sterilization cycles. The AccuFast BI is a self-contained biological indicator containing Geobacillus stearothermophilus inoculated on a paper carrier and placed in a plastic vial with a small glass ampule containing sterile culture medium. After activation, the plastic vial serves as a culture tube. The chemical indicator on the label changes from blue to black when exposed to steam, distinguishing exposed from unexposed units. This indicator does not indicate adequate sterilization. | Getinge Assure | 2 | August 21, 2013 | 2013 | Getinge USA Inc | AccuFast Biological Indicator Lot SR-430 may have been exposed to conditions affecting its viability. This could result in a false negative (or pass) when exposed to sub-lethal sterilization parameters. |  | GETINGE sent an DEVICE CORRECTION NOTICE letter dated July 31, 2013, and response forms were sent to customers on August 2, 2013, via Certified Mail Return Receipt Requested. The letters instruct customers to immediately discontinue use, quarantine, and destroy the affected lot. In addition, they instruct customers to: Check the appropriate box and record the quantity of affected unused product on the enclosed Correction Response Form OR If if they do not have any lots of affected unused inventory check the box indicating such. Destroy the unused product in accordance with your institution's policy and instructions for use. Record the method of destruction on the Correction Response Form. Be sure to include the Facility Name, Address, Contact Name and Phone Number. Sign, date and e-mail the form back to SR430responseform@getinge.com within 10 days or sooner of receiving this notice. Credit will be issued upon receipt of the Correction Response Form. Replacement product of unaffected lots is currently available. The warehouse shipping the affected lot is no longer being used (due to business decision, not as result of this recall). The current warehouse is maintaining the product in a controlled temperature segregated room away from any potential sterilants.If you have any questions regarding this notification please call 1-800-475-1400 ext. 5026. |  | 78 | Catalog # 61301606636: 78 Packs (3900 BIs); Catalog # 61301606637: 43 Packs (4300 BIs) | Worldwide Distribution - USA and the country of Canada. | Open |  |  | 121 | Not_Computer
Z-1992-2013 | 65888 | device, medical examination, ac powered | KZF | AC-powered medical examination light. | General Hospital | General Hospital | 510(K) Exempt | Stryker Visum 300 Ceiling-mounted Exam Light. | Stryker Visum 300 | 2 | August 21, 2013 | 2013 | Stryker Communications | Potential incomplete final weld on the drop tube flange. | PRODUCTION CONTROLS: Process Control | The firm decided to recall and notified consignees with an "Urgent Medical Device Correction/Removal" notification letter dated 07/29/13. The product issue was described and recommended actions were provided. Customers may contact Stryker Technical Support at 866-841-5663. |  | 1375 | 1,375 units | Worldwide Distribution: US natiowide and to the country of Canada. | Open |  |  | 1375 | Not_Computer
Z-1995-2013 | 65834 | drape, surgical | KKX | Surgical drape and drape accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Blue Belt Technologies, Inc. Monitor Drape REF: PFSD-V0001, Rx only, Sterile EO. Product Usage: The intended use of this device is t protect equipment from contamination during a variety of procedures throughout the clinical setting. | Blue Belt Technologies | 2 | August 21, 2013 | 2013 | Blue Belt Technologies MN | Blue Belt Technologies is recalling sterile Monitor Drapes used with their Navio PFS system during orthopedic knee surgery. The drapes are used to cover the Navio PFS system computer monitor and arm assembly and provide a sterile barrier in the surgical suite. The defect is evident when the computer monitor is being draped. The front section of the drape separates from the rest of the drape at | DESIGN: Packaging Design/Selection | Blue Belt Technologies contacted consignees by telephone using a script on July 19, 2013. The consignees were informed of the problem and the product involved in the recall. Consignees were advised to immediate stop using the product and to locate and quarantine all unused product. Additionally, an "Urgent Medical Device Recall" letter dated July 24, 2013 was issued on July 24, 2013 and August 12, 2013. The letter includes the description of the Issue, description of safety concerns, affected product, and action to be taken by user. For questions contact Blue Belt Technologies, Lindsey Bring (+1 763.452.4922) or Robin Gonzalez (+1 763.452.4910) and get a Return Authorization number. Lindsey or Robin will arrange a pickup of the product. BBT will send replacement product at no charge. |  | 60 | 60 cartons (20 pieces/carton) 40 in EU, 20 in US. | Worldwide Distribution and US Nationwide in the states of: CA and PA.United Kingdom. | Terminated | November 14, 2013 | 85 | 60 | Not_Computer
Z-1976-2013 | 65777 | appliance, fixation, spinal interlaminal | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Oasys Midline Occiput Plate-1) Small 2) Medium 3) Large 4) Large Long5) Mini.Allows for positioning directly on the midline keel of the occiput. | Stryker Spine Oasys Midline Occiput Plate | 1 | August 22, 2013 | 2013 | Stryker Spine | Stryker has received reports from customers indicating post-operativefracture of the pin that connects the tulip head to the plate body. |  | Stryker Spine sent a 'Urgent Medical Device Removal Letter'/'Customer Response Form' to Branch Managers, Agent Principals, Quality Contacts, Hospital OR Risk Managers, OR supervisor on 5/30/2013 via Fed Ex. Inventory should be examined and returned.The attached form should be completed and returned as well.Stryker Spine sent 'Urgent Medical Device Removal Letters' to Spinal Implant Surgeons on 6/20/2013 via Fed Ex. |  | 1536 | 1536 units (US) 880 (Foreign) | Nationwide Distribution. | Open |  |  | 1536 | Not_Computer
Z-1997-2013 | 65832 | reagent, occult blood | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Hemoccult ICT (Immunochemical Fecal Occult Blood Test) is a rapid, visually read, qualitative immunochemical chromatographic method for detection of human hemoglobin in fecal samples. | Hemoccult ICT Test Hemoccult ICT (Immunochemical Fecal Occult Blood Test | 2 | August 22, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has received a customer complaint that suggests trending ofpositive results has increased from approximately 4% to 1 0%. As a result of ourinvestigation, Beckman Coulter has confirmed an increase in sensitivity at theborderline positive level." This testing also demonstrated that the issue does not affect results reported asnegative. |  | Beckman Coulter sent an Urgent Product Correction letter dated July 9, 2013, to all affected customers. The letter was accompanied with a cover letter and response form, dated July 9, 2013. Customers were informed of the product under recall with product description, actions to take, contact information, the issue, impact, and resolution. Beckman Coulter conducted a risk assessment for this issue. Per the risk documentation, there is only a remote probability that use of the product will cause any adverse health consequences,and the product is not likely to lead to serious adverse health consequences.If you have any questions please call (714) 961-3777. |  | 68116 | 68,116 units | Worldwide Distribution - USA and the country of Canada. | Open |  |  | 68116 | Not_Computer
Z-2000-2013 | 65413 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00597201101FemurCR POROUS FEM COMP SIZE AML Rx, Sterile;00597201102FemurCR POROUS FEM COMP SIZE AMR Rx, Sterile; 00597201201FemurCR POROUS FEM COMP SIZE BML Rx, Sterile; 00597201202FemurCR POROUS FEM COMP SIZE BMR Rx, Sterile; Used in total knee arthroplasty and indicated for patients with severe pain. | NexGen Complete Knee Solution | 2 | August 22, 2013 | 2013 | Zimmer, Inc. | Zimmer continuing investigation and monitoring of the NexGen Cruciate Retaining (CR) Complete Knee System found continued high rates of components used in combinations that are not approved (micro with standard) were implanted. This included the usage of CR Micro Femurs, CR Micro Articular Surface, CR Micro Patellae, CR Micro instruments and provisionals combined with standard components. Alth | DESIGN: Labeling Design | Beginning June 7, 2013 all distributors will be sent Urgent Device Recall notices via electronic mail, with distributors, hospital risk managers with affected inventory also receiving the notice via courier. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product is to be returned to Zimmer Product Service Department , 1777 West Center St. Warsaw, IN 46580 or internation return at Zimmer International Logistics GmbH, Max - Immelmann Allee 12 79427 Eschbach Germany."Communications outside of the United States will occur approximately two weeks after the United States communications. |  | 192355 | 192,355 all devices | Worldwide Distribution-USA (nationwide) and the country of Japan. | Open |  |  | 192355 | Not_Computer
Z-1996-2013 | 65738 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II CRUCIATE RETAINING ARTICULAR INSERT, TRIAL, SIZE 3-4, A/P 48 MM, M/L 68 MM, QTY 1, NON-STERILE, REF 71430490, Product Usage:Orthopedic | GENESIS(R) II CRUCIATE RETAINING ARTICULAR INSERT, TRIAL | 2 | August 22, 2013 | 2013 | Smith & Nephew Inc | The products were produced using the wrong color material, resulting in the devicesbeing red (used for deep dished inserts) instead of blue (used for cruciate retaining inserts). | PRODUCTION CONTROLS: Labeling Mix-Ups | Smith & Nephew sent an Urgent Recall Notification letter dated July 8, 2013 to all affected customers. Surgeons were notified by letter beginning on July 16, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, locate unused devices, quarantine immediately, and return affected product to the address provided in the notification. For questions email to FieldActions@smith-nephew.com. | Contact the recalling firm for information | 9 | 9 units | International only: Germany, France, United Arab Emirates, and United Kingdom. | Open |  |  | 9 | Not_Computer
Z-2050-2013 | 65841 | enzyme immunoassay, cannabinoids | LDJ | Cannabinoid test system. | Toxicology | Toxicology | 510(k) | SARKEN STAT! 1- Panel Test Kit-THCProduct Usage- Labeled as FOR FORENSIC USE ONLY (Employee, Student or Probation Testing) . | SARKEN STAT! 1 Panel Test KitTHC | 2 | August 23, 2013 | 2013 | Sarken, Inc. | These products are being recalled because they were distributed without adequate labeling and instructions for use. Inadequate labeling could lead to misuse of the Test Kits and incorrect test results. | PREMARKET APPROVAL: No Marketing Application | Firm notified customers via telephone, followed up with MEDICAL DEVICE RECALL letter dated July 29, 2013 sent 08/02/13, via USPS regular mail to all affected customers. The recall notification letter is accompanied by a return response, acknowledgement and receipt form. The letter notify customers of the voluntary recall, product description (with prodcut images), reason for recall, potential impact, instructions/actions to follow, and contact information. |  |  | various quantities | Worldwide Distribution - US Nationwide in the country Virgin Islands | Open |  |  |  | Not_Computer
Z-2045-2013 | 65810 | intracranial aneurysm flow diverter | OUT |  |  | Neurology |  | Pipeline Embolization Device (PED)The Pipeline Embolization Device is indicated for the endovascular treatment of adults (22 years of age and older) with large or giant wide-necked intracranial aneurysms in the internal carotid artery from the petrous to the superior hypophyseal segments. | Pipeline Embolization Device | 2 | August 23, 2013 | 2013 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Covidien Neurovascular has initiated the recall of several lots of Pipeline Embolization Device (PED) because the Instructions For Use (IFU) do not include the following contraindication: Patients in whom a pre-existing stent is in place in the parent artery at the target aneurysm location. Additionally, the designation of aspirin and clopidogrel was missing from the following contraindication " | TRAINING: Employee Error | Covidien sent an Customer Notification letter dated July 19, 2013, to all affected customers who received affected lots of Pipeline Embolization Device (PED). The customer notification letter informed the customers of the problems identified and the actions to be taken. Customer are instructed to reply and fax response form to Covidien Customer Service at (763) 398-7001. Customers are instructed that Customer Service will issue an RGA number for the return of the affected product back to Covidien. Customers are also instructed that upon receipt of a Returned Goods Authorization (RGA) number, to write theRGA number on the FEDEX slip and send the product back to:CovidienAttention: Product Recall14005 13th AvePlymouth, MN 55441Prepaid and preaddressed FEDEX slips are enclosed for their convenience. Once Covidien receives the returned units, product replacement units will be shipped directly to customers. Local Covidien Neurovascular representative is also available to assist them in expediting the exchange of product.Customers with any questions, are instructed to call (949) 839-3700. |  | 62 | 62 units | US Distribution including the states of AZ, KY, MO, NC, GA, WI, OR, FL, NY, CA, IL, TN, NV, TX, CO, MA, IA, UT and VA. | Terminated | October 17, 2013 | 55 | 62 | Not_Computer
Z-2047-2013 | 65609 | [6 empty] | PSA Assays for use in the IMMULITE/IMMULITE 1000 Systems Analyzers, Catalog Numbers LKPS1, LKPS5 and LKPTS1(D); and in the IMMULITE 2000/2000XPi Systems Analyzers, Catalog Numbers L2KPS2, L2KPS6, L2KPTS2(D) and L2KPTS6(D). IVD Immunoassay kits in the following sizes: 100, 200, 500 and 600 test kits; For in vitro diagnostic use with the IMMULITE/IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi Systems Analyzers. PMA # P930027/ Device Listing # D002374.Product Usage:For the quantitative measurement of prostate-specific antigen (PSA) in human serum, as an aid in the detection of prostate cancer when used in conjunction with digital rectal examination (DRE) in men aged 50 years or older. This assay is further indicated as an adjunctive test to aid in the management of prostate cancer patients. For in vitro diagnostic use with the IMMULITE/IMMULITE 1000 and the IMMULITE 2000/ 2000 XPI Systems Analyzers | IMMULITE | 2 | August 23, 2013 | 2013 | Siemens Healthcare Diagnostics | Customers complained of a positive bias in patient results on the IMMULITE/ IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi PSA assay when compared against the IMMULITE /IMMULITE 1000 3rd Generation PSA assay. Siemens has confirmed an overall average positive bias of approximately 20% to 23% across the assay range relative to the WHO 96/670 with the IMMMULITE Systems PSA Assay. This positive bias i |  | Seimens sent an Urgent Medical Device Recall Letters to US consignees via FedEx and an Urgent Field Safety Notices were sent to the foreign consignees starting on June 26, 2013. In addition, Support Bulletins were distributed to Siemens Healthcare Diagnostics service personnel describing the issue and instructing them how to deal with customer questions. Customers were informed to discontinue use, discard the kits remaining in inventory and there are no replacement PSA kits available at this time. Siemens is instructing customers to contact their local Siemens representative for assistance with determining appropriate PSA testing solutions for your laboratory. |  | 24361 | 24,361 kits | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: Albania, Afghanistan, Algeria, Argentina, Austria, Bangladesh, Belgium, Bolivia, Bosnia- Herzegovinia, Brazil, Bulgaria, Camaroon, Canada, Canary Islands, Chile, China, Colombia, Coast Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Nicaragua, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, South Africa, Rep. of Korea (S), Sweden, Switzerland, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, Uruguay, White Russia, Venezuela, and Vietnam. | Open |  |  | 24361 | Not_Computer
Z-2085-2013 | 65770 | chromogenic in situ hybridisation, nucleic acid amplification, her2/neu gene, breast cancer | NYQ |  |  | Immunology | PMA | HSV 2 Primers, Catalog #18081020. Analyte Specific Reagent. The HSV-2 primer pair is specific for HSV-2 (Herpes Simplex Virus 2). For use in Clinical laboratories when developing their own diagnostic tests. | HSV2 Primers | 2 | August 26, 2013 | 2013 | Life Technologies Corporation | Firm received complaints that HSV-2 primers failed to amplify HSV-2 DNA. | DESIGN: Process Design | Life Technologices sent an Urgent Product Recall letter dated May 13, 2013 by certified mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use and dispose all affected units. Customers were advised to complete the Customer Response Sheet included for product/quantities consumed, and send e-mail to techsupport@lifetech.com. For questions about the details of this product issue, were instructed to contact Technical Support at 800-955-6288 option 4. | Contact the recalling firm for information | 15 | 15 kits | USA Nationwide Distribution in the states of DE, IA, MN, and NC. | Terminated | August 28, 2013 | 2 | 15 | Not_Computer
Z-2094-2013 | 65513 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | VERSA-FX II FEMORAL FIXATION SYSTEM COMPRESSION TUBE/PLATE, standard and short - low profile, different hole/length combinations (4/86, 5/102 and 6/118), different degree angles (130, 135, 140, 145 and 150). Item codes starting with 00-194-130, 00-1194-135, 00-1194-140, 00-1194-145, 00-1194-150)., orthopedic implant. | Free Lock Femoral Fixation System | 2 | August 26, 2013 | 2013 | Zimmer, Inc. | Single packaging configuration used for the VERSA-FX-II Femoral Fixation System failed the leak test compromising the sterility of the product. | PRODUCTION CONTROLS: Packaging Process Control | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 11, 2013 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect the device according to the insert instructions; not use the device if damage to the corner of the inner tray is observed, and contact the firm for return and questions. Customers should also keep this notification in mind and inform Zimmer of product with this condition using zimmer.per@zimmer.com.If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. | Contact the recalling firm for information | 794 | 85, 794 units in total | Worldwide distribution: US Nationwide (including PR and USVI), and countries of: Israel, Argentina, Brazil, France, Mexico, Saudi Arabia, Beirut, China, Korea, San Salvador, Honduras, Canada, Australia, Switzerland, Germany, Arab Emirates, UK, Belgium, Egypt, Spain, France, Italy, Iran, Lebanon, Jordan, Iraq, Canary Islands. | Open |  |  | 794 | Not_Computer
Z-2056-2013 | 65824 | drills, burrs, trephines & accessories (manual) | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Leyla Ball Joint Clamp, also as included in Leyla Table Mounting Hardware. The Leyla Retraction System, as part of the Ruggles brand of instruments is intended for use as a self-retaining retractor which is a self-locking device used to hold the edges of a wound open during neurosurgery. | Integra Leyla Ball Joint Clamp | 2 | August 26, 2013 | 2013 | Integra Limited | As a result of complaint investigations it was identified that there is a potential that the Leyla Ball Joint Clamp (R2383) may not provide the level of stability needed to maintain retraction in all cases. | DESIGN: Device Design | Integra sent an Urgent Medical Device letter dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to review your inventory and determine if you have any of the affected products. If so, stop using them immediately and remove them from your service. When your form is received, Customer Service will you an RMA number and directions for returning the products.Should you have any questions regarding these instructions, please contact Integra Customer Service at 1-800-997-4868. | Contact the recalling firm for information | 65 | 65 | Worldwide Distribution - USA including Puerto Rico and the states of AZ, CA, FL, GA, IL, IN, KY, MA, MI, MS, NE, NM, NY, OH, PA, TN, TX, UT, VA, WA, WI and WV., and the countries of Spain, Singapore, South Africa, Canada, Libya,Pakistan and Cayman Islands. | Open |  |  | 65 | Not_Computer
Z-2061-2013 | 65800 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | TRIAL STEM for MP(R) RECONSTRUCTION PROTHESIS, REF 99-0142/14, Stem size 001, prox. Stem dia. 14 mm, Length 210 mm, WALDEMAR LINK GmbH & CoKG, Barkhausenweg 10 . D . 22339 Hamburg | TRIAL STEM for MP(R) RECONSTRUCTION PROTHESIS | 2 | August 26, 2013 | 2013 | Waldemar Link GmbH & Co. KG (Corp. Hq.) | The trial stems were manufactured using the wrong design. | DESIGN: Device Design | The firm notified their consignees of their product recall by parcel post on 07/16/2013. | Contact the recalling firm for information | 5 | 5 units | US Distribution to the state of Tennessee only. | Open |  |  | 127 | Not_Computer
Z-2096-2013 | 65874 | reagent, general purpose | LDT | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | OLA BufferFor Laboratory Use;Contents: 1 tube with buffered solution containing MgCl2 (80 ÂŽM) and KCL (300 ÂŽM. 10.0ml.)PN: 50003300;Store at d - 15ÂŽC.Product Usage: General Purpose Reagent for laboratory use. | Oligonucleotide Ligation Assay Buffer | 3 | August 26, 2013 | 2013 | Celera Corporation | The package insert for FG OLA Buffer, GPR displayed the wrong concentration of ÂżM instead of the correct mM unit. | PRODUCTION CONTROLS: Error in Labeling | Celera sent a customer notification letter dated July 12, 2013 by e-mail and mailed on July 19, 2013 via first class mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The information is being provided for records and assures customer that there were not quality issues with the product and it may be used as usual in laboratory test procedures. For questions call 1-810-749-4200 or contact your local authorized Celera representative. |  | 43 | 43 units | USA Nationwide in the state of CA | Terminated | October 01, 2013 | 36 | 43 | Not_Computer
Z-2083-2013 | 65773 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BRAINLAB; FRAMELESS SRS QA TARGET POINTERRobotics is a device used to compensate rotational patient misalignment (roll and pitch) in a linear accelerator environment for stereotactic radiosurgery or radiotherapy procedures. The Frameless Radiosurgery Components are a device used for fixation, localization and repositioning of the patient's: head and neck; and head, neck, and shoulders, in a linear accelerator environment for stereotactic radiosurgery/radiotherapy procedures. | 43410 FRAMELESS SRS QA TARGET POINTER | 2 | August 26, 2013 | 2013 | Brainlab AG | The Frameless SRS QA Target Pointer - Pointer Cap with the engraved cross hairs may become loose even when only a moderate force is applied to the Pointer Cap. | DESIGN: Device Design | BrainLab sent an "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" dated July 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue use of the SRS QA Target Pointer. Consignees can instead use the Phantom Pointer feature until Brainlab can provide replacement hardware. Brainlab anticipates being able to provide all impacted facilities with replacement hardware by the end of July 2013.We sincerely apologize for any inconvenience and thank you in advance for your co-operation. If you require further clarification, please feel free to contact your local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33. | Contact the recalling firm for information | 6 | 6 units | Worldwide Distribution - US Distribution only to AR., and the countries of Australia, Germany and Japan. | Open |  |  | 6 | Not_Computer
Z-2081-2013 | 64852 | susceptibility test discs, antimicrobial | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs F300 (Nitrofurantoin), Cat. #CT0036B, packaged in cartridges of 50 discs, 5 cartridges/pack. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, EnglandProduct is used in the semi-quantitative agar diffusion test method for in-vitro susceptibility testing. | Oxoid Antimicrobial Susceptability Testing Discs F300 (Nitrofurantoin) | 3 | August 26, 2013 | 2013 | Remel Inc | Use of affected lots may give false indication of susceptibility to Nitrofurantoin. |  | ThermoFisher sent an Medical Device Recall letter dated March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please notify any personnel who needs notified of the potential for a false indication of susceptibility to Nitrofurantoin. Please review your inventory for affected product lots and discard the affected product. Return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913- 888-0939 (International). | Contact the recalling firm for information | 60 | 60 packs | US Distribution including the states of AL, AR, CA, FL, GA, IL, KS, LA, MA, MD, NY, OK, PA, RI, TX, VA, and WA. | Open |  |  | 60 | Not_Computer
Z-2057-2013 | 65556 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm VA-LCP Two Column Volar Distal Radius Plate and Torque limiting attachment, 1.2NmSynthes VA-LCP Distal Radius System is intended for fixation of complex intra- and extra-articular fractures and osteotomies of the distal radius and other small bones. | Synthes 2.4mm VALCP Two Column Volar Distal Radius Plate and Torque limiting Attachment, 1.2Nm | 2 | August 26, 2013 | 2013 | Synthes USA HQ, Inc. | The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard & Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm | DESIGN: Device Design | Synthes sent an Urgent Notice: Medical Device Labeling Correction letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard & Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm VA-LCP Two Column Volar Distal Radius Plate when used with a 1.2Nm Torque Limiting Attachment. A screw that passes through the VA-LCP plate is not likely to result in a serious adverse event although should it occur, additional medical intervention during the index procedure could be needed.If you have any questions, please call (610) 719-5450, or contact your Synthes sales Consultant. | Contact the recalling firm for information | 1010 | 119, 891 | Nationwide Distribution including the states of AK. AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MD, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI SC, SD, TN, TX, UT, VA, VT, WA WI and WV. | Open |  |  | 119891 | Not_Computer
Z-2058-2013 | 65631 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Nail System. The devices are indication for bone fixation. | Synthes Trauma Nail System | 2 | August 26, 2013 | 2013 | Synthes USA HQ, Inc. | Recall was initiated due to the possibility that the outer pouch was compromised. The product is packaged in two pouches; a sterile, internal pouch within an external pouch. It is the outer, external pouch that may be compromised. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | The firm initiated their recall of this product on November 2, 2012 by sending an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter to all consignees. The letter explained the problem and gave instructions for consignees to examine their inventory and urged consignees to return the identified devices. The letter stated If returning the product would adversely impact their ability to provide necessary medical care to patients, they should re-sterilize the product per Pre-Vacuum Sterilization instructions contained in the Instruction for Use. | Contact the recalling firm for information | 1011783 | 1,011,783 | Nationwide distribution. | Open |  |  | 1011783 | Not_Computer
Z-2087-2013 | 65890 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 VA-LCP 2 CLMN VLR DSTL Radius PL 7H HD/3H Shaft/RightProduct Usage: Intended for fixation of complex intra-and extra-articular fractures and osteotomies of the distal radius and other small bones. | Synthes 2.4 VALCP 2 CLMN VLR DSTL Radius PL 7H HD/3H Shaft/Right | 2 | August 26, 2013 | 2013 | Synthes USA HQ, Inc. | The possibility exists that 2.4 mm VA-LCP 2 column distal radius plates from a specific lot are etched with "R" for right plate when they should be etched with "L" for left plate. | PRODUCTION CONTROLS: Process Control | Synthes GMBH sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected product from stock. Customers were advised to complete the Verification Section at the end of the letter. For questions call 610-719-5450 or contact your Synthes Sales Consultant. | Contact the recalling firm for information | 11 | 11 | USA Nationwide Distribution in the states of FL, GA, NC, OH, KY, NJ, ID, IN, and MN. | Open |  |  | 11 | Not_Computer
Z-2082-2013 | 65080 | dialyzer, high permeability with or without sealed dialysate system | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 MeDS Pump, Venofer(R) PumpProduct Usage: The 2008 MeDs pump is an optional accessory for use with the Fresenius 2008 Series Hemodialysis Machines and is intended to administer Venofer to treat iron deficiency anemia in patients with chronic kidney disease undergoing hemodialysis, where Venofer, in conjunction with hemodialysis, is prescribed by a physician. | 2008 MeDS Pump, Venofer(R) Pump | 2 | August 26, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. dba Renal Solutions | Product is being recalled because the combination of the current 50 mg vial, the 2008 MeDS pump, and the 2008 MeDS pump tubing Set for CombiSet Bloodlines can lead to a situation where Venofer(R) delivery does not meet dosage accuracy requirements. This may lead to a slight over or under delivery of Venofer(R) that does not pose an acceptable risk to the patient. Please note that this recall | DESIGN: Device Design | Renal Solutions sent an Urgent Notification letter dated April 1, 2013 via certified mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to cease further use of the device. The letter informed customers that a Fresenius Medical Care Technical Service/Customer Service representative will contact customers to arrange for removal of the device. For questions contact Fresenius Medical Care Customer Service at 800-227-2572. | Contact the recalling firm for information | 132 | 132 | USA Nationwide Distribution in the states of CA, MA, and NY. | Open |  |  | 132 | Not_Computer
Z-2054-2013 | 65787 | methyl methacrylate for cranioplasty | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Stryker Part Numbers 79-43903 (3cc), 79-43905 (5cc), 79-43910 (10 cc), 79-43915 (15cc), 397003 (3cc), 397005 (5cc), 397010 (10 cc), 397015 (15 cc) HydroSet Injectable HA Bone Substitute Sterile R/EO, nonpyrogenic, Product of IrelandStryker Injectable Cement is a self-setting, calcium phosphate cement intended for use in the repair of neurosurgical burr holes, contiguous craniotomy cuts and other cranial defects as well as in the augmentation or restoration of bony contour in the craniofacial skeleton. The powder is mixed with the sodium phosphate/PVP solution provided in the kit and loaded into the empty cement delivery syringe to be applied directly to the site for void filling or augmentation. | Hydroset | 2 | August 26, 2013 | 2013 | Stryker Craniomaxillofacial Division | Stryker received reports of the desiccant bag in the kit leaking silica gel granulate. This could increase exposure to anesthesia during surgical delay up to 60 minutes or inadvertent implantation of silica gel granulate which could cause foreign body reaction including inflammation and might lead to the need for revision surgery. | DESIGN: Packaging Design/Selection | Stryker sent an Urgent Medical Device Recall Notices with attached Product Correction Bulletin dated, July 30, 2013, describing the product, problem and possible adverse events associated with the recall were sent to all affected customers. Customers were instructed to pass the notice and bulletin to all individuals who may be using the device, provided instructions on procedures which would help mitigate risk and return the customer response form to our Customer Service Team by email, fax or letter.Email: CMF-custserv@stryker.comFax: 1-877-648-7114Address: 750 Trade Centre Way, Suite 200Portage, MI 49002 Customers responses will be tracked and non-responders will be contacted.Further questions please call (877) 534-2464. |  | 12719 | 12,719 | Worldwide Distribution - USA (nationwide) including the countries through Stryker Distribution: Austria, Belgium, Denmark, Estonia, Finland, France, Germany, Ireland, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Serbia, Spain, Sweden, Switzerland, United Kingdom, Argentina, Australia, Brazil, Jordan, Kingdom of Saudi Arabia, Kuwait, Latin America (others), New Zealand, Russia, Singapore, South Africa, and United Arabian Emirates. | Open |  |  | 12719 | Not_Computer
Z-2055-2013 | 65677 | prosthesis, shoulder, non-constrained, metal/polymer cemented | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00 4309-049-00 Trabecular Metal Reverse Shoulder Glenosphere Distractor Rx, Non sterile (sterilize before use)The Glenosphere Distractor must be able to attach, connect, or mate with the Glenosphere and remove it without dislodging or loosening the base plate from the glenoid, or damaging the glenoid. Pulling the trigger releases a weight inside the Distractor. This force is transmitted to the bottom of the glenosphere dislodging it from the base plate. | Trabecular Metal | 2 | August 26, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Reverse Glenosphere Distractor due to the potential of the device not functioning properly should it be utilized without proper lubrication. As a result, there is a potential for the device to fail to provide impact to the Glenosphere when the trigger is pulled. In addition, there is a potential for the foot of the device to fract | DESIGN: Device Design | Zimmer sent a Urgent Medical Device Recall Notification letter August 1, 2013, and via E-mail and courier to all affected customers. The initial recall notice identified the reason for the recall, possible adverse risks and injuries and informed customers that a replacement will be available in the coming months. In addition, this letter will remind that the instrument is to be lubricated before each use and to inspect the foot. Customers are instructed to inspect any recalled product before and after use for cracks at the foot and ensure the foot is intact. Also ensure all moving parts are lubricated prior to each use according to the Recommendations for Care, Cleaning, Maintenance, and Sterilization Manual 97-5000-170-00. If you find a device which does not function properly or has a damaged foot, provide it to your Zimmer sales representative for return to Zimmer on a Product Experience Report (PER). Questions and concerns should be addressed to the Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. Phase 2: As devices become available for distribution, a 2nd letter will be sent to the distributors and risk managers. Distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices will be included. | Contact the recalling firm for information | 771 | 771 | Worldwide Distribution - USA (nationwide) including the countries of Singapore, Brazil, Taiwan, Australia, Canada, Germany(with further distribution to UAE, Austria, Belgium, Switzerland, Czech Republic, Denmark, Spain, Finland, France, Greece, Israel, Italy, Malta, Norway, Portugal, Sweden and South Africa, | Open |  |  | 771 | Not_Computer
Z-2095-2013 | 65848 | accelerator, linear, medical | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component Tabletop, 550TxT, TT-A s/n 1001-1233 and Tabletop 550TxT, TT-S s/n 1001-1083. Product Usage: The intended use of the SIEMENS branded, ARTISTE , ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | PRIMUS, ONCOR, ARTISTE Linac systems | 2 | August 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a customer complaint concerning the rail threaded holes of the 550TxT treatment tabletop. Due to a manufacturing error, the screws holding the side rails in position on the tabletop might break off. | DESIGN: Device Design | Siemens released Updated Instruction containing the Medical Device Correction Field Safety Notice of Improvement letter, which was delivered by the Siemens Customer Service Group to all affected customers. The release of this UI TH002/12/S to distribute the update instruction began on July 18, 2013. This document includes completion protocol, which will be completed at each facility and returned to Siemens. A copy of the notification should be kept in the Digital Linear Accelerator System owners manual under Safety Advisory letters.UPDATE: a secnd update intruction, TH003/12/S has been released to begin the replcement of tabletops identified through the inspection carried out through TH002/12/S. |  | 179 | 179 active devices. | Worldwide Distribution - USA Nationwide in the states of NJ, IN, PA, FL, OH, MN, MO, CA, AL, NE MI, LA. and countries of:Argentina, China, South Africa, Pakistan, Spain, Switzerland, Italy, Algeria, Croatia, Turkey, Canada, Germany, France, Poland, Japan, Russian Federation. Lebanon, Egypt, Slovakia, Latvia, Kenya, Brazil, Kuwait, Saudi Arabia, Greece, Belgium, Taiwan, Portugal. | Open |  |  | 179 | Not_Computer
Z-2086-2013 | 65672 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Illico Ti Cannulated Polyaxial Screw 7.5mm X 40mm. The Illico MIS Posterior Fixation System is intended to facilitate the surgical correction of noncervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. | lllico Ti Cannulated Polyaxial Screw | 2 | August 27, 2013 | 2013 | Alphatec Spine, Inc. | The firm initiated this recall of certain lots of Illico Ti Cannulated Polyaxial Screws 7.5mm x 40mm implants, due to an off-center cannulation hole at the proximal end of the bone screw. There is a potential risk that when the screw is mated to a polyaxial screwdriver, the guide wire may not pass through the screwdriver. | PRODUCTION CONTROLS: Equipment Maintenance | The firm, Alphatec Spine, sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" with response sheet dated June 28, 2013 via a combination of e-mail, Fedex, and telephone contact to its customers. The notification described the product, problem and actions to be taken. The customers were instructed to immediately return any and all affected lots of the illico Ti Cannulated Polyaxial Screw 7.5 mm x 40 mm implants that have been sent to them; if the device has previously been implanted, there is no further or additional risk post-implantation; immediately contact Alphatec Spine Customer Service Department - customerservice@alphatecspine.com on how to replace the implants; and complete and return the last page of the notification via mail to Alphatec Spine, Inc. 5818 El Camino Real, Carlsbad, CA 92008; email: sarrhenius@alphatecspine or fax to: (760) 431-0289.If you have any questions or need any additional information, please contact Director of Compliance at 760-494-9636 or email at dmcgrath@alphatecspine.com. | Contact the recalling firm for information | 333 | 333 screws | Worldwide distribution: US (nationwide) including states of:US states: AZ, CA, FL, GA, IL, MD, NJ, NY, OK, PA, TN, TX,and WA ; and countries of: Spain, France, Japan.International: Spain, France, Japan. | Open |  |  | 333 | Not_Computer
Z-2099-2013 | 65891 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker 1) Asnis III Cannulated Screw REF 6046404.0x40mm TL 13.5mm2) Asnis III Cannulated Screw REF 604642S4.0x42mm TL 14mm, sterileProduct Usage: The Asnis III Cannulated Screw System is intended for fracture fixation of small and long bones and of the pelvis. The system is not intended for spinal use. | Asnis III Csnnulated Screw System | 2 | August 27, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report that the length of the cannulated screw asnis III, article # 604640 Lot R11482, was incorrect. | PRODUCTION CONTROLS: Process Control | Stryker sent an Urgent Medical Device Removal letter dated July 8, 2013 via Fed Ex to all affected customers. The letter identified the affected product, product issues, potential hazards, risk mitigation and actions to be taken. Customers were instructed to return all affected products available to Stryer Osteosynthesis or contact Stryer customer service and complete and sign the enclosed Business Reply Form and fax a copy to: 1-865-251-3635. Questions can be directed to Stryker at 201-972-2100. |  | 23 | 23 units | USA Nationwide Distribution in the states of: FL. OH, TX, & VA | Open |  |  | 23 | Not_Computer
Z-2107-2013 | 65914 | needle, aspiration and injection, disposable | GAA | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | SpineNet Bone Marrow Aspiration Needle Kit (which is the Ranfac Bone Marrow Aspiration Needle #74150-01M).Use to aspirate bone morrow. | SpineNet Bone Marrow Aspiration Needle Kit | 2 | August 27, 2013 | 2013 | SpineNet | SpineNet was marketing this device to be used with SpineNet ACC device without FDA clearance or approval. |  | SpineNET sent a "VOLUNTARY PRODUCT RECALL" letter dated 02/04/2013 to Halifax Medical Center. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 407-539-2483 for questions regarding this notice. |  | 100 | 100 units | Distributed to one customer in Florida. | Terminated | November 15, 2013 | 80 | 100 | Not_Computer
Z-2103-2013 | 65917 | intervertebral fusion device with bone graft, cervical | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineNet Anterior Cervical Cage, Lordotic, Peek (originally submitted as Daytona Anterior Cervical Cage)Product Usage: Is indicated for use in skeletally mature patients with degenerative disc disease of the cervical spine with accompanying radicular symptoms at one disc level | SpineNet Anterior Cervical Cage | 2 | August 27, 2013 | 2013 | SpineNet | The recall decision has been made because because it has recently been determined that the instructions for use (IFU) had certain statements that should have been removed prior to initial product marketing. | PRODUCTION CONTROLS: Error in Labeling | SpineNET sent a Customer Notification letter dated 02/05/2013 to their only customer (Halifax Medical Center). The letter identified the affected product, problem and actions to be taken. Customers were informed to remove any copies of the original IFU-ACC-01 form from their records and replace with the IFU-ACC-02 form attached to the letter. For questions contact SpineNET office directly. | Contact the recalling firm for information | 400 | 400 | USA Nationwide Distribution in the state of Florida | Open |  |  | 400 | Not_Computer
Z-2100-2013 | 65913 | crown, preformed | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Primary Stainless Steel Crowns (First Primary UL4), Product code 902150 (Kit- 112 crowns/kit) and 900114 (Refill- 5 crowns/box).Product Usage:Stainless Steel Crowns are designed to fill the need for long term provisional coverage on adult molars and bicuspids as well as pediatric posterior and anterior teeth. | 3M ESPE Unitek Primary Stainless Steel Crowns | 2 | August 27, 2013 | 2013 | 3M Company/3m Espe Dental Products | 3M ESPE is recalling Unitek Primary Stainless Steel Crowns (First Primary UL4) because these products have been manufactured with the incorrect marginal contour edge. | PRODUCTION CONTROLS: Process Control | 3M ESPE sent a 3M "Voluntary Recall" letter dated August 1, 2013 to all affected customers. Separate letters were addressed to distributors and doctors. The letters described the product affected by the recall, problem and actions to be taken. The doctors were advised to complete the attached inventory checklist and call the 3M ESPE Dental Products Customer Care Center at 1-800-634-2249. Select option #1 to arrange for return and replacement of your product.Distributors were advised to quarantine and return the product and not to further distribute it. For product return they should contact 3M ESPE Dental Products Customer Service at 1-800-237-1650. If the product was further distributed they ask that they inform them of the recall. A letter was prepared that can be used for their communication to customers. for additional questions they should contact 651-733-7767. |  | 3409 | 3409 crowns (2949 US and 460 OUS) | Worldwide Distribution - USA Nationwide in the states of CA, FL, GA, IN, IA, MO, NV, NJ, NY, PA, SC, TN, TX, WA. and the country of CANADA. | Terminated | December 11, 2013 | 106 | 3409 | Not_Computer
Z-2097-2013 | 65995 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | E-Z CleanÂŽ Disposable Laparoscopic Electrode Model Number 0100L and 0100LS. Expiration dates of 05/2018 to 07/2018. Sterile. RX ONLY. This product is packaged individually in a chevron TyvekÂŽ - polyester peel pouch, 6 pouches per box.The E-Z Clean laparoscopic electrode is intended to conduct monopolar electrosurgical energy from an electrosurgical generator to target tissue during laparoscopic surgical procedures. This device is intended to be used whenever monopolar electrosurgical cutting and coagulation are indicated. | EZ Clean Disposable Laparoscopic Electrode | 2 | August 27, 2013 | 2013 | Megadyne Medical Products, Inc. | Megadyne Medical Products is recalling certain lots of E-Z Clean, L-Wire laproscopic Electrode Standard Stem and Split Stem due to potential breach of insulation. | PRODUCTION CONTROLS: Process Control | Megadyne sent an Urgent: Medical Device Recall Field Safety Notice dated August 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were also notified with follow-up phone calls and email beginning on August 7, 2013. Customers were instructed to examine their stock immediately to determine if you have any of the products from the affected lots on hand. If so, please remove form service/discontinue use and contact Megadyne Customer Service at 1 (801) 576-9669 or (800) 747-6110 (USA) for return authorization and replacement or credit. |  | 1816 | 1816 | Class II Recall - Worldwide Distribution - US including the states of CA, HI, IL, MN, NC, NJ, OH, PA, TX, UT and VA. | Terminated | January 28, 2014 | 154 | 1816 | Not_Computer
Z-2106-2013 | 65831 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion System with Trillium BioSurface (a polymer containing non-leaching heparin). Model Numbers: TL4X17R8, TL5D01R, TL5D02R, TL5D02R3, TL5D02R5, and TL6A65R7. Sterilized by Ethylene Oxide, Do no Reuse. Product Usage: Usage:This product is indicated for use in the extracorporeal circuit during cardiopulmonary surgery. | Medtronic Custom Perfusion System with Trillium BioSurface | 2 | August 27, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic was notified that Edwards Lifesciences has initiated a product recall for the Rigid Suction Wand, model S099B. The Wands are included within certain Perfusion Tubing Packs that are manufactured and distributed by Medtronic. | PRODUCTION CONTROLS: Process Control | Metronic sent an "Urgent Field Safety Notice" letter dated July, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken by User. Customerss were advised to remove and destroy all affected products. At the same time customers were informed that replacement products can be ordered by contacting their Medtronic representative. Customers were requested to complete the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. For any questions they can contact Medtronic Lifeline Technical Services at 877-526-7890 or their Medtronic representative. |  | 758 | 758 | Worldwide Distribution - USA Nationwide CA, IL, LA, MI and the country of CANADA | Terminated | February 04, 2014 | 161 | 14198 | Not_Computer
Z-2101-2013 | 65850 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ZIMMER PERIARTICULAR LOCKING PLATE SYSTEMDISTAL MEDIAL HUMERAL LOCKING PLATE, SHORTRIGHT 11 HOLES 178 MM LENGTHSTERILE RX ONLYProduct Usage: Usage:The Periarticular Locking Plate System is indicated for temporary internal fixation and stabilization of osteotomies and fractures, including: Comminuted fractures, supracondylar fractures, intra-articular and extra-articular condylar fractures, fractures in osteopenic bone, nonunions, and malunions. | Zimmer Periarticular Locking Plate System, Distal Medial Humeral Locking Plate, Short | 2 | August 27, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific voluntary recall of the Distal Medial Humeral Locking Plate, Short due to the incorrect information printed on the Chinese label. The original label indicates that implants are RIGHT sided, but the Chinese label indicates that the implants are LEFT sided. | PRODUCTION CONTROLS: Labeling Mix-Ups | Zimmer sent an Urgent Medical Device Recall Notification letter dated July 30, 2013 to Distributors and Hospitals in China. The letter described the affected product, problem, labeling issue, directions on how to respond if affected product is discovered, and potential health risks associated with the labeling error. An inventory return certification form is included with the recall notification. Questions or concerns are directed to Zimmer Shanhai at 021-2211 5196, Monday to Friday, between the hours of 9 a.m. and 5 p.m. | Contact the recalling firm for information | 11 | 11 units | Internationally Distributed in China only. | Open |  |  | 11 | Not_Computer
Z-2109-2013 | 65886 | lancet, blood | FMK | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | ACCU-CHECK FastClix Lancing Device. It is intended for use in conjunction with a blood glucose monitoring system. | ACCUCHECK FastClix | 2 | August 28, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Internal Roche inspections have revealed that after the lancet is used for a finger stick, the needle may not fully retract back within the lancet drum or back into the cap. This allows the needle to protrude inside the cap or outside of the cap of the ACCU-CHEK FastClix lancing device. | PRODUCTION CONTROLS: Equipment Maintenance | On 8/19/ 2013, Roche sent Distributor/Retailer URGENT MEDICAL DEVICE CORRECTION notifications which described the recall issue and health consequences, provided product description, and provided actions required of consignees. Roche sent a CUSTOMER BULLETIN to their direct customers. URGENT MEDICAL DEVICE CORRECTION faxback form for consignee acknowledgement of recall is to be faxed to 1-888-240-2804. Consignees may contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 if they have questions about the information contained in the Distributor/Retailer Letter. |  | 36332 | 36,332 | Nationwide Distribution | Open |  |  | 36332 | Not_Computer
Z-2114-2013 | 65852 | counter, cell, automated (particle counter) | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | ADVIA 120 Hematology Analyzer, Catalog Numbers SMN 10360958 and 10360959 .The ADVIA 120 is a hematology system that utilizes the principles of flow cytometry in order to provide complete blood counts. | ADVIA 120 | 2 | August 28, 2013 | 2013 | Siemens Healthcare Diagnostics | The ADVIA 120 optics cover or hood is located on the very top of the ADVIA 120 analyzer. When the optics cover is raised, it is supported by two gas struts (SMN 10309266) attached to the cover on the left and right sides. Over time, it is possible that these struts may lose their effectiveness and fail to support the optics cover in the open position. This could cause the cover to fall during m | PRODUCTION CONTROLS: Equipment Maintenance | Siemens sent an Urgent Device Correction Letter dated June 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the ADVIA 120 system displays a problem when raising and supporting the optics cover, immediately contact your service provider to arrange to have the cover struts serviced. If the optics cover will not stay in the up position, as required, do not perform maintenance procedures on the optics area until the hood struts have been replaced. Siemens field support will be proactively replacing the optics cover hood struts on a regular preventive maintenance cycle based on the age of the system.Please forward this notification to whomever you may have distributed this productWe apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support.This response form is to confirm receipt of the enclosed Siemens Healthcare Diagnostics Urgent Device Correction dated June 2013 regarding Optics Cover Struts (SMN 10309266).Please read each question and indicate the appropriate answer. Fax this completed form to Siemens Healthcare diagnostics at the fax number indicated at the bottom of this page.PLEASE FAX THIS COMPLETED FORM TO THE TECHNICAL SOLUTIONS CENTER AT (302) 631-7597. |  | 950 | Domestic: 950 units; Foreign: 3369 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Australia, Belgium, Bosnia - Herzegovina, China, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Ireland, Iraq, Israel, Italy, Japan, Kenya, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Martinique, Malaysia, Netherlands, Norway, New Zealand, Philippines, Poland, Portugal, Romania, Russian Republic, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Open |  |  | 4319 | Not_Computer
Z-2118-2013 | 49516 | slide stainer, automated | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark XT/LT and Discovery XTModel Numbers: N750-BMK-FS, N750-BMKXT-FS, N750-BMKLT-FS, N750-DIS-FS, N750-DISXT-FS.Product Usage: Usage:The Benchmark product lines are intended to automatically stain histological or cytological specimens on microscope slides with specific immunohistochemistry or in situ hybridization reagents for in vitro diagnostic use. The Discovery Systems are research use only instruments with similar functionality. | Benchmark & Discovery | 2 | August 28, 2013 | 2013 | Ventana Medical Systems Inc | Ventana Medical System is initiating the recall of the BenchMark and DISCOVERY series instruments because there is a possibility that the heating element could separate from the underside of the heater cap and cause a non-staining or inconsistent staining event. Additionally, the system may not detect this anomaly and therefore it may not generate an error message for the slide position. | DESIGN: Device Design | Ventana sent an Urgent Medical Device Correction Notification letter dated June 16, 2008 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to carefully read and perform all of the actions outlined in the notification letter concerning potential safety issue. Customers with concerns regarding the staining results on any patient specimen are instructed to follow their internal Quality Procedures regarding the review of patient reports. For question contact your local Customer Service Center. |  | 1217 | 1217 units total | Worldwide Distribution - USA Nationwide and the countries of Canada, Europe, Asia, Japan. | Terminated | August 29, 2013 | 1 | 1217 | Not_Computer
Z-2113-2013 | 65764 | tubing/tubing with filter, insufflation, laparoscopic | NKC | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | DeRoyal Insufflation Tubing sets:1) REF 28-0206 Insufflation Tubing, Laparoscopic Rx Only2) REF 28-0214 Hi-Flow Insufflation Tubing Rx Only 3) REF 28-0216 Tubing, Insuf w/.3 Micr Filter Rx Only 4) REF 28-0217 Insufflation Tubing, Laparoscopic Rx Only 5) REF 28-0208 Insufflation Tubing w/Vent Rx Only 6) REF 28-0212H Heated Hi-Flow Insufflation Tubing Rx Only.Used during laparoscopic surgical procedures to aid in the insufflation of the abdomen. | Insufflation Tubing Sets | 2 | August 28, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal Industries is recalling their product Insufflation Tubing Sets due to materials incompatibility which could possibly result in occlusion at or near the filter housing which could prevent the flow of gas through the device. | PRODUCTION CONTROLS: Packaging Process Control | The firm notified customers via letters dated 7/15/2013 sent UPS and via DeRoyal Representatives. Customers were asked to notify DeRoyal of any remaining inventory to return for replacement product or credit. Non-responders will receive additional notifications. Effectiveness checks will be performed by contacting non-responders by phone. | Contact the recalling firm for information | 524486 | 506,632 sets in US; 17,854 sets International | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Dubai, Arab Emirates, Quebec, Canada, Japan, and Ireland. | Open |  |  | 524486 | Not_Computer
Z-2130-2013 | 65902 | clamp, surgical, general & plastic surgery | GDJ | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Triathlon tibial alignment ankle clamp extramedullary (EM) referencing instrument (Cat. No. 6541-2-609), for use in Total Knee Arthroplasty (TKA) when the surgeon elects to use Stryker's Triathlon Total Knee System in conjunction with EM referencing. Non-Sterile Instrument. | Stryker Orthopaedics | 2 | August 29, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Cracks or fracture of the Triathlon Tibial Alignment Ankle Clamp, an instrument. | DESIGN: Device Design | Stryker Orthopaedics sent an Urgent Medical Device Correction letter dated July 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by t he customers. Customers are advised of the issue, potential hazards and risk mitigation factors. Customers are asked to complete the attached form and fax to 855-251-3635. Any questions should be directed to (201) 972-2100.All letters included a Product Correction Bulletin. |  | 10537 | 10,537 units | Nationwide Distribution | Open |  |  | 10537 | Not_Computer
Z-2136-2013 | 65823 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SC2316, Distal Femoral Augment with ScrewTotal knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis and/or post-traumatic degenerative problems and revision of failed previous reconstructions. | SC2316, Distal Femoral Augment with Screw | 2 | September 03, 2013 | 2013 | Stelkast Co | The firm became aware of an incident relating to a breach of sterility in the sterility barrier packaging of SC2316, Distal Femoral Augment with Screw. | PRODUCTION CONTROLS: Process Control | Stelkast called and emailed all customers on June 24, 2013, to notify them of the recall. Customers were asked to recover all affected products from their inventory and return them to Stelkast. Customers were instructed to contact Stelkast Customer Service for a Return Authorization (RA) Number prior to shipment to Stelkast. Customers with questions were instructed to call 1-888-273-1583. For questions regarding this recall call 724-941-6368. | Contact the recalling firm for information | 11 | 11 | Nationwide Distribution including TX, VA, OK, and PA | Open |  |  | 11 | Not_Computer
Z-2135-2013 | 65868 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry A1c_3 Calibrators, Catalog Number 10491408, IVD --- COMMON/USUAL NAME: ADVIA Chemistry A1c_3 CalibratorsShipping or unit package: Level 1: 2 x 0.5 g; Level 2: 2 x 0.5 g; Level 3: 2 x 0.5 g; Level 4: 2 x 0.5 g -- Siemens Healthcare Diagnostics, Inc. Tarrytown, NY 10591 --- System Information: The ADVIA Chemistry systems are automated, clinical chemistry analyzers that can run tests on human serum, plasma, or urine in random access and batch modes for in vitro diagnostic use.For in vitro diagnostic use in the calibration of A1c_3 (A1c_3 and A1c_3M) and Total Hemoglobin_3 (tHb_3 and tHb_3M) methods on ADVIAÂŽ Chemistry systems. | ADVIA Chemistry A1c_3 Calibrators | 2 | September 03, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens confirmed that the ADVIAÂż Chemistry Systems HbA1c method exhibits a positive bias of up to 12% for patient and CAP survey samples due to over-recovery when using A1c_3 Calibrator lots 1MD014, 1BD063, and 2GD014. The established ranges of CAP Survey Samples GH2-04, GH2-05 and GH2-06 and IFCC sample recoveries use Âż 7% from the established mean as their acceptance criteria. This is the crite |  | Siemens sent an Urgent Medical Device Correction letter dated July 30, 2013, to all affected customers via FedEx (for US consignees). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and discard any of the affected reagent kit lots. Customers werealso requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days.Field service personnel were sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions.Customers were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative for questions.For questions regarding this recall call 914-524-2955. | Contact the recalling firm for information | 1546 | US: 1546 units; Foreign: 518 units | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Argentina, Austria, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Hungary, Indonesia, Italy, Kazakhstan, Lithuania, Norway, Poland, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, South Africa, South Korea, Thailand, Turkey, Uruguay, Vatican City, and Vietnam. | Open |  |  | 2064 | Not_Computer
Z-2139-2013 | 65485 | occluder, internal vessel, temporary | OBC |  |  | Cardiovascular |  | LeGooÂŽ Endovascular Occlusion Gel-0.25 mL LeGoo with 3 Cannula PouchProduct Number: CS-25 LeGoo is a water-soluble, biocompatible polymer with reverse thermo-sensitive properties allowing it to be injected as a soft viscoelastic gel that forms an elastic plug within the vessel. LeGoo is indicated for temporary endovascular occlusion of blood vessels up to 4 mm in diameter below the neck. | LeGoo Endovascular Occlusion Gel0.25 mL LeGoo with 3 Cannula Pouch | 2 | September 03, 2013 | 2013 | Genzyme Corporation, A Sanofi Company | Product contains particulates (fibers) | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Genzyme sent a Medical Device Recall Notification letter dated June 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were avised to immediately discontinue use of all lots of LeGoo Endovascular Occlusion Gel and return all LeGoo and return product with the shipping label and packing slip enclosed. syringes.Complete the reply card via fax (414) 459-8791. Customer with questions were instructed to call GENCO at 877-674-2087.For questions regarding this recall call 508-270-2285. | Contact the recalling firm for information | 30 | 30 units | Worldwide Distribution - USA (nationwide) and internationally to Austria, Belgium, Canada, Denmark, Egypt. Europe, Finland, France, Germany, Hong Kong, Ireland, Italy, Netherlands, Poland, Saudi Arabia, Slovena, Spain, Sweden, Switzerland, The Netherlands, Turkey, and UK. | Open |  |  | 5312 | Not_Computer
Z-2134-2013 | 64937 | cement, bone, vertebroplasty | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | CONFIDENCE SPINAL CEMENT SYSTEM KIT, Product Code: 2839-13-000 The CONFIDENCE Spinal Cement System is intended for percutaneous delivery of CONFIDENCE Spinal Cement, which is indicated for fixation of pathological fractures of the vertebral body during vertebroplasty or kyphoplasty procedures. | CONFIDENCE SPINAL CEMENT SYSTEM KIT | 2 | September 03, 2013 | 2013 | DePuy Spine, Inc. | During injection of cement, the water in the hydraulic pump leaks past the piston within the pump body resulting in the loss of pressure and inability to continue to inject cement. | PRODUCTION CONTROLS: Process Control | DePut Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated April 15, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the con signee. Consignees were instructed to review their inventory and quarantine for return any products listed in the recall notice., complete the enclosed Business Reply Form, indicating level of inventory of the affected product and return the form to Depuy via fax to 888-943-4897or email to DePuy7862@stericycle.com and to return the product with the enclosed UPS label to Stericycle.For questions regarding this recall call 508-977-6606. | Contact the recalling firm for information | 3896 | 3896 units (Overall quantity distributed appears greater than quantities manufactured due to the distribution pathways) | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV, WY; Worldwide: Algeria, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, France, Germany, Hong Kong, India, Israel, Italy, Latvia, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Poland, Portugal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UAE, and UK | Open |  |  | 3896 | Not_Computer
Z-2137-2013 | 65919 | catheter, angioplasty, peripheral, transluminal | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular |  | Boston Scientific, Small Peripheral Cutting BalloonTM, Sterilized using irradiation: The Peripheral Cutting Balloon Catheters are indicated for Percutaneous Transluminal Angioplasty (PTA) of obstructive lesions of synthetic or native arteriovenous dialysis fistulae.1) 4.00mm x 15mm Small Peripheral Cutting BalloonTM MonorailTM Microsurgical Dilatation Device, UPN #001BPM4015140F0, Catalogue # BPM4015140F & UPN # M001PCBM4015140F0, Catalogue # PCBM4015140F.2) 4.00mm x 15mm Small Peripheral Cutting BalloonTM Over-the-Wire Microsurgical Dilatation Device, UPN # M001PCBO4015140F0, Catalogue # PCBO4015140F. | Boston Scientific, Small Peripheral Cutting BalloonTM | 2 | September 03, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of the 4.0mm x 15mm size of Small Peripheral Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilatation Devices and Flextome Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilation Devices. Boston Scientific has determined that users may experience significant difficulty or inability to remove the protector cap from the devic | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated July 30, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised that further distribution or use of the product should cease immediately, and to segregate the product and return it to Boston Scientific. Consignees were asked to complete and return the "Recall Removal Reply Verification Tracking Form". For questions consignees were instructed to contact their local Sales Representative.For questions regarding this recall call 763-494-2042. | Contact the recalling firm for information | 1331 | 1,331 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV,NH,NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC,SC, TN, TX, UT, VA, WA, WV, and WI. Internationally to Algeria, Finland, Japan, Spain, Australia, France, Latvia, Sweden, Austria, Germany, Netherlands, Switzerland, Belgium, Great Britain, Norway, Taiwan, Brazil, Greece, Panama, Turkey, Canada, Hong Kong, Poland, Unit.Arab Emir., China, India, Russian Fed., Colombia, Ireland, Saudi Arabia, Czech Republic, Israel, South Africa, Denmark, Italy, and South Korea. | Open |  |  | 1671 | Not_Computer
Z-2158-2013 | 66029 | pin, fixation, smooth | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Steinmann Pin, packed in a clear tube with end caps, tubes packed in a sterile barrier pouch, 6 pouches/sales unit. Pouches are labeled in part ***MicroAire Surgical Instruments 3590 Grand Forks Blvd Charlottesville, VA LLC 22911***Models:1620-109 1636-5091620-509 1640-5091620-609 1648-5091620-709 1648-509T1624-109 1648-6091624-509 1648-609T1624-509T1628-5091632-5091632-609A smooth or threaded wire / pin that may be used for fixation of bone fractures, for bone reconstructions, as a guide pin for insertion of other implants, or it may be implanted through the skin so that a pulling force (traction) may be applied to the skeletal system. | Steinmann Pin | 2 | September 04, 2013 | 2013 | MicroAire Surgical Instruments, LLC | A medical device used in surgical procedures was packaged in a pouch in which a defective seal may compromise sterility | PRODUCTION CONTROLS: Packaging Process Control | MicroAire Surgical LLC sent an Urgent Medical Device Recall letter dated August 12, 2013, to all affected customers. The letter advised consignees recover and quarantine the affected product so that it does not make it into the surgical process. Customers were instructed to either return affected product to MicroAire for inspection or they could inspect the product for the defect by looking at the seal area on the inside sterile package (short non-chevron side). The recalling firm further advised distributors to contact all of their customers to recover and quarantine the product so that it does not make it into the surgical process. Customers with questions were instructed to call 434-975-8370.For questions regarding this recall call 434-975-8370. |  | 15786 | 15786 total units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Ireland, Canada, Netherlands, Sweden, Japan, and United Kingdom. | Terminated | February 12, 2014 | 161 | 15786 | Not_Computer
Z-2162-2013 | 65905 | thermometer, clinical color change | FQZ | Clinical color change thermometer. | General Hospital | General Hospital | 510(K) Exempt | Tempa DOT Single-Use Clinical Thermometer Medical IndicatorsMade in U.S.A. by Medical Indicators, Inc. 16 Thomas J Rhodes Industrial Drive, Hamilton, NJ 08619 USA Phone 1.888.737.1601 www.medicalindicators.comPer 3M website (this product line was recently bought by Medical Indicators from 3M) and labeling in 806, TempaDOT thermometers are individually wrapped, single-use, disposable clinical thermometers for oral or axillary use. | Tempa DOT | 2 | September 06, 2013 | 2013 | Medical Indicators, Inc. | TempaDot products may have been mislabeled. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Medical Indicators Inc.sent an "URGENT" Recall letter dated July 10, 2013 and Acknowledgement Form to all distributors/customers via email and Priority Mail with delivery confirmation. The letter described the product, problem and actions to be taken. The customers were instructed to check full & partial cases of the tempadot; open and inspect all 5122NS cases; pull out and quarantine any of the lots found in the 5122NS cases, pending return to medical indicators; if individual Tempadot 100-count boxes of 5122 and 5122NS are stored in different warehouse follow the instructions provided in the letter; if Tempadot 100-count boxes of 5122 were sold to a customer with the affected lots, contact the customer and have them return the product; and complete and return the Recall Return Response Acknowledgement and Receipt Form via mail to: 16 Thomas J Rhodes Industrial Drive, Hamilton NJ, 08619; fax: (609) 587-5636 or email to: www.medicalindicators.com or General Manager at: pbaker@medicalindicators.com.Contact information for any question(s) related to this recall may be addressed to General Manager of Medical Indicators, Inc. at 609-737-1600, extension 204 or email to: pbaker@medicalindicators.com, Mondays through Fridays from 8:00am to 5:00pm EST. |  | 111800 | 111,800 boxes (5,590 full boxes) | Nationwide distribution: USA including states of: CA, GA, FL, IN, NY, NV OH, PA, TN,TX, VA, and WI. | Terminated | January 16, 2014 | 132 | 111800 | Not_Computer
Z-2161-2013 | 65970 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Tec 6 continuous Flow Vaporizer and Tec 6 Plus Anesthesia Vaporizer. The Datex-Ohmeda Tec 6 Plus and Tec 6 Plus NAD Variant is an electronic vaporizer which delivers the anesthetic agent desflurane. The Tec 6 Plus attaches only to Datex-Ohmeda anesthesia machines equipped with the proprietary Selectatec manifold. The Tec 6 Plus NAD Variant attaches only to NAD anesthesia machines equipped with the NAD exclusion type manifold. | Tec 6 Plus Vaporizer | 2 | September 06, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the pre-operative check, low pressure leak test at 1% dial setting of the Tec 6 and Tec 6 Plus Vaporizers. The user manual pre-operative check, low pressure leak test may not detect the full range of leaks from seal wear degradation in the vaporizers. If the leak is not detected, it will result in fresh gas leaking to | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | GE Healthcare sent an Urgent Medical Device Correction letter dated August 1, 2013, to all affected customers. The letter described the issue, asked that the Tec 6 and Tec 6 Plus Vaporizers user manual be updated with the attached addendum. Also Users are to perform the user manual recommended pre-operative check which includes the low pressure leak test on every vaporizer with the dial turned to the 12% setting instead of the specified 1% setting as per the attached addendum. If the device does not pass the new low pressure leak at 12%, it should be removed from service and consignee should ontact a GE Healthcare Representative for a repair/replacement. If the vaporizer passes the new low pressure leak test at the 12 % setting, it can continue to be used, however, as detailed in the user manual, continue to routinely perform the pre-operative check which includes the lowpressure leak test.A Medical Device Correction Confirmation form was asked to be sent back to GE.Customers with questions or concerns were instructed to call 1-855-700-4350 or send an email to vaporizer.admin@ge.com. For countries outside the US, customers were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. |  | 7496 | 31,443 devices (23,947 US, 7496 OUS) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COSTA RICA, CROATIA local name-Hrvatska, CYPRUS, DENMARK, ECUADOR, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA-REPUBLIC OF, KUWAIT, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA, | Open |  |  | 31443 | Not_Computer
Z-2166-2013 | 65993 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook Safety Infusion System Burette Set.For intravenous fluid administration. | B. Braun Outlook Safety Infusion System Burette Set | 2 | September 06, 2013 | 2013 | B. Braun Medical, Inc. | Reports were received of an incorrectly assembled Outlook Safety Infusion System Burette Set (Product Code 375038), and the investigation determined that one of the ULTRASITE injection sites and corresponding low pressure backcheck valve was assembled upside down. | PRODUCTION CONTROLS: Process Control | On August 4, 2013, a letter was sent to direct accounts asking them to check their inventory for the affected product. These customers were also instructed to notify their customers if the product was further distributed and to return all affected product and/or the product removal acknowledgement form. | Contact the recalling firm for information | 900 | 900 | Distributed in the states of FL, NY, NJ, and TX. | Open |  |  | 900 | Not_Computer
Z-2165-2013 | 66082 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa Vial Adapter, 14mm - Luer Lock, Order No. REF: 90, Quantity: 50 each, Single Use. Product Code H93890.Used as a needleless access device for reconstitution and drug transfer. | Baxa Vial Adapter, 14mm Luer Lock | 2 | September 06, 2013 | 2013 | Baxa Corporation | Baxter Healthcare Corporation is voluntarily recalling one lot of Baxa Vial Adapters, 14mm - Luer Lock, due to incorrect expiry date. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | An Urgent Product Recall notification was mailed to one customer via USPS First Class mail on 8/22/13. The one customer affected by this action was instructed to locate and remove the affected product from inventory and to return to Baxter for credit and/or replacement. Any customer returns received will be discarded per the Corporate Hold. |  | 34 | 34 cases | Distributed to One Customer in NC. | Open |  |  | 34 | Not_Computer
Z-2167-2013 | 64965 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Xlumena brand NAVIX Access Device, Model/Catalogue Number: NVX-10-03, Part Number: FG-00666, Product is manufactured and distributed by Xlumena, Inc., Mountain View, CAThe NAVIX Access Device is a multi-lumen balloon catheter with a trocar, anchor balloon, and dilation balloon that creates and dilates a tract across tissue planes. The NAVIX Access Device is indicated for use as an accessory to cannulate the transgastric or transduodenal wall and into a pancreatic pseudocyst, when it is visibly bulging into the gastrointestinal tract during endoscopic procedures. | Xlumena brand NAVIX Access Device | 2 | September 09, 2013 | 2013 | Xlumena, Inc. | A complaint investigation has found that product fractures can occur at the distal end of the catheter under load. | DESIGN: Process Design | Xlumena sent an Urgent Medical Device Recall letter dated April 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected product. Customers with questions were instructed to contact dhasker@xlumena.com. For question regarding this recall call 650-961-9901. | Contact the recalling firm for information | 148 | 148 units | Worldwide Distribution - USA including AZ, CA, FL, GA, IN, NY,SC, WA, PA, IL, CO, and internationally to Netherlands, Germany, Spain, Denmark, and Switzerland. | Terminated | September 23, 2013 | 14 | 148 | Not_Computer
Z-2178-2013 | 65978 | saline, vascular access flush | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Monoject 0.9% Sodium Chloride Flush Syringe, 12 mL Syringe with 10 mL FillProduct ID: 8881570121Product Usage: Both 0.9% Sodium Chloride Flush syringes and 10U/mL or 100U/mL Heparin Lock Flush syringes are sterile, single-use products which are intended for use in flushing and locking compatible intravenous administration sets and indwelling intravenous access devices | Covidien | 2 | September 10, 2013 | 2013 | Covidien LLC | Monoject prefill flush syringes may contain non-sterile water and labeled as saline or heparin | PRODUCTION CONTROLS: Process Control | Covidien sent an Urgent Medical Device Recall letter dated August 16, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested that customers immediately quarantine and discontinue use of the affected products, return any affected product by following the instructions provided in the letter, and complete and return the Recalled Product Return Form to Covidien. The letter states that if the affected products have been distributed, assure that all customers are notified and affected product is removed from the distribution channel as well as from customer stock. Questions or concerns, please do not hesitate to contact your Covidien representative or Covidien Customer Service, Monday through Friday, 8am- 6:30pm ET, at (800) 962-9888, option 1, and then option 2.Firm issued Press on 8/19/13. |  | 3235860 | 3,235,860 units | Worldwide Distribution - USA Nationwide and the country ofBermuda | Open |  |  | 3565980 | Not_Computer
Z-2174-2013 | 65957 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Cassettes, Product Code: 10144The STERRAD Sterilization System is a low temperature general purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | STERRAD | 2 | September 10, 2013 | 2013 | Advanced Sterilization Products | The recall of the STERRAD 100NX Cassettes was initiated because Advanced Sterilization Products (ASP) does not have adequate data to support the effectiveness of the packaging's leak indicator for the entire duration of the STERRAD 100 NX System cassette's 15 month shelf life. |  | ASP sent an Urgent Medical Device Field Safety Notification letter dated July 25, 2013, to all affected customers. The recall letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to contact Stericycle directly at (866) 562-5937 to report any complaints or suspected problems with their STERRAD 100NX System cassette. For questions regarding this recall call 949-453-6400. |  | 1923 | 1,923 units | Worldwide Distribution - USA (nationwide) and internationally to RU, VE, KW, SA, HU, GB, IT, HK, IL, MX, ZA, PH, GR, AE, IN, CO, KR, AU, IN, ES, SG, MY, TH, TR, TW, SI, BE, PL, EC, CN, JP, PA, DE, VN, CL, and Switzerland, United Kingdom. Israel. Slovenia, Qatar, Panama, Venezuela. | Terminated | November 27, 2013 | 78 | 1923 | Not_Computer
Z-2172-2013 | 65806 | acid, folic, radioimmunoassay | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Folate (2500 tests), REF SMN 10340209, Cat No. 09132781In vitro diagnostic use in the quantitative determination of folate in serum or red blood cells. | ADVIA Centaur Folate | 2 | September 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibration failures and/or significant negative shifts in quality control (QC) and patient results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Siemens sent an Urgent Medical Device Correction letter on July 26, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. an Urgent Field Safety Notice was sent to all domestic customers via Federal Express. On July 25, 2013 this notice was sent to all OUS customers via hard copy, email, and/or fax as determined by each country's local regulations and procedures. These notices inform the customer that a portion of Folate ReadyPacks in kit lots ending in 218 are exhibiting calibration failures and/or significant negative shifts in quality control (QC) and patient results. If customers have an alternate lot of the ADVIA Centaur or ADVIA Centaur CP Folate assay, they are instructed that this lot should be used for future testing. However if no alternative lot is available, quality control (QC) results should be obtained for each individual ReadyPack as a QC or calibration is an indication of an affected pack. An affected pack must be replaced with a different Folate ReadyPack where a valid calibration and acceptable QC values are obtained. These notices should be reviewed by the facility's laboratory or medical director, though no look back is required.Customers were instructed that Fedex Return Receipts will be used to confirm that the customers received the UMDC. Customers with questions were instructed to contact their local Technical Support Provider or Distributor.For questions regarding this recall call 508-668-5000. |  | 44 | 44 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA,MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada and Mexico. | Open |  |  | 6486 | Not_Computer
Z-2173-2013 | 66036 | kit, dna detection, human papillomavirus | MAQ |  |  | Microbiology | PMA | KIT cobas 4800 HPV AMP/DET 240T / 960T US-IVD, CE-IVDRoche Molecular System, Inc. 1080 US highway 202 South Branchburg, NJ 08876 USARoche Diagnostics 9115 Hague Road Indianapolis, IN 46250-0457 USAThe cobasÂŽ HPV Test is a qualitative in vitro test for the detection of Human Papillomavirus in patient specimens. The test utilizes amplification of target DNA by the Polymerase Chain Reaction (PCR) and nucleic acid hybridization for the detection of 14 high-risk (HR) HPV types in a single analysis. The test specifically identifies types HPV16 and HPV18 while concurrently detecting the rest of the high risk types (31, 33, 35, 39, 45, 51, 52, 56, 58, 59, 66 and 68). The cobasÂŽ HPV Test is indicated: (a) To screen patients 21 years and older with ASC-US (Atypical squamous cells of undetermined significance) cervical cytology test results to determine the need for referral to colposcopy. (b) To be used in patients 21 years and older with ASC-US cervical cytology results, to assess the presence or absence of high-risk HPV genotypes 16 and 18. This information, together with the physicians assessment of cytology history, other risk factors, and professional guidelines, may be used to guide patient management. The results of this test are not intended to prevent women from proceeding to colposcopy. (c) In women 30 years and older, the cobasÂŽ HPV Test can be used with cervical cytology to adjunctively screen to assess the presence or absence of high risk HPV types. This information, together with the physicians assessment of cytology history, other risk factors, and professional guidelines, may be used to guide patient management. (d) In women 30 years and older, the cobasÂŽ HPV Test can be used to assess the presence or absence of HPV genotypes 16 and 18. This information, together with the physicians assessment of cytology history, other risk factors, and professional guidelines, may be used to guide patient management. Cervical specimens that may be tested with the cobasÂŽ HPV Test include the following liquid based collection media and collection device: " ThinPrepÂŽ Pap TestTM PreservCytÂŽ Solution " Endocervical Brush/Spatula | KIT cobas 4800 HPV AMP/DET 240T / 960T USIVD, CEIVD | 2 | September 10, 2013 | 2013 | Roche Molecular Systems, Inc. | During an internal investigation, a formulation error that occurred was identified during the manufacture of the cobasÂż 4800 HPV Master Mix reagent that was packaged in four (4) cobasÂż 4800 AMP/DET kits, US- and CE-IVD (240 and 960 test kit configurations). Examination of the batch production record of the material revealed there was an error in the tare weight used for the carboy during the fina | DESIGN: Device Design | Roche Molecular Diagnostics sent an Urgent Medical Device Correction letter dated July 25, 2013, to all affected customers via UPS return receipt. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. Immediately discontinue use of the product lots listed above.2. Discard affected product according to local guidelines.3. To receive replacement inventory, complete the entire attached faxback form and fax it to 1-866-918-8741.4. The faxback form must be completed and returned even if you do not have affected inventory.5. File this Urgent Medical Device Correction (UMDC) for future reference.Please contact Roche Molecular Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-526-1247, if you have questions about the information contained in this UMDC. | Contact the recalling firm for information | 350 | US-IVD: 350 Kits; CE-IVD: 441 KITS | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Brazil, Czech Rep, Canada, Croatia, Columbia, Denmark, Dominican Republic, Germany, Ecuador, France, Japan, Greenland, Mexico, Italy, New Zealand, Netherlands, Peru, Norway, Singapore, Romania, South Africa, Spain, South Korea, Sweden, Switzerland and UK. | Open |  |  | 791 | Not_Computer
Z-2187-2013 | 65968 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | StaXxÂŽ IB Convex Expandable Implant (25 x 12 x 9 mm), Intervertebral Fusion Device, Lumbar; Catalog Number 11 -1902.Indicated for intervertebral body fusion procedures in skeletally mature patients. | StaXx IB Convex Expandable Implant | 2 | September 11, 2013 | 2013 | Spine Wave, Inc. | A component Superior endplate manufactured from a different polymer than specified | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Spine Wave notified surgeons verbally on 5/14/13 and followed by written notification on June 6, 2013 of the incorrect polymer used for the endplate and the health risks. |  | 17 | 17 units | Distributed in the states of FL and OR. | Terminated | February 10, 2014 | 152 | 17 | Not_Computer
Z-2122-2013 | 66058 | system, test, automated, antimicrobial susceptibility, short incubation | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Microscan Synergies Plus Negative Urine Combo 2Product Usage:MicroScanÂŽ Synergies plusÂŽ Panels &#38; MicroScanÂŽ Rapids plus (Negative) - For use in determining quantitative and/or qualitative antimicrobial agent susceptibility and/or identification for gram-negative bacteria. | MicroScan Synergies plus Panels &38; MicroScan Rapids plus (Negative) | 1 | September 11, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Confirmed increase in false susceptible and false intermediate misreads for Imipenem and/or meropenum on Synergies plus Negative and rap ID/S plus Negative panels read on Walk Away System. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Siemens sent an Urgent Field Safety Notice letters dated August 2013 to all US customers via over night delivery, receipt requested. Letters for foreign customers are being translated and will be sent as soon as possible. The letter identified the affected product, reason for field correction, risk to health problem and actions to be taken by customer. Customers were instructed to suppress all reporting of sensitive and intermediate results for imipenem and meropenem. Siemens also recommends that the customer consider the need to review previous test results, conduct patient follow-up, and/or repeat testing. Customer were requested to complete the attached Field Correction Effectiveness Check form and fax it to 302-631-8467. For questions contact your Siemens Technical Support Provider or Distributor. |  | 27804 | 1,324 boxes, 26,480 panels | Worldwide Distribution - USA Nationwide in the states of: AK, SC, IA, MO GA, WI, TX, KS, CA, IL, CO, AZ, SD, IN, MD, FL, MN and the countries of:Japan, France, Algeria, Germany, Austria, Romania, Russian Federation, South Africa, Kazakhstan, Belgium, Great Britain, Turkey, Spain, Italy, Saudi Arabia, Qutar, Poland, Iraq, Unite Arab Emrites, Algeria, Oman, Greece, Portugal, and Ireland. | Open |  |  | 80292 | Not_Computer
Z-2192-2013 | 66031 | lancet, blood | FMK | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | BD Microtainer Quikheel Lancet 1.0 mm incision Depth 2.5 mm Incision Length REF 368101 Rx Only Becton, Dickinson and Company, Franklin Lakes, NJ 07417 USA Made in USA.Used to perform heel sticks for the purpose of collecting skin puncture blood. | BD Microtainer Quikheel Lancet | 2 | September 11, 2013 | 2013 | Becton Dickinson & Company | The Lancet will not retract after activation thus creating a potential for exposure to a used sharp. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | BD notified their customers and distributors by phone, e-mail,and/or letter/notice of return/ customer packing list on 7/10/2013. | Contact the recalling firm for information | 58200 | US - 39,800 units ; Canada - 18,400 units | Distributed in the states of CT, KY, IL, NY, PA, TX, and VA and the country of Canada. | Open |  |  | 58200 | Not_Computer
Z-2193-2013 | 66123 | needle, catheter | GCB | Introduction/drainage catheter and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | *** 1) Cat. 4341B; Qty. 10; Thoracentesis Tray; With Catheter; Sterile; CareFusion; 1500 Waukegan Road; McGaw Park, IL 60085 USA; *** 2) Cat. 4341BSDF; Qty. 10; Safe-T Thoracentesis Tray; With Catheter and Sharps Injury Protection; Drug Free; CareFusion; 1500 Waukegan Road; McGaw Park, IL 60085 USA.Indicated for specimen retrieval and therapeutic drainage of the pleural cavity. | 1) Thoracentesis Tray; 2) SafeT Thoracentesis Tray | 2 | September 11, 2013 | 2013 | Carefusion 2200 Inc | Catheters included in Thoracentesis Tray and Safe-T Thoracentesis Tray - Drug Free may occlude, preventing drainage of excessive fluid from the pleural cavity leading to a potential delay in the medical procedure and potentially serious patient complications. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Recall notification letters dated 8/27/13 were distributed via courier to direct accounts. Consignees were/are requested to confirm receipt of the notification, examine their inventory for the recalled products and to destroy the products. After the products are destroyed, the consignees are requested to submit a completed response form to CareFusion for replacement with functional devices. Consignees who cannot destroy the affected products are requested to contact CareFusion for additional information. |  | 197913 | *** 1) 174,580 units; *** 2) 23,330 units | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MD, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV and Puerto Rico, and the country of Iceland. | Open |  |  | 197913 | Not_Computer
Z-2189-2013 | 65984 | transducer, blood-pressure, extravascular | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | "***Flo Trac Sensor Model MHD85***"Sterile, single use kit that monitors pressures when attached to pressure monitoring catheters. | FloTrac Sensor | 2 | September 11, 2013 | 2013 | Edwards Lifesciences, LLC | The Edwards Lifesciences FloTrac Sensor is being recalled due to a manufacturing non-conformance related to incorrect packaging configuration where the IFU was placed on top of the Tyvek instead of being placed at the bottom of the box. There is no evidence available at this time that could validate the sterility of this incorrect packaging configuration. | PRODUCTION CONTROLS: Packaging Process Control | Edwards Lifesciences initiated a US recall via an Urgent Product Recall letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that an Edweards Lifesciences representative would come to their facility to retrieve the affected device. A device replacement would be provided to the customer. Customers were asked to complete the enclosed Confirmation Form and give it to the Edwards Lifesciences representative that will come to their facility. For customers with questions were instruct4ed to call Edwards Customers Service at 1-800-424-3278.For questions regarding this recall call 972-410-7100. |  | 35 | 35 units | Nationwide Distribution including MO and MI. | Open |  |  | 35 | Not_Computer
Z-2190-2013 | 65854 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HemoCueÂŽ Glucose 201 Microcuvettes, 25 Microcuvettes Individually Packaged. The HemoCue Glucose 201 Microcuvettes are designed for use with the HemoCue Glucose 201 Analyzer and the HemoCue Glucose 201 DM Analyzer. Quantitative determination of glucose in whole blood using a specially designed analyzer, the HemoCue Glucose 201 with Plasma conversion. The quantitative determination of the instant blood glucose concentration in circulation supplements the clinical evidence in the diagnosis and treatment of patients with diabetes as well as monitoring of neonatal blood glucose levels. | HemoCue Glucose 201 Microcuvettes | 2 | September 11, 2013 | 2013 | HemoCue AB | During a complaint investigation HemoCue AB discovered punctured HemoCue Glucose 201 single pack pouches. | PRODUCTION CONTROLS: Packaging | HemoCue sent an Urgent Field Safety Notice dated July 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine any affected product and Return the Field Safety Notice Verification Form to e-mail: 2013glufsn@hemocue.com or fax 562-668-5794 Customers would be contacted to coordinate return and replacement of the affected microcuvettes. Customers with questions were instructed to call 800-426-7256.For questions regarding this recall call 440-925-3485. |  | 100 | 22,215 boxes of 100 cuvettes distributed in US, 34, 841 distributed outside the US | Worldwide Distribution - USA (nationwide) and internationally to AUSTRALIA, BELGIUM, CANADA, COTE D'IVOIRE, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, IRELAND, ITALY, KENYA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIAN FEDERATION, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, VIET NAM & YEMEN.. | Open |  |  | 61056 | Not_Computer
Z-1993-2013 | 66046 | system, laser, transmyocardial revascularization | MNO |  |  | Cardiovascular | PMA | Cardiogenesis Corporation - Cardiogenesis Transmyocardial Laser.The device is indicated for use in the treatment of patients with angina (Canadian Cardiovascular Society Class 4) refractory to medical treatment and secondary to objectively demonstrated coronary artery atherosclerosis and with a region of the myocardium with reversible ischemia not amenable to direct coronary revascularization. | Cardiogenesis Transmyocardial the SoloGrip III (HP5G3) and the PEARL 5.0 (HPPRL5). | 2 | September 11, 2013 | 2013 | Cardiogenesis Corporation | Following acquisition of Cardiogenesis, we began receiving complaints of sparks, burns, or smoke incidents related to the handpieces. | DESIGN: Device Design | Cardiogenesis is responsible for the correction of all noncompliant products. CDRH understands that the corrective action plan (CAP), involving 2,437 Sologrip III handpieces and 61 PEARL 5.0 handpieces, consists of the following:1. Notification to purchasers of the handpiece removal and replacement procedures,2. Replace existing hospital inventory of handpieces with a redesigned handpieceCDRH approves the CAP subject to the following conditions:1. Notification of all distributors and purchasers will be successfully implemented, and2. Corrections will be made at no cost to the purchasers by October 1, 2013.For further questions please call (949) 420-1800. |  | 2498 | 2498 | US Distribution | Open |  |  | 2498 | Not_Computer
Z-2195-2013 | 66110 | retainer, screw expansion, orthodontic | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Mini RPE Screw Assembly , a component ofHerbst, MARA and Fixed Metal devices.Product Usage:AOA Mini RPE Screw Assembly is a component of Herbst, MARA and Fixed Metal devices; removable functional appliances that are to modify the growth of the jaws in a prescribed growth pattern, however, this category is generally attached to the patients upper and lower molars or bicuspids by means of temporary stainless steel crowns or orthodontic bands. Retainer, Screw Expansion, Orthodontic. | AOA Mini RPE Screw Assembly | 2 | September 12, 2013 | 2013 | Sybron Dental Specialties | The firm initiated the recall of AOA Mini RPE Screw Assembly, because it may be missing a complete weld joining one side of the arm with the screw body. Herbst, MARA and Fixed Metal devices may be affected, because they contain an AOA Mini RPE Screw Assembly. | DESIGN: Device Design | AOA sent a recall notification letter titled "Urgent-Medical Device Correction", on June 6, 2013 via USPS 1st class mail. The letter identified the affected product, reason for recall, images, and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form to AOA. For questions contact the AOA Customer Support Department at 1-800-262-2551. |  | 1091 | 1,091 | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, and EU. | Terminated | September 13, 2013 | 1 | 1091 | Not_Computer
Z-2194-2013 | 66173 | implant, endosseous, orthodontic | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Brand Name: Vector TAS 8mm Screw.The Device Listing Number is E670535.Vector T AS 8mm Screw is intended to provide a fixed anchorage point for attachment of orthodontic appliances to facilitate the orthodontic movement of teeth. It is used temporarily and removed after orthodontic treatment has been completed. It is surgically placed in the bone of the upper or lower jaw arches to provide support for orthodontic devices. | Brand Name: Vector TAS 8mm Screw. | 2 | September 12, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one lot of 8 MM Vector TAS Screws because some screws of this lot were machined undersized at the top of the threads may potentially cause a weakening of the screw. Use of the undersized product may result in breakage of the screw during insertion or removal when torqued, possibly requiring surgical removal of any remaining pieces. | DESIGN: Device Design | Ormco sent an "Urgent Medical Device Recall letter dated July 12, 2013, via USPS 1st class mail, along with a return form to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed of the recall, with product description, product codes, ship dates, contact information and instructions to return any affected product identified in stock, Additionally, doctors are instructed to proceedwith treatment if the screw has already been placed and to take special care not to over-torque the screw upon removal. Periodic effectiveness checks will be conducted in order to follow up with non-responders.Please fax this acknowledgement form to the the following number to confirm receipt of this notification whether or not you have any affected products. (714) 288-4603. | Contact the recalling firm for information | 4786 | 4,786 | Worldwide Distribution - USA (nationwide) including the countries of Canada, Australia, New Zealand, EU and ROW consignees. | Open |  |  | 4786 | Not_Computer
Z-2196-2013 | 65980 | urinalysis controls (assayed and unassayed) | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | DiaScreen Liquid Urine Controls: Positive and Negative Assayed Liquid Quality Controls for use with Dipsticks in Routine Urinalysis; Product Number D76004 (4x12 mL bottles) consists of: 2x12mL Positive controls [and] 2x12mL Negative controls; Manufactured for ARKRAY USA, Inc.; Minneapolis, MN 55439Diascreen liquid urine control is a synthetic control designed to react with commercial urine dipsticks to register listed responses on color reagent pads. It is intended for use in a clinical laboratory as a urine control for qualitative procedures used in chemical determinations. The control is intended to be used exactly as if it were a patient sample to assist in the assessment of the listed analytical procedures. | DiaScreen Liquid Urine Controls | 2 | September 13, 2013 | 2013 | Assuramed | The products may have been stored at temperatures outside their required storage conditions. | PRODUCTION CONTROLS: Storage | Independence Medical (im) sent an Urgent Product Recall letter dated July 16, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to identify and immediately cease the distribution of the recalled products, quarantine the recalled products, and contact the recalling firm at 800-860-8027 for instructions on how to return the recalled products. Consignees who further distributed the recalled products were requested to notify their customers of the recall. Consignees were also instructed to complete the Product Recall Acknowledgement Form and return the document by email to recall@indemed.com indicating receipt of the notification.Customers with questions were instructed to contact the product manufacturer for ProTime Disposable Cuvettes at 800-579-2255, DiaScreen Liquid Urine Controls at 800-818-8877, and Hemoccult II SENSA elite Dispensapak at 800-854-3633.For questions regarding this recall call 800-860-8027. | Contact the recalling firm for information | 10 | 10 units | Nationwide Distribution including AR, AZ, CA, CO, FL, IL, KY, LA, MI, NJ, NY, OH,OK, OR, SC, TX, UT, and WA. | Open |  |  | 154 | Not_Computer
Z-2199-2013 | 66057 | instrument, cutting, orthopedic | HTZ | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | AEQUALIS Reversed II or AEQUALIS Reversed Fracture Impactor. These products are contained within the Aequalis reversed II Instrument set tray no.YKAD83 and the Aequalis reversed fracture instrument set tray no.YKAD95. The items affected are MWD21, MWD023, MWD024 and MWD025.The impactor is an instrument used during shoulder arthroplasty; it is not implanted. The impactor is used to set the glenoid sphere, poly insert, spacer, and hemi adapter into place during surgery. The impactor is a two piece instrument which consists of a tip and a handle. The tip is made of plastic and metal that is pressed together for assembly. | AEQUALIS Reversed II or AEQUALIS Reversed Fracture Impactor. | 2 | September 13, 2013 | 2013 | Tornier, Inc | Following the identification of the disassociation of an impactor tip during surgery, Tornier is initiating a voluntary recall of all lots of Aequalis Reversed II and Reversed Fracture impactors. Upon review of the event, Tornier determined that a redesign of the instrument was warranted. These products are contained within the Aequalis Revered II instrument set tray no YKAD83 and the Aequalis R | DESIGN: Device Design | Tornier sent an "Urgent Medical Device Recall" letter dated August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were advised to cease further use and distribution of the product and to contact all end users. Also requested the completion and return of the attached questionnaire via fax to 952-236-4007, e-mail to CustomerServices@Tonier.com or using the envelope provided.We apologize for the incovenience and thank you in advance for your cooperation in this matter. If you have any questions please contact our Customer Service Department at 1-888 494-7950. |  | 2298 | 2,298 (1045 US, 1253 OUS) | .Worldwide Distribution - USA including Puerto Rico and the states of AL, AZ, CA, CO,CT, FL, GA, IL, KY, MD, MA, MI, MN, NE, NJ, NC, OH, OR, PA and TX., and the countries of CANADA, MEXICO, HONG KONG, SOUTH AFRICA, ARGENTINA, COLOMBIA, ISRAEL, AUSTRALIA, BELGIUM, FINLAND,FRANCE, GERMANY, ITALY, NETHERLANDS, SLOVENIA, SPAIN, SWITZERLAND and UNITED KINGDOM | Open |  |  | 1253 | Not_Computer
Z-2200-2013 | 66010 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T8 REMOVAL HEAD, REF 8910641, QTY: 1 EA, Product Usage:The non-sterile T8 Removal Diriver is used with the Spire and Spire Z Spinal System in conjunction with the set screwdriver to remove the implant during removal surgery. The set screw driver is attached to the end of the removal head and used to disengage the remaining portion of the set screw within the device. | T8 Removal Driver for the CD Horizon Spire and Spire Z Spinal System | 2 | September 14, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | One lot of the T8 Removal Head had incorrect hardness specification, which could result in the T8 Removal head failing during removal of the CD Horizon Spire and Spire Z Spinal System implant. There is a potential for tissue damage to occur and/or minor surgical delay to remove the remaining fragments of the tip. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Medtronic sent an Urgent Voluntary Market Removal Recall letter dated August 9, 2013 with Recall Questionnaires and Confirmation forms included to Risk Managers and Sales Representatives. The letter identified the affected product, problem and action to be taken. Customers were instructed to remove the affected product, discontinue use and contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product. For questions call your Medtronic Sales Representative or Global Quality Department at 800-876-3133. |  | 9 | 9 units | USA Nationwide Distribution in the state of GA, AK, WA, and CA. | Terminated | December 27, 2013 | 104 | 9 | Not_Computer
Z-2205-2013 | 66143 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | C2244.4308 CONELOG Titanium Base CC ÂŽ4.3GH2 0.8 inkl HSProduct Usage:Endosseous dental implant abutment CONELOG Abutments are intended to be used to fabricate prosthetic restorations in conjunction with CONELOG implants to support prostheses in the maxillary and/or mandibular arch. | CONELOG Titanium Base | 2 | September 16, 2013 | 2013 | Camlog Usa | Product not cleared for sale/use in the United States. | PREMARKET APPROVAL: No Marketing Application | Camlog Usa notified four US consignees of the recall by telephone on August 5, 2013. The notification identified the recalled product, which did not have proper marketing clearance for distribution in the United States and asked customers to return all unused product. Cusotmers were sent an overnight UPS label via e-mail and asked customers to return the device to CAM:LOG USA Attn: Debby Callahan, 1822 Aston Avenue Carlsbad, CA 92007. |  | 6 | 6 to US customers | USA Nationwide Distribution in the states of 4 US customers in NH. MD, CO, and IL. | Terminated | January 29, 2014 | 135 | 12 | Not_Computer
Z-2209-2013 | 66155 | system, thermal regulating | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III, Model Number 233, Cincinnati Sub-Zero Products, Inc., 12011 Mosteller Road, Cincinnati, OH 45241Hyper-Hypothermia device used to lower or to raise a patients temperature and/or maintain a desired patient temperature through conductive heat transfer. | Blanketrol III, Model Number 233 | 2 | September 16, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a label review it was discovered that both the Blanketrol III Operations (57201-N) & Operations & Technical (56201-Q) manuals had verbiage added back into the manual during ECN M111-4260 (released 12/29/2011) which was previously removed for clarification of the intended use and subsequent 501(K) approval. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Cincinnati Sub-Zero (CSZ) sent an Urgent Medical Device Field Action letter dated September 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CSZ at 513-772-8810 or 1-800-989-7373 to request new versions of the manuals. At that time customers are instructed to provide the quantity and serial numbers of the affected devices that are at their facility. Upon receipt of the new versions of the manual, immediately collect and destroy/discard all the affectred devices Operation and Technical Manuals their facility is currentlyin possession of and replace the discarded manuals with the new versions of the manuals provided by CSZ. After the manuals have been discarded and replaced, customers were asked to complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that they have performed and completed the requested actions. Customers should return the form by fax to 513-772-9119 scan and email the information to B3manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, Ohio 45241. Customers with questions were instructed to contact their Blanketrol III Model 233 field action coordinator at 1-800-989-7373 or 513-772-8810.For questions regarding this recall call 513-772-8810, ext 5295. |  | 1126 | 115V - 396 units; 230V - 385 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, , NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI & WV. Internationally to ARGENTINA, AUSTRALIA, AZERBAIJAN, BELGIUM, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, GERMANY, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, KOREA, KUWAIT, MEXICO, NEW ZEALAND, PERU, SAUDI ARABIA, SINGAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, THE PHILIPPINES, UNITED ARAB EMERITES & UNITED KINGDOM. | Open |  |  | 783 | Not_Computer
Z-2202-2013 | 66027 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Sonogage EyeScan, A-Scan Biometer, Product Usage: A-Scan Biometer diagnostic device for measurement of axial eye length. | AScan Biometer | 2 | September 16, 2013 | 2013 | Sonogage Inc | During an FDA inspection on 7/15-26/2013, it was discovered that the firm lacked registration and 510(k) clearance for the device. | PREMARKET APPROVAL: No Marketing Application | Sonogage sent an Urgent: Medical Device Recall letter dated August 9, 2013 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine inventory and quarantine the affected product for return. Customers were asked to complete and return the enclosed response form. For question call 216-464-1119 or 800-798-1119. |  | 564 | 564 units | Worldwide Distribution - USA Nationwide in the following states: AK, AZ, CA, FL, GA, IL, KY, LA, MA, MD, MN, NE, NJ, NM, NY, OH, OR, PA, TX, VA and WA. Product was also shipped to Canada and Mexico. | Open |  |  | 564 | Not_Computer
Z-2207-2013 | 66204 | implant, endosseous, root-form | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite 2 Certain Prevail 5mm(D)X 4.1mm (P) X X 11.5mm (L) Dental ImplantProduct Usage: Surgical placement in the upper or lower jaw to provide a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | Osseotite, 2Certain, Prevail | 2 | September 16, 2013 | 2013 | Biomet 3i, LLC | One of the dental implants do not have the internal hex feature. One of the dental implants do not have the internal thread feature. | DESIGN: Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated June 10, 2013 to affected customer via email, facsimile, or postal mail. The letter identified the affected product, problem, and actions to be taking. Customer were instructed to check their respective inventory for the affected product and return to Biomet 3i. Biomet 3i will issue replacement product. For questions or concerns call 1-800-342-5454. |  | 215 | 215 | Worldwide Distribution - USA Nationwide in the states of MA, CT, NJ, DE, NY, FL, OK, TX, PA, DC, IL, KS, LA, MD, CA, RI, TN, MN, and SC and the countries of Canada, Japan, Australia, Italy Germany, France, Switzerland, Apain, Brazil, Uruguay, Singapore, Dominican Republic, Taiwan, Austria, Portugal, and the United Kingdom. | Open |  |  | 393 | Not_Computer
Z-2210-2013 | 65691 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 6, M/N 728256; Brilliance CT 16 M/N 728246; Brilliance CT 40, M/N 728235 & Brilliance CT 64 M/N 728231, Philips Healthcare, Cleveland, OH 44143.Intended to produce cross-sectional images of the body by computer reconstruction. | Brilliance CT 6, Brilliance CT 16, Brilliance CT 40, & Brilliance CT 64 | 2 | September 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the footswitch cover impinges on the footswitch. | DESIGN: Component Design/Selection | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Care Center at 1-800-722-9377, option 5 for information concerning this issue. | Contact the recalling firm for information | 4 | M/N: 728256, 4 units; M/N: 728246, 89 units; M/N: 728235, 9 units & M/N: 728231, 207 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, DE, FL, GA, IN, KY, MA, MD, MI, MN, MS, NC, NM, NV, NY, OH, OK, PA, SC, TN, TX, VT, WI & WV, and the countries of Australia, Austria, Bosnia & Herzegovina, Chile, China, Croatia, Denmark, Equatorial Guinea, France, Germany, Greenland, Iceland, Indonesia, Israel, Japan, Kazakhstan, Netherlands, Nigeria, Norway, Poland, Romania, South Africa, South Korea, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom & Vietnam. | Open |  |  | 309 | Not_Computer
Z-2215-2013 | 66037 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIORAPTOR 2.9 mm Suture Anchor with 2 ULTRABRAID Suture; Part Number: 72200775Biodegradable suture anchor | BIORAPTOR" 2.9 mm Suture Anchor | 2 | September 17, 2013 | 2013 | Smith & Nephew, Inc. Endoscopy Division | Pinholes detected in a small number of primary pouches that may constitute a breach of the sterile barrier. | PRODUCTION CONTROLS: Packaging | Smith&Nephew; sent an Urgent Product Recall Notification letter dated August 6, 2013, via Federal Express to all affected customers. Addendum customer letters for 4 additional part numbers were sent via Federal Express on August 19, 2013, The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to inspect their inventory, locate unused devices and quarantine them. Customers were also asked to complete the last three columns in the inventory Return Certification form indicating the quantities that need to be returned, and include their contact information in the spaces provided. Customers were instructed to contact Smith & Nephew's Demo Service Departmentat 800-343-5717 for a Return Authorization (RA#), and provide the information response form. Replacement product would be provided. For questions regarding this recall call 978-749-1073. | Contact the recalling firm for information | 43385 | 43385 US | Worldwide Distribution - USA (nationwide) Canada, Europe: Austria, Belgium, Cyprus, Czech Republic, Denmark, Germany, Greece, Finland, France, Italy, Ireland, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine and United Kingdom, Asia-pacific: Australia, China, Dubai, Hong Kong, India, Japan, Malaysia, New Zealand, Philippines, Sri Lanka, Singapore, South Korea, Taiwan, Thailand, Argentina, Bermuda, Brazil, Chile, Colombia, Costa Rica, Israel, Mexico, Puerto Rico, South Africa, and Venezuela. | Open |  |  | 224222 | Not_Computer
Z-2214-2013 | 66098 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cypher MIS Screw System, Biomet Spine Parsippany, NJ 07054. The Cypher MIS Screw System is a non-cervical spinal fixation device.1)Screw Tower - REF 14-5016602) Rod Reduce - REF 14-5016613) Counter Torque - REF 14-501662 | Multiaxial Translating MIS Screw | 2 | September 17, 2013 | 2013 | Ebi, Llc | Several complaints were reported to exhibit screw head splay in the Cypher MIS Screw System. | DESIGN: Device Design | Biomet Spine & Bone Healing Technologies sent "Urgent Medical Device Recall Notice" letters with Fax-Back Response Forms dated 8/13/2013 to distributors and OR managers via Fed Ex. Follow-up e-mails with a copy of the Recall Notice were sent to distributors on 8/13/2013. Customers with questions may contact 973-299-9300 x2162 Monday to Friday from 8:00 am to 5:00 pm EST. | Contact the recalling firm for information | 427 | 427 units | Nationwide Distribution (US), including the states of IN, MD, MS, NC, OH, PA, SC, and TX. | Open |  |  | 427 | Not_Computer
Z-2211-2013 | 66208 | prosthesis, chin, internal | FWP | Chin prosthesis. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Device Name: Terino Square Chin-Style II-Medium; Catalog Number: TSCII-M.To augment or reconstruct underdeveloped or traumatized mandibular region. | Terino Square chinStyle IIMedium | 2 | September 17, 2013 | 2013 | Implantech Associates Inc | Implantech initiated a voluntary recall of certain lots of Terino Square Chin-Style II-Medicum (Lot # 848488 and 850156), after determining that the devices from the lot numbers do not conform to the dimensional requirements for the TSCII-M. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Recall was initiated by sending recall notification letters to customers on 08/30/2013 via a combination of 1st class mail, email, and telephone. The recall notification letter, dated August 30, 2013, informs customers of the reason for recall, product description and codes, instructions on what to do with the recalled product and contact information. Within the recall notification letter, a section is provided for responses and signatures. |  | 8 | 8 | Distributed in the states of NY, KY, and CO, and the country of Israel. | Terminated | February 06, 2014 | 142 | 8 | Not_Computer
Z-2212-2013 | 66176 | apparatus, suction, ward use, portable, ac-powered | JCX | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems:The Neptune Waste management system is intended to be used in the Operating Room, Surgical Centers and Doctor's Offices to collect and dispose of surgical fluid waste as well as collect smoke generated from electrocautory or laser devices. ***The single-port manifold for the Stryker Neptune 1 (part # 700-015-000) can be used with the: Neptune 1 Gold Rover or the Neptune Bronze Rover as an accessory to Class 2 Air-handling apparatus for a surgical operating room (Sec. 878.5070); or with the Neptune 1 Silver Rover as an accessory to Class 2 Powered suction pump. ***four-port manifold for the Stryker Neptune 1 (part # 700-020-000) can be used with the: Neptune 1 Gold Rover or the Neptune Bronze Rover as an accessory to Class 2 Air-handling apparatus for a surgical operating room (Sec. 878.5070); or with the Neptune 1 Silver Rover as an accessory to Class 2 Powered suction pump. ***single-port manifold for the Stryker Neptune 2 (part # 702-025-000) can be used with the Neptune 2 Rover Ultra as an accessory to Class 2 Powered suction pump. ***four-port manifold for the Stryker Neptune 2 (part # 702-020-000) can be used with the Neptune 2 Rover Ultra as an accessory to Class 2 Powered suction pump. | Stryker Neptune 1 and Neptune 2 Waste Management Systems | 2 | September 17, 2013 | 2013 | Zeppessis Reprocessing, LLC | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems are recalled due to lacking 510K and and without a validated cleaning protocol for reprocessing the used manifolds. | PRODUCTION CONTROLS: Reprocessing Controls | Zeppessis Reprocessing, LLC sent an Urgent Medical Device Recall letter dated September 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to stop using the Zeppessis reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems immediately and instructed to destroy the products in a proper manner.Customers should fill out the Customer Response form and fax or email it back to Zeppessis.For any questions about this recall, please call Zeppessis Reprocessing, LLC at 208-559-7064. |  |  | undetermined the distributed amounts | Distributed in the states of California, Idaho, Oregon, and Washington. | Terminated | November 08, 2013 | 52 |  | Not_Computer
Z-2252-2013 | 66060 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Central Dressing Change Kit; AMS-8464CP- 6. Vygon Central Dressing Change Kit contains: 1 Pair Nitrile Gloves W/Wallet, Powderfree 1 Face Mask 1 Polylined Drape, 13" X 18" 1 ChloraPrep***, 3ml Applicator W/Insert 1 Pkg. Alcohol Swabsticks (3/Pkg.) 1 14" Extension Set W/Bionector*** 2 Dressing Sponges, 2" X 2" 1 Foam Strip, 8/5cm X 1cm 1 3M*** Tegaderm *** IV Dressing W/CHG, 2 1/2" X 4 1/2" *** Vygon MFG Dover, NH 03820***" | Microbore Extension Set with Bionector | 2 | September 18, 2013 | 2013 | Churchill Medical Systems, Inc. | Microbore tubing disconnecting from male luers on the extension sets. |  | "Urgent Medical Device Recall" letters were sent to each customer via courier service on August 12, 2013. Consignees were provided with a description of the issue and instructions on how to handle product that they received and that was further distributed by them. Questions were directed to customer service at 1-800-473-5414. |  | 3040 | 3040 | Nationwide Distribution (US), including the states of NY, MN, NJ, PA, FL, MD, and VA. | Open |  |  | 16140 | Not_Computer
Z-2254-2013 | 66223 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | POWERWAND Safety Introducer with Extended Dwell Catheter;Catalog no. 92006.Used to gain access to the vascular system to sample blood and administer fluids intravenously. | POWERWAND Safety Introducer with Extended Dwell Catheter | 2 | September 18, 2013 | 2013 | Access Scientific Inc | Access Scientific is recalling the POWERWAND Saftey Introducer with Extended Dwell Catheter because it is mislabeled. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | On 8/16/13, Access Scientific visited 2 consignees and initiated a field correction. This is the communication given to the two(2) accounts by our field personnel: "Access Scientific has inadvertently shipped mislabeled devices to your facility. The products affected were 5 Fr. POWERWANDs, Catalog # 92006, lot # I666473."The consignees were instructed to "Quarantine the devices until Access Scientific personnel arrived with the appropriate labeling and performed the corrective action of applying the correct labels to the devices and inserting the DFU into the boxes.". | Contact the recalling firm for information | 160 | 160 units | Distributed in MI and TX. | Terminated | September 24, 2013 | 6 | 160 | Not_Computer
Z-2242-2013 | 66052 | bandage, cast | ITG | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | 3M Scotchcast, Wet or Dry Cast Padding, Catalog Numbers WDP2, WDP3, WDP4, and WDP6.Intended for use in constructing casts for either wet or dry immobilization. | 3M Scotchcast | 2 | September 18, 2013 | 2013 | 3M Company - Health Care Business | 3M is conducting a voluntary product recall of selected lots of 3M" Scotchcast" Wet or Dry Cast Padding. Upon investigation of a health care provider feedback, and confirmed with internal testing, it was found that the cast padding material does not consistently repel water. Lack of repellency may result in prolonged cast dry times after exposure to wet conditions. There have been no repo | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | 3M sent a Voluntary Product Recall Notification letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to examine their inventory and set aside any affected lots subject to the recall notice. Customers were asked to inform all departments of any affected lots and provide them with a copy of the Notification letter that instructs them to contact their 3M distributor for alternative options. All affected products were to be returned to the 3M distributor they were purchased from. Customers with questions were instructed to contact the 3M Customer Helpline at 1-800-228-3957. For questions regarding this recall call 651-733-1110. |  | 1207 | 1207 cases (4 puches per case, 20 rolls per pouch) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO,CT, FL, GA, HI, IL, IN,KS,KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OR,PA, TN, TX, UT,VA, WA, and WI. Internationally to GERMANY, CANADA, BELGIUM, NETHERLANDS, HONG KONG, NEW ZEALAND, AUSTRALIA, INDIA, KOREA, SINGAPORE AND UNITED ARAB EMIRATES, JAPAN, JAMAICA.. | Open |  |  | 1207 | Not_Computer
Z-2248-2013 | 65893 | instrument, coagulation, automated | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP 700 CTS; PN 0000280050Bench top, fully automated, random access analyzer designed specifically for in vitro diagnostic clinical use in the hemostasis laboratory for coagulation and/or fibrinolysis testing in the assessment of thrombosis and/or hemostasis. The systems provide results for both direct hemostasis measurements and calculated parameters. | ACL TOP Family | 2 | September 18, 2013 | 2013 | Instrumentation Laboratory Co. | Potential carryover issue that can can cause shortened APTT clotting times on the instrument. | DESIGN: Device Design | Instrumentation Laboratory sent an Important Product Correction letter on July 19, 2013, to all affected customers. On July 23, 2013, the same communication was circulated by the IL International Group in Milan, Italy to the affected international distributors and dealers. Consignees were advised of a potential carryover issue and a workaround was provided. Consignees were asked to fax the Mandatory Response Tracking Form to 781-861-4207 or email to ra-usa.ilww.comFor technical questions customers were instructed to contact the IL Technical Support Center at 1-800-678-0710, option 2. For product availability questions, customers were instructed to contact their IL Customer Service Representative at 1-800-955-9525, Option #2.For questions regarding this recall call 781-861-4467. |  | 185 | 185 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. Internationally to Canada, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dubai, Finland, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Luthuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Open |  |  | 4799 | Not_Computer
Z-2243-2013 | 64406 | pin, fixation, smooth | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Cannulated Distal Femoral Nail (Trauma Nail), Part Number 450.851S, 420MM-Sterile. Orthppedic implant device. | Synthes Cannulated Distal Femoral Nail (Trauma Nail) | 2 | September 18, 2013 | 2013 | Synthes USA HQ, Inc. | Firm discovered that a nail that was within the scope of a previous Trauma Nail recall had been distributed. | PRODUCTION CONTROLS: Packaging Process Control | A letter dated April 18, 2013 was sent to one consignee who was identified as receiving the affected device. A review of inventory and return of any affected product on hand was requested. Negative responses were also requested. |  | 1 | 1 | 1 device was distributed in MA. | Open |  |  | 1 | Not_Computer
Z-2258-2013 | 65999 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | VisuALIF,System consisting of a 'U' shaped PEEK block in two footprint configurations, four heights and three lordosis angles., part number: 0301-1410.Intended to be implanted via an open anterior approach. | VisuALIF | 2 | September 19, 2013 | 2013 | Spine Smith Holdings, LLC | The product has the potential to be laser marked as a Medium when it actually is a Small. | PRODUCTION CONTROLS: Process Control | The firm decided to remove this lot and verify all of the markings. A stock sweep for mis-marked parts was performed at the time and it is not known why this lot is identified, however an additional stock sweep of all PEEK products from all product lines was performed beginning 07/31/2013 and no additional lots were identified. |  | 32 | 32 | Distributed in FL. | Open |  |  | 32 | Not_Computer
Z-1538-2013 | 62624 | high-power laser light show projector | REA |  |  |  |  | Laserworld RGB & Piko Nano RGB ModelsAll models are laser light show projectors, and will be used as such. | Laserworld | 2 | September 19, 2013 | 2013 | Laserworld Usa Inc | It was discovered potentially lack of certification missing laser laser safety warning and administrative labels, potentially incomplete Performance Feature set and incomplete User Information. | RADIATION CONTROL FOR HEALTH AND SAFETY ACT: Radiation Control for Health and Safety Act | CDRH approves the CAP subject to the following conditions:1. Laserworld shall expand the scope of the CAP to include all projectors distributed between 2008 and 20l2, including, but not limited to, the Proline Series AG, Revolution Series AG, Club Series AG, Club Series, Pure Light Series, and Revolution Series AG, and Laserworld will continue to evaluate whether these additional projector models have similar defects.2. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.3. Laserworld will maintain records to document the number of projectors returned under the CAP, and will document the specific items of non-compliance addressed for each product tested, repaired, relabeled and/or replaced.4. Laserworld will provide CDRH with quarterly status updates on the progress of the CAP.5. The correction of the non-compliant products must be made at no cost to thecustomers.You may proceed with implementation of the CAP (subject to the conditions. | Contact the recalling firm for information | 500 | Total distribution of more than 500 units between July1, 2008, and June 30, 2012 | US Distribution including the states of CA, CO, FL, KS and PA. | Open |  |  | 500 | Not_Computer
Z-2255-2013 | 66185 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LT-D Oropharyngeal AirwaysClass 1, 510(k) exemptThe King LT-D is intended for use in adult patients (in excess of 25 kg) for controlled ventilation during anesthesia for procedures that are short in duration and when the patient is considered to have a low risk of aspiration of stomach contents. | King LTD Oropharyngeal Airways | 3 | September 19, 2013 | 2013 | King Systems Corp. | King Systems is initiating a voluntary recall of an affected lot of its KLTD products shipped on April 30, 2013. King Systems has received a report that size 5 KLTD airways were packaged in pouches indicating a size of 2.5 and shipped to your facility. The correct size of the KLTD airway is clearly marked on the actual airway product and the physical size difference between 2.5 and 5 is distinct a |  | King Systems notified one consignee who received the affected product verbally on August 7, 2013, via email and an Urgent Medical Device Recall - Immediate Action Required notification on September 3, 2013. The Urgent Medical Device Recall - Immediate Action Required notification described the product and issue as well as includes a response form for the consignee to complete and send back to King Systems via FAX to 317-776-5175, or email to wnet@ambu.com. Questions regarding the notice are directed to: King Systems customer service representative at 800.642.5464, Monday through Friday 8:00am- 5pm EDT.For questions regarding this recall call 317-776-6823. | Contact the recalling firm for information | 54 | 9 cases; 45 products | Nationwide Distribution including NY | Open |  |  | 45 | Not_Computer
Z-2260-2013 | 66186 | clip, tubal occlusion | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Cranial Access (CRAK) Kits, Product / Catalogue No.s: INSHITH, INSHITHND, INSHITHRZNThe kits allow for access to the subarachnoid space or the lateral ventricles of the brain. The kits are intended to be used with an external drainage and monitoring system in selected patients to reduce intracranial pressure (ICP), to monitor CSF, to provide temporary drainage of CSF, and to monitor ICP. | Integra Cranial Access Kit | 2 | September 19, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra is recalling certain lots of Cranial Access Kit due to a possibility that the seal integrity of the sterile package may become compromised during shipping. | PRODUCTION CONTROLS: Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated September 4, 2013, to all affected customers that have been shipped unexpired affected product lots of IntegraÂŽ Cranial Access Kits, Catalogue No.'s INSHITH, INSHITHND, INSHITHRZN.The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review their inventory and immediately stop the use or distribution of any affected product. Customers were instructed to identify and report to Integra if they do or do not have any of the affected product lots. Customers were advised how to return any affected product lot numbers in their possession.Effectiveness Checks will be conducted to determine the level of customers who:- Integra has verified, have been notified, and / or acknowledged the recall either by returning the acknowledgement form.- Indicate they will be returning / have returned unexpired affected product lots of kits.Any affected product lot numbers that are returned from consignees will be forwarded to the Integra Salt Lake City, UT facility for final disposition per its site QMS processes.Customers with questions were advised to contact Customers Service at 1-855-532-1723.For questions regarding this recall call 801-886-9505. |  | 8648 | 8,648 | Nationwide Distribution including CA, CO, FL, IN, KS, NC, NV, NY, OK, TN, TX, VA, WA, and WI. | Terminated | January 28, 2014 | 131 | 8648 | Not_Computer
Z-2259-2013 | 66276 | material, tooth shade, resin | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | NX3 Try-In Gel. The product is used as a tooth shade resin material. | NX3 TryIn Gel | 2 | September 19, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is voluntarily recalling one lot of NX3 Try-In Gel, because some of the NX3 Try-In Gel syringes in the affected lot contain a different product material. The material in the affected syringe does not match the shade of the cement as it is labeled. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Kerr Corporation initiated the voluntary recall by sending out notifications on 09/13/2013 via USPS 1st class mail. The recall notification letter, titled "URGENT: MEDICAL DEVICE RECALL", informed customer of the recall's product description, reason for recall, potential risk, instructions to customers and contact information. |  | 197 | 197 units | Worldwide Distribution, including Nationwide in the US and the countries of Canada, Great Britain, Australia, France, Germany, Italy, Denmark, Hong Kong, China, Brazil and Saudi Arabia. | Terminated | February 11, 2014 | 145 | 197 | Not_Computer
Z-2261-2013 | 65981 | filter, blood, cardiopulmonary bypass, arterial line | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Terumo Pall AL6X Arterial Blood Line Filter, labeled in part ***Terumo Cardiovascular Systems Corporation Elkton, MD 21921***, packaged as part of a Terumo Cardiovascular Procedure Kit or as a single, sterile filter unitSingle/Sterile TerumoÂŽ Pall AL6X Arterial Filter: The Terumo Pall AL6X Arterial Filter is indicated for use in cardiopulmonary bypass procedures for the removal of micro-emboli greater than 40 microns in size, including gas emboli, fat emboli and aggregates composed of platelets, red blood cells and other debris from the arterial line and where the flow rate will not exceed 8 liters per minute. Cardiovascular Procedure (Convenience) Kit that includes the TerumoÂŽ Pall AL6X Arterial Filter: The Cardiovascular Procedure Kits are indicated for use only in the extracorporeal circuit for the cardiopulmonary bypass procedures for which the user designed. The sterile cardiovascular procedure kit is intended to be used one time for periods up to 6 hours. | Terumo Pall AL6X Arterial Blood Line Filter | 2 | September 20, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Medical device used in cardiovascular procedures is defective. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo Cardiovascuolar Services sent recall letters to Chief of Perfusion, Director of Operating Room Services, Director of Biomedical Services, and Risk Management departments on July 26, 2013. The notifications advised that filters are being recalled because some of the filters may exhibit a partial or full occlusion of an outlet or purge port. Terumo CVS advised users to stop using Terumo AL6X arterial filters from the affected lot numbers for replacement. A supply of replacement filters for customer's immediate use accompanied the notification. For users ordering the AL6X filters as part of a Cardiovascular Procedure Kit, Terumo CVS is providing a preliminary supply of replacement Terumo AL6X arterial filters in which users may follow the instructions to replace the filter or users may return the Cardiovascular Procedure Kit and Terumo CVS will replace the filter for them. Customers with questions were advised to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 800-521-2818. | Contact the recalling firm for information | 15817 | 15,817 filters | Nationwide Distribution including WA, IN, OK, TX, MI, CA, NY, CT, AZ, FL, WV, NC, VA, PA, TN, GA, AL, MA, SC, OH, MS, MO, SD, WI, AR, AND NV. | Open |  |  | 15817 | Not_Computer
Z-2262-2013 | 66049 | oxygenator, cardiopulmonary bypass | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000. The QUADROX oxygenators are intended for use in extracorporeal circulation during cardiopulmonary bypass in cardiac surgery. | QUADROX Oxygenator | 2 | September 20, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The manufacturer has conducted a review of the results of product testing performed during the last twelve months and has identified specific lots of product which may not meet specifications for bond strength on the inlet and/or outlet connector. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Maquet sent a Urgent Product Recall letter/Response Form, dated 8/15/2013 via Fed Ex. On 8/26/2013 Maquet sent a letter to their customers informing them that all shipments of oxygenators and stainless steel hold clamps have been placed on a temporary hold due to an agency review in connection with FDA clearance. Customer service may be contacted at 1-888-627-8383, (option 2 followed by option 2) Monday though Friday between the hours of 8:00 am and 6:00 pm EST.Maquet issued a Certificate of Medical Neccesity (CMN) on September 17, 2013 to notify Quadrox customers that they may continue to receive and use their current inventory of Quadrox oxygenators and clamps by acknowledging the conditions provided in the CMN. | Contact the recalling firm for information | 1182 | 1,182 units | Worldwide Distribution, including Nationwide (US) and foreign countries. | Open |  |  | 1182 | Not_Computer
Z-2263-2013 | 66247 | accessories, germicide, cleaning, for endoscopes | NZA | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reliance Endoscope Processing System (REPS), STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060-1834 USA www.steris.com.The intended use of the REPS is washing and high level disinfection of up to two manually pre-cleaned, immersible, reusable, heat-resistant, semi-critical devices. | Reliance Endoscope Processing System (REPS) | 2 | September 20, 2013 | 2013 | Steris Corporation | The air pipe that delivers unfiltered ambient air from the procedure room into the water bottle on the listed FUJINON G5 & G8 type flexible GI endoscopes does not connect to the main air/water channel within the light guide (LG) connection head and, as a result, the air pipe does not receive forced flow during the Reliance EPS processing cycle. |  | Steris sent an "URGENT PRODUCT SAFETY NOTICE" dated September 5, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 266 | 266 units | Worldwide Distribution-USA (nationwide) including the states of AK, CO, FL, GA, HI, ID, IL, KS, LA, MD, MN, MO, MT, NE, NM, OH, PA, SC, UT, & WY and the countries of Australia, Canada & Spain. | Open |  |  | 266 | Not_Computer
Z-2265-2013 | 66063 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | ENDOPLUS; Precision Laparoscopic Instruments; REF 345VA; Description 5 mm Traumatic Grasper S/A, Standard; Non-sterile; Manufactured by Endoplus; 750 Tower Rd., Suite A, Mundelein, IL 60060 USA.Intended for grasping, dissecting, retraction, clamping in conjunction with the laparoscopic during laparoscopic surgery. | Endoplus 5 mm Traumatic Grasper S/A, Standard | 2 | September 21, 2013 | 2013 | Endoplus, Inc. | Multiple complaints reporting that the moveable jaw has broken during laparoscopic procedures. | PRODUCTION CONTROLS: Process Control | Letters were sent on 7/15/2013 to all customers notifying them of the recall. The letters requested that customers return the recalled devices via private courier using an enclosed account number. Replacement devices will be made available upon customer request. |  | 9 | 9 units | Distributed in the states of FL, TX, and VA. | Open |  |  | 90 | Not_Computer
Z-2267-2013 | 65883 | enzyme immunoassay, cocaine and cocaine metabolites | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | DRI Cocaine Metabolite Assay, labeled as Multigent Cocaine Assay for Abbott Laboratories.100 ml, IVDMfg by:Microgenics Corporation,46360 Fremont Blvd.Fremont, CA 94538The DRI Cocaine Metabolite Enzyme Immunoassay is a homogeneous enzyme intended for the qualitative and semi-quantitative determination of benzoylecgonine (Cocaine Metabolite) in human urine with either 300 ng/mL or 150 ng/mL as a cutoff calibrator. The assay provided only a preliminary analytical test result. A more specific alternative chemical method must be used to obtain a confirm analytical result. | DRI Cocaine Metabolite Assay, Multigent Cocaine Assay | 2 | September 23, 2013 | 2013 | Microgenics Corp | Surveillance of reagent stability for specific lots of DRI Cocaine Metabolite Reagent show an unexpected decreasing change in the reaction rate, indicating a faster than expected reagent degradation. |  | Thermo Fisher Scientific sent an Urgent Medical Device Recall letters dated July 23, 2013, except Abbott labs who was mailed a letter on August 2, 2013. Abbott had previously been notified by telephone of the recall.The letter identified the product the problem and the action needed to be taken by the customer.Customers advised users on reagent stability issue with specific lots and the fact that it indicated a faster than expected degradation of the reagent. Microgenics asked that all use of the affected lots be stopped. Customers were informed that since each patient run is monitored with the controls at +/-25% of the cutoff, a possibility of reporting incorrect patient results is minimal to nonexistent. The use of appropriate controls addresses an risks of incorrect results.In addition to stopping use of the affected reagent, customers are asked to contact Thermo Fisher Scientific for replacement and to keep a copy of the recall letter for laboratory files. If the affected lots of reagent were forwarded to any other laboratories, they should be contacted immediately.. The enclosed acknowledgement form is to be completed and returned indicating that the letter was received, how much product, if any was on hand and destroyed.Questions and concerns may be addressed to Microgenics Technical Service at 1-800-232-3342 or 510-979-5417. |  | 1473 | 1473 units in distribution. | US Distribution including the states of DC, MA, NJ, TX. FL, CA, WA, NY, TX, CT, KY, LA, AR, TN, AL, GA, CT, VA, NC, KS, MD, OH and IL | Terminated | December 02, 2013 | 70 | 1473 | Not_Computer
Z-2270-2013 | 65348 | wheelchair, powered | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Power Wheelchair, one per boxTo provide mobility to a person restricted to a sitting position. | Invacare Power Wheelchair | 2 | September 23, 2013 | 2013 | Invacare Corporation | Unintended slowing or deceleration, and on rare occasion,unanticipated and unattended acceleration of power wheelchairs equipped with SPJ and MPJ type joysticks. | PRODUCTION CONTROLS: Process Control | Invacare sent an Urgent Patient Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their wheelchair povider to have their joystick replaced. Customers with questions were instructed to call their provider at Invacare's recall support line at 877-413-6008. Go to www.invacare.com/joystickrecall for more information regarding this recall. | Contact the recalling firm for information | 166230 | 166,230 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada, Australia, Bermuda, Brazil, China, Costa Rica, Germany, France, United Kingdom, Guam, Hong Kong, Israel, India, Kuwait, New Zealand, Puerto rico, Qatar, Saudi Arabia, Singapore, Thailand, Trinidad, Tobago, and South Africa. | Open |  |  | 166230 | Not_Computer
Z-2269-2013 | 66234 | port & catheter, implanted, subcutaneous, intravascular | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Smiths Medical, Deltec:REF 21-8011-24 Port-A-Cath Dual-Lumen. Titanium Venous Access System.REF 21-8052-24 & REF 21-8053-24 , Port-A-Cath II Dual-Lumen, Polysulfone /Titanium Venous Access System.Sterile EO, Rx Only.PORT-A-CATHÂŽ and PORT-A-CATHÂŽ II implantable venous and arterial access systems are designed to permit repeated access to the vascular system for the parenteral delivery of medications, fluids, and nutritional solutions and for the sampling of venous blood. Introducer sets are indicated for the introduction of catheters into the subclavian vein. | PORTACATH and PORTACATH II implantable venous and arterial access systems | 2 | September 23, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a voluntary recall of PORT-A-CATHÂż and PORT-A-CATHÂż II Implantable Venous and ArterialAccess Systems and Introducer Sets.Smiths Medical has become aware of a small number of customer reports regarding the introducer sheath and the introducer opening being too small for the width of the catheter with the use of these Sets.Smiths Medical has received no reports | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Smiths Medical ASD, Inc. sent an Urgent Field Safety Notice dated September 6, 2013, to all affected customers. The letter described the problem and the product involved in the recall. The letter included the "Advice on Action to be Taken by the User" which instructed consignees to inspect and remove affected devices and to complete and return the "Urgent Field Safety Notice Confirmation Form. The letter also included information on the "Transmission of this Urgent Field safety Notice" which included information where customers should report issues related to the product to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcomplaints@smiths-medical.com. For questions on the information provided they should contact 800-258-5361 (Smiths Medical's Customer Service Department).For questions regarding this recall call 1-800-332-0178. | Contact the recalling firm for information | 372 | 372 | Worldwide Distribution - USA including AZ, AR, CA, GA, IL, IN, KY, MD, MA, MI, NV, NJ, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WA, WV, WY.Internationally to DENMARK, BELGIUM CANADA. | Open |  |  | 562 | Not_Computer
Z-2278-2013 | 65009 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 3.5 MM LCP(R) Distal Humerus SystemThe Synthes 3.5 MM LCP Distal Humerus System is intended for fixation of fractures of the distal humerus, olecranon, and ulna in adults and adolescents (12-21) in which the growth plates have fused. Specifically, distal humerus plates are indicated for intra-articular fractures, comminuted supracondylar fractures, osteotomies, malunions and non-unions of the distal humerus. | Synthes 3.5 MM LCP(R) Distal Humerus System | 2 | September 25, 2013 | 2013 | Synthes USA HQ, Inc. | Recall is being initiated due to the part being mislabeled (Part Number 241.267 was mislabeled with Part Number 241.276). | PRODUCTION CONTROLS: Labeling Mix-Ups | Synthes sent an Urgent Notice Medical Device Recall letter dated March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by the checking the appropriate box indicating affected product located. Return the Verification Form with the product to: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. | Contact the recalling firm for information | 2 | 2 | Nationwide Distribution including CA,and VA.. | Open |  |  | 2 | Not_Computer
Z-2283-2013 | 65922 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The FL23SE Bed is an AC-powered adjustable hospital bed with four built-in electric DC motors and remote controls that is intended for medical purposes. It can be operated by the patient or the caregiver to adjust the height and surface contour of the bed. The bed includes movable and latchable siderails. The FL23SE Bed is intended for use with patients for procedures, therapy, recovery in healthcare environments, and to transport patients between bays and procedural rooms.The FL23SE Bed is intended for use with patients for procedures, therapy, recovery in healthcare environments, and to transport patients between bays and procedural rooms. | FL23SE MedSurg Bed | 2 | September 26, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | The nurse call cable (Part Number: QDF23-0571) on the FL23SE Bed is located right under the opening of the bed head box. If an accessory (i.e. IV pole) is inserted into this slot, the circuit directing the nurse call system can be damaged. | DESIGN: Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated August 19, 2013, to all affected consignees. The notification described the affected product, reason for recall, potential hazards, and recall customer actions and reply instructions. A Stryker Representative will contact the consignees to arrange a time to repair the unit(s) and eliminate the risk of the potential interference . Urgent questions or concerns are directed to Blaine Burnett (269)389-6954. Stryker Medical business hours are Monday-Friday 8 a.m-5 p.m. (ET).For questions regarding this recall call 269-324-6609. |  | 61 | 61 units | Worldwide Distribution - USA including MI, WI, VA and internationally to Canada. | Terminated | January 29, 2014 | 125 | 61 | Not_Computer
Z-2280-2013 | 66209 | cresolphthalein complexone, calcium | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Calcium CP Reagent, Model # A11A01633Calcium plays an essential role in many cell functions: intracellularly in muscle contraction and glycogen metabolism, extracellularly, in bone mineralization, in blood coagulation and in transmission of nerve impulses. Calcium is present in plasma in three forms: free, bound to proteins or complexed with anions as phosphate, citrate and bicarbonate. Under physiological conditions, calcium balance is determined by the relationship between calcium intake and calcium absorption and excretion. Urinary excretion is an important determinant of calcium retention in the body. Decreased total calcium levels can be associated with diseases of the bone apparatus (especially osteoporosis), kidney diseases (especially under dialysis), defective intestinal absorption and hypoparathyroidism. Increased total calcium can be measured in hyperparathyroidism, malignant diseases with metastases and sarcoidosis. Calcium measurements also help in monitoring of calcium supplementation mainly in the prevention of osteoporosis. | ABX PENTRA Calcium CP Reagent, Model A11A01633 | 3 | September 26, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all PENTRA 200 and PENTRA 400 customers of a potential for reportingincorrect calcium results when using the ABX PENTRA CALCIUM CP reagent. |  | Horiba Medical sent an Urgent Safety Notice dated September 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter dated September 6, 2013 will be sent to all PENTRA 200 and PENTRA 400 customers of a potential for reporting incorrect calcium results when using the ABX PENTRA CALCIUM CP reagent (Reference #: A11A01633) Lot#: 014664503. The recall letter informs the customers that the initial functional tests indicate that there is no interference identified with regards to the patient results. However, with time, there may be an increase in the quantity of precipitates and this could potentially interfere with the performance of this reagent. Customers are informed of the actions to be taken and to complete the customer response form within 10 days.Further questions please call (949) 453-0500 Ext. 208. | Contact the recalling firm for information | 1240 | 2,658 units ( 1240 US 1418 OUS ) | Worldwide Distribution - USA including the country of Canada. | Open |  |  | 2658 | Not_Computer
Z-2279-2013 | 66150 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model#72313EThe SmartSite Low Sorbing Infusion Sets are used to administer fluid and medications through a needle or catheter inserted into the patient's artery or vein. The SmartSite Low Sorbing set is comprised of components commonly found on intravascular administration sets and extension sets. It includes a 1.2 micron filter, needleless SmartSite connector, low sorbing tubing segment, pinch clamps, and luer connectors. | SmartSite Low Sorbing Infusion Set | 2 | September 26, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model 72313E, Lot Number13016314, because of an incorrect filter was assembled on this lot. The correct filter is a 1.2 micronfilter (blue color) but a 0.2 micron filter (clear color) may have been added instead. | TRAINING: Employee Error | CareFusion sent an Urgent Medical Device Recall Notification dated August 30, 2013, to all affected customers. The letter informed the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers were instructed to immediately check their inventory for the affected product, if any affected product is encountered, customers were instructed not to use. Customers were instructed to complete the customer response form and return the form to CareFusion with the affected product for replacement.. A distributor letter was also sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors were instructed to contact the Customer Support Team, return customer response form, and if they have questions then they were instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 600 | 600 units | CA, FL, AZ, and TN | Terminated | October 28, 2013 | 32 | 600 | Not_Computer
Z-2282-2013 | 64236 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Oracle Spacer System Slap Hammer, Part Number 03.809.690, Lot Number 3723847The Oracle Spacer is indicated for use in patients with degenerative disc disease (DDD) at one or two contiguous levels from L2 to S1 whose condition requires the use of interbody fusion combined with supplemental fixation. The Oracle Spacer is indicated as a vertebral body replacement device intended for use in the thoracolumbar spine (T1-L5) to replace a diseased vertebral body resected or excised during partial and total vertebrectomy procedures or the treatment of tumor or trauma (i.e., fracture), to achieve anterior decompression of the spinal cord and neural tissues, and to restore the height of the collapsed vertebral body. The Slap Hammer, 03.809.690, is a dual sided device with two separate attachment options. The Slap Hammer allows a user (surgeon) to remove a previously placed trial implants. | Synthes Oracle Spacer System Slap Hammer, Part Number 03.809.690 | 2 | September 26, 2013 | 2013 | Synthes USA HQ, Inc. | This recall is being initiated in response to complaints received where it was reported that the proximal end of the Slap Hammer broke during use. | DESIGN: Device Design | SYNTHES sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees examine their inventory, remove the recalled product from use and return them to Synthes. Customers were also instructed to completed the attached Verification Section at the end of the letter indicating whether or not they had any of the affected product and the number of devices found. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. | Contact the recalling firm for information | 27 | 27 | Worldwide Distribution - USA including CO, PA, TN, VA, TN, WA, and internationally to Czech Republic. | Open |  |  | 27 | Not_Computer
Z-2284-2013 | 66355 | replacement, ossicular prosthesis, total | ETA | Total ossicular replacement prosthesis. | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Universal Titanium Prosthesis, Round Head, with Flex H/A titanium shoe. The ear prosthesis is implantable and indicated for hearing loss. | Universal Titanium Prosthesis, | 2 | September 26, 2013 | 2013 | Medtronic Xomed, Inc. | One lot of the product was assembled using the Cam Head version of the same prosthesis instead of the Round head Shaft. | PRODUCTION CONTROLS: Process Control | An "URGENT PRODUCT RECALL" letter was faxed to customers following the initial telephone notification. The letter described the product issue and provided recommended actions. Questions were directed to customer service at 800-874-5797. | Contact the recalling firm for information | 60 | 60 | Nationwide Distribution (US), including the states of CA, NE, GA, and IL. | Terminated | November 07, 2013 | 42 | 60 | Not_Computer
Z-2274-2013 | 66070 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | BD Epilor SyringeBD Franklin Lakes, NJ 07417BD EDC Laagstraat 57, B-9140 Temse-BelgiumThe Epilor syringe is used in conjunction with an epidural needle with the purpose of finding the epidural space. | BD Epilor Syringe | 2 | September 26, 2013 | 2013 | Becton Dickinson & Company | BD has received several reports indicating that the plunger of the 7 mL BD Epilor plastic BD Luer-Lok Loss of Resistance Syringe may stall or stick within the syringe preventing the user from detecting a loss of resistance upon entering the epidural space. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | BD sent a product advisory notice to distributors/customers dated August 23, 2013 via UPS 2nd day mail. Consignees will be contacted by phone email and letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware of the issue with the product and to follow the standard routine procedure to ensure that there is no resistance in the syringe prior to use. Any syringe that shows sticking or stalling should be discarded. Customers who distributed this product, should identify their customers and notify them at once of the product advisory. Customers with questions were instructed to call BD Customers Service at 1-800-237-2762.For questions regarding this recall call 201-847-6800. | Contact the recalling firm for information | 1335885 | 1,335,885 units | Worldwide Distribution - USA including FL, VA, TX, NY, IL, AZ, MI, RI, NJ, SC, MO, TN, CA, MN, UT, GA, NE, OH, MD, PA, MT and internationally to Brazil, Australia, Canada, and Puerto Rico.. | Open |  |  | 1335885 | Not_Computer
Z-2281-2013 | 66152 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion Gravity Set, Model #44000-07The CareFusion Gravity Sets are used to administer fluid and medication through a needle or catheter inserted into the patient's artery or vein. The CareFusion Gravity Set is comprised of components commonly found on intravascular administration sets and extension sets. It includes a check valve, injection ports, 4-way stopcock, anti-siphon valve, roller clamp and tubing. | CareFusion Gravity Set | 2 | September 26, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Gravity Set (Model 44000-07) because of an incorrect expiration date. The affected lot numbers of the Gravity Set are labeled witha 5 year expiration instead of three (3) years. | TRAINING: Employee Error | CareFusion sent an Urgent Medical Device Recall Notification letter dated August 30, 2013, to all affected customers to inform them that CareFusion is recalling the Gravity Set, Model 44000-07, Lot Numbers 12086930, 12106215, 13016408, 13016834, 13025446 and 13025672 as a result of an incorrect expiration date. The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and if they have questions then they are instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 2850 | 2,850 units | IL | Terminated | December 26, 2013 | 91 | 2850 | Not_Computer
Z-2285-2013 | 66151 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model# 10015048. Intravascular administration set. | SmartSite | 2 | September 26, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model# 10015048, because there is a potential for separation at the tubing sleeve below the drip chamber. | DESIGN: Process Design | A customer letter dated August 30, 2013 will be sent to customers to inform them that CareFusion is recalling the SmartSiteÂŽ Low Sorbing Infusion Set Models 10015048, Lot Numbers 130153883 and 13025342 as a result of a potential for separation at the tubing sleeve below the drip chamber.The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on 8/30/13 to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and they are instructed to contact the CareFusion Support Center at 1-800-562-6018 with any questions. . |  | 15860 | 15860 units | Nationwide in US: PA, OK, TX, FL, NC, CA, IN, WI | Open |  |  | 15860 | Not_Computer
Z-2291-2013 | 66242 | system, nuclear magnetic resonance imaging | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750 3.0T, Optima MR450w 1.5T, and Discovery MR450 1.5T .scanner designed to support high resolution, high signal-to-noise ratio, and short scan times. The GE SignaÂŽ MR 750 System is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images, spectroscopic images, parametric maps, and/or spectra, dynamic images of the structures and organs of the entire body, including but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. The images produced by the GE SignaÂŽ MR 750 System reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images when interpreted by a trained physician yield information that may assist in diagnosis. The indications for use for the 3.0T GE SignaÂŽ MR750 System are similar to those for the SignaÂŽ HDx Magnetic Resonance System and the Siemens MAGNETOM Verio MR System. K113490: The Optima" MR450w is a whole body magnetic resonance scanner designed to support high resolution, high signal-to-noise ratio, and short scan times. It is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images, spectroscopic images, parametric maps, and/or spectra, dynamic images of the structures and/or functions of the entire body, including but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. Depending on the region of interest being images, contrast agents may be used. The images produced by the Optima" MR 450w reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and/or spectra when interpreted by a trained physician yield information that may assist in diagnosis. K083147: The Discovery MR 450 is a whole body magnetic resonance scanner designed to support high resolution, high signal-to-noise ratio, and short scan times. It is indicated for use as a diagnostic imaging device to produce axial, sagittal, coronal, and oblique images spectroscopic images parametric maps, and/or spectra, dynamic images of the structures and/or functions of the entire body, including, but not limited to, head, neck, TMJ, spine, breast, heart, abdomen, pelvis, joints, prostate, blood vessels, and musculoskeletal regions of the body. Depending on the region of interest being imaged, contrast agents may be used. The images produced by the Discovery MR450 reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and/or spectra when interpreted by a trained physician yield information that may assist in diagnosis. | GE Healthcare, Discovery MR750 3.0T, Optima MR450w 1.5T, and Discovery MR450 1.5T . | 2 | September 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to exposed yellow foam on the end bells of Discovery MR450, Discovery MR750, and Optima MR450w products.The yellow foam strip used between the end bell covers and patient bore of certain GE MR Products may extend past the surface of the covers. The foam itself is not hazardous. However, any fluid or body substance absorbed | DESIGN: Component Design/Selection | GE Healthcare sent and Urgent Medical Device Correction letter dated August 13, 2013, to all affected consignees. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and contact Information. Customers were instructed that the system may continue to be used as normal even with the yello foam protruding in the bore. Customers were told not to attempt to remove the yellow foam. Customers with questions were instructed to contact their local service representative if they have any questions concerning the notification.For questions regarding this recall call 262-513-4122. | Contact the recalling firm for information | 828 | 828 (443 US, 385 OUS) | Worldwide Distribution USA including DC, all states except HI, ND, RI and WY and Internationally to Australia, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Korea (Republic of), Kuwait, Lebanon, Martinique, Mexico, Netherlands, New Zealand, Norway, Oman, Poland, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, .OUS: AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Open |  |  | 828 | Not_Computer
Z-2292-2013 | 65043 | intervertebral fusion device with bone graft, lumbar | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Solitaire-35 Peek Solitaire Anterior Spinal System. The Solitaire-35 is an intervertebral body fusion device. | Solitaire35 | 2 | September 27, 2013 | 2013 | Ebi, Llc | Several complaints were recorded regarding one of the three screws driving through the cage of the Solitaire 35 implant. | PRODUCTION CONTROLS: Process Control | Biomet LLC sent "Urgent Medical Device Recall" letters/fax back response forms on Tuesday April 9, 2013 via Fed Ex to all distributors. The notification identified the problem, product involved, and risk factors. In addition, the notification provides instructions to the representatives on how to return the product to EBI, LLC. Surgeons that have implanted the product were contacted on April 12, 2013. |  | 1730 | 1730 units | Nationwide distribution: US including states of: CA, GA, IN, KY, MO, NY, OR, TX, and WI. | Open |  |  | 1730 | Not_Computer
Z-2272-2013 | 66184 | circuit, breathing (w connector, adaptor, y piece) | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flex2 Breathing CircuitClass I 510(k) ExemptAn anesthesia breathing circuit is a device that is intended to administer medical gases to a patient during anesthesia. It provides both an inhalation and exhalation route and may include a connector, adaptor, and Y-piece. | Universal Flex2 Breathing Circuit | 1 | September 27, 2013 | 2013 | King Systems Corp. | King Systems received a customer complaint indicating that the customer had ordered a breathing circuit with a Latex breathing bag. The product they received contained a Latex breathing bag, but was labeled as non-latex. | PRODUCTION CONTROLS: Labeling Mix-Ups | Ambu, Inc. sent a letter dated June 24, 2013, informing the customer that they have decided to discontinue the manufacture and sales of the King Systems latex breathing bags and to conver all customers to their King Systems latex-free breathing bags. The customer was advised to adjust their inventories accordingly and transition to the latex free part number as their inventory levels reach their minimums. For questions the customer was advised to call 317-403-8677.For questions regarding this recall call 317-776-6823. | Contact the recalling firm for information | 41 | 1 case; 40 units | Nationwide Distribution - NE only. | Terminated | November 18, 2013 | 52 | 41 | Not_Computer
Z-2302-2013 | 66334 | catheters, transluminal coronary angioplasty, percutaneous | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular |  | NC Sprinter; Rapid Exchange Balloon Dilatation Catheter;Sterile; 2.25 mm; Model NCSP22512X.Manufacturer: Medtronic, Inc. 710 Medtronic Parkway Minneapolis, MN 55432The NC Sprinter Rapid Exchange Balloon Dilatation Catheter is a Percutaneous Transluminal Coronary Angioplasty (PTCA) Rapid Exchange System. The balloon at the distal end of the catheter can be inflated to a defined diameter at a specific pressure (see labelling). The proximal end of the catheter has a female luer for attachment to an inflation device. The catheter provides a lumen which enables the use of a guide wire to position the catheter. Radiopaque balloon marker(s) enable accurate placement. Shaft markers for brachial and femoral techniques are in place. | NC Sprinter; Rapid Exchange Balloon Dilatation Catheter | 2 | September 30, 2013 | 2013 | Medtronic Vascular, Inc. | Compliance chart included in lot 207002011 of the NC Sprinter RX 2.25 x 12 mm products lists the size of the device as 2.5mm instead of the correct 2.25mm size. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Medtronic sent an Urgent Medical Device Recall letter beginning September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that your facility has received affected NC Sprinter RX Balloon Dilatation Catheters as noted in the attached customer notification detail report. Consequently, Medtronic is asking you to take the following actions:1. Remove and quarantine all affected catheters that remain in your inventory.2. Return the affected catheters to Medtronic. A Medtronic representative can assist in facilitating the return and replacement of this product as necessary.3. Complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.Medtronic has taken the necessary steps to prevent any future shipment of mislabeled product. Regulatory agencies will be notified about this recall as applicable.Please share this notification with others in your organization as appropriate. If any NC Sprinter RX Balloon Dilatation Catheters within the scope of this recall have been forwarded to another facility, please notify that facility accordingly and facilitate the retrieval of the affected product.We appreciate your cooperation and apologize for the inconvenience that this issue may cause. Should you have any questions, please contact your local Medtronic representative. |  | 122 | 122 | Worldwide Distribution - USA and the countries of Australia, Belgium, Canada, Hong Kong, Malaysia, New Zealand, Singapore, Spain, Switzerland, Taiwan and Thailand. | Terminated | February 10, 2014 | 133 | 122 | Not_Computer
Z-2303-2013 | 66108 | acid, lactic, enzymatic method | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lactate Membrane units for E7077 Electrode Radiometer Medical ApS. The intended use is in vitro testing of samples of whole blood. | Lactate Membrane units | 2 | September 30, 2013 | 2013 | Radiometer America Inc | RADIOMETER became aware that some membranes may have enzyme residue on the outer membrane. The enzyme residue may cause an initial negative bias on the reported Lactate result upon replacement of the Lactate membrane. | PRODUCTION CONTROLS: Process Control | On 11/07/2012, the firm sent "Urgent Field Safety Notices to their customers. The product issue was described and recommended actions were provided. Customers may contact 1 (800) 736-0600, option 4 if questions. |  | 13636 | 13,632 boxes/4 membranes | Nationwide Distribution, including the states of: AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 24, 2013 | 85 | 13636 | Not_Computer
Z-2306-2013 | 65697 | prosthesis, vascular graft, of 6mm and greater diameter | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Advanta SuperSoft Grafts | Advanta SuperSoft Grafts | 2 | September 30, 2013 | 2013 | Atrium Medical Corporation | Inner tray label may adhere to wall of outer tray which make it difficult to remove the tray onto the sterile field | DESIGN: Packaging Design/Selection | Atrium Medical Corporation issued an Urgent Medical Device Safety Alert letter on July 16, 2013, via UPS to customers and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. In addition, it provides instructions how to handle product that may be affected.For questions customers were instructed to contact their Atrium Medical Customer Service at 603-880-1433 Prompt 1 or by e-mail at returns@atriummed.com. | Contact the recalling firm for information |  |  | Worldwide Distribution - USA (nationwide) and internationally to the following countries:American Samoa AntillesArgentina Australia Austria Bangladesh Belgium Bolivia Brazil Canada Chile Colombia Egypt El Salvador Finland France Germany Great Britain Greece Guatemala Hong Kong India Indonesia Ireland Italy Jamaica Japan Korea Lebanon Libya Malaysia MauritiusMexico Neal New Zealand Nicaragua Norway PakistanPanama Peru Philippines Poland Portugal Puerto Rico Romania Russia Saudi Arabia Singapore Slovenia South Africa Spain Switzerland Taiwan Thailand the Netherlands Trinidad Tunisia | Open |  |  |  | Not_Computer
Z-0002-2014 | 66189 | airway, oropharyngeal, anesthesiology | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LTS-D Oropharyngeal AirwaysProduct Usage: The King LTS-D is intended for use in adult patients (in excess of 25 kg) for controlled ventilation during anesthesia for procedures that are short in duration and when the patient is considered to have a low risk of aspiration of stomach contents. | King LTSD | 2 | October 01, 2013 | 2013 | King Systems Corp. | On December 13, 2012, King Systems initiated a voluntary recall of one (1) case(s) (5) individual products) from one (1) lot [(IV1V3) 62 cases, 310 products]of King LTSD Oropharyngeal Airways because the package label indicated that the airway was a size 5 when in fact the airway in the package was a size 4. When King manufacturers a lot, all of the components are the same size, part number etc. I | PRODUCTION CONTROLS: Error in Labeling | King Systems notified the only one (1) consignee who received the affected product of the problem via a customer complaint. Five (5) of the five (5) affected products were destroyed by consignee. | Contact the recalling firm for information | 6 | 1 case; 5 units | USA Nationwide Distribution in the state of NY | Terminated | November 08, 2013 | 38 | 6 | Not_Computer
Z-0004-2014 | 66357 | tape and bandage,adhesive(with disinfectant) | MXE |  |  | General & Plastic Surgery | 510(k) | SHOPKO Waterproof Adhesive Pads Antibacterial.Used for covering minor cuts, scrapes and burns. | SHOPKO Waterproof Adhesive Pads Antibacterial | 2 | October 01, 2013 | 2013 | ASO, LLC | No antibacterial bandages were packed in boxes labeled as antibacterial. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | A letter was sent via e-mail and mail to the only customer on August 23, 2013 that a recall of two lots was in progress and to examine their inventory and stop distribution of the product. Product was to be either place in quarantine or destroyed. A credit for all returned product or by returned completed certificate of destruction. |  | 37 | 37 cases (72 retail boxes in each case) | Distributed in the states of ID, WI, IN, and NE. | Open |  |  | 37 | Not_Computer
Z-0006-2014 | 66128 | cuff, tracheal tube, inflatable | BSK | Inflatable tracheal tube cuff. | Anesthesiology | Anesthesiology | 510(k) | H&H; Emergency Cricothyrotomy Kit (Part No. HHECT01/ HHECTK01, NSN 6515-01-573-0692), vacuum sealed and packaged in a 5 mil poly bags, labeled in part ***H & H Associates, Inc. CAGE 1NNH6 www.gohandh.com Ordinary, VA 23131 (800) 326-5708***. The product is used to maintain fracture alignment. | H&H; Emergency Cricothyrotomy Kit | 1 | October 01, 2013 | 2013 | H & H Associates | Medical device tubing used for endotracheal airway application may become weakened and rendered ineffective. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | H&H; Medical Corporation notified their customers by "Urgent Medical Device Recall" letter dated August 27, 2013 and advised that the cuff balloon wall may become weakened and rendered ineffective. The firm advised their customers to discontinue use and take steps to remove affected units from further use. The manufacturer will replace the current endotracheal airway and cuff within the kits identified with a new cuffed airway in its own sterile packaging. The recaller included two response forms for customers to return: The Acknowledgement of Recall Form is provided for those who can quarantine the Cricothyrotomy Kits indicated in the recall. The form confirms notification of this recall and discontinued use of affected devices and provides details regarding the accounting and quarantine of the product lots in question for return and correction. The Product Recall Replacement Part Request Form is only required when inventory quarantine is not feasible. This form acknowledges the receipt of this notification and the undersigned accepts the responsibility for affixing the replacement airways along with the instructions that H&H; Medical Corporation will provide. The undersigned also assumes the responsibility to inform users regarding the subject of this recall, requires the use of the replacement airways over the airways provided in the affected kits and requires the disposal of the airways within the kits upon opening. For questions regarding the corrective request, please contact Ms. Michelle Morgan at 804-642-3663 or at mmorgan@gohandh.com during the hours of 8:00 a.m.-4:30 p.m. Eastern Standard Time, Monday through Friday. | Contact the recalling firm for information | 6619 | 6,619 kits | Nationwide distribution. | Open |  |  | 6619 | Not_Computer
Z-0003-2014 | 66026 | catheter, percutaneous | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis ADROIT 6F Guiding CatheterProduct Usage:For intravascular introduction of interventional and/or diagnostic devices into the coronary or peripheral vascular system. | Cordis ADROIT 6F Guiding Catheter | 2 | October 01, 2013 | 2013 | Cordis Corporation | Due to a labeling error related to the conversion of the inner diameter of the catheter from inches to millimeters. The English dimension is correct and is the primary dimension. | PRODUCTION CONTROLS: Error in Labeling | Cordis sent an Urgent Medical Device Recall letter dated August 12, 2013 Acknowledgement Form to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected product in a manner that ensures the product will not be used. Customers were asked to review, complete, sign and return the enclosed Acknowledgement form directly to Cordis to the fax number on the form. For questions call the Customer Service number at 800-561-7683. |  | 1238 | 1238 units | Worldwide Distribution - USA Nationwide in the states of AZ, CA, FL, ID, LA, MA, MD, MI, MN, MO, NC, NH, NY, OH, OK, PA, SC, TX, VA, WI, and WV and the countries of: Austria, Belgium, Czech Republic, France, Germany, Greece,Hungary, Israel, Italy, Latvia, Netherlands, Poland, Portugal, Russian Federation, Slovenia, Spain, Sweden, Switzerland, United Arab, and United Kingdom | Open |  |  | 1238 | Not_Computer
Z-2319-2013 | 66347 | apparatus, suction, ward use, portable, ac-powered | JCX | Powered suction pump. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | DeVilbiss Vacu-Aide QSU Suction Unit (Disposable Suction Container/Filter), 800 ML DISP Container, Unassembled, 800 ML DISP Container w/6' patient tubing, and 800 ML DISP Container w/6' patient tubing, assembled.DeVilbiss Vacu-Aide QSU Suction Unit (Disposable Suction Container/Filter) is used to remove fluids from the airway or respiratory support system and infectious materials from wounds. It is available for use by consumers by physician order. | DeVilbiss Healthcare | 2 | October 01, 2013 | 2013 | DeVilbiss Healthcare LLC | End users of the DeVilbiss Disposable Suction Container/Filter had exposed the bacterial filter cartridge to fluid, thereby occluding the filter and rendering it unusable. |  | DeVilbiss Healthcare sent End users a field corrective action letter dated March 7, 2013, to all affected customers to reinforce the proper cleaning method.This letter is to inform you that we have come across cases where end users had exposed the bacterial filter cartridge in the disposable suction container to fluid, thereby occluding the filter and rendering it unusableSpecifically what do you have to do when you receive a correction/recall notice?" The US FDA states in its correction/recall regulations that: "Consignees that receive a recall communication should immediately carry out the instructions set forth by the recalling firm and, where necessary, extend the recall to its consignees.""Enclosed you will find two Urgent Medical Device Correction letters addressed to DeVilbiss Dealer/Provider" and "DeVilbiss User" that we ask you read carefully."It is important that you follow the actions outlined in these letters to ensure we can effectively disseminate this information to end-users."To reach your direct accounts we ask that you send a copy of the "DeVilbiss User" letter to each of these end-users.We sincerely appreciate all of your efforts to ensure the safety of your customers that depend on DeVilbiss products. We thank you in advance for your assistance and apologize for any inconvenience that you may experience as a result of this situation. Should you have any questions or concerns, please contact DeVilbiss Healthcare Customer Service at 1-800-338-1988. |  | 112849 | 112,849 | Worldwide Distribution - USA and the countries of Mexico and Canada. | Open |  |  | 112849 | Not_Computer
Z-0005-2014 | 64454 | powered laser surgical instrument | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | The product is labeled in part "***Surgical Laser Technologies, Inc. Montgomeryville, PA***SLT Fiber Delivery System***Sterile***Non-Pyrogenic***. The product is packaged in a flexible pouch and includes the following laser surgery accessories - SLT-SSRH 8, Catalog/Product Number 0041-3082; SLT-SSRH 11, Catalog/Product Number 0041-3112; SLT-SSRH 4-SMA, Catalog/Product Number 0042-0075; SLT-SSRH 5-SMA, Catalog/Product Number 0042-0076; SLT-SSRH 6-SMA, Catalog/Product Number 0042-0077; SLT-SSRH 7-SMA, Catalog/Product Number 0042-0078; SLT-SSRH 8-SMA, Catalog/Product Number 0042-0079; SLT-SSRH 9-SMA, Catalog/Product Number 0041-4931; and SLT-SSRH 10-SMA, Catalog/Product Number 0042-0081.Indicated for the incision, excision, coagulations, and vaporization of soft tissue. | Surgical Laser Technologies, Inc. (SLT) SSRH | 2 | October 01, 2013 | 2013 | PhotoMedex, Inc. | Firm became aware that some products within the specified lot may not have been sealed appropriately, thus compromising the sterile barrier. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | The firm initiated their recall of these products beginning on January 4, 2013 by sending notification letters to their consignees via electronic mail and/or facsimile. Customers were asked to return the affected product for a replacement. |  | 549 | 549 | Worldwide Distribution-USA including the states of AR, KS and OH, and the countries of Germany, South Africa and the United Kingdom. | Open |  |  | 549 | Not_Computer
Z-0001-2014 | 66365 | cardiovascular surgical instruments tray (kit) | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Ultima OPCAB System, Sterile, Rx Only, Product Usage:Intended for use during performance of minimally invasive cardiac surgery through a sternotomy incision approach. The ULTIMA Access RailPlatform in combination with the Activator II Drive Mechanism is used to spread the sternum, providing access and direct visualization to the thoracic cavity. The AccessRail Platform also allows for the organization of pericardial sutures. The ULTIMA Stablizer isolates and provides local immobilization of the target vessel on the beating heart. | ULTIMA OPCAB System | 2 | October 01, 2013 | 2013 | Maquet Cardiovascular, LLC | MAQUET has identified a potential issue with the ULTIMA OPCAB (part numbers OM-2003S & OM-2001D). Labeling for the 2 blade lots had been incorrectly applied. The standard blades from Lot #25052176 were incorrectly labeled as deep blades, and the 20 deep blades from Lot #25052200 were incorrectly labeled as standard blades. | PRODUCTION CONTROLS: Error in Labeling | MAQUET sent an "Urgent Medical Device Recall" letter and "Field Recall Response Forms" dated September 20, 2013 to customers via Federal Express. The notification identifies the affected product; informs the customers of the nature of the issue; and advises of potential adverse patient consequences that may occur. The customer is asked to complete and return the Field Recall Response Form. If the customer has questions they are directed to contact their field representative or MAQUET Technical Support (888.627.8383 between 8 am and 5 pm EST Monday through Friday). |  | 20 | 20 units (Lot 25052176); 12 units (Lot #25052200) | USA Nationwide Distribution in the states of: Florida, Maryland, New Jersey, New York and Texas. | Open |  |  | 20 | Not_Computer
Z-0007-2014 | 66374 | stent, superficial femoral artery | NIP |  |  | Cardiovascular |  | Bard LifeStent Solo Vascular Stent, Item Numbers: EX062001CL, EX072001CL, EX062003CL, EX072003CL.The LifeStent Solo Vascular Stent System is intended to improve luminal diameter in the treatment of symptomatic de-novo or restenotic lesions up to 240 mm in length in the native superficial femoral artery (SFA) and in the proximal popliteal artery with reference vessel diameters ranging from 4.0  6.5 mm. | LifeStent Solo Vascular Stent | 1 | October 10, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling specific lots of Bard Lifestent Solo Vascular Stent because they have determined a higher incidence rate of partial stent deployment associated with their use. | DESIGN: Device Design | The firm, Bard Peripheral Vascular, sent a "Urgent: Medical Device Recall Notification" letter dated September 30, 2013 via FedEx to all the customers who received the specific lots on recall. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Recall and Effectiveness Check Form and fax to: BPV at 1-800-994-6772, even if you no longer have possession of the recalled product; call Bard Peripheral Vascular's Recall Coordinator at (800) 321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email silvia.carrillo@crbard.com. The Recall Coordinator will issue you either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product; and to return the product to Bard Peripheral Vascular, Inc.,1415 W. 3rd Street,Tempe, AZ 85281. BPV will provide replacement product for your returned product.If you have any questions, call 1-800-321-4254 Option #2 Ext. 2727. |  | 4112 | 4112 units | US Nationwide Distribution. | Terminated | February 05, 2014 | 118 | 4112 | Not_Computer
Z-0008-2014 | 66025 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-Mused to administer anesthesia and ventilation to patients during surgical procedures. | Draeger Fabius | 1 | October 10, 2013 | 2013 | Draeger Medical, Inc. | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch. | DESIGN: Device Design | The firm, Drager, sent an "Urgent Medical Device Recall" letter dated August 2013 to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that in the event of a failure, manual ventilation mode should be used and that Draeger reps would be contacting customers to schedule the replacement of the power supply. The customers were instructed to notify personnel accordingly.If you have any questions, contact Director, Quality at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your Draeger anesthesia machine please contact DragerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1127 | 1127 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Azerbaijan, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Columbia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Latvia, Libya, LithuaniaMalaysia, Morocco, Myanmar, Nepal, Philippines, Poland, Rumania, Russian Fed., Saudi Arabia, Senegal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, Vietnam, White RussiaTurkmenistan. | Open |  |  | 1127 | Not_Computer
Z-0011-2014 | 65898 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 2120/2120i Hematology Systems with ADVIA Autoslide connected to it, SMN 10361162 and SMN 10361798 --- CLASSIFICATION NAME: Counter, Differential Cell, Automated (Particle Counter)The ADVIA 2120/2120i are hematology systems that utilizes the principles of flow cytometry in order to provide complete blood counts. | ADVIA 2120/2120i Hematology Systems | 2 | October 17, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens determined that if an autosampler rack jam error occurs during operation of an ADVIA 2120/2120i connected to an ADVIA Autoslide, it is possible that the next slide processed by the Autoslide could be labeled with the wrong sample identification information. | DESIGN: Software Design | Siemens sent an Urgent Medical Device Correction Letter dated August 8, 2013, and Effectiveness Checks Forms were sent to customers on August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised if they do not have an ADVIA Autoslide connected to their system, no action is required. To avoid the possibility of mislabeling a slide when an ADVIA 2120/2120i rack jam error occurs: Eject the sample rack. Troubleshoot and correct the reason for the rack jam error. Rerun the sample rack. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions.If you need additional assistance, please contact Global Product SupportTarrytown, New York @ osggps.hschematology.healthcare@siemens.com. |  | 311 | 311 units total (US: 51 units; Foreign: 260 units) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Portugal, South Korea, Singapore, Slovenia, Spain, Switzerland, Thailand, Turkey, United Kingdom and Vietnam. | Open |  |  | 311 | Not_Computer
Z-0019-2014 | 66237 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-Slice CT Scanner System.Used as a whole-body computed tomography X-Ray system. | MX 16Slice CT Scanner System | 2 | October 17, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | A pinch point at the end of the table top can cause injury. |  | Field Safety Notice (FSN 72800601_72800603, will be sent to all US consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 326 | 326 | Worldwide Distribution-USA including the states of AL, AR, CA, FL, GA, IL, IN, KY, MS, NC, NJ, NV, NY, OH, PA, and VA; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, China, Ghana, Greece, India, Indonesia, Ireland, Italy, Jordan, Latvia, Lebanon, Libya, Lithuania, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, Uganda, Ukraine, Uzbekistan, Venezuela and Vietnam. | Open |  |  | 326 | Not_Computer
Z-0012-2014 | 66277 | system, x-ray, fluoroscopic, image-intensified | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D, Legacy, RFX, and SFX X. This table is intended for use in a diagnostic X-ray system to support patients during general purpose radiological procedures in the horizontal, vertical and trendelenburg positions. - The Expedio 500D is designed to perform radiographic and fluoroscopic x-ray examinations. - The Precision 500D R&F; X-ray System is intended to be used as a universal diagnostic imaging system for radiographic and fluoroscopic examinations, including general R&F; and pediatric examinations. The Precision 500D R&F; X-ray Systems are intended for use by a qualified/trained doctor or technician on both adult and pediatric subjects for taking diagnostic radiographic and fluoroscopic exposures of the whole body, skull, spinal column, chest, abdomen, extremities, and other body parts. Applications can be performed with the patient sitting, standing, or lying in the prone or supine position. | GE Healthcare, Precision 500D, Legacy, RFX, and SFX X. | 2 | October 17, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the spotfilm device power assist handle.There is a potential for the internal bolts on the power assist handle to be loose or to loosen and/or fall out. This issue, if present, can result in a loss of calibration and offset of the power assist handle and/or, under extreme condition, sticking of the handle in the ON p | TRAINING: Employee Error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to inspect your power assist handle for either of the two conditions listed below. Immediately contact your local GE Healthcare Service Representative if any issue is found.1. If you observe that the power assist handle plastic grip is loose relative tothe mount, discontinue use.2. If plastic grip is secure, but you observe erratic performance, (power assistunresponsive, drifting of spotfilm device with no command, or lock releasealways ON), you may continue to use the device with care. Use all safetyprecautions relative to application of the myelographic stop and the inhibit,table horizontal stop OFF.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 055 919For other countries, please contact your local GE Healthcare Service Representative. |  | 810 | 810 (783 US, 27 OUS) | Worldwide Distribution - USA (nationwided) including DC., and the countries of AUSTRALIA, CANADA, CHILE, GERMANY, CHANA, GUATEMALA, ICELAND, ISRAEL, ITALY, JAMAICA, JAPAN, RUSIAN FEDERATION, UNITED ARAB EMIRATES and UNITED KINGDOM. | Open |  |  | 810 | Not_Computer
Z-0014-2014 | 65955 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Midfoot Fusion Bolt 6.5 mm. It is indicated for fracture, fixation, osteotomies, nonunions, and fusions of large bones in the foot and ankle. | Synthes Midfoot Fusion Bolt | 2 | October 17, 2013 | 2013 | Synthes USA HQ, Inc. | A post market safety review following several customer complaints indicated additional clarification was necessary in the Midfoot Fusion Bolt Technique Guide to reduce potential for bolt back-out or migration by emphasizing the need for supplemental fixation. | DESIGN: Labeling Design | An "Urgent Notice: Medical Device Recall" letter dated March 8, 2013 was sent to customers asking them to check their inventory, remove all affected product technical guides from use, and return them to Synthes along with the Verification form. Customer questions were directed to (610) 719-5450. |  | 62 | 62 | Nationwide Distribution, including the states of MI, PA, SC, IA, MA, CA, WI, IL, and KS. | Open |  |  | 62 | Not_Computer
Z-0017-2014 | 64808 | turbidimetric, total protein | JGQ | Total protein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista (R) Urinary/Cerebrospinal Fluid Protein (UCFP) Flex (R) Reagent Cartridge Lot 11349BBThe Urinary / Cerebrospinal Fluid Protein method is an in vitro diagnostic test for the quantitive measurement of protein in urine and cerebrospinal fluid on the Dimension Vista System. | Dimension Vista (R) (UCFP) Flex (R) Reagent Cartridge catalog number K3026 | 2 | October 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed an increase on absorbance errors when calibrating Dimension Vista Urinary/Cerebrospinal Fluid Protein (U CFP) Flex Reagent Cartridge lot 11349 BB. |  | The firn, Siemens, sent an "Urgent Medical Device Correction" letter (dated December 2012) was provided to all Dimension Vista (R) UCFP customers who have received UCFP lot 11349BB. The letter described the product, problem and actions to be taken. The customers are instructed to discontinue use and discard any remaining inventory of Dimension Vista UCFP lot 11349BB; provide instructions for product replacement; and complete and return the attacted REQUEST FOR REPLACEMENT FORM via fax to (302) 631-8467. Siemens will replace any un-used inventory of this lot, at no charge.If you have any questions, please contact your Siemens Technical Solutions Center. your local Siemens Technical Support Representative or 800-441-9250. |  | 3395 | 3395 | Worldwide distribution: US (nationwide) including sates of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, ID, KY, KS, LA, MA, MD, MI, MN, MO, MS, MT, PA, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and PR; and country of: UK. | Open |  |  | 3395 | Not_Computer
Z-0020-2014 | 66047 | phosphomolybdate (colorimetric), inorganic phosphorus | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | Enforcement Discretion | Clinical Chemistry Phosphorus Reagent The Phosphorus assay is used for the quantitation of phosphorus in human serum, plasma, or urine. | Phosphorus Reagent | 2 | October 17, 2013 | 2013 | Abbott Laboratories, Inc | Reagent lots 64736UN12, 03991UN13, 51611UN13, 74805UN12, and 03992UN13 have incorrect expiration dates assigned by Supplier. | PRODUCTION CONTROLS: Error in Labeling | The firm, Abbott Laboratories, sent a "PRODUCT CORRECTION" letter dated August 12, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to update existing inventory and manually track the implanted lots with the corrected expiration dates listed in the letter; retain a copy of letter and if you have forwarded any of the lots listed to another laboratory, provide a copy of the letter to them.If you or any of your healthcare providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (422-2688). Customers outside the U.S., please contact your local area customer service. |  | 7213 | 7213 | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Austria, Australia, Azerbaijan, Bangladesh, Bahamas, Belgium, Bosnia & Herzegovina, Brazil, Bulgaria, Burkina, Canada, Cayman Islands, Chile, China, Columbia, Croatia, Curacao, Cyprus, Czech Republic, Denmark, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Mali, Mauritania ,Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Taiwan, Tanzania, Thailand, Trinidad & Tobago, Tunisia, Turkey, Turks and Caicos, United Arab Emirates, United Kingdom, Uruguay and Vietnam. | Open |  |  | 7213 | Not_Computer
Z-0022-2014 | 66268 | syringe, piston | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Harvest Graft Delivery System; Model Number: GDP-10Product Usage: The Harvest Graft Delivery System is intended for the aspiration of bone marrow, autologous blood, plasma, or other body fluids. The system is indicated for die delivery of allograft, autograft, or synthetic bone graft materials to an orthopedic surgical site. In addition, it is designed to facilitate pre-mixing of bone graft materials with the aspirate(s), I.V. fluids, blood, plasma, platelet rich plasma, bone marrow or other specific blood component(s) as deemed necessary by the clinical use requirements. | Graft Delivery System | 2 | October 17, 2013 | 2013 | Harvest Technologies Corporation | Potential for leur connector to leak or have cracks. | DESIGN: Device Design | Harvest Technologies sent a notification letter dated September 10, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form to confirm the number of kits received, used, and intending to return. They are asked to return the kits with the form. For questions contact Michelle Sullivan at 508-732-7500. |  | 730 | 730 cases (3 kits per case) | Worldwide Distribution - USA Nationwide: TX, NJ, MI, AZ, GA, LA, ME, MD, AR, MO, VA, MN, GA, MA, NY, PA, CO, CA and in the countries of Australia, Brazil, China, Singapore, Vietnam, Dominican Republic. | Terminated | February 10, 2014 | 116 | 730 | Not_Computer
Z-0018-2014 | 66000 | disinfector, medical devices | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 46-Series Medical Washer-Disinfectors, Listing # E208196Product Usage: The washer disinfector is used in healthcare facilities where reusable items such as instruments, utensils, anesthesia sets, and glassware are handled for decontamination purposes. | Getinge 46Series | 2 | October 17, 2013 | 2013 | Getinge Disinfection | Getinge Disinfection AB initiated a voluntary field safety correction for the Getinge 46-Series electrically-heated Washer Disinfectors due to possible overheating. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Getinge USA, Inc. sent a Device Correction Notices letter dated July 26, 2013 to all affected customers via US Postal Certified Mail within the U.S. and Canada. The letter identified the affected product, problem and actions to be taken. The letter indicate that all customers will be contacted by a Getinge representative for guidance if any machine within the serial number range of affected product is within their possession. If you should have any questions or concerns regarding this field correction, please call 1-800-475-9040, press "1" and enter extension 5036. | Contact the recalling firm for information | 143 | 143 units (142 domestically & 1 unit internationally) | USA Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and internationally in Canada. | Open |  |  | 143 | Not_Computer
Z-0015-2014 | 66294 | reagents, clostridium difficile toxin | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Product Name: OSOM C. Difficile Toxin A/B TestCatalog Numbers: 173 (for United States) and 173 E (for Europe).The Test is a qualitative assay (in vitro diagnostic) that employsimmunochromatographic, dipstick technology. If C. difficile toxin A or toxin B is present, a blue/gray line will appear in the test line region indicating a positive result. A red control line must appear for the results to be valid. If C. difficile toxins are not present, only the red control line will appear. An invalid test occurs when no control line appearsThe Test is intended for the qualitative detection of Clostridium difficile toxins A and/or B in human stool samples. This test is intended as an aid in the diagnosis of C.difficile associated disease (CDAD) in patients with symptoms of CDAD. | OSOM C. Difficile Toxin A/B Test | 2 | October 17, 2013 | 2013 | Sekisui Diagnostics LLC | Sekisuki Diagnostics has issued a Product Correction for certain lot numbers of the OSOM C. difficile Toxin A/B Test Kit, because it has been determined that there is potential for false negative results. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Sekisui Diagnostic sent an Customer Recall Notification letter dated September 16, 2013, via email and/or FedEx. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter informs the customers that Sekisui Diagnostics has issued a Product Correction for certain lots of the OSOM C. difficile Toxin A/B Test Kit (P/N 173). The customer notification lists the reason for recall, instructions to the customer, product description, required actions, and contact information. Customer Notification Letter lists the specific lot numbers being recalled for that catalog number. For each of these two catalog numbers, there is one letter directed to a distributor and one letter directed to an end user. Customer letters will be sent by email and/or FedEx. The letters have REQUIRED ACTIONS outlined in a box. If no response is received within 30 days, there will be directed follow-up communication. For further questions please call ( 858) 777-2600. |  | 391 | 391 kits | Worldwide Distribution - US (nationwide) and the countries of UK, Saudi Arabia, Dubai, Kuwait, Poland, Spain, India and Romania | Open |  |  | 391 | Not_Computer
Z-0027-2014 | 66346 | system, imaging, pulsed echo, ultrasonic | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | P21x/5-1 MHz transducer, Part Number P07698-23Product Usage: The P21x/5-1 MHz phased array transducer is designed for cardiac, abdominal and obstetric, TCD and orbital exams. It supports 2D with SonoMB Multi Beam Technology and Auto Gain capabilities, SonoHD Imaging Technology, Tissue Harmonic Imaging, M-Mode, Velocity Color Flow Doppler, Color Power Doppler (CDP), Pulsed Wave (PW), TDI Pulsed Wave, and Continuous Wave Doppler. | P21x/51 MHz transducer | 2 | October 18, 2013 | 2013 | SonoSite, Inc. | FUJIFILM SonoSite Inc. remove P21x/5-1 Transducer because the firm discovered the Biopsy Bracket Sensor is not working as intended. | PRODUCTION CONTROLS: Error in Labeling | SonoSite started calling their consignees on September 13, 2013. For customers who could not be reached via phone, a letter Urgent -Medical Device Removal will be sent on September 17, 2013. Affected customers will be offered a replacement transducer (P07698-24) free of charge. Consignees in the U.S. can call 1-877-657-8118, Monday-Friday, 6:00 a.m. - 5:00 p.m. (PST) or send E-mail to service@sonosite.com. Customers outside the US can call 425-951-1200. |  | 10 | 10 units | Worldwide Distribution - USA Nationwide in the states of Florida, Georgia, Oklahoma, Pennsylvania, Tennessee, Texas, West Virginia, and in the country Japan and Italy. | Terminated | December 11, 2013 | 54 | 10 | Not_Computer
Z-0029-2014 | 66105 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Model BP7507 OPTI LION E-Plus Cassettes 25 Cassettes with Sample Probes, OPTIMedicalCassettes are consumable used in the OPTI LION Electrolyte Analyzer intended to be used for the measurement of sodium, potassium, chloride, ionized calcium and pH in samples of whole blood, serum, and plasma. | Model BP7507 OPTI LION EPlus Cassette | 2 | October 18, 2013 | 2013 | OPTI Medical Systems, Inc | OPTI LION E-Plus Cassettes were labeled with an incorrect barcode. The incorrect barcode could perfect the performance of one or more parameters as the product ages. Also a Customer complaint from a distributor reported that a customer was getting QC failures. | PRODUCTION CONTROLS: Error in Labeling | The firm, OPTIMedical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2013 via email and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the product; return any unused product to OPTI Medical System; distributors outside the U.S. may destroy unused product and certify the instructions on the enclosed Recall Response Card; contact your distributor or OPTI Medical System Technical Support at 1-800-490-6784 or technicalsupport@optimedical.com to obtain authorization number and shipping instructions; and complete and return the Recall Response Card to technical support via fax to 1-770-510-4447; mail to: OPTI Medical Technical Support, OPTI Medical Systems, Inc., 235 Hembree Park Drive, Roswell, GA 30076 or email: technicalsupport@optimedical.com Please contact OPTI Medical Systems, Technical Support department at 1-800-490-6784 if you have questions related to this notice. |  | 361 | 361 boxes of 25 cassettes each | Worldwide distribution: US (nationwide) including states of: CA, MA, MN, and VA; and countries of: China, Guatemala, India, Iraq, South Korea, and Untied Arab Emirates. | Terminated | December 12, 2013 | 55 | 361 | Not_Computer
Z-0030-2014 | 66310 | tubes, vials, systems, serum separators, blood collection | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer 226 Sample Collection Device, Serial numbers: 06132903 - 06198900.The PerkinElmer 226 Sample Collection Devices are intended to be used as a medium to collect and transport whole blood specimen spots to a laboratory, in newborn screening. The device includes a tear-apart form for the collection of demographic information. | PerkinElmer 226 Sample Collection Device | 2 | October 18, 2013 | 2013 | Perkin Elmer Health Sciences, Inc. | Some devices may be contaminated on the back side of the device with glue in the blood application area. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Perkin Elmer, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated September 5, 2013 via email to its consignees. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately notify all submitting facilities and personnel that have received the PerkinElmer 226 Sample Collection Device of the potential glue contamination and to distribute the enclosed record of actions form to all submitting facilities/personnel and encourage them to perform the inspection activities as listed in the letter; and complete and return the enclosed Recall Response Form return it to PerkinElmer Health Sciences via fax to: 864-299-8797 or email to: Rick.Haines@perkinelmer.com.Please contact PerkinElmer Health Sciences Greenville facility Quality Manager for further information 864-299-8787 (ext 104). |  | 65984 | 65,984 devices | US distribution in OH | Open |  |  | 65984 | Not_Computer
Z-0028-2014 | 66325 | system, blood culturing | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BBL (tm) MGIT (tm) Mycobacteria Growth Indicator Tube, part # 245122, packaged in 100 ct cartons and labeled in part Product Usage: The Mycobacteria Growth Indicator Tube supplemented with Bactec MGIT Growth Supplement and BBL MGIT PANTA antibiotic mixture is intended for the detection and recovery of mycobacteria using the BACTEC MGIT 960 and BACTEC MGIT 320 Systems. Acceptable specimen types are digested and decontaminated clinical specimens (except urine), and sterile body fluids (except blood). | Mycobacteria Growth Indicator Tube | 2 | October 19, 2013 | 2013 | Becton Dickinson & Co. | An in-vitro diagnostic reagent may be labeled with an incorrect barcode, which when scanned on associated diagnostic instrument may cause a false negative result. | PRODUCTION CONTROLS: Error in Labeling | BD sent an Urgent Product Recall letter dated September 2013 to customers identified as microbiology lab managers. The letter identified the affected product, problem and actions to be taken. The letter advised that referenced lot numbers may contain an overlabel barcode that may begin to lift, exposing the original incorrect barcode pictured in the notification. Customers were instructed to complete the attached form and fax to Regulatory Compliance 410-316-4258. For further assistance regarding replacements contact BD Customer Service Department at 1-800-675-0908 and all other inquiries contact BD Technical Services Department at 1-800-638-8663. |  | 187000 | 187,000 tubes | USA Nationwide Disatribution in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, MA, ME, MN, MO, NC, ND, NJ, NV, NY, OH, PA, TX, UT, VA, WA, WI | Terminated | October 29, 2013 | 10 | 187000 | Not_Computer
Z-0031-2014 | 65952 | retractor | GAD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | UNITRAC Pneumatic Retraction Arm.A support arm used to hold surgical instruments during surgical procedures. | UNITRAC Pneumatic Retraction Arm | 2 | October 21, 2013 | 2013 | Aesculap, Inc. | Changes were implemented to the UNITRAC instructions for use to ensure safe use and potentially extend the life of the product. This included recommending the use of Sterile Cover JG901 and new Quick Release Coupling RT02OR to eliminate the need to sterilize the UNITRAC and a functional check was added prior to use to ensure optimal functionality of the UNITRAC system. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | A letter dated July 16, 2013 was sent requesting users to review the new instructions for use in their entirety, sign the attached Hospital Acknowledgment form and return to Aesculap. An Aesculap representative would be in contact to provide an in-service regarding the process for set-up of the Sterile Cover and required Functional Check. Discontinuation of use and repair of the device was recommended if the joints did not move freely or if the set position drifts during the Functional Check. | Contact the recalling firm for information | 217 | 217 | Nationwide Distribution. | Open |  |  | 217 | Not_Computer
Z-0051-2014 | 65969 | cannula, ophthalmic | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Synergetics 25 gauge Soft Tip Cannula, packaged in 6 units per box. Catalog number: 30.08.25.Single-use 25 gauge cannula used in vitreoretinal surgery. | Synergetics 25 gauge Soft Tip Cannula | 2 | October 22, 2013 | 2013 | Synergetics Inc | One lot of 25 gauge Soft-Tip Disposable Cannulas are co-mingled with 23 gauge cannulas; the product is labeled as 25 gauge cannulas. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm sent recall notification letters to their consignees on 7/30/13. |  | 54 | 48/6-cannula boxes | Nationwide Distribution-USA including the states of MO, CO, KY, TN, NC, PA, IL CA, NY, IN, MI, LA, NJ and OH. | Open |  |  | 54 | Not_Computer
Z-0044-2014 | 66317 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX CassettesThe STERRAD Sterilization System is a low temperature general purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | STERRAD 100NX Cassettes | 2 | October 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the three lots (13F049, 13F050, 13F051) STERRAD 100NX Cassettes because ASP has determined that the leak indicator on the cassettes is effective for only 12 months of the cassette's labeled 15-month shelf-life. |  | Advanced Sterilization Products (ASP) sent an Urgent Medical Device Recall letter on September 18, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to examine their inventory and return all affected cassettes using the enclosed prepaid UPS return lave to Stericycle. Customers were asked to count their inventory and record the data on the enclosed Business Reply Card and Packing Slip. Customers were asked to provide the notice to anyone in their facility that needs to be informed. Customers were also asked to maintain a copy of the notice with the affected product. Customers with questions or suspected problems were instructed to contact Stericycle directly at 1-877-257-7119.For questions regarding this recall call 949-453-6400. |  | 2995 | 2,995 units | Nationwide Distribution | Terminated | January 13, 2014 | 83 | 2995 | Not_Computer
Z-0043-2014 | 66328 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer (Single Door), Part Number: 10201.STERRAD 200 Sterilizer (Double Door), Part Number: 10202.The STERRAD Sterilization System is a low temperature general purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | The STERRAD Sterilization System | 2 | October 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 Sterilization System because it may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. |  | A customer notification letter was sent on September 18, 2013, to all customers who received the STERRAD 200 Sterilization System. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to indicate if they have received a STERRAD 200 System and to report any complaints or suspected problems to Advanced Sterilization Products (ASP) Professional Services directly at (888) 783-7723 and select option #2.An alternate design for the oil mist filter is being developed and qualification of new vacuum pump oil is in process. This action is to modify the PM interval to service the filter and oil more frequently. This will serve as an immediate/interim solution to reduce incidence of odor/smell complaints until the new oil mist filter design and vacuum pump oil are released.A Customer Letter will be sent to all STERRAD 200 customers to notify them of the issue and the corrective actions we are taking to resolve this issue. |  | 651 | 1,305 units total (654 units in US , 651 OUS ) | Worldwide Distribution - USA (nationwide) and the countries of EA (Europe Middle East and Asia), APAC (Asia Pacific), Japan and Canada | Open |  |  | 1305 | Not_Computer
Z-0045-2014 | 66400 | catheter, intravascular, diagnostic | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Performa Angiographic Catheter 5F Multipack. Part numbers 7560-43 and 7560-G3. | Performa Cardiac Multipack Catheter | 2 | October 22, 2013 | 2013 | Merit Medical Systems, Inc. | 4 lots of Performa Cardiac Multipacks contain a JR4.0 catheter with a slight variation in tip shape. | PRODUCTION CONTROLS: Process Control | Consignees were notified on 9/13/13 via Urgent Product Recall Notice, email, and product retrieval forms. Actions required of customers:1. Please immediately determine if any of the devices are within your facility, quarantinethem and discontinue use.2. Ensure that all personnel to whom the devices were distributed are made aware of this field action.3. Contact your Merit Sales Representative in order to arrange product return and replacement or credit.4. If you have any questions concerning this communication, please don't hesitate to contact Guido Sandulli at 1-801-826-4046. |  | 200 | 200 | US Distribution including the states of CO, VA, and MS. | Terminated | January 03, 2014 | 73 | 200 | Not_Computer
Z-0054-2014 | 65742 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Brain Tissue Oxygen Monitoring; 1 x Complete Brain Probe KitREF IP1.PRx Only, Manufacturer GMS- Gesellschaft fur medizinische, Sondentechnik mbH Dofstrasse 2 24247 Melkendorf, Germany Distributed by Integra Neurosciences 311 Enterprise Drive Plainsboro, NJ 08536 USA.Used in direct monitoring of the partial pressure of oxygen. | Integra Licox Brain Tissue Oxygen Monitoring Kit | 2 | October 23, 2013 | 2013 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has determined that there is a possibility that some lots of Licox CMP Brain Tissue Oxygen and Temperature Catheter Kits contain a hex wrench that does not loosen the set screw on the drill bit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Integra LifeSciences Corporation sent a recall letter/return response form on 6/11/2013 by traceable courier service, traceable e-mail or facsimile. |  | 21 | 21 kits | Worldwide distribution: US (nationwide) and countries of: AR, BR, CH, CL,CR. DE, ES, FI, FR, GB, IL, IT, MX, PA, PL, PT, RU, TW, and ZA. | Terminated | February 18, 2014 | 118 | 1663 | Not_Computer
Z-0064-2014 | 66248 | bone cement, antibiotic | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SmartSetÂŽ Endurance GMV, Gentamicin, REF 3105-040, Medium Viscosity Antibiotic Bone Cement, Sterile.SmartSet GMV Endurance Gentamicin, bone cements are self-curing, radiopaque, polymethyl methacrylate based cements, containing antibiotic, used for securing a metal or polymeric prosthesis to living bone in arthroplasty procedures where infection by gentamicin-sensitive organisms is a potential risk. | SmartSet GMV Endurance Gentamicin | 2 | October 23, 2013 | 2013 | DePuy Orthopaedics, Inc. | The SmartSetÂż GMV Endurance Gentamicin Bone Cement, product ref: number 3105-040 is intended for use in indications that have not been cleared by the US FDA. The intended use stated in the IFU for the SmartSetÂż GMV Endurance Gentamicin Bone Cement is for the fixation of prostheses to living bone in arthroplasty procedures of joints in which infection by gentamicin-sensitive organisms is a potentia |  | The firm DePuy Synthes, sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated September 16, 2013 to its consignees via E-mail, hand delivery or fax depending on the most appropriate delivery method for each consignees. The Urgent recall notice described the recalled product, explained the reason for the recall, and directed consignees to cease further distribution, complete the attached Reconciliation form and quarantine and return all affected products through the DePuy Orthopaedics Sales consultant, Fax 1-787-287-7681 or Email to vbaez1@its.jnj.com. Consignees were directed to maintain a copy of the notice and forward the notice if product was further distributed. Clinical questions should be addressed to DePuy Orthopaedics, Inc's Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). Recall questions or concerns should be addressed to the Quality & Compliance Supervisor, at 1-787- 272-1900 x3114 or vbaez1@its.jnj.com (MF; 8 am- 5 pm EST). | Contact the recalling firm for information | 16040 | 16,040 units | Distributed in Puerto Rico and US Virgin Islands. | Open |  |  | 16040 | Not_Computer
Z-0053-2014 | 66394 | drug mixture control materials | DIF | Clinical toxicology control material. | Toxicology | Toxicology | 510(k) | Alere TriageÂŽ TOX Drug Screen Control 1Product Generic Name: Quality Control MaterialDescription of the product:Good Laboratory Practice suggests that external controls should be tested with each new lot or shipmentof test materials, or every 30 days, and as otherwise required by a laboratorys standard quality controlprocedures. The Alere TriageÂŽ Drug Screen Controls, when used in accordance with the test proceduresdescribed in the Alere TriageÂŽ TOX Drug Screen package insert, provide this capability. The resultsobtained for the controls are to be compared with the assigned values given on the Expected Values Card,accompanying the package insert, to determine if the procedure is within control limits.Alere TriageÂŽ TOX Drug Screen Control 1 or Alere TriageÂŽ TOX Drug Screen Control 2 consist ofhuman urine matrix containing parent compounds and/or major metabolites/analogues of drugs of abuse.510(k)/IDE/PMA number: K060788Type of Packaging: 5 vials per level in a labeled kit box.Indications for use: The Alere TriageÂŽ TOX Drug Screen Controls are to be used with the Alere TriageÂŽ TOX Drug Screen tests and Alere TriageÂŽ Meters to assist the laboratory in monitoring test performance. | Quality Control Material | 3 | October 23, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego, Inc. initiated a voluntary recall of Alere Triage TOX Drug Screen Control 1 (PN 94413, Lot C2681A) because the CODE CHIP that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere Triage TOX Drug Screeen products. The Alere Triage TOX Drug Screen test devices and patient results are not affected. |  | Alere San Diego, Inc. sent an Urgent Medical Device Notice dated September 23, 2013, to all affected customers via FedEx and outside US (OUS) consignees on September 24, 2013, via email. The customer notification letter titled "URGENT MEDICAL DEVICE NOTICE", informed customers of the recall, product description with part number and lot code (Alere TriageÂŽ TOX Drug Screen Control 1 PN 94413 Lot C2681A), reason for recall and impact (The purpose of the letter was to inform customers that the CODE CHIP" that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere TriageÂŽ TOX Drug Screen products. Customers were instructed to discontinue use of and discard the CODE CHIP contained within the Alere Triage TOX Drug Screen Control 1 PN 94413 Lot C2681A kit box. Customers were also instructed to provide a copy of the notice to their customers and instruct them to contact Alere for a replacement code chip. Customers were asked to complete and fax the enclosed Verification Form to confirm their receipt of the notice. Customers with questions were instructed to contact:(Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121, U.S.A.Phone: 877 308 8287FAX: 858 805 8457E-mail: Responses.ts@alere.com).For questions regarding this recall call 858-805-2000, ext 3015. | Contact the recalling firm for information | 38 | 38 | Worldwide Distribution - USA (nationwide) and countries of: South Africa and Canada. | Open |  |  | 38 | Not_Computer
Z-0065-2014 | 66121 | system, applicator, radionuclide, remote-controlled | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VARISOURCE Breast Template System for high-dose rate Brachytherapy. ***INSTRUCTIONS FOR USE*** Breast Bridge & Template System AL13011000 ***Manufacturer: Varian Medical Systems, Inc., 3100 Hansen Way, Palo Alto, CA 94304 USA |  | 2 | October 24, 2013 | 2013 | Varian Medical Systems, Inc. | The recall was initiated after Varian Medical became aware the Lexan templates of the Breast Bridge Template System could become bent and turn opaque after steam sterilization. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Urgent Medical Device Correction, Urgent Field Safety Notice were mailed by FedEx tracked delivery to all consignees on August 12, 2013. |  | 3 | 3 sets | Worldwide distribution: US (nationwide) distribution in CA and in the country of India. | Open |  |  | 3 | Not_Computer
Z-0066-2014 | 66341 | image-intensified fluoroscopic x-ray system, mobile | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Brivo OEC715/785/865 C-Arm, Models: 715 Prime, 785 Essential, and 865 Plus.The QEC Brivo Mobile C-Arm X-Ray Products are designed to provide digital spot film imaging and fluoroscopic image guidance for all adult and pediatric populations for orientations between patient anatomy and surgical instruments. The product is used for general surgical applications and musculoskeletal procedures to visualize, for example, implant localizations or needle positions for aspirations,injections or biopsy. | GE Brivo OEC | 2 | October 24, 2013 | 2013 | GE Healthcare Beijing , Hangwei Medical Systems, Co. LTD | GE is recalling certain lots of GE Brivo OEC715/785/865 C-Arm due to the potential for the wires in the high voltage cable bundle to fail, resulting in loss of system functionality. | DESIGN: Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated September 18,2013, to all affected customers to inform them of the issue and provide instructions for immediate mitigation of the issue. Customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions. Customers with questions were instructed to call 1-800-874-7378 option 8.For questions regarding this recall call 801-295-4952. |  | 29 | 29 | Worldwide Distribution - USA (nationwide) and Internationally to China, India, Japan, Malaysia, Australia, Ireland, France, Germany, Spain, Hungary, Belgium, and Switzerland. | Terminated | January 27, 2014 | 95 | 29 | Not_Computer
Z-0083-2014 | 66322 | system, endovascular graft, aortic aneurysm treatment | MIH |  |  | Cardiovascular | PMA | Valiant Thoracic Stent Graft with Captivia Delivery System,Talent Thoracic Stent Graft with the Captivia Delivery System. | Valiant Thoracic Stent Graft with Captivia Delivery System, Talent Thoracic Stent Graft with the Cap | 2 | October 25, 2013 | 2013 | Medtronic Vascular, Inc. | Firm received reports of difficulty in flushing the graft cover of the delivery system during device preparation as required by IFU. | DESIGN: Component Design/Selection | A Medical Device Information letter was sent to all affected customers on September 19, 2012. Additionally, all Medtronic Field Representatives were notified of the issue via a Field Memo. |  | 1000 | ~1000 | International distribution:Austria, Belgium, Cyprus, Czech Republic, Denmark, Finland, Germany, Greece, Hungary, Iceland, Ireland, It lay, Latvia, Lithuania, Luxembourg, Netherlands, Norway, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | October 25, 2013 | 0 | 1000 | Not_Computer
Z-0086-2014 | 65918 | automated urinalysis system | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HemoCueÂŽ Urine Albumin Microcuvettes, 25 Microcuvettes HemoCue AB Angelholm, SwedenThe quantitative, rapid, turbidimetric immunoassay of albumin in human urine using a specially designed analyzer, the HemoCue Albumin 201 Analyzer. The system can be used for the quantitative determination of low levels of albumin in urine for the purpose of screening for, diagnosing, monitoring and to supplement the clinical evidence in the treatment of microalbuminuria. The system is designed for testing using spot samples or timed collections. A quantitative result is obtained within 90 seconds. HemoCue Urine Albumin Microcuvettes are for in vitro diagnostic use only. The HemoCue Albumin 201 Analyzer is only to be used with HemoCue Urine Albumin Microcuvettes | HemoCue Urine Albumin Microcuvettes | 2 | October 25, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCueÂż Urine Albumin Microcuvettes single pack pouches. | PRODUCTION CONTROLS: Packaging | HemoCue sent an Urgent Field Safety Notice dated July 24, 2013. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to :1. Quarantine any HemoCueÂŽ Urine Albumin (110608); lot numbers 1208848 to 1304847 you may have in your possession2. Return the Field Safety Notice Verification Form to: fax: 562-668-5794We apologize for the inconvenience this advisory has caused you. Please feel free to contact Technical Support at 800-426-7256 with any questions. |  | 10406 | 10,356 boxes/50 microcuvettes | Worldwide Distribution - Nationwide Distribution including the states of AL, AR, AZ, CA, CA ,CT, FL, GA, HE, ID, IL, IN, LA, MA, ME, MI, MN, MO, MS, NC, NH, NM, NY, OH, PA, SC,TX, UT, VA, VT, WA and WI., and the countries of Australia, Bahrain, Belarus, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Indonesia, Italy, Kenya, Lesotho, Malawi, Mauritius, Mexico, Myanmar, Nigeria, Norway, Omen, Poland, Qatar, Romania, South Africa, South Africa, South Korea, Spain, Sweden, Taiwan, Tanzania, Ukraine, Yemen and Zambia. | Open |  |  | 10406 | Not_Computer
Z-0081-2014 | 66483 | filter, intravascular, cardiovascular | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Vena Cava FilterIs indicated for use in the prevention of recurrent pulmonary embolism (PE) via percutaneous placement in the vena cava in the following situations: - Pulmonary thromboembolism when anticoagulants are contraindicated, - Failure of anticoagulant therapy for thromboembolitic disease, - Emergency treatment following massive pulmonary embolism where anticipate benefits of conventional therapy are reduce, and - Chronic, recurrent pulmonary embolism where anticoagulant therapy has failed, or is contraindicated. | Cordis OPTEASE Vena Cava Filter | 2 | October 25, 2013 | 2013 | Cordis Corporation | Cordis identified a printing error on one unit of the Optease Vena Cava Filter, in which the orientation arrow for the femoral approach was printed in the incorrect direction. The error resulted in the filter being implanted upside down, requiring and additional percutaneous procedure to retrieve the filter. | PRODUCTION CONTROLS: Error in Labeling | Cordis sent an Urgent Medical Device Recall letter dated October 8, 2013, to all affected customers. The letter identified the product (Cordis OPTEASEÂŽ Vena Cava Filter Catalog Numbers 466F220A and 466F220B), the problem, and the action to be taken by the customer. Customers were instructed to:1) Read this Urgent Medical Device Recall letter.2) Immediately identify and set aside all product listed below in a manner thatensures the affected product will not be used.3) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form.4) Return any affected product with reference to the RA# on the form. ContactCustomer Service (see Available Assistance section) if assistance is needed for return of the affected product. Credit will be provided or consignment levels reduced.5) Share this letter with others in your facility that need to be made aware of thisrecall.6) Contact any other facility to arrange the return of OPTEASEÂŽ if any productlisted below has been forwarded to them.7) Maintain awareness of this notice until all affected product has been returnedto Cordis.8) Keep a copy of this notice with the affected product.Please contact the special Customer Service number (800) 551-7683,Monday through Friday from 8:30 AM to 6:30 PM ET.o Return instructions are provided on the Acknowledgment Form.Questions related to the Acknowledgement Form" Please contact Cordis QA at (786) 313-8730.For questions regarding this recall call 1-800-551-7683. |  | 29395 | 29,395 units | Worldwide Distribution - US (nationwide) including Hawaii and Puerto Rico and states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Canada and Panama. | Open |  |  | 29395 | Not_Computer
Z-0074-2014 | 65799 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂŽ Femoral Component (CR-FLEX GSF POR FEM, C-L; CR-FLEX GSF POR FEM, C-R; CR-FLEX GSF POR FEM, C-L MINUS; CR-FLEX GSF POR FEM, C-R MINUS; CR-FLEX GSF POR FEM, D-L; CR-FLEX GSF POR FEM, D-R CR-FLEX GSF POR FEM, D-L MINUS; CR-FLEX GSF POR FEM, D-R MINUS CR-FLEX GSF POR FEM, E-L; CR-FLEX GSF POR FEM, E-R; CR-FLEX GSF POR FEM, E-L MINUS; CR-FLEX GSF POR FEM, E-R MINUS; CR-FLEX GSF POR FEM, F-L ; CR-FLEX GSF POR FEM, F-R; CR-FLEX GSF POR FEM, F-L MINUS; CR-FLEX GSF POR FEM, F-R MINUS; CR-FLEX GSF POR FEM, G-L CR-FLEX GSF POR FEM, G-R; CR-FLEX GSF POR FEM, G-L MINUS CR-FLEX GSF POR FEM, G-R MINUS; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE C, LEFT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE C, RIGHT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE D, LEFT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE D, RIGHT; LPS-FLEX GSF POR FEM, E-L; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE E, RIGHT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE F, LEFT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE F, RIGHT; NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE G, LEFT; and NEXGEN GSF POROUS FEMORAL LPS FLEX SIZE G, RIGHT).Provides increased flexion capability. | NexGen Femoral Component | 2 | October 25, 2013 | 2013 | Zimmer, Inc. | The low density polyethylene (LDPE) bag used to package implants adheres to the highly polished surface of the devices. | DESIGN: Packaging Design/Selection | August 30, 2013 Urgent Recall Notices were sent to Distributors, hospital risk managers and surgeons identifying the issue and directing consignees to make all affected personnel aware of the contents of the notification. The notice identified the products involved, the issue, possible adverse events associated with the issue and recommended that implants found with LDPE bag adhesion not use the product and return it to Zimmer sales representatives. Questions or concerns should be addressed to Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. |  | 11694 | 11694 | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, India, Israel, Jamaica, Japan, Korea, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Taiwan, Thailand, and Venezula. | Open |  |  | 204220 | Not_Computer
Z-0042-2014 | 66433 | instrument, biopsy | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Brand: The ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT), MASCT System Kits, Clarity System Kits, MASCT Patient Sample Kits & Clarity Patient Sample Kits Instrument, Biopsy.The MASCT System Kit, Patient Sample Kits, or the Nipple Aspirate Fluid Laboratory Kit has the following codes:PRODUCT CODE: AG-MASCT; AG-FC5; DTG-MASCT, DTG-FC5, and NRLBH-5.PART NUMBER: 9002513; 9002513MD; 9002528; 9002528MD; 9002587; 9002614; and 9002717MD.THE MASCT SYSTEM KIT CONSISTS OF:1. MASCT Breast Pump;2. Instructions for Use (IFU);3. Heating Pad;4. Timer;5. Saccomono's Fixative;6. Nu Prep Gel;7. Welcome and Training Materials Coversheet;8. MASCT System Order Form;9. ForeCYTE Training Video DVD;10. MASCT System Instructions with pictures - pink;THE PATIENT SAMPLE KIT CONSISTS OF:1. Two (2) flower assemblies (e.g., filter, filter retainer, and filter holder);2. Instructions for Use (IFU);3. Two (2) specimen collection devices;4. Two (2) breast bags for transporting the specimen collection container;5. Barcode labels;6. Saccomono's Fixative;7. Nu Prep Gel;8. MASCT System package insert;9. ForeCYTE Test Requisition Form;10. Patient Information Sheet;11. Atossa Patient Information;12. FED EX Clinical PAK;13. MASCT System Order Form;14. FED EX Pouch;15. FED EX Airbill;Product Usage: The MASCT Device is intended for use in the collection of nipple aspirate fluid for cytological testing. The collected fluid can be used in the determination and/or differentiation of normal versus premalignant versus malignant cells. | ForeCYTE Breast Health Test and the Mammary Aspiration Specimen Cytology Test (MASCT) | 1 | October 28, 2013 | 2013 | Atossa Genetics, Inc. | Atossa Genetics is recalling ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT) due to concerns raised by the U.S. Food and Drug Administration (FDA) in a warning letter received by Atossa in February 2013. The FDA raised concerns about (1) the current Instructions for Use (IFU); (2) certain promotional claims used to market these devices; and (3) the need for FDA cl | PREMARKET APPROVAL: No Marketing Application | Atossa Genetics Inc. issued a press release to announce a voluntary recall on October 4, 2013.The firm started sending out the URGENT - VOLUNTARY PRODUCT RECALL NOTIFICATION ForeCYTE Breast Health Test; Mammary Aspiration Specimen Cytology Test (MASCT) RESPONSE REQUIRED letter to consignees on October 7, 2013. The letter identified the affected product, problem and actions to be taken. The firm initiated a voluntary recall to remove the ForeCYTE Breast Health Test and the Mammary Aspiration Specimen Cytology Test (MASCT) device from the market. The MASCT device has not been cleared by the FDA for the screening or diagnosis of breast cancer. In addition, the ForeCYTE Breast Health Test has not been cleared or approved by the FDA for any indication. The ForeCYTE Breast Health Test and the MASCT device are not a replacement for screening mammograms, diagnostic imaging tests, or biopsies. Patients should follow the recommendations and instructions of their physician with respect to breast cancer screening and diagnosis.Distributors are instructed to notify downstream customers and physicians are instructed to notify patients who have been tested using the MASCT and/or ForeCYTE systems. Consignees should return the Response Form back to Atossa.Distributors and customers should call 800-219-4629 to obtain a Returned Materials Authorization (RMA) number to return affected products. Customers with questions about this voluntary recall or Atossa's products, or would like assistance with the recall documentation may contact Atossa at 1-888-219-4629.Any problems experienced with the use of this product may be reported to the FDA at http://www.fda.gov/Safety/MedWatch/HowToReport/default.htm,Online (form available to fax or mail), or call FDA 1-800-FDA-1088. | Contact the recalling firm for information | 577 | 577 MASCT System kits, 101 Lab Kits, and 2557 Patients Kits (5 test per kit) | USA Nationwide Distribution in the states of: AL, CA, FL, GA, IL, IN, KY, MD, ME, MI, MO, NC, ND, NJ, NV, NY, OK, PA, SC, TN, TX, UT, VA, WA, Washington DC, and Puerto Rico. | Open |  |  | 577 | Not_Computer
Z-0087-2014 | 66290 | [6 empty] | chromID Salmonella Agar, Reference No. 43621.A selective isolation and differentiation medium for the detection of Salmonella in human specimens. | chromID Salmonella Agar | 2 | October 28, 2013 | 2013 | bioMerieux, Inc. | Complaints for Salmonella strains growing on this medium producing an uncharacteristic white colored colony instead of the expected pale pink to mauve colored colonies. This could potentially lead to Salmonella colonies not being detected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Recall letters were issued to the US customers by Biomeriux, Durham, NC on 9/6/2013. |  | 90 | 90 | Nationwide Distribution. | Open |  |  | 90 | Not_Computer
Z-0092-2014 | 66352 | driver, wire, and bone drill, manual | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | REF 5100-088, TPS Micro Driver, For Use with Stryker Total Performance System Caution Federal Law (USA) restricts the device to use by or on the order of a physicianDriver, Wire and Bone Drill; The Stryker Total Performance System (TPS) is intended for use in the cutting, drilling, reaming, decorticating and smoothing of bone and other bone related tissues for a variety of applications such as ENT, dental, orthopedic, maxillofacial, spinal, and plastic surgery. The instruments are also used in the placement of screws, wires, pins, and other fixation devices. | TPS Micro Driver | 2 | October 28, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | It was discovered that the silicone potting failed to properly cure after undergoing the prescribed curing possess. The potting material is used to insulate, stabilize and provide a moisture barrier to the TPS MicroDriver main Printed Circuit Board Assembly (PCBA). Inadequate cure of the silicone potting could potentially allow it to leak to the outside of the handpiece or cause an electrical ma | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Stryker sent an Urgent Medical Device Correction Notification letter via overnight mail and interntion consignees via E-mail on October 2, 2013. The Urgent Notice identified the recalled product, the reason for the recall, potential adverse risks to health associated with the recall. Customers were instructed to:1. Immediately review this Recall Notification2. Locate, and quarantine, the units listed in this notification. Do not use the recalled TPS MicroDriver until it has been repaired. A list, of units shipped to your facility, is included on the attached Business Reply Form.3. Return the enclosed Business Reply Form to confirm receipt of this notification and identify how many affected units are currently in your inventory.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, Stryker will send you a pre-paid shipper to send your affected device(s) back to Stryker to be repaired. Unit will be repaired and returned to your facility7. Send back all affected devices using the pre-paid shipper provided to you by Stryker.8. If you require a loaner, please contact the Stryker Service Department at 888-308-1983.For questions regarding this recall call 269-389-4518. |  | 67 | 67 | Worldwide distribution: US (nationwide) and countries of: Australia, Canada, France, Germany, Italy, Japan, Spain, Switzerland, and United Kingdom. | Open |  |  | 67 | Not_Computer
Z-0088-2014 | 66434 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | da Vinci Si Surgical System IS3000 with the da Vinci Surgeon Console which contains a high Resolution Stereoscopic Viewer.Intended to assist in the accurate control of Intuitive Surgical Endoscopic instruments. | da Vinci Si Surgical System IS3000 | 2 | October 28, 2013 | 2013 | Intuitive Surgical, Inc. | Improper restraints during transportation of the da Vinci system could cause the HRSV Monitor to become loose from its mount. |  | Advisory letters were sent to all affected customers on 10/17/11 informing them to contact Customer Service prior to moving the system outside of the hospital. |  | 1090 | 1,090 consoles | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Belgium, Canada, Chile, China, Denmark, Egypt, Finland, France, Greece, Germany, India, Indonesia, Israel, Ireland, Italy, Kuwait, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | October 28, 2013 | 0 | 1090 | Not_Computer
Z-0084-2014 | 66379 | insufflator, automatic carbon-dioxide for endoscope | FCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PENTAX Endoscopic Accessories. OF-B194. Made in Japan.PENTAX Upper/Lower G.I. Endoscopes are provided with Air/Water Valve. As an alternative, the optionally available Gas/Water Valve may be used in place of the standard Air/Water Valve to deliver a non-explosive gas for insufflation. | PENTAX Endoscopic Accessories | 1 | October 28, 2013 | 2013 | Pentax Medical Company | PENTAX Medical has become aware of a product issue associated with the OF-B194 Gas/Water Feeding Valve. A manufacturing defect may cause improper stoppage of CO2 gas during an endoscopic procedure with the potential for serious hazard to and/or user. |  | PENTAX Medical sent an "Urgent Device Recall" notification letter with /Response forms dated September 16, 2013, via UPS Certified delivery to all affected customers. . The notification identified the affected product; informed the customers of the nature of the issue; advised of potential adverse patient consequences that may occur. The customer was asked to immediately examine their inventory and quarantine product subject to the recall. The customer was instructed to identify and notify their customers if they further distributed the product. The notification instructs users to return the OF-B194 Valves for replacement with corrected product. Customers were asked to complete the attached response form and fax back to Pentax Medical at 1.201.799.4063. An enclosed UPS Airway Bill is to be used to return recalled product. A contact phone number and email address is provided if the customer has questions (1.800.431.5880 xt2064).For questions regarding this recall call 800-431-5880.Pentax issued a press release on 10/31/2013.Pentax posted their press release on their webisite at www.pentax.com. | Contact the recalling firm for information | 76 | 76 units | Nationwide Distribution including Arizona, California, Florida, Idaho, Indiana, Kansas, Maryland, Missouri, Missippi, North Carolina, New York, Pennsylvania, Virginia, Wisconsin and Wyoming. | Open |  |  | 76 | Not_Computer
Z-0101-2014 | 65920 | system, hemoglobin, automated | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HemoCueÂŽ Hb 201 Microcuvettes, carton/4 boxes/25 pouches of Microcuvettes, HemoCue AB Angelholm, Sweden | HemoCue Hb 201 Microcuvettes | 2 | October 29, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCue Hb 201 Microcuvettes single pack pouches. | PRODUCTION CONTROLS: Packaging | On 8/13/2013 the firm sent Field Safety Notices to their customers. |  | 42359 | 42,330 packages/4 boxes/25 microcuvetes | Worldwide distribution which includes the following states: AK, AL, AS, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY,MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, WA, WI, VT, WV & WYand the countries of Armenia, Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, El Salvador Finland, France, Germany, Great Britain, Guatemala, Hong Kong, Israel, Ivory Coast, Netherlands, New Zealand, Norway, Peru, Poland, Russia, South Africa, Sweden, Switzerland and Taiwan. | Open |  |  | 42359 | Not_Computer
Z-0099-2014 | 66238 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter CLEARLINK System Non-DEHP Three Lead Extension Set. 6.0" (15 cm), 1.0 mL, Luer Activated Valve, male Luer Lock Adapter, Sterile, nonpyrogenic, Baxter Healthcare Corporation, Deerfield IL 60015.For use with a vascular access device for the administration of drugs and solutions. The Clearlink Luer Activated Valve is an in-line injection site, which can be connected to standard male Luer adapters for continuous or intermittent fluid administration or the withdrawal of fluids. | CLEARLINK System NonDEHP Microbore Three Lead Catheter Extension Set with LAV | 2 | October 29, 2013 | 2013 | Baxter Healthcare Corp. | The firm is issuing a voluntary recall due to the potential for pinholes in the packaging. Pinholes are a breach in the sterile barrier and could lead to potential contamination. |  | Baxter sent an Urgent Product Recall letter dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by Customer/User:1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use. The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and return it to Baxter by either faxing it to 224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.Please have the product code and lot number(s) available. |  | 3200 | 3,200 units | US Distribution including the states of OK, GA, CA, MN, TX and HI. | Open |  |  | 3200 | Not_Computer
Z-0098-2014 | 66362 | prosthesis, elbow, hemi-, radial, polymer | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations (SBi) rHead Radial Stem, rHead Recon Radial System, rHead Radial Stem Lateral Assembly, rHead Radial Stem 6mm Collar, rHead Recon Radial Stem 6mm Collar, rHead Radial Stem Lateral Assembly 6mm Collar, rHead Standard Extended Stem, rHead Radial Stem Plasma Coated, rHead Recon Radial Stem Plasma Coated, rHead Radial Stem 6mm Collar Plasma Coated, rHead Recon Radial Stem 6mm Collar Plasma Coated, uHead Ulnar Stem Standard, uHead Ulnar Stem 20mm Collar packaged in a sterile pouch inside a shrink-wrapped box. The recall was expanded on December 20, 2013 to include the following devices: rHead lateral assembly, radial head implant, sizes 2 through 4, Uni-elbow lateral assembly, radial implant assembly, sozes 2 through 4, ReMotion Left, large distal radial component, x-small, small, medium, and large, ReMotion Right, distal radial component, x-small, small, medium, and large. | Small Bone Innovations (see also Product Description) | 2 | October 29, 2013 | 2013 | Small Bone Innovations, Inc. | The recall was initiated due to an increased risk of implants breaching the sterile pouches inside the shrink-wrapped box. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | Email correspondence and a notification letter, dated September 6, 2013, were sent to users requesting the immediate return of affected devices. Another notification letter, dated December 20, 2013 was also sent for recall of the additional devices after the recall was expanded. |  | 2036 | 2036 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada and Mexico. | Open |  |  | 2036 | Not_Computer
Z-0117-2014 | 66243 | device, monitoring, intracranial pressure | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Brand Names: IntegraÂŽ Camino 110-4 Intracranial Pressure (ICP) MonitoringCatheters.Product / Catalogue No.s: 110-4B, 110-4BC, 110-4G, 110-4H, 110-4HC, 110-4L, 110-4LC.The fully packaged 110-4B, 110-4H, and 110-4L catheters arealso provided as a component in kits that contain other cranialaccess devices. These kits are Catalogue No. 110-4BC, 110-4HC, and 110-4LC respectively.Common Name: Intracranial Pressure and Pressure-Temperature Monitoring KitsClassification Name: Device, Monitoring, Intracranial Pressure510(k) Number: K102875The Camino 110-4 Intracranial Pressure monitoring catheters are sterile transducer-tipped pressure monitoring catheters that are used by neurosurgeons to rapidly determine and continuously monitor intracranial pressure or pressure/temperature.  The Camino 110-4B catheter is used for Intracranial Pressure Monitoring in the parenchyma or subarachnoid space.  The Camino 110-4G catheter is used for Post Craniotomy Subdural Pressure Monitoring.  The Camino 110-4H catheter is intended for Intracranial Pressure Monitoring and Cerebrospinal fluid drainage.  The Camino 110-4L catheter is intended for use with the LicoxÂŽ Brain Oxygen Monitoring System to measure intracranial pressure in the parenchyma.  The fully packaged 110-4B, 110-4H, and 110-4L catheters are also provided as a component in kits that contain other cranial access devices. These kits are Catalogue No. 110-4BC, 110- 4HC, and 110-4LC respectively. | Integra Camino 1104 Intracranial Pressure (ICP) Monitoring | 2 | October 30, 2013 | 2013 | Integra Limited | Integra LifeSciences has identified through an internal evaluation that in order to comply with US FDArequirements, certain CaminoÂż Intracranial Pressure Monitoring Kits require additional labelling to indicatethey are UNSAFE in an MR environment.Only the following CaminoÂż Intracranial Pressure Monitoring Kits are affected:110-4B, 110-4BC, 110-4G, 110-4H, 110-4HC, 110-4L, 110-4LC. |  | Integra sent an Urgent Medical Device Correction Field Safety Notice dted September 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and determine if they have any of the affected product. If so, customers were asked to ensure that the information contained in the letter be distributed to the appropriate personnel. Complete the attached form return it by email or fax as indicated on the form. Customers with questions were instructed to contact Customer Service at 1-855-532-1723.For questions regarding this recall call 1-800-997-4868, outside of the US 888-980-7742. | Contact the recalling firm for information | 79887 | 79,887 | Nationwide Distribution | Open |  |  | 79887 | Not_Computer
Z-0112-2014 | 66144 | gastroscope and accessories, flexible/rigid | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujinon/Fujifilm Ultrasonic EndoscopeIntended to provide ultrasonic images of submucosal and peripheral organs of the upper gastrointestinal tract for observation, diagnosis and endoscopic treatment. | Fujinon/Fujifilm | 2 | October 30, 2013 | 2013 | Fujinon Inc | A limited number of endoscopes were shipped to FMSU-ESD with a manufacturing defect which included missing set screws used to more securely attach the Forceps Inlet (biopsy inlet port) to the endoscope. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | The firm, FUJIFILM Medical Systems, USA, Inc. (FMSU-ESD) sent an "URGENT-MEDICAL DEVICE CORRECTION (UMDC)" and Customer Acknowledgement forms dated July 31, 2013 to their customers. The letter described the issue with the affected product; instructed customers on how to identify affected units; instructed customers to immediately cease use of the affected units; return affected units to FMSU-ESD for repairs and quality control inspection; and complete and return the attached "Customers Acknowledgement" form to FMSU-ESD, Attn: Regulatory Compliance, via fax to (973) 686-2616 or email: gwalljasper@fujifilm.com.If you have any questions regarding this notice, please do not hesitate to contact Director, Quality and Regulatory Compliance, at (973) 686-2636 or by email: gwalljasper@fujifilm.com. |  | 12 | 12 Units | Worldwide distribution: US (nationwide) including states of: NJ, PA and VA; and country of Canada. | Open |  |  | 12 | Not_Computer
Z-0110-2014 | 66524 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Single Site Bipolar Maryland ForcepsIntuitive Surgical, Sunnyvale, CA. 94086The Intuitive Surgical da Vinci Single-Site Instruments and Accessories used with the da Vinci Si Surgical System (IS3000) are indicated for use by trained physicians in an operating room environment for endoscopic manipulation of tissue, grasping, cutting, blunt and sharp dissection, approximation, clip-ligation, electrocautery and suturing during single incision laparoscopic cholecystectomy, benign hysterectomy and salpingo-oophorectomy with the da Vinci Single-Site Instruments and Accessories, including graspers, dissectors, needle drivers, scissors, suction irrigators, monopolar cautery, bipolar cautery, 5 mm curved cannulae, 5 mm and 10 mm straight cannulae, flexible blunt obturators, and the 5mm Single-Site Port. 2. | Single Site Bipolar Maryland Forceps | 2 | October 30, 2013 | 2013 | Intuitive Surgical, Inc. | A limited number of FDA-cleared instruments for da Vinci Single Site use, Bipolar Maryland and curved needle Driver, were shipped without the correct revision of the user manual which provided added instruction for use of these additional Single Site instruments including indication, warnings and precautions. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, Intuitive Surgical, sent an " Urgent Medical Device Correction" letter dated October 2, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure all affected personnel are fully informed of this notice; after receiving the updated manual, locate and discard all previous versions of the Single Site Instruments manual and replace with the new manual; complete and return the attached Acknowledgement Form via FAX to Intuitive Surgical, Inc., ATTN: REGULATORY COMPLIANCE, U.S. (408) 716-3040, or Scan and Email to: isi.compliance@intusurg.com.; Inform affected personnel when the correction has been completed; and Retain a copy of this Notice and the Acknowledgement Form for your files.If you have any questions, contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" North and South America: 800-876-1310 Option 3 (6 am to 5 pm PST)" Japan: 0120-56-5635 or 003-5575-1362 (9 am to 6 pm JST)" South Korea: 02-3271-3200 (9 am to 6 pm KSTJ)" Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). | Contact the recalling firm for information | 286 | 286 boxes ( 6 instruments per box) for a total of 2,220 affected clinical product | Worldwide distribution: US (nationwide) and countries of:Australia, Belgium, France, Italy and United Kingdom. . | Open |  |  | 286 | Not_Computer
Z-0109-2014 | 66222 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Vented Spike Adapter Product Code 2C0471, Sterile, nopyrogenic. Manufactured by an affiliate of Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield IL 60015 USA.An IV transfer device, used to transfer fluids in pharmacy compounding system. | Vented Spike Adapter | 2 | October 30, 2013 | 2013 | Baxter Healthcare Corp. | Potential tears in the pouch of six lots of Vented Spike Adapter Product Code 2C0471 |  | The firm, Baxter, sent an "Urgent Product Recall" letter dated September 10, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from your facility; If you distributed the products to other facilities or departments, forward them a copy of the letter; complete and return the attached Customer Reply Form via fax to 1-224-270-5457 or scan and email to fca@baxter.com; and contact Baxter Healthcare Center for Service at 1-888-229-0001 between hours of 7:00 am and 6:00pm Central Time to return affected product and receiving credit.For additional information contact the Center for One Baxter at 1-800-422-9837, Monday thru Friday during the hours of 8:00am to 5:00pm Central Time. | Contact the recalling firm for information | 49344 | 49,344 units | Nationwide distribution. | Open |  |  | 49344 | Not_Computer
Z-0115-2014 | 66410 | vehicle, motorized 3-wheeled | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Breeze C Scooter is a motorized electric scooter for elderly people. Brand name Breeze, model C. It is intended to provide increased mobility for elderly and/or disabled person and it is also an aid to independent living.Products can be identified by their color; the recalled products are either metallic blue or metallic silver. The products have two rear wheels and have either one or two wheels in the front. The scooters have an adjustable swivel seat and a front storage basket; some scooters may be fitted with an optional larger rear basket. | Breeze C Scooter | 2 | October 30, 2013 | 2013 | S.A.E. AFIKIM | Breeze C Scooter, motorized 3-wheeled electric scooter for elderly people, is recalled because it may become unresponsive during operation due to an electrical contact which may fault due to extensive wear, which potentially could result in the driver not being able to stop the scooter during operation except by turning off the ignition key. | PRODUCTION CONTROLS: Process Control | Afikim Electric Vehicles sent an URGENT: MEDICAL DEVICE RECALL letter, dated September 27, 2013, to all affected customers and distributors. .Customers are advised to remove the ignition key and contact Afikim Electric Vehicles via telephone Monday through Friday at 1-800-809-3010 between the hours of 9AM and 5PM (EST). Consumers may also contact the company via e-mail at dov@afiscooters.com.Customers are instructed to return the Recall Response Form by fax it to 800-466-1171 or email to ofer@afiscooters.com.Consumers should report any adverse event or quality problems to FDA's MedWatch Adverse Event Reporting program either by online, fax, or regular mail. |  | 69 | 69 units | Nationwide Distribution: US including the states of California, Florida, Georgia, Illinois, Maine, Maryland, Massachusetts, Michigan, Minnesota, New Jersey, New York, Ohio, Oklahoma, Pennsylvania, Texas, Virginia, and Wisconsin. | Open |  |  | 69 | Not_Computer
Z-0114-2014 | 65997 | cerclage, fixation | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument Sternal Zip Flex Systemintended for use in primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion | Synthes Application Instrument Sternal Zip Flex System | 2 | October 30, 2013 | 2013 | Synthes USA HQ, Inc. | When using the 1st generation Zip Fix Application Instrument it is possible to cut the implant while the tensioning trigger is being squeezed which could lead to loosening of the implant. | DESIGN: Device Design | Synthes sent a Urgent Medical Device Notification dated August 13, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: This notification is to make you aware of the issue and the importance of completely releasing the tensioning trigger before cutting the implant. Please take the following action:"To ensure your receipt and understanding of this information, Please complete the verification section of this letter and return to Synthes by either of the following methods:" Fax: 1-888-352-9616 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 319 | 319 | Worldwide Distribution - US (nationwide) and country of Mexico. | Open |  |  | 319 | Not_Computer
Z-0123-2014 | 66266 | filler, bone void, calcium compound | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Globus Medical MicroFuse Putty, 2.5cc and 10cc.Product Usage:MicroFuse Bone Void Filler is indicated only for bony voids or gaps that are not intrinsic to the stability of the bony structure. MicroFuse Bone Void Filleris a bone graft extender to be used with autogenous bone marrow aspirate or autograft | MicroFuse Bone Void Filler | 2 | October 31, 2013 | 2013 | Globus Medical, Inc. | The sterility of this product to a Sterility Assurance Level (SAL) of 10(-6) cannot be assured. | PRODUCTION CONTROLS: Process Control | Globus sent an Urgent: Medical Device Recall Notification letter, dated August 28, 2013 to customers/users. The letter identified the affected units, issue, potential risk, product, and actions to be taken. Customers were requested to complete the attached response card and return of the affected product to Globus Medical. For questions and support contact Globus Medical by fax 1-610-300-1342 or email:recall@globusmedical.com. |  | 193 | 193 | USA Nationwide Distribution. | Open |  |  | 193 | Not_Computer
Z-0126-2014 | 65469 | system, x-ray, mobile | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-ray System.Produces radiographic images. | DRXRevolution | 2 | October 31, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. is recalling the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. | DESIGN: Device Design | On June 6, 2013 Carestream Health Inc. initiated a recall for the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. Carestream Health Inc. will replace the power plug in all potentially impacted DRX-Revolution Mobile X-Ray Systems. Your Carestream service representative will contact you to answer any questions and schedule a visit to replace the plug of your Mobile X-Ray System as soon as possible. If you have any questions or concerns, or if you experience any issue with your plug prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time. |  | 180 | 180 units | Nationwide Distribution including the states of NY, TX, PA, OH, WI, CA, IL, CO, FL, IN, OK, NJ, GA, LA, AL, KS, AZ, MA, SC, MI, NC, IA, UT, VA, TN, and MN. | Open |  |  | 180 | Not_Computer
Z-0091-2014 | 66305 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences EMBOL-X Glide Protection System; Model Numbers: EXGF24D, EXGF24LLD, EXGF24MMD, EXGF24SSD, EXGF24XLD, EXGF24XSD, EXGFXS2D, EXGF24SS2D, EXGF24MM2D, EXGF24LL2D and EXGF24XL2D.The EMBOL-X access device is indicated for the perfusion of the ascending aorta during short-term (less than or equal to0ÂŽ 6 hours) cardiopulmonary bypass (CPB) surgery where procedures may require the hemostatic introduction and removal of compatible intravascular devices into the vascular system. | Edwards Lifesciences EMBOLX Glide Protection System | 1 | October 31, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of EMBOL-X Glide Protection System due to deformation of the cannula tip. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL-ACTION REQUIRED" letter dated September 16, 2013 via Fed-Ex to its consignees/customers. Consignees were asked to review inventory for the cannulae and remove the cannulae from any kits received; quarantine affected material at site and return product to Edwards LIfesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020, Attention: RECALL, RGA #; call Edwards Customer Service at 800-424-3278 to obtain RGA# and replacement product (note: the filter is not affected and does not have to be returned with the cannulae); and complete and return the attached acknowledgment form via fax to Edwards Customer Service at 800-422-9329 within three days of receipt of this Field Safety Notice.If you have any questions, please call Edwards Customer Service at 800-424-3278 from the hours 6:00am - 4:30pm PST or contact your Edwards' sales representative concerning the recall. |  | 12034 | 12,034 | Worldwide distribution: US (nationwide) and country of: Canada. | Open |  |  | 12034 | Not_Computer
Z-0128-2014 | 66423 | prosthesis, laryngeal (taub) | EWL | Laryngeal prosthesis (Taub design). | Ear Nose & Throat | Ear Nose & Throat | 510(k) | Blom-Singer Adjustable Tracheostoma Valve IIModels: BE8025H and BE8025H-R3 | BlomSinger Adjustable Tracheostoma Valve II | 3 | November 01, 2013 | 2013 | Helix Medical LLC | InHealth Technologies (IHT) is recalling the Blom-Singer Adjustable Tracheostoma Valve II (ATSV II) because the body component of the device is missing from the packaging. The ATSV II Valve product cannot be used as intended without the product body. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | A phone script and a customer notification letter are provided. The recall was initiated on 10/1/13 and customers will be notified that InHealth Technologies, a division of Helix Medical, LLC has initiated a recall of the following 3 lots of the Blom-Singer Adjustable Tracheostoma Valve (BE8025H and BE8025H-R3): 1057760, 1057919 and 1059222. The phone script provides the problem identified and the actions to be taken. Customers are instructed to complete the Customer response form and return by email to belinda.jackson@helixmedical.com, fax to (805) 566-5381 or by mail to InHealth Technologies, 1110 Mark Avenue, Carpinteria, CA 93013. Customers are also instructed to Contact UPS at 1-800-742-5877 and provide In Health Technologies UPS account no., 846718. UPS will come to your site, package the recalled product(s) and ship it to InHealth Technologies at InHealth Technologies expense. Customers are instructed to respond within 10 days. |  | 47 | 47 units | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Croatia, Germany, Ireland, and Italy. | Terminated | January 21, 2014 | 81 | 47 | Not_Computer
Z-0129-2014 | 66196 | washer, bolt, nut, non-spinal, metallic | NDG | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ti-Magna Fx WasherTi-Versa-FXÂŽII Femoral Fixation System S/C PlatesThe compression tube/plates of this trauma system are used with Ti-Versa-FxÂŽ II Femoral Fixation System lag screws for the internal fixation of proximal femoral and supracondylar fractures. The Ti-Magna Fx System is an osteosynthesis item for orthopedic use in open internal fixation of fractures of the femoral neck, condyle, or tibia, etc. The washer prevents the screw head from being buried in the bone. | TiMagna Fx WasherTiVersaFXII Femoral Fixation System S/C Plates | 2 | November 01, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing, to provide statisti | PRODUCTION CONTROLS: Packaging Process Control | The firm, Zimmer Inc., sent a letter entitled "Voluntary Recall of Ti-VFxII Condylar Tube Plates and Ti-Washers" dated September 12, 2013 to its customer. The letter described the product, problem and action to be taken. The customer was instructed to immediately discontinue use of the listed products. A Sales representatives from Zimmer K.K. or agencies will be in contact with you. Zimmer Inc. is planning to resume product manufacturing upon the completion of package validation review. Zimmer K.K. will contact you when a detailed schedule has been determined.If you have any questions, contact Zimmer K.K. Safety Office (9:00 to 17:00; closed weekends and public holidays) TEL: 0120-015-180. |  | 2632 | 2,632 devices | International distribution: Japan only. | Open |  |  | 2632 | Not_Computer
Z-0113-2014 | 66383 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTERÂŽ LH 750and LH 780 Analyzers.The COULTER LH 750 Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The COULTER LH 750 Hematology Analyzer provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids. The COULTER LH 780 Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The COULTER LH 780 Hematology Analyzer also provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids. | Coulter LH 750 Hematology Analyzer Coulter LH 780 Hematology Analyzer | 2 | November 01, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for the Coulter LH 750 Hematology Analyzer (PN 6605632 and A85570) and Coulter LH 780 Hematology Analyzer (PN 723585 and A90728) because the check valves may fail. A failure may result in an air or liquid leak at the connection to the Vacuum Overflow Tank waste line or at the Backwash Tank drain line. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Beckman Coulter sent a Urgent Product Correction letter dated September 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . The customer notification letter dated September 12, 2013, titled "URGENT PRODUCT CORRECTION", provided product description with codes, the issue (reason for recall), impact, action, resolution, and contact information.Customers were instructed to:If you experience a "Backwash Tank Not Full" error message, review manual mode sample results since the last acceptable QC run. System safeguards that assist with detection of erroneous results include Startup, QC, system messages and flagging (e.g., R), delta checks, flagging from limit sets (H, L, a, c), definitive messages and decision rules that would trigger further review.Avoid physical contact with a possible leak by continuing to use personalprotective equipment as recommended in the Instructions for Use and by national regulations.If you have any questions concerning this notification, please contact our Customer Support Center: Via our website, http://www.beckmancoulter.com/customersupport/supportVia phone, call 800-526-7694 in the United States and Canada.Outside US and Canada, contact your local Beckman Coulter RepresentativeWe apologize for the inconvenience this issue has caused your laboratory. |  | 1697 | 1,697 units | Worldwide Distribution - Nationwide Distribution and the country of Canada.. | Open |  |  | 1697 | Not_Computer
Z-0085-2014 | 66368 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Baylis Medical TorFlex Transseptal Guiding Sheath, Rx Only, non-pyrogenic, sterile.Product Usage: The TorFlex Transseptal Guiding Sheath kit (K102948) is a single use device consisting of three components: a sheath, a dilator, and a J-tipped guidewire. The kit is designed for safe and easy catheterization and angiography of specific heart chambers and locations. | TorFlex | 1 | November 01, 2013 | 2013 | Baylis Medical Corp * | The TorFlex Transseptal Guiding Sheath Kits are being recalled due to the possibility that the sheath could remove particulate of less than 5 mm from the dilator. The particulate could potentially migrate into a patients bloodstream. | DESIGN: Component Design/Selection | Baylis Medical sent an Urgent Device Recall Notification letter dated September 23, 2013, via overnight courier to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form provided and fax to 905-602-5671 ATTN: Quality Department to arrange for a free of charge replacement of the affected product. The form requires consignees to return the affected product. For questions contact Quality Assurance at 800-276-4416. |  | 370 | 370 devices | Worldwide Distribution - USA Nationwide in the states of : NE, NH, NC, FL, PA, VA, MI, CA, AL, and TN and the countries of China, Canada, Saudi Arabia, and Italy. | Open |  |  | 370 | Not_Computer
Z-0127-2014 | 66456 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dekompressor Percutaneous Discectomy Probe.The Dekompressor is a single use disposable discectomy probe that passes through and works in conjunction with an introducer cannula to remove intervertebral disc nucleus pulposus material. | Dekompressor Percutaneous Discectomy Probe | 2 | November 01, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Stryker has identified a potential for the auger and/or cannula to fracture when the Dekompressor is used contrary to the warnings in the Instructions For Use (IFU). Specifically, manually bending the cannula can weaken the device and may result in a fracture of the auger and/or cannula. This could result in the need for additional medical intervention to remove the broken piece, pain or loss of m | DESIGN: Device Design | On 11/4/13, an URGENT MEDICAL DEVICE Recall Notification was sent to all domestic consignees via mail and foreign by email. An advisory was sent to every physician that uses any of the products listed. Every physician thatuses any of these products must print their name and sign below which means you acknowledge that youve received a copy of the attached advisoryand that you understand the information regarding the proper usage of the Decompressor.Once the form has been signed by all physicians, the form mustbe completed and signed by the recall coordinator. Please fax to866-521-2762 or email to kara.spath@stryker.com. |  | 2004 | 20,385 total distributed since 2004 | Nationwide Distribution including the states of AR, AZ, CA, CO, FL, GA, IL, KS, MS, ND, and NE. | Open |  |  | 20385 | Not_Computer
Z-0134-2014 | 66112 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Fem-Flex Femoral Arterial Cannula with Duraflo Coating, DIITFA02225, Sterile, RX only For Single Use Only, Product Usage:Edwards Lifesciences femoral access cannulae are intended for use in situations which require rapid femoral venous and arterial access for short-term ( | Duraflo coated FemFlex Arterial Cannula | 2 | November 01, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling two lots of Duraflo Cannulas due to reported complaints involving cracks in the t-connector of the devices which were confirmed during product evaluation. | PRODUCTION CONTROLS: Process Control | Edwards sent an Urgent Field Safety Notice Product Recall letter dated August 26, 2013 to all affected customer, via Fed-Ex. The product is sold directly to hospitals and distributors. The distributors involved will communicate with their customers directly. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their entire inventory, quarantine and return the affected products. An acknowledgement form is included for completion to assist in the assessment of the inventory. For questions contact Edwards Customer Service at 800-424-3278. |  | 80 | 80 units | Nationwide Distribution. | Open |  |  | 220 | Not_Computer
Z-0130-2014 | 66545 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, NON-POROUS TIBIAL BASE, SIZE 5, LEFT, QTY: (1), REF 71420168.Product Usage: orthopaedic | GENESIS(R) II, NONPOROUS TIBIAL BASE, SIZE 5, LEFT | 2 | November 01, 2013 | 2013 | Smith & Nephew Inc | Swapped product. Non-porous tibial base size 5 L was mis-packaged as size 6 R and vice versa. | PRODUCTION CONTROLS: Process Control | Smith & Nephew sent an Urgent - Product Recall 1st Notification letter to all affected sales staff via email and FedEx on October 7, 2013. The hospital was notified by sales rep on October 9, 2013 and the notification letter was sent to the hospital via Fed Ex on October 11, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, discontinue use, quarantine affected products immediately and return to Smith & Nephew. Customers were asked to complete the attached Inventory Return Certification Form and fax to 901-566-7975. | Contact the recalling firm for information | 4 | 4 units | USA Nationwide Distribution in the states of: AZ, NC, IA, PA, OH, MD, and TN. | Open |  |  | 4 | Not_Computer
Z-0136-2014 | 66101 | assay, glycosylated hemoglobin | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bayer brand A1CNow+ MULTI-TEST A1C SYSTEM TEST KITS (PART NUMBER 3021/08842610); Hemoglobin A1c Monitor; Glycosylated Hemoglobin Assay; Product is manufactured and distributed by Bayer HealthCare LLC, Sunnyvale, CA The A1CNow+ test provides quantitative measurement of the percent of glycated hemoglobin (%HbA1c or %A1C) levels in capillary (fingerstick) or venous whole blood samples. The test is for professional use to monitor glycemic control in people with diabetes. | Bayer brand A1CNow MULTITEST A1C SYSTEM TEST KITS | 2 | November 02, 2013 | 2013 | Bayer Healthcare Llc | Test results from these lots may produce a high-bias reading (by up to 20%) that is outside of limits defined for stability and product accuracy when product is stored at room temperature four (4) to five (5) months after release. |  | Bayer Healthcare sent an Urgent Medical Device Recall letter dated August 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and quarantine any affected product. Complete the attached Recall Response Form and fax to Bayer Customer Service. Customers were instructed to use the attached letter to notify affected next level customers of the recall. For questions regarding this recall call 408-524-2255, ext 290. |  | 138 | 138 Kits in US (Total 2452) | Worldwide Distribution - USA (nationwide) and Internationally to China, Argentina, Uruguay, Japan. | Terminated | January 14, 2014 | 73 | 138 | Not_Computer
Z-0146-2014 | 66570 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | da Vinci Vision System Cart, a component of theda Vinci Si Surgical System IS3000;Y1903 Illuminator (Illuminator + Lamp Module); Y1902 Replacement Lamp Modules.Intuitive Surgical, Sunnyvale, CA.As a component of the IS3000 da Vinci Si Surgical System, the 3DHD Vision System provides a high resolution image for the surgeon (at the 3D viewer) and the patient-side assistant (at the touchscreen). The 3DHD endoscope assembly (endoscope and camera head) can be used manually (that is, handheld, as with a traditional endoscopy cart) or can be mounted on the camera arm to assist with preoperative exploration. Also integrated in the Vision System Cart is the Illuminator. The illuminator has a single light source attached to the endoscopy assembly by the light guide cable. It provides illumination inside the body for vision of the surgical field. The Illuminator monitors the life left on the lamp module and notifies the customer when replacement is recommended. | da Vinci Vision System Cart | 3 | November 04, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive has identified that the metalized coating on the IS3000 HD lamp module may be incompatible with the illuminator control board, leading to an erroneous red Preventive Maintenance Message that shows up on monitors. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to open the illuminator and replace any metal-finish lamp modules with plastic finish lamp modules, if customers do not have a replacement plastic-finish lamp module available, they may continue to use the metal-finish lamp module until they receive a replacement. If the message "illuminator lamp module error: Please reseat or replace lamp module" is displayed on the patient side monitor and the surgeon console, contact intuitive Surgical Customer Service. Customers were also instructed to examine any remaining or spare inventory to identify any metal-finish lamp modules and return them to Intuitive Surgical for credit. Customers were instructed to ensure all affected personnel are fully informed of this Urgent Medical Device Corrrection letter. Customers should complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed. Customers with questions should contact Intuitive Surgical Customer Service.For questions regarding this recall call 408-523-2244. |  | 161 | Standalone: 161 units; With Vision Starter Kit: 240 units; With IS3000 Vision Cart: 283 units | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Austria, Belgium, Brazil, *Bulgaria, Canada, Chile, China, CzechRepublic, Denmark, Finland, France, Germany, Great Britain, India, Israel, Italy, Mexico, Monaco, Netherlands, Norway, Pakistan, Panama, Qatar, Romania, Russian Fed., Saudi Arabia, Singapore, South Korea, Sweden, Switzerland, Taiwan, and Turkey. | Open |  |  | 684 | Not_Computer
Z-0144-2014 | 66040 | punch, corneo-scleral | HNJ | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | MLS Torpedo Mini-Light Ophthalmic single-port illumination.Sutureless 25- gauge illuminator for vitreo-retinal surgery when full field illumination of bimanual surgery if desired. | MLS Torpedo MiniLight Ophthalmic singleport illumination | 2 | November 04, 2013 | 2013 | Insight Instruments, Inc. | The product was registered and distributed as a Class I medical device and was subsequently determined to be a Class II medical device which requires FDA market clearance. | PREMARKET APPROVAL: No Marketing Application | Insight Instruments, Inc. sent a letter dated April 8, 2013, via USPS priority mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue using and dispensing all affected product, perform a physical count of their inventory and complete the Return Response Form and fax to 772-219-9342. Customers were also instructed to call 800-255-8354 to arrange for return and prompt reimbursement.For questions regarding this recall call 727-219-9393. |  | 10212 | 10212 | Worldwide Distribution - USA (nationwide) including AZ, CA, FL, IN, MA, MD, TX, and WA and Internationally to Czech Republic, Germany, Netherlands, Spain, Switzerland, and Turkey. | Open |  |  | 10212 | Not_Computer
Z-0147-2014 | 66382 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis SystemsThe UniCel DxH 800/DxH 600 analyzers are quantitative, multiparameter, automated hematology analyzers for in vitro diagnostic use in screening patient populations found in clinical laboratories. The DxH 800 is available as an analyzer for use on a tabletop or with a floor stand. The DxH 600 is available as an analyzer for use on a tabletop only. | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems | 2 | November 04, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for UniCel DxH 800 Coulter Cellular Analysis System (PN 629029) and UniCel DxH 600 Coulter Cellular Analysis System (PN B23858) because the company's internal investigation indicates that specific lots of check valves used in several locations within the DxH 800 and DxH 600 systems may fail. A failure may result in an air or liquid leak at the VCS Waste Chamb |  | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated September 12, 2013, its customers. The letter includes the products descriptions and codes, issue, impact, action, resolution, instructions to the customers, and contact information. The customers were instructed to review the three samples analyzed prior to error message,if you experience a "Vent/overflow chamber detected overflow' error; to avoid physical contact with a possible leak by continuing to use personal protective equipment as recommended in the Instructions for Use and by national regulations; share this information with your staff and retain this notification as part of your Quality System documentation; if you have forwarded any of the affected products, provide them a copy of this letter; and to complete and return the enclosed RESPONSE FORM within 10 days via fax to: 786-639-7500 /7501/7502/7504; email: Regulatory.Notifications@beckmancoulter.com (note: please include "DxH system check valve letter dated 09/12/2013" in email subject field) and/or by mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami Fl 36116-9015. .Your service representative will contact you to schedule an on-site visit for check valve inspection and replacement of any affected valves.If you have any questions concerning this notification, please contact our Customer Support Center: via our website, http://www.beckmancoulter.com/customersupport/support, via phone, call 800-526-7694 in the United States and Canada, and Outside US and Canada, contact your local Beckman Coulter Representaitive. |  | 1025 | 1,025 units | Worldwide distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Belgium, Canada, China, Croatia, Czech Republic, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Macao, Netherlands, New Zealand, Panama, Poland, Portugal Qatar, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Thailand, Trindad and Tobago, Turkey, United Arab Emirates, and United Kingdom.. | Open |  |  | 1025 | Not_Computer
Z-0145-2014 | 66457 | resectoscope | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Individual units labeled in part: REF: 4630022; for S(a)line resectoscopes; Vaporization electrode, sterite; 4mm telescope 30 25 / 12: only for continuous irrigation sheath 24,5 Fr.; LOT 426111; 2011/06; 2016/06; STERILE; RICHARD WOLF; Cartons of 3 units labeled in part: REF 46300223; Vaporization electrode, sterile; 4mm telescope 30/25/12; only for continuous irrigation sheath 24,5 Fr.; Lot 426111; 2011/06; 2016/06; STERILE; RICHARD WOLFS(a)line Resectoscopes are used for endoscopically controlled removal (ablation) of tissue using 0.9% NaCl solution (saline) as the irrigation medium. It can also be used for monopolar resection if necessary. | BiPolar Vaporization Electrodes | 2 | November 04, 2013 | 2013 | Richard Wolf Medical Instruments Corp. | Richard Wolf Medical Instrument Corporation recieved a complaint stating that a 4630024 vaporization electrode was packaged in a 4630022 pouch. | PRODUCTION CONTROLS: Process Control | The firm, Richard Wolf, sent an "Urgent: Medical Device Recall" letter dated September 16, 2013, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory of vaporization electrodes from lot 426111, contact Richard Wolf Medical Instrument Corporation at 800-323-9653 x 344 or ext 225 or ext. 380 to inform Richard Wolf Medical Instrument Corporation if the customer has or does not have any devices from this lot, return the affected devices with the return authorization provided by Richard Wolf Medical Instrument Corporation for credit, and complete and return the attached Recall Response form via fax to: 847-913-0924; email: dclark@richardwolfusa.com; or mailed to: Richard Wolf Medical Instrument Corporation, 353 Corporate Woods Parkway, Vernon Hills, IL 60061-3110.Customers who require urgent replacement of the affected devices can contact their customer service representative for replacement.If you have any questions, call 800-323-9653 x 344. |  | 87 | 87 cartons (261 units) | Nationwide distribution: US including states of: CA, GA, IL, KY, LA, MI, MD, NY, OR, TX, VA and Puerto Rico. | Open |  |  | 87 | Not_Computer
Z-0139-2014 | 66526 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Patient Cart used in conjunction with the da Vinci S surgical and SI System.Intuitive Surgical1266 Kifer RoadSunnyvale, CA 94086The Patient Side Cart (PSC) is a robotic platform used for da Vinci Surgery. It is the operative component for both da Vinci systems and its primary function is to support he instrument arms and camera arm. It uses remote center technology. The remote center is a fixed point in space around which the Patient Cart arms move. Remote center technology enables the system to maneuver instruments and endoscopes in the surgical site while exerting minimal force on the patient body wall. The Patient Cart works in the sterile field, assisting the Surgeon Console operator by exchanging instruments and endoscopes, and by performing other patient-side activities. To help ensure patient safety, the action of the Patient Cart operator take precedence over actions of the Surgeon Console operator. | Patient Cart used in conjunction with the da Vinci S surgical and SI System | 2 | November 04, 2013 | 2013 | Intuitive Surgical, Inc. | Increased number of complaints regarding fluid entering the patient card base of the da Vinci S and SI systems. | DESIGN: Device Design | Intuitive Surgical Inc sent an Product Notification to all affected customers via email on November 30, 2012.E-mail notifications advised users of an addendum to the User Manual for Covering the Patient Cart column that provided a recommended method for covering the da Vinci patient cart column to prevent splashing on the cart during a procedure. The addendum was attached to the e-mail and is available online.If you have any questions, please contact Customer Service at 1-800-876-1310. Outside the United States call +41 21 821 2020 or +800.0821.2020. | Contact the recalling firm for information | 1754 | 1754 patient carts | Worldwide Distribution - USA including Puerto Rico and the countries of Australia, Austria, Argentina, Belgium, Bulgaria, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Israel, Japan, Kuwait, Mexico, Netherlands, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay and Venezuela. | Terminated | November 04, 2013 | 0 | 1754 | Not_Computer
Z-0167-2014 | 66461 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Aquatec Ocean VIP, Aquatec Ocean Dual-VIP, Aquatec Ocean E-VIP mobile shower and toilet commode with tilting function. Mobile shower and toilet commode with tilting function | Aquatec Ocean VIP | 2 | November 05, 2013 | 2013 | Invacare Corporation | The snap-on fixtures of the backrest cushion on the Mobile Shower and Toilet Commode Aquatec Ocean VIP and Ocean Dual VIP do not withstand a required pull out force of 500 N, when the product is used in a reclined position and or tilted position. This may result in the snap-on fixture coming unlatched. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge their receipt and understanding of the urgent recall notice and indicate the number of affected product in their inventory on the enclosed Provider Response Card. Customers were instructed to examine their existing stock and quarantine any affected product and fax the completed Provider Response Card immediately to Invacare Corporation at 440-326-3544. Customers were responsible for ensuring that the recall was conducted to the customer level. Correction kits were provided to the customers with the communication. Customers were instructed to document the correction details on the enclosed Product Tracking Sheet and fax to Invacare Customer Service at 440-326-3544.For questions regarding this recall call 440-329-3544. |  | 538 | 538 (US only) 3317 (Total) | Worldwide Distribution - USA including PA and Internationally to Canada. | Open |  |  | 538 | Not_Computer
Z-0148-2014 | 66339 | system, peritoneal, automatic delivery | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius LibertyÂŽ CyclerPart Number: RTLR180111Product Usage:The Liberty cycler is a portable device designed to perform acute and chronic peritoneal dialysis. | Fresenius | 2 | November 05, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Fluid leaking may be observed inside the pump compartment of the Liberty Cycler and may result in dialysate becoming contaminated, potentially causing peritonitis | DESIGN: Device Design | Fresenius Medical Care contacted all affected consignees by telephone on 10/14/13 to inform them of the product safety alert. Follow-up Urgent Safety Alert letter was sent on 10/18/13 and 10/23/13 by Certified Mail, Return Receipt requested. The letter identified the affected product, problem and actions to be taken. Customers were instructed to report any events associated with this safety alert to Customer Service Team at 1-800-323-5188. |  | 16600 | 16,600 | USA Nationwide Distribution | Open |  |  | 16600 | Not_Computer
Z-0151-2014 | 66465 | implant, intragastric for morbid obesity | LTI |  |  | Gastroenterology/Urology |  | Adjustable Gastric Band, packaged alone in transparent PETG inner and outer blisters with Tyvek lids.Product Usage:Intended for use in surgical treatment of morbid obesity | SAGB QuickClose Swedish Adjustable Gastric Band including Velocity Injection Port and Applier | 2 | November 05, 2013 | 2013 | Ethicon Endo-Surgery Inc | During aging studies, the firm determined that the gastric belts were unable to pass the test requirements established for Force to Lock. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Ethicon sent an Urgent: Medical Device Recall letter dated September 30, 2013 to affected consignees via overnight, UPS to allow tracking of the receipt. The letter identified affected products, problem and actions to be taken. Each consignee is requested to respond to a provided Business Reply Form (BRF), indicating whether they have affected product, how much and whether they will/have returned the product or whether they no longer have any of the recall product. For questions contact your local sales representative or call our Customer Support Center 1-877-384-4266 option 6. |  | 1250 | Product Code: BD2XV - 1,248 units | Worldwide Distribution: USA (nationwide) including the states of: Alabama, Connecticut, Indiana, Michigan, North Carolina, South Carolina, Alaska, Delaware, Kansas, Minnesota, Ohio, Tennessee, Arizona, Florida, Kentucky, Mississipp,i Oklahoma, Texas, Arkansas, Georgia, Louisiana, Nebraska, Oregon, Virginia, California, Idaho, Maryland, Nevada, Pennsylvania, Washington, Colorado, Illinois, Massachusetts, New Jersey, Rhode Island, and West Virginia; and the countries of: Argentina, France, Malaysia, Spain, Australia, Georgia, Martinique, Sweden, Belgium, Germany, Mexico, Switzerland, Brazil, Greece, Netherland, Taiwan, Chile, Hungary, Panama, Tunisia, China, India, Philippines, United Arab Emirates, Czech Republic, Israel, Poland, United Kingdom, Egypt, Italy, Portugal, Venezuela, Estonia, Kenya, Russia, Finland, Kuwait, and Saudi Arabia. | Open |  |  | 6893 | Not_Computer
Z-0153-2014 | 66253 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ISE REF Valve, Part No. MU7638 for the AU480, AU680, and AU5800 Chemistry Analyzer.Automated chemistry analyzers that measures analytes in samples. | Beckman Coulter AU480, AU680, AU5800, and AU5400 Clinical Analyzers | 2 | November 05, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that some AU Clinical Chemistry Analyzer ISE units can experience premature ISE REF (Reference) valve failure. Premature valve failure can be detected by the presence of air bubbles in the ISE REF solution tubing on the ISE unit. If air bubbles are present, accuracy of the ISE test results may be affected. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A customer notification letter dated 9/9/13 was sent to customers who received the ISE REF Valve, Part Number MU3226 for the AU400, AU640, AU2700, AU5400 and the ISE REF Valve, Part Number MU7638 for the AU480, AU680, AU5800 Chemistry Analyzers. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center via http://www.beckmancoulter.com/customersupport/support, via phone at (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. | Contact the recalling firm for information | 4584 | 4584 units total (1173 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, Vietnam, Zambia, and Zimbabwe. | Open |  |  | 4584 | Not_Computer
Z-0164-2014 | 66351 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | PortexÂŽ Saddleblock Tray with Drugs  25g Quincke; 4795-20The Regional Anesthesia family of products is comprised of sterile (unless otherwise indicated in the product scope), single-use devices designed to perform epidural, spinal, combined epidural/spinal, nerve block, lumbar puncture and regional anesthesia procedures. The spinal products are comprised of spinal needles, introducer needles and accessories required to perform a spinal procedure. | Spinal Anesthesia Trays | 2 | November 05, 2013 | 2013 | Smiths Medical ASD, Inc. | Visual particulate in the glass ampules of 5% Lidocaine HCL and 7.5% Dextrose Injection, USP, 2 mL Ampoule, NDC # 0409-4712-01, Hospira Lot Number 23-227-DK. These ampules are included in certain Portex Spinal Anaesthesia Trays. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | Smiths Medical sent an Urgent Recall Notice dated September 23, 2013, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and remove from use. Complete the Confirmation Form and return by fax to 603-358-1017 or by email to spinal@smiths-medical.com. Upon receipt of the completed form, a customer service representative would contact them to arrange for exchange of their unused affected devices for credit or replacement. Customers were also instructed to forward the notice to all personnel who need to be aware within their organization. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at 800-258-5361.For questions regarding this recall call 800-258-5361. | Contact the recalling firm for information | 222 | 222 | Nationwide Distribution including AK, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SC, TX, UT, VA, VT, WA, WI, WY. | Open |  |  | 7549 | Not_Computer
Z-0165-2014 | 66425 | pump, infusion, insulin | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPod Insulin Management System Starter Kit; Model Reference Number: SKT-UST400 (Starter Kit)Product Usage: The Starter Kit (SKT-UST400) is provided to users as an introduction to the OmniPod Insulin Management System prior to initiating care. It includes the Personal Diabetes Manager (PDM) with two AAA alkaline batteries, a USB cable, a Users Guide, a Soft Carrying Case, a Troubleshooting guide, a Warranty Registration Card, and two Alcohol Pads. The OmniPod Insulin Management System is intended for subcutaneous delivery of insulin at set and variable rates for the management of diabetes mellitus in persons requiring insulin and for the quantitative measurement of glucose in fresh whole capillary blood (in vitro) from the finger. The glucose measurements should not be used for the diagnosis of or screening for diabetes. The PDM glucose meter is intended for single-patient use and should not be shared. Abbott FreeStyle test strips are used with the built-in FreeStyle meter for the quantitative measurement of blood glucose in fresh whole capillary blood from the finger, upper arm and palm. Abbott Freestyle Control Solutions are used to verify that the meter and test-strips are working together properly and that the test is performed correctly. | OmniPod Insulin Management System Starter Kit | 2 | November 05, 2013 | 2013 | Insulet Corporation | Alcohol Prep Pads contained within the kit are not properly labeled. | DESIGN: Labeling Design | Insulet sent an Important Corrective Action Notification letter dated September 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to discontinue use and properly dispose of the affected product. For questions contact Customer Care at (866) 941-0294. |  | 1719 | 1,719 starter kits | Worldwide Distribution - USA Nationwide and the country of Israel. | Open |  |  | 6769 | Not_Computer
Z-0154-2014 | 66532 | coronary drug-eluting stent | NIQ |  |  | Cardiovascular |  | Resolute Integrity Zotarolimus-eluting Coronary Stent:Stent ID 3.0 mm; Stent length 34 mm; Rx, Sterile.Mfg in Medtronic Ireland, Galaway, Ireland. Distributed by: Medtronic Inc.Minneapolis, MN.Indicated for improving coronary luminal diameters in patients. | Resolute Integrity Zotarolimuseluting Coronary Stent | 2 | November 05, 2013 | 2013 | Medtronic Vascular, Inc. | A small number of units in lot number 0006573585 of the Resolute Integrity Zotarolimus-Eluting Coronary Sten System may contain a compliance chart incorrectly referencing a 2.25 mm diameter instead of the correct 3.0mm. | PRODUCTION CONTROLS: Process Control | Urgent Medical Device Recall Letters were sent on October 7, 2013 via UPS 3-day service. |  | 201 | 201 units | Nationwide Distribution. | Terminated | February 11, 2014 | 98 | 201 | Not_Computer
Z-0174-2014 | 66597 | catheter, percutaneous, cardiac ablation, for treatment of atrial flutter | OAD |  |  | Cardiovascular |  | US Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and Cool Path Duo Ablation Catheter, MediGuide Enabled REF A700244, Rx only, Sterile EO. The catheters are intended for use with a compatible external irrigation pump and the IBI-15500T9-CP RF Generator at a maximum of 50 Watts. MediGuide Enabled Ablation catheters are used with the MediGuide Technology to enable real-time tip positioning and navigation and in conjunction the the EnSite Velocity System. The MediGuide Technology is indicated for use as an adjunct to fluoroscopy. | US Safire Duo Ablation Catheter MediGuide Enabledand Cool Path Duo Ablation Catheter, MediGuide Enab | 2 | November 06, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batches of Coolpath Duo Ablation Catheter MediGuide Enabled products and Safire Duo Ablation Catheter MediGuide Enabled. These units were inadvertently shipped with the incorrect version of the Instructions for Use (IFU). | PRODUCTION CONTROLS: Error in Labeling | The firm, St. Jude Medical (SJM), sent a "Urgent Medical Device Notice" dated 11 October 2013 to its consignees. The letter described the problem, the product involved in the recall and the actions to be taken. The consignees were instructed to discard the IFU received with the batches and replace it with the IFU received in conjunction with the notification letter. SJM has determined that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU.If you have any questions, contact Vice President, Quality, at 651-756-5402. |  | 59 | 59 | US distribution: KS, MA, UT,TX, KY, OH. | Open |  |  | 59 | Not_Computer
Z-0141-2014 | 66138 | [6 empty] | Harmony Equipment Management System, Model #603X/80XX, STERIS Corporation, Mentor, OHHarmony Equipment Management Systems (EMS) are used in hospital facilities to provide ready access to medical gasses, electrical power and audio/visual data services, in addition to managing surgical support equipment such as endoscopic equipment, cautery devices and insufflators. | Harmony Equipment Management System, Model 603X/80XX, STERIS Corporation, Mentor, OH | 2 | November 06, 2013 | 2013 | Steris Corporation | DrÂżger Medical GmbH has notified STERIS of the potential for a break in the drive screw contained in the motorized ceiling supply unit, a component of the Harmony EMS system. The ceiling supply unit controls the motorized adjustment of the height and rotation of the supply column. Should a break occur in the drive screw, there is potential for the adjustable arm to drop down, causing the EMS sys | DESIGN: Component Design/Selection | STERIS sent an Urgent Voluntary Field Correction Notice dated August 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS action- Drager Medical has provided STERIS instructions for inspection of the alignment of the mounting plate of the motor unit. If it is found that the mounting plate is bent, the drive screw within the motor unit must be replaced. STERIS will visit your facility to inspect the motorized ceiling supply unit of your Harmony EMS system(s) and if damage is observed, will replace the drive screw within the motor unit. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 127 | 127 units | Nationwide distribution: USA including the states of Alabama, Arkansas, Florida, Georgia, Illinois, Indiana, Kansas, Kentucky, Louisiana, Maryland, Massachusetts, Michigan, Minnesota, Missouri, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Virginia, Washington and Wisconsin. | Open |  |  | 127 | Not_Computer
Z-0173-2014 | 66132 | radioimmunoassay, parathyroid hormone | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH Calibrator (6 pack); Catalog Number: SMN 10492387For in vitro diagnostic use in calibrating Intact PTH assays on the ADVIA Centaur Systems. | ADVIA Centaur Intact PTH Calibrator | 3 | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | PTH Calibrator packaged with only high calibrator. | PRODUCTION CONTROLS: Packaging | Packaging error was brought to Siemens' attention through escalations received from (3) customers in the United States. Internal Investigation confirmed the issue was limited to the five (5) kits distributed to these customers. As the customers were aware of the issue and non affected product was provided to the customers to use during Siemens' investigation, no further communication was warranted. |  | 595 | 595 kits total (5 kits with packaging error) | Nationwide Distribution including MD, PA, and CT. | Terminated | February 13, 2014 | 99 | 595 | Not_Computer
Z-0143-2014 | 66249 | port, protector/cushion | OBK | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Dual Luer Lock CapThe device is intended for use as a cap for male or female Luer ports on medical devices such as manifolds, stopcocks or sets. | Dual Luer Lock Cap | 2 | November 06, 2013 | 2013 | Baxter Healthcare Corp. | The firm is recalling lots 10043 and 10044 due to loose particulate matter found in the packaging. Particulate matter entering the fluid path from the Luer lock caps may result in embolic events. |  | Baxter sent an Urgent Product Recall notification dated September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to :1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use.The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and returning it to Baxter by either faxing it to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 1-888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling. Please have the product code and lot number(s) available.4.If you are a dealer, wholesaler,or distributor/reseller that distributed any product to other facilities, please forward this communication. Baxter distributed this product to customers between June 19, 2013 and August 20, 2013.The Center for One Baxter is available at 1-800-422-9837, Monday throughFriday during the hours of 8:00 am to 5:00 pm Central Time. |  | 628992 | 628,992 units | Worldwide Distribution - USA (nationwide) and the country of Canada. | Open |  |  | 628992 | Not_Computer
Z-0178-2014 | 66331 | azo dye, calcium | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Arsenazo, CALA-0250, For the Quantitative Determination of Total Calcium in Human Serum. The kit label shows CALCIUM ARSENAZO, CALA-0250, R: 12x20 mL. The vial's label shows CALCIUM ARSENAZO, CALA-5220, 20 mL, .CALCIUM ARSENAZO is intended use for the quantitative in vitro diagnostic determination of total calcium in human serum and plasma on ELITech Clinical Systems Selectra Pro Series Analyzers. Calcium measurements are used in the diagnosis and treatment of parathyroid disease, a variety of bone diseases, chronic renal disease and tetany (intermittent muscular contractions or spasms). It is not intended for use in Point of Care setting. | Calcium Arsenazo, CALA0250 | 2 | November 07, 2013 | 2013 | SEPPIM SAS | Calcium Arsenazo Reagent, reference CALA-0250, is recalled due to complaint of plasma specimens report high results than serum. | DESIGN: Device Design | ELITech Group of Bothell, WA sent an Urgent: Product Recall Notice dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:" Laboratories must not use plasma for the determination of total calcium when using this product. The product can still be used to determine total calcium with serum sample." ELITech Clinical Systems SAS has revised the package insert and kit box label for the CALCIUM ARSENAZO reagent reference CALA-0250. All mentions of the use of the plasma will be removed from the documentation. If you have inventory of CALCIUM ARSENAZO, reference CALA- 0250. All lots. in your laboratory, please discard the insert that you have and replace it with the insert enclosed with this notice." If you have transferred any of this product to another location, please forward a copy this bulletin to that location." All users receive CALCIUM ARSENAZO, CALA-0250, from distributors. To be certain everyone who is possibly affected by this change is contacted, we are sending this communication to all registered users. As confirmation of receipt of this notice, please return the attached receipt confirmation to Technical Support by fax (401-642-9003) within 48 hours of receipt.We apologize for the inconvenience that this situation can cause in your laboratories; we remain at your disposal for further information. If you have any questions, please contact Technical Support at 1-855-ELITECH (855-354-324)We thank you for your cooperation and the attention which you will give to this bulletin. |  | 47 | 47 kits in US and 48 kits in Sweden | Worldwide Distribution - US including the states of KY and UT and the country of Sweden. | Terminated | December 12, 2013 | 35 | 47 | Not_Computer
Z-0181-2014 | 66267 | enzyme immunoassay, phenobarbital | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit 2000 Phenobarbital Assay and Syva Emit 2000 Phenobarbital Assay | Emit 2000 Phenobarbital Assay and Syva Emit 2000 | 2 | November 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Imprecision with certain lots | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated August 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product. Siemens will replace unused affected product at no charge. For replacements customers were instructed to incude this information on the Effectiveness Check form included with the letter. The Effectiveness Check form must be returned even if no replacement reagents are needed. The recall letter was to be retained with customers laboratory records and forwarded to those who may have received this product. Customers with questions were instructed to contact Siemens Technical Solutions or their Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 4963 | 4963 | Worldwide Distribution - USA (nationwide) and Mexico. | Open |  |  | 4963 | Not_Computer
Z-0176-2014 | 66191 | stimulator, muscle, powered, for muscle conditioning | NGX | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Pinook Mini Massager Distributed by www.pinookusa.com.Used to strengthen the abdominal muscles. | Pinook Mini Massager | 2 | November 07, 2013 | 2013 | Pinook-Usa | Failure to submit a premarket submission and gain approval of a medical device. | PREMARKET APPROVAL: No Marketing Application | On October 18, 2013 the firm sent the following notification to their customers:This is to inform you of a product recall involving: Pinooks Mini-massager and Pinooks Micro Massager model number PT-11 See enclosed product label (for ease in identifying the product at retail/user level). This recall has been initiated due to failure to submit a premarket submission and gain approval.We began shipping this product on 11/27/2012. Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Pinook USA is removing the products (mini and micro massager) from the market until receiving FDA approval to market them again under a new 510(k).If you have stock on your end please return them to our office:901 C FL PKWY STE A6, Orlando FL, 32824.Before returning, please contact Sheva Ben-Shushan: accounting@pinookusa.com in order to receive an RMA number.Once you have RMA number, place that number on the package label.This recall should be carried out to the retail level. Your assistance is appreciated and necessary to prevent patient harm. If you have any questions, call Dvir Lev-Ran, telephone number: 407-854-0055 Ext 304. This recall is being made with the knowledge of the Food and Drug Administration. |  | 28994 | 28,994 units (Mini and Micro) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the states of AZ, CA, CO, FL, IL, MA, MD, MO, MT, NC, ND, NJ, NV, NY, TX, VA, and WA and the country of Canada. | Open |  |  | 28994 | Not_Computer
Z-0175-2014 | 66475 | ph buffer | JCC | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | pH 7.0 Buffer Solution Pint and Quart Bottles, Part Number 02.0030 for Pints and 02.0031 for Quarts. Packaged in Plastic Bottles. Can be sold individually or in cartons with 6 bottles per box (for quart bottles only).Used in the hemodialysis setting to calibrate pH meters. | pH 7.0 Buffer Solution | 2 | November 07, 2013 | 2013 | Mesa Laboratories, Inc. | Mesa Laboratories, Inc. is recalling certain lots of pH 7.0 Buffer Solution due to mold growth. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | An email letter was sent to consignees on October 2, 2013 asking all customers to discard the product and a replacement product and tubing will be issued. |  | 58466 | 58,466 bottles | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open |  |  | 58466 | Not_Computer
Z-0191-2014 | 66145 | arthroscope | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek FMS Intermediary Tubing with One-Way valveProduct Code: 281142 Product Usage:The Intermediary Tubing for FMSÂŽ Fluid Management Systems connects the Inflow Tubing to the arthroscopic sheath during arthroscopic procedures. The Polyvinyl Chloride (PVC) tubing integrates several functional components including: luer locks for attaching the tube to the Inflow tubing and sheath, clamps, and a one-way valve. The FMS Intermediary Tubing is sterile and intended for single use with the FMS Fluid Management Systems in a surgical setting by personnel trained in arthroscopy. | DePuy Mitek FMS Intermediary Tubing with OneWay valve | 2 | November 08, 2013 | 2013 | DePuy Mitek, Inc., a Johnson & Johnson Co. | Pillow valve included in the FMS Duo outflow tubing set and the FMS Solo intermediary tubing set may not perform as intended. may result in backflow of irrigation fluid into the One Day Set which may potentially lead to patient cross-contamination | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | DePuy Mitek inititaed telephone notification on Ausgust 29, 2013 and sent an Urgent Voluntary Product Recall letter dated September 5, 2013 to US and OUS affected customers/affiliates. The customers were notified of the affected products, problem and actions to be taken. The customers were instructed to immediately check all inventories to locate and return affected product following the enclosed instructions. Stericycle is cooridianting the returns. If you have any questions or concerns in regards to this recall, please contact Stericycle directly at 1-866-737-1928. |  | 368 | 368 pk | Worldwide Distribution - USA (Nationwide) and in the countries of: Australia , Austria , Belgium , Canada , Czech Republic , Denmark , Finland , France, Germany , India, Ireland , Israel, Italy , Latvia , Netherlands , Norway , Nepal, Poland , Portugal , Spain , Sweden , Switzerland , United Kingdom. | Open |  |  | 29648 | Not_Computer
Z-0189-2014 | 66392 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library component (SHARP LCD screen, 55074-2 revision B)Product Usage: Intended to be used as a visual aid during the controlled administration of intravenous fluids. These fluids may include pharmaceutical drugs, blood, blood products and mixtures of required patient therapy. | Sigma Spectrum Infusion Pump component (LCD screen) | 2 | November 08, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation is recalling the LCD display screen on the SIGMA SPECTRUM Infusion Pump with Master Drug Library for not meeting standards for withstanding an electrostatic discharge. |  | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 24, 2013 to all affected customers via first class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate the affected product, remove them from use and acknowledge your receipt of the recall notification by completing the provided Customer Reply Form and fax to 224-270-5467. Customers were asked to contact Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday during the hourse of 7:00 am and 7:00 pm Eastern Time to order replacement device(s). For technical questions contact Baxter Healthcare Medina at 800-356-3454, (choose option1) and for clinical questions contact Medical Information Services at Baxter 800-933-0303. |  | 93 | 93 units | Worldwide Distribution - USA Nationwide in the states of: AR, AZ, CA, CO, CT, FL, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NM, NY, OH, OK, OR, PA, TN, UT, VA, WA and WV and the country Canada. | Open |  |  | 93 | Not_Computer
Z-0182-2014 | 66598 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Prostiva RF Therapy Hand Piece Kit, REF 8929, Rx only, sterilized using Ethylene Oxide.The Prostiva RF Model 8929 Hand Piece Kit is the delivery system component of the Prostiva RF Therapy System. Prostiva RF Therapy delivers low-levels of radiofrequency (RF) energy directly into the hyperplastic tissue for the treatment of benign prostatic hyperplasia (BPH). The RF energy delivered by the system provides selective thermal ablation of prostate tissue to help reduce urethralconstriction and relieve BPH voiding symptoms.The Prostiva RF Therapy System consists of the following major components:" Radio Frequency Generator" Hand Piece Kit" TelescopeThe Model 8929 Prostiva RF Therapy Hand Piece Kit contains the following single-use, disposable components:" One sterile Hand Piece with radiofrequency cable (Model 8929)" One sterile Tubing System (Model 6101)" One non-sterile Return Electrode (Model 8934)The hand piece has two insulated electrode needles that deliver the radio frequency energy interstitially to the prostate. These needles are covered by insulating shields that localize the lesions within the prostate, protect the urethra from thermal damage, and eliminate collateral thermal damage. The Hand Piece stainless steel (SS) tube is inserted into the urethral canal up to the prostate with the aid of a telescope for visualization during the therapy. The Hand Piece features two insulated Nitinol needles inside the SS tube. The needles can be deployed (extended), at approximately a 90-degree angle from the SS tube, to a preset depth of between 12 mm and 22 mm (at 2 mm intervals). The needles inside the SS tube are deployed into, or retracted from the prostate, using the handle lever. RF energy is delivered through the needles to the prostate tissue. The RF energy creates molecular friction eventually resulting in tissue heating and cell death (therapy intent). | Prostiva RF Therapy Hand Piece Kit | 2 | November 08, 2013 | 2013 | Urologix, Inc. | Urologix is voluntarily recalling Prostiva RF Therapy Hand Piece Kits, Part Number 8929, Lot Number 215251 due to a Hand Piece not meeting an in-process force test specification. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Consignees were sent a Urologix Urgent Medical Device Recall letter dated October 21, 2013. The letter described the problem and the device involved in the recall. Advised consignees to examine their inventory and remove the recalled product. Requested consignees to complete and fax the "Customer Acknowledgement Form" to Hope Przybilla at 763-475-1443 or call her for questions. The consignees local representatives will be personally assisting them with the return process and ensuring that replacement product is offered and received promptly. |  | 142 | 142 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, KY, MA, MN, NC, NJ, NV, NY, OH, OR, TN, TX, and UT. | Open |  |  | 142 | Not_Computer
Z-0195-2014 | 66624 | catheter, cannula and tubing, vascular, cardiopulmonary bypass | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns" Soft-FlowÂŽ Aortic Cannula without Suture Flange, straight tip, wire-reinforced with luer port, aortic cannula, 8.0 mm (24 Fr) OD with 3/8" connector, 14" (36 cm) long.Indicated for use in perfusion of the ascending aorta during cardiopulmonary bypass surgery. | Sarns" SoftFlow Aortic Cannula without Suture Flange | 2 | November 08, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | During an in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of loose fiber particulate that exceeded finished product specifications on certain product lots of Sarns" Soft-FlowÂż Aortic Cannulae and Sarns" Venous Return Cannulae. | TRAINING: Employee Error | On 10/23/13, Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE RECALL LETTER to their consignees. The letter identified the reason for the device removal, identified the affected product and associated potential hazards, and provided recall instructions to their customers regarding affected product return. Questions regarding this recall are directed to CVS Customer Service M-F, 8am-6pm at 1-800-521-2818. |  | 30 | 30 units | Worldwide Distribution-USA including DC and the states of MO,TX,OK, LA, MA, CA, WI, and MI and the country of Canada. | Open |  |  | 230 | Not_Computer
Z-0188-2014 | 66726 | monitor, physiological, patient(with arrhythmia detection or alarms) | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare elance Vital Signs Monitor, Model 93300, with Option S. It is a patient monitor with arrhythmia detection or alarms. | Spacelabs Healthcare elance Vital Signs Monitor, Model 93300, with Option S | 2 | November 08, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare elance Vital Signs Monitor with Option S, model 93330, is recalled because the monitor will not permit connection of any of the Spacelabs TruLink line of SpO2 sensors. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare elance Vital Signs Monitor with Option S" letter, dated 01 November 2013, to the US consignee on November 1, 2013.Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013.USERS ARE ADVISED THAT:-They may not be able to monitor their patient's oxygen saturation because the monitor will not allow the TruLink SpO2 sensor to be connected. The clinician will be required to find an alternate method to monitor oxygen saturation.-They will be contacted at the earliest possible date to schedule a convenient time and Spacelabs will correct all of your facility's affected monitors at no cost.Customer in the US can call Spacelabs Healthcare, Inc. of Snoqualmie, Washington at 800-522-7025.If you have any questions about this corrective action program, please contactSpacelabs Healthcare at +86 512 87178304 for Technical Support. |  | 89 | 89 units total (1 in US and 88 International) | Worldwide Distribution-Distributed in Puerto Rico and the countries of France, Indonesia, Korea, Nicaragua, Philippines, and Taiwan. | Terminated | February 25, 2014 | 109 | 89 | Not_Computer
Z-0190-2014 | 66546 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Sterrad 100NX, Product Code: 10104.Used to sterilize heat and moisture sensitive reusable medical devices. | STTERRAD 100NX Sterilizer | 2 | November 08, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the Sterrad 100NX Sterilizer some healthcare workers detected odors/smells emanating from the 100NX Sterilizers, and it may be related to normal breakdown of the oil used in the Sterrad 200 vacuum pumps. |  | A recall letter was sent to all customers on 10/18/13 who purchased the Sterrad100NX Sterilizer. The letter informed the customers of the problems identified and the actions to be taken. Customers are instructed to contact ASP Professional Services directly at (888) 783-7723 and select option #2. Customers are instructed to report any complaints or suspected problems with their Sterrad 100NX System to Stericycle at (877) 598-5705. |  | 2106 | 2106 units total (1224 units in US) | Worldwide Distribution-USA (nationwide) and the countries of EMEA, APAC, Japan, Canada. AE, AU, BE, BH, CA, CH, CL, CN, CO, DE, EC, ES, GB, GR, HK, HU, IL, IN, IT, JP, KR, KW, MX, MY, NX, PA, PH, PK, PL, PR, QA, RU, SA, SG, SI, TH, VE, VN, ZA, and Tripoli. | Terminated | January 16, 2014 | 69 | 2106 | Not_Computer
Z-0186-2014 | 65739 | hemodialysis system for home use | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Combiset Blood Line Hemodialysis useCatalog Number: 03-2622-3.Intended for use as the extracorporeal blood circuit during Hemodialysis and is indicated for use with conventional and high flux negative pressure hemodialyzer equipment. | Combiset Blood Line | 2 | November 08, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Updated Instructions for Use: Bloodline connections not tightened per instructions in the operator's manual may disconnect and lead to blood loss | DESIGN: Labeling Design | Fresenius issued a Field Notification Letter dated 8/29/13 to Patient Clinics/ associated at home patients and provides additional detail instructions to ensure the connection to the transducer protector bloodline female luer interface is secure.This field advisory notice sent by Certified Mail, Return Receipt Requested. Accounts requested to return both the receipt and the completed reply form (enclosed) to confirm receipt of the notice.If you have any questions regarding this field advisory, please contact FMCNA Customer Service Team at 1-800-323-5188. Customer Service will be available Monday-Friday 7am-6pm CST. |  | 508029 | 508,029 cases | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open |  |  | 540759 | Not_Computer
Z-0187-2014 | 66574 | catheter, percutaneous, cardiac ablation, for treatment of atrial flutter | OAD |  |  | Cardiovascular |  | Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and A700238, Rx only, Sterile EO.Intended for creating endocardial lesions during cardiac ablation procedures (mapping, stimulation and ablation) for the treatment of typical atrial flutter. | Safire Duo Ablation Catheter MediGuide Enabled | 2 | November 08, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batched of Safire Duo Ablation Catheter MediGuide Enabled. These units have a cable connector configuration that is different than intended. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Consignees were sent a St. Jude Medical "Urgent Medical Device Notice" dated 11 October 2013. The letter was addressed to Cath lab manger / Risk Manger. The letter described the problem and the product involved in the recall. The letter stated that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU. |  | 46 | 46 | Nationwide Distribution including the states of KY, OH, TX, MA, and NC. | Open |  |  | 46 | Not_Computer
Z-0200-2014 | 65863 | drills, burrs, trephines & accessories (simple, powered) | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Product Catalog (part number 20-0020) utilized with the Electric and Pneumatic Motor Systems including: MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, and Emax2Plus Motor Systems, all models and serial numbers.Product Usage: Indication for use is cutting and shaping bone. | BlackMax Motor Systems, MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, Emax2Plus M | 2 | November 13, 2013 | 2013 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL initiated a correction of its medical Device labeling for its Pneumatic Systems Operating Manual, Cutting Burr Chart and Product Catalog due to a recent review of their regulations and compliance requirements that determined the claim that the device is capable of dissection of biometals, bioplastics and cement has been verified, but is is not part of th | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | The firm, Anspach Effort, Inc., sent an "Urgent Medical Device Labeling Correction" letter dated June 27, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Destroy or delete any items in their inventory immediately and replace with the attached revised Operating Manual, Product Catalog and/or Cutting Bur Chart. 2) Complete and return the attached Customer Reply Form indicating disposition of the labeling materials and confirming their receipt of this letter, and fax to 1-800-327-6661. 3) If they distributed any of the products to other services or facilities, please forward this information as appropriate. For additional copies, contact Anspach Customer Support at (800) 327-6887. |  | 2020 | 2,020 | Worldwide Distribution: USA Nationwide and the countries of: Egypt, India, Germany, Turkey, Canada, Saudi Arabia, Greece, United Kingdom, Cayman Islands, Kazakhstan, Republic of Korea, Taiwan, Hong Kong, China, Lebanon, Argentina, Denmark, Belgium, Netherlands, Mexico, France, Australia, Panama, Italy, Guam, United Arab Emirates, Israel, Polan, Venezuela, Austria, Hungary, Lithuania, Norway, Malaysia, Colombia, Costa Rica, Ecuador, Jamaica, Peru, Thailand, Italy, Czech Republic, Cameroon, South Africa, Portugal, Finland, Sweden, Nigeria, Bahamas, Cambodia, Dominican Republic, Switzerland, Chile, Singapore, Spain, Syrian Arab Republic, Jordan, Japan. | Open |  |  | 2020 | Not_Computer
Z-0272-2014 | 66689 | saliva, artificial | LFD |  |  | Dental | 510(k) | Biotene Moisturizing Mouth Spray, with xylitol, For dry mouth symptom relief, 44.3ml/1.5 FL. OZ., Sugar Free Sweetened with xylitol and sorbitol, GlaxoSmithKline, Consumer Healthcare, L.P., Moon Township, PA 15108.Biotene Mouth Spray is used for the treatment of dry mouth | Biotene Moisturizing Mouth Spray | 2 | November 13, 2013 | 2013 | GlaxoSmithKline, LLC. | Label indicates, "Sweetened with xylitol and Sorbitol", however product does not contain sorbitol. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | GlaxoSmithKline sent an Urgent Medical Device Recall letter dated October 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product. If product was further distributed customers were instructed notify their customers of the recall and the actions required. Customers were asked to complete and return the enclosed business replay card even if they have no affected product in their inventory. Customers should complete the return goods packing slip, enclose it with their return goods shipment and affix the return label to the package. Customers with questions should call Customers Service at 1-800-743-4014.For questions regarding this recall call 919-269-1194. |  | 483204 | 483,204 bottles | Nationwide Distribution . | Open |  |  | 483204 | Not_Computer
Z-0210-2014 | 66449 | holder, head, neurosurgical (skull clamp) | HBL | Neurosurgical head holder (skull clamp). | Neurology | Neurology | 510(k) | Composite Series Skull Clamp, one clamp per tote. The MAYFIELDÂŽ Skull Clamp (A3059) is a cranial stabilization device, designed to provide rigid skeletal fixation. The Skull Clamp is placed on the patient's skull to hold their head and neck securely in a particular position when rigid fixation is desired. The clamp is indicated for use in open and percutaneous craniotomies as well as spinal surgery when rigid skeletal fixation is necessary | MAYFIELD Composite Series Skull Clamp | 2 | November 13, 2013 | 2013 | Integra LifeSciences Corp. | An investigation of an adverse trend of complaints for the Skull Clamp index locking knob not easily unlocking or disengaging and malfunctions of the 80 lbs. torque knob was initiated. None of the complaints resulted in a report of patient injury, adverse health consequence, or a delay / prolongation of the surgical procedure. | DESIGN: Device Design | Integra sent an Urgent: Voluntary Medical Device Recall Notification dated August 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your local Integra NeuroSpecialist will set up a time with you to discuss replacement options but in the interim, Integra is asking that you do the following:1. Review your inventory and determine if you have any MAYFIELDÂŽ Composite Series Skull Clamps (A3059). If so, stop using them immediately and remove them from service to prevent any inadvertent use of the Composite Series Skull Clamp.Either while the Integra NeuroSpecialist is on-site or independently;2. Complete the attached form. If you do not have affected product, check the box: I do not have Composite Series Skull Clamps.3. If you do have affected product, check the box: I do have Composite Series Skull Clamps. Record the total quantity and Serial No.(s) of all Skull Clamps in your inventory. The Serial No. is located on the clamp upright comprised of two letters (SS) followed by six numerals.4. Complete the other information and return by email or fax as indicated on the form, or give it to the Integra NeuroSpecialist who will return it for you.For assistance returning and exchanging these products or any other questions that you may have, please contact Integra Customer Service at 1-855-532-1723.We sincerely apologize for any inconvenience this may have caused and thank you for your cooperation. |  | 37 | 37 units | US Distribution including the states of AL, AR, CA, CT, FL, GA, IA, IL, MS, NY, OH, SC, TN, and TX. | Terminated | January 14, 2014 | 62 | 37 | Not_Computer
Z-0202-2014 | 66485 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 8.5 MM Medullary Reamer Product Usage:An orthopedic manual surgical instrument is a nonpowered hand-held device intended for medical purposes to manipulate tissue, or for use with other devices in orthopedic surgery. This generic type of device includes the cerclage applier, awl, bender, drill brace, broach, bmT, corkscrew, countersink, pin crimper, wire cutter, prosthesis driver, extractor, file, fork, needle holder, impactor, bending or contouring instrument, compression instrument, passer, socket positioner, probe, femoral neck punch, socket pusher, reamer, rongeur, scissors, screwdriver, bone skid, staple driver, bone screw starter, surgical stripper, tamp, bone tap, trephine, wire twister, and wrench. | Synthes 8.5 MM Medullary Reamer | 2 | November 13, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received for intraoperative reamer head breakages which can allow for unretrieved fragments of non-implant grade material. | DESIGN: Device Design | Synthes sent an Urgent Notice: Voluntary Medical Device Recall Notification letters, dated October 3, 2013, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected lots from stock. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 2979 | 2979 | Worldwide Distribution: USA (Nationwide) and countries of: Canada and Mexico. | Open |  |  | 2979 | Not_Computer
Z-0283-2014 | 64263 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Hyperbaric Infusion System is a cassette based multifunction infusion system that allows two lines in and one line out and can be used for standard, piggyback, or concurrent delivery. The Plum A+ Hyperbaric Infusion System is designed to deliver parenteral, enteral, or epidural infusions over a broad range of infusion rates from multiple fluid container types. The Plum A+ Hyperbaric Infusion System includes an infuser, disposable IV sets (hereafter called a set), optional accessories, and this operator's manual.The Plum A+ Infusion System with Hospira MedNet Software is indicated for use in parenteral, enteral and epidural therapies and the administration of whole blood and blood products. | Plum A Hyperbaric Infusion Pump | 2 | November 14, 2013 | 2013 | Hospira Inc. | Plum A+ infusers have the potential for the distal (occlusion) press sensor pin to break. The distal pressure sensor pin is a part of the overall subsystem that measures the pressure within the distal line of the administration set and indicates the presence of a full or partial distal occlusion. A broken distal pin can only be detected via a visual inspection of the distal pressure pin. A broken | DESIGN: Component Design/Selection | Hospira sent an Urgent Device Field Correction letter dated February 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter recommends that facilities immediately visually inspect their Plum a+ devices to determine if the distal pressure pin is broken or damaged via instructions included in the letter. If a broken or damaged distal pressure pin is observed, remove the device from service and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM, CST) to report the issue. If the return of defective pumps significantly impairs a facility's ability to operate normally, Hospira will provide loaner pumps. The letter also encourages customers to insert cassettes into pumps in accordance with the instructions found in the operator's manual. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to call 1-800-441-4100 or 1-800-241-4002, option 4.The letter also asked direct accounts that further distributed the pumps notify their customers of the URGENT DEVICE FIELD CORRECTION and ask them to contact Stericycle at 1-866-201-9068 to receive a reply form.. For questions regarding this recall call 800-441-4100,. |  | 179 | 179 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Costa Rica and Dominican Republic. | Open |  |  | 272558 | Not_Computer
Z-0280-2014 | 66444 | mixer, cement, for clinical use | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Acetabular Pressurizer; single use and six units per box.The Acetabular Pressurizer 0206-522-000 is an accessory to theAdvanced Cement Mixer (ACM). The Acetabular PressurizerTiplet P/N 0206-522-012 is sealed in a two pouches, sterilized, and labeled as part number 0206-522-000.This Acetabular Pressurizer is an accessory used to pressurize the cement applied to the prosthesis to be implanted during the surgery. The Acetabular Pressurizer is an injection molded component made with resin material DOW HDPD 10462N which is a rigid plastic once it is molded. | Acetabular Pressurizer 0206522000 | 2 | November 14, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | On March 14, 2013 a Work Order was opened for 120 units (20 boxes) of the Acetabular Pressurizer (P/N 0206-522-000 Rev. H). On March 15, 2013, when the product was going to be placed in the pouch to be sealed, the manufacturing operator noticed that some assemblies were cracked. As a result of the injection pressure of the cement gun, a cracked tiplet may release from the Hub during cement applica | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Stryker Instruments, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/01/2013, to all domestic consignees/customers and an email was sent to the International Subsidiaries and Sales Representatives. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the Business Reply Form even if youdon't have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in this Notification.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Acetabular Pressurizers to Stryker.7. Upon receipt of recalled Acetabular Pressurizers, credit will be applied to your account.For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2921 or email: kellyjo.whipple@stryker.com. |  | 87 | 87 boxes (522 each) | Worldwide distribution: US (nationwide) including states of: TX, FL, IL, IA, ND, WI, VA, and NJ.; and countries of: CANADA, POLAND and SWITZERLAND. | Terminated | February 10, 2014 | 88 | 87 | Not_Computer
Z-0291-2014 | 66397 | analyzer, chemistry (photometric, discrete), for clinical use | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System, Part Number: 81600NThe Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access 2 Immunoassay System | 2 | November 14, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Immunoassay System because they determined that the fuses in the system may be of the wrong amperage. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | A phone call was made on 9/12/13 to customer who received the Access 2 System. The phone script explains the problem identified of a possible wrong fuse and may result in failure of the fuse to operate properly. Customer was instructed that their Field Service Engineer, will call them back the following week to check on their system and replace the fuses if necessary. If you have any questions concerning this information, please call 952-393-8546. |  | 8 | 8 units total (1 unit in the US) | Worldwide distribution: US (nationwide) in the state of: TX and countries of: Colombia, Hong Kong and Korea. | Terminated | November 14, 2013 | 0 | 8 | Not_Computer
Z-0277-2014 | 66199 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂŽ Cemented Revision Calcar (CRC) Hip System Femoral StemsVerSys Heritage Hip System Revision Femoral Stems.Indicated for total hip arthroplasty. | VerSys Cemented Revision Calcar (CRC) Hip System | 2 | November 14, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged in one of the clean rooms at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing (seal integrity and seal strength) to provid | PRODUCTION CONTROLS: Packaging Process Control | Removal of the subject devices from the field including distributor and hospital inventories. Distributors and Risk Managers and Surgeons at hospitals that received a direct shipment from Zimmer or its Distributors received a notification on 09/19/2013. |  | 12866 | 12866 units | Worldwide distribution: US (Nationwide) including Puerto Rico and Virgin Islands; and countries of: Australia, Algeria, Belgium, Bolivia, Canada, Canary Islands, Chile, Croatia, Finland, France, Germany, Greece, Guatemala, Honduras, India, Israel, Italy, Japan, Jordan, Lima-Peru, Malaysia, Morocco, Nicaragua, Portugal, Saudi Arabia, Seoul Korea, Serbia, Singapore, South Africa, Spain, Switzerland, Turkey, Utd.Arab.Emir., United Kingdom, and Venezuela. | Open |  |  | 12866 | Not_Computer
Z-0282-2014 | 66537 | table, operating-room, ac-powered | FQO | Operating tables and accessories and operating chairs and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Draeger Movita 603x/DVE 803x Ceiling Supply Unit. It is intended for use in operating theaters and intensive care | Draeger Movita 603x/DVE 803x Ceiling Supply Unit | 2 | November 14, 2013 | 2013 | Draeger Medical, Inc. | Cases were encountered with the Draeger Movita 603x/DVE 803x Ceiling Supply Units where the drive screw in the motor unit broke and the lift arm fell down approximately 60 cm (23.6 inches). | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | An "Urgent Medical Device Recall" notification letter was sent to users in September 2013. The issue was described and recommended actions were provided. Customers may contact (800) 543-5047 and select 1, then 2. |  | 139 | 139 | Worldwide distribution: US (Nationwide)and in the country of Canada. (note: the ceiling supply units are not classified as a medical device by Health Canada). | Open |  |  | 139 | Not_Computer
Z-0204-2014 | 64091 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Infusion Pump; List Number: 11971; Product Usage:The Plum A+ Infusion Systems are a multipurpose family of devices that provides precise delivery of multiple therapies across the general spectrum of clinical care | Plum A Infusion Pump | 2 | November 14, 2013 | 2013 | Hospira Inc. | Hospira has received reports of distal pressure sensor calibration drift on Plum A+ Infusers. The following error codes may be displayed during setup or infusion signifying possible calibration drift: E180/N180, E181/N181, E186/N186, E187/N187 or E346. Additionally, if the pressure sensor calibration has drifted, users may experience false distal occlusion alarms or distal occlusions that may not | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Hospira sent an URGENT DEVICE FIELD CORRECTION letter dated February 5, 2013 was sent to all direct accounts notifying them of the issue. For device users who experience these issues, Hospira recommends the following actions be taken: 1) Remove the pump from service and perform a Performance Verification Test (PVT) Distal Occlusion Test to determine if the distal pressure sensor is performing correctly. 2) If the device does not pass the PVT, it may indicate that the distal pressure sensor has drifted. Perform the appropraite troubleshooting and repair activities defined by the user's facility, whihc may include returning the device to Hospira for further diagnosis and servicing. The remediation for this issue is to perform a recalibration of the distal pressure sensor. Hospira also intends to release in 2013 a recommendation of a yearly test to determine if the distal pressure sensor's calibration has drifted. Customers who suspect that their device(s) may be affected by this issue are urged to contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue. Hospira will provide loaner devices if necessary. Customers are requested to complete and return the reply form included with the notification letter per the instructions on the reply form.For questions regarding this recall call 224-212-2000. |  | 22640 | List No. 11971: 10,669 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Croatia, Dominican Republic, Turks and Caicos. | Open |  |  | 10669 | Not_Computer
Z-0274-2014 | 65760 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Extension ShimsExtension shims are manufactured from grade 5 titanium and are curved, rectangular slabs which have four protrusions used for ratcheting. They come in straight and pointed varieties, either with or without a tab containing an additional bend. Extension shims are designed to slide down retractor blades, and are optional retractor components utilized during lumbar interbody fusion procedures. They are intended to manipulate tissue and prevent tissue from interfering inside the surgical field. | SLIFT Extension Shims | 2 | November 14, 2013 | 2013 | SpineFrontier, Inc. | Fit variation among all lots of S-LIFT extension shims with corresponding retractor blades. The firm's investigation revealed that extension shims may not fit down all retractor blades. | DESIGN: Device Design | SpineFrontier sent an Urgent Voluntary Recall letter dated August 28, 2013, to all affected customers. The letter identified teh product, the problem, and the action to be taken by the customer. Customers were asked to initiate the immediate return of the parts noted and complete an acknowledgement form and return it to SpineFrontier via fax or email. Customers with questions were instructed to call 978-232-3990. For questions regarding this recall call 978-232-3990. |  | 67 | 67 | Nationwide Distribution including FL, TX, OK, MO, IL, OH, VA, OR, CO, KS, CT | Open |  |  | 67 | Not_Computer
Z-0276-2014 | 66450 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5MM LOCKING SCREW, T15 HEXALOBE, SELF TAPPING, 34MMThe OrthoPediatrics Proximal Femur Plating System combines implants and instruments in one convenient system. This system offers the advantages of the osteotomy blade plates and cannulated instrumentation. Osteotomy plates provide ease of reduction and good rotational stability while maintaining bone stock. Cannulated instruments work over a guide wire for precise placement and safety. The subject device is a Bone Screw designed for use with this system. | 3.5MM LOCKING SCREW, T15 HEXALOBE, SELF TAPPING, 34MM | 2 | November 14, 2013 | 2013 | OrthoPediatrics Corp | The firm identified, through complaints that the 3.5MM LOCKING SCREW packaging material was marked as a 3.5mm screw and the part number was correct but that the description indicated a 7.0mm screw. Upon investigation, it has been determined that a lot number 009U2 was packaged and mislabeled internally by OrthoPediatrics. | PRODUCTION CONTROLS: Labeling Mix-Ups | OrthoPediatrics sent an Urgent Medical Device Recall letter dated September 26, 2013, to all affected customers.The letter identified the product the problem and the actions needed to be taken by the customerThe sales representatives and distributors who have been identified via shipping records as having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on 18-September-2013. If you or any of your customers have any of these devices in the original packaging please take possession of them and quarantine them in preparation for returning them to OrthoPedlatrics. They are misbranded. If the Screw has already been removed from the original packaging and the original packaging has been discarded and the Screw placed in a sterilization tray, it need not be returned.For further questions please call ( 574) 268-6379. |  | 7 | 7 units shipped | Worldwide Distribution - US including Michigan and the countries of Ireland and Australia | Terminated | December 26, 2013 | 42 | 7 | Not_Computer
Z-0275-2014 | 66608 | sponge, ophthalmic | HOZ | Ophthalmic sponge. | Ophthalmic | Ophthalmic | 510(k) | DeRoyal (R) Surgical Eye Spears, REF 30-049, Sterile EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, MADE IN GUATEMALAcleaning of fluids from surgical site in or near the eye | DeRoyal (R) Surgical Eye Spears | 2 | November 14, 2013 | 2013 | DeRoyal Industries Inc | One lot of eye spears was labeled as being x-ray detectable, but contained non x-ray detectable product. | PRODUCTION CONTROLS: Error in Labeling | DeRoyal Industries sent an Urgent Voluntary Recall letter dated August 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease usage and return the product to:DeRoyal RecallsRA#: Eye Spears1755 Hwy 33 SouthNew Tazewell, TN 37825Ship via UPS Ground using Shipper #312780 Distributors were requested to notify their consignees of the same. The recall is extended to the user level. Customers with questions were instructed to call 1-800-251-9864.For questions regarding this recall call 865-362-2334. |  | 726 | 726 cases (7242 units) | Nationwide Distribution | Open |  |  | 726 | Not_Computer
Z-0332-2014 | 66547 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | ALTRUSÂŽ THERMAL TISSUE FUSION, Tissue Fusion Handpiece, 5 mm O.D., 36 cm Shaft, REF 60-9522-001, STERILE -- CONMED CORPORATION 525 FRENCH RD, UTICA, NY 13502 --- The ConMed AltrusÂŽ Thermal Tissue Fusion System is comprised of a dedicated energy source and disposable handpiece used to ligate (seal) and divide (cut) blood vessels and tissue bundles that fit into the jaws of the handpiece. The AltrusÂŽ System utilizes a thermal energy platform to achieve the desired clinical effect. The ConMed AltrusÂŽ Thermal Tissue Fusion handpiece is a single use device which is provided sterile. The device uses scissors to parallel jaw closure mechanism with one flat jaw and one crowned jaw in which the vessels/tissues are grasped and through which pressure and heat are applied. It is a multi-functional device capable of vessel sealing, grasping and dissecting during open general surgical procedures. The two jaw sizes allow for applications in a variety of clinical environments for access in confined areas as well as large open areas. | ALTRUS | 2 | November 15, 2013 | 2013 | ConMed Corporation | The affected lots were exposed to a higher than acceptable level of particulate matter during manufacture and subsequently distributed. ConMed evaluated the potential risk that this issue could present to the patient and concluded the risk of injury to be negligible; however exposure could cause a minor, local, self correcting infection. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, ConMed, sent an "URGENT: VOLUNTARY DEVICE RECALL" letter dated October 23, 2013 and Business Reply Forms to the domestic consignees/customers via USPS Priority Mail. Distribution of the recall letter to the foreign consignees will follow once the letter has been translated into the different foreign languages. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop the use of these devices; review your inventory for any of the devices listed; contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to; complete and return the ATTACHMENT I Effectiveness Check Business Reply Form via fax to: ConMed Recall Coordinator at 315-624-3225.If you have any of the affected devices, please complete Attachment I and return it with the devices to: ConMed Corporation, 525 French Road, Utica, NY 13502Attn. Ed Kovac, Return via: FedEx Account # 487553646. Please indicate with your returned devices, if you are requesting credit or replacement of the devices. Please do not return used devices. If you do not have any devices to return, please complete Attachment I, indicating you have no devices and fax it to 315-624-3225, Attn: ConMed Recall Coordinator. For questions, please contact ConMed Recall Coordinator at 315-624-3237 or e-mail altrus@conmed.com. |  | 147 | 147 handpieces (Domestic: 135 devices; Foreign: 12 devices) | Worldwide distribution: US (Nationwide) and countries of: Canada, Germany and Italy. | Open |  |  | 212 | Not_Computer
Z-0335-2014 | 66601 | system, x-ray, tomography, computed | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, Product Usage:The Brilliance iCT is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Brilliance iCT Computed Tomography XRay System, | 2 | November 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the couch, with the footrest extension attached, is positioned where the table travel is within the bore of the gantry, and the user begins to lower the couch, the footrest extension can contact the gantry cover or come in contact with the user due to the length of the accessory. |  | Philips sent an Urgent - Medical Device Correction letter dated October. 11, 2013, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue the use of the footrest extension. Philips also, stated they will replace the old footrest extension with the current one in use. For questions contact your local Philips representative or local Philips Healthcare office. |  | 9 | 9 units | Worldwide distribution: USA (nationwide) Distribution in the states of DC, GA, IL, IN, NY and OH and the countries: Germany, Israel and Netherlands. | Open |  |  | 9 | Not_Computer
Z-0268-2014 | 66567 | occluder, catheter tip | DQT | Catheter tip occluder. | Cardiovascular | Cardiovascular | 510(k) | ZINGERÂŽ Steerable Guidewire; Models: LVZRLS180J, LVZRLS180S, LVZRMS180J, LVZRMS180S, LVZRXT180J, LVZRXT180SMedtronic guide wires are steerable guide wires that are used for the introduction and placement of diagnostics or interventional devices in the coronary and peripheral vasculature, and may be used to reach and cross a target lesion. Medtronic guide wires are not intended for cerebral vasculature. Medtronic steerable exchange guide wires are used to facilitate the substitution of one diagnostic or interventional device for another. | ZINGER Steerable Guidewire | 1 | November 15, 2013 | 2013 | Medtronic Vascular | Potential for PTFE (polytetrafluroethylene) coating to delaminate and detach from Guidewire. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | The firm, Medtronic, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 21, 2013 (in the US via UPS overnight service, 1-day delivery) to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately remove and quarantine all potentially affected guidewires that remain in inventory; return the affected product to Medtronic; if replacement product is needed, your Medtronic representative can assist you with identifying suitable replacement product, and to complete and return the Customer Confirmation Certificate confirming receipt of the notification via fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality (100% customer confirmations will be required).Should have any questions, please contact your Medtronic Representative or call 763-526-2513.Notifications to customers concerning the additional affected product began December 19, 2013 (in the US via UPS overnight service, 1-day delivery). Customers are requested to return a Confirmation Certificate confirmingtheir receipt of the notification. 100% customer confirmations will be required from the additional consignees. |  | 5 | 5 | Worldwide distribution: US (Nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong,Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Mexico, Montenegro, Nepal, Netherlands, New Zealand, Nepal, Nicaragua, Norway, Oman, Poland, Portugal Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Uruguay, and Yemen. | Open |  |  | 14431 | Not_Computer
Z-0305-2014 | 65572 | differential rate kinetic method, cpk or isoenzymes | JHS | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products CK Slides packaged as 300 slides/pack catalog number 8479396 and packaged as 90 slides/pack catalog number 8478034VITROS CK Slides quantitatively measure creatine kinase (CK) activity in serum and plasma using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. | VITROS Chemistry Products CK Slides | 2 | November 15, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling Slide Dispense cartridges used with some VITROS Chemistry Product Slides due to slide dispense issues. | PRODUCTION CONTROLS: Equipment Maintenance | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 12, 2013, to all affected customers via FedEx regarding Slide Dispense issues with various VITROS Chemistry Product Slides. VITROS ALT, CL-, GGT, LDH, K+, Na+ and TP slides use a different ABP, they are not affected by this issue. OCD provided instructions that may allow their customers to reposition the ABP if a jam occurs. OCD will provide you with a credit for any cartridge where the ABP becomes misaligned within the cartridge body in such a way that you are unable to reposition the ABP successfully to resolve the slide dispense issue. In addition to the Confirmation of Receipt form, OCD also enclosed three (3) copies of a Request for Credit form for you to complete when you require credit for affected cartridges. Complete and return the Confirmation of Receipt by June 21, 2013. If you experience slide dispense condition codes more than once a week, call our Customer Technical Support at 1-800-421-3311.For questions regarding this recall call 585-453-4224. |  | 86503 | 86,503 units (30,378 units domestically & 56,125 units internationally) | Worldwide Distribution - USA (nationwide) including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY including Puerto Rico and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, England, Germany, France, Italy, and Spain. | Open |  |  | 86503 | Not_Computer
Z-0339-2014 | 66763 | dura substitute | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Dura Guard, Dural Repair Patch with Apex Processing.For use as a dura substitute for the closure of dura mater during neurosurgery.Product Code:DG0209SN DG0209SN DG0209SN DG0209SN DG0209SN DG0209SN DG0209SN DG0209SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0404SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN DG0608SN | DuraGuard Dural Repair Patch | 2 | November 18, 2013 | 2013 | Synovis Surgical Innovations, Inc. | Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because the shipping configurations in which the 6 fluid ounce units containing a specific revision of the Instructions for Use (IFU) shipped in a divided cell shipping | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Packaging Change Control | A customer notification letter dated 10/28/13 was sent to all customers who purchased the Vascu Guard and the Dura Guard. The letter informs the customers of the problem identified and the actions to be taken. Customers with any questions regarding the recall are instructed to contact Customer Service unit at (800) 423-2090. If customers are distributors or if they have have distributed the recalled products to any other health care provider, they are instructed to notify their accounts and provide them with a copy of the recall notification. If customers would like Baxter to perform the recall task for them, then they are instructed to contact Stericycle at 1-866-912-8417. |  | 446 | 446 units | Nationwide Distribution. | Open |  |  | 6460 | Not_Computer
Z-0354-2014 | 66558 | computer assisted hair harvesting system | ONA | Stereotaxic instrument. | Neurology | General & Plastic Surgery | 510(k) | ARTAS Disposable Kit ( Contains Skin Tensioner FP-45384; Inner Needle with Vacuum Tube FP-45385; Dissection Punch 4.0mm FP-23304; Saline Nose Cone FP-45387; Follicle Trap - FP-23306 and Non-sterile Disposable Clinical Kit FP -23267):and ARTAS Disposable Kits, 18 Gauge Epsilon (Contains Disposable Kit, 18 Gauge Epsilon FP 45547 and Non-Sterile Disposable Kit FP-23267).RX. Product Usage:Skin Tensioner is applied to patients scalp during follicular harvesting during hair restoration procedure. Product applies tension to scalp to facilitate follicular extractions. | ARTAS System Disposable kit | 2 | November 18, 2013 | 2013 | Restoration Robotics | Reports of discoloration (residue) on the pins of the skin tensioner have been received by several customers. Discoloration due to insufficient cleaning of stainless steel pins during passivation process. The residue hasbeen identified as a compound of iron and chlorine. | PRODUCTION CONTROLS: Process Control | Restoration Robotics sent an Urgent Recall letters dated October 13, 2013 to all distributors via first class mail. Direct accounts were notified by phone as well as fax. Customers were informed of the affected product, problem and actions to be taken. Replacement products are being distributed to each site and will be delivered to each customer. For questions call 408-883-6760. |  | 515 | 515 US, 315 International (all packaging configurations) | Worldwide Distribution - USA Nationwide in the countries of: Korea, Japan, Singapore, Thailand, United Kingdom, Canada, and Taiwan. | Terminated | December 10, 2013 | 22 | 515 | Not_Computer
Z-0357-2014 | 66439 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trinica Anterior Lumbar Plate (ALP) system, part 07.01058.001. Non Sterile, Rx only. The Trinica Anterior Lumbar Plate System is a temporary supplemental fixation device consisting of a variety of shapes and sizes of plates and screws. The Trinica Anterior Lumbar Plate System is used as an implant for the correction and stabilization of the spine. This system provides temporary stabilization and augments the development of a solid spinal fusion. Additionally, this system provides the surgeon with the ability to supplement an interbody device with anterior plate fixation. The Trinica Anterior Lumbar Plate System components can be locked into a variety of configurations and each construct may be customized to individual cases. The plates are low profile and anatomically designed to provide optimal fit from either anterior or anterior-lateral approach. This system also features anti-migration locking caps to help secure the fixation screws. All Trinica Anterior Lumbar Plate System implant components are made from titanium alloy (Ti-6AI-4V).The Trinica Anterior Lumbar Plate System is indicated for use via the lateral or anterolateral surgical approach above the bifurcation of the great vessels or via the anterior surgical approach, below the bifurcation of the great vessels. This system is indicated in the treatment of lumbar or lumbosacral (L1-SI) spine instability as a result of fracture (including dislocation and subluxation), tumor, degenerative disc disease (defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies), pseudoarthrosis, spondylolisthesis, scoliosis, lordotic deformities of the spine, spinal stenosis, or a failed previous fusion. | Trinica Anterior Lumbar Plate (ALP) system | 2 | November 18, 2013 | 2013 | Zimmer, Inc. | Zimmer Spine, Inc. is initiating a correction regarding the Trinica ALP Instrument Tray, part 07.01058.001 (all lots). The Ratchet Handles, part 07.00438.001 may not be effectively sterilized if left in their designated location(s) in the instrument tray. This is due to the current bracketing in the tray potentially interfering with complete steam penetration into the cannulated handle. | DESIGN: Device Design | Zimmer Spine sent a Medical Device Correction letter dated October 9, 2013 via FedEx A Medical Device Correction letter, dated October 09, 2013 via FedEx to distributors, Hospitals and Surgeons who have received and/or used the device. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notificaiton and ensure affected personnel were aware of its contents. Only sterilize the Ratchet Handles external to the instrument tray. Customers were to be notified when replacement trays were available. Customer with questions were instructed to call the Zimmer Customer Call Center at 866-774-6368.For questions regarding this recall call 800-613-6131. | Contact the recalling firm for information | 94 | 94 Tray sets | Worldwide Distribution - USA (nationwide) and Internationally to Nationwide and OUS countries include: Australia, Austria, Belgium, Canada, France, Germany, Italy, New Zealand, Spain, Switzerland, and the UK. | Open |  |  | 94 | Not_Computer
Z-0340-2014 | 66712 | catheter, recording, electrode, reprocessed | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Diagnostic Electrophysiology (EP) Catheters and Diagnostic EP Catheter Cables, Model Numbers: D6-DR-252-RT, F6-QF-252RT, 201101, 201104, 401442, 401860, D7-A20-131-RT, 401661, 401972, 401977, 401981, 401982, C6-MR/MST4-SA, CB3434CT.Product Usage: Reprocessed diagnostic EP catheters are indicated for temporary intracardiac sensing, recording, stimulation and electrophysiological mapping of cardiac structures. Note: These devices do not support or sustain life. Diagnostic EP Catheter cables are designed as electrode cables with a multi-pin connector on the distal end and the appropriate number of tails on the proximal end. The cables interface an EP catheter with the appropriate external stimulation or recording equipment or serve as an extension cable between an EP catheter and equipment out of immediate reach. Reprocessed diagnostic electrophysiology catheter cables are indicated for use with the appropriate electrode catheter during electrophysiology studies. | Eletrophysiology Catheter; BARD,Daig,BW CristaCath,Response,Supreme,Qwik Cable,BW Webster, Carto 3 | 2 | November 18, 2013 | 2013 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Reprocessed Diagnostic Electrophysiology (EP) Catheters Cables due to a potential sterility breach in the package seal. | PRODUCTION CONTROLS: Process Control | Stryker sent an Urgent Medical Device Removal Notification letter to affected customers. The letter informed customers of the affected product, problem identified and the actions to be taken. Customers are instructed to complete the "Recall Effectiveness Form" and fax to Moira Barton-Varty, Senior Director Regulatory Affairs at (480) 809-4957. Customers with any serious adverse events or product problems are instructed to call Stryker Sustainability Solutions at (888) 763-8803. |  | 46 | 46 units total | US Distribution: in the states of: Arizona, California, and Ohio. | Terminated | December 06, 2013 | 18 | 46 | Not_Computer
Z-0342-2014 | 66688 | table, surgical with orthopedic accessories, manual | JEB | Manual operating table and accessories and manual operating chair and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | TruSystem Surgical Table (TS7000)  Leg Section One Part Light HUUse of the Leg Section One Part Light HU is permitted only with the TRUMPF TruSystem 7000 operating table. The operating table TruSystem 7000, in conjunction with operating table components is intended for the use of patient positioning during surgery, ranging from anesthesia induction to actual surgery to recovery from anesthesia. The Leg Section One Part Light HU is attached to coupling point L on the operating table and is used to position the patients legs during pre-op and post-op surgery. | TruSystem | 2 | November 18, 2013 | 2013 | Trumpf Medical Systems, Inc. | The coupler point (hood) of the Leg Section One Part Light may become detached from the leg section frame. | PRODUCTION CONTROLS: Packaging | TRUMPF sent an Urgent Safety Information letter dated October 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to inspect Leg Section One Part Light for damage to the coupler point (hook) between each case. If any damage is noticed to the coupler, they were instructed to stop use immediately and contact TRUMPF Technical Service. If they drop or collide the Leg Section One Part Light into anything then stop use immediately and contact TRUMPF Technical Service. Review the User Instruction for proper use and care of the device. The leg section should be used only as intended as part of the TRUMPF TS7000 surgical table. The intended use is to support and position the patients legs during the pre-operative and post-operative stages of surgery. The leg section should be removed during surgery. TRUMPF will be exchanging the hooks, counter plates and screws of the coupler point to increase the area that the load is distributed over. This will allow the screws securing the hood to the leg frame to withstand external forces. TRUMPF anticipates parts to be available the first week of November. TRUMPF trained technicians will complete the TRUMPF Medical Systems, Inc. for the part exchange. Completion is anticipated by November 11, 2013.Further questions please call (843) 534-0606. | Contact the recalling firm for information | 15 | 15 components | US Distribution including the states of CA, IL, NJ, NY, NC, TX, WA and SC. | Open |  |  | 15 | Not_Computer
Z-0347-2014 | 65584 | suture, surgical, absorbable, polydioxanone | NEW | Absorbable polydioxanone surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Polyglycolic-Lactic Acid Absorbable Surgical SutureLabel reads in part "Coated VILET *** (Polyglacin 910) Suture UNDYED BRANDED V423 *** Riverpoint ***"PGLA surgical suture is indicated for use in general soft tissue approximation and/or ligation, including use in opthalmic procedures, but not for use in cardiovascular and neurological tissues | PolyglycolicLactic Acid Absorbable Surgical Suture | 2 | November 18, 2013 | 2013 | Riverpoint Medical, LLC | Riverpoint Medical is recalling PGLA Surgical Suture (Vilet) because the label indicates the suture is undyed and it is actually a violet colored suture. | TRAINING: Employee Error | Riverpoint sent an Urgent Medical Device Recall letter in December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to review their inventory and immediately quarantine the product in question, if found. In addition, if they have further distributed the recalled product, the consignees were advised to identify their customers and notify them of this product recall. All affected product is to be returned to Riverpoint Medical. Consignees were asked to call 503-517-8001 to return recalled product and request a credit or replacement.For questions regarding this recall call 503-517-8001. |  | 1932 | 1932 pieces | US Distribution in the states of: AZ, FL, and TX. | Terminated | November 21, 2013 | 3 | 1932 | Not_Computer
Z-0341-2014 | 66529 | retractor | GAD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small ExtendibleOffset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion. | Synthes Soft Tissue Retractor Small Extendible | 2 | November 18, 2013 | 2013 | Synthes USA HQ, Inc. | The potential exists for the locking screw nut to break on the Soft Tissue Retractor if it is overt-tightened. | DESIGN: Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and immediately remove the affected lots from stock, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the letter by checking the appropriate box indicating affected product has been located and return the Verification Form with the product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have the identified product, they were instructed to complete the Offset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion.e attached Verification Section and return to Synthes. | Contact the recalling firm for information | 374 | 374 | Worldwide Distribution - USA (Nationwide) and Canada. | Open |  |  | 374 | Not_Computer
Z-0336-2014 | 66670 | prosthesis, vascular graft, of 6mm and greater diameter | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Vascutek Gelsoft Vascular Graft (peripheral)Labeling reads in part, "***Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis***STRAIGHT***STERILE***EO***Item No: 631508***Diameter 8mm***Usable Length 15cm***Manufactured By***VASCUTEK Ltd. a TERUMO Company Newmains Avenue Inchinnan, Renfrewshire PA4 9RR Scotland***"The Vascutek Gelsoft Vascular Prosthesis is a gelatin-sealed, woven polyester graft indicated for abdominal and peripheral vascular repair, i.e. replacement or bypass in aneurysmal and occlusive disease of arteries | Vascutek Gelsoft | 2 | November 18, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) became aware of a customer complaint that a 15mm x 8mm Gelsoft graft was received which the customer believed was a 15mm x 6mm Gelsoft graft unit. The manufacturer confirmed a batch of product code 631508 was incorrectly labeled. | PRODUCTION CONTROLS: Error in Labeling | Terumo CVS called affected consignees (via telephone)on 10/23/2013 of this device removal. Terumo CVS advised users to discontinue use and return all affected grafts in inventory. Terumo will replace or issue credit for returned grafts. Terumo provided consignees with the reason of the removal, potential hazard, and description of affected product. Terumo addressed any questions from the consignees. Consignees with questions were instructed to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 734-741-6173. |  | 9 | 9 | US Distribution in the states of: NC and CA. | Terminated | February 27, 2014 | 101 | 9 | Not_Computer
Z-0365-2014 | 66396 | diazo colorimetry, bilirubin | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Chemistry Calibrator (CAL) used with the ADVIA Chemistry Direct Bilirubin_2 Reagent, Siemens Material Number (SMN) 10312279, REF 09784096, Part Number T03-1291-62; IVD Product usage: Usage:DBIL_2: For in vitro diagnostic use in the quantitative determination of direct bilirubin in human serum and plasma on the ADVIA Chemistry systems. Such measurements are useful in the screening of liver function disorders or in the diagnosis of jaundice. | Siemens Chemistry Calibrator used with the ADVIA Chemistry Direct Bilirubin 2 Reagent | 2 | November 19, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics is conducting a field correction for Siemens Chemistry Calibrator Lots 090445B and 090445C on the ADVIA 1200, 1650, 1800, and 2400 Chemistry Systems. The System Specific Values (SSVs) for the Direct Bilirubin Assay (DBIL_2) have been reassigned for these lots. Siemens has determined that the DBIL_2 SSVs currently assigned to these lots may cause an average shift of |  | Siemens Healthcare Diagnostics sent an Urgent Medical Device Correction Letters dated September 24, 2013 and Response Forms to the US consignees via FedEx. Urgent Field Safety Notices were also distributed to customers outside the US. Customers are instructed to use the reassigned System Specific Values for the affected lots listed in the letter. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within thirty (30) days. The questionnaire is to be faxed to the Technical Solutions Center at 302-631-7597. On September 30, 2013, Siemens sent a revised customer notifications to correct the Part Number. |  | 1191 | Domestic: 1191 units; Foreign: 3052 units | Worldwide Distribution - USA Nationwide in the states of: Argentina, Austria, Belgium, Bulgaria, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, France, French Polynesia, Germany, Guadeloupe, Guatemala, Great Britain, Greece, Korea, Indonesia, Israel, Italy, Kazakhstan, Latvia, Liechstenstein, Lithuania, Luxembourg, Macedonia, Martinique, Mexico, Netherlands, Norway, Peru, Poland, Portugal, Russian Federation, Serbia, Singapore, Slovenia, Slovakia, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, Uruguay, Vatican City, and Vietnam. | Open |  |  | 4243 | Not_Computer
Z-0362-2014 | 66609 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Fixed Offset Adapter, A-642100.The Offset Adapter is a fixed modular component which sets the tube 14.5 mm (0.57 inch) offset from the center line of the pyramid. It is made of a high-strength alloy with a pyramid receiver and is intended where there is a need to offset the tube horizontally in relation to other components. It is fully compatible with all standard 30mm prosthetic pylons. (3) Offset: 14,5mm (0.57inch). | Fixed Offset Adapter, A642100 | 2 | November 19, 2013 | 2013 | Ossur North America Inc | Ossur inititated a voluntary recall of the A-642100 Fixed Offset Adapter because it can develop fatigue cracks in normal use over an extended period of time and ultimately separate completely. | DESIGN: Device Design | Ossur sent an "Safety Notice Voluntary Product Recall" letter dated October 29, 2013, via regular mail. Follow-up telephones are also made. The letter identified the product the problem and the actions to be taken by the customer.ÂŽssur is committed to providing superior customer service and we will do our utmost to minimize any inconvenience this may cause your practice. To ensure patient safety we kindly ask you to take the following action:" Please identify the users of each of the A-642100 adapters, see attached list." Also, please check your inventory for any A-642100s you may have on hand but have not fitted on users. A follow-up phone call will be made a few days later to coordinate replacements. A new replacement component will be provided to the customer. For further questions please call (949) 360-3634. |  | 66 | 66 units | US Distribution including the states of AL, TX, NY, CA, VA, GA, WI, NC, ID and NJ. | Open |  |  | 66 | Not_Computer
Z-0358-2014 | 66391 | disinfectant, medical devices | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | Ready to Use Thymo-Cide, Germicidal Detergent, containing 0.051% thymol, packaged in 1-qt. bottles. The firm name on the label is Wexford Labs, Inc., Kirkwood, MO.For disinfection of hard, non-porous surfaces (floors, walls, tables, etc.) in operating rooms, intensive care, nurseries, emergency areas, dental operatories, police, EMS, and all health care facilities. | Ready To Use ThymoCide | 2 | November 19, 2013 | 2013 | Wexford Labs Inc | Expired active ingredient was used in the manufacture of the product. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Wexford Labs, Inc. sent undated letters by e-mail to affected customers beginning September 10, 2013, and follow-up phone calls were initiated requesting return of the affected product. E-mail letters dated September 26, 2013, were issued requesting return of product and asking the customers to notify their downstream customers. Customers with questions were instructed to call 800-506-1146.For questions regarding this recall call 314-966-4134,. |  | 7021 | 7,020/1-qt. bottles (1755 gallons) | Nationwide Distribution - USA including FL, SC, TX, MO, CA, OH, OK, MA, IN, and GA. | Open |  |  | 7021 | Not_Computer
Z-0361-2014 | 66534 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap Arcadius XP L Trial Insertion Instrument (ME020R/ME020R-US) | Aesculap Arcadius Trial Insertion Instrument (ME020R/ME020RUS) | 2 | November 19, 2013 | 2013 | Aesculap, Inc. | The trial insertion instrument may disassemble during use due to a potentially weak weld on then trial inserter shaft block area located near the thumb wheel. | DESIGN: Device Design | An Important Product Removal Notification dated September 20, 2013 was sent to customers. |  | 30 | 30 | Worldwide distribution: USA (nationwide) GA, TN, CA, WY, NV, TX, OK, OH, and IL; and internationally to Canada. | Open |  |  | 30 | Not_Computer
Z-0360-2014 | 66720 | analyzer, gas, carbon-dioxide, gaseous-phase | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs EtCO2 Sample line 15' (4.6 M) tube, male luer connectors.Rx only; Made in Mexico; Manufacturer:Salter Labs, 100 West Sycamore Road,Arvin, CA 93203.CO2 sample line. | Salter Labs EtCO2 Sample line | 2 | November 19, 2013 | 2013 | SALTER LABS | It has been determined that three lots of model number 4515-15-25 CO2 Sample Line 15' tube, male Luer lock connectors may contain a 30' tube instead of the 15' length. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Salter Labs called the five customers that received product from the three lots, then follow-up immediately by sending a recall notification letter directly to each customer. Each letter will identify all shipments of subject product that they received.All customers letters were sent certified mail, return receipt requested. |  | 91 | 91 cases (25 units per case) | Distributed in the states of MS, TX FL,, WI, MI, VA, NY, MO, NC, and MS. | Terminated | December 17, 2013 | 28 | 91 | Not_Computer
Z-0364-2014 | 66639 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Juggerknotless Drill Bit, REF 110003173, Sterile, Single UseProduct Usage:The Juggerknotless Soft Anchors are indicated for soft tissue reattachment procedures in the shoulder, foot/ankle, hand/wrist, and hip. | Juggerknotless Soft Anchors | 2 | November 19, 2013 | 2013 | Biomet, Inc. | Investigation identified that the design of the drill bit may make the suture anchor difficult to insert into the bone. The implant may skip across the hole or miss the hole during insertion. | DESIGN: Device Design | Biomet sent an Urgent Medical Device Recall Notice letter dated October 10, 2013 to all consignees. The letter identified the affected product, reason for recall, and actions to be taken. The letter stated customers will be contacted on an individual basis in an attempt to reconcile the affected product. Customers were directed to locate and remove implicated devices and fax a copy of the response form to 574-372-1683 prior to returning product. For questions call 574-372-1570. |  | 25 | 25 | USA Nationwide Distribution in the states of VA, CA, Ky, and IN | Open |  |  | 25 | Not_Computer
Z-0363-2014 | 66137 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Model FL14E Rose BedsThe Rose Bed is an AC-powered adjustable long-term care bed intended for medical purposes that consists of a bed with a built-in motor and controls that can be operated to adjust the height and surface contour of the bed. The device includes movable and latchable siderails | Rose Bed | 2 | November 19, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | It was reported that the siderail may appear latched when it is not, due to lack of proper preventive maintenance. This might result in bruises and/or lacerations to a patient who is leaning against the siderails if the siderail releases. Additionally, the siderail may not effectively constrain a patient to prevent them from rolling or falling off the bed. | PRODUCTION CONTROLS: Equipment Maintenance | Stryker Medical sent a Medical Device Notification letter dated October 23, 2013, to all affected consignees. The letter described the recalled product, potential injuries associated with the side rail sticking, as well as providing preventive measures to be taken to mitigate the risk. Consignees were directed to locate the beds listed in this notice, ensure that the siderails have been inspected and adjusted, as required, continue to routinely inspect the siderails and functionality using the enclosed, enhanced product maintenance information. Problems and concerns should be addressed to Stryker at (1-800-327-0770, option 2). Consignees were asked to Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosed return form and directed to forward the notice if product was further distributed and advise Stryker of the beds new location.For questions regarding this recall call 269-324-6609. |  | 4126 | 4126 units USA; 20449 inits Internationally | Worldwide Distribution - USA (nationwide) and Internationally to Canada; Japan, South America, Latin America and China. | Open |  |  | 4126 | Not_Computer
Z-0359-2014 | 66300 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIOMET Spine, Parsippany, NJ 07054. 5.5 TI 5.5x50mm Translation Screw Polaris 5.5 TI Spinal System. Rx Only. Non-Sterile.Intended for immobilization and stabilization as an adjunct to fusion as a pedicle screw fixation system | Polaris Translation Screw System | 2 | November 19, 2013 | 2013 | Ebi, Llc | Polaris Translation Screw Failures were recently reported to Biomet Spine. These failures are generally described as dissociation of the screw head from the shaft of the Polaris Translation Screw. | DESIGN: Device Design | Urgent Product Information notices dated 8/22/2013 were mailed via FedEx to distributors on 8/23/2013. Follow-up emails with a copy of the UPI notice were sent to distributors on 8/23/2013. The UPI notice informed the customers of the issue with the product and how to identify the affected product. Customers were advised that the Polaris Translation Screw System and the SpF Implantable Spine Fusion Stimulator not be used in conjunction with any implants containing cobalt chrome. Firm contacts were provided in case customers had any questions. EBI, LLC had a conference call with their domestic distributors on 8/23/2013 informing them they are required to notify their customers. |  | 7092 | 7,092 units | Nationwide distribution including Puerto Rico. | Open |  |  | 7092 | Not_Computer
Z-0352-2014 | 66721 | gas-machine, anesthesia | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | CAS I/II Absorbers that are parts of the BleaseSirius Anesthesia Workstation, the BleaseFocus Anesthesia Workstation with the following part numbers: 12200900; 12200901; 12200902; 12200903; 14200100, and service kits part number 050-0659-00 and 050-0901-00. BleaseFocus Anesthesia Workstation: The Blease Frontline Plus Range, Anesthesia Machines are intended for use in the hospital environment and locations not requiring portability. It may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of a mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. This device is intended for use only by a suitably qualified physician. BleaseSirius Anesthesia Workstation: The Spacelabs BleaseSirius Anesthesia Workstation is intended for use in the hospital environment and operating room. It may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. The device is intended for use only by a suitably qualified physician. | BleaseSirius Anesthesia Workstation, BleaseFocus Anesthesia Workstation | 1 | November 20, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs Healthcare Ltd. of United Kingdom recalls , and Service Kits, Part Number 050-0659-00 and 050-0901-00. The Bag to-Vent switch in CAS I/II Absorbers may fail due to loose fastening hardware in an absorber. | PRODUCTION CONTROLS: Process Control | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated November 1, 2013, to all affected customers. Spacelabs Healthcare also sent a customer letter (translated as necessary) to all international subsidiaries and distributors of record via e-mail on November 6, 2013. The letter informed customers of a potential defect in CAS I/II Absorber products employed in the Spacelabs Healthcare BleaseSirius Anesthesia Workstation, and Service Kits, Part Number 050-0659-00 and 050-0901-00.The Bag-to-Vent switch in CAS I/II Absorbers may fail. In the worst case, this would result in a large leak from the absorber circuit and could impede the ability to provide ventilation.Customers were instructed to immediately advise their staff of the problem. Customers with questions were instructed to contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025. |  | 37 | 37 units (19 in the US and 18 International) | Worldwide Distribution - USA (nationwide) including Oklahoma and North Carolina and Internationally to Libyan Arab Jamahiriya, Finland, Jordan, Mexico, Papua New Guinea, and Taiwan. | Open |  |  | 37 | Not_Computer
Z-0366-2014 | 66640 | laryngoscope, rigid | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Green Rusch - Lite Disposable Metal Laryngoscope Ridge Blade to visualize upper airway and aid in placement of tracheal tube. | Green RuschLite Disposable Metal Larynogoscope Blade. | 2 | November 20, 2013 | 2013 | Teleflex Medical | Labeling inconsistency | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Recall letters were sent to customers and distributors on 10/15/2013 via courier. The letter indicated that the consignee was to immediately discontinue use and quarantine any products with the affected catalog numbers. The letter also included a Recall Acknowledgment Form which was to be faxed back to the firm. |  | 6200 | 6,200 ea. | Nationwide distribution: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV and WI. | Open |  |  | 6200 | Not_Computer
Z-0369-2014 | 66552 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical BladeSynthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blade is used to treat stable and unstable fractures of the proximal femur including pertrochanteric fractures, intertrochanteric fractures, basal neck fractures, and combinations thereof. The long TFN is additionally indicated for subtrochanteric fractures, pertrochanteric fractures associated with shaft fractures, pathologic fractures (including prophylactic use) in both trochanteric and diaphyseal regions, long subtrochanteric fractures, proximal or distal non-unions, malunions, and revisions. | Synthes Titanium Trochanteric Fixation Nail (TFN)Helical Blade | 2 | November 20, 2013 | 2013 | Synthes USA HQ, Inc. | A potential exists for non conforming Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blades to be distributed. | PRODUCTION CONTROLS: Process Control | Synthes sent a Field Safety Notice dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the options provided in the Field Safety Notice to clarify and mitigate the potential risk. Customers were asked to prepare a list of the affected part and lot numbers and contact the Complaint Unit at 610-719-5596 to receive a Return Authorization Number. Complete and return the Verification Form with the product and Return Authorization number. Customers were also asked to forward the Field Safety Notice to anyone in their facility that needs to be informed. Customers were asked to maintain a copy of the Field Safety Notice.Customers with questions were asked to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 350195 | 350195 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada. | Open |  |  | 350195 | Not_Computer
Z-0372-2014 | 66263 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | *** 1) SYMBIQ One-Channel Infuser; *** 2) SYMBIQ Two-Channel Infuser;*** E09 batteries can be identified by the "E09" that is printed beneath the bar code and at the end of the serial number along the bottom of the label found on each battery; *** E10 batteries can be identified by the "E10" that is printed beneath the bar code and at the end of the serial number along the bottom of the label found on each battery.Symbiq is a general-purpose infuser designed to deliver fluids, solutions, medications, agents, nutritionals, electrolytes, blood and blood products for parenteral, enteral, intra-arterial, subcutaneous, or irrigation routes of administration. | SYMBIQ OneChannel and TwoChannel Infusers | 2 | November 20, 2013 | 2013 | Hospira Inc. | Hospira shipped the incorrect replaceable battery pack (Part No: HSP3566-E10) instead of the correct E09 battery pack (Part No. HSP3566-E09) for the Symbiq Single and Dual Channel Infusers. | TRAINING: Employee Error | The firm, Hospira, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 28, 2013 to its customers. Hospira contracted with Stericycle, Inc., in Indianapolis, IN for customer notification, response, tracking, returns, effectiveness checks and destruction of returned product starting on August 28, 2013. The letter described the product, problem and actions to be taken. Customers were instructed to: 1) complete and return the attached reply form via fax to 1-866-470-3291 or e-mail to Hospira2921@stericycle.com, even if you do not currently have the affected batteries; 2) contact Stericycle at 1-877-779-7810 (M-F, 8am - 5pm ET) if they require additional response forms; and, 3) return any affected batteries upon receipt of the new (correct) replacement batteries to Stericycle. Customers are also instructed to forward the recall information to any accounts that they may have shipped the affected batteries to and have them contact Stericycle to recieve a reply form and return label.For further inquiries, please contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 6, availabel 24 hours a day/7 days per week. |  | 59 | 59 batteries | US Distribution in states of: : KY, MN, ND, NJ, NY, PA, and SD. | Open |  |  | 59 | Not_Computer
Z-0371-2014 | 66056 | lift, patient, non-ac-powered | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medcare Products Max-500 Stand-N-Weigh, Total Support, patient stand, model #400007. The responsible firm name listed in the Operations Manual is Medcare Products, Burnsville, MN.Used for assisting patients/residents to a standing position. | Medcare Products Max500 Stand, Total Support | 2 | November 20, 2013 | 2013 | Ergosafe Products LLC DBA Prism Medical Services USA | The lift might fail with the labeled 500-lb. load rating. | DESIGN: Device Design | The firm, Medcare Products, sent an "URGENT FIELD CORRECTIVE ACTIONS" letter dated October 21, 2013 via certified mail on October 30, 2013 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to peel off the boom stickers and serial number label and replace with the new ones supplied; review the revised Manual and dispose of the original; and complete the action report form once the products have been relabeled. Please refer to your Operations Manual or visit our website at www.medcarelifts.com to view the manual online.If we can answer any questions that you may have regarding the testing standards or the relabeling of the products, please do not hesitate to contact us at 1-800-695-4479. |  | 233 | 233 units | Nationwide distribution. | Open |  |  | 330 | Not_Computer
Z-0349-2014 | 64832 | pump, infusion | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. | Gemstar Infusion System | 1 | November 21, 2013 | 2013 | Hospira Inc. | The proximal and distal pressure sensor calibration can drift resulting in the pump failing the Proximal or Distal Occlusion Operational Test, as described in the GemStar Technical Service Manual, or reporting one of the following errors during device setup or infusion: 1) Cassette Check - D; 2) Cassette Check - P; 3) Proximal Occlusion; 4) Distal Occlusion; 5) Pressure Calibration Err | DESIGN: Device Design | Hospira sent an Urgent Device Field Correction letter dated March 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter asks that customers immediately perform proximal and distal occlusion tests as defined i nthe GemStar Technical Service Manual. If the device fails either of the tests, remove it from clinical service and contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM) to report the issue and arrange for the return of the device for recalibration. The letter further reminds customers that the proximal and distal occlusion tests should be performed on an annual basis as a part of the yearly GemStar maintenance schedule. Customers were also requested to complete and return the attached reply form and return it to Hospira by fax to 1-888-714-5077 or by e-mail to Hospira7906@stericycle.com.Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. For questions regarding this recall call 800-441-4100. |  | 15 | 1) List Number 13000: 2,162 pumps; 2) List Number 13100: 15 pumps; 3) List Number 13150: 1,955 pumps | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open |  |  | 43855 | Not_Computer
Z-0376-2014 | 66359 | dental cement w/out zinc-oxide eugenol as an ulcer covering for pain relief | MZW | Dental cement. | Dental | Dental | 510(k) | ExciTE F DSC Soft touch Single Dose, Dual-curing Regular Ref. 50 dental adhesive for restorations | Excite F. DSC | 2 | November 21, 2013 | 2013 | Ivoclar Vivadent, Inc. | Ivoclar Vivadent, Inc. is recalling a lot of ExciTE F DSC Dental Adhesive found to produce a blue green discoloration when used in combination with cements. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | The firm, Ivoclar Vivadent, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 19, 2013 via first class mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check your stock and segregate any material marked batch R59595; contact the Ivoclar Vivadent Customer Service at 800-533-6825 as soon as possible between the hours of 9:00am and 7:00pm EST Monday through Friday and a representative will arrange to have your stock replaced; both unopened and opened packages can be returned; and complete and return the Medical Device Recall Return Response Acknowledgment and Receipt Form.Please communicate locally with Ivoclar Vivadent Customer Service at 800-533-6825 for any questions. |  | 63 | 63 refill packs (50 single use applications per refill) | US Distribution: including states of PA, WA, WI, NY, LA, TX, NJ, CA, OR, IA, SC, IN, IL, and AZ. | Open |  |  | 63 | Not_Computer
Z-0375-2014 | 66657 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker1) XIA Torque Wrench2) XIA 3 Torque Wrench3) XIA Elegance Short Torque WrenchManufactured by: Stryker Spine SAS Z.I. Marticot -33610 -CESTAS -FRANCE +33 (0) 5.57.97.06.30http://www.stryker.comDistributed in the USA by: Stryker Spine 2 Pearl Court Allendale, NJ 07401-1677 USA +1-201-760-8000XIA 3 Spinal System- Intended for use in the noncervical spine. When used as an anterior/anterloateral and posterior, non cervical pedicle and on-pedicled fixation system, the XIA 3 Spinal system is intended to provide additional support during fusion autograft or allograft in skeletally mature patients in the treatment of the following acute and chronic instabilities or deformities; degenerate disc disease (DDD) ; spondylolisthesis; trauma spinal stenosis; curvatures; tumor; pseudoarthrosis and failed previous fusion. XIA - Intended for anterior/anterolateral and posterior, non cervical pedicle and non-pedicle fixation of the spine to provide immobilization and stabilization of spinal segments in skeletally mature patients as an adjunct to fusion for the following indications: degenerative disc disease; spondylolisthesis, trauma, spinal stenosis, curvatures, tumor, pseudoarthrosis and failed previous fusion. | Stryker Spine Torque Wrench | 2 | November 21, 2013 | 2013 | Stryker Spine | Multiple complaints have reported that the hex tip of the torque wrench fracturing during final tightening. | DESIGN: Process Design | Stryker sent an Urgent Medical Device Removal Notification letter dated September 25, 2013, with product accountability forms to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify any affected product, reconcile any Torque Wrenches with Catalog Number with Lot Numbers referenced in the letter. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 855-632-9049 or you email it to Spine-RegulatoryActions@Stryker.com. Customers were asked to return the affected product using the enclosed mailing label to: RECALL RA 2013-105Stryker Spine59 Route 17 South,Allendale, New Jersey 07401branches/agencies, the field and surgeons via Fed Ex on starting on 9/24/2013.For questions regarding this recall call 201-760-8298. |  | 144 | 35 units US; 109 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Canada, Chile, France, Canada, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Netherlands, New Zealand, Spain, Iraq, Egypt, and Turkey. | Open |  |  | 144 | Not_Computer
Z-0379-2014 | 66663 | glucose oxidase, glucose | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Product Brand Name(s): Alere Cholestech LDXÂŽ Lipid Profile"GLU Test Cassette, Model 10-991.Product Generic Name: Alere Cholestech LDXÂŽ Lipid Profile"GLU Test Cassette.Description of the product: Lipid. Glucose panel Test cassettes.Product packaging: 10 individually pouched cassettes in a labeled kit box.Intended for the quantitative determination of total cholesterol, HDL (high-density lipoprotein) cholesterol, triglycerides, and/or glucose in whole blood, serum, and plasma. | Alere Cholestech LDX Lipid Profile"GLU Test Cassette | 3 | November 22, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated a recall of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 320167, becuase an incorrect lot number was printed on the outer 10 pack kit box label. The incorrect Lot # number printed was 321067. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Alere initiated this recall by sending out notification letters to customers via fax, email, or direct mail. The letter titled "URGENT MEDICAL DEVICE NOTICE", dated October 23, 2013 is accompanied by a response form/Verification form. The notification informed customers of the reason for recall, product description with codes, required actions by the customers, and contact information. |  | 3621 | 3,621 kits | Worldwide Distribution: USA (nationwide) and Internationally to Canada. | Open |  |  | 3621 | Not_Computer
Z-0377-2014 | 66561 | clip, aneurysm | HCH | Aneurysm clip. | Neurology | Neurology | 510(k) | Mizuho Sugita Titanium Standard Aneurysm Clip II , Temporary and permanent occlusion of cervical segment of cerebral aneurysms and peripheral blood vessels. They are designed for Single Use. | Mizuho Sugita Titanium Standard Aneurysm Clip II Temporary Mini Type No 56 | 2 | November 22, 2013 | 2013 | Mizuho America, Inc. | Specific item numbers form Lots 27-J and 29-J of Sugita T2 Aneurysm Clips. Mini Temporary Type, are discolored. | PRODUCTION CONTROLS: Process Control | Mizuho sent a Product Recall Notice dated October 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to verify serial number before opening package and conduct a physical examination of the product for discolored Clips from the affected lots. Customers were also asked to fill out the Confirmation of User Receipt, and return to Misuho America. If customers should find affed clips they were instructed to quarantine them an call Mizuho America Customer Service at 800-699-2547 for return instructions. The Clips will be exchanged at no charge. Customers were also asked to ensure that all appropriate persons within the Hospital have been notified of the Recall Notice.For questions regarding this recall call 510-324-4500, ext 122. |  | 32 | 32 | Nationwide Distribution | Terminated | January 13, 2014 | 52 | 67 | Not_Computer
Z-0380-2014 | 66411 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4 and 2.7 MM Locking Screw SLF-TPNG with T8 Stardrive Recess 12 MMFixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments, particularly in osteopenic bone. | Synthes 2.4 and 2.7 MM Locking Screw SLFTPNG with T8 Stardrive Recess 12 MM | 2 | November 22, 2013 | 2013 | Synthes USA HQ, Inc. | Specific lots of the SS 2.4mm and SS 2.7mm Locking Screw (self-tapping) due to a mislabeling issue. The SS 2.7mm Screw was labeled as SS 2.4 mm Screw. | PRODUCTION CONTROLS: Labeling Mix-Ups | Synthes sent an URGENT NOTICE: MEDICAL DEVICE RECALL letter dated July 31, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Nuber, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Fax the documents to Synthes at 1-866-792-6446 or email to FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 98 | 98 | Nationwide Distribution including PA, CA, FL, MO, KS, IN, IL, WV, KY, LA, NM, SC, IA, TX, AZ, GA, NC, OK, NY, TN, MI, CO, and MT. | Open |  |  | 98 | Not_Computer
Z-0384-2014 | 66415 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix Threaded PersuaderDeformity is a comprehensive thoracolumbar pedical screw system designed to provide flexibility biomechanical performance and a solution to complex posterior pathological challenges. | Synthes Matrix Threaded Persuader | 2 | November 22, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes became aware that during a procedure using the Matrix Threaded Persuader, the Reduction Insert malfunctioned causing pieces to break off. Once there is a breakage to the Reduction Insert it will not allow the instrument to function as required. | DESIGN: Labeling Design | The firm, Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 31, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from stock; If you do have the affected devices, contact Synthes at 1-800-479-6329 to obtain the RA# and return the product to with completed form to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; complete and sign the Return Receipt Requested form via fax at 1-888-731-7955 or email to: FieldAction@synthes.com, even if you do not have the affected product.If you have questions, please call 610-719-5450. | Contact the recalling firm for information | 634 | 634 | Worldwide distribution: US (nationwide) and countries of Canada and Switzerland. | Open |  |  | 634 | Not_Computer
Z-0382-2014 | 66569 | abutment, implant, dental, endosseous | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Tri-Nex Healing Abutment, 0.50 mm x 5 mm height, Catalog No. HA-L-50-5.The NSI Implant System is intended to be implanted in the upper and lower jaw arches to provide support for fixed or removable dental prostheses in a single tooth, partially edentulous prostheses, or full arch prostheses. It further adds the option for immediate loading on single and splinted multiple unit restorations when good primary stability is achieved and with appropriate occlusal loading, to restore chewing function. | TriNex Healing Abutment | 2 | November 22, 2013 | 2013 | Southern Implants, Inc | Southern Implants is recalling the Tri-Nex healing abutment because the package was incorrectly labeled and packaged as the Tri-Nex healing abutment, 05.0 mm x 5 mm height (HA-L-50-5). The package contains the Tri-Nex healing abutment 05.0 mm x 6.0 mm Flare x 5 mm Height (HA-L-50W-5). | DESIGN: Labeling Design | Southern Implants, Inc. sent an Urgent Product Action letter dated September 11, 2013, to all affected customers. The letter identiied the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out the customer response form and fax/email the completed form to Customer Support at (866) 903-9272 or customersupport@keystonedental.com.For questions regarding this recall call 866-903-9272. |  | 197 | 197 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Guatemala, Taiwan. | Terminated | January 21, 2014 | 60 | 197 | Not_Computer
Z-0383-2014 | 66739 | urinalysis controls (assayed and unassayed) | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(k) | MAS Urinalysis Control Level 1, Multi-analyte control (Assayed and Unassayed), Catalog/Part Numbers: 024222; Product is manufactured and distributed by Microgenics Corporation, a subsidiary of Thermo Fisher Scientific, Clinical Diagnostic Division, Fremont, CAMAS Urinalysis is intended for use in the clinical laboratory as a control for qualitative and semi-quantitative procedures used in routine urinalysis testing. Assay values are provided for specific systems listed. The user can compare observations with expected ranges as a means of assuring consistent performance of reagent and equipment. The product is suitable for use as a control material for physiochemical, chemical, and microscopic methods of routine urinalysis. Urinalysis may be used in conjunction with commercially available standardized systems for urine microscopic analysis. | MAS Urinalysis | 3 | November 22, 2013 | 2013 | Thermofisher Scientive | Via one (1) consumer complaint and in-house testing has confirmed that the product does not meet the 30 day open vial stability specifications for Lot: UA1514011. |  | ThermoFisher Scientific, sent an Urgent Medical Device Field Correction letter dated October 29, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use, segregate and discard the affected product. Customers with questions were instructed to contact mas.controls@thermofisher.com. For question regarding this recall call 1-800-232-3342. | Contact the recalling firm for information | 914 | 914 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and internationally to Hong Kong. | Open |  |  | 914 | Not_Computer
Z-0381-2014 | 66792 | guide, surgical, instrument | FZX | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | REF 200068 Security (TM) Enclosed Carpal Tunnel SystemBlade ( Single) ABS /Stainless SteelSterile Device, Single Use Only. Designed to release the transverse carpal ligament. | Security Enclosed Carpal Tunnel Security Blade | 2 | November 22, 2013 | 2013 | Biomet, Inc. | Security Blades were manufactured oversized and potentially won't advance or may be difficult to advance through the mating Security Clip. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | 10/25/2013 Urgent Medical Device Recall Notifications were sent to each Distributor and Risk/Recall Manager at each direct account identifying the recalled device, possible adverse health consequences and instructing consignees to immediately locate and remove devices from circulation, following instructions on the FAX back form and returning it within 3 business days to FAX 574-372-1683 and returning product to Biomet Return Goods, 56 East Bell Drive Warsaw In 46580. Questions should be directed to 574-372-1570 M-F, 8 a.m. to 5 p.m. |  | 1130 | 1130 | Worldwide Distribution: USA (nationwide) and Internationally to Japan. | Open |  |  | 1130 | Not_Computer
Z-0388-2014 | 66420 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Tensioning Device The Synthes Osteogenesis System, MR Conditional is intended for fracture fixation (open and closed), pseudoarthrosis or non-unions of long bones, limb lengthening by epiphyseal or metaphyseal distraction, correction of bony or soft tissue deformities and correction of segmental bony or soft tissue defects in adult and pediatric patients. | Synthes Tensioning Device | 2 | November 25, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of the Synthes Tensioning Device were recalled due to the potential for the reaction fork to come loose from the main body of the device and not allow the wire to release, nor the device to apply tension. | DESIGN: Device Design | Synthes sent an Urgent: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer."Please remove this product from your inventory and return it to Synthes. A replacement product will be sent to you when available. If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page 3 of this letter) with the product to:" Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the document to Synthes by: Fax: 877-272-3517"Email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or email Fieldaction@synthes.comWe apologize for any inconvenience this may cause and thank you in advance for your attention and cooperation. |  | 369 | 369 | Nationwide Distribution. | Open |  |  | 369 | Not_Computer
Z-0386-2014 | 66542 | drug specific control materials | LAS | Clinical toxicology control material. | Toxicology | Toxicology | 510(k) | Single Analyte Urine DAU Control (5 m LNorbuprenorphine, Level II controlFor In Vitro Diagnostic Use Only. store at 2 - 8 degrees C.Contents: Contains 13 mg/Ml of Norbuprenorphine in human urine with sodium azide as a preservative.Manufacturer:Lin-Zhi International, Inc.670 Almanor Ave.Sunnyvale, CA 94085The Norbuprenorphine Drug of Abuse (DUB) Positive control is used as assayed quality control material to monitor the precision of the Lin-Zhi International, Inc. Buprenorphine Enzyme immunoassay on the Synchron clinical analyzer. It in intended for the qualitative and semi-quantitative determination of norbuprenorphine in human urine. | Single Analyte Urine DAU Control (5 m L | 3 | November 25, 2013 | 2013 | Lin-Zhi International Inc | Product catalog # A68825 Norbuprenorphine 13 ng/mL Level 2 Controls gives a decreased concentration reading due to the degradation of the NBUP analyte in solution. |  | Beckman Coulter sent a Manufacturer Field Action Letter dated November 2013, to the one distributor. The letter identified the product the problem and the action needed to be taken by the customer.Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.If you have any questions regarding this notice, please contact our Customer Support CenterÂŽÂŽ Via our website, http://www.beckmancoulter.com/customersupport/supportÂŽÂŽ Via phone, call 1-800-854-3633 in the United StatesÂŽÂŽ Outside the United States, contact your local Beckman Coulter Representative.We apologize for any inconvenience. |  | 60 | 60 items | US Distribution in CA. | Terminated | December 26, 2013 | 31 | 60 | Not_Computer
Z-0387-2014 | 66421 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Large Buttress Compression Nut For Blade Guide Sleeve TFN The Synthes Large Buttress Compression Nut For Blade Guide Sleeve TFN is intended to treat stable and unstable pertrochanteric fractures, intertrochanteric fractures, basal neck fractures, and combinations there of. | Synthes Large Buttress Compression Nut For Blade Guide Sleeve TFN | 2 | November 25, 2013 | 2013 | Synthes USA HQ, Inc. | The Large Buttress Compression Nut For Blade Guide Sleeve TFN was recalled because of complaints that the compression nut did not thread completely onto the blade guide sleeve and became stuck. | PRODUCTION CONTROLS: Process Control | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: Our records indicate you received this product. Synthesis requesting that the product be returned:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form {page 3 of this letter) with the product to:"Credit/Returns, Synthes, .1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by." Fax: 1-888-731-7952 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 135 | 135 | US Distribution including the states of TX, PA, NY, GA, CO, IN, AI, AR, IL, MI, TN, CA, AZ, MS and VA. | Open |  |  | 135 | Not_Computer
Z-0393-2014 | 66551 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Material Mix (Adjustable Cervical Distractor-Right, Matrix Distractor Rack, T-PAL Spacer Remover, Implant Inserter)Synthes Material Mix: Adjustable Cervical Distractor-Right is used to distract spine segment. Matrix Distractor Rack is used to distract the interbody disc space. T-PAL Spacer Remover is used to remove T-PAL spacer. Implant Inserter is used to insert the ProDisc implant. | Synthes Material Mix | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | One material batch used for production of the affected parts of the Synthes Material Mix were mixed in with an incorrect material with a lower than specified hardness and corrosion resistance. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product was located. Customers were also asked to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers should also return the Verification Form if no affected product was found in their inventory. Return all documents to Synthes by fax: 888-679-4875 or Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 44 | approximately 44 | Worldwided Distribution: USA (nationwide) including IL, MO, TX, WA and Internationally to Canada and Switzerland. | Open |  |  | 44 | Not_Computer
Z-0397-2014 | 66882 | shunt, central nervous system and components | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CSF-Unitized Shunt, Contoured Regular, Low Pressure, Catalog No. 46022;CSF-Unitized Shunt, Contoured Regular, Medium Pressure, Catalog No. 46024; CSF-Unitized Shunt, Contoured Regular, High Pressure,Catalog No. 46026.CSF-Unitized Shunt Kits are designed as integrated CSF-Flow Control Shunts for use in shunting cerebrospinal fluid from the lateral ventricle of the brain into the peritoneal cavity. | CSFUnitized Shunt | 2 | November 26, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the CSF-Unitized Shunts, Delta Shunt Assembly Kit, and the Strata II Shunt Assembly Kit because they were found to have label discrepancies. | OTHER/UNDETERMINED: Pending | Medtronic Neurosurgery sent an Urgent Medical Device Recall letter dated November 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to provide a copy of the letter to all those who need to be aware of the matter within their organization or to any organization where the potentially affected product may have been transferred. Customers with questions were instructed to call 805-571-8725. |  | 288 | 288 units total (147 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to US and worldwide: Austria, Belgium, Bosnia, Herzegovina, Canada, Croatia, Germany, Japan, Poland, Portugal, Taiwan, United Arab Emirates, United Kingdom. | Terminated | February 06, 2014 | 72 | 327 | Not_Computer
Z-0404-2014 | 66581 | nail, fixation, bone | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar is used as part of the Synthes Tibial Nail System EX to stabilize fractures of the proximal and distal tibia and the tibia shaft, open and closed tibia shaft fractures, certain pre- and post- isthmic fractures, and tibial malunions and non-unions. | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received reporting that a product packaged as a 14.5 mm Outer Protection Sleeve for Suprapatellar contained a 12.0 mm Outer Protection Sleeve for Suprapatellar. | PRODUCTION CONTROLS: Labeling Mix-Ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Number, complete the Verification Section at the end of the letter and return the Verification Form with the affected product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should complete the attached Verification Section and return all documents to Synthes by:Fax: 866-229-7778Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 55 | 55 | Nationwide Distribution | Open |  |  | 55 | Not_Computer
Z-0401-2014 | 66723 | condylar plate fixation implant | JDP | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5MM TI VA-LCP CRVD Condylar Plate/6H/159MM/RT-STERSynthes 4.5MM TI VA-LCP CRVD Condylar Plate/6H/159MM/RT-STER is used for buttressing multi fragmentary diatal femur fractures including: supra-condylar, intra-articular and extra articular condylar fractures, periprosthetic fractures, fractures in normal or osteopenic bone, nonunions and malunions. | Synthes 4.5MM TI VALCP CRVD Condylar Plate/6H/159MM/RTSTER | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | A plate was inadvertently released to a sales consultant that was restricted for sale. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Release of Mat./Comp. prior to Test Results | Synthes sent a letter dated November 7, 2013, to the implanting surgeon to follow up on a complaint received. The firm requested any additional information the surgeon might have about the use of the plate as well as a status on the patient following use. For questions the surgeon was instructed to call 800-620-7025.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | US Distribution in Colorado. | Open |  |  | 1 | Not_Computer
Z-0394-2014 | 66580 | washer, bolt nut | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Spiked Washers 13.5/6.5Spiked Washers 13.5/6.5 are intended to be used in ligament reattachment or fixation, specifically readaptation of torn or avulsed ligaments. | Synthes Spiked Washers 13.5/6.5 | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | During a retrospective assessment, Spiked Washers 13.5/6.5 were discovered to be mislabeled as Spiked Washers 13.5/5.5. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Finished Device Change Control | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product located. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes Avenue, Monument, CO 80132Customers were also asked to return the Verification Form indicating they did not have the affected product also. All documents were to be returned to Synthes by:Fax 1-877-270-5855Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales ConsultantFor questions regarding this recall call 610-719-5000.. |  | 3087 | 3087 | Nationwide Distribution | Open |  |  | 3087 | Not_Computer
Z-0390-2014 | 66666 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm and 5.0mm Dynamic Locking Screw (DLS) Product Usage: 3.7mm and 5.0mm Dynamic Locking Screw (DLS) in combination with Synthes Locking Compression Plates (LCP) are intended for use in long bone fractures, the fixation of osteopenic bone, the fixation of osteotomies, and for the fixation of non-unions and malunions. | Synthes 3.7mm and 5.0mm Dynamic Locking Screw (DLS) | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of 3.7mm and 5.0mm Dynamic Locking Screw (DLS) were recalled due to complaints of breakages at the distal tip of the pin identified after successful healing and during planned removal of the device. | MISBRANDING: Labeling False and Misleading | Synthes sent an Urgent Notice: Medical Device Recall letter to Sales Reps and End Users on July 22, 2013. The letter identified the affect product, problem and actions to be taken. For questions call the Recall center at 610-719-5450. |  | 40248 | 40248 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Open |  |  | 40248 | Not_Computer
Z-0400-2014 | 66836 | intervertebral fusion device with integrated fixation, lumbar | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX Timberline MPF Lateral Modular Plate Fixation System Straight Split Tip Driver - Short ("Driver"). Model Number 8630-0107.The Timberline MPF device is an intervertebral body fusion device consisting of a PEEK intervertebral spacer, titanium plate, and screws. The PEEK spacer has a generally rounded shape with various heights and footprints and has a hollowed out central area to accommodate autogenous bone graft. The upper and lower surfaces have a series of transverse grooves formed to improve stability and fixation once the device is inserted. The titanium plate has holes for receiving bone screws and a central hole for receiving a cover plate to prevent screw back-out. The Timberline MPF system is available in a variety of sizes and configurations to approximate anatomical variation in different vertebral levels and/or patient anatomy. The Timberline MPF system is provided nonsterile. The Driver is used with the Timberline MPF system to drive titanium screws through the plate and into the vertebrae thus locking plate to the vertebrae. | Timberline MPF Straight Split Tip Driver Short ("Driver") | 2 | November 26, 2013 | 2013 | Lanx, Inc. | Lanx, Inc. is recalling various lots of Timberline MPF Straight Split Tip Driver - Short ("Driver") because the width of the most distal end from the Driver handle (the tip of the hex portion of the Driver) is oversized. | PRODUCTION CONTROLS: Process Control | LANX sent an URGENT: Recall letter dated November 5, 2013, to all affected customers via Federal Express, Overnight delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product to the firm. Customers with questions were instructed to contact Field Inventory Management.For questions regarding this recall call 303-501-8402. |  | 25 | 25 units | Nationwide Distribution including AZ, PA, WI, NY, TX, PR, and TN. | Open |  |  | 25 | Not_Computer
Z-0402-2014 | 66815 | sterilizer, chemical | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX, Product Code: 10104The STERRAD Steri lization System is a low temperature general purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | STERRAD 100NX | 2 | November 26, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX and STERRAD 100NX sterilization systems because ASP has determined that the bronchoscopes may contain a polyurethane material in the suction channel location that is not listed in the current claims for interior channels (lumens) for processing in STERRAD systems. | DESIGN: Labeling Design | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notificaiton letter on November 11, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. Customers were instructed to provide the notice to anyone in their facility that needs to be informed and post this notice where appropriate. Customers were instructed to contact ASP at 1-888-783-7723 for any suspected problems. ASP is collaborating with KARL STORZ Endoskope to determine the sterilization impact of the suction port lumen material. Customers are instructed to contact KARL STORZ Endoskope Technical Support at (800)421-0837, ext 5350 for alternative reprocessing instructions.For questions regarding this recall call 949-453-6400. |  | 4381 | Total 4,381 units | Nationwide Distribution | Terminated | January 22, 2014 | 57 | 4381 | Not_Computer
Z-0391-2014 | 66416 | instrument, bending or contouring | HXP | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix 5.5.mm Left and Right Coronal Bender Product Usage: The Matrix 5.5.mm Left and Right Coronal Bender are used together to bend the 5.5 Matrix Rods in the coronal plane, and is intended for stabilization of the thoracic, lumbar, sacral, and/or iliac spine through screw and/or hook and rod fixation. | Synthes Matrix 5.5.mm Left and Right Coronal Bender | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | A complaint was received of the tips of the Matrix 5.5.mm Left and Right Coronal Bender breaking while bending a spinal rod component of the Matrix System. | DESIGN: Component Design/Selection | Synthes sent an Urgent Notice: Medical Device Recall letter dated June 22, 2013 to users and sales reps. The letter identified the affected product, problem and actions to be taken. For questions call 610-719-5450. |  | 209 | 209 | Worldwide Distribution - USA (nationwide) including GA, FL, CA, UT, IA, AZ, OH and Internationally to Canada. | Open |  |  | 209 | Not_Computer
Z-0395-2014 | 63230 | orthosis, spinal pedicle fixation, for degenerative disc disease | NKB | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Minimally Invasive Reduction Screw System Polyaxial Screw Body, Part Number 04.631.500Non-cervical spinal fixation devices intended for posterior pedicle screw fixation, posterior hook fixation, or anterolateral fixation. | Synthes TI MIRS Polyaxial Screw Body, Part Number 04.631.500 | 2 | November 26, 2013 | 2013 | Synthes USA HQ, Inc. | This action is being initiated following a detailed investigation in response to reported complaints of post-operative locking cap loosening. Synthes has concluded that if the construct is mis-aligned during final tightening, the potential exists for locking caps to loosen post-operatively in certain situations. |  | Synthes sent an Urgent: Medical Device Product Removal letter on September 7, 2012, to all affected consignees. Consignees were requested to immediately cease using the product, examine their inventory and remove the parts from inventory. If they had the parts, they were further instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the verification section at the end of the letter and return the verification form. If they did not have the parts, they were requested to complete the verification section at the end of the letter and return the verification form. Consignees with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 757 | 757 | Nationwide Distribution including CA, CO, CT, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Open |  |  | 1502 | Not_Computer
Z-0392-2014 | 66389 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter, VIAL-MATE Adaptor, 2B8071, Rx Only, STERILE R, Baxter Healthcare Corporation, Deerfield, IL 30015 USAThe VIAL-MATE is reconstitution device labeled for use with 20 mm drug vials and 50mL, 100mL, and 250mL Single Pack Mini Bags. | Baxter, VIALMATE Adaptor | 2 | November 26, 2013 | 2013 | Baxter Healthcare Corp. | Incomplete foil seal on one lot of sterile product. | PRODUCTION CONTROLS: Packaging Process Control | Baxter Healthcare sent an Urgent Product Recall letter dated September 25, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected product from their facility, contact Baxter's Center for Service at 1-888-229-0001 to arrange for return and credit. Complete the attached Customer Reply Form and return it to Baxter by fax or email. Customers with questions were instructed to call 1-800-422-9837. For questions regarding this recall call 224-948-4770. |  | 38400 | 38,400 units | Nationwide Distribution including WV, KY, IA, AR, OH, IL, IN, MI, KS, TN, WI, MN, MO. | Open |  |  | 38400 | Not_Computer
Z-0428-2014 | 66509 | crown, preformed | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE, Unitek Stainless Steel Permanent Molar CrownsStainless Steel Crowns, are designed to fill the need for long term provisional coverage on adult molars and bicuspids as well as pediatric posterior and anterior teeth. | 3M Unitek" Primary Stainless Steel Crowns | 3 | November 27, 2013 | 2013 | 3M Company - Health Care Business | 3M is recalling specific lots of 3M ESPE Unitek Primary Stainless Steel Crowns - Series UR4 contain LL3 crowns which are incorrectly labeled as UR4 crowns. | PRODUCTION CONTROLS: Process Control | 3M Company sent a Voluntary Recall letterdated June 4, 2013, to all affected consignees. The letter described the problem and the product involved in the recall. Advised consignees to review their inventory and quarantine the product and do not further distribute it. For replacement they should call Customer Service at 3M ESP Dental Products at 1-800-237-1650. If they have further distributed the product they requested they contact their customers to inform them of the recall. 3M prepared a letter to be used in their communications with their customers. The customers should contact 1-800-634-2249.For questions regarding this recall call 651-733-5718. |  | 1419 | 1419 (1325 US, 94 OUS) | Worldwide Distribution - USA including IN, TX, PA, FL, NV, SC, and TN and Internationally to Canada and Saudi Arabia. | Terminated | December 11, 2013 | 14 | 1419 | Not_Computer
Z-0405-2014 | 66577 | orthopedic manual surgical instrument | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Implant Holder for Synfix (TM)-LRthe Synthes Implant Holder for Synfix (TM)-LR is a non-powered hand-held device intended for medical purposes to manipulate tissue, or for use with other devices in orthopedic surgery | Synthes Implant Holder for Synfix (TM)LR | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received which described the SynFix LR implant holder breaking at the interface between the implant and the holder. | OTHER/UNDETERMINED: Unknown/Undetermined by the firm | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 548 | 548 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Open |  |  | 548 | Not_Computer
Z-0420-2014 | 66864 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SIGMAÂŽ LCSÂŽ HIGH PERFORMANCE"MBT PUNCH CEMENTEDSize 4-7HP M.B.T. Keel Punch Knee Instrument is used to prepare the proximal tibial canal geometry to accept the final or definitive tibial tray used during total knee arthroplasty. | SIGMA LCS HIGH PERFORMANCE" | 2 | November 27, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary device correction of the HP M.B.T. Keel Punch Knee Instrument because of the potential for tabs on the connection end to fracture and be left in the patient. | DESIGN: Component Design/Selection | DePuy Orthopaedics sent an Urgent Information - Medical Device Correction letter to all affected customers on November 18, 2013. Immediately after initiation on November 18, the company halted initiation with our US distributors and international affiliates because the surgical technique referenced and images from the surgical technique were not correct. No surgeons received the November 18 surgeon letter. The recall was re-initiated on November 20, 2013. The communication to the surgeons who used the affected devices will be delivered by Distributors either by mail or in person and includes the surgeon letter.The device correction includes awareness, reiteration of surgical techniques and instrument care, and visual check of the device as outlined in the Field Notice and Surgeon letter. The devices are not being removed from the market, but the company is investigating potential design changes to the HP M.B.T. Keel Punch Knee Instrument to reduce the potential for the tabs to fracture. All DePuy United States Distributors and DePuy affiliate contacts for Canada, Latin America and International will be notified via email. The US Distributors will identify surgeons who use the affected instruments on Distributor Response Card #1. The US Distributor will notify surgeons either by mail or in person with the written communication. Once all device correction activities are complete, the US Distributor will complete and return Distributor Card #2 to DePuy Orthopaedics, Inc.The US distributors are the US consignees and are responsible for notifying all new HP M.B.T. Keel Punch Knee Instrument surgeon users of this device correction.All territories are required to ensure the HP M.B.T. Keel Punch Knee Instrument Device Correction is promptly initiated. Customers were asked to review the instructions included, along with the attached documents, whichcontain important information regarding notification of surgeons affected by this Device Correction.Customers. |  | 2008 | 29,235 units in total / 2008 through November 2013 | Worldwide Distribution - USA (nationwide) including AR,AZ,CA,CO,FL,GA,HI,IA,IL,IN,KY,LA,MA,MD,ME,MI,MN,MT,NC,NV,NY,OH,PA, TX,VA,WA,WI and Internationally to: AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CZECH REPUBLIC DENMARK EGYPT FRANCE GERMANY GREECE HOLLAND HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA MALAYSIA NEW ZEALAND NORWAY P R CHINA PAKISTAN PANAMA POLAND PORTUGAL PUERTO RICO RUSSIA SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TURKEY UAE UK | Open |  |  | 29235 | Not_Computer
Z-0410-2014 | 66414 | injector, vertebroplasty (does not contain cement) | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Vertebroplasty Needle Kit 10g Diamond Tip Intended to be used to inject PMMA cement into vertebral bodies during Vertebroplasty procedure. | Synthes Vertebropasty Needle Kit 10g Diamond Tip | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received of inconsistencies in the expiration dates shown on the packaging of the Vertebroplasty Needle Kit 10g Diamond Tip device. | EXPIRATION DATING: Incorrect or No Expiration Date | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 40 | 40 | Nationwide Distribution including IN, MT, HI, and FL. | Open |  |  | 40 | Not_Computer
Z-0411-2014 | 66654 | screwdriver | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Shortened Plug Driver; Low Volume Specialty Device; Qty:1; Rx Only; Non Sterile; Biomet Spine, Parsippany, NJ 07054The Polaris Spinal Systems is a non-cervical spinal fixation device intended for immobilization and stabilization as an adjunct to fusion as a pedical screw fixation system, a posterior hook and sacral/iliac screw fixation system, or as an anterior or anterolateral fixation system. Pedicle screw fixation is limited to skeletal mature patients for use with autograft and/or allograft. The device is indicated for all the following indications:degenerative disc disease (defined as discogeneic back pain with degeneration of the disk confirmed by history and radiographic studies), spondylolisthesis, and or lordosis) tumor, stenosis, pseudoarthrosis, or failed previous fusion | Biomet The Polaris Spinal System; Shortened Plug Driver | 2 | November 27, 2013 | 2013 | Ebi, Llc | Biomet received a complaint on 5/20/2013 reporting that the tip of the instrument sheared off during a removal case. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Biomet sent an Urgent Medical Device Recall Letters/Fax Back Response Form dated September 24, 2013, via Fed Ex. The customer is asked to examine their inventory and return the recalled product following the attached instructions. The Fax Back Response Form should be completed and returned to 973-887-1347 prior to the return of the product. Any questions please call (973) 299-9300 x 2162. |  | 2 | 2 units | US Distribution to CA | Open |  |  | 2 | Not_Computer
Z-0409-2014 | 66418 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5 mm LCP Curved Condylar Plate Intended for buttressing multifragmentary distal femur fractures. | Synthes 4.5 mm LCP Curved Condylar Plate | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes 4.5 mm LCP Curved Condylar Plate was recalled because a Straight Condylar Plate is etched with a Curved Condylar Plate part number. | PRODUCTION CONTROLS: Process Control | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 13 | 13 | Worldwide Distribution - USA including ID, FL, TN, NY, LA and Internationally to Canada. | Open |  |  | 13 | Not_Computer
Z-0406-2014 | 66413 | plate, fixation, bone | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 mm LC-DCP Plate 6 Holes/51mmProduct Usage: Intended for fractures, osteotomies, and replantations of small bones including the foot, ankle, and hand. | Synthes 2.4 mm LCDCP Plate 6 Holes/51mm | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | A complaint was reported regarding a screw being found in a package labeled for 2.4 mm LC-DCP Plate. | PRODUCTION CONTROLS: Labeling Mix-Ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 13 | 13 | Worldwide Distribution - USA (nationwide) including IN, NY, WA, NE, MD, TX, AL, MI, and Internationally to Canada. | Open |  |  | 13 | Not_Computer
Z-0407-2014 | 66419 | nail, fixation, bone | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Cannulated Humeral Nail Intended to aid in the alignment and stabilization of humeral fractures | Synthes Titanium Cannulated Humeral Nail | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | The packaging incorrectly depicts the Synthes Titanium Cannulated Humeral Nail as a 3 distal hole nail instead of a 2 distal locking hole nail. |  | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have any affected product they should still complete Verification Form and return to Synthes by fax to 1-888-731-7954 or Scan/email: FieldAction@synthes.com.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-519-5000. |  | 14 | 14 | Worldwide Distribution - USA (nationwide) including SC, IA, NC, TX, KY, VA, AL, PA, and Internationally to Canada. | Open |  |  | 14 | Not_Computer
Z-0425-2014 | 65356 | set, i.v. fluid transfer | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 13mm addEASE(TM) Binary Connector with 17 Ga. Needle for use with PAB(R) Containers, BC1300Product is intended for fluid transfer between a Partial Additive Bag (PAB(R)) container and a drug vial. | B. Braun 13mm addEASE(TM) Binary Connector with 17 Ga. Needle for use with PAB(R) Containers, BC2000 | 2 | November 27, 2013 | 2013 | B. Braun Medical, Inc. | Possible solution leakage at the interface of the addEASE(TM) Binary Connector (Product Codes BC2000 and BC1300) and the PAB(R) bag. Based on extensive root cause investigation, it has been determined that the leakage is the result of a misaligned cannula (needle) on the addEASE(TM) Binary Connector, which pierces the PAB(R) bag stopper to access the solution. The misaligned cannula may result i | PRODUCTION CONTROLS: Equipment Maintenance | B Braun sent an Urgent - Medical Device Removal Notification letter dated April 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by he customer. Customers were instructed to review the Product Recall Notification in its entirety and ensure that all users in their organization are informed about this voluntary product recall. Determine their current inventory of the affected lots within their distribution center. Contact their customers that have received the affected lots and inform them of this voluntary product recall. Combine both their inventory, and the inventory reported by their customers, and utilize their specific "Product Removal Acknowledgement" form to record the total number of individual units (within partial cases) and the number of full-unopened cases. If no inventory remaining enter zero (0). Return the completed Acknowledgement Form to BBMI Quality Assurance in the self addressed envelope or fax to 610-849-1197. Customers should call Customer Support Department at 800-227-2862 to arrange for eturn and replacement product or for questions.For questions regarding this recall call 800-227-2862. |  | 4000 | 4,000 units | Nationwide Distribution | Open |  |  | 1750079 | Not_Computer
Z-0408-2014 | 66412 | bit, drill | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Slipped Capital Femoral Epiphysis System Product Usage: Intended for fracture fixation of large bones and large bone fragments and for slipped capital femoral epiphysis, pediatric femoral neck fractures, intercondylar femur fractures and sacroiliac joint disruptions. | Synthes Slipped Capital Femoral Epiphysis System | 2 | November 27, 2013 | 2013 | Synthes USA HQ, Inc. | There is a possibility for the Guide Wires of the Slipped Capital Femoral Epiphysis System to stick in the cannulation of the drill bit and the drill bits to break during surgery. | DESIGN: Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 166 | 166 | USA Nationwide Distribution | Open |  |  | 166 | Not_Computer
Z-0429-2014 | 66530 | retractor | GAD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible.Offset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion. | Synthes Soft Tissue Retractor Small Extendible | 2 | November 29, 2013 | 2013 | Synthes USA HQ, Inc. | The small extendible Soft Tissue Retractor was recalled due to locking nut (clamping sleeve) malfunction. | DESIGN: Device Design | Synthes sent an Field Safety Notification letter dated August 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following actions:Exchange the old surgical technique guide with the new guide version F provided with this notification.Review the revised surgical technique guide.Forward this Field Safety Notification to anyone in your facility that needs to be informed.If the technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all technique guides have been exchanged.Maintain a copy of this notice.Please complete and return this page by either of the following methods: Fax: 877-567-9320 or Scan/email: FieldAction@synthes.comThank you for your attention to this issue. |  | 374 | 374 | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Mexico. | Open |  |  | 374 | Not_Computer
Z-0436-2014 | 66866 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Assembly, Patient Side Manipulator on da Vinci Surgical Systems:da Vinci S Surgical System;da Vinci SI Surgical Systemda Vinci Si-e Surgical System;Manufactured by Intuitive SurgicalSunnyvale, CA 94086.The Intuitive Surgical Endoscopic Instrument Control System (Intuitive Surgical da Vinci Surgical System Model IS2000/3000) is intended to assist in the accurate control of Intuitive Surgical Endoscopic Instruments including rigid endoscopes, blunt and sharp endoscopic dissectors, scissors, scalpels, ultrasonic shears, forceps/pick-ups, needle holders, endoscopic retractors, stabilizers, electro cautery and accessories for endoscopic manipulation of tissue, including grasping, cutting, blunt and sharp dissection, approximation, ligation, electro cautery, suturing, and delivery and placement of microwave and cryogenic ablation probes and accessories, during urologic surgical procedures, general laparoscopic surgical procedures, gynecologic laparoscopic surgical procedures, general cardiovascular and non-cardiovascular thoracoscopic surgical procedures and thoracoscopically-assisted cardiotomy procedures. The system is indicated for adult and pediatric use | Patient Side Manipulator | 2 | December 03, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of friction within certain instrument arms can interrupt smooth instrument motion. This can be felt by the surgeon as resistance in the movement of the master.In this situation, the instrument can stall momentarily and then suddenly catch-up to the master position if the surgeon pushes through the resistance. | DESIGN: Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Recall" letter dated November 11, 2013 all affected customers via traceable mail. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions: 1. Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to all Surgeons and OR staff who perform da Vinci surgical procedures, Risk Manager, OR Director, Purchasing, and Biomedical Engineering staff. 2. Complete the attached Acknowledgement Form and return it to Intuitive Surgical via Fax to: Intuitive Surgical, Inc., ATTN: REGULATORY AFFAIRSU.S. +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com.3. An Intuitive Surgical Representative will contact your site to schedule the inspection of the instrument arms on your system(s) as indicated above. Please work with the Intuitive Surgical Representative to schedule this activity.4. Please inform appropriate personnel when the correction has been completed.5. Please retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this medical device correction, then please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below. ÂŽ North and South America: 800-876-1310 Option 3 (6 am to 5pm PST) ÂŽ Japan: 0120-56-5635 or 03-5575-1362 (9am to 6pm JST) ÂŽ South Korea: 02-3271-3200 (9am to 6pm KSTJ) ÂŽ Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1386 | 1,386 | Worldwide Distribution: US (nationwide) and to countries of: Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Denmark, France, Germany, India, Israel, Italy, Japan, S. Korea, Luxembourg, Mexico, Monaco, Netherlands, Norway, Romania. Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | February 18, 2014 | 77 | 1386 | Not_Computer
Z-0435-2014 | 66758 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | VersaSafe Extension Set, Model No. 21000M-007 | VersaSafe Extension Set | 2 | December 03, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the VersaSafe Extension IntravascularAdministration Set because it may have excessive adhesivewhich will prevent the male spin collar to operate as intended. | TRAINING: Employee Error | A customer letter and a distributor letter will be distributed on 11/8/13 to all customers who purchased the VersaSafe Extension Set, Model 21000M-07, Lot Number 13066550. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. |  | 3500 | 3500 units | Nationwide Distribution. | Open |  |  | 3500 | Not_Computer
Z-0434-2014 | 66502 | wrap, sterilization | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | STERITITE Containers: STERITITE Container Kit, 1/4 size, 3" high, product code: 13902; STERITITE Mini Mid Size, 3" high, product code: 13903; STERITITE Container Kit, 1/2 size, 4" high, product code: 13904; STERITITE Container Kit, 3/4 size, 4" high, product code: 13905; STERITITE Container Kit, Full size, 4" high, product code: 13906.The ASP STERITITE sealed container system provides an efficient and cost-effective way to package and protect surgical instruments for sterilization, transportation, storage, and aseptic presentation of contents. | STERITITE Containers | 3 | December 03, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the SteriTite containers because they do not have consistent data to show that sterility assurance level of 10(-6) is achieved during the half-cycle test runs on new, unprocessed containers. | DESIGN: Device Design | Advanced Sterilization Products (ASP) sent a Urgent Medical Device Field Safety Notification letter dated October 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers who received the SteriTite Containers.Customers were instructed to:Provide this notice to anyone in your facility that needs to be informed.Maintain a copy of this notice with the affected product.Maintain awareness of this notice.Customers are instructed to contact Stericycle at (888) 912-7092 for additional questions related to the recall letter. Customers are instructed to report any complaints or suspected problems related to the recall by contacting ASP at (888) STERRAD (888-783-7723). |  | 267 | 267 units total (252 units in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Vietnam, Thailand, Malaysia, China, Korea (South), Mexico, Spain, Singapore, Hong Kong and Portugal. | Terminated | January 13, 2014 | 41 | 267 | Not_Computer
Z-0450-2014 | 66759 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set with Texium Close Male Luer, Model No. 24301-0007TThe SmartSite Low Sorbing Infusion Sets with Texium Male Luer are used to administer fluid and medications through a needle or catheter inserted into the patient's artery or vein. | SmartSite Low Sorbing Infusion Sets | 2 | December 04, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer because it is not assembled with low-sorbing tubing. | TRAINING: Employee Error | The firm, CareFusion, sent an "URGENT: Medical Device Recall Notification" letter dated November 8, 2013 to all customers who purchased the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card and return all recalled product directly to the distributor form whom it was purchased. Customers with questions are instructed to contact their distributor directly or call the CareFusion Support Center at (888) 562-6018.If you have recall related questions or need support, please call CareFusion Support Center at 1-888-562-6018; hours: 7am to 4pm PST or Technical Support at 1-888-812-3229; hours 7am to 5pm PST. |  | 480 | 480 units | US Distribution in CA, WY, OH, UT, and MA. | Open |  |  | 480 | Not_Computer
Z-0441-2014 | 66828 | endoscopic access overtube, gastroenterology-urology | FED | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Navigator HD l3F/15F x 36cm Ureteral Access Sheath Set Product No M0062502280, RDF Catalog No 250-228, Sterile.Indicated for use in endoscopic procedures to facilitate the passage of endoscopes, urological instruments and for the injection of fluids into the urinary tract. | Navigator HD | 2 | December 04, 2013 | 2013 | Boston Scientific Corporation | The 13F/15F Navigator HD Ureteral Access SheathSet are being recalled after receipt of 5 reports where the tip of the dilator has separated fro the body of the dilator. | DESIGN: Component Design/Selection | Urgent recall notification letters were sent via Federal Express Priority mail on November 5, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it toBoston Scientific. The notification also requests the completion of a Reply Verification Tracking Form asa method of documenting the presence or absence of affected product in their inventory. Follow-up communications will be made to the accounts that do not respond to this inquiry. |  | 939 | 939 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Ireland, Denmark, Italy, France, Canada, Great Britain, Switzerland, South Africa, Germany, and the Netherlands. | Open |  |  | 1662 | Not_Computer
Z-0451-2014 | 66757 | test, time, prothrombin | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Abbott i-STAT PT/INR cartridges Abbott Point of Care Inc., Princeton, NJ 08540The i-Stat PT a prothrombin time test is useful for monitoring patients receiving oral anticoagulation therapy such as Coumadin or Warfarin. The cartridge is to be used with the i-Stat 1 Analyzer and as part of the i-Stat System, the PT/INR test is to be used by trained health care professionals in accordance with a facility's policies and procedures. | Abbott iSTAT PT/INR cartridges | 2 | December 04, 2013 | 2013 | Abbott Point Of Care Inc. | Abbott Point of Care Inc (APOC) has determined that i-Stat PT/INR cartridges have the potential to exhibit incorrectly elevated INR. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Abbott Point of Care Inc. (APOC) issued an Urgent Recall Notice on October 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from PT/INR lots in this letter. All unused cartridges would be credited and are to be returned as instructed on the enclosed Business Reply Card. If customers provided cartridges from the lot number listed above to another institution they were instructed to provide a copy of this letter to them.Customers with questions were instructerd to contact Abbott Point of Care Technical at 1-800-366-8020, Option 1.For questions regarding this recall call 800-366-8020. |  | 914400 | 794,928 units US; 119,472 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Chile, China, Finland, France, Germany, Greece, Hong Kong, Ireland, Italy, South Korea, Kuwait, Mexico, Netherland, New Zealand, Nigeria, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, south Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE, and United Kingdom, | Open |  |  | 914400 | Not_Computer
Z-0448-2014 | 66333 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, PS HIGH FLEXION ARTICULAR INSERT, SIZE 5-6, 11 MM QTY: (1), REF 71421516, STERILE EO, Smith & Nephew, Inc. Memphis, TN 38116 USAOrthopedic | GENESIS(R) II, PS HIGH FLEXION ARTICULAR INSERT | 2 | December 04, 2013 | 2013 | Smith & Nephew Inc | The Tyvek(TM) lid on one, or both, of the two trays that form adouble barrier around the insert may not have been adequately sealed. | DESIGN: Packaging Design/Selection | Smith & Nephew sent an Urgent Recall Notification letter dated August 27, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. They were instructed to cease use and distribution of the product and quarantine for return. Customers were instructed to contact Smith & Nephew via email at FieldActions@smith-nephew.com or fax 901-566-7975 to obtain a return authorization (RA) number. Customers with questions were also instructed use this email address. Customers with affected product were instructed to return it along with the completed inventory Return Certification Form to the address listed on the form. Non-responding consignees were notified again on September 27, 2013, by email.For questions regarding this recall call 901-396-2121. |  | 46 | 46 pieces | Worldwide Distribution - USA including Nevada and Internationally to United Arab Emirates, Italy, Spain, Belgium, Bulgaria, Poland, France, Japan, Netherlands, Australia, Turkey, Russian Federation, Germany and United Kingdom. | Open |  |  | 227 | Not_Computer
Z-0455-2014 | 66538 | intervertebral fusion device with integrated fixation, cervical | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P VA Implant 7 MM Height Parallel SterileThe Synthes Zero-P VA Implant 7 MM Height Parallel Sterile is a stand alone anterior cervical interbody fusion device indicated for use in skeletally mature patients with degenerative disc disease (DDD) with accompanying radicular symptoms at one level from C2-T1. | Synthes ZeroP VA Implant 7 MM Height Parallel Sterile | 2 | December 05, 2013 | 2013 | Synthes USA HQ, Inc. | A potential non conformance was identified with the Zero-P VA plate where the screw could back out if the blocking mechanism is not visually confirmed to properly cover the head of the screw. | DESIGN: Process Design | Synthes sent an email message dated March 23, 2012, to the sales representative who received the affected product. The email identified the product, the problem, and the action to be taken. The letter requested the sales rep examine inventory, stop using the product, and quarantine the affected product to ensure that it will not be used. Synthes instructed the customer not to remove any product from their account or return any product without further instruction. For questions they were instructed to call 610-719-5364.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | Nationwide Distribution - only AZ. | Open |  |  | 1 | Not_Computer
Z-0453-2014 | 66896 | wheelchair, mechanical | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Brand Zippie Voyage, Chair with Casters, Early Intervention Device (Stroller), Model: EIZ14; Product is manufactured and distributed by Sunrise Medical (US) LLC., Fresno, CA.Used by a person to move from one place to another while sitting. | Sunrise Brand Zippie Voyage | 2 | December 05, 2013 | 2013 | Sunrise Medical (US) LLC | There is a remote possibility for the plastic latch to deform over the lifetime of use. When this happens there is the possibility for the seat to become dislodged from the base. And when not using the positioning strap, there is a risk that the occupant could slide out of the seat and sustain injury. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Urgent Medical Device Field correction letters were sent to Sunrise Suppliers/Dealers and Zippie Customers on November 2013 by first class mail. Foreign notification was supplied to suppliers and regulatory affairs director in Canada. |  | 316 | 316 units in US, 29 units international | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Hong Kong, and China. | Open |  |  | 316 | Not_Computer
Z-0432-2014 | 66407 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | TB-0545PC: THUNDERBEAT 5 mm , 45 em, Pistol Grip TB-0535PC: THUNDERBEAT 5 mm, 35 em, Pi stol Grip TB-0545IC: THUNDERBEAT 5 mm, 45 em, In line Grip TB-0535IC: THUNDERBEAT 5 mm, 35 em, Inline Grip TB-0520IC: THUNDERBEAT 5 mm, 20 em, In line Grip TB-051 OIC: THUNDERBEAT 5 mm, I 0 em, In line GripTB-545FC: THUNDERBEAT 5 mm, 45 em, Front Actuated GripTB-0535FC: THUNDERBEAT 5 mm, 35 em, Front Actuated GripThe Olympus Thunderbeat Surgical Tissue Management System ("Thunderbeat") is intended to be used for open, laparoscopic (including single-site surgery), and endoscopic surgery to cut (dissect) or coagulate soft tissue or to ligate (seal and cut) vessels. | THUNDERBEAT | 2 | December 05, 2013 | 2013 | Olympus America Inc. | Complaints of damage to various models of Thunderbeat Hand Instruments during surgical procedures, including cracking, breaking, or deformation of the components of the probe tip or jaw. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Olympus sent an Urgent Medical Device Field Correction Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the actions needed to be taken by the customer.Please make the attached Procedural Tips and Techniques instructions available to all users of the THUNDERBEAT hand instruments. Please also maintain a copy with the Instruction Manual(s) fot"the device.OCA requires that you indicate on the enclosed questionnaire your acknowledgement of receipt of this corrective action and the contact information for your facility.This additional user instruction is also being added to the Instruction Manual(s) for the THUNDERBEAT hand instruments. If you would like a copy of the updated Instruction Manual(s), please contact OAI Customer Service at 1-800-848-9024.For further questions please call (484) 896-5688. |  | 8709 | 8709 | Worldwide Distribution - USA (nationwide) including Puerto Rico and Internationally to Canada. | Open |  |  | 8709 | Not_Computer
Z-0452-2014 | 66417 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Titanium Trochanteric Fixation Nail (TFN)SystemIntended to treat stable and unstable fractures of the proximal femur. | Synthes Titanium Trochanteric Fixation Nail (TFN)System | 2 | December 05, 2013 | 2013 | Synthes USA HQ, Inc. | Two lots of Titanium Trochanteric Fixation Nail (TFN) System are missing the second bend in the nail. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affeccted product to obtain a Return Authorization Number, Complete the Verification Section of the letter by checking the appropriate box indicating the affected product has been located. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132.Customers should complete and return the Verification Frorm even if they do not have any of the affected product and return by fax to 1-888-386-2077 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Trauma Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 17 | 17 | Nationwide Distribution including IL, CA, AZ, FL, TX, UT, VA, NY, NM, and MT. | Open |  |  | 17 | Not_Computer
Z-0457-2014 | 66431 | apparatus, suction, operating-room, wall vacuum powered | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Product is labeled in part: LIST NO. 43449; RECEPTAL CANISTER; DO NOT FILL 900 ABOVE THIS LINE; Product of China; Hospira, Inc.,; Lake Forest, IL 60045 USA.Used to isolate suction waste during surgeries, or to clear patient airways, throat or wounds to allow treatment. | RECEPTAL CANISTER | 2 | December 06, 2013 | 2013 | Hospira Inc. | Hospira has recieved reports that the vacuum needed to create suction with the 1L Receptal System cannot be achieved due to the canister and liner not properly seating. | PRODUCTION CONTROLS: Process Control | An URGENT DEVICE RECALL letter dated July 17, 2013 was mailed to direct consignees. The letter requested the consignees: 1) quarantine any affected products, 2) complete and return the reply form included with the letter, 3) notify healthcare professionals within the consignee's organization of the recall, 4) return the affected products to Stericycle in accordance with the instructions in the letter. Consignees are further requested to forward a copy of the notification materials to any customers that they may have distributed the products to. |  | 7576 | 3,560 units distributed from June 2004 - December 2012 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NY, PA, SD, TX, UT, VA, and WA and countries of: Canada and EMEA. | Open |  |  | 3560 | Not_Computer
Z-0458-2014 | 66627 | suture, surgical, absorbable, polydioxanone | NEW | Absorbable polydioxanone surgical suture. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device Ethicon LLC Guaynabo, Puerto Rico 00969.Intended for general soft tissue approximation where use of an absorbable suture is appropriate. | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device | 2 | December 06, 2013 | 2013 | Ethicon, Inc. | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device has a small number of tab failures and fascial dehiscences in lower abdominal incisions. |  | Ethicon Inc. sent recall notifications/business reply forms on 9/25/2013 via UPS next day certified mail. |  | 37164 | 37,164 eaches | Worldwide Distribution-USA (nationwide) and countries of: Caribbean, Chile, Singapore and Malaysia. | Open |  |  | 37164 | Not_Computer
Z-0461-2014 | 66809 | template | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | NexGenÂŽ Complete Knee SolutionCR Articular Surface Provisional-Purple 9mm HeightNonsterileDevices are used in total knee arthroplasty and indicated for patients with severe pain and disability due to conditions such as arthritis, collagen disorders, post-traumatic loss of joint configuration, flexion deformities, and in revision of previous surgical attempts. | NexGen Complete Knee Solution | 2 | December 06, 2013 | 2013 | Zimmer, Inc. | This is an expansion of the June 2013 recall related to NexGen Micro components due to a component compatibility issue. Since the initiation of this recall, five (5) additional instruments have been identified that are associated only with the NexGen Micro components. Given this, all lots of these additional instruments are now being recalled as well. There is the potential for these devices | DESIGN: Labeling Design | Zimmer sent an Urgent Medical Device Recall Notice dated November 19, 2013, to all affected customers. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product will be removed by a Zimmer Sales representative, Question should be directed to 1-877-946-2761 8a.m - 5 p.m. EST.For questions regarding this recall call 800-613-6131. |  | 638 | 638 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Denmark, Ecuador, Germany, India, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Peru, Singapore, South Africa, Spain, Switzerland, UAE, and UK. | Open |  |  | 3027 | Not_Computer
Z-0466-2014 | 66948 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Vasoview Hemopro Endoscopic Vessel Harvesting System Maquet Cardiovascular LLC 45 Barbour Pond Drive Wayne, NJ 07470 Indicated for use in minimally invasive surgery allowing access for vessel harvesting and is primarily indicated for patients undergoing endoscopic surgery for arterial bypass. It is indicated for cutting tissue and controlling bleeding through coagulation, and for patients requiring blunt dissection of tissue including dissection of blood vessels, dissection of blood vessels oft he extremities, dissection of ducts and other structures int he extraperitoneal or subcutaneous extremity and thoracic space. Extremity procedures include tissue dissection/vessel harvesting along the saphenous vein for use in coronary artery bypass grafting and peripheral artery bypass or the radial artery for use in coronary artery bypass grafting. Thorascopic procedures include exposure and dissection of structures external to the parietal pleura including nerves blood vessels and other tissues of the chest wall. | Vasoview Hemopro Endoscopic Vessel Harvesting System | 2 | December 06, 2013 | 2013 | Maquet Cardiovascular, LLC | Maquet has received an increase in the number of complaints for "self-activated/remains activated/overheats" for the Hemopro VH-30000, "Hemopro 1" | DESIGN: Component Design/Selection | Maquet Inc. sent an Urgent - Medical Device Recall (Removal) Immediate Action Required Letter dated October 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please examine your inventory immediately to determine if you have any of the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System (HEMOPRO 1) VH-3000 with the lot number(s), as specified above. The VH-3000 catalog number and lot number are located on the package label. Please remove the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System and place in a secure location.NOTE: VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 lot numbers not listed are not affected by this recall and should not be removed. HEMOPRO 2 is not affected by this recall.Your MAQUET Cardiovascular Sales Representative will contact you shortly to assist with the completion of the enclosed Acknowledgement Form ( Fax Back Form), the return of any affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 product and arrange for necessary replacement.We apologize for any inconvenience this may cause. If you have any questions, please contact your local MAQUET Cardiovascular Sales Representative or MAQUET Customer Service at 1-888-880-2874, Monday through Friday, between the hours of 6:00 a.m. and 5:00 p.m. PST.Thank you for your cooperation and immediate assistance. |  | 12259 | 12,259 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Czech Republic, Denmark, France, Hong Kong, Italy, Middle East, Netherlands, Singapore, Switzerland, and the United Kingdom. | Open |  |  | 12259 | Not_Computer
Z-0465-2014 | 66823 | laparoscope, general & plastic surgery | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Axillent Debakey Grasping Forceps, Part No. 33821DYDIn general the surgeons use 33821DYD Debakey forceps to grasp tissue. | Axillent Debakey Grasping Forceps | 2 | December 06, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ is recalling the Axillent Debakey Grasping Forceps because the 33821WTD Wavy grasper was mislabeled as 33821DYD Axillent Debakey Grasping Forceps. | PRODUCTION CONTROLS: Labeling Mix-Ups | Karl Storz Endosocopy sent a Safety Alert letter dated November 6, 2013, to all customers who received the 33821DYD Axillent Debakey Grasping Forceps. The letter informed the customers of the problems identified and the actions to be taken. Customers were asked to set aside the mislabeled product and return the mislabeled instruments with the RMA (Return Material Authorization) number issued to them. Customers were instructed to fax or email completed recall response form to (424) 218-8519. Customers with any questions were instructed to call (800) 421-0837 ext 8289.For questions regarding this recall call 746-708-8440. |  | 7 | 7 units | Nationwide Distribution including states of: TX, PA, and NY. | Terminated | January 13, 2014 | 38 | 7 | Not_Computer
Z-0464-2014 | 66830 | [6 empty] | Centurion Sterile # 84 Rubber Bands Reorder EB84, Caution: This product contains natural rubber latex which may cause allergic reactions. LATEX FREE , CAUTION: FEDERAL lAW (USA) RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIANSingle Use Onlybands items together and podiatry office uses bands as tourniquet on toe during in grown toenail procedure | Centurion Sterile 84 Rubber Bands | 2 | December 06, 2013 | 2013 | Centurion Medical Products Corporation | Package labeling indicates both "latex free" and "contains natural rubber latex" . The rubber bands do contain natural rubber latex. This could cause a significant risk to users with latex allergies. | MISBRANDING: Labeling False and Misleading | Centurion sent a Urgent Recall Notification letter via Certified Mail October 31, 2013, return receipt to all affected customers. The affected Centurion Medical Products Corporation sales representatives were notified via email on October 28, 2013. Customers were instructed to destroy all implicated product and complete the accountability record included with the notice and fax to 517-546-3356. Customers were asked to forward a copy of the notice if product was further distributed. Additional notices will be mailed to non-responsive customers via Certified Mail ReturnReceipt, and will be documented in the recall file. For further questions please call (517) 546-5400. |  | 500 | 500 lots | US Distribution including the states of GA, LA and NY. | Terminated | January 29, 2014 | 54 | 500 | Not_Computer
Z-0467-2014 | 66761 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 1) Accolade II 127 Neck Angle Hip Stem, Size # 32) Accolade II 127 Neck Angle Hip Stem, Size #5Howmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430 USATelephone # +1 201-831-5000Stryker FranceZAC Satolas Green PusignanAv de Satolas Green69881 MEYZIEU Cedex FranceIntended for cementless application in total or hemiarthroplasty procedures. | Accolade II 127 Neck Angle Hip Stem | 2 | December 07, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that an Accolade II Hip Stem Size #3 was identified in packaging with an Accolade II Hip Stem Size #5. | PRODUCTION CONTROLS: Packaging Process Control | Stryker sent an e-mail to their domestic locations on 10/11/2013. Notification letters and product accountability forms were sent via Fed Ex on 10/25/2013 with return receipt to branches/agencies and hospital risk management, chief of orthopaedics and surgeons. |  | 20 | 20 units | Nationwide Distribution including states: AL, GA, IL, and WA. | Open |  |  | 20 | Not_Computer
Z-0468-2014 | 66586 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arrow Transradial Artery Access productsArrow Transradial Artery Access products are used for percutaneous introduction of devices into the radial artery for diagnostic and therapeutic purposes. | Arrow Transradial Artery Access products | 2 | December 09, 2013 | 2013 | Arrow International Inc | There is a risk that the introducer needles packaged within the kits can be partially or fully occluded, which can cause delayed treatment or an adverse event when used on a patient. | PRODUCTION CONTROLS: Process Control | Arrow International sent an Urgent Medical Device Recall Notification letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product,to return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, ATTN: Customer Service. Customers were asked to completed the Recall Acknowledgement Form even if they have none of the affected product.Customers with questions were instructed to contact their local sales representative or Customer Service at 1-800-343-2935.For questions regarding this recall call 610-378-0131. |  | 294430 | 294430 | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Brazil, Canada, China, Ireland, New Zealand, Singapore, and Thailand. | Open |  |  | 294430 | Not_Computer
Z-0427-2014 | 66674 | infrared hematoma detector | OPT | Near Infrared (NIR) Brain Hematoma Detector. | Ophthalmic | Ophthalmic | 510(k) | Cutera brand truSculpt infrared System (optional), Massager, Vacuum, radio frequency induced heat, Product is manufactured and distributed by Cutera, Inc., Brisbane, CAThe truSculpt massage device is intended to provide a temporary reduction in the appearance of cellulite. The truSculpt infrared (optional) and RF energies are intended to provide topical heating for the purpose of elevating tissue temperature for the treatment of selected medical conditions such as relief of pain, muscle spasms, and increase in local circulation. | truSculpt | 2 | December 09, 2013 | 2013 | Cutera, Inc. | Cutera received two adverse events associated with the edge of the electrode, when it was pressed against neck skin over the thyroid cartilage and due to the absence of adequate underlying conductive tissue, increased skin heating occurred. Both patients developed linear welts and blisters consistent with partial thickness thermal burns to the skin. | TRAINING: Employee Error | Cutera sent an IMPORTANT SAFETY INFORMATION letter dated October 16, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. It is important that you share this correspondence and the updated Operator Manual andvtreatment guidelines with all staff at your office that use the truSculpt, so they are aware of this new safety information. Please immediately remove and dispose of the current truSculpt Operator Manual and clinical guidelines and replace them with the new and revised truSculpt Operator Manual and clinical guidelines.Customers with questions were instructed to call 415-657-5500 or 415-657-5568. |  | 451 | 451 | Worldwide Distribution - USA (nationwide) and Internationally to Canada | Terminated | December 09, 2013 | 0 | 451 | Not_Computer
Z-0473-2014 | 66582 | appliance, fixation, nail/blade/plate combination, multiple component | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes TomoFix PlatesThe Synthes J5606-C TomoFix Osteotomy System is used for open and closed wedge osteotomies of the medial proximal tibia, lateral proximal tibia, medial and lateral distal femur, treatment of bone and joint deformities, fractures, and malalignment caused by injury or disease such as osteoarthritis. | Synthes TomoFix Plates | 2 | December 10, 2013 | 2013 | Synthes USA HQ, Inc. | A labeling correction was initiated related to the Surgical Technique Guide J5606-C "TomoFix Osteotomy System" due to the potential for a surgical delay if the surgeon inserts the superior screws at a lesser angle (4 or 6 degrees) and is not aware that those angles are not present. | MISBRANDING: Labeling False and Misleading | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter dated October 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exchange the old surgical technique guide version J5606-C with the new version J5606-D provided with this notification. Review the revised surgical technique guide (Pages 25 & 26). Forward this Field Safety Notification to anyone in their facility that need to be informed.If the surgical technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all surgical technique guides have been exchanged.Maintain a copy of this notice. Customers with questions were instructed to contact their SynthesSales Consultant. For questions regarding this recall call 610-719-5000. |  | 2912 | 2912 | Worldwide Distribution - USA (nationwide) Canada and Mexico | Open |  |  | 2912 | Not_Computer
Z-0470-2014 | 66826 | counter, differential cell | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH 750 Hematology Analyzer, Part No. 6605632,A85570Product Usage:The Coulter LH 750 Hematology Analyzer is a quantitative, automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The CoulterÂŽ LH 750 Hematology Analyzer provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids | Coulter LH 750 Hematology Analyzer | 2 | December 10, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the LH 750 and LH 780 Hematology Systems because the optical degradation may potentially cause a delay in reporting results for nucleated red blood cells (nRBC), white blood cell differential and/or reticulocytes. |  | Beckman Coulter sent an customer notification letter dated 11/8/13 to all customers who purchased the LH 750 and LH 780 Hematology Systems. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Support Center at http://www.beckmancoulter.com/customer/support or call (800) 526-7694 in the US and Canada. Outside the US and Canada, customers are instructed to contact Beckman Coulter Representative.On 2/17/14, Beckman Coulter expanded the recall to inform Coulter LH 750 and Coulter LH 780 Hematology Analyzer customers of the issue related to LH Cleaner and Coulter Clenz. |  | 4155 | 4155 units total (1676 units in US) | Worldwide Distribtion - USA Nationwide including the countries of Algeria, Andorra, Angola, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bosnia and Herzegovina, Bulgaria, China, Colombia, Croatia, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Virgin Islands, U.S., and Zimbabwe. | Open |  |  | 4155 | Not_Computer
Z-0474-2014 | 66409 | rod, fixation, intramedullary and accessories | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trochanteric Fixation Nail As part of the Synthes Trochanteric Fixation Nail (TFN) System, the TFN Nail is intended to treat stable and unstable fractures of the proximal femur. | Synthes Trochanteric Fixation Nail | 2 | December 10, 2013 | 2013 | Synthes USA HQ, Inc. | Certain lots of Trochanteric Fixation Nail were distributed to the field with the incorrect labeling of the length and diameter of the nail on the packaging. | TRAINING: Employee Error | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Call Syntehs at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section checking the appropriate box indicating affected product has been located.Return the Verfication Form with the product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If no affected product found customers were instructed to take the following steps:Complete the attached Verification Section at the end of the letter by checking the appropriate box indicating no affected product has been located.Return the documents to Synthes by fax to 610-251-9005 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | Nationwide Distribution including WA, FL, LA, TX, GA, IN, OH, CA, MI, and PA. | Open |  |  | 12 | Not_Computer
Z-0477-2014 | 66717 | prosthesis, ankle, semi-constrained, cemented, metal/polymer | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular Metal Total Ankle Tibial Base Component Sizes 1-6Total ankle arthroplasty is intended to provide a patient with limited mobility by restoring alignment, reducing pain and preserving the flexion/extension movement within the ankle joint.The Zimmer Trabecular Metal Total Ankle is indicated as a total ankle replacement in primary or revision surgery for patients with: " Rheumatoid arthritis. " Post-traumatic arthritis. " Degenerative arthritis. This device is intended for cemented use only. | Zimmer Trabecular Metal" Total Ankle Tibial Base and Talar Components | 2 | December 10, 2013 | 2013 | Zimmer, Inc. | As part of routine manufacturing process monitoring by Zimmer, it was discovered that the Tibial Base and Talar Components from the Total Ankle system have the potential of containing elevated levels of manufacturing materials residue on the implants. The specific material was found to be mass media utilized during the manufacturing process to provide the uniform surface finish. This mass media | PRODUCTION CONTROLS: Manufacturing Material Removal | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification letter dated November 12, 2013 to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to locate and remove the affected product from their inventory and notify their Zimmer representataive. The Zimmer representative will remove the recalled product from their facility. For patients that previously had the affected product implanted, consignees were instructed to continue their operative follow up routine. Consignees with questions were instructed to call 1-877-946-2761.For questions regarding this recall call 800-613-6131. |  | 390 | 390 units (US) 129 units (outside of US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, China, Finland, Germany, and Italy. | Open |  |  | 924 | Not_Computer
Z-0483-2014 | 66908 | screw, fixation, bone | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone ScrewHowmedica Osteonics Corp.325 Corporate DriveMahwah, NJ 07430 USAStryker FranceZAC Satolas Green PusignanAv de Satolas Green 69881 Meyzieu Cedex FranceBoth the 6.4 mm diameter, 25mm length, Cancellous Bone Screws (Cat no 2030-6525-1) and the 5.5mm diameter, 40 mm length Cancellous Bone Screws (Cat no 2029-5540-1) are cleared for use in Total Hip Arthroplasty (THA), when implanting Stryker's Hemispherical Cluster Hole Acetabular Shell Systems and Total Knee Arthroplasty (TKA) when implanting Stryker's Series 7000 Standard Tibial Tray systems. The use of bone screws, during both THA and TKA are intended as supplemental fixation. | Stryker Orthopaedics | 2 | December 11, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a 5.5 mm Cancellous Bone Screw 40 mm was identified in packaging associated with a 6.5 mm Cancellous Bone Screw 25 mm. | PRODUCTION CONTROLS: Labeling Mix-Ups | Stryker sent notification letters/accountability forms via Fed Ex on October 1, 2013, with return receipt. to branches/agencies and hospital risk management, chief of orthopedics and surgeons.Domestic branch locations were notified by e-mail on 9/16/2013.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax the Product Recall Acknowledgement Form to 855-251-3635, return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430. Customers were instructed to attach the fluorescent organge PRODUCT REMEDIATION sticker to their return, indicating the Produc Remediation # RA 2013-145. Mark the outer box with the words "Product Recall."Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5000. |  | 4 | Catalog Number 2030-6525-1- 4 units; Catalog No: 2029-5540-1-7 units | Nationwide Distribution | Open |  |  | 11 | Not_Computer
Z-0478-2014 | 66874 | sterilizer, steam | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCOÂŽ EvolutionÂŽ steam sterilizer Transfer Carriage. Product Usage:Mobile, wheeled transportation unit utilized to carry, load, &#38; unload the Loading Car into and out of steam sterilizer units. | AMSCO Evolution steam sterilizer Transfer Carriage | 2 | December 11, 2013 | 2013 | Steris Corporation | The firm discovered that unlatching can occur when overloaded transfer carriages are pulled or maneuvered incorrectly or with excessive force. | DESIGN: Device Design | Steris sent an Urgent Voluntary Field Correction Notice letter dated October 18, 2013 to their customers. The letter identified the affected product, problem, and actions to be taken. The letter informed customers that a Steris Service Representative will contact affected facility to arrange for Steris to upgrade the Transfer Carriage(s). For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 936 | 936 units | Worldwide Distribution - USA Nationwide in the states of: AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MI, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and the District of Columbia including the countries of: Brazil, Canada, Chile, Guam, Mexico, Saudi Arabia, United Arab Emirates & Venezuela. | Open |  |  | 936 | Not_Computer
Z-0480-2014 | 66714 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂŽ Complete Knee Solutions Stemmed Tibial Component PrecoatNexGen Tibial Baseplates are part of the NexGen systems of semi-constrained, non-linked condylar knee prosthesis. The device is indicated for patients with severe knee pain and disability due to: Rheumatoid arthritis, osteoarthritis, traumatic arthritis, polyarthritis; Collagen disorders, and/or avascular necrosis of the femoral condyle; Post-traumatic loss of joint configuration, particularly when there is patellofemoral erosion, dysfunction or prior patellectomy; Moderate valgus, varus, or flexion deformities. | NexGen Complete Knee Solutions Stemmed Tibial Component Precoat | 2 | December 11, 2013 | 2013 | Zimmer, Inc. | Two complaints were received in which it was reported that the stem would not lock into the tibia. Evaluation of the returned device found that the stemmed portion of the tibial plate would not hold the stem as the connection was too loose. The taper diameter was found to be oversized by .010 . | DESIGN: Device Design | All distributors will be notified via electronic mail. Hospital risk managers and distributors with affected inventory will also be notified via courier. Surgeons that have implanted this manufacturing lot will be identified and notified by the distributor. Distributors will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying and notifying surgeons that have implanted this manufacturing lot. Distributors will return on-hand affected products to Zimmer and ensure all of their products are accounted for using Attachment 1 provided in the letter. Distributors are to complete and return Attachment 2 indicating which surgeons have implanted the product and to which they have provided notice to. Hospitals that received direct shipments from Zimmer will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative.Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitalso 100% of the FedEx notifications will be accounted foro An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.o Additional notifications via email and FedEx will be sent to those who have not completed the required certification formo Accounts will be deemed unresponsive after 3 attempts.Further questions please call 800-613-6131. |  | 128 | 128 units | Worldwide Distribution - USA (nationwide) including the states of CA, IN, OH, GA, WI, IA, NE, NJ, PA, NY, MD, NC, AZ, NM, TX, CO, IL, AL, MI, WA, FL and KS, and the countries of China, Japan and Taiwan. | Open |  |  | 128 | Not_Computer
Z-0482-2014 | 66834 | system, automated platelet aggregation | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow IIb/IIIa Test, 10-Test Kit, Catalog No. 85310,10 tests/box; 690 total tests.The VerifyNow IIb/IlIa Test is a semi-quantitative, whole blood platelet function test used to measure glycoprotein (GP) lIb/IlIa receptor blockade in patients treated with abciximab or eptifibatide. VerifyNow lIb/IlIa Test results should be interpreted in conjunction with other clinical and laboratory data available to the clinician. The VerifyNow System is a turbidimetric based optical detection system which measures platelet-induced aggregation. The system consists of an instrument, a disposable test device and quality control materials. The VerifyNow lIb/IlIa Test is a semi-quantitative, whole blood platelet function test used to measure glycoprotein (GP) lIb/IlIa receptor blockade in patients treated with abciximab or eptifibatide. | VerifyNow IIb/IIIa Test, 10Test Kit | 2 | December 11, 2013 | 2013 | Accumetrics Inc | Accumetrics is recalling the VerifyNow IIb/IIIa 10-Test because it contains incorrect information for the amount of time in which to perform the test following collection of the patient sample. The correct instruction is to perform the test within 15 minutes of collecting the patient sample. | PRODUCTION CONTROLS: Error in Labeling | Accumetrics sent an Urgent Medical Device Correction letter dated November 14, 2013 to end users for the VerifyNow IIb/IIIa 10-Test. The letter informs the customers of the problems identified and the actions to be taken. The customers are instructed to call Accumetrics Customer Support at (800) 643-1640 option 2 or email at support@accumetrics.com. Customers are instructed to complete and return the enclosed Customer Account Tracking Form as soon as possible.For questions regarding this recall call 858-643-1600. |  | 69 | 69 units | Nationwide Distribution including NY, PA, DC, CT, IN, NJ, AR, TX, AZ,and NM. | Terminated | December 23, 2013 | 12 | 69 | Not_Computer
Z-0481-2014 | 66978 | bracket, ceramic, orthodontic | NJM | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Opal Orthodontics by Ultradent. Avex CX2 Ceramic Orthodontic BracketsAvex CX2 ceramic brackets are intended for use in orthodontic treatment. | Opal Orthodontics by Ultradent | 3 | December 11, 2013 | 2013 | Ultradent Products, Inc. | Ultradent Products Inc. is recalling various lots of Avex CX2 Orthodontic Brackets because the design control process was not complete prior to distribution. | PRODUCTION CONTROLS: Process Control | Ultradent Products notified consignees via email on November 19, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product. Sales representatives will pick up remaining unplaced brackets or returns will be sent back using pre-paid envelopes back to Ultradent recall coordinator for quarantine. Consignees were asked to complete evaluation forms to be returned by fax to 800-793-5206.For questions regarding this recall call 801-553-4420. |  | 217 | 217 units | Nationwide Distribution including UT, CA, WA, VA, OH, and TN. | Open |  |  | 217 | Not_Computer
Z-0479-2014 | 66533 | electrosurgical, cutting & coagulation & accessories | GEI | Electrosurgical cutting and coagulation device and accessories. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Aesculap GNI161 Bipolar Foot ControlProduct Usage:The Aesculap Bipolar Coagulator is intended for use in surgery to generate electrical power for bipolar instruments. Bipolar coagulators are used in neurosurgery, ENT surgery, urology, laparoscopy, and plastic surgery. | Aesculap GNI161 Bipolar Foot Control | 2 | December 11, 2013 | 2013 | Aesculap, Inc. | Complaints were received for the GN161 Bipolar Foot Control reporting that the bipolar energy did not stop after release of the foot control. | DESIGN: Device Design | Aesculap sent an Important Recall Notification letter dated September 23, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected product and receive replacement product or credit. An Inventory Sheet is attached and must be completed. For questions call 610-984-9265 or 610-984-9414. |  | 85 | 85 | Worldwide Distribution - USA Nationwide and countries of: Austria, Germany, Japan, and Taiwan. | Open |  |  | 85 | Not_Computer
Z-0484-2014 | 66790 | light, surgical, ceiling mounted | FSY | Surgical lamp. | General & Plastic Surgery | General & Plastic Surgery | 510(k) | Helion S Exam Light or H300The device is intended to locally illuminate an examination area of the patient's body with high intensity light in a clinic or doctor's practice. | Helion S Exam Light or H300 | 2 | December 12, 2013 | 2013 | Trumpf Medical Systems, Inc. | Possible fatigue failure of the plastic joint may occur after an average use of seven years. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | TRUMPF sent an Urgent Recall Notice dated February 19, 2013, to all US affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The recall notice instructed customers to verify if they still have these light systems, and if they do have the light system verify if they have a metal or plastic ring at the joint of concern (per the manufacturer's Safety Notice). TRUMPF will directly work with affected customers to replace the affected arm system.For further questions please call ( 888 ) 474-9359. |  | 287 | 287 | US Distribution including the states of AR, CA, CO, KS, MA, MI, MO, NJ, NY, NC, OH, OK, PA, SD, TX and WA. | Open |  |  | 287 | Not_Computer
Z-0502-2014 | 66888 | device, neurovascular embolization | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Stryker brand GDC-10 ULTRA SOFT 2MM X 3CM, Catalog Number: M0033432030; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | GDC10 ULTRASOFT 2MM X 3CM | 2 | December 13, 2013 | 2013 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015. |  | Stryker Neurovascular sent an Urgent Medical Device Recall letter dated November 4, 2013, all affected users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected products. Customers with questions were instructed to contact via NeuroComplaints@stryker.com. Or Fax the response form to 1-800-876-4355.For questions regarding this recall call 510-413-2500. |  |  |  | Worldwide Distribution - USa (nationwide) and Internationally to:Austria, Belgium, Czech Rep, Germany, France, Great Britain, Hungary, Italy, Lithuania, Netherlands, Norway, Poland, Slovakia, Spain, Turkey, Australia, Canada, Chile, Hong Kong, India, Korea, Malaysia, Russia and Japan | Open |  |  |  | Not_Computer
Z-0492-2014 | 66536 | prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Aesculap Columbus Revision CoCr 15mm distal and posterior spacers, sizes F5-F7 | Aesculap Columbus Revision CoCr 15mm distal and posterior spacers, sizes F5F7 | 2 | December 13, 2013 | 2013 | Aesculap, Inc. | The Columbus Revision CoCr 15 mm distal and posterior implant spacers, sizes F5-F7, were not compatible. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Component Change Control | Domestic consignees were contacted by phone and a recall notification letters was also sent to these Users on/about September 26, 2013. |  | 24 | 24 | Distributed in FL and MD. | Open |  |  | 24 | Not_Computer
Z-0518-2014 | 66859 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Kimberly-Clark Single Shot Epidural Pain Management Trays containing the BD LOR Syringe (BD Code 405198)Product Usage:The device is designed to provide regional anesthesia. The device may contain syringes, needles and drugs | KimberlyClark Single Shot Epidural Pain Management Trays | 2 | December 16, 2013 | 2013 | Kimberly-Clark Corporation | Kimberly-Clark Single Shot Epidural Pain Management Trays contain the BD 7 mL Epilor plastic Luer-Lok LOR Syringes under recall by Becton Dickinson (BD) because the product may stall or stick when traveling within the barrel of the syringe. | DESIGN: Component Design/Selection | Kimberly Clark Corporation sent a Product Advisory Notice letter to affected consignees. The letter identified the affected product, problem and actions to be taken. For questions call 770-587-8393 or via email. |  | 389300 | 389,300 units | U.S. Nationwide Distribution | Open |  |  | 389300 | Not_Computer
Z-0516-2014 | 66894 | plasma, coagulation control | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | HemosIL High Abnormal Control 3 UNASSAYEDPart Number: 0020003320 | HemosIL High Abnormal Control 3 UNASSAYED | 2 | December 16, 2013 | 2013 | Instrumentation Laboratory Co. | HemosIL Low, Abnormal, Normal Controls UNASSAYED (no-value assigned) lacks 510(k) approval | PREMARKET APPROVAL: No Marketing Application | Instrumentation Laboratory issued Important Correction Notice via USPS to US customers dated November 2013 advising users in the US that unassayed controls are not acceptable for use in their unassayed (non-value assigned) format. IL will no longer sell or distribute controls without assayed values for the US market as of November 15, 2013, and are converting US customers to an assayed control format. Additional contact with customers via phone calls and a reminder email.An enclosed response tracking form provided to be returned. For technical questions, contact the IL Technical Support Center at 1-800-678-0710, Option # 6 (ACL Coagulation Products).For replacement product if value assignments are not available for your product lot, please contact your IL Customer Service Representative at 1-800-955-9525, Option # 6.For questions regarding this recall call 781-861-4467. |  | 84561 | 84561 kits | Nationwide Distribution | Open |  |  | 171384 | Not_Computer
Z-0520-2014 | 66918 | system,surgical,computer controlled instrument | NAY | Endoscope and accessories. | Gastroenterology/Urology | General & Plastic Surgery | 510(k) | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000.Manufactured by Intuitive Surgical, Sunnyvale, CA.EndoWrist Instruments, including scissors, scalpels, forceps, needle drivers and electrocautery are intended for endoscopic manipulation of tissue, including: grasping, cutting, blunt and sharp dissection, approximation, ligation, electrocautery and suturing. Mega Needle Driver and Large Needle Driver instruments are part of the EndoWrist instrument family and are intended for use in endoscopic manipulation and suturing. | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000 | 2 | December 16, 2013 | 2013 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers. | DESIGN: Software Design (Manufacturing Process) | Intuitive Surgical sent a Field Safety Notice dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure all appropriate personnel are fully informed of the Notification. Discontinue use of affected product, Use the attached instructions to identify and segregate affected product. Return affected product to Intuitive Surgical following the established Returned Materials Authorization (RMA) process. Complete the Acknowledgement Form and return to Intuitive Surgical. Customers with questions were instructed to contact Intuitive Surgical Customer Service. Customers were asked to retain a copy of this letter and the acknowledgement form for their files.For questions regarding this recall call 800-876-1310, option 3.Japan 0120-56-5635 or 03-5575-1362South Korea 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 +41 21 821 2020For questions regarding this recall call 408-523-2100.Urgent Medical Device Recall notification letters were sent to all affected sites on November 15, 2013. |  | 236 | 15, 236 UPDATED 12/19/13 to 11,235 | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Finland, Belgium, Netherlands, UK, Czech Republic, Sweden, Germany, Austria, Switzerland, Norway, Denmark, South Korea, France, Ireland, and Slovakia. | Open |  |  | 15236 | Not_Computer
Z-0513-2014 | 66626 | fluorescence polarization immunoassay, tobramycin | LFW | Tobramycin test system. | Toxicology | Toxicology | 510(k) | Thermo Fisher Scientific/Microgenics brand CEDIA Tobramycin II Assay, Catalog Number 100018, Product is manufactured and distributed by Microgenics Corporation, located at Fremont, CA., part of Thermo Fisher ScientificThe CEDIA Tobramycin II Assay is an in vitro diagnostic medical device intended for the quantitation of Tobramycin in human serum or plasma. | CEDIA Tobramycin II Assay | 2 | December 16, 2013 | 2013 | Microgenics Corp | Microgenics Corp., part of Thermo Fisher Scientific, is recalling CEDIA Tobramycin II Assay (Lot Number: 60169764) due to calibration stability problem relating to two complaints received. |  | Microgenics Corp., part of Thermo Fisher Scientific sent an Urgent Medical Device Correction letter dated Octobere 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard. Customers with questions were instructed to contact mgc@thermofisher.com. For question regarding this recall call 1-800-232-3342. |  | 370 | 370 units | Worldwide Distribution - USA including WI, UT, ND, MT, WA, CA, OK and Internationally to Australia, Canada, and Germany. | Terminated | February 03, 2014 | 49 | 370 | Not_Computer
Z-0517-2014 | 65051 | enzyme immunoassay, tracrolimus | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension(R) TACR Flex(R) reagent cartridge (DF107)The TACR method is an in vitro diagnostic test intended to quantitatively measure Tacrolimus in human whole blood on the Dimension(R) clinical chemistry system. Measurements of Tacrolimus are used as an aid in the management of Tacrolimus therapy in kidney and liver transplant patients | Dimension(R) TACR Flex(R) reagent cartridge (DF107) | 2 | December 16, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the TACR method may demonstrate reduced on-board stability which may result in imprecise and inaccurate QC and patient results. | DESIGN: Component Design/Selection | Siemens sent an Urgent Medical Device Correction letter dated April 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A recall notification letter was sent to consignees in April 2013. Customers using the affected product should load Flex reagent cartridges on board immediately prior to use and remove the Flex after 8 hours. Do not load more Flexes than the lab will use in 8 hours and do not pre-hydrate TACR Flexes. Customers were asked to complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at 302-631-8467. Customers were asked to forward the communication to anyone to whom may have distributed the product. Customers with questions were instructed to contact the Siemens Technical Solutions Center or their local Siemens technial support representative. For questions regarding this recall call 800-441-9250. |  | 13751 | 13,751 | Nationwide Distribution including AK, AL, AR, AZ, CA, DC, FL, GA, IA, IL, LA, MD, MI, MO, MS, NC, NH, NJ, NM, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, and WI. | Open |  |  | 13751 | Not_Computer
Z-0527-2014 | 66862 | aligner, bracket, orthodontic | ECQ | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Broncho Flexible Fiberscope, models: 11001BN1, 11002BD1, 11004BC1, and 11009BC1." Examination and treatment of the tracheobronchial tree Contraindications for Bronchoscopy " Uncontrolled angina, uncontrolled heart failure, or serious uncontrolled ventricular arrhythmias " Uncontrolled hypertension " Bleeding disorder or use of anticoagulants. | Broncho Flexible Fiberscope | 2 | December 17, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ Endoscopy America, Inc. is recalling the Flexible Broncho Fiberscope or a Flexible Intubation Fiberscope with a suction port because the material used in the suction channel portion was thermoplastic polyurethane (TPU), which is not listed in STERRAD systems claims for interior endoscope lumens. |  | The firm, KARL STORZ, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated November 8, 2013 to all their customers on November 13, 2013. The letter describes the product, problem and actions to be taken.The letter informs the customers that the endoscopes incorporated a lumen material that is only approved for use as an external surface material and not included in the STERRAD NX and 100NX FDA-cleared claims for use in flexible endoscope lumens. The customers were instructed to follow the recommended processing methods included in the letter.Customers with questions about the letter are instructed to contact the undersigned and the Technical Services at (800) 421-0837, ext 5350. |  | 141 | 141 units | Nationwide Distribution. | Open |  |  | 1657 | Not_Computer
Z-0490-2014 | 66735 | [6 empty] | Nephros Dual Stage Ultra Filter (DSU) BrochureDocument 60-3003Revisions: 0, 1, 2Product Usage: A disposable point of use water ultrafilter that delivers microbiologically pure water for washing and drinking. | Nephros Dual Stage Ultra Filter | 1 | December 17, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | MISBRANDING: Labeling False and Misleading | Nephros Inc., sent an Urgent Medical Device Recall letter dated October 29, 2013 to all affected customers via UPS 2nd Day Air. The letter identified the affected product, problem and actions to be taken. Customer are asked to provide a customer information list, immediately remove and discard all affected product labeling materials and any independent labeling materials they have created for the affected products. Customers are instructed to complete and sign the attached Response Form to acknowledge receipt, understanding and compliance with the instruction. For questions contact the Nephros representative at 201-343-5202 x 100.The letter was also posted at : http://www.nephros.com/download/news/DSU_Recall_Letters.pdf. |  | 3367 | 3,367 untis with affected labeling | USA Nationwide Distribution | Open |  |  | 3367 | Not_Computer
Z-0485-2014 | 66886 | system, test, blood glucose, over the counter | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Test Strips;For in vitro diagnostic testing.50 countUPC 6 99073 12050 2;Freestyle Blood Glucose Test Strips; for Institutional Use only50 count:UPC 6 99073 70792 5Manufactured by Abbott Diabetes Care Inc. Alameda, CA.The FreeStyle Blood Glucose test strip is intended for use in the quantitative measurement of glucose in capillary whole blood from the finger, upper arm and palm. It is intended for use by healthcare professionals and people with diabetes mellitus at home as an aid in monitoring the effectiveness of a diabetes control program. It is not intended for the diagnosis of or screening for diabetes mellitus, and it is not intended for use on neonates or arterial blood. | FreeStyle Blood Glucose Test Strips | 1 | December 17, 2013 | 2013 | Abbott Diabetes Care, Inc. | Certain lots of FreeStyle and FreeStyle Lite Blood Glucose Test Strips produce erroneously low blood glucose results when using FreeStyle Blood Glucose Meters, FreeStyle Flash Blood Glucose Meters and the FreeStyle blood glucose meter built into the OmniPod system. Erroneously low results that are not recognized may pose significant risks to your health. |  | Abbott Diabetes Care sent an Urgent Product Recall letter in November 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who were affected were instructed to take the following steps:Call Abbott Diabetes Care Customers Service immediately at 1-888-736-9869 for replacement of the affected product. While waiting for replacement strips to arrive, use an alternate method to measure their blood glucose.If only test strips available to customers were from affected lots, customers should not stop testing their blood glucose. Customers were instructed to dispose of the test strips in accordance with their local and state disposal regulations. Customers with questions were instructed to call Abbott Diabetes Care Customers Service at 1-888-736-9869.For questions regarding this recall call 510-239-2775.Abbott issued a press release at the request of CDRH on november 27, 2013. |  | 141219 | 141,219 cartons for human use - all lot numbers | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Australia, Austria, Belgium, Denmark, Egypt, Finland, France, Germany, Jamaica, Ireland, Israel, United Kingdom, Luxembourg, Netherlands, Norway, Sweden, Switzerland, UAE, and Canada.**Center Recommended Depth - Consumers/User** | Open |  |  | 141219 | Not_Computer
Z-0525-2014 | 66496 | test, epithelial ovarian tumor-associated antigen (ca125) | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROS CA 125 II CalibratorsFor use in the calibration of the VITROS ECi/ECiQ Immunodiagnostic Systems, VITROS 3600 Immunodiagnostic System and VITROS 5600 Integrated System for the quantitative measurement of OC 125 defined antigen in human serum and plasma (EDTA or heparin). | VITROS Immunodiagnostic Products CA 125 II Calibrators | 2 | December 17, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products CA 125 II Reagent Pack and Calibrators due to a positive shift in patient results. |  | The firm, Ortho Clinical Diagnostics, sent a "IMPORTANT PRODUCT CORRECTION NOTIFICATION" letter dated October 9, 2013 to their clients/customers informing them of a shift in results that may be observed when performing patient sample correlation tests between VITROS CA 125 II Calibrators and Reagent Packs. The letter described the product, problem and actions to be taken. The customers were informed that it is acceptable to continue using any remaining inventory until its expiration date providing that quality control results are within acceptable limits (Note: Lots 1220 and 1230 are expired). The customers were also instructed to consider the need to review patient results using the affected lot(s). If you suspect that a previously reported result may have been affected, provide this information to your Laboratory Medical Director and the requesting physician or health care provider so that appropriate actions may be taken for the patients; Complete and return the attached Confirmation of Receipt form no later than October 18, 2013 via FAX TO: Ortho Clinical Diagnostics at 1-888-557-3759 or 1-585-453-4110, and Forward the information in this notification if you have distributed this product outside of your facility.If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 601 | 601 units | Worldwide Distribution: US (nationwide) including states of: AK, CA, CT, CO, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV, WY, and Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy & Spain. | Open |  |  | 601 | Not_Computer
Z-0530-2014 | 65477 | kit, identification, yeast | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | remel RapID Yeast Plus System, 20 tests, Ref: R8311007, IVD, containing a 15-ml bottle of RapID Yeast Plus Reagent A and a 10-ml bottle of RapID Yeast Plus Reagent B. Product Usage: For the biochemical identification of medically important yeast and yeast-like organisms. | remel RapID Yeast Plus System | 2 | December 17, 2013 | 2013 | Remel Inc | The two reagents supplied in the RapID Yeast Plus System kit (Reagent A and Reagent B) are mislabeled. The labels were reversed. | MISBRANDING: Labeling False and Misleading | Thermo Fisher Scientific sent a Medical Device Recall letter dated April 3, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm instructed customer to review their inventory for affected lot and discard the product. A Medical Device Recall Return Response was enclosed for the customer to indicate they understood the instructions and to report the amount of product they discarded. For questions call Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 150 | 150 kits | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CT, FL, ID, IN, KY, MA, MI, MS, NJ, OH, PA, TX, WI and the countries of Canada, UK, and Japan. | Open |  |  | 150 | Not_Computer
Z-0533-2014 | 66885 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Integra Disposable, Convenience Kit Single Shot Epidural Tray Reorder Number 340269 Rx OnlyProduct Usage:The Integra kits which contain a BD LOR syringe are intended to provide licensed medical professionals with a kit containing drugs and other medical devices that are used to support the administration of general, local or spinal anesthetics. The BD LOR syringe is intended for use, in conjunction with an epidural needle, to verify the needle tip placement is in the epidural space by use of the Loss of Resistance Technique as detailed in medical textbooks and medical journal articles. The LOR Syringe is not intended for injection or aspiration. | Integra Kits for Epidural Anesthesia Administration | 2 | December 17, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corporation is recalling certain lots of Integra Kits for Epidural Anesthesia Administration as a result of a Product Advisory Notice from BD Medical for the LOR syringes indicating the plunger of the syringe may stall or stick when travelling within the syringe. | DESIGN: Component Design/Selection | BD Medical sent a Product Advisory Notice letter dated August 23, 2013 to Iaffected consignees by via traceable courier. The letter identified the affected product, problem and actions to be taken. Distributor have been asked to forward copies of the notification to their customers. For questions call BD Customer Service at 1-888-237-2762. |  | 359250 | 359,250 | USA Nationwide Distribution and one foreign consignee in Mursalat. | Terminated | January 29, 2014 | 43 | 359250 | Not_Computer
Z-0528-2014 | 65734 | susceptibility test discs, antimicrobial | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, CTO425B Ciprofloxacin CIP 5 mcg, 50 discs per cartridge, 5 cartridges per box. The firm name on the box label is Oxoid Ltd., Basingstoke, UK.Used in the semi-quantitative agar diffusion test method for in vitro susceptibility testing to assist in the determination of a suitable antibiotic agent for treatment purposes. | Oxoid Antimicrobial Susceptibility Test Discs, CTO425B Ciprofloxacin CIP 5 mcg. | 3 | December 17, 2013 | 2013 | Remel Inc | Boxes labeled as CIP5 may contain cartridges of AMP5. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Mix-Up of Materials/Components | Thermo Fisher sent a Field Safety Notice dated November 13, 2012, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to check stocks of the affected product if necessary and contact Customer Service or their local Oxoid supplier who will issue a replacement. The notice should be passed on to all who need to be aware within their organization. Customers were also instructed to discard any product labeled as AMP5. For questions, customers were instructed to call +44 (0)1256 841144.discard any product labeled AMP5.For questions regarding this recall call 913-895-4185. |  | 115 | 60/5/50-disc cartridge boxes | Nationwide Distribution including AL, CA, CO, FL, IA, KS, LA, MA, MS, NE, NH, NJ, NY, OK, OR, RI, TN, TX, VA, WA WI, and Puerto Rico. | Open |  |  | 60 | Not_Computer
Z-0523-2014 | 66964 | forceps | HTD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | outer box: Ref 906768, CurvTek Eye Needle 22 MM X-large Stainless Steel, Single Use Gamma Sterilized, Rx inner incorrect label: CurvTek 7 MM Medium, Single Pack Quantity -3, REF 906760 To pass the CurvTek needle, place the jaws of the instrument utilized to grasp the needle on the sides of the needle and orient and needle parallel to the long axis of the instrument. Position the needle to follow the path of the tunnel created by the cartridges. This aligns the needle with the radius of the transosseous tunnel. Advance the needle while rotating the instrument away from the entry point. Gently feed the needle through the length of the tunnel. | CurvTek Eye Needle | 2 | December 17, 2013 | 2013 | Biomet, Inc. | Inner label of PN: 906768 Lot: 055450 CurvTek Eye Needle 22mm X-Large states incorrectly PN: 906760 7mm Medium. | PRODUCTION CONTROLS: Labeling Mix-Ups | Biomet sent an Urgent Medical Device Recall Notice dated November 26, 2013, to all affected customers. The notice identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm by FAX 574-372-1683. Customers with questions were instructed to call 574-372-1570.For questions regarding this recall call 574-372-1570. |  | 21 | 21 | Nationwide Distribution including MI, IA, IL, IN, and TN. | Open |  |  | 21 | Not_Computer
Z-0522-2014 | 65915 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | KayserBetten model IDA bedsProduct Usage: BED, AC-POWERED ADJUSTABLE HOSPITAL | KayserBetten | 2 | December 17, 2013 | 2013 | mobility unlimited inc | Some joints on the doors of a limited number of IDA beds may loosen or separate over time which can cause the doors to drop outward and cause injury. | DESIGN: Device Design | Mobility Unlimited sent a notification letters dated April 12, 2013 to affected customers The letter identified the affected product, problem and actions to be taken. The letter informed customers that a Retrofit Kits along with illustrated instructions will be sent. For questions call 1-800-574-7880. |  | 33 | 33 beds | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, FL, GA, IL, MN, MO, NY, PA, TN, TX, VA, and Canada. | Open |  |  | 33 | Not_Computer
Z-0524-2014 | 66992 | system, image processing, radiological | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrakÂŽ 3500 Plus, ENTrakTM Plus, InstaTrakÂŽ 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), | InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM | 2 | December 17, 2013 | 2013 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. |  | GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.Customers with questions were instructed to contact the service team at 800-874-7378.For questions regarding this recall call 801-536-4952. |  | 654 | 654 | Nationwide Distribution | Open |  |  | 654 | Not_Computer
Z-0487-2014 | 66736 | [6 empty] | Nephros SafeSpoutManufacturer: Nephros Inc. 41 Grand Avenue, River Edge, NJ 07661Assembled at : MEDICA 41036 Medolla, ItalyFilters are packaged in a Tyvek pouch and then placed inside a cardboard box.15 filters per boxThe Nephros SafeSpout and SafeShower are intended to filter water for washing and drinking. They are intended to be installed at the end of a standard sink faucet or shower hose | Nephros SafeSpout | 1 | December 17, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | DESIGN: Device Design | Nephros sent an Urgent Medical Device Recall letter dated October 28, 2013, with a return response form to all affected customers via UPS 2nd Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease shipment of the affected product immediately. Obtain an RGA number to return the new and used filters for credit. In addition, customers were instructed to immediately remove and discard all copies of the labeling materials that pertain to the affected product. Remove all electronic copies of these materials from any publicly accessible locations (e.g. website). Customers with questions were instructed to contact their Nephros representative or call 201-343-5202, ext 100.For questions regarding this recall call 201-343-5202.The letter was also posted at:http://www.nephros.com/download/news/SafeSpout_SafeShower_Recall_Letters.pdf. |  | 3742 | 3742 units | Nationwide Distribution | Open |  |  | 8946 | Not_Computer
Z-0535-2014 | 66944 | material, impression | ELW | Impression material. | Dental | Dental | 510(k) | 3M ESPE Imprint II Regular Body , Catalog Number 9379, and Paradigm Regular Body VPS, Catalog Number 5315, Impression Material that are intended to reproduce the structure of a patients teeth.Impression Materials are composed of hydrophilic vinyl polysiloxane impression materials, designed for making precise crown and bridge, inlay, onlay, veneer preparations, removable denture and partial denture impressions. | Imprint II and Paradigm impression material | 2 | December 18, 2013 | 2013 | 3M Company - Health Care Business | Specific lots of Imprint II and Paradigm impression material refill packs are being recalled because the moisture displacement feature in both is greatly reduced which may impact the quality of the dental impression. | TRAINING: Employee Error | 3M ESPE sent an Voluntary Recall Notification Letter dated November 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to:1) Review your inventory and identify any products from these lot numbers2) Quarantine products from these lots and do not further distribute3) If you have shipped any products from these lots to your customers, inform them of this voluntary recall using the attached Dear Doctor Letter and Inventory Checklist4) To return and replace product from these lots, contact 3M ESPE Dental ProductsCustomer Service at 1-800-237-1650If you have any questions, please contact the above number for assistance. Once again, we sincerely apologize for any inconvenience this may cause you. |  | 1977 | 1977 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, France, Germany, Hong Kong, Italy, Japan, Korea, Mexico, United Arab and Emirates. | Open |  |  | 1977 | Not_Computer
Z-0537-2014 | 65476 | discs, strips and reagents, microorganism differentiation | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Streptex* kit, Ref R30950501 ZL50, Rapid latex test for Streptococcal grouping, IVD. The kit contains reagent groups A, B, C, D, F, G, a control, and an extraction enzyme. The firm name on the label is Remel Europe Ltd., Dartford, England.Streptex is a rapid latex test system for use in the qualitative detection and identification of the Lancefield group of streptococci. Reagents are provided for groups A, B, C, D, F and G covering the majority of clinical isolates, group E streptococci are rarely isolated. | Streptex kit, Ref R30950501 ZL50, Rapid latex test for Streptococcal grouping | 2 | December 18, 2013 | 2013 | Remel Inc | The product may give weak or false negative test results. | PRODUCTION CONTROLS: Process Control | ThermoFisher sent a Medical Device Recall letter dated April 26, 2013, or May 2, 2013, to all affected customers via first class mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be notified of the potential for a false negative test result, review their inventory for affected product lots and discard the affected product. Customers were also instructed to return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 768 | 768 kits | Nationwide Distribution | Open |  |  | 877 | Not_Computer
Z-0546-2014 | 66589 | anesthesia conduction kit | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Continuous Peripheral Nerve Block Catheter Kit and SetContinuous Peripheral Nerve Block Catheter Kit and Set permits placement of catheters next to nerves and nerve plexuses for continuous nerve block anesthesia or analgesia techniques for periods not exceeding 72 hours. | Continuous Peropheral Nerve Block Catheter Kit and Set | 2 | December 19, 2013 | 2013 | Arrow International Inc | A labeling inconsistency was discovered in which the lidstock states that the kit contains a stimulator cable with 1.5mm socket, however, the kit contains a stimulator cable with 1.5mm pin. | PRODUCTION CONTROLS: Labeling Mix-Ups | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated October 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of listed affected products/lots you have in stock; return affected products, customer service representative will contact you with RGA# and instructions for the return of the product, and complete enclosed Recall Acknowledgement form and fax to 1-800-343-2935, Attn: Customer Service, even if you have not affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 121 | 121 | Worldwide distribution: US (nationwide) including states of: MA, LA, CA, GA, KS, TX, PA, NY, MO, WI, NJ, MS, CT, ID, AL, CO, and PR; and Singapore. | Open |  |  | 121 | Not_Computer
Z-0541-2014 | 66363 | urinary homocystine (nonquantitative) test system | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista Homocysteine Flex Reagent CartridgeThe HCYS method is an in vitro diagnostic test for the quantitative measurement of total homocysteine in human serum, heparinized plasma and ETDA plasma on the Dimension Vista System. Measurements of homocysteine aid in the diagnosis and treatment of patients suspected of having hyperhomocysteinemia and homocystinuria. | Dimension Vista HCYS | 2 | December 19, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista HCYS lot 13064MA shows erratic recovery of QC and patient samples. |  | Siemens sent a Urgent Medical Device Recall letter dated September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions to be taken by Customer:Please discontinue use and discard any remaining inventory of Dimension VistaÂŽ HCYS FlexÂŽ reagent cartridge (Catalog# K7044) lot 13064MA.Please complete the attached form and fax it to (302) 631-8467 to facilitate your no charge replacement needs.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center- Technical Solutions or your local Siemens technical support representative. For further questions please call (800) 441-9250. |  | 302 | 302 | Worldwide Distribution - US Distribution including the states of OH, MD, CO, NJ, IA, NC, VA, WI and DC., and the countries of Austria, Denmark, France, Italy, Netherlands and Norway. | Open |  |  | 302 | Not_Computer
Z-0545-2014 | 66983 | unit, operative dental | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | SEAC Advanced Dental System, Model A591561The SEAC Advanced Dental System is a self-contained dental delivery system with a built-in air compressor,vacuum system and water supply. It is designed operate dental handpiece attachments and suction instruments. The SEAC Advanced Dental System can also be upgraded to include a Micadent air abrasion system mounted inside the unit with the handpiece mounted in a holder along side the other handpiece connections. The Micadent is used to perform intraÂŽ oral air abrasion procedures including etching and preparing surfaces for composite restoration. The SEAC Advanced Dental System is substantially similar in design and functions as the SEAC Mobile Dental System manufactured by ASI Medical, Inc. The Micadent is manufactured by Medidents and will be supplied to ASIas a completed product. The Micadent is substantially the same as it is provided and is attached to an accessory air line inside the SEAC unit for operation. It is intended for use to provide general dental restorative care and hygiene procedures in both traditional dental office settings and/or mobile applications by or under the supervision of a licensed dentist or a hygienist if permitted by applicable law. | SEAC Advanced Dental System | 2 | December 19, 2013 | 2013 | ASI Medical, Inc. | ASI Medical is initiating a field correction for the SEAC Advanced Dental System because the miniature plastic air tank can potentially rupture. | DESIGN: Device Design | ASI Medical sent an Field Correction Action and Safety Notice letter dated November 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.REQUIRED FIELD CORRECTIONOur field correction will consist of sending you a warning label to be applied to the back service access panel of the cart and an ASME safety pop off valve that will open and exhaust if pressure should reach over 100 psi. This safety pop off valve is designed be easily installed in line between the air compressor and air tank in your system. We are requesting your cooperation in making this field correction. |  | 333 | 333 | Worldwide Distribution - USA (nationwide) including the countries of Austria, Aruba, Saudi Arabia, Canada, Germany, Spain and West Indies. | Open |  |  | 333 | Not_Computer
Z-0544-2014 | 66698 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw; Model: 01-80115-55The PedFuse Reset pedicle screw implants intended to be used to help provide immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic, lumbar, and/or sacral spine. | PedFuse Reset Screw | 2 | December 19, 2013 | 2013 | SpineFrontier, Inc. | During routine servicing, it was discovered that one 8mmx55mm Reset screw was assembled with an 8mm tulip head and a 9mm screw. The implant should haveconsisted of an 8mm tulip head and an 8mm screw. A recall was initiated to remove all remaining screws of the same lot from the field to investigate the nonconformity. | DESIGN: Software Design (Manufacturing Process) | SpineFrontier sent an Voluntary Recall Notification dated November 5,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.All affected implants are to be returned to SpineFrontier for inspection. SpineFrontier is requiring confirmation that you have received this notice and that you will initiate the immediate return of the implants listed above.IMMEDIATE RESPONSE REQUIRED.Please feel free to contact us at 978-232-3990 Monday through Friday8:30AM to 5:30PM EST with any questions. |  | 35 | 35 Devices | US Distribution to VA and TX. | Open |  |  | 35 | Not_Computer
Z-0543-2014 | 66844 | lens, contact, (disposable) | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) Toric 1 day (ocufilcon D), 30 daily disposable contact lenses for astigmatism, STERILE, Rx only --- CooperVision Hamble, SO31 4RF, UK - Scottsville, NY 14546 USA - Juana Diaz, PR 00795 USAdaily use disposable contact lenses for astigmatism | Proclear | 2 | December 19, 2013 | 2013 | CooperVision Inc. | Lots were labeled with incorrect expiration dates on secondary package (carton). The primary package (blister) is marked with the correct expiration date. | EXPIRATION DATING: Incorrect or No Expiration Date | CooperVision sent an Urgent Medical Device Recall Letters dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Our records indicate you received some of the affected product. Enclosed you will find a list of affected carton lot numbers, with the correct expiration date, that were shipped to your account." Please immediately examine your inventory," Stop any further distribution of carton(s) subject to the recall," Quarantine carton(s) with incorrect expiration date, and" CooperVision will make arrangements to remove affected, quarantine product.As soon as possible, please complete the attached response form and fax it to 1.866.484.1048 or scan and email to info@coopervision.com. If you have product to return please contact customer service for instructions and to schedule pick-up and removal of affected, quarantined product. Your account will be credited once the product is returned and processed.We appreciate your assistance and apologize for any inconvenience you or your patients experienced by this removal. If you have any further questions, please feel free to contact us at 800-341-2020. |  | 45 | 45 cartons | Worldwide Distribution - USA (nationwide) and the countries of Aruba, Canada, Colombia and Trinidad. | Open |  |  | 198 | Not_Computer
Z-0550-2014 | 66842 | system, test, vitamin d | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FastPack Free T4 Immunoassay FastPack Vitamin D Immunoassay Chemilunescence assay for the determination of Vitamin DThe FastPack Vitamin D Immunoassay is intended for the quantitative determination of FastPack Total PSA Immunoassay Chemilunescence assay for the determination of Total PSA FastPack Total PSA Immunoassay is a paramagnetic particle immunoassay for the in vitro quantitative determination of prostate-specific antigen (PSA) in human serum and plasma as an aid in the management of patients with prostate cancer. | FastPack Vitamin D Immunoassay | 2 | December 19, 2013 | 2013 | Qualigen Inc | Qualigen initiated this recall because the FastPack kit may cause lower than expected results due to a manufacturing defect. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Process Change Control | Qualigen, Inc., notified customers of the recall by calling customers via telephone (with a telephone script) on 11/08/2013. The telephone (script) included the reason for recall, product description with codes, instructions and contact information. The customers were instructed to examine your inventory and discontinue use of the kits immediately. Qualigen will send youreplacement kits and provide you with a shipping label to return all of the defective kits including any that were partially used.Feel free to contact System Support at 760.579.6900 if you have anyquestions or require assistance with returning your kits. |  | 6 | 6 | US Distribution to states of: AL, CA, KS, KY, IN, MA, NC, NY, NV, OK, SC, TN, TX, UT, VA, and WA. | Open |  |  | 223 | Not_Computer
Z-0557-2014 | 66900 | test, hepatitis b (b core, be antigen, be antibody, b core igm) | LOM |  |  | Microbiology |  | VITROS Anti-HBs Calibrators & VITROS Anti-HBs Reagent Packs | VITROS Immunodiagnostic Products AntiHBs Calibrators & AntiHBs Reagent Packs | 2 | December 20, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products Anti-HBs Calibrators and Anti-HBs Quantitative Reagent Packs due to false positive results. |  | On October 25, 2013, Ortho Clinical Diagnostics issued product correction notifications to their US customers via US Postal Service Priority Mail and email notices to their foreign affiliates informing them of the recall. Ortho Clinical Diagnostics advised their customers to review previously reported results between |  | 1920 | 1,920 kits | Nationwide distribution - all states. | Open |  |  | 1920 | Not_Computer
Z-0554-2014 | 66873 | orthosis, spinal pedicle fixation | MNI | Pedicle screw spinal system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw Inserters; PedFuse Screw Inserter 05-7 REspond Part# 11-80111The PedFuse Screw Inserter 05-7 REset, Screw Inserter 05-7 REspond, and Screw Inserter 08-9 REset are used as a screw capture, screw driver, and screw insertion tool in the PedFuse system. | PedFuse Reset Screw Inserters | 2 | December 20, 2013 | 2013 | SpineFrontier, Inc. | Screw inserters may not mate properly with hex interface of the screws. | DESIGN: Device Design | The firm, SpineFrontier, Inc., sent an "Urgent PedFuse Screw Inserter Voluntary Recall Notification" letter on November 20, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to initiate the immediate return of the recalled products and to complete a recall ACKNOWLEDGEMENT FORM via fax or email to: SpineFrontier, Inc., 500 Cummings Center, Suite 3500, Beverly MA 01915, Attn: Customer Service; Fax: 978-232-3991; email: customerservice@spinefrontier.com.Please feel free to contact at 978-232-3990 Monday thru Friday 8:30AM to 5:30PM Eastern with any questions. |  | 25 | 25 | US Distribution to states of:Texas, New York, Illinois, Ohio, Virginia, Washington D.C., and Florida.. | Open |  |  | 42 | Not_Computer
Z-0531-2014 | 66930 | aid, surgical, viscoelastic | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Viscocel and Viscocel Plus.These products are intraocular fluid devices consisting of non-gaseous fluid intended to be introduced into the eye to aid performance of surgery, such as to maintain anterior chamber depth, preserve tissue integrity, protect tissue from surgical trauma, or function as a tamponade during retinal reattachment. | Viscocel and Viscocel Plus | 1 | December 20, 2013 | 2013 | C.L.R. Medicals International, Inc. | CLR Medicals International Inc., initiated a nationwide recall of Viscocel and Viscocel Plus, in response to FDA inspection findings of deficiencies in the Quality System Regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human | PREMARKET APPROVAL: No Marketing Application | CLR Medicals initiated a nationwide recall on November 25, 2013, of Viscocel and Viscocel Plus, by issuing a press release,in response to FDA inspectional findings of deficiencies in the Quality System regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human and animal use, have been determined to be a public health concern due to distribution in the US without FDA approval, which evaluates the safety and efficacy of the product, and violations of the QSR regulations related to sterility, and biocompatibility that may result in the release of contaminated products. CLR Medicals International, Inc. sent a Medical Device Recall letter dated December 6, 2013, to all affected customers. The letter informed customers of the reason for recall, product description with manufacturing date range, possible health hazard, instructions on what to do with recalled product, and contact information for the firm and to report any adverse events to FDA MedWatch. The recall notification letter was accompanied by a response form. Customers with questions were instructed to call 909-398-1880.For questions regarding this recall call 909-398-1880. |  | 249592 | 249,592 units | Worldwide Distribution - USA (nationwide) and Internationally to Egypt, Mexico, and Canada. | Open |  |  | 249592 | Not_Computer
Z-0558-2014 | 66971 | wire, guide, catheter | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HT Connect Peripheral Guide WirePart Number Description1012587 018 HT CONNECT 145 CM1012588 018 HT CONNECT 195 CM1012589 018 HT CONNECT 300 CM1012590 018 HT CONNECT FLEX 145 CM1012591 018 HT CONNECT FLEX 195 CM1012592 018 HT CONNECT FLEX 300 CM1012593 018 HT CONNECT 250T 145 CM1012594 018 HT CONNECT 250T 195 CM1012595 018 HT CONNECT 250T 300 CM | HT Connect Peripheral Guide Wire | 2 | December 20, 2013 | 2013 | Abbott Vascular, Inc. | Abbott Vascular has initiated a recall of the HT Connect Peripheral Guide Wires due to a small number of devices exhibiting partial delamination of the PTFE coating. |  | Abbott Vascular sent an Urgent Device Recall letter dated November A recall communication letter dated November 25, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. The letter informs the customers of the 9 part numbers and 9 description of the products. Abbott Vascular is instructing the customers that the returned products will be replaced with similar products pending availability. Abbott Vascular has attached a customer response form for customers along with the customer notification letter.For questions regarding this recall call 951-914-2400. |  | 54895 | 54,895 units total (5,356 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, French Guyana, Macedonia, South Korea, Austria, Germany, Malaysia, Spain, Bahrain, Greece, Martinique, Sweden, Belgium, Hong Kong, Netherlands, Switzerland, Brazil, Indonesia, New Zealand, Taiwan, Bulgaria, Iran, Norway, Thailand, Canada, Iraq, Poland, Turkey, China, Ireland, Portugal,United Arab Emirates, Czech Republic, Italy, Saudi Arabia, United Kingdom, Denmark, Latvia, Singapore, Estonia, Libya, Slovakia, Finland, Lithuania, Slovenia, France, Luxembourg, South Africa. | Open |  |  | 54895 | Not_Computer
Z-0551-2014 | 66949 | shunt, central nervous system and components | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Cranial Access Kit, Model Number HITHSP10, Rx Only.Integra Cranial Access Kits (CRAK) allows for burr hole access to the subarachnoid space or the lateral ventricles of the brain. The kit is intended to be used with an external drainage and monitoring system in selected patients to reduce intracranial pressure (ICP), to monitor CSF, to provide temporary drainage of Cerebrospinal Fluid (CSF), and to monitor ICP. These kits contain all components and accessories required for burr hole access into the cranium. | Integra Cranial Access Kit | 2 | December 20, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is voluntarily recalling anyunexpired lots of HITHSP10 Cranial Access Kits due to identified labeling discrepancies. kit contains DuraPrep and Monocryl sutures whose labeling does not align with the intended use and labeling of the Cranial Access Kit. | PRODUCTION CONTROLS: Labeling Mix-Ups | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 26, 2013, to all affected consignees. The letter identified the product, the problem, the action to be taken by the consignee and included a Recall Acknowledgement and Return form. Consignees were instructed to review their inventory and determine if they have any affected product, stop using the affected product immediately and remove from service. Consignees were instructed to complete the attached form even if they do not have the affected product in their inventory and return by email or fax as indicated on the form. When the form is received Customer Service will send an RMA number and directions to return the product. Customers with questions were instructed to contact Customer Service at 1-855-532-1723 or their Integra sales representative at 909-201-4099.For questions regarding this recall call 609-936-6822. |  | 731 | 731 | Nationwide Distribution | Open |  |  | 731 | Not_Computer
Z-0559-2014 | 66535 | introducer, catheter | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | B. Braun Introcan Safety IV Catheter PUR 24G, 0.7X19mm-US, catalog number 4251601-02. Product Usage: The B. Braun Introcan Safety IV Catheter is a passive anti-stick device to provide venous or arterial access for the infusion of fluids, drugs, and/or blood components, or to facilitate the placement of vascular access devices such as guidewires, indwelling central venous catheters, peripherally inserted central catheters, and midline catheters into vascular system. | B. Braun Introcan Safety IV Catheter | 2 | December 21, 2013 | 2013 | B. Braun Medical, Inc. | One lot of Introcan Safety IV Catheters may not contain the anti-needlestick safety clip which is used for the prevention of needlestick injuries. | MISBRANDING: Labeling False and Misleading | B. Braun sent an Urgent Medical Device Recall Notification letter dated September 27, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to utilize the attached "Product Removal Acknowledgement Form" to record the number of individual units in their inventory and fax the completed form to 610-848-1197. To ensure proper credit for affected product, customers were instructed to return the product to B. Braun Medical Inc. For questions contact Customer Support Department at 800-227-2862. |  | 55200 | 55,200 units | USA Nationwide Distribution including Puerto Rico. | Open |  |  | 55200 | Not_Computer
Z-0560-2014 | 66934 | forceps | HTD | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | Edwards Lifesciences Ergonic Double Acting DeBakey Forceps with Narrow Straight Tip. Model Number ERGF01330. The Edwards Surgical Instruments are designed to perform specialized functions in specific cardiac surgical or endoscopic procedures. | Edwards Lifesciences Ergonic Double Acting DeBakey Forceps with Narrow Straight Tip | 3 | December 21, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling one lot of Ergonic instrument due to mislabeling. Ergonic scissors standard curved labeled as a Double Acting DeBakey forcep. Edwards control confirmed that the laser etching on the device was incorrect. | PRODUCTION CONTROLS: Labeling Mix-Ups | Edwards sent an " Urgent Field Safety Notice" to all affected customers dated November 20, 2013. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken by user:Our records show that you received one of these instruments. Our sales representative will be on-site to remove this instrument and provide a replacement. Please assist our sales representative in reviewing your entire inventory for the instrument listed in the recall letter. The form has been pre-populated with the inventory from our records. If you have additional inventory of this affected lot within your control, please add the quantity to the section on the form provided. Please remove the instrument from use, quarantine the device at your site and support the Edwards Sales Representative to return this product to Edwards. An acknowledgment form is included to assist you in the assessment of your inventory.While our sales representative is on-site, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 if you have questions about the return and crediting process.The Customer Service organization can answer questions about tracking of the replacement instrument provided by the Edwards Sales Representative. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST . |  | 3 | 3 | Distributed to 3 customers in FL only. | Terminated | February 10, 2014 | 51 | 3 | Not_Computer
Z-0569-2014 | 66713 | acid, folic, radioimmunoassay | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITE/IMMULITE 1000 Systems FOL Folic Acid, REF/Catalog Number LKFO1 (100 tests), Siemens Material Number (SMN) 10380902, IVD; for in vitro diagnostic use with the IMMULITE and IMMULITE 1000 Analyzers --- Siemens Healthcare Diagnostics Products Ltd. An IVD for the quantitative measurement of folic acid in serum, heparinized plasma or ascorbic acid-treated whole blood, as an aid in clinical diagnosis and treatment of anemia. | IMMULITE /IMMULITE 1000 Folic Acid | 2 | December 24, 2013 | 2013 | Siemens Healthcare Diagnostics | potential for an accelerated decline in counts per second (CPS) for three IMMULITE/IMMULITE 1000 Folic Acid kit lots. This may lead to a failed adjustment before the labeled expiration date is reached. |  | Siemens sent an Urgent Medical Device Correction notices dated November 7 2013, via FedEx to the US customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers are instructed that If the adjustment slope is less than or equal to 1.8 and quality control results are within the established ranges, the kit may be used to generate patient results. If the adjustment slope is greater than 1.8, customers are instructed to discontinue use and discard the kit lots. Ex-US distribution is determined at the country level. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Field service personnel were sent a support document describing the issue and instructing them how to deal with customer questions.For further questions please call CUSTOMER CARE CENTER NUMBER: (302) 631-7597. |  | 442 | Domestic: 442 kits; Foreign: 680 kits | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of Argentina, Bosnia-Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Denmark, Dominican Republic, Estonia, France, Germany, Great Britain, Greece, Guatemala, Hungary, India, Italy, Latvia, Mexico, Netherlands, Panama, Paraguay, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Tunisia, United Arab Emirates, Uzbekistan and Venezuela. | Open |  |  | 1112 | Not_Computer
Z-0561-2014 | 65131 | radioimmunoassay (two-site solid phase), ferritin | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Dimension(R) Chemistry Wash (RD701), Siemens Material Number (SMN) 10445052Product Usage: Used in the Dimension(R) system with the heterogeneous module | Siemens Dimension(R) Chemistry Wash (RD701) | 2 | December 24, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics received complaints of Chemistry Wash causing QC and patient sample result shifts on CTN1, LTN1, FT4, TSH, PBNP and LPBN when using Chemistry Wash lots RD23031, RD23111, RD23241, RD23311, RD23391, and RD23461. These lots shipped between 12/12/2012 and 3/22/2013. Internal investigation has confirmed shifts can be encountered in two scenarios: 1 - When an affected l | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated May 2013 to all affected customers. The letter identified the product and lots, the reason for the field action, the risk to health and the action(s) to be taken. Customers were instructed to discontinue use and discard inventory of the affected product. Siemens will replace any unused inventory of the affected lots at no charge. For questions contact your Siemens Technical Center or your local Siemens technical support representative. |  | 17719 | 17,719 | US Nationwide and OUS. | Open |  |  | 17719 | Not_Computer
Z-0579-2014 | 67003 | fastener, fixation, nondegradable, soft tissue | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics ACL Reconstruction System is a pediatric and small stature adult ACL reconstruction system containing Sleeve Installation Tool Set , Part 01-1008-0040/0047, 6mm; 6.5mm, 7mm, 7.5mm, 8mm, 8.5mm, 9mm, and 10mm.Used with ACL reconstruction system associated surgery. | ACL Reconstruction System Installation tools | 2 | December 26, 2013 | 2013 | OrthoPediatrics Corp | Part # 01-1008-0040/0047, ACL Sleeve Installation/ Insertion Tool, shipped November 4, 2013 - November 6, 2013 may be etched with the wrong size. | DESIGN: Labeling Design | November 7, 2013 all consignees were notified of the Urgent Device recall via telephone calls and E-mail messages describing the product and reason for the recall,. The notice included directions to identify and quarantine the recalled devices returning product to OrthoPediatrics with completed return form to OrthoPediatrics via mail, email or courier service, or byphotographing with a smart phone and emailing to logistics@orthopediatrics.com or by answering therecall specific questionsin an email to mfox@orthopediatrics.com. |  | 40 | 40 | Nationwide Distribution including the states of CA, KY, OH, RI, and TN. | Open |  |  | 40 | Not_Computer
Z-0577-2014 | 66935 | [6 empty] | ProPlan CMF Guide(s) Model(s)REF SD900.008 Description: Single Use Only : Patient-Specific Guide, Orthognathic, Intermediate Single Use Only Rx only.Used as surgical tools to transfer a pre-operative plan to surgery. | ProPlan CMF | 2 | December 26, 2013 | 2013 | Materialise USA LLC | Patient Specific Guide contained the incorrect case report. The case report contains patient specific surgical plan information. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Customer/ ordering physician was made aware of the issue11/08/2013. Revised accurate case report was immediately provided. |  | 1 | 1 | Distributed in Canada. | Terminated | January 07, 2014 | 12 | 1 | Not_Computer
Z-0573-2014 | 66914 | appliance, fixation, spinal intervertebral body | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Indus Invue Screws: IM71058-XX: 04.2mm, SelfTapping, Tapered.Used to secure the Invue Anterior Cervical Plate in position. | Invue Screws | 2 | December 26, 2013 | 2013 | SpineFrontier, Inc. | Mismarked and unmarked screws | PRODUCTION CONTROLS: Packaging Process Control | On 11/20/2013, consignees were notified of the email via letter. Consignees were instructed to return the screws listed in the recall notification letter and return the Recall Acknowledgement form to SpineFrontier via fax or email. |  | 993 | 993 | Nationwide Distribution- including DC and the states of OH, TX, KS, MA, IN, CO, NY, AL, and MI. | Open |  |  | 1915 | Not_Computer
Z-0572-2014 | 66916 | enzyme linked immunoabsorption assay, treponema pallidum | LIP | Treponema pallidumtreponemal test reagents. | Microbiology | Microbiology | 510(k) | Captia Syphilis-G Test KitProduct Usage: This is an enzyme immunoassay for the qualitative detection of IgG antibodies to T. pallidum in serum specimens, to be used in conjunction with non-treponemal testing to provide serological evidence of infection with T. pallidum ( the agent of syphilis). It is also intended for testing of serum or plasma specimens to screen blood and/or plasma donors to exclude a history of syphilis. | Captia Syphilis IgG | 2 | December 26, 2013 | 2013 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | Trinity Biotech recalled Trinity Biotech Captia Syphilis IgG kits due to false negative results. | DESIGN: Process Design | Trinity Biotech issued an Urgent Product Information Request letter to their customer on August 27, 2013. The letter identified the affected product, problem and actions to be taken. The request was to determine if their customer did or did not backfill microtiter-strips during the analytical process. If a customer backfilled, they were advised to return the completed request form. On September 11, 2013 Trinity Biotech issued a notice to their customers advising them to re-evaluate any negative patient results for those who perform backfilling. For questions contact Trinity Biotech Help Desk at 800-325-3424, option 2. |  | 38 | 38 kits | USA Nationwide Distribution in the state of WA | Open |  |  | 38 | Not_Computer
Z-0581-2014 | 67042 | prosthesis, hip, semi-constrained, metal/polymer, porous uncemented | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REFLECTION(R), MICROSTABLE(R), ACETABULAR LINER, 20o OVERHANG, 28 MM ID, 54-56 OD, QTY: (1), STERILE EO, REF 71742854, Smith & Nephew, Inc. Memphis, TN 38116 USA.Orthopaedic. | REFLECTION(R), MICROSTABLE(R), ACETABULAR LINER | 2 | December 26, 2013 | 2013 | Smith & Nephew Inc | One batch of the Reflection(TM) Polyethylene Acetabular Liners may be out of tolerance causing the liner not to fully seat into the shell. |  | All affected sales staff was notified of problem and the recall via email on 11/25/2013. There were no US hospitals affected by this recall. |  | 38 | 38 units | Distributed in China and Spain. | Open |  |  | 38 | Not_Computer
Z-0593-2014 | 67052 | automated external defibrillators (non-wearable) | MKJ | Automated external defibrillator. | Cardiovascular | Cardiovascular | 510(k) | Philips HeartStart MRx Monitor/DefibrillatorModels: M3536A Options A20 - A27, HeartStart MRx ALSMonitor (Grey)The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Philips Healthcare | 2 | December 27, 2013 | 2013 | Philips Medical Systems, Inc. | Error in the printing of the therapy energy setting labels on Philips HeartStart M3536A MRx Monitor/Defibrillators. The first energy setting is labeled 10 when it should be labeled 1-10 | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | Philips sent a Customer Information Letter dated June 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.During an EI of the firm, the firm had not reported the correction to FDA.The Customer Information letter informs the customers what the behavior is and under what circumstances. The correction consisted of a replacement label that was provided free of charge to all units affected by this issue. Philips Healthcare representatives have been contact customers with affected devices to arrange for installation of the replacement label.To correct this issue, Philips will replace the label free of charge. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the replacement label. This issue does not affect current product shipments.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconvenience caused by this problem. |  | 93 | 93 units | Worldwide Distribution - US Distribution including the states of IN, MS, TX and VA., and the countries of DENMARK, FRANCE, GERMANY, SPAIN and SWITZERLAND. | Open |  |  | 93 | Not_Computer
Z-0603-2014 | 67064 | valve, pressure relief, cardiopulmonary bypass | MNJ |  |  | Cardiovascular | 510(k) | Terumo Cardiovascular Custom Cardiovascular Procedure Kits containing an Overpressure Safety Valve, LN130B, in the Arterial Recirculation Line | Terumo Cardiovacular | 2 | December 27, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Over-Pressure Safety Valves not cleared with an indication for use on the arterial side of the extracorporeal bypass circuit. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Terumo Cardiovascular issued Customer notification letters on June 4, 2013. Accounts advised to: Review this notice and decide to continue to use or not use the Cardiovascular Procedure Kits as configured until a replacement kit with a one-way valve indicated for use on the arterial side of the perfusion circuit is available. If you have questions, contact Terumo CVS Customer Service: 800-521-2818. |  |  | 24,227. | NationwideCanada | Open |  |  |  | Not_Computer
Z-0587-2014 | 66499 | curette, surgical, general use | FZS | Manual surgical instrument for general use. | General & Plastic Surgery | General & Plastic Surgery | 510(K) Exempt | DIST LAT FEM LOCK PLT, LT | DIST LAT FEM LOCK PLT, LT | 2 | December 27, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged at a Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a packaging tray configuration which consists of an inner and outer sterile tray, resulted in failure | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Material/Component Contamination | On 11/5/2013, Urgent Medical Device Recall Notifications were sent to affected distributors, hospital Risk Managers, and doctors informing them of the failure. All distributors were notified via electronic mail. Hospital risk managers, surgeons and distributors with affected inventory within the expiration date of the device will also be notified via courier. The letter identifies the issue, health risks, and their responsibilities. Questions concerning the recall are directed to 1-877-946-2761Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitals/Surgeons100% of the notifications will be tracked to ensure delivery of the notifications. An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.Additional notifications via email and FedEx will be sent to those who have not completed the required certification form or returned affected product.Accounts will be deemed unresponsive after 3 attempts. |  | 65 | 42, 065 units total | Worldwide Distribution. | Open |  |  | 65 | Not_Computer
Z-0594-2014 | 66895 | set, administration, intravascular | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Asante Comfort Infusion Set with 60cm tubing & 13mm cannula packaged in 5-packs, singles,or inside a Combination Kit (FG-5123-4)Manufactured for:Asante SolutionsSunnyvale, CA;Manufactured by:UnomedicalLejre, Denmark | Asante Comfort Infusion Set | 2 | December 27, 2013 | 2013 | Asante Solutions, Inc. | Reports of the infusion tubing disconnecting from the connector, resulting in leakage of insulin from the infusion set and providing patient with less insulin than intended. | DESIGN: Device Design | Asante sent urgent medical device recall notifications letters to distributors and customers on November 19, 2013 by US mail, return receipt requested. |  | 834 | 834 kits | Nationwide Distribution. | Open |  |  | 6977 | Not_Computer
Z-0583-2014 | 66909 | impactor | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Ball Impactor Tip Product Usage: Used for ADM Surgery. | Stryker Orthopaedics Ball Impactor Tip | 2 | December 27, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker has received complaints associated with cracks and/or fracture of the Ball Impactor Tip and Rim Impactor Tip Instruments. | DESIGN: Device Design | Stryker sent an Urgent Medical Device Correction letter dated October 1, 2013 to all affected customers via Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and fax the attached Product Correction Acknowledgement form to 855-251-3635. |  | 1024 | 1,024 units | USA Nationwide Distribution | Open |  |  | 2714 | Not_Computer
Z-0584-2014 | 67012 | solid state x-ray imager (flat panel/digital imager) | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, a Stationary x-ray system, X-ray systemModel: 712070The Essenta DR is a multifunctional stationary X-ray system, in which the swivel arm rotation, the vertical movement of the swivel arm and the SID adjustment are motorized. It is designed for radiographic examination of the recumbent, standing or seated patient and also for lateral exposures. It enables the operator to quickly and safely move the unit to all the pre-programmed basic positions at the press of a button, using the autopositioning feature. Fine positioning of the tube/collimator and detector on the patient is easy to carry out using the command. | Philips Essenta DR | 2 | December 27, 2013 | 2013 | Philips Medical Systems, Inc. | X-ray system C-arm casting may loosen and fall | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | Philips Medical Systems issued a Urgent Field Safety Notice-Warning Label and Inspection, dated November 21, 2013, advising of potential of C-arm detaching and requesting a daily inspection and a safety warning label . If you need any further information or support concerning this issue, please contact your local Philips representative. 1-800-722-9377. |  | 424 | 424 units | Worldwide Distribution - USA (Nationwide) including Puerto Rico and the countries of : Afghanistan, Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hungary, India, Indonesia Iran,Italy, Jordan, Kazakstan, Kenya, Liberia, Lithuania, Morocco, Netherlands, New Zealand, Nigeria, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Thailand, Turkey, Ukraine, United Arab, United Kingdom and Viet nam. | Open |  |  | 424 | Not_Computer
Z-0605-2014 | 67036 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂŽ Dual Species Self-Contained Biological Indicator, Distributed by STERIS Corporation, Mentor, OH.Intended for use in installation testing and routine monitoring of steam sterilization and ethylene oxide sterilization processes. | VERIFY Dual Species SelfContained Biological Indicator | 2 | December 28, 2013 | 2013 | Steris Corporation | STERIS has identified that the population of the G. stearothermophilus spores that monitor STEAM sterilization cycles does not meet specifications for Lot # 131004 of the Verify Dual Species Self-Contained Biological Indicators. As a result the spore count on the affected product is lower than the labeled minimum and the product may incorrectly indicate that a 5-log spore kill was achieved. |  | On September 10, 2013, the firm sent Urgent Voluntary Recall Notices to their customers. |  | 1238 | 1,238 boxes | Worldwide Distribution-USA (nationwide) including the states of AK, Al, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY, and the countries of BAHRAIN, CANADA, COLUMBIA, GERMANY, ITALY, MALAYSIA, PHILIPPINES, ROMANIA, SPAIN, and THE NETHERLANDS. | Terminated | February 14, 2014 | 48 | 1238 | Not_Computer
Z-0604-2014 | 66981 | indicator, biological sterilization process | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI, Part No. 14324.Intended to be used as a standard method for frequent monitoring of the STERRAD Sterilization System cycles. | STERRAD Cyclesure Biological Indicator | 2 | December 28, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD Cyclesure Biological Indicator (BI) because ASP has identified a microscopic intermittent and low frequency flaw in the outer vial of the STERRAD Cyclesure 24 Biological Indicator. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A customer notification letter was sent to all customers on 11/26/2013 who purchased the STERRAD CYCLESURE Biological Indicator (BI). The letter informs the customer of the problems identified and the actions to be taken. The customer notification letter includes a table with the products on recall which includes the manufacturing dates, lot numbers, expiration dates, and quantity. |  | 4048 | 4048 units total (4014 units in US) | Worldwide Distribution. | Open |  |  | 4048 | Not_Computer
Z-0607-2014 | 66979 | system/device, pharmacy compounding | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | IntelliFill i.v. Pharmacy Compounding Device (PCD).Intended to fill syringes from 0.5mL up to 11.5mL using mixing instructions from its formulary to fill the orders. | IntelliFill i.v. Pharmacy Compounding Device (PCD) | 2 | December 29, 2013 | 2013 | Baxter Healthcare Corp. | Following a retrospective review for Baxter Daytona Complaint CR-954, it was identified that the IntelliFill i.v. Pharmacy Compounding Device (PCD) (product code: IFIV) has the potential to continue to run and fill syringes with the HEPA filtration disabled. | COMPONENT CONTROLS (GMP - GOOD MANUFACTURING PRACTICE): Nonconforming Material/Component | A Service Bulletin was sent to all IntelliFill i. v. customers February 22, 20 I 0 by USPS mail.The Service Bulletin informs customers that IntelliFill i.v. operators are to be trained to listen for the running HEPA filtration fan motor and observe/feel for airflow coming from the automation deck.The field service engineer dispatched to the site replaced the damaged sensor on 11-Feb-2010 and a Service Bulletin was sent to all customers regarding the function and use of the HEPA filter on 22-Feb-2010.A new HEPA sensor was designed and installed into all current IntelliFill i.v. devices between 21-July-2011 to 9-Feb-2012. |  | 28 | 28 devices | Nationwide Distribution including the states of CA, CO, CT, FL, IL, LA, MA, MD, MI, NC, NE, NY, OH, OR, PA, TN, TX, and WV. | Open |  |  | 28 | Not_Computer
Z-0606-2014 | 66902 | circuit, breathing (w connector, adaptor, y piece) | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | RTC 25-C Inline Aerosol Tee Adapter.Intended to be used to adapt an aerosolized medication delivery system having an elliptical mouthpiece to a ventilator circuit. | RTC 25C Inline Aerosol Tee Adapter | 2 | December 29, 2013 | 2013 | Instrumentation Industries Inc | Some packages of the RTC 25-C may have been shipped to customers with incomplete Installation and Usage Directions due to a printing error. | CHANGE CONTROL (GMP - GOOD MANUFACTURING PRACTICE): Labeling Change Control | A recall notification letter, dated November 13, 2013, was sent to Users. |  | 220 | 220 | Nationwide Distribution. | Open |  |  | 220 | Not_Computer
Z-0608-2014 | 66968 | prosthesis, hip, semi-constrained, metal/polymer, cemented | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CPT Hip System, Parts 00-8114-002-10 and 00-8011-004-00Hip joint metal/polymer semi-constrained cemented prosthesis.The CPT Hip System is indicated for cemented use in: Patients suffering from severe hip pain and disability due to rheumatoid arthritis, osteoarthritis, traumatic arthritis, polyarthritis, collagen disorders, avascular necrosis of the femoral head, and nonunion of previous fractures of the femur. Patients with congenital hip dysplasia, protrusio acetabuli, or slipped capital femoral epiphysis. Patients suffering from disability due to previous fusion. Patients with previously failed endoprostheses and/or total hip components in the affected extremity. Patients with acute femoral neck fractures. 510(k) K960658 | CPT Hip System | 2 | December 30, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the CPT Hip Stem, Versys Build-Up Block, and Trabecular Metal Reverse Glenosphere medical devices due to reports of the products being packaged in the incorrect outer carton. Product labels and patient labels are correct. | PRODUCTION CONTROLS: Labeling Mix-Ups | Zimmer sent an Urgent Medical Device Recall letter dated December 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of the contents. If affected product was found, it should be quarantines and their Zimmer representative notified. The Zimmer sales representative will remove the recalled product from their facility. Customers with questions were instructed to call 1-877-946-2761. . |  | 36 | 36 | Worldwide Distribution -USA including OH, FL, WI, MI NJ, SC, CA, NC, PA, OR, WA, TN, KY, OK, KS, AZ, TX, CO, NV and Internationally to Germany.There have been 71 distribution events. Note that this product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf. 7 units have been returned, including three through the complaint process, which leaves 64 affected units in the field subject to removal. | Open |  |  | 64 | Not_Computer
Z-0618-2014 | 66468 | bed, ac-powered adjustable hospital | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 Med/Surg BedSecure II and the 3002 S3 are the beds typically found in the Med-Surg units of the hospitals. The bed is designed for a 15 year expected service life under normal use conditions, and with appropriate periodic maintenance as described in the maintenance manual for each device. There are minimal differences between the functionality of the two models, each of which function as a highly complex system of a large number of individual components. | S3 Med/Surg Bed | 2 | December 31, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified a potential issue associated with Secure II and S3 Med/Surg beds built between April 2nd, 2005 and July 26th, 2011. The affected beds have the potential to experience damage to the Foot End Cover and/or Power Coil Cable due to reduced clearance between the Foot End Cover and Foot End Lift Header. | DESIGN: Device Design | Stryker Medical sent an Urgent Medical Device Recall Notification letter on November 26, 2013, to all affected customers. The notification included the products being recalled, the product description, product issue, potential hazards, and actions needed. Customers were instructed to locate and inspect the beds listed in the notice and ensure that no damage has occurred to the foot-end cover, power coil cable, or sensor coil cable. If the beds were not damaged customers may continue to use them. If damage occurred, customers wsere instructed to remove the beds from service and contact Stryker at 1-800-327-0770, Option 2, then Option 2 to arrange a service call. Continue to inspect the foot-end cover and coil cables as part of their routine preventive maintenance. Customers were instructed to fax the enclosed business reply form to 269-488-8691 or email productfieldaction@stryker.com. Customers were instructed to forward the letter if the product was loaned or sold to other customers. If any beds were disposed of and were no longer in use, customers should advised Stryker by providing them with the serial number. Questions or concerns are directed to Stryker Customer Service at 1-800-327-0770. Consignees are to fax (269-488-8691) or email (productfieldaction@stryker.com) the business reply form to confirm receipt of this notification.For questions regarding this recall call 269-324-6609. |  | 108769 | 108,769 Total | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, China, Argentina, Brazil, Germany, Korea, Mexico, Singapore, UK, Europe, Middle East, Africa, and Latin America. | Open |  |  | 108769 | Not_Computer
Z-0614-2014 | 67128 | catheter, continuous flush | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Catheter, Continuous FlushThe Merit Microcatheter is intended for general intravascular use, including peripheral and coronary vasculature. Once the subselective region has been accessed, the microcatheter can be used for the controlled and selective infusion of diagnostic, embolic or therapeutic materials into vessels. | Maestro Microcatheter | 2 | December 31, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is conducting a recall for 3 lots of Maestro Microcatheters with a graphicdiscrepancy on the carton/box label end panel. The catheter graphic on the front of the cartonbox (correct) and the graphic on the end panel (incorrect) do not match. | PRODUCTION CONTROLS: Packaging Process Control | Merit Medical Systems sent an Urgent Product Recall Notice dated December 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of the Maestro Microcatheters identified and isolate them in a quarantined area until the product is over-labeled by their Merit Sales Representative Customers were instructed to ensure that all personnel to whom the devices were distributed were made aware of the field action. Customers with questions were instructed to contact their sales representative or call 1-801-826-4062.For questions regarding this recall call 801-253-1600. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 27, 2014 | 27 | 59 | Not_Computer
Z-0615-2014 | 65007 | instrument, surgical, orthopedic, ac-powered motor and accessory/attachment | HWE | Surgical instrument motors and accessories/attachments. | General & Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Hand Switch - Electric Pen Drive, Part Number 05.001.012The Synthes Electric Pen Drive is indicated for screw insertion, pin and wire replacement, cutting of bone and metal, drilling, decorticating, shaping, and smoothing of bones and teeth in a wide variety of surgical procedures, including, but not limited to general orthopaedic trauma, foot, hand, maxillofacial, neurosurgical, oral, otolaryngolical, reconstructive, and spine surgery. | Synthes Hand Switch Electric Pen Drive | 2 | December 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes received a complaint from a customer in which the Hand Switch for the Electric Pen Drive was not labeled correctly. The "on" and "lock" labels were interchanged. | DESIGN: Process Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product, remove from inventory, and return them to Synthes immediately for replacement. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located. Return the Verification Form with the product to: Credit/Returns, Synthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should still complete the Verifcation Section at the end of the letter by checking the appropriate box and faxing the document to 610-251-9005 or email to: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 290 | 290 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada.. | Open |  |  | 290 | Not_Computer