## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0126-2008 | 45328 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Viasys Healthcare NicoletOne ICU Monitor, Modular Neurodiagnostic System (containing Axiomtek Panel PC computer system), released and installed between December 2004 and March 2007 (Axiomtek Panel PC Model Number: 2147-671) VIASYS NeuroCare Part Numbers: 842-673800, 842-673900, 982A0421 and 982A0422. |  |  | September 11, 2007 | 2007 | Nicolet Biomedical Div of Viasys Healthcare | Display lock-up: VIASYS Healthcare NicoletOne ICU Monitor freezes up during operation when using the digital video option. | Component design/selection | Certified direct mail delivery of URGENT MEDICAL DEVICE FIELD CORRECTION (RECALL) were made to consignees on 9/11/07 by Nicolet Biomedical Div of Viasys Healthcare. The letter describes the product and problem. VIASYS requests that customers cease use of the video option of the device until the monitor can be replaced by VIASYS authorized personnel. Letter also requests consignees to fax or mail back the response form. |  | 33 | 33 | Worldwide Dsitribution - USA, Germany, Netherlands, Kuwait, Japan, Ireland, and China. | Terminated | December 10, 2011 | 1551 | 33
Z-0246-2008 | 45153 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph 16, Biograph PET/CT Scanner, Catalog No. 872858, Manufactured by Siemens Medical Solutions USA, Inc., 810 Innovation Drive, Knoxville, TN 37932 USA, TEL: 865-218-2534 |  |  | July 23, 2007 | 2007 | Siemens Medical Solutions USA Inc | Visual asymmetric appearance; An asymmetry can be introduced into the attenuation-corrected PET images when using iterative reconstruction in combination with a specific number of subsets and reconstruction matrix sizes. (This problem only occurs when using protocols other than the default Biograph protocols) | Software design | A certified letter-Urgent Device Correction- dated June 15, 2007 was sent to all affected consignees on July 23, 2007 informing them of the problem (PET iterative reconstruction algorithm) and requesting as a short term fix that they discontinue the practice of operating using protocols other than the default Biograph. Software upgrades as the long term fix were promised in the Urgent Device Correction Notice. Updated information received October 4, 2007: a software patch to fix PETsyngo 5.0 was released, software patch to fix PETsyngo 4.x was released, and target date for the software patch to fix PETsyngo 3.2 and 3.4 is the end of November 2007. |  | 162 | 162 units | Worldwide-USA and countries of Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Czech Republic, Germany, Denmark, Spain, France, Great Britain, Hong Kong, Hungary, Ireland, India, Italy, Japan, South Korea, Kuwait, Lebanon, Mexico, Malaysia, Netherlands, Norway, Poland, Portugal, Sweden, Singapore, Slovak Republic, Thailand, Turkey, and Taiwan. | Terminated | September 11, 2009 | 781 | 415
Z-0301-2008 | 45446 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott Diagnostics CELL-DYN Sapphire Hemoglobin Reagent Analyzer Syringes, List Number 08H49-02, Abbott Laboratories, Abbott Park, IL (the syringe is packaged in a white box and sealed with a label wrapped around the center of the box-The main analyzer list number is 08H00-01, ) |  |  | July 25, 2007 | 2007 | Abbott Laboratories | Syringe Failure: Insufficient lubrication of the plunger tip may cause syringes to fail at installation or shortly thereafter | Nonconforming Material/Component | This Recall was initiated 07/24/07 via a Product Recall letter-Immediate Action Required-instructing users to discontinue use and destroy any units under recall. Replacements may be obtained by contacting the local Customer Support Center at 1-877-4Abbott |  | 302 | 302 units | Worldwide Distribution: USA including states of AR, AZ, CA, CT, FL, GA, ID, KY, LA, MD, MI, NC, ND, NV, NY, OH, OK, PA, SC, TX, UT, VA, and W I (4 distributors and 50 consignees)and countries of Australia, Brazil, Chile, Germany, and Hong Kong. | Terminated | November 17, 2009 | 846 | 302
Z-0460-2008 | 45504 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Cardinal Health Alaris Pump Module (formerly Medley Pump Module), Model 8100, Cardinal Health, San Diego, CA. (All Alaris Pump modules shipped prior to December 23, 2007 are subject to this recall) |  |  | October 29, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | Inaccurate flow rate; related to misassembled (missing, bent or broken) springs during the manufacturing or servicing of the mechanism assembly | Process control | The field notification (Urgent Medical Device Recall Letter) will consist of a mailing conducted through consignees (distributors in the affected countries as well as direct customers for the Alaris Pump module (a.k.a. Medley Pump module). The domestic letters are expected to be mailed on November 5, 2007 to all of the affected accounts by return receipt mailing. A Customer Response Card will be mailed with each letter along with a Service Bulletin detailing the protocol to conduct the Occluder Pressure Test. Four (4) letters will be mailed to each account to the following titles/job functions: Director of Nursing, Director of Biomed, Director of Materials Management, and Director of Risk Management. Distributors will be sent multiple copies of the Recall Notification, Customer Response Cards, and Service Bulletin, along with a cover letter outlining their responsibilities to complete the Recall notification plan and notify their customers as requested. The letters will be the first source of communication regarding this issue. The Recall Center will be contacting customers shortly to make arrangements for inspection of the devices at the users site. The international mailings are being coordinated out of the offices located in the affected countries. The letter will be translated to the appropriate language and mailed to all of the affected customers. The appropriate international regulatory authorities will be notified of this issue, and will work with the country managers.---------A web site for all information on this recall is listed as: www.cardinalhealth.com/alaris/indexmodulealert.asp******************January 23, 2008 the firm notified the Los Angeles Districe that they will be sending a recall notification update to 267 customers and distributors who are affected by this 806 Notification Amendment. Cardinal Health intends to notify these consignees by mail on January 30, 2008. This update will affect a total of 267 customers, 5,681 units and 798 |  | 207257 | 207,257 units | Worldwide: USA including the following forty-six (46) states, along with Washington, D.C. Puerto Rico and Guam: Alaska, Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, Nebraska, New Hampshire, New Jersey, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, and Wyoming and countries of Canada &amp; Saudi Arabia (Note: Devices were shipped to the Cardinal Health Distribution Center in the Netherlands. These devices were then shipped to Saudi Arabia.) | Terminated | June 06, 2011 | 1316 | 207257
Z-0457-2008 | 45799 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | bioMerieux, BacT/ALERTÂż FA Culture Bottles, bioMerieux, Inc., Durham, NC 27704 |  |  | November 05, 2007 | 2007 | bioMerieux, Inc. | Duplicate Bottle Indentification -- Three lot numbers of BacT/ALERTÂż FA Culture Bottles labels contained duplicate bottle identification (Bottle ID) barcode numbers. (Lots 1017685 and 1017788 contain the same series of Bottle ID's; Lot 1017693 contains duplicate Bottle ID's within the lot and may share Bottle ID's with both lots 1017685 and 1017788) | Process control | bioMerieux notified all Consignees with a Urgent Product Recall Notice letter on 11/05/2007. They were instructed to examine their inventory to determine if they have any remaining inventory of the 3 affected lots. If they have any of these bottles they were instructed to stop using them and destroy all remaining inventory. A Customer Acknowledgement Form was attached to be completed and returned to bioMerieux via fax or mail. There were also instructions attached for the use of Generic Bottle Barcodes and Creating a Unique Bottle ID. |  | 7025 | 7,025 | Worldwide Distribution-USA including countries of Australia, Austria, Bangladesh, Brazil, Bulgaria, Byelorussia, Chile, China, Colombia, Croatia, Denmark, Dijbouti, Estonia, Federation of Russia, Finland, France, Germany, Greece, Guadeloupe, India, Israel, Italy, Korea, Kuwait, Martinique, Mexico, Netherlands, New Caledonia, Norway, Panama, Poland, Portugal, Reunion, Saudi Arabia, Serbia Montenegro, Slovenia, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Terminated | December 08, 2009 | 764 | 7025
Z-0566-2008 | 45904 | No | MQX | Acupuncture needle. | General Hospital | General Hospital | 510(K) Exempt | HAENG LIM --Acupunture Needles- PRE-STERILIZED GAMMA-RAY; 100 PKGS (10 NEEDLES/PACKAGE-1,000 needles in each box). Sale Agency HAENG LIM SEO WON #45-9, Gunjang-ri, Ganam-myeon, Yeoju-gun, Kyungki-do, Korea |  |  | November 28, 2007 | 2007 | Marknew Products | Marketed without 510(k) | PMA | Marknew Products, LLC, sent an Important Recall Notice to their customers on November 27, 2007, advising them that all Haeng Lim Seo Won Acupuncture Needles that they have received, have not been cleared for distribution in the USA. All customers have been requested to return product back to Marknew Products for credit. A response form was included. If customers have questions concerning this product, they are asked to call 562-908-5000. |  | 32038000 | 320-380 boxes/100 packages (10 needles/package) | Nationwide Distribution | Terminated | January 18, 2012 | 1512 | 32038000
Z-0320-2008 | 45852 | No | HJM | Transilluminator. | Ophthalmic | Ophthalmic | 510(K) Exempt | Transilluminator, Model BFO-150, a Fiber Optic Light Source (Illuminator), available in 2 voltages: 115v for USA users &amp; 220v for foreign countries. --- Chiu Technical Corporation, 252 Indian Head Road, Kings Park, NY 11754. |  |  | November 09, 2007 | 2007 | Chiu Technical Corp | Device was marketed without a 510(k). | Process design | On 11/09/07, the recalling firm started faxing the Recall Letter and response form to its customers, informing them of the recall of the unapproved device to the user level. |  | 191 | Domestic: 170 units, Foreign: 21 units | Worldwide: USA, Canada and India | Terminated | October 17, 2012 | 1804 | 191
Z-0300-2008 | 45437 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Ventana Medical System's Symphony Staining System, KPA, Model #: 2200000, Catalog #: 900-SYM1 &amp; 900-SYM3, Ventana Medical Systems, Tucson, AZ |  |  | September 06, 2007 | 2007 | Ventana Medical Systems Inc | Shock hazard: Devices released prior to full testing; the configuration of the Symphony instrument during Hi pot testing procedures did not completely evaluate all components in the test loop. In specific circumstances, a shock hazard exists. | Process control | On 9/6/07, the firm began visiting customers to perform the HiPot test. Prior to performing the High Potential Test of Symphony Instruments already in the Field, the Field Service Engineer or FSE ensures that the customer fully understands the purpose of the test and why it is necessary to perform the test at their site. The FSE conveys the following points regarding the HiPot test:- The HiPot test is a safety verification for the electrical path within the instrument.- The HiPot test was performed incorrectly during the manufacturing process.- The HiPot test is used to verify that if a high voltage overload condition were to occur it would not lead to a short/arc within the instrument, which could pose a risk to anyone that touches the instrument- The HiPot test places a voltage of 2,000VDC onto the high voltage path of the instrument and if any of the insulating components are compromised, the tester will notify the FSE of the condition.- If the instrument fails the HiPot test, the faulty component will be replaced and the instrument will be retested until it passes.- Due to the high voltage used the during the test, no one will be allowed to touch the instrument until the test is complete.The FSE performing the test documents that the customer was informed of the reason for the High Potential Test on their Field Service Report. |  | 45 | 45 units | Worldwide: USA, Japan and France | Terminated | July 08, 2008 | 306 | 45
Z-0299-2008 | 45397 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT i System Assay CD-ROM - WW (excluding USA accounts), Versions prior to Version 24.0; for use on the ARCHITECT i System; List Number (LN) 6E59,Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA |  |  | September 18, 2007 | 2007 | Abbott Laboratories | The lower limit flag for auto-dilution protocol 3 (amniotic fluid) was set too low (15 ng/mL) for the ARCHITECT AFP assay file for the ARCHITECT AFP non-U.S. product. The assay labeling indicates that the limit flag should be set no lower than 20 ng/mL | Software design | Abbott affiliates were e-mailed copies of the Product Correction (Immediate Action Required) Recall Letter on 9/18/07 for translation where needed and dissemination to their ARCHITECT accounts who have the affected assay CD-ROMs. Each Abbott international organization will be responsible for contacting customers and determining the method of effectiveness checks based on their country requirements. Abbott sent Product Correction, Immediate Action Required letters dated 18 September 2007 to all active non-U.S. ARCHITECT customers who have the affected assay CD-ROMs. The end users were informed that current AFP file has been updated on a new ARCHITECT I Assay CD-ROM (LN 6E59, version 24). The Low-Linearity flag threshold for the Amino 1:50 dilution has changed from 15 ng/mL to 20 ng/mL for consistency with overall assay sensitivity. The accounts were advised to install the updated version of the current ARCHITECT AFP assay when they receive the version 24 ARCHITECT I Assay CD-ROM. The accounts were requested to keep a copy of the letter on file. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect testing, and fax it to Abbott by 9/28/07 at 1-800-777-0051. |  | 2490 | 2,490 | International Only: Australia, Canada, Chile, Costa Rica, Dominican Republic, Germany, Guatemala, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Thailand, Uruguay and Venezuela. There were no USA accounts. | Terminated | November 05, 2008 | 414 | 2490
Z-0310-2008 | 45461 | No | JDK | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VariLift Bone Plug w/ End Cap; 13mm Dia, 24mm long; Ref #: CIE1-13-24-BP; Sterile; Encore Orthopedics, Austin, TX 78758. |  |  | September 19, 2007 | 2007 | Encore Medical, Lp | Device marketed without label warning required in the device's 510(k) approval letter. (In part-Warning: This device is not intended for any spinal indications.....) | Labeling Change Control | The recalling firm notified consignees of its recall via an Urgent Device Recall letter on 09/19/07. Consignees notified of a labeling error regarding device usage and were told they would be notified by a customer service rep to facilitate return of the recalled devices. |  | 61 | 61 units. | Nationwide: including direct accounts in CO, KS and TX. | Terminated | March 10, 2008 | 173 | 175
Z-0297-2008 | 45511 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens, ACUSON Antares, Ultrasound System (Standard and Premium Edition), 5.0 Systems with software version 200.0.054, P/N : 10032747 and 10032746 standard, 10032746, 10037591, and 10038202 for premium edition, Siemens, Ultrasound Division, Mountain View, CA 94039 |  |  | July 10, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Image errors: A software problem results in on-screen indications (graphics) that lead the user to believe that the patient's right and left and the transducer orientation are oriented the same. This is not the case and the actual patient's right and left are reversed from the on-screen indicators. | Software design | The firm sent a recall letter, Urgent Medical Device Notification, to consignees on 07/10/2007, to be followed by a service visit by Siemens personnel for a mandatory software update to be installed. Short term the users were instructed to avoid increasing the SieClear compounding levels to an Advanced SieClear compounding setting after an L/R invert action with Biopsy guidelines on-screen. |  | 418 | 418 units (Z-0297-0298-2008) | Worldwide: Product was distributed to medical facilities Nationwide including the states of OH, FL, IL, NY, CA, NH, TX, NJ, PA, MD, MO, GA, WA, WI, NE, WV, HI, MA, SC, MI, VA, AZ, UT, NH, RI, KY, SD, and IN, and internationally to Canada, Germany, Australia, Sweden, Saudi Arabia, Chile, South Korea, Mexico, France, UK, Taiwan, Singapore, Denmark, Brazil, New Zealand, Ireland, Belgium, Norway, Serbia, Vietnam, UK, Poland, China, Spain, South Africa, Austria, Croatia, Russia, Greece, Trinidad and Tobago, United Arab Emirates, India, Japan, Malaysia, Turkey, Thailand, Norway, Slovakia, Romania, Netherlands, Switzerland, Portugal, Italy, and Venezuela. | Terminated | December 20, 2010 | 1259 | 418
Z-0305-2008 | 45811 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba Aplio XG Diagnostic Ultrasound System, Model: SSA-790A with version 1.1 software, Toshiba America Medical Systems, Inc., Tustin, CA. 92780 |  |  | October 02, 2007 | 2007 | Toshiba American Med Sys Inc | Panel Lock-up: When using dynamic image acquisition modes the panel may lock up, if the operator attempts to perform a still image capture. If the panel locks up the system will have to be rebooted to continue the examination which may cause a delay in completion of the patient examination. | Software design | Notification letters, Urgent Medical Device Correction, with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department.--A short term instruction included "Avoid using the Freeze Key when performing dynamic image acquisitions" |  | 69 | 69 | Nationwide: USA including the states of AK, AZ, CA, FL, IA, IL, IN, KS, MI, MN, MO, MT, NE, NV, NY, OH, OR, PR, VA, WI &amp; WV | Terminated | April 20, 2012 | 1662 | 69
Z-0447-2008 | 45460 | No | CBH | Tracheal tube fixation device. | Anesthesiology | Anesthesiology | 510(K) Exempt | Dale 242 Blue (Neonate, Infants and Small Children) Tracheostomy Tube Holder, Product Number: H84102421, Dale Medical Products, Inc., Plainville, MA 02762 |  |  | September 27, 2007 | 2007 | Dale Medical Products, Inc. | Revised Labeling : Instructions For use: "Do not trim fastener hook tabs. " Trimming can cause tabs to disengage at the trach plate | Labeling design | Dale Medical notified Distributors by letter dated 9/27/07, identified as an Urgent Medical Device Correction Letter. The letter requested distributors to notify all their customers by letter. The Correction Letter reminded users to take notice of new labeling " Do not trim fasterner hook tabs" |  | 65000 | 65,000 | Worldwide-USA and countries of Canada, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Iceland, Italy, The Netherlands, Norway, The UK, Spain, Switzerland, Australia, Hong Kong,India, Japan, Korea, Malaysia, New Zealand, The Philippines,Singapore, Taiwan, ROC, Brazil, Bermuda, Chile, Costa Rica,Ecuador, Mexico, Panama, Venezuela, Thailand, Saudi Arabia, and The United Arab Emirates. | Terminated | May 09, 2012 | 1686 | 665000
Z-0302-2008 | 45474 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Extron Electronics OEM, CAB, DVI-Male to DVI-Female Cable Assembly (14 inch.); Part #: 43-111-01 (Stryker Part #: 0100-224-131); Distributed by Stryker Communications, Flower Mound, TX 75028. (These cables are commonly used with the Switchpoint Infinity II) |  |  | October 03, 2007 | 2007 | Stryker Communications | Intermitten or Complete Loss of Signal: out-of-specification DVI cables causing intermitten or complete loss of signal on monitors used in surgery. | Nonconforming Material/Component | Firm notified consignees of recall via letter (Urgent Device Recall) on 09/21/07. Consignees were notified that a customer service representative would visit each facility to replace out-of-spec cables with properly shielded cables at no charge. |  | 2255 | 2,255 cables. | Worldwide: USA and Australia, Belgium, Canada, France, Germany, Great Britain, Hong Kong, Japan, Korea, Mexico, Netherlands, New Zealand, Portugal, South Africa, Spain, Sweden, Switzerland, Taiwan, and The U.A.E. | Terminated | June 16, 2008 | 257 | 3908
Z-0376-2008 | 45494 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Cepheid, Xpert GBS Assay Diagnostic Test Kit Part No: GXGBS-100N-10 (in foil pouch) for use with the GeneXpert DX System, Cepheid, Sunnyvale, CA 94089 |  |  | September 05, 2007 | 2007 | Cepheid | Labeling Error: Some cartridges may be mislabeled as "Xpert EV" instead of "Xpert GBS". (The information embedded in the barcode label on the cartridge is correct for Xpert GBS and brings up the GBS assay definition file) | Release of Material/Component prior to receiving test results | Consignees were sent a Correction and Removal letter on 09/04/2007, notifying them that there was a labeling error with some of the cartridges labels indicating Xpert EV instead of Xpert GBS on the cartridge label. They are asked to discontinue using it immediately and return to Cepheid. |  | 105 | 105 | Worldwide Distribution | Terminated | April 04, 2008 | 212 | 105
Z-0488-2008 | 45870 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Premier Edge, Bulk Microsurgical Knife, 3.0mm Slit Knife, Item Code B-PE3830, Oasis, Glendora, CA |  |  | July 18, 2007 | 2007 | Oasis Medical Inc | Sterility (package integrity) compromised: the device is puncturing the packing during handling. Once the packaging is damaged, the product is no longer sterile as labeled. THIS IS AN EXPANSION OF A PREVIOUS RECALL | Device Design | A recall notification, Urgent Medical Device Recall Expansion Letter, dated 7.18.2007, was sent to customers advising the letter is to inform them of a voluntary product recall expansion involving the OASIS Medical Premier Edge Knives. A product label for ease in identifying the product is included. The letter further advises this voluntary recall has been expanded due to the product puncturing the packing during rough handling. Once the packaging is damaged, the product is no longer sterile as labeled. The letter instructs customers to immediately examine their inventory and quarantine product subject to recall. Once the product has been identified, product is to be returned by FedEx to OASISMedical. The letter further instructs that if the customer has further distributed this product, to identify their customers and notify them at once of this product recall. This recall should be carried out to the user level. |  | 28319 | 28,319 individual knives (Z-0485-0516-2008) | Worldwide: USA, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Norway, South Africa, Saudi Arabia, Slovakia, Spain, Sweden, Switzwerland, Turkey, United Kingdom, Ecuador, Finland, France, Germany, Greece, Italy, Denmark, Columbia, Chile, Brazil, Belgium, and Australia | Terminated | September 25, 2009 | 800 | 28319
Z-0326-2008 | 38368 | [6 empty] | Actiwatch 64 AM, Bandless Part Number 198-0301-AM1, Respironics, Inc., Mini Mitter, a Respironics Company, Bend, Oregon |  |  | June 27, 2007 | 2007 | Mini-Mitter Company, Inc. | Communication Failure: the devices have a memory chip that has a limited number of read/write cycles which can trigger set up information and/or subject information anomalies or intermittent communication annomalies. If left unaddressed, the issues can lead to permanent communication failure. | Software design | A Voluntary Medical Device Recall letter dated June 26, 2007. The recall letter will instruct the customers to 1) fill out a business reply form to confirm contact information and devices serial numbers in their possession. Then FAX this info back immediately to Stericycle or Mini-Mitter. 2) Refer to the work scheduled rework date reserved for them and noted in their recall notification letter and return the device for service at that time; and 3) Package the device appropriately, insure it and return it via UPS using the provided UPS shipping label. Product returned to Mini-Mitter will be quarantined, decontaminated, and repaired. Min-Mitter will rework all the devices returned. Each will have a new memory chip installed, be re-calibrated, retested, and have a new battery installed if required. Min-Mitter will provide a new one year warranty. |  | 20 | 20 | Worldwide | Terminated | October 20, 2010 | 1211 | 1073
Z-0150-2008 | 45405 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datascope, Panorama Patient Monitoring Network; Panorma Telepack 608;Software Versions 8.1.X, 8.3, 8.3.1, 8.4.1, and 8.5.X, Part Number: 0998-00-0191-04, Datascope Corp., Mahwah, NJ |  |  | October 09, 2007 | 2007 | Datascope Corporation | Display problems: If an ECG cable, which has been damaged due to the ingress of liquid or by mechanical trauma, is utilized with the Telepack, it may cause the Telepack to switch between the 3-lead and 5 lead-input modes. During the switching process, the digital heart rate displayed at the Panorama Central Station will be frozen, and subsquent arrthymia alarms will not be announced. | Software design | Urgent Product Field Correction notices were sent to hospital administrators on October 9, 2007, by certified mail, return receipt requested. The letter describes the scenario that causes the problem and its remediation It also describes how the software upgrade will fix the problem and alerts hospitals that a Datascope Service Representative will be contacting them to arrange for the upgrade. |  | 338 | 338 sites | Nationwide | Terminated | November 04, 2008 | 392 | 338
Z-0295-2008 | 44941 | timer, general laboratory | JBS | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Nova 8 Analyzer Calibrator Pack (Catalog/REF # 15196), Nova Biomedical, Waltham, MA |  |  | September 17, 2007 | 2007 | Nova Biomedical Corporation | Elevated Normalized Ionized Calcium (nCa) and Normalized Ionized Magnesium (nMg) calculated values on patient samples | Device Design | Nova issued a Customer Advisory Notice notification by phoneon 9/17/07 to contact customers using a script . If the facility reports Normalized Ionized Calcium or Normalized Ionized Magnesium, Nova Biomedical will replace existing stock of Nova 8 calibrator packs. If the facility does not report Normalized Ionized Calcium or NormalizedIonized Magnesium, the account can continue to use the existing lot of calibrator packs.This contact will be documented within the Nova Biomedical complaint database. Intemational Customer Advisory Notice being sent to Nova Biomedical Subsidiariies and Distributors by Federal Express Envelope.The Federal Express envelopes tracking number will be used for delivery confirmation. |  | 2999 | 2,999 units | Worldwide, USA, Canada, Argentina, Bangladesh, Bolivia, China, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,Germany,Greece, Hungary, India, Indonesia, Ireland,Italy, Japan, Korea, Mexico, Peru, Poland, Scotland, South Africa, Spain, Sri Lanka. Switzerland, Syria ,Taiwan, Turkey and The United Kingdom | Terminated | March 16, 2012 | 1642 | 2999
Z-0128-2008 | 45431 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Biomet 3i, Certain PreFormance Temporary Cylinder, Catalog Number: IPFTC42, Dental Implant, Biomet 3i, Palm Beach Gardens, FL 33410 |  |  | October 08, 2007 | 2007 | Biomet 3i | Incorrect assembly: The product does not allow for the screw to pass through the access hole and engage the implant properly. Therefore, the provisional restoration will not be seated on the implant platform. | Process control | Beginning on 9/27/07, Biomet 3i contacted its US consignees by telephone using a prepared script. Following telephone contact, Biomet 3i sent a recall letter, titled Urgent Medical Device Recall, by fax. The international consignees were contacted by email followed by a faxed letter. Each customer was requested to examine their inventory for the affected part and corresponding lot number, and return them to Biomet 3i. |  | 114 | 114 | Worldwide Distribution - USA, Italy, France, Spain and UK. | Terminated | July 10, 2008 | 276 | 114
Z-0484-2008 | 45822 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SERFAS Energy Suction Probe, Model 279-351-100, 3.5mm 90-S, Manufactured by Stryker Puerto Rico |  |  | April 26, 2007 | 2007 | Stryker Endoscopy | Device Shut off/on: Device may not turn off after taking finger/foot off of activating switch and may activate without pushing a button. (Patient burns may result from these device malfunctions) | Mixed-up of materials/components | Consignees were mailed an Urgent Device Removal Recall notice, dated April 27, 2007. International subsidiaries will receive the recall notice and they will distribute to their consignees. Product is to be returned as a recall. |  | 6037 | 6,037 units | Worldwide Distribution: USA including states of AL, AR, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, WY, and WV; INTERNATIONAL: Affiliated Stryker Distribution Companies: Sweden (for Scandinavia area), Australia, Canada, France, Germany, Greece, South Africa, Iberia, Italy, Mexico, Portugal, Switzerland (2 separate sites), and The UK | Terminated | January 16, 2008 | 265 | 6037
Z-0519-2008 | 45913 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Medical Systems Information Technologies Clinical Information Center (CIC) Pro v5.0.3 &amp; v5.0.6 systems, GE Healthcare, Clinical Systems |  |  | November 08, 2007 | 2007 | General Electric Medical Systems Information Technology | Potential for Loss of Audible Alarming and ECG Tab Setting Changes; from Clinical Informaiton Center (CIC) Pro v5.0.3 &amp; 5.0.6 when used in combination with Telemetry and bedside monitors. | Software design | A "GE Urgent Medical Device Correction" letter was sent to consignees on 10/16/2007. The letter described the problem and product involved. The letter also gave short term recommendations. A long term solution is being worked on by GE and will be announced when determined. |  | 3116 | 3, 116 | Worldwide-USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MN, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, PR, and District of Columbia and countries of The United Kingdom, Turkey, Thailand, Tsiwan, Syria , Switzerland, Sweden, Spain, Slovenia, Singapore, Saudi Arabia, Russia, Poland, Qatar, Pakistan, Philippines, Palestinian Territory, New Zealand, Norway, Netherlands, Nigerie, Malaysia, Lebanon, Jordan, Italy, Israel , Ireland, India, Honk Kong, Greece, Germany, France, Finland, Egypt, Denmark, Czech Republic, Cyprus, Costa Rica, Colombia,China, Chile, Canada, Delgium, Bahamas, Austria, and Australia. | Terminated | December 10, 2011 | 1493 | 3116
Z-0733-2008 | 45906 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Hyperbaric Chamber Gurney, Standard Gurney &amp; Stretcher Model, P/N : 21502, Sechrist Industries, Inc., Anaheim, CA |  |  | October 31, 2007 | 2007 | Sechrist Industries Inc | Brake Problems: The stretcher brake cams may crack while attempting to lock the stretcher brake (Sechrist is supplied the stretchers by the manufacturer Stryker Medical) | Device Design | Urgent Medical Device Correction letters were sent to customers beginning on November 9, 2007, via FedEx &amp; DHL overnight. Customers were notified this action was being taken to providing notification to their customers as a continued implementation of the Stryker Medical field correction.The Notification letter instructed their customers to:1. Locate the gurneys with the serial numbers listed in this advisory.2. If any of the affected gurneys cannot be located, determine if they have been sold, transferred, or donated to another facility and notify us regarding their new location.3. Verify the brakes are working properly. The brake function may be checked by engaging the brake pedal and pushing on each end of the gurney to make sure that casters are locked securely.4. Remove the gurneys with non-functioning brakes from service until brake cams have been replaced and brakes are fully functional.5. Contract Sechrist for the replacement parts package.6. Complete the retrofit of the Gurney.7. Verify correct function of the brake assembly following retrofit.8. Complete and send the Correction Verification Form contained in the package to Sechrist. |  | 559 | 559 Units Total (483 in US) for all models | Worldwide-USA, Brazil, Bulgaria, Canada, Dubai, U.A.E., Estonia, India, Indonesia, Japan, Malaysia, Saudi Arabia, Sweden, Taiwan, Thailand, and Vietnam. | Terminated | April 19, 2012 | 1632 | 559
Z-0664-2008 | 45413 | No | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Tucker Embryo Catheter with URM Tip, Embryo Transfer Catheter, 20 cm, Order #: ftr120, Manufactured by: Fertility Technology Resources, Inc., Marietta, GA 30062 |  |  | September 13, 2007 | 2007 | Fertility Technology Resources, Inc. | Expiration Date for sterility; The 5 year dating could not be validated/supported for sterility assurance. | Process design | All customers were contacted by telephone starting 9/13/2007. They were asked to return all Tucker family of catheters and all Morton IUI catheters that are on the shelves. A UPS call tag was sent to each customer to facilitate the return of the products. A follow up Urgent Device Recall letter was sent to all customers on 10/9/2007 reminding them to return all outstanding inventory when the UPS call tag arrives. |  |  | not reported | Worldwide: USA including states of CA, CO, FL, GA, LA, MD, ME, MI, MS, NC, NH, NJ, NY, OH, TN, and WA and The Netherlands. | Terminated | February 06, 2008 | 146
Z-0120-2008 | 44937 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AmniSure ROM Test, Fetal Membrane Rupture Test for |  |  | October 03, 2007 | 2007 | AmniSure International LLC | Weak true postive/false negative results: weak true positive or false negatives in ruptured fetal membranes may occur due to a "hook effect" | Device Design | AmniSure notified users by Medical Device Correction (Recall) letter dated October 1, 2007 issued via US Mail. The letter provides additional labeling instructions and a checklist for trouble shooting, including a diluation procedure that allows users to determine and solve the problem of the hook effect with this device. |  | 319976 | 319,976 units | Worldwide, USA, Australia, Austria, Italy, Korea, Bulgaria, Russia, Spain, Kuwait, Switzerland, Luexemburg, Ireland, Denmark, Netherlands, Jordan, Israel, Chile, UK, Singapore, Basil, New Zealand, Turkey, France, Greece, and Equador | Terminated | February 06, 2008 | 126 | 319976
Z-0529-2008 | 45482 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Oxy Serve II Oxygen Conserving Regulator, Model Number: MCRB-870-6, Superior Products, Inc/SP Medical, LLC, Cleveland, Ohio 44144 |  |  | October 19, 2007 | 2007 | Superior Products Inc | Leakage: Device was leaking oxygen | Component design/selection | A Voluntary Medical Device Recall notification was delivered by certified mail on or about 10/19/2007, to all consignee's; "bill to" and "sent to." Consignee's are asked to contact their customers (if any), notify them of the recall, and return all implicated products to Superior Products. |  | 125 | 125 | Worldwide-USA including states of CA, FL, IA, IN, KY, OH, SC, TX, and WA. and countries of Singapore, Mexico City, MX, and Turkey | Terminated | April 25, 2012 | 1650 | 125
Z-0365-2008 | 39353 | No | LZS |  |  | Ophthalmic | PMA | Advanced Medical Optics (AMO), VISX WaveScan WaveFront System (software). Part Number 0070-1478. Advanced Medical Optics, Santa Clara, CA 95051. |  |  | May 30, 2007 | 2007 | Visx Inc | Erroneous Treatment Calculations - Two software caused errors; in combination, in the WaveScan System Software, will result in an erroneous treatment calculation (overcorrection) in patients. (1) Installation of software for a certain brand of computer (Dedicated Computing LLC) can cause a software registry setting for an algorithm that identifies image reflection to be erroneously set in the "off | Software design | AMO sent an Action Required notification on May 30, 2007, to AMO Certified Physicians. The letter included a CD software update (0070-1721) to be installed by users that provided correct registry parameter settings. Upon completion of installation, a confirmation form will print out. Customers are requested to fax back confirmation to AMO (408-773-7055). |  | 26068 | 196 systems are affected in the U.S. and 133 outside of the U.S. | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO,MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK,OR,PA, SC, TN, TX, UT, VA, WA, WI, and WV and countries of Canada and Mexico. | Terminated | April 17, 2008 | 323 | 26068
Z-0183-2008 | 45441 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE OEC, Insta Trak 3500 Plus with the Cranial Application for Image Guided Surgery, GE Healthcare, Surgery, Salt Lake City, UT 84116 |  |  | October 18, 2007 | 2007 | GE OEC Medical Systems, Inc | Crack, Fracture or Detachment: The Transmitter Arm for the Mayfield-compatible Transmitter may become detached from the unit and enter into the surgical field. | Device Design | Consignees were sent an Urgent Recall Notice via mail dated 10/11/07, notifying the customer to examine the plastic end of the transmitter arm and to not use the transmitter arm, if any cracks or damage are observed. Also to stop using Steam Sterilization method and to utilize the other Sterilization methods indicated in the Operators Manual. The root cause is still under investigation GEHC Surgery will provide, at no charge, a solution to this issue when it becomes available. |  | 169 | 169 units | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, MA, MD, MI, MS, NC, NY, OK, OR, PA, SC, TX, UT, VA, WI, and WV, and countries of Canada, China, France, Hong Kong, Japan, Mexico, Spain, and Switzerland. | Terminated | August 10, 2012 | 1758 | 169
Z-0481-2008 | 45528 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | StelKast Proven Knee - Tibial Half Block Augmentation and Screw. Part Number SC2287-3-5, Size 3, Thickness 5 mm, StelKast, Murray, PA 15317 |  |  | October 05, 2007 | 2007 | Stelkast Co | The screws may fracture | Device Design | The recalling firm telephoned their customers on 10/5/07 to inform them of the problem and the need to return the product. As a follow up to the telephone calls, the recalling firm issued a recall letter (Field Notification) on 10/9/07 via certified mail. |  | 41 | 41 units | Nationwide: shipped to medical facilities in IL, NC, PA, TN, TX, VA, and WV. | Terminated | January 23, 2008 | 110 | 207
Z-0588-2008 | 45919 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day INFUSOR Portable Elastomeric Infusion System, Product Code: 2C1071KJP; 2mL/hr with Tubing Flow Restrictor; 12 units per case; Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | November 29, 2007 | 2007 | Baxter Healthcare Corp. | Leaking: leaks at the tubing flow restrictor connection to either the coupler or the male luer during filling. | Process control | Baxter on 11/30/07 sent an Urgent Product Recall letter to affected customers, to the attention of the Director of Nursing with copies to the Materials Manager, via first class mail to advise them of the potential leaks during filling or use of the infusor. The accounts were instructed to immediately discontinue use and segregate the listed affected lot numbers, and contact Baxter Center for Service at 1-888-229-0001 to arrange for the return and credit for the affected product. Any clinical/technical questions were directed to the Baxter Healthcare Product Information Center at 1-800-933-0303. |  | 86192 | 86,192 units | Worldwide: USA, Canada, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Italy, Portugal, San Marino, Spain, Sweden, United Kingdom, South Africa, Colombia, Argentina, Hong Kong, Australia and New Zealand. | Terminated | April 29, 2010 | 882 | 86192
Z-0589-2008 | 45965 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Push Button Blood collection Sets with Pre-Attached Holder;Catalog Number : 367352: 21G x 3/4" x 12"; 0.8 x 19mm x 305mm, Wingset PBBCS, BD Diagnostics, Franklin Lakes, NJ |  |  | November 20, 2007 | 2007 | Becton Dickinson &amp; Company | Safety Mechanism Failure: the safety mechanism of the BD Vacutainer Push Button Collection Set with Pre-attached holder cannot be activated; thus the needle cannot retract into the rear barrel. | Material/Component Contamination | BD sent Recall notification letters, Urgent Class II Product Recall, to customers and distributors on 12/10/07 through a third party, West Inc., Appleton, Wi. Mailings were sent return receipt requested. Letters requested immediate inventory search and product return. Contact number for questions was given as 1-800-631-0174 |  | 105600 | 105,200 USA; 400 Canada | Worldwide-USA and Canada | Terminated | September 08, 2008 | 293 | 876100
Z-0591-2008 | 46149 | No | NWX |  |  | Cardiovascular | PMA | AngioSculpt PTCA Scoring Balloon Catheter, Catalog (REF) Number: 2027-2510, size 2.5mm x 10 mm, AngioScore, Inc., Fremont, CA. |  |  | November 13, 2007 | 2007 | AngioScore Inc. | Mis-labeling: Product labeling on product pouch indicates wrong size, the correct size is on the product carton | Process control | The firm initiated a field correction and issued Medical Device Field correction letters to its consignees on 11/09/2007, instructing consignees to review their inventory and return any mislabeled product or contact the firm. |  | 40 | 40 units | Nationwide-to 30 medical facilities in NC, SC, OH, IL, NY, PA, NJ, PA, LA, KS, FL, WI, MI, WA, and CA. | Terminated | March 25, 2008 | 133 | 40
Z-0115-2008 | 45855 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Straight Magnum Inserter Handle, a preparatory instrument used for acetabular component insertion; REF # or Catalog Number: S313141. (systems affected by this action include M2a-Magnum, Regenerex, Acetabular Revision Shell and M2a-38), Biomet Orthopedics, Inc., Warsaw, IN. 46581 |  |  | September 04, 2007 | 2007 | Biomet, Inc. | Binding: The instrument was manufactured incorrectly and its use may result in binding to the acetabular component and cause the implant to not release following impaction. | Other | Consignees were notified by Urgent Medical Device Correction Notice (letter) dated 9/4/07 to discontinue use of the instruments and to return them for evaluation. |  | 22 | 22 | Nationwide; including states of Alabama, Arizona, Arkansas, Georgia, Illinois, Indiana, Kentucky, Massachusetts, Michigan, Nevada, North Carolina, Ohio and South Dakota. | Terminated | September 22, 2009 | 749 | 22
Z-0532-2008 | 37989 | No | KRL | Cardiopulmonary bypass bubble detector. | Cardiovascular | Cardiovascular | 510(k) | Terumo Ultrasonic air sensor, 1/4" by 3/32" (6.4 mm x 2.4 mm), Gold, for use on Terumo perfusion systems 1, 8000 and 9000; Catalog No: 5791, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 19, 2007 | 2007 | Terumo Cardiovascular Systems Corp | False Alarm: The air sensor may malfunction and trigger false alarms, and may continue to alarm thus preventing the device from being reset. | Process control | Consignees were notified of the defect/issue and what to do if it were to occur (short term work around) via an Urgent Medical Device Removal letter dated 9/14/07. As the permanent fix users were informed the sensor will be replaced when replacements become available. |  | 175 | 175 | Worldwide: USA, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Costa Rica, Dominican Republic, Equador, Germany, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Latin America, Malaysia, Mexico, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, Uruguay, Venezuela and Vietnam. | Terminated | July 07, 2009 | 657 | 2826
Z-0258-2008 | 38312 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical EndoWrist, 8 mm, PK Dissecting Forceps for use with the da Vinci and da Vinci S Surgical systems, Part Numbers: 400214-02, 400214-03, 420214-02 and 420214-03, Intuitive Surgical, Inc., Sunnyvale, CA |  |  | April 13, 2007 | 2007 | Intuitive Surgical, Inc. | Mislabeling-electrical isolation requirements: devices were incorrectly labeled with a CF symbol not their proper BF Symbol on the instrument housing. (All instruments meeting the CF isolation requirements are suitable for direct cardiac application, while instruments meeting BF isolation requirements are not suitable for direct cardiac application) | Device Design | On January 9, 2008, via an Urgent Device Recall Letter the recalling firm notifed users of the defect (mislabeling with heart symbol) and to Immediately Discontinue any direct cardiac applications. The recalling firm will arrange with each consignee a time to relabel their devices. Customers on April 17, 2007 had earlier been warned of the device defect via letter but no correction was indicated. |  | 1136 | 1,136 units | Worldwide Distribution --- USA including states of MN, PA, AK, NY, CA, MD, IN, AZ, CO, AR, TX, TN, FL, VA, MA, NE, NC, IL, OH, ME, WI, MD, NJ, MN, MI, VA, IA, OR, MO, NC, MS, AL, WA, GA, KY, LA, AK, AR, SD, SC, DE, NV, HI, NH, WV, and CT, and countries of Canada, Australia, Belgium, Spain, Hong Kong, Italy, Switzerland, and Sweden. ( 192 USA hospitals, 3 foreign hospitals and 5 foreign distributors) | Terminated | August 01, 2008 | 476 | 1136
Z-0539-2008 | 45496 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark LT, Slide Staining System, Catalog Number: N750-BMKLT-FS, Ventana Medical Systems Inc, Tucson, AZ |  |  | September 29, 2006 | 2006 | Ventana Medical Systems Inc | Leaking Carboys/Safety Updates:The Ventana Staining Platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. | Device Design | The firm has mailed three customer letters, a) the first letter dated 9/30/2005 describes the upgraded Voila CD that contains safety information; b) a second letter, dated 6/27/2006, identified as a "Customer Alert" was mailed describing an upgrade to the software system version 9.3/10.3 and the CD that includes safety improvements, c) the third letter dated 9/29/2006 and identified as a Safety Alert advises that as part of their effort to continuously improve laboratory safety, they are updating the labeling and instructions for use regarding installation and cleaning for the carboy as the Ventana staining platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. Enclosed are pictogram labels that indicate "slip hazard" and "refer to instructions for use." The customer is instructed to apply these two labels to the carboys above the spigot as a reminder to laboratory personnel to keep the spigot in the full "off" position when not in use and to refer to the instructions for use for maintenance. In addition, the letter advises that there is an electrical shock warning for the power strip that is part of their system to indicate that it should be mounted off of the floor. This power strip is intended to connect the computer components of the staining system and should not be placed on the floor where fluid might contact it. Also included is a warning label for shock hazard to apply to their power strip. |  | 1712 | USA - 1,712 units; Foreign - 1,159 units (total for recalls-Z-0537-0541-2008) | Worldwide; USA, Australia, China, Canada, Germany, Great Britain, France, Italy, Hungary, Singapore, Korea, Taiwan, Spain, Poland, Greece, Austria, Sweden, Japan, Israel, Denmark, Turkey, Finland, Saudi Arabia, The Netherlands, Norway, Slovenia, Ireland, and Belgium | Terminated | June 06, 2013 | 2442 | 1712
Z-0127-2008 | 45395 | No | LOM |  |  | Microbiology | PMA | Abbott AxSYM AUSAB Reagent Pack, List No: 3C74-20, 100 tests, an in-vitro test for Antibody to Hepatitis B Surface Antigen (Anti-HBs); Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 01, 2007 | 2007 | Abbott Laboratories | Incorrect results: Reagents in these lots were manufactured with an incorrect ratio of 2 key components, which could result in both false reactive and false nonreactive specimens. | Process change control | Abbott sent a Product Recall letter, dated 10/01/07, to all AxSYM AUSAB customers that received the affected lots of product, informing them of the manufacturing error that resulted in both false reactive and false nonreactive specimens. The letter provided the assay bias for both lots. The accounts were requested to examine their inventory for the affected lots; destroy any packs found; record the number of packs destroyed on the Customer Reply Form; and fax the completed form back to Abbott at 1-800-777-0051 prior to 10/12/07. They were also requested to evaluate patient results generated with these lots in alignment with the information provided in the letter to determine if additional action is necessary. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the lots; and to retain a copy of the letter for their records. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT. |  | 1722 | 1,722 packs | Nationwide Distribution, including Puerto Rico. | Terminated | May 09, 2008 | 221 | 1722
Z-0542-2008 | 45510 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VNUS 7F ClosureFAST Catheter, Order #: CF7-7-60 (60 cm working length), Manufactured by VNUS Medical Technologies, Inc., 5799 Fontanoso Way, San Jose, CA 95138 |  |  | August 08, 2007 | 2007 | VNUS Medical Technologies Inc | Instructions for Use change: Instructions for Use have been changed to include a modified placement recommendation. | Labeling design | The Recalling firm is updating the Instructions For Use, training materials, and procedural guides to incorporate the treatment recommendations. The recalling firm sent an Update to customers who received the product, Important Update Regarding the Use of the VNUS ClosureFAST Catheter, dated August 6, 2007 with supplemental device instructions. Future product distribution will include the updated labeling. |  | 16716 | 16,716 units | Worldwide | Terminated | May 22, 2008 | 288 | 22937
Z-0569-2008 | 45951 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Stainless Steel Permanent Molar Crowns , Product Code: 900321, Made in U.S.A by 3M ESPE Dental Products St. Paul, MN 55144-1000. |  |  | October 17, 2007 | 2007 | 3M Company / Medical Division | Mislabeled: Unitek Stainless Steel Permanent Molar Crowns were placed in packaging which incorrectly identified the product as Unitek Stainless Steel Primary Molar Crowns. | Packaging process control | 3M ESPE Dental Products initially contacted Distributors and Doctors on 10/22/07 through telephone calls, and sent a Device Recall letter via Fed Ex - Overnight on 10/22/07.The letter informed the users that Unitek Stainless Steel Permanent Crowns Lot # P060928, Product Code 900321 may be mislabeled as Unitek Stainless Steel Primary Molar Crowns. Users were asked to stop using and/or distributing the product and contact ESPE Dental Products at 1-800-237-1650 to initiate product return and replacement. |  | 975 | 975 | Worldwide Distribution USA including states of TX, IN, FL, NV, KY, TN, MN, WA, NJ, PA, IL, CA, and IA, and country of CANADA. | Terminated | October 17, 2008 | 366 | 975
Z-0592-2008 | 46154 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Hoana LifeBed Patient Vigilance System, (Intelligent Medical Vigilance System),Model Number: 010135-01, Product distributed by Hoana Medical, Inc., 828 Fort St. Mall, Suite 620, |  |  | November 07, 2007 | 2007 | Hoana Medical | Bed Exit Failure: Failure to recognize "Bed Exit" when the feature is active due to a software anomaly. | Software design | On 11/7/07, all consignees were notified via an untitled letter and on-site visit by the firm's sales representative, informing them of the affected product and providing instructions on the recall. The recalling firm will arrange for installation of a software fix at each consignee. |  | 216 | 216 units | Nationwide including states of HI, CA, TX and MA | Terminated | February 01, 2008 | 86 | 216
Z-0675-2008 | 46250 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf, interactive wound and burn dressing, Organogenesis, Inc., Canton, Massachusetts 02021 |  |  | December 26, 2007 | 2007 | Organogenesis, Inc. | Contamination: Apligraf units were reported to have contamination in the agarose nutrient medium of some units retained at Organogenesis. Preliminary tests show a gram negative rod (bacillus) organism. | Material/Component Contamination | On 12/26/07, Organogenesis began to telephone the consignees to inform them of the recall. The firm sent a follow-up fax on 12/31/07 to all treating physicians who applied the product requesting patient status information. The Dear Doctor letter, dated December 27, 2007, was titled, "Subject: Potential contaminated Apligraf Units manufactured by Organogenesis, Inc." Consignees were instructed to return unused product, and monitor patients who received the Apligraf for adverse events. |  | 177 | 177 units | Nationwide Distribution. | Terminated | May 15, 2015 | 2697 | 177
Z-0435-2008 | 45478 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Depuy TK2 Compression Hip Screw trauma plate, Short Barrel, 135 DEG, 4H-92.6mm, sterile; REF 8315-35-104, DePuy Orthopaedics, Inc., Warsaw, IN 46581 |  |  | October 30, 2007 | 2007 | Depuy Orthopaedics, Inc. | Mis-etched/labeled as to barrel length: the lot with the short barrel was etched and labeled as a standard barrel and the lot with the standard barrel was etched and labeled as a short barrel. | Process control | Consignees were notified via Urgent Information, Recall Notice(letter) dated 10/30/07 to determine whether they have product on hand and to contact their sales representative to return the product. Dear Doctor letters will be issued to implanting physicians as they are identified. |  | 4 | 4 | Nationwide including the states of Florida, Massachusetts, Ohio, Pennsylvania, South Carolina, Tennessee and Virginia. | Terminated | April 21, 2008 | 174 | 13
Z-0546-2008 | 45809 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | ProTime Microcoagulation System Instrument, Catalogue Number: PROTIMEPRO, International Technidyne Corporation (ITC), Piscataway, NJ 08854 |  |  | November 19, 2007 | 2007 | International Technidyne Corporation | An increase in the frequency of non-conformance reports (NCMR) for certain displayed error messages. | Device Design | ITC sent to distributors, self test users and health care professionals on 11/19/07, an Urgent Medical Device Recall Letter requesting return of the devices. |  | 3482 | 3,482 units | Worldwide-USA and countries of China, Taiwan, Greece, Ireland, Chile, Finland, Italy, Spain, Norway, Germany, Switzerland., Singapore, Korea, Indonesia, Lichtenstein, Israel, Saudi Arabia, Latvia, Kuwait, and Slovenia. | Terminated | August 12, 2008 | 267 | 5910
Z-0549-2008 | 45871 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | ACAT 1 Series Intra-Aortic Balloon Pumps, Catalog # IAP-0100, Arrow International, Everett, MA 02149 |  |  | November 12, 2007 | 2007 | Arrow International Inc | Leak may occur in the helium drive system for the IABP | Material/Component Contamination | Arrow International notified customers by letter dated November, 2007 titled: Urgent Medical Device Field Correction. Customers are instructed to follow the Recommended Action included in the letter while awaiting replacement of the defective component to be performed by service representative. |  | 47 | 47 units | Worldwide- USA, including the states of CA, CO, FL, IN, KS, MA, MO, NV, NY, PA, SC, TN, TX, and WA and countries of Canada, Africa, Argentina, Austria, Bangladesh, Belarus, Belgium, Brazil, Canada, Columbia, Czech Rep, UK, Finland, France, Germany, Hungary, India, Indonesia, Italy, Japan, Korea, Mexico, Netherlands, New Zealand, Panama, Philippines, Portugal, Russia, Saudi Arabia, Italy, Spain, Sweden, Taiwan, and Turkey. | Terminated | January 11, 2010 | 791 | 272
Z-0551-2008 | 45873 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott, ACCELERATOR Device Manager (ADM); List Number: 08H74-01; Abbott Laboratories, Abbott Park, IL |  |  | October 25, 2007 | 2007 | Abbott Laboratories, Inc | Sample/Patient Mis-identification: Two issues identified related to the ACCELERATOR Decision Manager (ADM), to include: 1) Sample Identification (SID) and/or Patient Identification (PID) numbers that contain more than 12 characters are truncated to the final 12 characters by the ADM; 2) for SIDs using alpha characters that are case sensitive, SIDs with lower case characters are transmitted by the | Software design | Firm notified consignees via Product Correction, Immediate Action Required letter(s) on 10/25/07. Consignees were notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade. |  | 19 | 19 units. | Nationwide:Distributed to hospitals/laboratories in IL, PA, TN, TX, UT, VA and WV. | Terminated | April 01, 2008 | 159 | 19
Z-0552-2008 | 45888 | No | IWX | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | AADCO Rayshield Frame Mounted Overhead X-Ray Barrier; |  |  | November 01, 2007 | 2007 | Aadco Medical Inc. | Screws may become loose from the mounted shield compromising stability | Device Design | AADCO Medical Inc., issued a Product Safety Notification/recall letter to Hospital Administrators, Managers of Radiology/Cardiology and Radiologists and Cardiologists dated 11/1/07. AADCO will perform on site inspections with device tightening or replacement of parts as needed; at no cost the the users |  | 324 | 324 units | Worldwide; USA, Canada, Mexico, Brazil, Venezula, Chile, Bermuda, Spain, and Panama | Terminated | May 02, 2012 | 1644 | 324
Z-0553-2008 | 45891 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Symphony Staining System (High Volume Stainer), KPA, Model Number 2200000, Ventana Medical Systems Inc |  |  | October 30, 2007 | 2007 | Ventana Medical Systems Inc | Fluid Leaks-The T fitting located in the Symphony Clear system of the Symphony High Volume Slide Stainer needs to be larger to securely hold the tubing in place to eliminate any possible fluid leak due to the tubing movement. | Component design/selection | The end user customers were verbally contacted starting in November 2007 to set up a scheduled visit to replace the fitting. |  | 55 | 55 | Worldwide-USA, France and Japan | Terminated | April 14, 2008 | 167 | 55
Z-0477-2008 | 45485 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Blood Collection Reservoir, 40 micron,Catalog/Model # EL240, Medtronic, Inc., Minneapolis, MN 55432-5804 |  |  | October 01, 2007 | 2007 | Medtronic Perfusion Systems | Unnaproved Component Resin: Affinity NT Integrated CVR Membrane Oxygenators and EL 240 Autotransfusion Blood Collection Reservoirs contain components made from an unapproved resin | Nonconforming Material/Component | Medtronic sales force was notified and initiated calls with cutomers on 10/1/07. Additionally, certified Urgent Medical Device Recall letters dated 10/5/2007, were sent to the 32 customers. The letter explained the situaiton and requested that the impacted product be returned for replacement or credit. Medtronic is assessing patient use risks and will report those when determined. Standard post-surgical monitoring is advised at the present by Medtronic . |  | 413 | 413 | Nationwide including the states of IL, CT, OK, OH, MN, FL, NY, KS, MO, LA, PA, IN, CA, MI, KY, SC, NE, and AR. | Terminated | October 23, 2008 | 388 | 447
Z-0522-2008 | 45819 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | life+cel, Replacement battery for Cardiac Science FirstSave Survivalink; OEM P/N:9141-001, Part #: 2L561, Lithium Sulfur Dioxide, Voltage 12V, 7.5 amps., Mfg for Exocomm Technology Group; Mfg by Nordix, Inc. |  |  | September 03, 2007 | 2007 | Exocomm Tech Grp | Marketed without a 510K | PMA | The original recall communication was sent out on 9/3/07 by Phoenix Solutions, a consultant to Battery Zone, as a Dear Valued Customer notice. A second recall communication, Urgent Device Recall Letter, was issued by Exxocom Technology Group dated 10/25, was sent out to all customer (even those who had returned product) on 11/5/07. Both communication were sent by first class mail, requesting return of batteries |  | 266 | 266 units | Worldwide-Batteries were distributed to dealers or distributors in CT, GA, MN, NC, NY, OH, OK, OR, PA and TN plus one Canadian Consignee | Terminated | July 01, 2013 | 2128 | 312
Z-0527-2008 | 45856 | Yes | KPF | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter RenalSoft v.2.0 Patient Management Software Suite - HD Module; clinical data management software; Baxter Healthcare Corporation, McGaw Park, Illinois 60085 |  |  | November 14, 2007 | 2007 | Baxter Healthcare Renal Div | Pre-Treatment/Post-Treatment Report errors: 1) The Heparin Bolus Value is displayed on Pre and Post Treatment Reports for patients with "No Heparin" orders (RenalSoft v2.0, Renal Software Suite v.3.0 and v.3.1), and 2) Withheld Medications are incorrectly displayed as Administered on the Pre-Treatment report (RenalSoft v.2.0 only). | Software design | Baxter Healthcare Corporation sent an Urgent Device Correction letter, dated 11/14/07, to all of the affected customers, informing them of the Heparin Bolus and the Withheld Medication issues and provided them with interim mitigation instructions that they should implement while waiting for a software correction. A Baxter representative will contact the customer to schedule the software update when the upgraded software is released. Customers were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff, and indicating if they still have RenalSoft software and require an upgrade. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, option 4.A follow-up letter dated 3/6/08 was sent to the consignees on the same date, informing them that the upgraded RenalSoft Software v.2.0.1 was now available for installation. The consignees were requested to complete and fax back the customer reply form indicating that they received the letter, and were informed that they would be contacted by a Baxter representative to schedule an installation of the software upgrade at no charge. |  | 25 | 25 units | Worldwide Distribution-USA including Puerto Rico, and countries of Australia, Austria, Canada, Chile, Denmark, Finland, Greece, Ireland, Italy, Mexico, The Netherlands, New Zealand, Norway, Panama, Portugal, Spain, Sweden and the United Kingdom | Terminated | February 03, 2010 | 812 | 34
Z-0478-2008 | 45513 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Hallmark Anterior Cervical Plate System: Primary Bone Screw, Semi Constrained, REF # 62-3140,14mm; Blackstone Medical, 90 Brookdale Drive, Springfield, MA 01104 (The Hallmark Anterior Cervical Plate Systems consists of an assortment of implantable titanium alloy plates and bone screws that are provided non-sterile) |  |  | May 08, 2006 | 2006 | Blackstone Medical, Inc. | Incomplete Design Validation: the laser marking on the head of the screw lacked complete process qualification | Device Design | On May 8, 2006, a communication to the field was sent via phone, e-mail and fax. Instruction was made to retrieve and return all product in the field. |  | 1991 | 1,991 semi-constrained primary and rescue screws. | Nationwide | Terminated | November 06, 2008 | 913 | 1991
Z-0554-2008 | 45519 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Real-Time Position Management System (RPM), A Respiratory Gating System; Version 1.7.3, All Models; Varian Medical Systems, Palo Alto, CA 94304 |  |  | September 26, 2007 | 2007 | Varian Medical Systems Inc | Erroneously Resetting: RPM System version 1.7.3, software failure can occur that affects the gated radiotherapy treatment when phase-based gating is used. RPM 1.7 erroneously resets the phase-based gating thresholds to system default during the treatment session, if the user accesses the Session options during the treatment session. This can lead to difference between the planned and delivered | Software design | Varian sent an Urgent Medical Device Correction letter on September 26, 2007, to all its direct consignees, informing them that the RPM version 1.7 may reset the phase thresholds if Sessions options are accessed during treatment. A Service Representative will be scheduling a visit to user site to perform the modification requested to address the problem, when a corrected software version is available. |  | 59 | 59 units | Worldwide including countries of: Ireland, Italy, Spain, UK, Denmark, Finland, France, and Canada | Terminated | September 30, 2009 | 735 | 59
Z-0565-2008 | 45869 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ethicon Endopath 5 mm Curved Dissector with Monopolar Cautery, Product Code: 5DCD, Ethicon Endo-Surgery, Inc, Cincinatti, OH 45242 |  |  | November 06, 2007 | 2007 | Ethicon Endo-Surgery Inc | Sterility (package integrity) compromised; Damaged packaging may have compromised the sterility of the devices. | Process control | Endopath sent consignees an Urgent Device Recall notification letter on or about 11/15/07, informing them that the possibility exists that damage to the packaging of this product may have compromised the sterility of the product. The firms were told to examine their inventory, and remove all affected product. All consignees were instructed to complete the Business Reply Form; even if they did not have any product. Recalled products are to be returned to Ethicon for a credit. |  | 2844 | 2,844 | Worldwide, USA including states of AL, AR, AZ, CA, CO, DE, FL, GA, IL, KS, KY, LA, MA, MN, MO, MS, NC, NJ, NV, NY, OH, PA, RI, TN, TX, AND WA, and countries of Austria, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Greece, Hong Kong, Italy, Japan, Korea, Mexico, Portugal, Puerto Rico, Republic of Panama, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | December 09, 2009 | 764 | 2844
Z-0568-2008 | 45958 | No | LYJ |  |  | Neurology | PMA | Cyberonics, VNS Therapy System Model 250 Handheld programming system using Dell x5 handheld and preprogrammed with software v7.1 and higher; manufactured by Cyberonics, Inc., Houston, TX. |  |  | October 25, 2007 | 2007 | Cyberonics, Inc | Screen Freezes-- The Dell X5 Handheld PC screen will freeze caused due to incompatibility between the Microsoft 2002 OS and the model Dell X5 handheld computer. Once frozen, the handheld device becomes non-responsive to user input. | Software design | Cyberonics sent an initial notification (Important Safety Alert) to physicians via Certified Mail on 10/01/07. Letter notified physicians that the Dell X5 handheld computer screen will freeze following an interrogation. Once frozen the handheld device becomes non-responsive to user inputs. The firm informed the user that in the event that the computer becomes non-responsive, the user has two options, 1) Reinitialized the flash memory and, 2) Perform a software reset of the handheld computer. A second notification addressing the approval and addition of labeling was sent to physicians on 11/15/07. Physicians were asked to respond to both letters via an enclosed response card, acknowledging receipt of the notification. Additional information will be posted at www.VNStherapy.com. |  | 4201 | 4,201 (US) 199 (International) | Worldwide Distribution- USA and the countries of; Austria, Belgium, Denmark, France, Germany, Ile de la Reunion, Ireland, Israel, Italy, Japan, Marocco, Mexico, Netherlands, Ireland, Norway, Poland, Portugal, S. Korea, Slovenia, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 05, 2009 | 711 | 4201
Z-0708-2008 | 46142 | No | HCI | Aneurysm clip applier. | Neurology | Neurology | 510(k) | Caspar Titanium Aneurysm Clip Appliers, Catalog No/Part #: FT420T, Casper Applier F/Mini Ti-Clipsstr 146 mm, Aesculap, Inc., Center Valley, PA. 18034 |  |  | November 01, 2007 | 2007 | Aesculap Inc | Breakage: the inner rod of the applier may corrode and break | Device Design | The recalling firm issued an Import Recall Notice, on or about December 18, 2007, to all their customers to inform them of the problem and the need to return the devices . |  | 60 | 60 clips | Worldwide- USA hospitals in AZ, CA, CO, IL, KY, MI, MS, ND, NM, NY, OH, OK, PA, TX, VA, WA, and WI, to sales representatives in NC, OK, and MO and shipped to distributors in CA and Canada. | Terminated | March 28, 2008 | 148 | 144
Z-0124-2008 | 44824 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips PageWriter Touch Cardiograph, |  |  | September 12, 2007 | 2007 | Philips Medical Systems | Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. | Software change control | Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection. |  | 266 | 266 units | Worldwide Distribution - USA, Canada, Australia, Hungary, New Zealand, Nigeria, Norway, Qatar, Taiwan, and UAE. | Open, Classified |  |  | 266
Z-0828-2008 | 46182 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Sierra Pre-Filled, Heparin Lock Flush Solution USP, 100 units/ml in 0.9% Sodium Chloride 120/Case, Catalog #1003-02, NDC #64054-1003-02, 5 mL Fill in a 12mL Syringe, For Single Use only, Sterile Solution Pathway, For IV Flush, Sierra Pre-Filled Inc., 455 W. Depot St. Angier, NC 27501 |  |  | December 11, 2007 | 2007 | AM2 PAT INC | Lack of assurance of sterility by manufacturer | Process design | AM2PAT contacted their distributors by phone on or about January 17, 2008 and instructed them to stop distribution of their pre-filled syringes and to contact their customers to recall all AM2PAT manufactured lots and all sizes of prefilled, Heparin Lock Flush syringes and Normal Saline IV Flush syringes (B Braun and Sierra Brands) to the consumer level. A press release was issued on January 17, 2008. The firm had eariler on 12/20/2007 recalled a single lot of Heparin syringes, the January extension was to all lots/all sizes of its Heparin and Saline syringes. |  |  |  | Nationwide | Terminated | December 02, 2010 | 1087
Z-0659-2008 | 38315 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System 5 mm instrument cannula, Part #: 420011-01, Intuitive Surgical, Inc., Sunnvale, Ca. 94086 |  |  | March 09, 2007 | 2007 | Intuitive Surgical, Inc. | Devices may have a ridge on the side of the cannula, which has the potential to abrade instrument shafts and generate black particulate matter. | Nonconforming Material/Component | The recalling firm issued an Important Device Recall Notification on 3/9/2007, announcing the recall and exchange program for device replacement. A toll free number of 1-800-876-1310 was given in the recall notice for questions. |  | 11 | 11 units | Worldwide: 92 hospitals nationwide and abroad, as well as 6 distributors abroad, USA including the states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OR, PA, SD, TN, TX, VA, WA, WI, and WV and countries of Belgium, China, France, Germany, India, Italy, Japan, Mexico, The Netherlands, Spain and Sweden. | Terminated | April 19, 2011 | 1502 | 896
Z-0462-2008 | 44934 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems, GEMINI GXL 6, 10, &amp; 16 slice PET/CT Systems, diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems, Philips Medical Systems, Cleveland, OH 44143 |  |  | September 21, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Artifacts: Artifacts may appear in some images from head scans on the CT subsystem of the GEMINI PET/CT system. Reading images with these artifacts may lead to misdiagnosis or improper treatment. | Software design | Philips Medical notified domestic customers via Product Safety Notification dated May 30, 2007. The product safety notification includes a Technical Bulletin notifying customers of the issue and how to avoid it. This information will be used until the software correction becomes available. The software correction will be released in October 2008. Once software is available, the firm will begin correcting the field. |  | 126 | 126 | Worldwide Distribution - USA, Australia, Belgium, Canada, China, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Jordan, Korea, Netherlands, Poland, Portugal, Saudi Arabia, Slovenia, Spain, Sweden, Taiwan, Turkey, and Venezuela. | Terminated | March 02, 2012 | 1624 | 187
Z-0660-2008 | 44965 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Spacelabls Medical, Ultraview Multigas Analyzer, Model 91518,Spacelabs Healthcare, Issaquah, WA, 98027 |  |  | October 08, 2007 | 2007 | Spacelabs Healthcare, Incorporated | Inaccurate partial pressure readings; readings are not compensated for ambient pressure in mmHg or kPa. | Software design | The firm sent out their Urgent Medical Device correction/recall letters on 10/5/2007, and an updated versions on 11/30/07 and 12/7/2007, forwarding an attached set of work around directions for use and associated tables for calculations. This short term fix will be updated to a permanent fix when a new software revision is released and installed at no cost to each user as noted in the recall notification. Telephone contact was provided at 1-800-522-7025. |  | 722 | 722 devices | Worldwide | Terminated | April 12, 2011 | 1282 | 722
Z-0690-2008 | 45845 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | November 15, 2007 | 2007 | GE Healthcare | Frozen Image Problems: During an acquisition (fluoro and/or record), there have been cases reported where an image became "frozen" on the DL (digital leader acquisition system) live monitor screen. In such cases, the system continued to send out X-rays without reporting an error message. The result was that the live imaging screen displayed an older "frozen" image until the operator released the | Device Design | Customers were contacted by means of an Urgent Medical Device Correction Letter on November 9, 2007. The letter described the product and the safety issue. The letter also made short term recommendations and described a long term solution, which is pending as an expected hardware and/or software modification that will be made at no cost. Short Term Recommendation: Take extra care when working with motionless anatomy or devices where there is little evident motion to warn you of the frozen image occurrence. Always completely shut down and restart the system (not just reset) immediately after such an error occurs. Also, please immediately report this issue to your local GE Healthcare field service engineer. |  | 378 | 378 | Worldwide-USA including states of AR, CA, CO, FL, GA, IL, IN, LA, MA, MD, MO, NC, NJ, NM, NV, SC, TN, TX, WA, WI, AK, AL, AZ, ,CT, DE, HI, ID, KS, KY, ME, MI, MN, MS, MT, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SD, UT, VA, WV, and Puerto Rico and countries of Hungary, Korea, Algeria, Argentina, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Finland, France, Georgia, Germany, Great Britain, Greece, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Morocco, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Singapore, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Turkey, and Venezuela. | Terminated | January 29, 2016 | 2997 | 869
Z-0520-2008 | 45820 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Images ViTALConnect , Version 4.0 and 4.1, PET/CT Visualization option. Vital Images, Inc., 5850 Opus Parkway, Suite 300, Minnetonka, MN 55343. ( A web-based medical diagnostic tool that allows physicians to use PCs or notebook computers to gain remote access to 2D, 3D, and 4D advanced visualization. The software enables users to measure, rotate, and analyze images) |  |  | November 08, 2007 | 2007 | Vital Images, Inc. | Standard Uptake Values (SUV) display/update problems; in the way Standard Uptake Values (SUV) are displayed in 3D/MPR views. The problem arises when a user draws a selected region of interest in order to calculate the SUV statistics of the region. If the user then changes the slice shown by scrolling through additional slices, the original SUV measurements will remain on the screen and do not up | Process control | Customers were notified via telephone (11/7 -11/8/2007) and issued an Urgent Medical Device recall notification letter on 11/8/07. The letter described the problem and product involved. The letter also informed customers they they will provide a free software patch within the next 6-8 weeks. |  | 18 | 18 | Worldwide-USA including states of SC, MO, MI, WI, NY, FL, CO, CA, PA, and WA, and countries of England and The Netherlands | Terminated | December 17, 2011 | 1500 | 18
Z-0285-2008 | 45495 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Si emans Evansi, with patient table, Part No: 44 81 586, and without patient table Part No: 44-81-669; Siemens Medical Solutions USA, Concord, CA 94520 |  |  | February 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Loose Mounting Screws -- Cassette holder may become loose and result in the device dislodging and colliding with the patient. | Labeling design | Customers were sent a Customer Safety Advisory Notice on February 27, 2007, notifying them that the mounting screws on the cassette holders on the SIMVIEW 3000 and NT Simulators, as well as some MEVASIM Simulators, could loosen and can cause damage to the system, or personal injury. The firm requested that they check of tightness on the mounting hardware. A STOP Modification notice was issued on 06/01/2007, to all affected regional service organizations that are responsible for implementing updates due to new information. Updated letters and User instructions will be released in August 2007. |  | 287 | 287 units total, all recalls Z-0281-285-2008 | Worldwide Distribution - USA including states of IN, VA, MA, VT, NJ, NY, PA, DC, NC, SC, GA, AL, TN, MI, KY, OH, IN MI, MN, ND, MT, IL, MO, KS, NE, AR, OK, TX, AZ, CA, SD, LA, Puerto Rico, and countries of Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Brazil, Bulgaria, Canada, Colombia, Costa Rica, Croatia, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Greece, Hungary, India, Iran, Ireland, Italy, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, China, Israel/Palestine, Philippines, Poland, Portugal, Qatar, Republic of Korea, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, UK, and Vietnam. | Terminated | February 03, 2011 | 1437 | 287
Z-0695-2008 | 45889 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Smiths Medical, Add-On Kids Kit , Blood Sampling System, Product Code: MX4037, Smiths Medical ASD, Inc, Dublin, Ohio 43016 |  |  | November 21, 2007 | 2007 | Smiths Medical Asd Inc | Misbranded; The product codes are labeled incorrectly with a green promotional label for "New split septum sampling site. Only use needle-free access devices .......", however this product code contains a sampling site which is accessed with a shrouded needle, NOT with a needle-free access device | Packaging process control | Smiths Medical mailed certified letters (Urgent Product Advisory Notifications) dated November 19, 2007, to all customers that received the affected lot. Firms are instructed to make the field correction (extra label removal) or mail the product back to Smiths for the correction. |  | 1440 | 1440 units | Nationwide including the states of CA, DC, IL, IN, MD, MN, NC, NE, SC, and TN and Puerto Rico. | Terminated | September 30, 2010 | 1044 | 1690
Z-0891-2008 | 45857 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Argon Medical Devices, ArgoGuide Hydrophilic NiTi Guidewire, K011905, Catalog Number: 00-AG387324, .038/300cm, stiff/angled/standard, sold as a single hydrophilic guidewire in both sterile and bulk non-sterile packaging configurations. Argon Medical Devices, Inc. Athens, TX 75751 |  |  | November 19, 2007 | 2007 | Argon Medical Devices, Inc | Degradation -- Hydrophilic guidewires may exhibit degradation of blue Pebax cladding causing particulate to dislodge potentially resulting in embolism. | Process design | Argon Medical sent an Urgent Product Recall letter to consignees on 11/15/2007. They requested that users immediately quarantine and return any product remaining at their facility, utilizing instruction on response card. Distributors were directed to issue the recall to their customers and return all retrieved product to Argon. |  | 13 | 13 | Worldwide Distribution -- USA including states of FL, KS, NV, IL, CA, NY, CO, VA, MD, NC, UT, GA, MI, NJ, TN, MA, and CT, and countries of The Netherlands, Germany, and Turkey. | Terminated | June 05, 2008 | 199 | 2951
Z-0935-2008 | 45903 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Impulse Angiographic Catheter, Sterile EO. Catheters are single lumen catheters with a single proximal hub for attachment to a dye delivery system. Catalog No. 16599-301. 6F MULTI PIT 124 (Single Pack). Distal Curve FL4, FR4, PIG 145. 100 and 110 cm Length. Made in Mexico: Parque industrial Le Mesa, Tijuana, Baja California 22050 MEX. Boston Scientific Corporation, Natick, MA 01760-1537 USA. |  |  | November 20, 2007 | 2007 | Boston Scientific | Flash may be protruding from the lumen of the catheter shaft. If this flash is present in the catheter, there is the potential of an embolization resulting from the flash detaching from the device during the procedure. | Release of Material/Component prior to receiving test results | The firm sent an Urgent Medical Device Recall letter dated 11/20/2007, to all affected hospitals. The letter describes the defect with the catheters and Product Table information including UPNs, Catalog Numbers, Lot/Batch Numbers and Product Descriptions. The firm requests that further distribution or use of any remaining product affected by the recall should cease immediately. ****IMPORTANT INFORMATION: If it is a practice of the hospital to remove product from the outer carton and store on shelves in the inner-pouch only. Users are to review the Product Table carefully and consider both the inner and outer packaging product codes when search for affected/recalled product.**** |  | 211 | 211 single units | Nationwide Distribution including states of AL, CA, FL, GA, KS, KY, LA, MN, NV, OH, OK, PA, SD, TN, TX, and WI. | Terminated | October 23, 2008 | 338 | 211
Z-0567-2008 | 45957 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Datascope CS 100 Intra-Aortic Balloon Pump and Datascope |  |  | November 19, 2007 | 2007 | Datascope Corp | Defective Circuit Board -- A defect in the Printed Circuit Board's (PCB) in the CS 100 and CS300 Intra-Aortic Balloon Pump may cause the pump to exhibit intermittent malfunctions, which include failure to start-up or reset of the display screen during therapy. | Other | Datascope sent a Product Field Correction notification letters to Hospital Administrators, via certified mail, on November 19, 2007. The letter informed them of a defect in the circuit board. It described the possible malfunction and provided instructions should malfunctioning occur. Datascope will be replacing the defective circuit boards in all affected IABP units. A Service Representative will be contacting the hospital to arrange for the replacement of the defective board. |  | 70 | 70 units US | Worldwide Distribution | Terminated | May 20, 2008 | 183 | 70
Z-0706-2008 | 46141 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet CC Cruciate Tibial Tray, 67 mm.; Product No. 141232, Biomet, Inc., Warsaw, IN 46581 |  |  | December 12, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size: Package labeled as containing a 67 mm tray actually contains a 75 mm tray and vice versa. | Packaging process control | Consignees were notified via an Urgent Medical Device Recall Notice dated 12/12/07 to cease use of the product and to return it. |  | 8 | 8 | Worldwide: USA including states of Georgia, Idaho, Indiana, Missouri, Ohio, and Virginia and country of Turkey. | Terminated | June 11, 2008 | 182 | 8
Z-0669-2008 | 46160 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand da Vinci S, 5 mm Instrument Cannula |  |  | November 19, 2007 | 2007 | Intuitive Surgical, Inc. | Sharp Edges: the 5mm Cannula may have a sharp edge on the inner diameter of the Cannula. The defective cannulae may cause particulate shavings to be skive (scraping) from the instrument shafts during surgery. | Other | On 11/19/07, the firm issued Urgent Device Recall Notification letters via Federal Express to all its direct consignees, informing them of the affected product and providing instructions on the recall. |  | 89 | 89 units | Worldwide: USA including states of AZ, CA, CT, DC, GA, lA, MI, MN, NC, NJ, NV, OH, PA, SD, TX, and VA and country of Hong Kong | Terminated | August 27, 2008 | 282 | 89
Z-0280-2008 | 45454 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center, Central Station, Software Version 4.1. (CIC Pro v4.1). GE Medical Systems Information Technologies, Inc., Milwaukee, Wisconsin (System hardware consists of a processing unit and a monitor) |  |  | January 24, 2007 | 2007 | General Electric Medical Systems Information Technology | Two separate issues may occur with operating The CiC Pro v4.1 System with Telemetry. 1) Experience a loss of audible alarms when operating system memory resource reaches approximately 170MG. This will occur if CIC Pro continuously runs for average of 160 day, and 2) System may continuously reboot upon start up in an attempt to correct the issue. | Software design | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07. As a Short Term Solution, GE recommends that user reboot of system according to instructions and precautions in Service Manual at least once every three months. Also, take the necessary precautions to ensure that system is shut down properly. This will decrease the likelihood of a continuous reboot. If system continues to reboot, GE suggest that user contact Technical Support at 1-800-588-7044 (Press Option 2) and ask for Technical Support. As for a Long Term Solution, GE will provide a software update that will rectify the issue as soon as an upgrade is available. |  | 6217 | 6,217 Units (1,549 OUS, 4,668 within US) | Worldwide Distribution - USA (excluding DE, HI, ID, MT, ND, SD, VT), Peurto Rico, Washington DC and including the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungaria, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherland, New Zealand, Norway, Pakistan, Philippines, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom | Terminated | May 16, 2012 | 1939 | 6217
Z-0699-2008 | 45966 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Assure 3 Blood Glucose Test Strips, For in vitro Diagnostic use. Product Number: 555050, Control Solution Range (mg/dL). Manufactured for Hypoguard, Minneapolis, MN 55439 USA. Made in Taiwan. |  |  | November 16, 2007 | 2007 | ARKRAY USA INC. | Mislabeling: incorrect control range-The Level 1 control solution range printed on the 50-count test strip bottles is incorrect. The range printed on the bottle is 62-54 when it should be 62-94. | Process design | On Friday, November 14, 2007 the recalling firm notified their distributors via telephone and email to inform them of the device defect and replaced all devices still in inventory. The distributors were also instructed to inform all of their customers of the issue and offer replacement product. Additionally, replacement will be offered to any end user that calls customer service about the issue. To reach ARKRAY use 1-800-818-8877 extension 3121 for replacement product. |  | 4920 | 4,920 | Nationwide including the states of CA, FL, GA, IL, IN, MA, MI, MS, NC, NJ, OH, PA, and TX | Terminated | October 30, 2008 | 349 | 4920
Z-0738-2008 | 46292 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optiflux 180NR Advanced Fresenius Polysulfone, Catalog No.: 0500318E, Fresenius Medical Care North America, Waltham, MA 02451. (Hemodialyzer) |  |  | December 03, 2007 | 2007 | Fresenius Medical Care North America | Leaking: the Dialyzer may leak at the header resulting in small amounts of blood loss. | Process control | Consignees were notified by telephone and certified letter , Urgent Device Recall, on 12/03/2007. They were told to discontinue use and to return all unused affected dialyzers to Fresenius. |  | 2791 | 2,791 cases (33,492 dialyzers) | Nationwide including the states of AZ, CA, HI, IA, ID, IL, IN, KS, KY, MI, MN, MO, NE, NV, NY, OH, OR, PA, WA, and WI. | Terminated | February 08, 2008 | 67 | 2791
Z-0739-2008 | 36595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8626-10. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA.-- (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump motor stall due to gear shaft wear | Device Design | Beginning August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 23 | 23, 895 total for recalls Z-0739-0746-2008 | Worldwide: USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of HONG KONG, AUSTRALIA, CANADA, AUSTRIA, BELGIUM, CROATIA, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, IRAN, ICELAND, IRELAND, ITALY, THE NETHERLANDS, NORWAY, POTUGAL, SAN MARINO, SAUDI ARABIA, SPAIN, SWEDEN, SWITERLAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, IRAQ, BRAZIL, and MEXICO. | Terminated | June 06, 2011 | 1403 | 23
Z-0727-2008 | 45940 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Cardioblate Gemini-s, model 49260. Surgical Ablation Device. Sterilized using Ethylene Oxide. Two, unique jaw curvatures are provided: a standard curve (model 49260) and extra curve (Model 49261). The device is intended for intermittent operation. Manufacturer: Medtronic, Inc. Minneapolis, MN 55432-5604. |  |  | November 21, 2007 | 2007 | Medtronic Cardiac Surgery Technologies | Mechanism failure: Medtronic has identified an issue with four lots of the Cardioblate Gemini-s Surgical Ablation Devices in which the polycarbonate distal coil retainer can fail during the course of a procedure, potentially resulting in an inability to apply the necessary jaw closure force. | Nonconforming Material/Component | Certified letters were sent on 11/21/07 to Consignees. The letter identifies the issue with the device and requests that customers return the device back to the firm. Customers will be requested to complete a Field Correction Action certificate acknowledging the receipt of the information. For additional information you can contact a Medtronic Sales Representative at 763-391-9167. |  | 28 | 28 units | Nationwide Distribution including states of FL, IL, IN, LA, MI, PA, TN, TX, UT, and WI | Terminated | December 17, 2011 | 1487 | 28
Z-0557-2008 | 45862 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | FRAZIER Surgical Suction Instrument, 7 French with Control Vent and Obturator (Bend with Obturator in Place), Product Code: 0031070, 50/C. STERILE, EO, Rx ONLY. CONMED CORPORATION 310 Broad Street, Utica, NY 13501 |  |  | September 12, 2007 | 2007 | Conmed Corporation | Sterility (package integrity) compromised -- Frazier and Poole suction instruments may have an inadequate seal. The product pouch seal made during assembly of this product, may be incomplete or not present. | Packaging process control | URGENT-Device Recall Letters (dated 9/10/07) and FIELD CORRECTIVE ACTION CERTIFICATION forms were sent via Certified Mail to affected customers the week of 9/09/07 (mailed on 9/12/07). The recall letters informed the consignees of the recall due to potentially inadequate seals of the instruments, and instructed them to review their inventory of the instruments checking them for inadequate seals. The letters also instructed the consignees to complete the form and fax it back by 9/21/07. For those consignees that have any of the recalled instruments or may have questions, please contact ConMed Quality Assurance at 1-800-552-0138 ext. 5225 between the hours of 7 am - 6 pm (MST). International customers were notified on 9/18/07. A second recall letter (dated 10/09/07) was sent to non-respondents of the initial mailing. ConMed plans to send a third recall letter to the non-responders during the week of 12/10/07 which addresses sub-recall notifications to the distributors' customers. |  | 53691 | 53,691 instruments | Nationwide and Worldwide Distribution | Terminated | August 30, 2012 | 1814 | 4700864
Z-0948-2008 | 39366 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Functional Anaesthetic Discography (F.A.D.) Catheter |  |  | April 25, 2006 | 2006 | Kyphon Inc | Guidewire Breakages: during the Functional Anaesthetic Discography (F.A.D.) procedure. | Labeling design | A customer advisory letter was sent on 4/25/06 to consignees. In some cases some were issued in person. Follow up with field representatives was also conducted. Both contacts involved providing IMPORTANT CHANGES IN INSTRUCTIONS FOR USE that are to be attached to the on site users documents. The recalling firm is updating future distribution with the new instructions for use. |  | 23 | By Lot Number : J 6013102 = 147, Lot: J6033012 = 23 | Nationwide including states of AL, AZ, CA, FL, IN, NH, NY, DC, ME, MD, WI, and PA. | Terminated | June 10, 2008 | 777 | 66
Z-0693-2008 | 45847 | No | MHY |  |  | Neurology | PMA | Medtronic Kinetra 7428, Dual Program Neurostimulator for Deep Brain Stimulation, Rx only, Medtronic, Inc., Minneapolis, MN 55432-5604 USA. (The Medtronic Kinetra Model 7428 Neurostimulator is a multiprogrammable device and a component of the Kinetra Neurostimulator System for deep brain stimulation.) |  |  | October 01, 2007 | 2007 | Medtronic Neuromodulation | Cessation of Therapy: separation of internal connections (lifted wirebonds) between the electronic circuit and battery may lead to sudden cessation of therapy. | Packaging change control | A letter "Medtronic Urgent: Device Recall" dated October 2007 was sent to Healthcare Professionals. The letter described the device defect, product involved, root cause, probability of occurrence, recommendations and also provided information for assistance. At the time of the recall all of the product remaining in date for use (not expired) had been implanted, therefore Medtronic provided Important Patient Management Information for the implanting physicians to use in patient care and decisions regarding patient explants. |  | 1427 | 1,427 | Worldwide: USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, WA, WI, WY, UT, VA, VT, and WV and countries of Australia , Austria , Belgium , Canada , Cyprus , Czech Republic , Denmark , Finland , France , Germany , Greece , Hong Kong , Hungary , Iceland , India , Iran , Israel , Italy , Japan , Latin America, Lebanon , The Netherlands , Norway , Poland , Portugal , Russian Federation , Saudi Arabia , South Africa , Spain , Sweden , Switzerland , Turkey , and The United Kingdom | Terminated | May 03, 2012 | 1676 | 3079
Z-0697-2008 | 45936 | No | DJZ | Breath-alcohol test system. | Toxicology | Toxicology | 510(K) Exempt | Maximo Concepts Alcohol Breath Tester, Model Number: MAX91721, Sakar International, Inc., Edison, NJ (LED 3 Step Alcohol Breath Tester) |  |  | November 30, 2007 | 2007 | Sakar International, Inc. | Marketed without a 510(k) | PMA | Recall notification letters , Urgent Device Recall, were sent by first class mail on 11/30/2007 requesting return of product and notification of accounts sold product. |  | 72581 | 72,581 | Nationwide to retail distribution centers in CA, CO, IL, KS, MA, MO, NE, NJ, NY, PA, VA and WI. | Terminated | June 30, 2010 | 943 | 72581
Z-0698-2008 | 45949 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Quick-Cross Support Catheters, .035in x 90cm, REF#: 518-036, Sterile, Manufactured by: Spectranetics, Colorado Springs, Colorado 80907. |  |  | December 10, 2007 | 2007 | Spectranetics Corporation | Mislabeling: Catheters were placed into mis-labeled inner pouches as REF Number 518-037, 135 cm--the exterior box labeling is correct as 90 cm. | Packaging process control | Consignees were notified by Dear Healthcare provider letter on 12/10/2007 and were told that a sales rep would remove and replace all affected catheters. |  | 64 | 64 | Nationwide including states of FL, IL, IN, MO, NC, NY, TX, and UT. | Terminated | February 08, 2008 | 60 | 64
Z-0710-2008 | 46276 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Monitoring System;Horizon Cardiology Hemo External Control Unit;~100 - 230 VAC |  |  | November 12, 2007 | 2007 | McKesson Provider Technologies | Power supply reliability: McKesson received word from their manufacturer of the medical grade power supply that some power grade supplies may not provide reliable power output. | Nonconforming Material/Component | Recalled by letter sent to the firm's customers beginning on 11-12-2007, advising them that the field correction would be completed by the recalling firm service technicians. |  | 13 | 13 | Nationwide to hospital and medical centers in LA. CA, PA and IL, | Terminated | May 28, 2008 | 198 | 13
Z-0518-2008 | 45986 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetom Espree with Swiveling OR (Operating Room) Table, Magnetic Resonance Imaging System,Siemens Medical Solutions USA, Inc., Malvern, PA |  |  | November 05, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Table may experience a deadlock situation | Software design | The recalling firm issued a Customer Safety Advisory to affected customers per Update Instructions MR031/07/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. The recalling firm has also released the software update to correct the issue via Update Instructions MR034/07/S. Notifications were issued on or about November 5, 2007. |  | 4 | 4 units | The products were shipped to medical facilities in TX and NY. | Terminated | February 05, 2008 | 92 | 4
Z-0705-2008 | 45987 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet RINGLOC + Instrumentation Impactor Plate, Size 28; REF/Part#: 31-400628, Biomet Orthopedics, Inc., Warsaw, IN. |  |  | October 18, 2007 | 2007 | Biomet, Inc. | The instrument will not lock onto the impactor handle; the device was distributed without assembly of the Bal Seal component. | Process control | Distributors were notified of the device defect by letter (Urgent Field Correction Notice) dated 10/18/07, and provided with assembly instructions and parts, and instructed to make corrections at consignee locations. |  | 141 | 141 total for all recalls Z-0700-0705-2008 | Worldwide-USA, Germany, The Netherlands, Sweden and The United Kingdom. | Terminated | March 11, 2008 | 145 | 141
Z-0135-2008 | 44879 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | THSCF-35-260-3-AUS2, (Cook Vascular Wire Guide, PTFE coated curved Amplatz ultra stiff catheter exchange, Diameter .035 Inch, length 260 cm, curve radius 3 mm, Cook Incorporated, Bloomington, IN 47402 - not on label) |  |  | August 15, 2007 | 2007 | Cook, Inc. | The device is normally sold as sterile, however this lot was not sterilized and the device is not adequately labeled, as the label does not state whether the device is sterile or non-sterile and lacks the manufacturer's name and address. | Process control | Cook Incorporated notified consignees by Urgent Product Recall letter dated 8/15/07 instructing them to return the product. |  | 33 | 33 | Worldwide Distribution - USA, Australia, India, Trinidad and Tobago. | Terminated | April 21, 2008 | 250 | 33
Z-0525-2008 | 45918 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Beckman Coulter Flow-Count Fluorospheres - Part Number 7547053, Beckman Coulter, Inc., Fullerton, CA 92835 |  |  | May 21, 2007 | 2007 | Beckman Coulter Inc | Sporadic absolute count recovery failures with assayed control cell products when using Lot Number 754801F of the Flow-Count Fluorospheres. Additionally, this lot has experienced secondary fluorescent populations containing more than 20% of the total population, as stated in the Evidence of Deterioration section of the Package Insert. | Process design | A Product Corrective Action (PCA) letter was mailed on May 21, 2007, to all accounts that have ordered Flow-Count Fluorospheres, Part Number 7547053, Lot # 754801F. The letter advised that Beckman Coulter has confirmed absolute count recovery failures with assayed control cell products. Additionally, internal testing has indicated secondary fluorescent populations slightly above their indicated 20% of the total population as stated in the Evidence of Deterioration section of the Package Insert. The letter requests that users immediately discard any bottles of the product and call for a replacement. If no evidence of deterioration has been observed, the firm states that there should not be any negative impact to the results obtained. A response form is included and should be returned to the firm within 10 days. Questions regarding this recall can be addressed by contacting Beckman Coulter Customer Service at 1-800-526-7694. |  | 134 | 97 units in the US; 37 units Internationally | Worldwide Distribution --- including USA and country of Canada. ***CENTER COMMENTS -- The Center requests that the subject recall not be terminated until CDRH has determined that the firm has identified the root cause and performed an adequate CAPA to address the cause of the problem, i.e., quality control.*** | Terminated | October 15, 2012 | 1974 | 134
Z-0747-2008 | 46378 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | CORDIS "Dura Star" 2.25 x 10 Dilatation Catheter, Catalog # 70110225, Distributed by Cordis Corporation, Miami Lakes, FL 33014 |  |  | January 14, 2008 | 2008 | Cordis Corporation | Slow Deflation or No Deflation | Process control | Cordis initated the recall on 1/14/2008. For U.S. customers: An Urgent Medical Device Recall letter and an Acknowledgement Form was sent to the following 4 individuals in each account: Cath Lab Director, Operating Room Director, Materials Director, Risk Manager; to be followed up with direct Cordis representative contact to facilitate obtaining signature, faxing form, collecting and returning units. For non-US customers, email notification with return acknowledgement is used to notify the affiliate distributors, who then notify and reconcile the customers in those countries. All users were instructed to immediately set aside product and return it for replacement. |  | 133300 | 133,300 units total for recalls Z-0747-0824-2008 | Worldwide. | Terminated | October 20, 2009 | 645 | 133352
Z-0826-2008 | 45532 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI |  |  | January 04, 2008 | 2008 | Xoran Technologies, Inc. | Misbranding; These devices do not bear the manufacturer's name or address. | Labeling design | Consignees were notified via letter, Urgent Medical Device Correction, dated 1/4/08 of the need to apply proper labeling to the products, and were sent instructions and labeling to affix to their units. |  | 71 | 71 | Nationwide. | Terminated | October 22, 2010 | 1022 | 71
Z-0934-2008 | 45407 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Inotech brand Heparin Adsorbant, Catalog #: HA-1, in 10mm x 75mm polystyrene tubes, each containing 70mg of Heparin Adsorbent, 10 tubes per foil lined pouch, 10 foil pouches/box, Inotech Biosystems International, Inc., Rockville, MD 20855 |  |  | August 20, 2007 | 2007 | Inotech Biosystems Intl., Inc | Erroneous Coagulation Tests | Labeling design | The recalling firm notified consignees by phone on 08/20/07 and provided instructions for increasing the speed of centrifugation from 1,500 x g to 2,500 x g and decanting the supernatant aliquot carefully so as to not draw in resin fibers that may be stuck on the test tube walls during sample preparation for testing in the coagulation studies. The firm updated their recall notification on 10/01/07 as an Urgent Recall/Correction to alert healthcare practitioners to not use "lupus sensitive reagents" with the heparin absorbant product and that Platelet Factor 3 and any other negatively charged molecule in the plasma may interact with the resin causing false high readings due to the partial removal of the clotting initiating factors. |  | 37370 | 37,370 tubes | Worldwide; USA to Hospitals and Clinical Laboratories and to one domestic distributor for export to Canada. | Terminated | August 07, 2012 | 1814 | 37370
Z-0849-2008 | 45912 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NexframeÂż Stereotactic System Kits, Model Number(s): DB-1021-MR, DB-1031, DB-1040-BL, DB-1040-ST, DB-1041, |  |  | October 04, 2007 | 2007 | Medtronic Image Guided Neurologics, Inc. | Sterilty (package integrity) Compromised: Some failures were for damage to the outer pouch, while another set of failures were for the seals on this pouch. The seal between the inner tray and lid has not been compromised and the contents remain sterile, however, the sterility of the outer surfaces of the inner tray and lid cannot be assured. | Package design/selection | On 10/4/2007, consignees were notified by Medtronic field personnel, either by personal visits or telephone. All wholesalers of this product are OUS. Medtronic OUS Regulatory personnel communicated this issue directly to the wholesalers, who provided the letters to their customers. Medtronic field personnel are responsible for completing and returning reply cards confirming all notifications (100%). |  | 385 | 385 | Worldwide, USA, Taiwan, Thailand, Japan, Spain, Sweden and The Netherlands. | Terminated | July 13, 2012 | 1744 | 385
Z-0125-2008 | 44854 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medtronic Sofamor Danek COLORADO 2" SPINAL SYSTEM Break-Off Nut, non-sterile, RX; Catalog Number 8634111, Medtronic, Memphis, TN 38138. |  |  | September 10, 2007 | 2007 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | Assembly may be compromised: Two thread grooves are missing on the nut, which may cause damage to the bone screw and result in the need for the bone screw to be replaced. | Process control | Medtronic Sofamor Danek notified consignees by Urgent Device Recall letter dated 9/10/07 to immediately cease using the product and to return the devices with a Product Recall Questionnaire, recording quantity of devices and any patient injuries. |  | 2569 | 2,569 | Worldwide Distribution - USA, Australia, Hong Kong and the Netherlands. | Terminated | July 14, 2009 | 673 | 2569
Z-0844-2008 | 45846 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD Sterilizers ( Models 100S, 50, 200) Compatible Medical Device Reference List for US Customers, Document Numbers 15-53229 &amp; 15-53230; STERRAD NX Sterilizers Compatible Medical Device Reference List for NX Customers, Product Code 15-53231; STERRAD NX System Brochure, Product Code AD-53421-001, Advanced Sterilization Products, Irvine, CA |  |  | November 12, 2007 | 2007 | Advanced Sterilization Products | Outdated Compatibility Lists: Recalling firm has decided to discontinue dissemination of Compatible Medical Device Reference Lists and instrument assessment activities | Other | Notification letters, Urgent Product Correction, were sent to customers beginning on November 13, 2007 via UPS overnight. Customers were notified this action was being taken as a result of an investigation by the firm determined that all brand- and model-specific compatibility lists and associated instrument assessments provided to STERRAD Sterilization System customers are now out-of-date. User were instructed to discontinue their use and destroy them. |  | 9686 | 9,686 units | Worldwide: USA, France, Germany, Italy, The UK, Switzerland, Benelux,Uruguay,Venezuela,Argentina,Chile,Colombia,Ecuador,Belgium, Puerto Rico, Hong Kong,Indonesia,Mexico,Philippines,Australia,Austria,China,Greece,Ireland,Israel, Japan,Korea,Malaysia,Singapore,Taiwan,Thailand,Turkey, Middle East, Peru,Canada,The Czech Republic,Portugal,S Africa,Spain,Sweden,Hungary,Poland,Brazil,Russia,Egypt,India &Slovenia; and Canada | Terminated | December 19, 2011 | 1498 | 9686
Z-0967-2008 | 45902 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker LFIT COCR V40 Femoral Head; xx-Long (+16mm) Neck Length 32mm Diameter Sterile, Forged Vitallium Alloy Product Number 6260-9-532 |  |  | November 02, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Fracture of Stem Neck - The firm recently performed laboratory testing indicates that there may be a potential increased risk of stem neck fractures when a V40 +16mm offset femoral head is mated with certain Accolade TMZF femoral stems in patients classified as overweight or obese. | Other | Stryker Orthopaedics sent an Urgent Product Recall notification letters on November 9, 2007, by Federal Express with return receipt. The letter informed users of potential increased risk of stem neck fracture, when a V40 + 16mm offset femoral head is mated with certain Accolade TMZF femoral stems in patients classified as overweight or obese. The firm has asked that inventory be examined, removed and held in quarantine. A Stryker representative will contact the medical facilities to retrieve the product. If a representative has not contacted the facility to date, it is suggested that they contact Stryker at (201) 831-5972. |  | 5599 | 5,599 total units | Nationwide Distribution | Terminated | August 30, 2010 | 1032 | 5599
Z-0718-2008 | 38335 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | Hamilton Sample Carriers for the ML STAR line (SMP-CAR-24). Product Code: 866.2500; Part Number: 173400. Hamilton Company, Reno, NV 89520 (Accessory on the Microlab STAR series instruments.) |  |  | June 25, 2007 | 2007 | Hamilton Co | Incorrectly labeled-- Product has the wrong barcode labeling. The barcode error will not influence the results of testing. The instrument will discover the error and post an error notice. | Error in labeling | The firm sent an Urgent: Product Recall letter to consignees on June 25, 2007. The firm also issued an alert via their User news distribution. They instructed consignees to examine their stock and return any affected carriers. A fax back form was supplied with the notification. The firm will be providing replacements. |  | 179 | 179 units distributed. | Worldwide Distribution --- USA including states of OH, CA, NJ, NC, and countries of Japan, France, Switzerland, Spain, Italy, Germany, and the UK. | Terminated | April 15, 2008 | 295 | 179
Z-0676-2008 | 45329 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems -- Cardiac Viewer or Pulmonary Viewer software application version 2.0 in use with Precedence SPECT/CT System and all CT EBW workstations using software version 3.5. Philips Nuclear Medicine, Milpitas, CA 95035 |  |  | September 28, 2007 | 2007 | Phillips Nuclear Medicine | Incorrect Measurement: The firm discovered a software defect in Cardiac and Pulmonary Viewer software application options used in the Precedence CT Sub-System. The defect causes incorrect measurements in derived images which are zoomed and saved in a batch file. | Software design | A Customer Advisory Notification was mailed to affected consignees on September 28, 2007. The firm informed consignees that incorrect measurement can occur when the Cardiac Viewer or Pulmonary Viewer is used to zoom and save derived images in batch files. The firm provided instructions for operators to prevent incorrect measurements. Phillips is currently investigating methods to correct this issue and will provide updates, free of charge, when they become available. If you have questions concerning this recall, please contact Phillips Medical at 1-800-722-9377, (option5: Diagnostic Imaging, option 5:NM). |  | 19 | 19 | Worldwide Distribution --- USA including states of FL, NY, VT, MD, TX, and MO and country of Canada. | Terminated | February 04, 2009 | 495 | 19
Z-0847-2008 | 45851 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer, Product Code: 10201, Advanced Sterilization Products, Irvine, CA |  |  | November 12, 2007 | 2007 | Advanced Sterilization Products | User Guide Update: based on user reports about contact with residual hydrogen peroxide from instruments, pouches and trays after sterilization and subsequent light colored residue on these devices after sterilization updated user guidance has been issued. | Employee error | Urgent Product Correction letters were sent to customers beginning on November 13, 2007 via UPS overnight. Customers were notified this action was being taken to clarify instructions contained within the STERRAD System User Guides to reinforce appropriate use and to address two separate issues which have arisen from user reports: (1) User reports about contact with residual hydrogen peroxide from instruments, pouches and trays after sterilization. (2) User reports about light-colored residues observed on instruments, trays and pouches after sterilization. The Product Correction Letters instructed customers to refer to the revised excerpt of the User Guide (attached to the letter) for information about instrument preparation prior to sterilization, particularly with regard to the washing, rinsing and thorough drying of instruments prior to sterilization in the STERRAD Sterilization System. If customers have any questions they were instructed to call ASP Customer Care Center to answer additional questions at (888) 783-7723. |  | 9686 | 9,686 units total for recalls Z-0845-0848-2008 | Worldwide: USA, France, Germany, Italy, The UK, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, The Philippines, Switzerland, Australia, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Middle East, Peru, Canada, The Czech Republic, Middle East, Portugal, S Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India &amp; Slovenia | Terminated | March 14, 2012 | 1584 | 9686
Z-0688-2008 | 45910 | No | JFJ | Amylase test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche LIPC Lipase Colorimetric Reagent; Catalog Number: 03029590322. Roche Diagnostics Corp., Indianapolis, IN 46250 |  |  | December 03, 2007 | 2007 | Roche Diagnostics Corp. | False Results- The reagent of the LDL-Cholesterol assay shows a carryover effect on the Lipase assay. Results may be falsely elevated by up to 50 units per liter when used on COBAS INTEGRA 400 or COBAS INTEGRA 400+ analyzers. | Process design | Consignees were sent and Urgent Medical Device Correction notice dated 12/03/07. The notice provided instructions for conducting an additional wash cycle whenever both the LDL-Cholesterol assays and Lipases assay are used on the COBAS INTEGRA 400 or COBAS INTEGRA 400+ analyzers. Contact Roche Technical Support at 1-800-428-2336 for questions regarding these instructions. |  | 9817 | 9,817 kits. | Nationwide Distribution | Terminated | June 11, 2008 | 191 | 9817
Z-0950-2008 | 45595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pumps, Models 8626-10, Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump Motor Stall; pumps can stall due to gear shaft wear. | Device Design | Begining August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 81023 | 81,023 total for recalls Z-0950-0957-2008 | Worldwide-USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, DS, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and countries of HONG KONG, JAPAN, AUSTRALIA, CANADA, BELGIUM, CROATIA, AUSTRIA, CYPRUS, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN, ICELAND, IRELAND, ITALY, JORDAN, KWAIT, LEBANON, MALTA, The NETHERLANDS, NORWAY, POLAND, PORTUGAL, SAN MARINO, SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, The UNITED KINGDOM, CHILE, VENEZUELA, BRAZIL, and MEXICO. | Terminated | December 27, 2011 | 1607 | 81023
Z-0881-2008 | 46322 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | PERFEKTUM TOOMEY EVACUATING SYRINGES, REF/Part # 5047, DESCRIPTION: 30CC LOCK TIP, RX ONLY, NON-STERILE. Manufactured By: Popper &amp; Sons, Inc. New Hyde Park, NY. 11040. |  |  | January 02, 2008 | 2008 | Popper &amp; Sons Inc | Mis-labeling: the product name on the label was erroneously declared as PERFECTUM TOOMEY EVACUATING SYRINGES instead of POPPER INTERCHANGEABLE HYPODERMIC SYRINGES for one lot of the product. Note: the product information (part number, description, lot number) found in the white box portion of the label is listed correctly. | Process control | The recalling firm, Popper &amp; Sons, began contacting its customers by telephone on/about 1/02/08, informing them of the issue. This was followed up with formal recall letters, Urgent Recall, and response forms starting on 1/09/08 via Certified Mail, Return Receipt Requested to the US accounts, and Registered Mail to the two Canadian accounts. The firm has directed the customers to discontinue using or distributing the lot and return any remaining quantity of this lot of syringes. If the lot has been further distributed, the distributors were instructed to immediately contact their accounts, advise them of the recall situation and have them return their outstanding recall stock as indicated above. |  | 272 | Domestic - 259 units; Canada - 13 units | Worldwide-USA including the states of NY, NJ, PA, MI, IL, OH, FL, MO, TX, and CA, and country of Canada | Terminated | May 12, 2008 | 131 | 272
Z-0873-2008 | 46267 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RenalPureÂż Liquid Acid Concentrate R-259; One Gallon (3.78 liters); manufactured by Rockwell Medical Technologies, Inc., Grapevine, TX 76051. |  |  | December 03, 2007 | 2007 | Rockwell Medical Technologies, Inc | Foil Seal Deterioration; occuring on the tamper resistant foil seals on dialysis liquid concentrate and Dri-Sate product containers older than two years old. | Process control | Firm began notification of consignees by letter, Alert: Voluntary Recall, on 12/03/07, sent via e-mail, fax, and/or mail. Consignees asked to check inventory and to identify and call for the replacement of any product older than 2 years old. |  | 6 | 6 cases (4 one gallon containers/case). | Nationwide Distribution to the following states: AL, AR, FL, KS, LA, MO, MS, OK, TN and TX. | Terminated | March 31, 2008 | 119 | 49
Z-1040-2008 | 37183 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Anspach MicroMax Pneumatic Motor System; ENT Surgical Drill, including handpiece and Accessories Attached. All Models; Black Motors; Anspach Effort, Inc., Palm Beach FL 33410. |  |  | December 18, 2006 | 2006 | The Anspach Effort, Inc. | Incorrect Sterilization Process --The Operating Manual for the product contains an incorrect statement regarding the sterilization of the Autolube foot control. The manual states that the product can withstand ETO sterilization, but there is no data to support this process. | Labeling False and Misleading | Anspach issued an Urgent Device Correction letter on 11/27/06, to all manual holders advising them of the error and requested that they update the manual with the corrected instructions. A reply receipt confirming notification was included. |  | 7399 | 7,399 total units | Worldwide Distribution | Terminated | July 06, 2012 | 2027 | 7399
Z-0880-2008 | 46304 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Comfort T-Handle Hudson Connector with Impactor Cap; Part #1006-1; handle is component of the following consignment surgical kits: Harmony 2700-0001-PL, Harmony 2700-0009-PL, Harmony 2900-0001-PL, Harmony 2900-0002-PL, Cadence 1500-0002-PL and Cadence 1500-0008-PL; Distributed by Abbott Spine, Austin, TX 78727. |  |  | December 19, 2007 | 2007 | Abbott Spine | Dislodged Collar Spring; The collar spring binding on the shaft of generic handle used with multiple surgical instruments has the potential to release the device's internal bearings during use. | Device Design | All medical suppliers with consigned inventory from affected lots were contacted via phone beginning 12/19/07. Consignees notified of potential issue and were asked to inspect their inventory and contact Abbott Spine for return and replacement of the affected handles. |  | 230 | 230 handles. | Worldwide:Medical suppliers in the following states: AR, AZ, CA, CT, FL, IL, LA, MD, MI, MO, MN, NC, NJ, NM, NV, OH, OK, OR, PA, TX, VA and WI. and country of Spain. | Terminated | October 04, 2010 | 1020 | 230
Z-0879-2008 | 46288 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smiths CADD Yellow Medication Cassette Reservoirs with Flow Stop, Clamp and Female Luer, 100 ml, Reorder No. 21-7300-24, Nonvented stopper included, Sterile, Smiths Medical, St. Paul, MN 55112 |  |  | December 14, 2007 | 2007 | Smiths Medical MD, Inc. | Leakage associated with the CADD Medication Cassette Reservoirs for use with the CADD Ambulatory Infusion Pumps | Process control | Consignees were mailed the recall letters on December 21,2007. The recalled letter "Smiths Mecical Urgent Product Recall Notification" makes them aware of the potential issue. Customers will be requiired to return the affected devices to Smiths Medical. |  | 4644 | 4,644 | Worldwide Distribution --- USA including states of AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND,NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY, and countries of AN, AR, AU, BR, CA, CL, CN, CZ, DK, EE, ES, FI, FR, GB, GR, HK, HU, IE, IL, IT, JO, JP, KW, LB, MY, NL, NO, PA, PL, PT, SA, SE, SG, TAIWAN, TH, VI, ZA, CH, AT, and BE. | Terminated | January 11, 2011 | 1124 | 1266792
Z-0112-2008 | 39371 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Blood glucose meter, marketed with multiple brand names including: Precision Xtra New and Improved meter, Precision Xceed meter, Optium Xceed meter, Optium meter, ReliOn Ultima meter, Rite Aid meter, Kroger meter, Omron HEA-214 meter, Optium Xido meter, Boots meter. |  |  | August 02, 2007 | 2007 | Abbott Diabetes Care, Inc. | Damage could lead to no display. Meters manufactured after January 31, 2007 could exhibit meter display damage if dropped on a hard surface. These meters could exhibit unreadable lot number fields and date/time fields, in addition to complete blanking of the numerical reading portion of the display. | Component design/selection | Abbott sent Urgent Medical Device Correction letters, dated 31Aug07, to registered users/customers via FedEx or mail. The letter instructs users to verify that their display is correct, and if not, to discontinue use and call Abbott Customer Care at 1-877-844-4404 for a new meter. Warranty kits will have fliers added to them to alert users. Email and telephone contacts may be made as needed. U.S. consignees and registered users will be contacted by letter, telephone call or visit starting 23Aug07. Worldwide, customer lists will be created and communications tracked according to local requirements. Abbott issued a Press Release August 31, 2007 also. |  | 279289377839 | 882,601 total meters distributed, 316,439 in the U.S. and 566,162 outside of the U.S. including Mexico and Canada | Worldwide Distribution - USA, Canada, Mexico, Australia, Malaysia, Philippines, Germany, New Zealand, China, India, Singapore, Hong Kong, Taiwan, Argentina, Brazil, Chile, Colombia, Costa Rica, Ecuador, Mexico, Peru, Puerto Rico, Uruguay, Venezuela, and Japan. | Terminated | March 28, 2011 | 1334 | 279289377839
Z-0619-2008 | 45469 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Allogenix; DBM in a Lipid Carrier with Calcium Granules; Model Numbers: 02-3110, Interpore Cross International, a Biomet company (Allogenix DBM products contain human tissue (allograft bone) and are intended for transplantation) |  |  | October 15, 2007 | 2007 | Interpore Cross International Inc | Unapproved Testing: Interpore Cross International was recently informed that one of its human tissue suppliers, LifeLink Tissue Bank ("LifeLink"), utilized unapproved diagnostic testing (HBsAg &amp; HBc) on certain lots of distributed donor tissue instead of the FDA approved screening test. As a result, Interpore Cross International has initiated a recall on specific lots of products containing human | Nonconforming Material/Component | The Recall Notification letters will be sent via commercial carrier (UPS) with signature for verification of receipt required on Friday October 26, 2007, to all consignees, including hospitals and sales representatives, which summarizes the recall details and instructs the consignee to contact Interpore Cross via mail or facsimile with the enclosed Customer Response form, and to immediately return the affected product. |  | 34 | 34 | Nationwide | Terminated | September 25, 2008 | 346 | 1290
Z-0677-2008 | 46157 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Mevatron KD2 Medical Charged-Particle Radiation Therapy System; (Interlock #51 Jaw Field Position) Model Number:08515520, Siemens Medical Solutions, Concord, CA 94520 |  |  | August 07, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect field size -- Products manufactured prior to 1999 may experience jaw field size errors during interlock resulting in potential mistreatment if a light field check is not performed. This problem is present in machines with serial numbers less than 3094, as they do not contain redundant sensors for jaw position. | Device Design | The firm has sent consignees a Customer Safety Advisory Notice on August 7, 2007 informing them that a misplacement of the X and/or Y jaws can occur if interlock #51 Jaw Field Position is asserted and then cleared with the left-hand key rest. The firm recommends that if an interlock issue is encountered tha treatment of the patient be discontined until this issue is resolved. |  | 55 | 55 units | Worldwide Distribution | Terminated | December 01, 2010 | 1212 | 642
Z-0545-2008 | 45520 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Accelerator, Linear, Medical charged-particle radiation therapy system; |  |  | September 10, 2007 | 2007 | Varian Medical Systems Inc | The collimator drive chain may break or slip off of its drive track, allowing the collimator to rotate freely without motor control; if undetected resulting in a treatment with the wrong collimator angle. | Labeling design | On 9/10/07, the firm issued an Urgent Medical Device Correction letter to all its direct consignees, informing them of the device defect and providing instructions on the recall. Short term user instructions were provided and the firm will schedule follow-up visits to make the needed permanent device corrections at each device site. |  | 418 | 418 (total for Recalls Z-0544-0545-2008) | Worldwide- Canada, Mexico, and the USA | Terminated | November 27, 2010 | 1174 | 418
Z-0825-2008 | 45502 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OsteoStim Cervical Allograft (Lordotic Cervical Spacer) |  |  | October 23, 2007 | 2007 | EBI, L.P. | Unapproved Testing: Tissue supplier (LifeLink) utilized unapproved diagnostic testing on certain lots of distributed donor tissue instead of the FDA approved donor screening test. (Note: tissue samples were later re-tested by an approved method and found non-reactive) | Nonconforming Material/Component | Urgent Recall Notice letters were mailed on October 23, 2007, requesting location of all product including sub-accounts and return for credit. |  | 19 | 19 units | Nationwide to distributors/customers in NY, NC, MO, CA, FL and KS | Terminated | April 15, 2008 | 175 | 19
Z-0720-2008 | 45527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand VARiS 1.4g (Medical Charged Particle Radiation Therapy System, Record and Verify System); Linear Accelerator with RTP Exchange v6.2, v6.6, &amp; v8.0, All Models; Varian Medical Systems, Palo Alto, CA |  |  | October 02, 2007 | 2007 | Varian Medical Systems Inc | Incorrect Software Validation- The ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery that will not be recognized properly by the system. Under these circumstances, an arc plan can be administered to a patient whereby the gantry will not rotate, resulting in delivery to a single location rather than through the gantry arc. | Software design | An Urgent Medical Device Correction letter was distributed October 2, 2007, to all direct consignees, informing them that the ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery will not be recognized properly by the System. The firm provided corrective instructions and informed consignees that a product modification will be developed to resolve the issue. A Varian Service Representative will be scheduling a visit to upgrade the application. |  |  |  | Worldwide Distribution | Terminated | May 17, 2011 | 1323
Z-0721-2008 | 45530 | No | JTR | Quality control kit for culture media. | Microbiology | Microbiology | 510(K) Exempt | PML -- LyfoCults Haemophilus parainfluenzae - micro organism. Product label includes LyfoCults Haemophilus parainfluenzae ATCC 7901 Lot No.: 233851-1 Exp 05-17-09. PML Micrcobiologicals, 27120 SW 95th Avenue, Wilsonville, Oregon 97070 |  |  | October 15, 2007 | 2007 | PML, Inc. dba Pml Microbiologicals Inc. | Incorrect micro-organism-- Quality control In-vitro diagnostic test was manufactured with Cryptococcus neoformans (ATCC 76484) instead of Haemophilus parainfluenzae (ATCC 7901). | Mixed-up of materials/components | On 10/15/2007, firm notified consignees by telephone. A follow up letter dated 10/26/07, was faxed to customers instructing them to discontinue use and discard remaining inventory on site as a biohazardous material. Consignees were requested to fax back response indicating the amount of product they had on hand at time of notification. |  | 150 | 30 packs (each pack contains 5 vials of Haemophilus parainfluenzae) | Nationwide Distribution | Terminated | April 07, 2011 | 1270 | 151
Z-1107-2004 | 45411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemans- ACUSON Aspen Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system, Model Number: 8247887, Siemans, Mountain View, CA 94039 |  |  | September 18, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Software error: Error affects all Aspen customer systems using the PAL video standard configuration and have a Perspective PC installed. This can lead to a measurement error under this specific condition: On secondary capture images/clips stored from within any Perspective advanced display option (specifically: 3D Surface Rendering, 3D MultiPlanar Reconstruction, Color Doppler 3D MultiPlanar Reco | Software design | Siemens issued an Urgent Medical Device Notification to consignees 9/18/07, under Siemens update program US014/07/S. (Distribution on this product had ceased on 6/29/2005.) The firm has instructed the user to not take measurements or use measurements on these secondary captures under the conditions described in the notice, to place a copy of the instructions for use in the file, and to notify all users of the issue with the software. The Siemens Service organization has coordinated the implementation of the field correction with the distributors for all countries. Non-responding consignees will be visited or called by telephone for follow up. |  | 280 | 280 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain (UK), Greece, Hong Kong, India, Italy, Jordan, Malaysia, Norway, Poland, Portugal, Romania, Russian Federation, Sweden, Switzerland, Thailand, and Turkey. | Terminated | April 21, 2011 | 1311 | 280
Z-0724-2008 | 45815 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova - EZ CHEM Creatinine Meter: SRS Versions 5.1 and 5.2 |  |  | November 02, 2007 | 2007 | Nova Biomedical Corporation | Incorrect Results - Creatinine results were lower than the laboratory reference method. | Device Design | Nova Biomedical issued an Urgent: Medical Device Recall notification letter to consignees dated November 19, 2007. The letter addressed the issues with the product(s) and requested that if users had any remaining stock that it must be returned. The firm also provided a Customer Response form to be returned. Questions concerning this recall will be addressed by Gina Vogelsberg at 1-800-544-4624 or by email at gvogelsberg@ezem.com. |  | 133 | 133 units | Nationwide Distribution, including the state of NY | Terminated | December 09, 2008 | 403 | 719
Z-0728-2008 | 45953 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Invictus-- Bionime Rightest Blood Glucose Monitoring System |  |  | November 26, 2007 | 2007 | Invictus Scientific Inc | Abnormal Memory Recall: When the meter is fixed in 12 hour mode, the time frame display of all memory data will be shown incorrectly as AM or PM depending on when the user is recalling the memory stored in the meter. This will cause confusion as to when the last test was taken. | Software design | An Urgent: Device Recall Notification letter were sent to customers (direct accounts) beginning on November 20, 2007, via US Postal Service, with a return receipt requested. The distributors were instructed to inform consumers/users to fix the time format to the 24 hrs mode and to contact the distributors or the firm for replacement, if they prefer to use the time format in the 12 hrs mode. They were also instructed that if they have further distributed any of the recalled products, to please immediately contact those accounts and advise them of the recall situation, and have them return their outstanding recalled stocks. If customers have any questions, they were instructed to contact at 858-481-8485. |  | 21070 | 21,070 | Nationwide Distribution | Terminated | October 13, 2011 | 1417 | 21070
Z-0850-2008 | 45991 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Primus Linear Accelerator System, PRIMART MX Linear Accelerator System; PRIMUS Mid, Part Number:1940035 and PRIMART MX, Part Number: 5500371, Siemens Medical Solutions USA, Inc., Concord, CA |  |  | July 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Door Hinge failures; loose or falling off (Stationary Structure Doors) | Component design/selection | Stationary Door Hinges will be replaced on each device by the recalling firm's technician visits. |  | 1284 | 1,284 total for recalls Z-0850-0852-2008 | Worldwide- USA including the states of Missouri, Georgia, Delaware, Texas, Ohio, Alabama, Kentucky, Florida, California, North Carolina, New Jersey, New York, Illinois, Pennsylvania, Colorado, Tennessee, Arizona, Pennsylvania, Virginia, New Jersey, Michigan, Wisconsin, Indiana, Nebraska, Montana, Kansas, Mississippi, Rhode Island,Washington, South Carolina, Maryland, New Hampshire, Maine, Arkansas, Idaho, Missouri, Louisiana, Oklahoma, North Dakota, Minnesota, Massachusetts, and The District of Columbia and countries of Canada and Mexico. | Terminated | April 29, 2011 | 1372 | 1284
Z-0853-2008 | 46198 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | First Aid Kits, soft canvas, plastic, and metal containers holding various products to include ibuprofen and other first aid supplies, in various sizes. Kits include: |  |  | December 10, 2007 | 2007 | First Aid Only Inc | Failed USP Impurity testing: the Ibuprofen supplied by its own manufacturer, used in these first aid kits, tablet refills, and cabinets, is being recalled for failing its USP limits for impurities during stability testing. | Nonconforming Material/Component | The firm issued an Urgent Device Recall letter dated 10/10/07 to consignees receiving any of the products. The letter instructs consignees to return any recalled product to the firm. If any recalled prdoucts were further distributed, the consignee was requested to notify those consignees and have product returned. |  | 72904 | 72,904 kits | Worldwide: USA and to one distributor in Dubai, United Arab Emirates. | Terminated | April 08, 2011 | 1215 | 137211
Z-0079-2008 | 38311 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical, daVinci S Surgical System, Model Number: IS2000, A5.1 P1 and A5.0 P5 software level, Surgical Endoscopic Instrument Control System, Intuitive Surgical, Sunnyvale, CA 94086. |  |  | March 30, 2007 | 2007 | Intuitive Surgical, Inc. | System Lock-Up: Software anomalies could cause product failure during use, or on start-up. System transitions to a safe "soft-lock" state. | Software design | Intuitive Field Service Engineers installed a software upgrade beginning on 03/30/2007, and left a notification letter that the upgrade had been performed at each site. New software level is A5.1 P2. |  | 159 | 159 | Worldwide Distribution - USA, Belgium, China, France, India, Italy, Japan, Korea, Spain, Sweden, Switzerland, UK and Mexico. | Terminated | September 05, 2008 | 525 | 159
Z-0151-2008 | 45518 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand da Vinci S Surgical System, Endoscopic Instrument Control System, Model IS1200 with Isolation Transformer (P/N: 952012) and Vision Cart (Model VS1000), Distributed and Manufactured by: Intuitive Surgical, Inc. 950 Kifer Road Sunnyvale, CA |  |  | October 04, 2007 | 2007 | Intuitive Surgical, Inc. | Under-Rated Fuses--The Firm may have installed incorrect fuses. The daVinci Surgical System Vision Cart's isolation transformer requires a 10A fuse for 115V operation or a 5A fuse for 230V operation. Operation of this product with the use of an under-rated fuse will result in failure of the fuse and a loss of power to the vision cart, including the loss of power to any ancillary equipment that m | Mixed-up of materials/components | A product recall was initiated on October 4, 2007. Direct and Distributor consignees were sent an Important Product Notice, via electronic mail on October 5, 2007. The firm plans on sending a Field Service Engineer to inspect the system and make any necessary corrections. |  | 63 | 63 total units | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, IL, IN, KS, LA, MA, MD, MN, NC, NJ, NV, OH, PA, PR, SD, TN, TX, WA, and WI, and countries of Belgium, Ireland, Italy, Korea, Russia, and the United Kingdom. | Terminated | March 31, 2011 | 1274 | 63
Z-0583-2008 | 45807 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Three Birthing Bed, Model P3700A. |  |  | November 21, 2007 | 2007 | Hill-Rom, Inc. | Hole/opening on the labor bar -- Amputation of a finger could result, if a finger is placed into the opening, as the hole/opening creates a shear point when the mechanism is activated. | Device Design | Consignees were notified of the problem via an Urgent Medical Device Correction letter dated 11/21/07. Users were informed that they would be receiving a free upgrade kit for all of their Affinity Birthing Beds. The Kit includes two plugs per bed, an insertion tool, and full written instructions. An instruction video is available at www.Hill-Rom.com. |  | 32379 | 32,379 for all model numbers | Worldwide Distribution-- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Honduras, Hong Kong, Hungary, Iceland, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, The Netherlands, New Calidonia, New Zealand, Nigeria, Northern Mariana Islands, Norway, Oman, Panama, Philippines, Poland, Portugal, Qatar, Republic of Korea, Saudi Arabia, Singapore, South Korea, Romania, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguey, Venezuela, and Virgin Islands . | Terminated | September 22, 2009 | 671 | 32379
Z-0729-2008 | 46155 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Centurion Healthcare Products Laceration Tray, sterile; (Reorder SUT8805A) Tri-State Hospital Supply Co., Howell, MI 48844. The primary kit is labeled as containing 1 needle, 21G x 1-1/2", among other components, however some kits may contain a 21Gx1" needle. The label on the Kendall Monoject Magellan 21G Safety Needle itself is the correct size. |  |  | October 26, 2007 | 2007 | Tri-State Hospital Supply Corporation | Incorrect Needle size - Some of the kits contain a 1 inch 21G needle instead of the 1-1/2 inch 21G needle specified on the kit label. | Mixed-up of materials/components | An Urgent Product Recall Notice was sent on October 25, 2007 to Distributors, Distributors customers, and all third party consignee. The letter provided information concerning the product recall, and requested removal of all product inventory. A product return form requesting information on products still in the facility was enclosed. A Tri-State Representative will be in contact regarding retrieval of the product. |  | 40 | 40 kits. | Nationwide Distribution, including states of Massachusetts and Washington. | Terminated | March 13, 2008 | 139 | 436
Z-0732-2008 | 46318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Imaging System with VB30A version of software. Model Number: 10014063, Siemans Medical Solutions, Malvern, PA 19355 |  |  | November 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Inconsistent data - A buffer overflow may occur between information system and picture system if the system experiences a high load. (e.g. RIS/PACS communication is unusally high) | Software design | The recalling firm issued a Customer Safety Advisory, via certified mail. to the affected customers on 11/27/07. The letter informs the customer of the issue and also provides instruction to detect and avoid its occurrence. In addition, the firm is developing a software update to correct the issue, which will be installed on the affected systems when it is released. |  | 10 | 10 units | Nationwide Distribution, including states of AL, MO, NC, NY, OH, OR, TN, and TX. | Terminated | September 17, 2008 | 295 | 10
Z-0731-2008 | 46266 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView Model DR 3000 System, Stationary X-ray System; Catalog #s 6551360, 6551386, 6552152, 6552160, 6551345, 6552319, 6552335. The units are labeled under the following firms: EASTMAN KODAK COMPANY (now known as Carestream Heath Inc.) Rochester, NY 14608, or CARESTREAM HEALTH, INC., Rochester, NY 14608. |  |  | December 21, 2007 | 2007 | Carestream Health, Inc. | Unintended Positioner Movement - The U-Arm Positioner starts to move 1) without command from the operator and/or 2) can unexpectedly start movement if the system is powered down and re-energized after a collision has occurred with the Quantum Table. | Software design | On 12/21/07, Service Bulletins (dated December 2007) were sent to consignees informing them of design modification, and providing modification instructions. The firm followed up with telephone calls to the direct user sites to schedule appointments to visit each site, in order to make the necessary modifications. The firm will be installing Safety Adaptation Board and software into all the affected DR3000 systems. |  | 19 | Domestic - 19 units | Nationwide Distribution including states of NY, PA, MA, MI, IL, NC, GA, FL, KS, MO, TX, IA, and WA. | Terminated | April 10, 2008 | 111 | 19
Z-0131-2008 | 38377 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | American Medical Systems, Inc., (AMS) GreenLight HPS Fiber Optic (ADDStat) Part Number 0010-2090B. Fiber optic laser for treatment of BPH (Benign Prostatic Hyperplasia). AMS Innovative Center, San Jose, CA 95134 |  |  | June 18, 2007 | 2007 | AMS Innovative Center - San Jose | Control knob break: The knob used to finely control the direction of the laser energy may break. | Process design | American Medical Systems, Inc. (AMS) issued Important: Medical Device Recall AMS Greenlight ADDStat Fibers notification letters to customers who have purchased the product on June 28, 2007. The customer letter requests recipients that the product use be discontinued, and the fibers be returned. The firm will send replacements. |  | 611 | 611 units | Worldwide Distribution - USA and Singapore. | Terminated | March 23, 2010 | 1009 | 611
Z-0134-2008 | 44799 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical CADD-Sentry Pro Medication Safety Software - electronic pump protocol programming software - Administrator |  |  | August 14, 2007 | 2007 | Smiths Medical MD, Inc. | Incorrect dosage possible: A software anomaly in the CADD-Sentry Pro Administrator software (version B) where the values of a weight based protocol go to zero or round up to the next higher hundredths place, even if Administrator displayed the values correctly at data entry. | Software design | A Smiths Medical MD employee visited the consignees between August 14 and 31, 2007, to verbally notify them of the recall, collect the recall CD-ROM containing the affected software and replace with the updated software. |  | 3 | 3 CD-ROMs | UT, MD, CO | Terminated | October 31, 2008 | 444 | 3
Z-0674-2008 | 45878 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Orthofix TrueLok External Ring Fixation System, Telescopic Linear Distractor, Threaded in the following sizes: 75mm (Catalog Number: 50-1101), 100mm (Catalog Number: 50-1014), 150mm (Catalog Number: 50-1015) and 200mm (Catalog Number: 50-1016), distributed by Orthofix, Inc., McKinney, TX 75069. |  |  | September 28, 2007 | 2007 | Orthofix, Inc | Markings reversed: Graduated markings on telescoping bone distractors were reversed, causing the distractors to compress rather than distract the treatment site when adjusted. | Process control | Orthofix notified sales agents by phone starting on 09/28/07. Sales agents were to return product for replacements. |  | 69 | 69 units. | Nationwide Distribution - USA states: CA, FL, GA, IL, PA, MI, NY and WI. | Terminated | March 20, 2008 | 174 | 69
Z-0920-2008 | 45947 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Modular Replacement System Curved Cemented Stem; |  |  | November 15, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Curvature Discrepancy -- The radius of curvature may be out of specification to varying degrees causing a discrepency in the stem curvature between trials and implants. | Process design | On November 26, 2007, Stryker Sent Urgent Product Recall notification letters to Stryker Agencies/Branches and hospitals (OR Supervisor) by Federal Express. The notification requested that all inventory be examined and returned. Returning instructions were provided in the letter. Questions concerning this recall are being addressed by contacting the Regulatory Compliance Hotline at (201) 831-6633. |  | 803 | 803 total, all sizes | Nationwide Distribution | Terminated | September 22, 2008 | 312 | 803
Z-0930-2008 | 46146 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Abbott AxSYM Gentamicin Reagent - Is a reagent for in vitro diagnostic use. This reagent system is composed of one bottle of gentamicin antiserum, one bottle of pretreatment solution and one bottle of gentamicin flourescein tracer. Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 26, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect Results -- Through evaluation of stability, Abbott have found a decrease in the calibrator A to F span with these reagents lots, which may produce the following results: - Printed Error Code "SPAN LESS THAN MIN SPAN" - Controls out of range. When a control is out of range, patient results should not be reported. | Other | Letters were sent to customers by Federal Express Priority mail on October 26, 2007. Two unique customer letters were sent. One letter was delivered to TDxlTDxFLx customers who received shipment oflots 5ll87Q100 and 51189QI00 and a second letter was delivered to AxSYM customers who received lots 53190QI00 and 53190QI01. Two letters were written because the actions required by the customer differ for the AxSYM and TDxlTDxFLx platforms. TDx/TDxFLx customers "Customer Letter (Product Recall) "Customer Reply Form AxSYM Customers "Customer Letter (Product Correction) "Customer Reply Form The Customers were instructed the following: TDx/TDxFLx Gentamacin Reagents " Discontinue use of and destroy TDx/TDxFLx reagent lots 51187Q 1 00 and 51189QI00. " Document the number of kits discarded for your records. Abbott will provide reimbursement for your destroyed kits. " If customers have forwarded any of the affected lots to another laboratory, to provide the laboratory with a copy of the communication. AxSYM Gentamicin Reagents " Revise the expiration date for lots 53190Ql 00 and 53190QI0l from 4/12/08 to 11/30/2007. " Run each level of control with each AxSYM Gentamicin sample or batch of samples. " If customers have forwarded any of the affected lots to another laboratory, to provide the laboratory with a copy of the communication. |  | 3787 | 3787 kits | Worldwide, USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, Hawaii and Puerto Rico, an countries of Canada, Mexico, Guatemala, Germany, Australia, Japan, Singapore, Hong Kong, Taiwan, South Korea, and New Zeland. | Terminated | July 30, 2012 | 1739 | 5186
Z-0578-2008 | 46158 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TENET Medical Engineering Wrist Stabilization Kit, Medium; REF (Part) #7692; Sterile; manufactured by TENET Medical Engineering, Calgary, Alberta, Canada T2Z 4M3. |  |  | November 23, 2007 | 2007 | Tenet Medical Engineering Inc | Sterility compromised: Sterility seals on some of the firm's wrist stabilization kit pouches may be compromised. | Packaging process control | TENET notified US distributor consignee via Urgent Product Recall Notification letter sent FedEx on 11/23/07. Consignee asked to return product and to contact their consignees to also return product. |  | 360 | 360 units. | Nationwide Distribution. | Terminated | May 01, 2009 | 525 | 820
Z-1135-2008 | 46210 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Scales, 800 lb. scale weight limit; Part Number 700.00510, Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5; |  |  | December 19, 2007 | 2007 | Arjo, Inc. | Unintended Hanger Bar Detachment: The locknut on the scales may loosen during use and allow the hanger bar assembly of the patient lift to detach and fall. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 12/19/07 to the end user accounts who received the affected scales, advising them of the potential for the hanger bar assembly to detach. The accounts were given the option immediately stopping use of the lifts with the affected scales or removing the scales from the lift so the lift can continue to be used. Instructions for removal of the scales were provided with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating whether they will complete the correction upon receipt of the correction parts or if they require the assistance of an Arjo Service Technician. |  | 772 | 772 scales | Nationwide including the states of Alabama, Arizona, California, Delaware, Florida, Hawaii, Iowa, Illinois, Indiana, Massachusetts, Michigan, Minnesota, Mississippi, North Carolina, North Dakota, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, South Dakota, Tennessee, Washington and Wisconsin. | Terminated | November 18, 2009 | 700 | 798
Z-0961-2008 | 45853 | No | LGW |  |  | Neurology | PMA | Advanced Bionics Precision Implantable Pulse Generator Kit, Model Number SC-1100 with firmware version prior to Revision 3.02, Advanced Bionics Corportation, Sylmar, CA 91342 |  |  | October 05, 2007 | 2007 | Advanced Bionics Corp | Incorrect Data -- Corruption of internal memory component results in an inability for the physician to reprogram the IPG with firmware version prior to Revision 3.02. When this occurs, the IPG will report an error code of "10h0" or "00h0" through the Remote Control. Under this condition, the IPG will cease to log in some data that could be used for informational purposes. ** It should be noted t | Device Design | Advanced Bionics sent an Important Notification letter on October 5, 2007, via FED-Ex or US Postal Service Certified mail. (International notification activity is in progress.) The recall letter advises that the firm has become aware of an issue with Precision Implantable Pulse Generators (IPG model numbers SC-1100, Sc1110, and ONS-1110) manufactured prior to July 2007. The letter describes the issue in detail. The firm is requesting that customer immediately locate the affected IPGs and set them aside so that they are not used. A Boston Scientific/Advanced Bionics sales representative will contact your facility to confirm these IPGs and arrange for the exchange of these devices.The firm enclosed Reply Verification Tracking Form listing the IPGs that are at your Clinic according to the firm's records. In addition, the firm recommends that if error code 10h0 or 00h0 is encountered in Implanted IPGs, it may be corrected by upgrading to a newer version of IPG firmware (software that runs in the IPG). The new version of firmware corrects this condition and helps prevent it from occurring in the future. This firmware upgrade is performed non-invasively, through wireless RF telemetry.If any of patients encounter error code "10h0" or "00h0" on their Remote Control, consignees are advised to contact their Advanced Bionics representative or Technical Services team at (866) 566-8913 to help identify whether the condition can be corrected through a firmware upgrade and, if so, they will help arrange for the upgrade. Please note that some devices may not be upgradeable. The firm request that all consignees return the notification acknowledgement form is included with the letter. |  | 12000 | Over 12,000 units | Class II Recall - Worldwide Distribution --- including USA and countries of EU, Canada, Australia and Israel. | Terminated | February 29, 2012 | 1608 | 12000
Z-0570-2008 | 45447 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences MultiSET Flow Cytometry Software Versions 2.1, 2.1.1, and 2.2 with OS X operating systems, BD Biosciences, San Jose, CA 95131. |  |  | September 06, 2007 | 2007 | BD Biosciences | Inaccurate results: Software error results in inaccurate display result statistics. If the user adjusts the lymph gate or attractors in the Lab Report view without using the Manual Gate function (as described in the system manual), the statistic results will not be updated to match the adjusted gate. | Software design | Becton Dickinson sent an Important Product Information letter by mail on 09/14/2007 informing customers of the need to follow the instructions for adjusting the lymph gate or attractors. The next software revision will alleviate this limitation. |  | 1136 | 1136 units | Worldwide Distribution - USA, Belgium, Brazil, Peru, Singapore, Canada, Jamaica, Chile, Ecuador, Hong Kong-China, Thailand, Mexico, Venezuela, Barbados, Philippines, Puerto Rico, Guatemala, Japan, Uruguay, Argentina, China, Trinidad and Tobago, Honduras, Australia, New Zealand, and India. | Terminated | April 15, 2008 | 222 | 1136
Z-0576-2008 | 45984 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS 6000 Clinical Chemistry analyzer; Catalog Number: 04745914001, Roche Diagnostics Corporation, Indianapolis, IN 46250. |  |  | December 04, 2007 | 2007 | Roche Diagnostics Corp. | Incorrect results: CRP assays may be calculated using incorrect calibration parameters which could result in falsely high or falsely low patient results being reported. | Software design | Roche Diagnostics contacted the consignee by phone on 12/4/07 and by follow-up letter dated 12/4/07. The user was instructed to discontinue use of two CRP assays on this analyzer. The letter described a workaround and that a new revision of software, due out first quarter 2008, will fix this. |  | 1 | 1 | USA Distribution in the state of Nebraska. | Terminated | August 04, 2008 | 244 | 1
Z-0574-2008 | 45881 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Disposable StrykeProbe 32cm Spatula Tip, Ref: 250-070-551, Electrosurgical Suction Tip, Stryker Endoscopy, San Jose, CA 95138. |  |  | November 16, 2007 | 2007 | Stryker Endoscopy | Cannot connect: Tips have electrodes that are too large to fit monopolar cables and are unusable. | Device Design | Stryker notified their sales representatives by phone. Stryker sent all consignees a Product Advisory Notice dated November 19, 2007, and verbally notified them of the recall via telephone. A response form with return postage was sent with the notice. The letter requested the hospitals to quarantine and return the product to Stryker. |  | 48 | 8 boxes of 6 units (48 total) | Nationwide Distribution. | Terminated | March 26, 2008 | 131 | 48
Z-0865-2008 | 46206 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | Pulsavac Plus Low Pressure Fan Spray Kit, Catalog #00-5150-975-00, 10 kits/package, Zimmer Orthopaedic Surgical Products, Dover, Ohio |  |  | December 18, 2007 | 2007 | Zimmer Orthopaedic Surgical Products | Sterility (package integrity) compromised; The effect of a silicone stain produced during assembly operations on the sterility barrier properties of the Tyvek lids has not been validated by the firm. | Other | Zimmer notified their customers via an Urgent Device Recall letter dated January 8, 2008. The letter requests each customers to inventory their stock. If any recalled product is located the customer is to notify Zimmer via fax and a replacement part will be shipped to them. |  | 607 | 607 kits | Worldwide-USA, forty-two international and one Canadian account, including countries of Singapore, Hong Kong, Lebanon, Finland, Australia, Japan, Israel, Korea, Saudi Arabia, Dubai, Denmark, Czech Republic, Guatemala, Brazil, Taiwan, China, Nicaragua, Virgin Island, Switzerland, Austria, Spain, Belgium, The Netherlands, France, Germany, Portugal, England, Sweden, India, Argentina, Italy, Chile, Thailand &amp; Saudi Arabia | Terminated | May 03, 2012 | 1598 | 84644
Z-1182-2008 | 46877 | [6 empty] | System 83 Plus Washer-Disinfector, Endoscope use, DOS based system, Custom Ultrasonics Inc., Buckingham, PA |  |  | February 11, 2008 | 2008 | Custom Ultrasonics, Inc. | Inadvertent selection of the "wash" cycle | Labeling design | The recalling firm issued a Product Correction letter dated 2/11/08 to its customers informing them of the problem and the correction choices. The letter states that a key cover is available to reduce the error in the future and will be provided to each users upon return of the recall letter notification questionnaire. |  | 3190 | 3,190 units | Worldwide; USA medical facilities and physicians and Canada | Terminated | September 03, 2009 | 570 | 3190
Z-0947-2008 | 45861 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice PRO Automated Peritoneal Dialysis Systems; Catalog Numbers; 5C8310, 5C8310R, R5C8320, R5C8320R, T5C8300, and T5C8300R, Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A., Made in U.S.A |  |  | November 20, 2007 | 2007 | Baxter Healthcare Renal Div | Unrecorded Infusion; If the system is powered down or a power failure occurs during a fill cycle, the HomeChoice system may not record the last 1-4 pump strokes infused into the patient when the power is restored. This situation may create a potential unrecorded delivery of approximately 15-60 mL of fluid being infused into the patient. | Device Design | Urgent Product Recall letters dated 11/20/07 were sent via first class mail to all of the affected customers, informing them of the potential for overfill if the system is powered down or a power failure occurs during a fill cycle. A Baxter representative will contact the customer to schedule the software update which corects this anomaly. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1.Baxter sent a follow-up letter dated 3/17/08 to the consignees, accompanied by a copy of the 11/20/07 letter. The letter reiterated the information in the 11/20/07 letter, emphasizing the serious consequences of an overfill for neonates and young pediatric patients, and providing information describing the symptoms of excess fluid in the abdomen and steps to be taken if the user suspects overfill of a neonate or young pediatric patient. Baxter requested that Baxter be contacted if an overfill should occur so Baxter can perform a complete investigation into the event. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1. |  | 32229 | 32,229 units | Worldwide; USA, Australia, New Zealand, China, Hong Kong, the Philippines, Thailand, Asia, Canada, Algeria, Austria, Bahrain, Belgium, Bulgaria, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Lebanon, Lithuania, Malta, Morocco, The Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Senegal, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Ukraine, Latin America and Japan. | Terminated | October 26, 2009 | 706 | 103490
Z-1196-2008 | 46784 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Icy Hot Heat Therapy , Air Activated Heat Patch, Back and Large Areas, UPC Code 0-41167-08330-7, Product Code 08330, 4.1" x 8.4" (10.5 cm x 2.1 cm) each, 1 and 3 count patches, Distributed by Chattem, Inc. P.O. Box 2219, Chattanooga, TN 37409-0219 |  |  | February 12, 2008 | 2008 | Chattem Inc | Skin irritation and burns with product usage | Device Design | The firm initiated its recall on 02/08/2008 by issuing a press release warning the public of the hazard and directing them to stop its use and either return or discard the products. The firm initiated the recall at the retail level on 02/12/2008 by notifying its employee sales force by email, containing a Urgent Medical Device Recall Letter, with instructions to immediately email a copy of the recall letter to each Chattem consignee (wholesalers and direct retail accounts) with sub recall requests. |  | 666636 | 666,636 units (total for Recalls Z-1196-1197-2008) | Worldwide - USA, Canada, and Honduras | Terminated | October 05, 2009 | 601 | 712656
Z-0186-2008 | 44911 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Size 8 Pedicle Screws of the Allez Spine Laguna Pedicle Screw System, Screw Size (diameter x length) 8 x 35, Catalog Number LG-PR0835, Allez Spine, LLC., Irvine, CA |  |  | July 30, 2007 | 2007 | Allez Spine, LLC | Screw Shank Failure; separation of the screw shank portion from the pedicle head portion of the Size 8 Pedicle Screw, when torquing down the single piece locking nut during implantation. | Device Design | On July 30, 2007 Allez Spine notified all its field representatives and distributors via email of a potential for failure of the Size 8 Pedicle Screws when used in combination with the single piece locking nuts and instructed them to return all unused inventory of Size 8 Screws and the two piece Locking Nuts with saddle. This action was to be carried out to the retail level-hospital, clinic, lab, and physicians sold the product. The email also had an attachment of a list of expected frequently asked questions (FAQ) from surgeons and answers to help them respond. A second Notice to all users was sent by the recalling firm on Feb 29, 2008, to request follow-up of implanted units. |  | 317 | 317 | Nationwide: including the states of AZ, CA, CO, GA, NV, TX and UT | Terminated | April 09, 2012 | 1715 | 1578
Z-1193-2008 | 46726 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker VertePort, 10 G Cement Cannula, REF 306-410, Qty 18; For use with the Stryker PCD Precision System and ACM Delivery System, Rx only, sterile, Stryker Instruments, Kalamazoo, MI. |  |  | February 13, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Mis-labeling; The instructions for use include a use for which the device is not intended. (Device not approved to capture a bone biopsy) | Labeling Change Control | Consignees were notified via Medical Device Recall Notification letter dated 2/13/08 to read the enclosed instructions for use, remove all old instructions for use at the consignee location and replace them with the new instructions for use, and to discard the old instructions for use. |  | 953 | 953 boxes Recalls Z-1192-1195-2008 | Nationwide. | Terminated | December 30, 2008 | 321 | 953
Z-0572-2008 | 45859 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech 12/14 Biolox Alumina Femoral Head 28MM +3.5, hip prosthesis component, Exactech, Gainesville, FL 32653 |  |  | November 01, 2007 | 2007 | Exactech, Inc. | Mislabeled: A ten piece lot of 140-32-03, 12/14 BIOLOXforte Alumina Femoral Heads 32mm (1069020 - 1069029), from Ceramtec were incorrectly labeled as 140-28-03 12/14 BIOLOXforte Alumina Femoral Heads, 28mm | Process control | Exactech issued a recall letter, titled Important Product Recall Notice on 10/31/2007 to the Distibutors that received the recalled product. It requested that all device recipients be notified and all product be returned to the firm. |  | 10 | 10 | Distributed to NY, OH and TN. | Terminated | May 06, 2012 | 1648 | 10
Z-1206-2008 | 45901 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Smith &amp; Nephew, Inc., Exu-Dry Wound Burn Dressing, Product Number: 5999M36, Exu-Dry Pad 24"x36" Non-Permeable Heavy Absorbency, Smith &amp; Nephew, Inc.,Largo, Florida. |  |  | November 21, 2007 | 2007 | Smith And Nephew, Inc. Wound Management Division | Sterility (package integrity) compromised: Distribution of product that did not meet specifications; open seals. | Package design/selection | Consignees were notified of recall by certified mail, Urgent Device Recall letter dated 11/21/2007, which was sent to each consignee. The recalled product will be returned to the firm by pre-paid FedEX return labels. A second letter is to be issued to each distributor for recall to the retail level, as the firm's original recall strategy of distributor level recall, was deemed inappropriate by CDRH |  | 5100 | 5,100 Dressings | Worldwide-USA, England, Japan, South Africa, New Zealand, Australia and Canada. | Terminated | December 15, 2008 | 390 | 47554
Z-0455-2008 | 38465 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott - CELL-DYN 1800 System with CELL-DYN 22 Calibrators and the CELL-DYN Calibrators and CELL-DYN 16 Controls. Part/Model Number: 07H77-01. Abbott Laboratories, Santa Clara, CA 95054-1113 |  |  | October 27, 2006 | 2006 | Abbott Laboratories | Incorrect Results: Hemoglobin results in donor samples assayed on the CELL-DYN 1800 System vary more than expected across the range of claimed operating temperatures of 18 degrees Celcius to 30 degrees Celcius. | Software design | A Product Correction letter was sent to all customers with a reply form on October 27, 2006. Product information letters are being packed with the instrument's accessory kit and with calibrators and quality control products used with the CELL-DYN 1800 System. Correction letters with reply forms have been resent to nonresponding customers. |  | 4224 | 4,224 Systems | Worldwide Distribution, USA and countries of Argentina, Bolivia, Brazil, Canada, Chile, Columbia, Costa Rica, Dominican Republic, El Salvador, Germany, Guatemala, Honduras, Hong Kong, China, Japan, Mexico, Paraguay, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | January 06, 2011 | 1532 | 4224
Z-0933-2008 | 46293 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Linemaster Wireless Footswitch Accessory, Model SP-9970214-005112010 and Model SP-9970214-005112232. Sound Surgical Technologies, LLC, Louisville, CO 80027. |  |  | December 28, 2007 | 2007 | Sound Surgical Technologies, LLC | Failure to shut-off: Surgical equipment, activated by a wireless foot switch with VASER system, may remain powered on when switch is no longer depressed. | Device Design | Consignees were sent an Advisory Notice on 12/28/2007. The notification describes the defect. Sound Surgical will be contacting facilities to arrange a service call to install new cables between the footswitch receiver and the VASER system. Facilities may contact the Technical Service Center at 1-888-677-8765. |  | 127 | 127 switches | Worldwide Distribution including USA states of AL, CA, CT, FL, GA, IN, KY, LA, MI, NJ, NY, OH, OK, PA, TX, and VA. Foreign Distribution to Colombia, Costa Rica and Switzerland. | Terminated | April 01, 2008 | 95 | 127
Z-0931-2008 | 46207 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Bayer Contour TS blood glucose monitoring system with 10 blood glucose test strips enclosed; Bayer Healthcare, Mishawaka, IN. Products 1802 and 1802N (U.S.), 1810 (Korea), 1811 (India, Korea and Mexico) and 1825 (Mexico). |  |  | December 21, 2007 | 2007 | Bayer Healthcare LLC | InaccurateTest Strip Results: Results in blood glucose readings with a positive bias are outside the product specifications. Patient test results may demonstrative results 5 percent to 17 percent higher. | Process design | The firm issued a Press Release on 12/21/07. A Test Strip Recall letter was also sent to the firm's consignees and health care providers instructing them that use of the product should cease and that health care providers contact the firm for replacements. Additional information can be found at www.bayerdiabeties.com or by contacting the Bayer at 1-800-348-8100. |  | 97000 | 97,000 bottles of test strips | Worldwide Distribution --- including USA and countries of Austria, France, India, Korea, Mexico, Philippines and Turkey. | Terminated | July 07, 2009 | 564 | 97000
Z-0717-2008 | 46349 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Mobile Fluoroscopy X-Ray System, OEC, Model 7700, GE Healthcare, Surgery, Salt Lake City, Utah |  |  | September 28, 2007 | 2007 | GE OEC Medical Systems, Inc | Unnecessary operator exposure; X-ray machine may allow unwanted x-ray exposure to the operator. | Other | On or about 9/6/2007, by Customer Notification Letter, users were notified that all affected units will be corrected by GE personnel with the addition of an additional shield. |  | 309 | 309 systems (Domestic), 101 (Foreign) | Worldwide- USA including states of CO, DE, FL, NJ, MO, NY, PA, and TX (with VA and Military facilities) and countries of Australia, Austria, Belgium, Brazil, Chile, China, The Czech Republic, Denmark, France, Germany, Hungary, Israel, Italy, Norway, Spain, Sweden, Switzerland, Thailand, and The UK. | Terminated | November 01, 2011 | 1495 | 466
Z-1051-2008 | 46742 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Draeger TI500. Isolette Infant Incubator. Catalog Number: MU20505. Draeger Medical, Inc., Telford, PA 18969 |  |  | January 31, 2008 | 2008 | Draeger Medical, Inc. | Heating failure- The power board which controls the incubator heater may not regulate the temperature properly resulting in a high temperature alarm and a loss of temperature control within the patient compartment. | Nonconforming Material/Component | The firm notified consignees on 2/1/08, by an Important Safety Notice. The letter instructed users to remove any units from use until a Draeger Service Representative can replace the power board. A Service Representative will be in contact to schedule a service visit. For questions concerning this recall please contact 215-660-2418 or Anbu.arasu@draeger.com. |  | 18 | 18 units | Worldwide Distribution-- USA including states MO, OH, and PR, and countries of Australia, Canada, Columbia, Israel, UAE, France, China, Italy, Denmark, Switzerland, and UK. | Terminated | May 08, 2008 | 98 | 24
Z-0571-2008 | 45798 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 |  |  | June 26, 2007 | 2007 | Ventana Medical Systems Inc | Incomplete upgrade: The software upgrade from February 1, 2007 may not have beencompleted. This software upgrade adds the algorithm (Slidetype) for thePATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. | Software change control | Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have beencompleted on their system. This software upgrade adds the algorithm (Slidetype) for thePATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketingand is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to beperformed remotely with the support of the Ventana Medical Systems Customer Support Center(CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule aTechnical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAYanti-HER-2/neu (4B5). |  | 77 | 77 devices | Nationwide and Canada | Terminated | March 26, 2012 | 1735 | 77
Z-0603-2008 | 44704 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew, CALAXO Bioabsorbable lnterference Screw, 9 mm x 25 mm, |  |  | August 21, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Patient Reaction: Post operative condition sterile fluid pocket has been identified, including the potential for graft failure and premature material degradation. | Nonconforming Material/Component | Smith &amp; Nephew issued an Urgent Product Recall letter via Federal Express to User Facilities, Sales Reps and Distributors on 8/21/07. A second letter was issued on 8/23/07 to correct two errors in model numbers listed in the original letter. Letter issued by FedEx and/or email to OUS contacts. |  | 6925 | 6925 units | Worldwide Distribution - USA, Austria. Belgium, Germany, Denmark, Spain, Finland, France, England, Italy, Portugal, Netherlands, Norway, New Zealand, Sweden, Turkey, Switzerland,China, Hong Kong, Malaysia, Taiwan, Singapore, Thailand, Australia, Canada, Mexico, Colombia, South Africa, Dubai, Venezuela, and Chile. | Terminated | May 09, 2012 | 1723 | 66263
Z-0620-2008 | 45479 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trilogy Acetabular System Shell with Cluster Holes, Porous, 60mm O.D., Tivanium TI-6AL-4V Alloy, Hip Prosthesis Component, REF 6200-060-22, Catalog # 00-6200-060-22, Zimmer, Inc., Warsaw, IN, 46580. |  |  | October 19, 2007 | 2007 | Zimmer Inc. | Parts missing and incomplete etch: The units may not contain the locking ring or the etched alignment marks on the rim. | Process control | Zimmer sent consignees a Product Recall Notification letter dated 10/19/07, requiring them to return the product. |  | 31 | 31 | Worldwide Distribution - USA (Arizona, California, Delaware, Indiana, Iowa, Kansas, Massachusetts, New York, Ohio, Pennsylvania, South Carolina, Texas, and Wisconsin), Australia, Canada, Germany, Italy, and Spain. | Terminated | May 21, 2008 | 215 | 31
Z-1108-2008 | 46362 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel Dxl Access Immunoassay Systems Reaction Vessels for Use with UniCel Dxl Access Immunoassay Systems, Part Number : 386167, Beckman Coulter, Inc, Brea CA 92822 |  |  | July 19, 2007 | 2007 | Beckman Coulter Inc | Erroneous Results - UniCel Dxl Access Immunoassay Systems Reaction Vessels have been found to be defective and there is a potential for an increase in signal leading to an erroneous result. | Process design | A Product Corrective Action (PCA) letter was mailed on July 19, 2007, to the Customers that received UniCel Dxl Accesslmmunoassay System Reaction Vessels lot 1201291 and1201431. The customers were informed that a smallpercentage of the reaction vessels (Lot #s: 1201291and1201431 ) have been found to be defective and there is apotential for an increase in signal leading to an erroneousresult. The customers were instructed to discontinue use anddiscard these lots. A response form was included. |  | 273 | 264 units in the US; 9 units in Canada | Nationwide &amp; Canada | Terminated | May 04, 2012 | 1751 | 273
Z-1109-2008 | 46387 | No | CGS | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Clinical Chemistry Creatine Kinase, Model LN 7D63-30, Lot Number: 52044HW00; Expires October 19, 2007, RCR#2018433-10/25/07-006R Abbott Laboratories, Inc., South Pasadena CA 91030 |  |  | October 12, 2007 | 2007 | Abbott Laboratories Inc. | Three potential issues with Creatine Kinase Lot Number: 52044HW00: (1) Decrease in quality control and/or patient results (2) An iIncreased imprecision, or (3) Error Code 1054 (Unable to calculate results, Reaction Failure) on Architect cSystems. | Nonconforming Material/Component | On October 12, 2007, A Product Recall letter was provided to all Clinical Chemistry Creatine Kinase Reagent customers that have received lot number 52044HW00 . Customer reply form were included in all US customers letter. The letter instructed customers to discontinue use of the kits and destroy them following their laboratory procedure, and document the number of kits discarded. Abbott will provide reimbursement for destroyed kits. |  | 498 | 334 Distributed Worldwide; 164 Domestically | Worldwide Distribution --- including USA and countries of Canada, Chili, Uruguay, Germany, Hong Kong, and New Zealand. | Terminated | March 28, 2012 | 1629 | 498
Z-1054-2008 | 46743 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex-Free Needle Guide, Sterile transrectal needle guide for use with Shimadzu UB1OR-065U transducer, Part #677-057, packaged 1 needle guide per tyvek pouch, 20 pouches per box. The label shows the product is Mfg by CIVCO, Kalona, IA. |  |  | January 16, 2008 | 2008 | Civco Medical Instruments Inc | Incorrect Needle Guide- The individual pouches contained V53W Endocavity Needle guide(p/n/034-054) instead of the Shimadzu T/R Needle guide (p/n134-054). | Packaging | On 1/16/08, the recalling firm contacted their customers and notified them of the problem. Pictures of the problems were provided for identification. The firm requested that all products be returned. |  | 207697 | Lot M102250 - 5 boxes; Lot M105440 - 2 boxes | Nationwide Distribution- including states of TX and GA | Terminated | April 03, 2008 | 78 | 207697
Z-0130-2008 | 38336 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS System, Surgical Laser System and Accessories. Part Numbers 0010-0070 and 0010-0075 (international units), AMS Innovative Center, San Jose, CA 95134. |  |  | April 17, 2007 | 2007 | AMS Innovative Center - San Jose | Component missing: Some products were issued without thermal protection switches. | Process design | AMS issued Important: Medical Device Correction letters, dated April 17, 2007, to its US and international consignees by mail, instructing them to return an acknowledgement form. The letter informs customers that some HPS systems may lack a motor overload switch, that the firm believes that there is no immediate safety issue, and installation of the MOL switch will be conducted during normal preventive maintenance. AMS expects to complete installation of the switches by end of August 2007. |  | 169 | 169 units | Worldwide Distribution - USA, Italy, France, Germany, Spain, UK, Denmark, Sweden, Australia, India, Japan, Greece, Switzerland, Saudi Arabia, Thailand, Turkey, Singapore, Taiwan/Hong Kong (China), and Austria. | Terminated | March 23, 2010 | 1071 | 169
Z-0646-2008 | 45854 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics; UHR Universal Bipolar Component; 28 x 56 mm, Hip Prosthesis Component, Stryker Howmedica Osteonics Corp., Mahwah, NJ 07430 |  |  | October 26, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Incorrect part: Device may have an incorrect locking ring assembled with the UHR bipolar head. | Process control | Stryker sent Urgent Product Hold letters to US branches/agencies and hospitals on October 26, 2007 advising them to quarantine the product from these lots, and Urgent Product Recall letters on November 2, 2007 advising them to return the devices to Stryker. International distributors were sent an advisory on November 6, 2007. |  | 360 | 360 total | Worldwide Distribution - USA, Canada, China, India, Australia, Korea, Singapore, Taiwan, Sweden, Germany, Netherlands, France, Spain, Italy, Poland, South Africa, and England. | Terminated | January 27, 2010 | 824 | 360
Z-1142-2008 | 46685 | No | LKK |  |  | General Hospital | PMA | Medtronic - SynchroMed EL Implantable Infusion Pump, Model Number: 8626-10. The contents of the inner package have been sterilized by ethylene oxide gas. 10 mL pump with reservoir valve. Medtronic, Inc., 710 Medtronic Parkway NE, Minneapolis, MN 55432-5604, USA. The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. |  |  | January 16, 2008 | 2008 | Medtronic Neuromodulation | Device/Drug Interaction - The company updated the labeling for the devices to include current patient management and treatment recommendations. The company received reports of inflammatory mass formations at or near the distal tip of intrathecal catheters which infuse opioids, baclofen, or chemotherapy drugs into patients. On January 16, 2008, Medtronic sent a letter to doctors who implant these d | Device Design | An Urgent Medial Device Correction letter was sent January 16, 2008, to Health Care Professionals. The letter describes the incidences, symptoms and recommendations for patient management. Excerpts from the approved Medtronic Professional Labeling are also included with the letter. The firm has requested unreported inflammatory mass in a patient with a Medtroinc device to the firm and to the FDA MedWatch Program by phone at 1-800-FDA-1088. Additional assistance may be obtained by contacting Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 102792 | ~102,792 worldwide. 90,330 within US and 12,462 OUS | Worldwide Distribution -- USA including states of Washington D.C., Puerto Rico, and countries of Aruba, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Faroe Islands, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Mexico, Netherlands, Netherlands Antilles, New Caledonia, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State. | Terminated | June 06, 2011 | 1237 | 925128
Z-1044-2008 | 46245 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Harmony sygno. Model Number: 7104693, Siemens Medical Solutions USA, Malvern PA, 19355 |  |  | December 01, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Magnet Quench, unintended- magnet of a mobile MRI system quenched into the exam room. | Device Design | The firm issued an Update Instruction package to consignees 12/07. As an immediate measure, the recalling firm is initiating field checks of mobile MRI systems per Update Instruction MR039/07/S. This update defines a safety inspection for mobile systems, including checks for potential ice formation and proper configuration of venting assemblies. |  | 10 | 10 units | Nationwide Distribution | Terminated | October 01, 2012 | 1766 | 210
Z-1367-2008 | 45950 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro cartridge blood tubing set used with Gambro Dasco Phoenix and Cobe Centrysystem 3 (and 3+) Hemodialysis System, Gambro, Lakewood, CO 80401 |  |  | September 10, 2007 | 2007 | Gambro Renal Products, Inc. | Kinked tubing on hemodialysis device may cause hemolysis. | Device Design | On September 10, 2007, the firm sent an Advisory Notice to all consignees alerting them that a kink in the tubing can result in physical hemolysis. The firm included with the Notice a Quick Reference Guide that provides drawing of the proper installation of the Phoenix and describes how to prevent bloodline kinking. A customer reply form was also included. The firm is in the process of making labeling changes that will be released with the updated software revisions. |  | 41000000 | 41,000,000 | Nationwide, including 17 VA facilities and 1 military facility. Foreign distribution to Canada, Korea, Mexico. | Terminated | January 03, 2012 | 1576 | 41000000
Z-1297-2008 | 46746 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Locking Screw, T7, Catalog Number: 40-27622, Size: 2.7X22MM, Non Sterile,Titanium Alloy; Manufactured and Distributed by: Stryker Leibinger GmbH &amp; Co. KG;Boetzinger Strasse 41; 0-79111 Freiburg, Germany |  |  | January 28, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Screws do not have locking threads | Process control | Recall notification letters (Urgent Product Recall) were sent on January 28, 2008 to all Stryker Branches/Agencies by Federal Express. All product distributed was requested to be located and returned, including from hospitals. |  | 100 | 100 units | Nationwide-USA including states of MA, PA, GA, OH, and TX (hospital consignees) and Stryker Branches/Agencies in MA, NJ, PA, MD, VA, NC, GA, TN, KY, OH, MI, LK, OK, TX, and CA. | Terminated | June 04, 2008 | 128 | 100
Z-1298-2008 | 46780 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | AlloCraft DBM with Syringe Assembly, 5CC, Catalog/Reorder Number: 653005, LifeCell Corporation, Branchburg, NJ 08876 , Distributed by Stryker Spine |  |  | October 17, 2007 | 2007 | LifeCell Corporation | Improper donor screening test; diagnostic test kits were used in lieu of a donor screening test kits for HBsAG and HBcAB (later re-testing with FDA-approved donor screening test found samples non-reactive for both Hepatitis B Surface Antigen and Hepatitis B Core IgM/IgG Antibody) | Release of Material/Component prior to receiving test results | Consignee was notified by telephone, followed by a Recall letter sent by FAX on October 17, 2007 requesting return of any unused product. |  | 46 | 46 kits | All syringe assemblies were distributed to Stryker Orthopaedics, Mahwah, NJ. | Terminated | May 14, 2008 | 210 | 156
Z-1352-2008 | 46307 | No | MCB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Triage Micro ClostrIdium difficile (C.difficile) Panel, Catalog No: 96000, BIOSITE INCORPORATED, San Diego, CA 92121 |  |  | December 04, 2007 | 2007 | Biosite Inc | False Lower Results: Toxin A results in the Triage Micro C. difficile Panel have shown falsely lower visual signals (color strength) in certain lots compared to results observed with previous lots. | Process design | Consignees located in the United States received an Urgent Medical Device Recall Letter notification by United Parcel Service on 12/4/2007 and 12/14/2007. Consignees outside of the United States were notified by Federal Express on 12/4/2007 and 12/14/2007. The Customer Recall Letter with attached fax-back Verification form informed customers of the recall and instructed customers to dicontinue use, discard appropriately all affected kits and complete and FAX the enclosed Verification form. Contact information was also provided form the recalling firm. (USA-(800) 640-0640 ) |  | 4549 | 4,549 kits | Worldwide- USA, Argentina, Australia, Belgium, Canada, Switzerland, Colombia, The Czech Republic, Germany, France, Greece, Guatemala, Hong Kong, Croatia, India, Italy, Mexico, Portugal, Sweden, Turkey &amp; The United Kingdom | Terminated | March 20, 2012 | 1568 | 4549
Z-1351-2008 | 46214 | No | JJM | Osmometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Osmocoll, COP Control Reference, SS-038, Contents: 6 unit packages of 1 mL vials (3 each High and Low), Wescor, Inc., 459 South Main Street, Logan, Utah 84321. |  |  | December 26, 2007 | 2007 | Wescor, Inc | Inaccurate readings; The control does not provide an accurate calibration value/readings when performing Colloid Osmotic Pressure (COP) testing. | Nonconforming Material/Component | Consignees were notified by Recall Notification letter on 01/11/2008, and asked to arrange for return of affected product. |  | 75 | 75 vials | Worldwide- USA, Argentina, Australia, Brazil, Canada, Denmark, England, France, Germany, Greece, Hungary, Japan, Malaysia, Mexico, The Netherlands, New Zealand, Paraguay, Poland, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and The United Kingdom. | Terminated | March 27, 2008 | 92 | 1025
Z-1360-2008 | 46783 | No | CJE | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Liquid Alkaline Phosphatase Reagent Set; Catalog No. A7516-150-SAM, In vitro diagnostic; for the quantitative determination of alkaline phosphatase in serum, Sold by Pointe Scientific, Inc |  |  | February 07, 2008 | 2008 | Pointe Scientific, Inc. | May be contaminated with microorganisms. | Other | Consignees were notified via recall letter, Urgent Medical Device, dated 2/7/08, to cease using the product, conduct a recall to the user level, destroy product on hand and to notify the firm via the response form of the amount destroyed for replacement purposes. International consignees were notified via email. |  | 6 | 6 | Worldwide-USA, Canada, Chile, Cyprus, Ecuador, Greece, India, Italy, Jamaica, Korea, Nigeria, Pakistan, Philippines, Poland, Turkey and The West Indies. | Terminated | June 12, 2009 | 491 | 302
Z-1361-2008 | 46706 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Vanguard Width Checker, Part Number: 32-487062, 67.5/75 Sizing Wing, Biomet Orthopedics, Inc., Warsaw, IN; The device is a preparatory instrument used for femoral knee insertion. |  |  | February 12, 2008 | 2008 | Biomet, Inc. | The instrument's sizing line is in the wrong place. | Process control | Distributors were notified via Urgent Medical Device Removal Notice letter dated 2/12/08 to locate the instruments in their possession and at hospital sites and to return them to Biomet. Hospitals with the instruments on hand will be provided with a recall letter dated 2/12/08 which describes the nature of the problem and the removal of these instruments. The Puerto Rico consignee was notified by Biomet UK, Ltd. via letter on 2/26/08. |  | 35 | 35 (31 to U.S. and 4 international). | Worldwide: USA, Canada, China and Finland. | Terminated | June 26, 2009 | 500 | 36
Z-1311-2008 | 46869 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Cardiovascular Procedure Kit, P/N 64789-01, Wisconsin Cardiovascular Services Pump and Table Pack, Terumo Cardiovascular, Ashland, MA. (the Pump &amp; table pack contains a weighted flexible sucker, Catalog No. 804113) |  |  | January 29, 2008 | 2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient. | Nonconforming Material/Component | Consignees were notified via recall letter, Urgent Medical Device Removal, dated 1/29/08 via fax or email to cease using the weighted sucker in the packs and to use an alternative sucker until replacement product arrives from Terumo. |  | 110 | 110 | Nationwide-USA including states of Missouri, Tennessee and Wisconsin-Custom Kits for 3 hospitals. | Terminated | July 07, 2009 | 525 | 1121
Z-1365-2008 | 47341 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i Certain MicroMiniplant Straight Healing Abutment, Model/Item Number: ISMHA32, 3.4mm(D) X 0mm(P) X 2mm(H), BioMet 3i, Palm Gardens, FL. |  |  | February 06, 2008 | 2008 | Biomet 3i, Inc. | Seating Problems: The Micro-miniplant Straight Healing Abutments do not fully seat onto the restorative platform of the implant leaving a gap between the abutment and the implant. | Device Design | USA customers were contacted by telephone with a follow-up fax. International customers were first contacted by E-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i. |  | 270 | 270 | Worldwide | Terminated | July 07, 2008 | 152 | 270
Z-1050-2008 | 46351 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Physio-Control LIFEPAK 20 defibrillator/monitor, AC powered external defibrillator with battery backup. Part Numbers: 3202487-073, 3202487-079, 3202488-036, Physio Control, Inc, Redmond, WA 98052 |  |  | January 08, 2008 | 2008 | Physio Control, Inc. | Inoperable Device - Potential for delay in therapy or prevention of defibrillation therapy due to corrosion of the Printed Circuit Board Assemblies (PCBA). | Nonconforming Material/Component | The firm contacted Customers by phone and by Urgent-Medical Device Recall letters sent via return receipt requested to their consignees on 1/8/08. The letter advised consignees that the Printed Circuit Board Assemblies (PCBA) might not meet quality requirements. This manufacturing deviation could lead to eventual PCBA corrosion that could render the device inoperative and not ready for use in an emergency under battery power. The firm recommended keeping the unit in service until it could be replaced. Representatives will visit each consignee and replace units. |  | 103 | 103 total devices - 84 domestic, 19 international | Worldwide Distribution --- USA including states of OK, SC, TX, and WA, and country of Germany. | Terminated | September 04, 2013 | 2066 | 103
Z-1366-2008 | 47342 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i, EP Healing Abutment, REF/Item Number: THA53, 4.1mm(D) X 5mm(P) X 3mm(H), Sterile, BIOMET 3i, Palm Beach, FL |  |  | February 14, 2008 | 2008 | Biomet 3i, Inc. | Mis-labeling: Labeled THA53 device packaging may contain ITHA53 devices (The ITHA53 is designed to be used with an internally hexed implant--The THA53 is used with an external hex platform) | Process control | Each USA customer was first contacted by telephone and then with a faxed, Urgent Medical Device Recall letter, dated 2/25/2008. Each International customer was first contacted by E-mail and then with a faxed letter. They were instructed to return the affected product to Biomet 3i. |  | 254 | 254 | Worldwide | Terminated | July 10, 2008 | 147 | 254
Z-1349-2008 | 46889 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa IMPAX CardioVascular Suite Results Management Nuclear Reporting Module, Outbound Report, Model/Catalog No.: L7W1P, Software Versions: 2.04.17, 2.04.19 and 2.04.20, AGFA HealthCare, Westerly, RI |  |  | February 22, 2008 | 2008 | AGFA Corp. | Reporting Differences: The Outbound Report software produces an Outbound Report (text version) which does not contain all the clinical content elements present in the PDF version of the report, also generated on the Nuclear Cardiology Reporting Module. | Software design | All five (5) sites utilizing this product were notified about the issue directly via telephone. An "Urgent Safety Notice" was sent via FED-EX to all the sites using the product on February 22, 2008. The letter describes the potential issue and mitigation. The recalling firm will also visit each site and make configuration changes to resolve this issue. |  | 5 | 5 units | Nationwide to five hospitals in DC, NY, OH, TN, and VA. | Terminated | July 27, 2010 | 886 | 5
Z-1383-2008 | 46882 | No | ODW | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | Warfarin Metabolism Nucleic Acid Test Cartridge; Part Number: 20-006-002; an in vitro diagnostic drug metabolizing enzyme genotyping system; Nanosphere, Inc., 4088 Commercial Ave., Northbrook, IL 60062 |  |  | February 25, 2008 | 2008 | Nanosphere, Inc. | Aberrant Results: A warfarin 2C9*2 mutant capture signal was found to be high enough to result in an aberrant result (i.e., wild-type called heterozygous). | Other | Nanosphere sent Urgent Device Recall letters dated 2/25/08 to the direct accounts on the same date, informing them of the potential for aberrant results with the use of the affected lots of cartridges. The accounts were requested to examine their stocks for the affected lots of cartridges, segregate those found and return them to Nanosphere for replacement. Any questions can be directed to Dr. Gregg Shipp at 847-400-9115. |  | 624 | 624 units | Nationwide; USA including states of Kentucky, Michigan, Tennessee, Texas, Louisiana, Connecticut, Illinois, and California. | Terminated | May 17, 2010 | 812 | 624
Z-0617-2008 | 45444 | No | HDD | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity 4 Birthing Bed; Model P3700B, Hill-Rom, Inc., Batesville, IN 47006 |  |  | September 06, 2007 | 2007 | Hill-Rom, Inc. | Brake failure: The brakes may not hold or lock properly. | Component design/selection | Hill-Rom notified consignees via Urgent Medical Device Correction letter dated 9/6/07. The letter instructed facilities to continue using the beds if they pass a test of the brake mechanism; and a Hill-Rom representative will contact the facility to schedule a replacement of the detent mechanism and adjustment of the casters. |  | 1796 | 1,796 | Worldwide Distribution - USA, Argentina, Australia, Bahamas, Belgium, Belize, Brazil, Canada, China, Denmark, Dominican Republic, France, Germany, Guam, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Jordan, Kenya, Malaysia, Mexico, Netherlands, Norway, Philippines, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Syria, Turkey, United Arab Emirates and United Kingdom. | Terminated | October 15, 2010 | 1135 | 1796
Z-1107-2008 | 46350 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker ACL Workstation slider, Model Number: 234-010-503, Stryker, San Jose, CA 95138 |  |  | December 07, 2007 | 2007 | Stryker Endoscopy | Slider Lock Failure - The ACL Workstation slider, a component of the ACL workstation, does not consistently lock into the ACL Workstation Base. | Component design/selection | On January 9, 2008, sales representatives were notified via telephone. All consignees were mailed an Urgent: Device Removal notice, dated January 10, 2008. The firm requested that any affected inventory in stock be quarantined and returned to Stryker Endoscopy for destruction. A response form with return postage was included with notification mailed to all consignees. |  | 142 | 142 sliders were distributed. | Worldwide Distribution - including USA states of South Carolina, Oregon, Michigan, Nebraska, Maryland, California, New Mexico, Ohio, Texas, West Virginia, Massachusetts, and Florida and countries of Canada, Mexico, Hong Kong, Argentina, Brazil, France, Germany, Greece, Spain, India, Italy, Columbia, Switzerland, Poland, Sweden, Australia, and Netherlands. | Terminated | December 29, 2008 | 388 | 142
Z-1400-2008 | 46940 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncor Expression Medical Linear Accelerator, Material Number: 07360717, with control console software versions 6.5/7.2, 8.0, 9.0 and 9.1, Manufactured by Siemens Medical Solutions USA, Inc, Concord, CA |  |  | January 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Linear accelerator gantry may rotate in an unexpected direction and may cause patient injury on impact. | Software design | This is a mandatory software upgrade for consignees. The firm's representatives have contacted consignees via Urgent Device Notification letters dated February 14, 2008, notifying them of the new update and making arrangements to install it. |  | 5 | 5 units | Worldwide-USA including states of AL, AK, AZ, AR, CA, CO, DE, DC, FL, GA, IL, IN, KS, KY, LA, MD, ME, MI, MN, MS, MO, MN, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI, and WY and countries of Angola, Australia, Austria,Belgium,Brazil,Canada,Chile,Colombia,Costa Rica, Croatia,Cyprus,The Czech Republic,Denmark,DPR of Korea,El Salvador,France,Germany,Greece,Hungary,India,Iran,Ireland,Italy,Japan,Jordan,Malaysia,Mexico, The Netherlands,New Zealand,Norway,P.R. China,Pakistan,Peru,Philippines,Poland,Portugal,Republic of Korea,Romania,Russian Fed.Saudi Arabia,Singapore,Slovakia,South Africa,Spain,Sweden,Taiwan,Thailand,Trinidad,Tobago,Turkey, U.A.E.,Ukraine, The United Kingdom, and Venezuela | Terminated | May 06, 2011 | 1211 | 3292
Z-0673-2008 | 45800 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Parks Medical Dual Frequency Doppler Vascular Flow Detector, Model 915-BL (finished assembly without accompanying probes), Parks Medical Electronics, Inc., Aloha, OR 97007 |  |  | November 08, 2007 | 2007 | Parks Medical Electronics Inc | Premature battery failure: Wrong acid flux used on circuit board which may lead to premature battery failure of ultrasonic non-directional doppler flowmeter. | Process control | 811-B: On 9/28/07 Parks Medical Electronics notified their related U.S. distributor, Parks Medical Electronics Sales, Inc., Las Vegas, NV by email instructing them to stop further shipments. On 10/1/07 firm followed up with phone call instructing all units be returned for rework. On 11/8/07 firm sent replacement units accompanied with a Warranty Exchange Letter to their other U.S. customers notifying them of affected units. Letters were customized for each customer with a fill-in-the-blank of the model and serial numbers. Letters and replacement units were sent UPS 2nd Day Air and informed customer about the assembly error and instructions for returning the defective units. On 11/8/07 emails were sent to Brazilian and Australian distributors informing about assembly error and affected units. Rework instructions and a schematic were provided by email for each distributor to perform their own rework. On 11/28/97 Canadian customers were sent Warranty Exchange Letters and replacement units via UPS Ground.Firm is retrieving units from US and Canadian firms and will rework them. Firm is waiting to hear whether Brazilian or Austraian distributors need to sub-recall. Canadian distributors instructed to conduct subrecall if needed.All returned units will be reworked by disassembly of unit, removal of affected printed circuit board, washing board to remove acid flux, and testing. |  | 5 | 4 units distributed to 1 U.S. hospital | Worldwide Distribution - USA, Brazil, Australia, and Canada | Terminated | October 13, 2011 | 1435 | 3176
Z-1364-2008 | 47340 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i LOCATOR ABUTMENT, Catalog/Item Number: LOA003, 3.0mm, BioMet 3i , Palm Beach, FL. |  |  | February 07, 2008 | 2008 | Biomet 3i, Inc. | Mispackaging: The Locator Abutment labeled as LOA002 may contain Catalog/Item # LOA003 product and visa versa | Other | Each USA customer was contacted by telephone and then by letter, Urgent Medical Device Recall, dated 2/15/2008, sent by fax. International customers were first contacted by E-mail then by fax letter also. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory for product return. They were also instructed to return the affected product to Biomet 3i. |  | 97 | 97 | Worldwide. | Terminated | July 10, 2008 | 154 | 197
Z-0713-2008 | 45907 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | Coloplast Elefant Suction/Irrigation Cannula for Laparoscopy |  |  | November 12, 2007 | 2007 | Coloplast Corp | Lack of sterilization assurance: Sterilization of Elefant Suction-Irrigation Devices is not performed according to ISO 11135 standard. | Process design | Coloplast sent customers an "Urgent: Recall Notification" letter on 11/12/07. The letter described that the sterilization of the products was not performed according to ISO 11135 standard. The notification requested the customers to return the product immediately and to send a fax or contact the company via email in regard to inventory at hand. A phone script for distributors also included an advice to discontinue dispensing the product and to fax or email in regard to inventory at hand. |  | 154 | 154 | Nationwide, including Puerto Rico. | Terminated | November 04, 2008 | 358 | 717
Z-1384-2008 | 46912 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Baxter HEP LOCK (Heparin Lock Flush Solution, USP) 10 u/mL-1mL vial, 25 Pack, NDC 0641-0392-25 (NDC 0641-0392-21; on vial), Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | February 28, 2008 | 2008 | Baxter Healthcare Corp. | Due to an increase of adverse events in Baxter single and multi-dose Heparin for injection products, which used the same active pharmaceutical ingredient (API) source as the HEP LOCK products | Nonconforming Material/Component | A Press Release was issued on February 28, 2008. Recall notification letters , Urgent Product Recall, were sent on February 29, 2008, by overnight mail to customers/wholesalers/ distributors and dialysis center/renal home patients. Recall letters were also sent on March 7, 2008, to health care professionals |  | 335860 | 335,860 vials | Worldwide Distribution --Worldwide: USA including Puerto Rico, and countries of Germany, Guam, America Samoa, Marshal Islands, Northern Mariana Island/Virgin Islands, and Qatar . Accounts included wholesalers, pharmacies, hospitals, and clinic/renal home patients | Open, Classified |  |  | 4603297
Z-0711-2008 | 38631 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ConMed Linvatec, Ultrapower Burs, Linvatec Corporation, Largo, FL 33773 |  |  | August 09, 2007 | 2007 | Linvatec Corp. | Compromised sterility: ConMed Linvatec medical devices packaged within a certain Peel Pouch Sterile Package may have compromised sterility due to an improper seal. | Packaging process control | Conmed Linvatec sent their consignees Urgent New Recall Medical Device Recall Notification letters On August 9, 2007 by UPS overnight delivery. The letter requests the customer return the devices to Linvatec for credit. Domestic Distributors / Sales Representatives were notified by telephone and email. |  | 1181 | 1181 units | Worldwide Distribution. | Terminated | October 01, 2009 | 784 | 1181
Z-0712-2008 | 45503 | No | JXK | Cranial electrotherapy stimulator. | Neurology | Neurology | 510(k) | CES Ultra, Cranial Electrotherapy Stimulator, Serial Numbers: 1000 through 3999, Manufactured by Arco Electronics Limited, Distributed by Neuro-Fitness LLC, Fall City, Washington 98024 |  |  | October 29, 2007 | 2007 | Neuro-Fitness LLC | No FDA clearance for the .35Hz/.45Hz frequency option on device. | Finished device change control | Consignees were notified of this recall by an Urgent Device Recall Letter sent via U.S. First Class mail on 10/29/07. The letter instructs users to contact the recalling firm to make arrangements to return the device. The letter also instructs users to return an enclosed response card. Consignees who do not respond will receive a follow-up telephone call two weeks after initial mailing of recall notice. For more information, contact firm at 1-866-937-4237. |  | 3000 | 3,000 units | Worldwide Distribution | Terminated | September 21, 2011 | 1423 | 3000
Z-0689-2008 | 46353 | No | MIH |  |  | Cardiovascular | PMA | Medtronic brand AneuRx AAAdvantage Endovascular Stent Graft with the Xcelerant Delivery System; |  |  | December 10, 2007 | 2007 | Medtronic CardioVascular | Sterility compromised: Device sterility may be compromised as evidenced by a loss of outer package integrity. | Package design/selection | Medtronic sent Urgent Communication - Medical Device Correction Advisory Notice letters to all its direct consignees on 12/7/07, via Federal Express, informing them of the affected product and providing instructions on avoiding contamination of the sterile surgical field with these devices. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 28, 2008 | 231 | 7000
Z-1317-2008 | 46398 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan OneTouch Data Management Software Version 1.0, Part Number: 021-722-01(for OneTouch Select and UltraMini (Download) Blood Glucose Meters), LifeScan Inc, Milpitas, CA 95037. |  |  | December 18, 2007 | 2007 | Lifescan Inc | Meter Temporarily Freezes-- A software compatibility issue may cause the blood glucose meter to cease operations and freeze temporarily. | Software design | The firm sent an Urgent Recall letter to Health Care Professionals on 12/18/2007. The letter advised user to uninstall software drivers and provided options on how to uninstall. For further assistance, please contact LifeScan Customer Service at 1-866-679-9250. |  | 260 | 260 units | Nationwide Distribution including the states of AL, CA, IN, IO, MA, MD, MI, MS, NE, NY, NC, OR, PA, TN, TX, UT, VT, VA, WV, and WI. | Terminated | July 18, 2008 | 213 | 260
Z-0585-2008 | 46251 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific SlimLine 200 Micron Holmium Fiber Delivery Device, Catalog No. 840-840, Sterile, Made In Israel, Manufactured By LUMENIS, P.O.Box 240, Yokneam, ISRAEL. |  |  | December 18, 2007 | 2007 | Lumenis, Inc. | Incorrect Label: The Units outer pouch is labeled 200 micro fiber however, the inner pouch's label affixed to the product indicated that it is a 365 micron fiber. The micron fiber contained in the inner pouch is 200 micro fiber. The inner product is mislabeled with the wrong size. | Labeling mix-ups | The firm notified Consignee by certified letter on 12/18/2007, and asked that all affected products be return. Any questions concerning this recall can be addressed by contacting the Manager, Regulatory Compliance, at 1-877-586-3646 ext. 3355. |  | 59 | 59 units. | MA | Terminated | April 09, 2008 | 113 | 59
Z-1449-2008 | 46792 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Hemovac Wound Drainage Device Infection Control Kits, Catalog Number/ REF 00-2550-002-10, 400 ml-INF. Cntrl. Kit w/ 1/8" drain, 3.2 mm, 10Fr, PVC, 107 cm, medium, sterile, latex free, Rx only; Sold in packages of 10, which contain packages of 1 and 5; Zimmer U.K. Ltd., Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio |  |  | January 31, 2008 | 2008 | Zimmer Inc. | Leakage: The kits may disassemble at the fluid collection port, which would present a risk of exposure to blood borne pathogens to health care providers should it occur during use. | Process control | Consignees were notified via Urgent Infection Control Recall letter dated 1/31/08 to examine their inventory and return the recalled lots. Custom kit packers were instructed in the same letter to recall these products/lots from their own customers. |  | 34354 | 34,354 (Recalls Z-1448/1458-2008) | Worldwide-USA, Australia, Canada, Colombia, Denmark, Dubai, Italy, Japan, Mexico and Singapore. | Terminated | November 12, 2009 | 651 | 34354
Z-0526-2008 | 45880 | [6 empty] | BD FACS Sample Prep Assistant II, Part Number: 337170; |  |  | October 02, 2007 | 2007 | BD Biosciences | Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance. | Software change control | On 10/1/07, all consignees were notified via Important Product Information letter and also an on-site visit by the firm's sales representative. The letter informed users of the affected product and provided instructions on the recall. For additional concerns or questions contact BD Customer Support Center at 1-877-232-8995. |  |  |  | Worldwide Distribution --- USA including states of: NJ, IL, TX, PA, NY, MN, and AZ, and countries of Australia, Belgium, Hong Kong, Trinidad and Tobago.Australia, Belgium, Hong Kong, Trinidad and Tobago. | Terminated | December 13, 2010 | 1168
Z-0587-2008 | 46317 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | MyelotecÂż Steerable Video Guided Catheter, 3.0mm, Label Insert contained in LB0020-0, LB0020-1, LB0020-2, LB0020-3, and LB0020-4 (Product Code: 2010). Myelotech, Roswell, GA 30076. |  |  | December 19, 2007 | 2007 | Myelotec, Inc. | Incorrect Labeling information: Product may contain inappropriate information in its label insert that refers to the product's use with energy delivering instrumentation. | Error in labeling | Consignees were first notified by phone on 12/19/2007. A follow up Urgent Label Insert Recall Notification letter was sent via certified mail on/about 12/28/2007. Consignees were instructed to remove the subject label inserts from the affected product in inventory, replace with the enclosed revised label inserts and to complete the attached Recall and Accountability Form. The incorrect label inserts were to be returned, via package provided to Myelotec. |  | 1252 | 1,252 shipping containers | Worldwide Distribution --- USA including states of AL, CA, CT, FL, GA, ID, IN, ME, MN, NH, NJ, NY, OH, and OR, and countries of Aurba, Korea, Japan, Switzerland, South Africa, UK. | Terminated | December 03, 2008 | 350 | 1252
Z-0916-2008 | 46380 | No | DYH | Bracket adhesive resin and tooth conditioner. | Dental | Dental | 510(k) | Grengloo (Unidose Tips), Bracket adhesive and tooth conditioner, Part Number 740-0322, Lot Number 2792584, Ormco Corporation, Glendora, CA 91740 |  |  | October 31, 2007 | 2007 | Ormco Corporation | Mislabeled: Product mislabeled with incorrect expiration date. Expiration date labeled as 2010/07 instead of 2009/07. | Process design | Ormco sent fourteen (14) consignees Urgent Device Recall Notification via U.S. First Class mail on October 31, 2007. There were two (2) additional letters sent to the consignees in France that were translated into French. The consignees were instructed to complete the Return Form and return any affected product in their inventory. Periodic effectiveness checks will be conducted in order to follow up with non-responders. For more information, contact firm at 1-800-854-1741. |  | 4 | 4 kits (Total US and Internationally) | Worldwide Distribution-USA states of IL, NY, LA, CA, NJ, NV and SC and countries of Canada, France and South Africa. | Terminated | March 22, 2012 | 1604 | 4
Z-0843-2008 | 45812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Nonin Model 2500 PalmSAT handheld pluse Oximeter , Nonin Medical, Inc., Plymouth, MN USA |  |  | October 11, 2007 | 2007 | Nonin Medical, Inc | Incorrect Label: The label on the back of the device may show Model 2500A rather than Model 2500. | Error in labeling | An Urgen Field Action letters was sent (10/11/07) to Nonin distributors who received the potentially mislabeled product. The letters provided labeling differences and included instructions for notifying their customers, with a response form to be returned to Nonin. For additional information please contact Kim Aves at 763-577-3196. |  | 240 | 240 | Worldwide Distribution ---- USA including states of ME, SC, SD, CA, MO, NY, MI, FL, IL, PA, NC, CT, and TX, and countries of Canada, Greece, Russia, Italy, Sweden, Australia, &amp; Japan. | Terminated | December 18, 2011 | 1529 | 240
Z-0892-2008 | 46151 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Peel-Away Introducer set, 7.0 FR; Catheter introducer, includes sheath: 7.0 FR., 15.5 CM long, Wire: .035" DIA., 50 CM long, 3 MM J needle: 18 ga., 7.0 cm long. Accepts devices up to 7.0 French,sterile, Catalog Number: C-PLI-7.0-35. Cook Inc., Bloomington, IN 47404 |  |  | December 07, 2007 | 2007 | Cook, Inc. | Packaging Problem: The peel-away sheath does not peel uniformly or completely. | Package design/selection | Cook Inc. initially notified consignees via FEDEX by an Urgent: Product Recall on 12/7/07. The recall was then expanded and consignees were notified on 12/18/08 by another Urgent: Product Recall. The letters instructed any distributors to inform their customers of recall, and all users to return the recalled products and enclosed Recall Response Form to the firm. For additional information, contact 1-800-346-2686. |  | 1881 | 1,881 of total products | Worldwide Distribution- including USA and countries of Argentina, Australia, Canada, Chile, Denmark, Japan, New Zealand and Singapore. | Terminated | April 15, 2008 | 130 | 1881
Z-0915-2008 | 46347 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 ml Syringe Luer Lok Tip, Reorder #3039605, BD, Franklin Lakes, NJ 07417. |  |  | January 07, 2008 | 2008 | Becton Dickinson &amp; Company | Defective Seal: The convenience trays may have open seals, which can adversely impact tray sterility. | Packaging | BD notified distributors by an Urgent Product Recall Letter sent via UPS on January 7, 2008. In addition, the firm sent an Urgent Product Recall Letter on January 11, 2008 to customers identified by the distributors. The letter instructed the users to remove the product from its inventory and to return affected products to the firm using an enclosed customer packing list. For additional information, contact 1-888-237-2762. |  | 2926560 | 2,926,560 | Nationwide, including the states of CA, NY, MO, MA, IL, OH, TN, TX, MI, LA, MO, ID, GA, NE, SC and FL. | Terminated | May 19, 2008 | 133 | 4680660
Z-0668-2008 | 46275 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker GMRS Proximal Femoral Surgical Protocol; |  |  | December 18, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The GMRS Proximal Femoral Surgical Protocol states " The GMRS Proximal Femoral components are fully compatabile with all Stryker V40 femoral heads." This statement is incorrect and conflicts with the Alumina Head label and packaging insert. | Labeling design | Stryker sent Urgent Product Recall notification letters and Product Accountability forms via FedEx on 12/21/2007, return receipt requested. An Urgent Product Recall letter dated 1/7/2008, was sent to surgeons who are GMRS users, advising that the GMRS Proximal femoral components are not compatible with Alumina ceramic heads. |  | 5869 | 5869 | Worldwide Distribution - USA, Canada, Mexico, China, Australia, Hong Kong, Singapore, Sweden, Germany, Brussels, the Netherlands, Italy, Poland, Dubai, United Kingdom, Spain, and South Africa. | Terminated | October 15, 2008 | 302 | 5869
Z-0575-2008 | 45892 | No | NVI | Amphetamine test system. | Toxicology | Immunology | 510(k) | BioRad BioPlex 2200 Software, for BioPlex 2200 Instrument, an automated immunodiagnostic analyzer, Bio-Rad Laboratories, Inc., Hercules, CA 94547 |  |  | November 13, 2007 | 2007 | Bio-Rad Laboratories Inc | Incorrect values: An error was found in the BioPlex 2200 Software, Version 2.0 that results in assignment values being utilized for the BioPlex assay Calibrator Set that may be slightly different from those printed on the Value Assignment Data Sheet. Some values may be identical between the lot assignment combinations with the ANA Screen with MDSS kit. | Software design | Bio-Rad sent a Voluntary Field Correction notice dated November 1, 2007. Consignees were notified by phone call on Nov. 13, 2007, followed up by facsimile (domestic). Sub-recall to be by e-mail message for consignees outside the US. Distribution of the product ceased on Nov. 2, 2007. Bio-Rad will provide Calibrator Set CDs with only one lot of calibrator assignment values until a new version of software is available. |  | 7333800 | 85 units of calibrator lot 86280 and 16 units of calibrator lot 86290 were distributed. | Worldwide Distribution - USA, Canada, France, and Germany. | Terminated | July 23, 2008 | 253 | 7333800
Z-0988-2008 | 46341 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker Round Fluted Bur (6.0mm) Extra-Long for use with TPS Universal Drill Extra Long "R" Attachment, sterile, Stryker Instruments, Kalamazoo, MI; REF 5160-10-60. |  |  | January 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Lack of assurance of sterility. | Package design/selection | Stryker sent a Medical Device Recall Notification letter dated 1/10/08 to consignees. The letter instructed them to check their facility for product, cease using it, and to return the product using a mailing label provided. International subsidiaries were notified via email. |  | 10588 | 10,588 of all products | Worldwide Distribution - USA, Australia, Brazil, Canada, England, France, Germany, Hong Kong, India, Italy, Japan, Mexico, the Netherlands, New Zealand, Portugal, Saudi Arabia, Poland, Romania, South Africa, Spain, Sweden, Switzerland and United Arab Emirates. | Terminated | December 01, 2009 | 691 | 10588
Z-0661-2008 | 45410 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Sequoia Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system. Siemens Medical Solutions USA, Inc., Mountain View, CA |  |  | September 25, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Measurement Error affecting all Sequoia systems using the PAL video standard configuration and have a perspective PC installed: 1) The Sequoia" system is operating at a software revision below 8.0. 2) The Sequoia" system is configured to PAL video mode (normal for systems running at 240V 50 Hz.). 3) While using any of the Perspective" advanced display options listed above, an image or clip is t | Software design | Consignees were notified via a Urgent Medical Device Notification letter/customer letter sent out on 9/25/2007 under Siemens update program US013/07/S. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. Consignees who do not respond to the recall communication will be visit ed and/or telephoned for follow up. |  | 280 | 280 | Worldwide Distribution including the countries of: Australia, Austria, Belgium, Brazil, Brunei, China, Denmark, Finland, France, Germany, Great Britain (UK), Greece, Hungary, India, Iran, Ireland, Italy, Latvia, Netherlands, New Zealand, Romania, Russian Federation, Singapore, Spain, Switzerland, Thailand. U.S customers were not affected. | Terminated | December 20, 2010 | 1182 | 280
Z-1073-2008 | 45545 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Non-continuous ventilator accessories under the following brand names: a) RP-P ComfortGel Cush and Flap model number 1031391 and b) RP-S ComfortGel Cush and Flap model number 1031392. Filter, mask, and/or accessories for sleep therapy and ventilation devices, Respironics, Inc., Murrysville, PA 15668. |  |  | October 15, 2007 | 2007 | Respironics, Inc. | Foreign material: Glass fragments may present in the plastic bag material used to ship components. | Environmental control | The recalling firm issued a Recall Notification dated 10/18/07 to its customers. Replacement products were to be sent. |  | 20950 | 20950 units | The products were shipped nationwide to medical facilities. | Terminated | September 01, 2009 | 687 | 20950
Z-1165-2008 | 46368 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Hemoglobin Reagent Syringe used on the CELL-DYN Sapphire Automated Hematology Analyzer. Abbott Labs, Abbott Park, IL. |  |  | January 15, 2008 | 2008 | Abbott Laboratories | Count may be out of specification--Hemoglobin background count may be out of specification (high) after installation of new CELL-DYN Sapphire Hemoglobin Reagent Syringes with packaging dates between May 8, 2007 and November 29, 2007. | Process control | On January 11, 2008, an Urgent Product Recall Letter and Customer Reply Form were sent to all currently active CELL-DYN Sapphire customers (end-users). The firm provided information for verification of replacement syringes and instructions if hemoglobin background is not within specification. A second Urgent Product Recall letter was sent on January 15, 2008. If this second Customer Reply Form is not faxed back to Abbott, customers will be contacted via an approved telephone protocol. For further information contact customer support at 1-877-4ABBOTT. |  | 2746 | 2,746 syringes were distributed. | Worldwide Distribution-USA including states of AR, AZ, CA, GT, FL, GA, ID, IL, KY, LA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OM, OR, PA, SC, TN, TX, UT, VA, WI, and WV and countries of Canada, Mexico, Australia, Brazil, Canada, China, Chile, Columbia, Germany, Hong Kong, Japan, Mexico, Peru, Singapore, and South Korea. | Terminated | December 23, 2010 | 1073 | 2746
Z-1102-2008 | 46150 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Seradyn QMS Vancomycin reagents, Part Number: 0373589. Seradyn, Inc, Indianapolis, IN. |  |  | December 20, 2007 | 2007 | Seradyn, Inc. | False negative test results due to Interfering substances (heterophile antibodies) in the patient's blood may produce erroneously low results. | Device Design | Consignees were notified of the problem via Urgent Medical Device Correction letter dated 12/20/07. The letter provided information on the issues concerning the recall. The firm will revise the package insert to address the potential occurrence of false negative bias. If inconsistencies with Vancomycin level exist, confirm with an alternative test method. For further information please contact 1-800-428-4072. |  | 3682 | 1,819 kits plus bulk material for 1,863 kits. | Worldwide Distribution --- USA including states of California and New York, and countries of Australia, Belgium, Canada, Finland, Germany, Ireland, Japan, Spain and United Kingdom. | Terminated | December 30, 2008 | 376 | 3682
Z-1038-2008 | 46360 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Leonardo Workstation Picture Archiving and Communication System, Model Number: 7129534, Software Version: VB30B, Siemens Medical Solutions, Inc., Malvern, PA 19355 |  |  | December 21, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect display: The orientation labels will be incorrectly displayed on the reconstructed InSpace 3-D image if the orientation was not originally HFS on the acquisition system. | Software design | Siemens Medical Solutions issued a Customer Safety Advisory Notice dated 12/21/07 to its customers to inform them of the problem and instructions to avoid the problem until the firm's representative is available to upgrade the software. |  | 47 | 47 Units | Nationwide Distribution. | Terminated | September 30, 2008 | 284 | 47
Z-0958-2008 | 45945 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 MPL Flow Cytometry System with MXP Software, Software Version 2.1, Automated differential cell counter, Part Number: 626554, Beckman Coulter, Fullerton, CA 92834 |  |  | June 29, 2007 | 2007 | Beckman Coulter Inc | Data generation error: If the cytosettings is not refreshed or restarted during data importation, the output will include old and new data. | Software design | Consignees were notified by an Urgent Product Corrective Action letter sent on 6/29/07. The letter provided details on the data error and instruction on how to avoid the error. Users were asked to retain notification as part of the Quality Systems documentation and return a response form acknowledging receipt. For additional information, contact 1-800-526-7694. |  | 589 | 516 units in the US; 73 units in Canada | Worldwide Distribution- USA and country of Canada. | Terminated | December 13, 2012 | 1994 | 589
Z-0973-2008 | 46200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Coherence Impression Therapist system, Part No: 7341410, medical charged-particle radiation therapy system with software version 2.1. Siemens AG, Concord, CA. |  |  | October 02, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Image Orientation Incorrect (Image Shifts) - 1) if the reference image has been calibrated for centering and the reference image is used for patient positioning, the image will shift when the positioning tools are used. 2) This problem occurs only when non-square reference images are used, such as when images are cropped. At this point, the images will refresh, and the center of the reference im | Software design | On 10/03/2007, the firm sent Consignees a Customer Safety Advisory Notice informing them of the two potential issues with the display of reference images used for position correction. Consignees will be informed when a solution to this issue has been validated. |  | 16 | 16 units | Worldwide Distribution-USA including 48 consignees in the US and 39 outside the US. | Terminated | December 24, 2008 | 449 | 94
Z-0905-2008 | 46330 | No | EZN | Ureteral dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pathway Balloon Expandable Ureteral Access Sheath- 11/14/28, Ureteral dilator, Catalog Number: 255100; Product Number: M0062551000, Onset Medical, Boston Scientific Corp., Natick, MA 01760 |  |  | November 30, 2007 | 2007 | Boston Scientific Corporation | Sheath Removal Difficulty: Users experience difficulty in removing the sheath. | Device Design | Consignees were sent an Urgent Medical Device Recall letter on November 29, 2007. The letter instructed users to check their inventory, remove any affected product and place them in quarantine. It also included instructions for returning the product. Users were asked to return a Reply Verification Tracking Form. Follow-up communication will be made to users who do not return the tracking form. |  | 89 | 89 | Nationwide Distribution | Terminated | October 22, 2008 | 327 | 3996
Z-1037-2008 | 46358 | No | EGG | Precision attachment. | Dental | Dental | 510(K) Exempt | Straumann NC Closure Screw, Titanium, Precision attachment, Article Number: 024.2105.04, Straumann, Andover, MA 01810 |  |  | January 10, 2008 | 2008 | Straumann Usa, Llc | Mislabeled package: Narrow Connection (NC) Closure Screws contain closure screws for the regular connection (RC) bone level implant. The RC closures screw is larger (3.5mm diameter) and therefore, will not fit into the NC implant (3.0mm diameter) narrow connection. | Mixed-up of materials/components | Consignees were notified by telephone on 1/10/08 followed by an Urgent Device Recall letter on January 10, 2008. The letter instructs users to return any affected product to the firm along with a recall acknowledgement form. For additional information, contact 978-747-2514. |  | 26 | 26 units | Nationwide Distribution including states of CA, DE, FL, IL, ME, NY, OH, OR, and TX. | Terminated | November 06, 2009 | 666 | 26
Z-1106-2008 | 46340 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Advanced Cement Mixer (ACM) Bowl and Base without 180 Gram Cement Cartridge for use with Stryker High Vacuum Cement Injection System, sterile; REF 306-554, Stryker Instruments, Kalamazoo, MI 49001. |  |  | January 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Sterility may be compromised, as the packaging may have channels in the packaging seal. | Nonconforming Material/Component | Stryker notified consignees via Medical Device Recall Notification letter dated 1/10/08 to quarantine the product and to return it. |  | 150 | 25 boxes of 6 each | Nationwide Distribution - Illinois, Michigan, New York, Ohio and Wyoming. | Terminated | July 31, 2008 | 203 | 150
Z-1104-2008 | 46306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | December 28, 2007 | 2007 | GE Healthcare Integrated IT Solutions | Report missing characters: Four special characters, when entered into the exam notes, are not transferred to the preview panel or hard copy printout (greater than [&gt;], less than [&lt;], quotation mark ["], and apostrophe [']). This may not be evident to the end user, as the symbols only appear when the Exam Note window is open. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 12/28/07 to all customers who have the GE Centricity PACS RA1000 Workstation software versions 2.1.X, or 3.0.X, informing them of the potential patient safety issue involving the exam notes window, which may be used to provide various patient information including lab results. To mitigate this problem, the users were instructed not to use the following four special characters (greater than sign &gt;, less than sign &lt;, double quotation mark " and apostrophe ') in the exam notes field in order to prevent any adverse patient events. For example, if a lab result was entered as GFR &lt;60, it should be typed as GFR less than 60. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the problem and the recommended actions provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 557 | 557 units | Worldwide Distribution - USA, Australia, Belgium, Canada, Denmark, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Kuwait, Malaysia, New Zealand, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 17, 2010 | 871 | 557
Z-1118-2008 | 45911 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Fisher Healthcare Sure-Vue Serum/Urine hCG-STAT; in vitro diagnostic test; Catalog Number: FHC-A202; Product Codes: FHC-A202-KFI3O &amp; FHC-A202-KFI5O; Inverness Medical Professional Diagnostics, San Diego, CA 92121. |  |  | October 01, 2007 | 2007 | Innovacon Inc | Incorrect results: The Sure-Vue Serum/Urine hCG-STAT test, when interpreted at extended read times, may exhibit sensitivity to patient samples containing hCG at levels well below the cut-off, potentially resulting in a false positive interpretation by the user. | Labeling design | Inverness Medical sent an Urgent Medical Device Alert letter, dated October 1, 2007, to all Innovacon customers for the product via US Postal Service return receipt requested. The field correction notice removed the extended read time of 10 minutes. Customers were instructed to note that although the package insert currently allows interpretation up to 10 minutes, do not read at times greater than 3 minutes for urine and 5 minutes for serum. |  | 1825480 | 1,825,480 units | Nationwide | Terminated | August 12, 2011 | 1411 | 1825480
Z-1128-2008 | 46693 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 IntraStent LD Mega Large Diameter Biliary Stent, Model Number: S17-36, |  |  | January 18, 2008 | 2008 | Ev3, Inc | Incorrect information on label: Specific lots of ev3 IntraStent Unmounted Balloon Expandable Stents have the two symbols for length and diameter on the side and end flaps reversed. | Labeling Change Control | ev3 notified all consignees via "Medical Device Recall" letter on 1/18/08. The letter was addressed to the Hospital and it described the device affected and the problem. The letter provided recommendations to segregate the affected devices and requested the return of the "Device Recall Field Action Reconciliation Form." |  | 70 | 70 | Worldwide Distribution - USA, ARGENTINA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, DENMARK, GERMANY, GREECE, HONG KONG, IRAN, IRELAND, ITALY, JORDAN, KUWAIT, LEBANON, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | November 04, 2008 | 291 | 942
Z-1119-2008 | 46346 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ADVIA Centaur HAV IgM, 100 Test Kit, used on the ADVIA Centaur or XP, REF: 05004126, in vitro diagnostic Hepatitis A test, Siemens Medical Solutions Diagnostics, Tarrytown, NY 10591. |  |  | January 12, 2008 | 2008 | Siemens Medical Solutions Diagnostics | Incorrect results: False reactive specimens on the ADVIA Centaur HAV IgM assay. | Nonconforming Material/Component | Siemens issued a Urgent Field Safety Notiifcation (078D0709-01) to all Siemens Diagnostics Branches on 1/12/08 for communication to their customers. This Customer Bulletin/Urgent Field Safety Notification advises the customer of the problem and the actions to be taken to mitigate risk. A Confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory action. |  | 49533 | 49,533 | Worldwide Distribution - USA, Puerto Rico, Argentina, Australia, Brazil, Canada, Chile, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Israel, Italy, Malaysia, Mexico, the Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | June 20, 2012 | 1621 | 49533
Z-1036-2008 | 46348 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Diapact CRRT Assembled Kit HF/HD with Manifold and Standard Spike, item number 7210517. The product is shipped 4 units per shipping carton. Hemodialysis kit, B. Braun Medical, Inc., Allentown, PA 18109 |  |  | December 10, 2007 | 2007 | B. Braun Medical, Inc. | Faulty tubing does not prime machine as intended. | Process control | B. Braun Medical sent an Urgent Medical Device Reall letter dated 12/12/07. Customers were told to discontinue use of the tubing. |  | 956 | 956 units | Nationwide Distribution. | Terminated | October 28, 2008 | 323 | 2500
Z-0517-2008 | 46044 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Medtronic Bravo pH Capsule with Delivery System, Stomach pH electrode, Product Numbers 9012B1011 and 9012B1001. Medtronic, Inc., Minneapolis, MN 55432 |  |  | December 03, 2007 | 2007 | Medtronic Neuromodulation | Unable to detach from source: The capsule may not detach from the delivery system following attachment to the esophageal wall. | Labeling Change Control | Consignees were notified by an Urgent Device Recall in December 2007. The letter emphasizes to users the current Instructions for Use regarding detachment of the capsule from the deliver system. It advises that users may continue to use the Bravo capsules and delivery systems. It instructs users to discard all copies of the Âż2005 Bravo pH Monitoring System Interactive Training CD and to use the enclosed Âż2007 Bravo pH Monitoring System Interactive Training. A "Bravo Delivery System Disassembly Procedure" was also provided with the letter. The letter requests consignees to complete and return the reply form included with the notification. For additional information, contact 1-800-707-0933. |  | 169365 | 169,365 individual Bravo capsule and delivery systems have been sold worldwide | Worldwide Distribution -- USA states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY and countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan , Korea, Kuwait, Lebanon, Liechtenstein, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Panama, Portugal, Russian Federation, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | May 03, 2012 | 1613 | 169365
Z-1305-2008 | 46785 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Mach 1 peripheral Guide Catheter 8F, RDC Renal double Curve, .091 in (internal dia), Shaft lenght 55cm, Catalog Number: 19-866, UPN M001198660, Sterile EO, Availmed S.A. de C.V. (Mexico), Boston Scientific Corp., Maple Grove, MN 55311 |  |  | February 07, 2008 | 2008 | Boston Scientific Corporation | Mislabeled: Seven lots/ batches of Boston Scientific Mach 1 Guide Catheter may be mislabeled with the incorrect device length. | Labeling mix-ups | Consignees were sent a Boston Scientific Urgent Voluntary Medical Device Recall letter on 2/7/08. The letter included the customer steps for recall, a complete listing of all product batches being recalled, the Reply Verification Tracking Form(s) and a return shipping label. |  | 4 | 4 | Worldwide Distribution including USA states of PA, AZ, NY, KY, CA, WV, OH, WA, MI, TN, IN, IL, NJ, KY, FL, AR, FL, NC, CO, TX, WV, WI, OR, AL, MN and ND and countries of BE, FR, GB, DE, IE, IT, JP, AU, SOUTH KOREA, NZ, TR, TW and ZA. | Terminated | November 03, 2008 | 270 | 94
Z-1240-2008 | 46832 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LIAISONÂż 25-OH Viatamin D Kit, Vitamin D test system, Model Number 310900, DiaSorin Inc, Stillwater, MN 55082 |  |  | February 01, 2008 | 2008 | Diasorin Inc. | Inaccurate measurements: Kits may sporadically recover kit and external control values out of range high. | Nonconforming Material/Component | Consignees were notified by a Diasorn Customer Notification Letter sent by email on 2/1/08. The letter instructs users to stop using the affected kits, destroy any kits remaining in inventory and to return a response form in order to receive replacement product. For additional information contact 1-800-328-1482. |  | 2267 | 2,267 kits | Worldwide Distribution including USA states of AR, CA, GA, IA, IL, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OR, PA, TN, UT and WA and country of Austria. | Terminated | October 31, 2008 | 273 | 2267
Z-1179-2008 | 46802 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Active Life Little Ones One-Piece Custom Urostomy Pouch 5/16 - 1 inch, 8-25mm; Ostomy pouch and accessories, Product Order Number 020917, Nypro Dominican Republic Inc., ConvaTech, Skillman, NJ 08558 |  |  | February 14, 2008 | 2008 | ConvaTec | Mislabeled: The market unit carton label reads 5/8 - 1 inch, and it should read 5/16 - 1 inch. The metric units (8-25mm) on the label are correct. | Labeling mix-ups | Consignees were notified by an Urgent Product Recall Letter on 2/14/08. Updated letters were issued on 2/21/08 and 4/10/08 with expanded lot numbers. The letters instruct users to return any affected products and a product recall response form. For additional information, contact (908) 904-2199. |  | 3986 | 3986 market units | Nationwide Distribution. | Terminated | October 21, 2008 | 250 | 3986
Z-1141-2008 | 46414 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS image management system; Picture Archiving and Communications System; Radiology software package, Philips Healthcare Informatics, Foster City, CA 94404. |  |  | January 04, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Incorrect display of one of the patient's images. This malfunction could cause an image to be incorrectly displayed, for example flipped left to right and anterior to posterior. The original image is still presented in the correct orientation next to the incorrect image when displayed on the monitor. | Software design | Philips Healthcare Informatics mailed customers notification: Urgent Product Correction regarding FCO#2954704-01/04/08-001-C. Firm released a corrected version of software (iSite PACS v.3.5.69) on January 21, 2008. Firm is scheduling customers to upgrade to the corrected version. |  | 197 | 197 units | Worldwide Distribution: USA, United Arab Emirates, Germany, the Netherlands, Denmark, Australia, France, Italy, Belgium, Ireland, Italy, and Canada. | Terminated | October 07, 2009 | 642 | 197
Z-1543-2008 | 47378 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medifil Heparin I.V. Flush Syringe 100 units/mL; a polypropylene luer locking syringe containing Heparin Lock Flush Solution, USP, a sterile non-pyrogenic, isotonic injection; each mL in a Heparin I.V. Flush Syringe contains 100 units of heparin sodium derived from porcine intestinal mucosa and 9 mg of sodium chloride in water for injection; Medefil, Inc., Glendale Heights, IL 60139 |  |  | March 20, 2008 | 2008 | Medefil Incorporated | The heparin lock flush solution was manufactured from contaminated Heparin Sodium USP Active Pharmaceutical Ingredient (API). The FDA has received reports of serious injuries and/or deaths in patients administered finished heparin injectable products manufactured from heparin API containing this contaminant. | Material/Component Contamination | Medefil telephoned the direct accounts on 3/20/08, e-mailed them a copy of the 3/20/08 recall letter on 3/21/08, and sent the accounts a hard copy of the recall letter on 3/24/08. The accounts were informed of Medefil's recall of 19 lots of Heparin I.V. Flush Syringes produced from the active pharmaceutical ingredient (API) Heparin Sodium, USP, lot 1035-0773, that has been recalled by the manufacturer, Scientific Protein Laboratories, due to the presence of a contaminant in the recalled lot of API. The accounts were instructed to stop use of the 19 lots of Heparin I.V. Flush Syringes immediately, place them under quarantine, and contact all customers to whom they further distributed the product, telling them to stop use of the product and return the product to Medefil through their supplier. The accounts were requested to complete and return to Medefil via fax the enclosed response form, indicating the quantity on hand of each recalled lot. Any questions were directed to Pradeep Aggarwal, President at 630-682-4600. |  | 3151080 | 3,151,080 syringes | Class II Recall - Nationwide Distribution --- including states of Arizona, Florida, Illinois, Massachusetts, Pennsylvania, and Texas. | Terminated | April 15, 2010 | 756 | 3713040
Z-1602-2008 | 47436 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Tyco Healthcare Monoject Prefill 100UlmL Heparin |  |  | March 28, 2008 | 2008 | Covidien LP | Scientific Protein Laboratories (SPL), disclosed that two lots of Heparin Sodium, USP Active Pharmaceutical Ingredient has a heparin-like contaminant | Material/Component Contamination | Tyco Healthcar notifed customers by letter sent via US Postal Service beginning 3/28/08, through 4/1/08. The letter requests that they immediately stop using and return the product. Distributors are requested to in-turn notify their customers. An instruction sheet was attached for product return and replacement.All customers are required to respond to the notice regardless of remaining inventory.. |  | 21600 | 21,600 units | Nationwide Distribution | Terminated | June 25, 2012 | 1550 | 11054984
Z-0987-2008 | 46294 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | AirLife Infant nCPAP System Driver (Catalog No: 006900;) a nasal continuous positive airway pressure device, which is part of a larger system that includes an Infant Nasal CPAP generator, fixation device and patient interface (prongs and/or mask); Cardinal Health, McGaw Park, IL 60085 |  |  | January 04, 2008 | 2008 | Cardinal Health | Oxygen fluctuations - The AirLife Infant Nasal CPAP System exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxillary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. | Device Design | Cardinal Health sent an Urgent Medical Device Correction Notice dated 1/4/08 with a list of catalog number and affected serial numbers of the nCPAP Driver, and informed them that the nCPAP Driver may exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxiliary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. Customers were instructed to discontinue use of the PTO/Auxiliary port and were provided with a sticker reading "WARNINIG! DO NOT USE PTO/AUXILIARY PORT" that must be affixed to each unit in their possession. Sticker placement instructions and photo showing proper placement were provided with the letter. The accounts were informed that future upgrades to this device will eliminate the PTO/Auxiliary port and are expected to be completed in the field by June1, 2008. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503. |  | 314 | 314 units | Nationwide Distribution | Terminated | January 09, 2009 | 371 | 314
Z-0982-2008 | 46215 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | VIASYS AVEA Ventilator, Catalog numbers: 17210-00 (USA Comprehensive Version) |  |  | December 17, 2007 | 2007 | Viasys Respiratory Care, Inc.dba Bird Products | Alarm Failure -- There is a potential of alarms not being activated subsequent to specific electronic faults that may cause INOP condition. | Component design/selection | On December 18, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (registed mail) and a product notification letter was mailed to all international customers via Federal Express. The recall letter informed the customers of the problem and also inform them that a Cardinal Health (Viasys) service representative will be contacting them immediately to arrange &amp; service their ventilator. For futher information please contack Andrea von Muller at 760-778-7323. |  | 2 | 2 | Worldwide Distribution --- USA including states of AL, CA, MA, NM, NV, and NY and country of Italy. | Terminated | May 13, 2008 | 148 | 41
Z-1105-2008 | 46331 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon Baseplate Impactor Extractor, Orthopedic manual surgical instrument, Catalog Number 6541-4-805, Non-Sterile, Stryker Howmedica Osteonics Corp, Mahwah, NJ 07430 |  |  | February 23, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Insertion/Removal Difficulties: The Baseplate Impactor/Extractor may not assemble/disassemble easily to the baseplate. | Device Design | Consignees were notified via an Important Market Withdrawal Letter sent 2/25/05. The letter instructed users to fax a Product Accountability Form to the recalling firm identifying any affected products and to return them to the firm. The letter informed users that the products will be reworked and returned at a later date. The letter also identified other products that could be used as replacements. For additional information, contact 201-831-5825 or 201-831-5110. |  | 293 | 293 total units | Worldwide Distribution | Terminated | May 30, 2008 | 1192 | 293
Z-1103-2008 | 46296 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | White Blood Cell (WBC) Reagent-Part A used with the CELL-DYN 4000 and CELL-DYN Sapphire Systems, Red cell lysing reagent, List Number 01H77-01, Supplied in 3.8L bottles, Abbott Laboratories, Abbott Park, IL 60064 |  |  | December 19, 2007 | 2007 | Abbott Laboratories | Microbial Contamination: The WBC (White Blood Cell) Reagent Part A was confirmed positive for Pseudomonas contamination. | Material/Component Contamination | Consignees were notified by a Product Recall letter sent 12/19/07. The letter instructs users to discontinue use of affected product and destroy any remaining inventory. Users were requested to return a customer reply form. For additional information, contact 1-877-4ABBOTT. |  | 1378850 | Product was distributed in the follow amounts: 1,199 botttles of lot code 51749i2; 1,150 bottles of lot code 1574i2, 1,150 bottles of lot code 56675i2 | Worldwide Distribution including USA states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI and WV, and countries of: Argentina, Australia, Brazil, Canada, Germany, Hong Kong, Japan, Singapore, South Korea and Venezuela. | Terminated | July 02, 2008 | 196 | 1378850
Z-1462-2008 | 46740 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 |  |  | January 21, 2008 | 2008 | Abbott Laboratories | Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result. | Software change control | Consignees were sent a Product Correction-Immediate Action Required letter on 1/21/08. The letter provided users with the steps to take to temporarily correct the problem until the recalling firm is able to send a corrected assay file. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 1270 | 1,270 disks | Worldwide Distribution including USA territory of Puerto Rico, and countries of Australia, Canada, Chile, Colombia, Germany, Mexico, New Zealand, South Korea, Taiwan, Thailand and Venezuela. | Terminated | December 01, 2009 | 680 | 1270
Z-1624-2008 | 46801 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel RapID Inoculation Fluid, Microorganism differentiation and identification device, Packaged in 1-ml tube, Catalog Number: R8325102, 20/pack, Remel, Lenexa, KS 66215 |  |  | October 23, 2007 | 2007 | Remel Inc | Failure to produce reaction: Use of the product with various RapID Identification systems panels may exhibit poor or no reactions. | Other | Consignees were notified by an Important Medical Device Customer Advisory Notice on 10/23/07. The letter instructs users to review inventory and discard any affected products. The letter also instructs users to return a product inventory checklist by 11/23/07. For additional information, contact 1-800-447-3641. |  | 500 | 500/1-ml packs | Worldwide Distribution including USA and countries of Taiwan, Japan, and U.K. | Terminated | July 16, 2008 | 267 | 1054
Z-1461-2008 | 47401 | No | HNR | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon/Grieshaber Morris ILM Forceps, Direct Action; Manual ophthalmic surgical instrument, Catalog Number: 641.22; Non-sterile/Reusable; Alcon Laboratories, Inc., Fort Worth, TX 76134. |  |  | March 07, 2008 | 2008 | Alcon Research, Ltd | Device corrosion: Forceps are prone to corrosion and possibly premature fracture and/or malfunction. | Device Design | Consignees were notified by an Urgent Medical Device Recall that was sent via fax on 03/07/08. Users were instructed to cease using the affected products and to return a Recall Acknowledgement Form to the recalling firm. The sales representative will pick up the affected product and discuss replacement alternatives. For additional information, contact 1-800-757-9780. |  | 24 | 24 units | Nationwide Distribution including states of AL and CA. | Terminated | July 18, 2008 | 133 | 24
Z-1138-2008 | 46373 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo US Workplace Picture Archiving and Communication System, software version 2.0, Siemens Medical Solutions, Inc., Mountain View, CA 94043 |  |  | December 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Inaccurate results: A software bug may cause inaccurate wall motion abnormality scoring results to be displayed. | Software design | Siemens Medical Solutions sent an Urgent Medical Device Notification letter, dated November 2007, to its consignees. The letter described action required by users, and that Siemens will contact the user with a new software version correcting this error. A field correction to update software is planned to be underway, designated to start on 12/21/2007, and the correction is expected to be completed within 45 days. |  | 17 | 17 units | Worldwide Distribution: South Korea, Malaysia, Lithuania and China, and 1 unit to New Jersey in the USA. | Terminated | December 20, 2010 | 1110 | 17
Z-1616-2008 | 47205 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 9051 Electrode Adapter; Automated external defibrillator, Used with QUIK-COMBO System; Cardiac Science Corp, Bothell, WA 98021. |  |  | January 23, 2008 | 2008 | Cardiac Science Corporation | Adaptor Cable Incorrectly Manufactured: The red and white connector ends of the adapter are switched. Therefore, the cable will not connect to the electrodes as indicated in the instructions. | Process control | Consignees were notified by an Urgent Medical Device Recall letter sent on 01/23/2008. The letter instructs users to check each cable of the affected lot to ensure that it properly connects to the electrode. Users were also instructed to return any defective product to the recalling firm for a replacement. For additional information, contact 1-888-466-8686 x2041. |  | 58812 | Lot W029275 - 24; Lot W029475 - 38 | Nationwide Distribution including states of CT, IA, IN, NC, NY, OH, PA, WA, and WI. | Terminated | July 21, 2011 | 1275 | 58812
Z-1614-2008 | 46905 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Fiducial Markers (0.9 x 3 mm) Soft tissue gold markers, Medical charged-particle radiation therapy system; Part Number: MT-NW-887-809, CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 28, 2008 | 2008 | Med Tec Inc dba CIVCO Medical Solutions | Mislabeled: Fiducial markers were incorrectly labeled. They are marked (0.9x3mm cylindrical) Soft tissue gold markers, but contain (2mm spheres) Bone gold markers. | Labeling mix-ups | Consignees were notified via telephone beginning 2/28/08 and a Recall Notification letter sent on 2/29/08. Users were requested to return any affected products for a replacement. They were also requested to return a recall notification form indicating the amount of affected product in their inventory. For additional information, call 800-842-8688. |  | 33 | 33/3-marker pouches | Nationwide Distribution including states of WA, PA, DE, AL, CT, WI, FL, TX, MN, and IL. | Terminated | July 02, 2008 | 125 | 69
Z-1608-2008 | 46863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MT-APSD-2.4 Type-S Thermoplastic Mask, Medical charged-particle radiation therapy system; Disposable non-sterile (43.2cm) standard perforated thermoplastic head, neck and shoulder mask (2.4mm), 1 device per package; CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 25, 2008 | 2008 | Civco Medical Instruments Inc | Mislabeled: The mask is mislabeled with the incorrect part number but the outer packaging is correct. | Labeling mix-ups | Consignees were notified by a Recall Notification letter on 2/25/08. The letter informed users that the recalling firm will replace the affected product. It also requested that users return a recall notification acknowledging notification of recall. Users were instructed to contact their sales representative to arrange servicing or for additional information at 1-800-842-8688. |  | 160 | 160 | Worldwide Distribution including USA states of IL, PA, NC, TN, SD, CA, MI, WI, MA, and CO; and country of Japan. | Terminated | July 02, 2008 | 128 | 160
Z-1132-2008 | 46741 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Keeler All Pupil II Indirect Ophthalmoscope, Keeler Instruments, Inc., Broomall, PA 19008 |  |  | January 08, 2008 | 2008 | Keeler Instruments Inc | Good Manufacturing Practices (GMP) deficiencies may compromise the safety and effectiveness of the device. Firm is on import alert 89-04 and released product. | Process design | Keeler Instruments issued an email dated 1/8/08 to the distributor requesting that they return the product. |  | 25 | 25 units | The products were shipped to a distributor in NJ. | Terminated | August 12, 2008 | 217 | 25
Z-1110-2008 | 45909 | No | CKA | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid ALT (SGPT) reagent set for the quantitative determination of alanine aminotransferase in serum, 150 ml, Catalog no. A7526-150, in vitro diagnostic, Pointe Scientific, Inc.; Canton, MI 48188 |  |  | December 14, 2007 | 2007 | Pointe Scientific, Inc. | Failure of the reagent to produce test results. The R1 reagent may be contaminated with microorganisms. | Process control | Pointe Scientific notified consignees via Urgent Medical Device Recall letters dated 12/14/07, instructed to dispose of the product and to send back a reply form to request replacement product. |  | 955 | 955 for all products. | Worldwide Distribution - USA, Bolivia, Canada, Chile, China, Cyprus, Greece, Indonesia, Israel, Korea, Mexico, Pakistan, Panama, Peru, Philippines, Republic of Yemen, Russia, South Africa, Trinidad, Ukraine and West Indies. | Terminated | September 15, 2008 | 276 | 955
Z-1094-2008 | 45960 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX20 Clinical Chemistry System, Part Numbers: 466200,473032, Beckman Coulter, Inc. |  |  | April 05, 2007 | 2007 | Beckman Coulter Inc | Intermittent Failure of Stirrer Motor-Beckman Coulter has determined that incorrect Glucose results on LX20 orDxC instruments have been related to an intermittent failure of the stirrermotor.-Stirrer motors can stall without any flags or motion errors.-If a stall occurs during a GLUm test near the time of sample inject,results may be affected.-Affected GLUm results may be high or low. | Software design | GLUm stir motors greater than one year old were replaced. Replacements began April 5, 2007. Based on the Risk Assessment, BCI decided that Customer Notification is not needed. GLUm stir motors greater than one year old will be replaced on high volume LX/DxC systems by 3rd quarter of 2007. Replacements began April 5, 2007. A new stirrer design is being developed and is expected to be available before 4th quarter of 2007. Implementation Strategy of Stock Control: 1. A Stop Ship was not necessary based on risk assessment. The possibility of stirrer motor failure may increase with time and duty cycles. Systems running a high volume of tests may have a greater likelihood of exhibiting stirrer motor failure with motors greater than one year old. 2. New Customers were placed on a preventative maintenance plan to have their motor replaced at one year cycle. Completed 3. A Technical Update was issued to notify Hotline and Field Service Personnel of the issues. |  | 370 | 344 in the US; 26 in Canada | Worldwide Distribution - USA including Canada. | Terminated | January 15, 2013 | 2112 | 2599
Z-1166-2008 | 46389 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry ICT Calibrator, List Number: 1E46-02, Abbott Labs, Abbott Park, IL 60064 |  |  | September 26, 2007 | 2007 | Abbott Laboratories, Inc | Quality Control (Out of Range)- ICT Calibrator lot number: 0505017, may contain an interferent. Not all vials within a kit are affected.The investigation determined that the interferent affects Potassium (K+) only. Use of the affected calibrator vials will generate an acceptable, but low calibration slope. When Quality Control (QC) is run to verify the calibration, K+ QC results may be bel | Component design/selection | An Immediate Action Product Recall dated 9/26/07, was provided to all AEROSET and ARCHITECT cSYSTEM customers that have received Lot Number 0505017. The effectiveness check will be completed through a customer reply form for 100% of US customers. Outside of the U.S. each individual country will determine the methods of effectiveness check.Necessary Actions:If the ICT calibration generates an acceptable slope for (K+), and the calibration is verified by generating QC results within the acceptable range, then the customers may continue to use the lot until a new one is available.If the ICT calibration generates an acceptable slope for (K+) with results below QC range, and (Na+) and (Cl-) are not affected, then they should discontinue use of the Calibrator vials that generate the unacceptable calibration and destroy them following their laboratory procedures.If (Na+) and/or (Cl-) calibration slopes or QC results are not acceptable, additional troubleshooting is required.Questions may be addressed to Abbott Customer Support at 1-877-4ABBOTT. |  | 1796 | 1,377 Distributed Worldwide; 419 Domestically | Worldwide Distribution --- including USA, Puerto Rico and countries of Mexico, Ecuador, Brazil, Canada, Hong Kong, New Zealand, Venezuela, Chile, , Germany, Thailand, Trinidad/Tobago, Columbia, Argentina, El Salvador, Singapore, Australia, and the Bahama. | Terminated | March 27, 2012 | 1644 | 1796
Z-1644-2008 | 47526 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | EVMS (Enterprise Visual Medical System); Picture archiving and communications system; Model Numbers: 5.30.4, 5.30.5, 5.30.6, 5.30.7, 5.40.0 and 7.0.0.1. Emageon, Inc., Hartland, WI 53029. |  |  | January 01, 2008 | 2008 | Emageon | Miscalculation: The calculation of the standard uptake value does not include the required calibration factor/decay correction factor. The resulting value is incorrectly calculated, which may lead to incorrect interpretation of tissue density. The calculation error exists in both the software and user guide. | Device Design | Consignees were notified by Urgent Device Correction letters that were sent on 1/9/08 and 1/18/08. The letters instructed users not to use the standard uptake value in EVMS for diagnostic interpretation when viewing PET studies unless they received service bulletin 445, 446 or 7.0.1. Users were also instructed to replace pages in the user's guide with the pages provided and to affix the enclosed warning sticker on all workstations. Users will be contacted by the recalling firm to schedule a software patch. |  | 99 | 99 | Nationwide Distribution including states of: MI, WI, CA, MN, AZ, CO, MN, MD, DC, MO, TX, IL, GA, FL, WI, SC, PA, NY, KS, TN, AL and OK. | Terminated | June 11, 2009 | 527 | 99
Z-1546-2008 | 47386 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with Software v. 3.3.1.1; Full field digital, system, x-ray, mammographic; Model Number: SEL-00002; Hologic, Danbury CT 06810 |  |  | February 28, 2008 | 2008 | Lorad, A Hologic, Inc. | Marker Misalignment: The new software version contained a magnification factor that shows CAD markers misaligned with the identified indications. The marked area is misaligned up to 2cm from the indication. | Software design | Consignees were notified by an Urgent Recall Notice sent on 3/21/08. The letter informed users that representatives were dispatched to all locations and they have corrected the magnification factor problem. The letter instructed users to re-evaluate any cases read using CAD from the period of installation until the date of system correction. The letter also instructed users that only images read with CAD markers are affected. The letter requests users to return a confirmation of notification. For additional information, contact 1-877-371-4372. |  | 73 | 73 units | Worldwide Distribution including USA and countries of Canada and the Republic of Korea. | Terminated | February 03, 2009 | 341 | 73
Z-1137-2008 | 45959 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva Flat Detector X-Ray System; Part Number: SW 9896 001 43361; Philips Medical Systems, |  |  | November 29, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Delay in Imaging: MultiDiagnost Eleva with Flat Detector x-ray system cannot perform fluoroscopy procedures while certain image file housekeeping operations are performed. Fluoroscopy cannot be performed and new exposures are not able to be acquired until the housekeeping operations are complete which may take several minutes. The loss of system availability could occur during an interventional | Software design | On 11/29/07 Philips Medical Systems sent a letter dated 11/28/07 to customers stating that if users perform certain image file housekeeping operations, fluoroscopy cannot be performed and new exposures cannot be acquired until those operations are complete. It also instructs that the unavailability of fluoroscopy and exposure can be avoided by not pressing the "Delete All Images" button until any pending Autopush operation is complete. FCO 70800067 will be issued to service representatives to install new software. |  | 18 | 9 units were distributed to hospitals in the US, 6 were installed, 3 were not installed or in use. | USA Distribution - CO, MN, NC, NJ and TX. | Terminated | September 17, 2010 | 1023 | 18
Z-0580-2008 | 44725 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur TnI-Ultra Assay, Troponin Immunoassay, 500 tests - Part/Catalog # 02790309, and 100 tests - Part/Catalog # 02789602. For in vitro diagnostic use, Siemens Medical Solutions Diagnostics, Tarrytown, NY 10591 |  |  | June 07, 2007 | 2007 | Siemens Medical Solutions Diagnostics | Incorrect results: Falsely elevated Troponin values have been reported which are inconsistent with the patient's clinical picture and test negative by other Troponin assays. | Material/Component Contamination | Two Support Bulletins (078N0361-01 &amp; 078N0361-02) were sent by e-mail to all affected Siemens Diagnostics Branches to inform them of this issue. The first one was sent on 6/07/07 and the second one was sent on 7/05/07. Two Customer Bulletins (078D0647-01 &amp; 078D0651-01) were sent by e-mail to the Siemens Diagnostics Branches for communication with the affected customers. The first one was sent on 6/07/07 and the second one was sent on 7/05/07. These bulletins state that if a false positive is suspected, the sample should be diluted and retested. |  | 58818 | Domestic: 31,497 kits; Foreign: 27,321 kits | Worldwide Distribution - USA, Argentina, Austria, Australia, Brazil, Canada, Chile, Colombia, Egypt, France, Germany, Greece, Hong Kong, India, Israel, Italy, Malaysia, Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela | Terminated | August 28, 2012 | 1909 | 58818
Z-1139-2008 | 46393 | No | LZS |  |  | Ophthalmic | PMA | AMO VISX STAR Excimer Laser System, manufactured by VISX, Santa Clara, CA 95051 |  |  | December 13, 2007 | 2007 | VISX INCORPORATED, A SUBSIDIARY OF AMO INC | Firm received two complaints (including one injury complaint) associated with the movement of the VISX chair while in the patient loading ("fully swung out" unlocked) position under an IntraLase FS laser. | Component design/selection | On December 13, 2007, firm issued an Urgent Safety Notice to all VISX and IntraLase laser owners related to recall. The Safety Notice contains information related to the measures that can be taken to prevent unintentional movement of any chair during other procedures. AMO is initiating a field corrective action, whereby Field Service Engineers will visit each practice that owns both systems to disable the VISX chair back switch. For additional information, please contact AMO at (714) 247-8656. |  | 1853 | 1,853 units | Worldwide Distribution - USA, including the states of: AKALARAZCACOCTDCDEFLGAHIIAIDILINKSKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPAPRRISCSDTNTXUTVAVTWAWIWVand countries of: ArgentinaAustraliaAustriaBahrainBelgiumBoliviaBrazilBulgariaCanadaChileChinaColombiaCosta RicaCyprusCzech RepublicDjiboutiDominican Rep.EcuadorEgyptFinlandFranceGermanyGreat BritainGreeceHong KongIndiaIndonesiaIrelandIsraelItalyJapanJordanKasachstanKuwaitLebanonMalaysiaMartiniqueMexicoMongoliaNetherlandsNew ZealandNorwayOmanPakistanParaguayPhilippinesPolandPortugalRussian Fed.Saudi ArabiaSingaporeSlovakiaSouth AfricaSouth KoreaSpainSwedenSwitzerlandTaiwanThailandTrinidad,TobagoTurkeyUnit.Arab Emir.VenezuelaVietnam | Terminated | May 05, 2010 | 874 | 1853
Z-1140-2008 | 46394 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | B-K Medical Intraoperative Transducer Type 8814, diagnostic ultrasound transducer, B-K Medical Systems, Inc., Wilmington, MA 01887 |  |  | January 24, 2008 | 2008 | B &amp; K Medical Systems, Inc. | Crack: Incomplete glue joint in the device housing may affect electrical safety or sterilization. | Process control | B-K Medical issued letter notification on 1/23/08 to accounts via Fed Ex, advising them not to use the unit until a representative can inspect and exchange, if unit is found defective. Instructions are also provided to continue use of the device with a sterile cover until site visit. |  | 7 | 7 | USA Distribution: MA, MI, PA, TX, and OR. | Terminated | June 23, 2008 | 151 | 7
Z-0842-2008 | 46859 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a Radiographic X-ray System, Model Number 2259976; previously marketed as the Silhouette FC, GE Healthcare, Waukesha, WI 53188. |  |  | November 01, 2007 | 2007 | GE Healthcare | The warning label that is required by 21 CFR 1020.30(j) was not on the control console of the Proteus XR/a Radiographic X-ray System, Model 2259976. | Process control | GE issued a Field Modification Instruction to their field service representatives, released in December 2007, with instructions to replace the labels on the console with the appropriate labeling. Consignees will be visited by GE field service representatives beginning December 2007. All consignees will have the required labels affixed to their units by July 31, 2008. |  | 400 | 400 Units | Worldwide Distribution, including USA, Canada, Mexico, France, Hong Kong, Italy, Korea, China, Poland, UK, New Zealand, Saudi Arabia, Taiwan, Turkey, and Australia. | Terminated | December 18, 2011 | 1508 | 400
Z-1711-2008 | 47335 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March 26, 2004: (1) Alaris PC point-of-care unit ("PC Unit") (Model 8000)* (a.k.a. Medley PC Unit), |  |  | March 05, 2008 | 2008 | Cardinal Health 303 Inc DBA Alaris Products | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors with the Alaris System due to the Inter-Unit Interface (IUI) connectors on all Alaris System modules (manufactured on or before March 26, 2004). Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March | Device Design | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support. |  | 77515 | 77,515 Units | Class II Recall - Worldwide Distribution --- USA including states of AR, AZ, CA, CT, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD MI, MN, MO, MS, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WY, Guam and country of Canada.****CDRH Comments: "Note for District Office to follow-up" The firm has not provided a grand total of devices affected by these problems. It is possible that devices suffering from U19 &amp; U9 problems may also experience IUI problems. This recall has 3 problems with 6 modules. One problem affects only 2 modules, one problem affects 6 modules. For simplicity and expediency, CDRH agrees to keep the file as is, rather than separate into 6 products and have complicated explanations of the problems.77,515 devices affected by IUI problem.85,286 devices affected by U19 problem.8,286 devices affected by U9 problem. **** | Terminated | September 13, 2011 | 1287 | 171087
Z-1178-2008 | 46795 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource Standard Sterile OR Scissors; straight scissors, S/B, 5.5", made in Pakistan; individually packaged in rigid styrene roll stock with a tyvek top, 50 units per case; Cardinal Health, McGaw Park, IL 60085; catalog #SSI-0003 |  |  | February 12, 2008 | 2008 | Cardinal Health | The product has the potential of being non-sterile. | Process control | Cardinal Health notified the accounts who ordered the affected scissors via letter dated 2/12/08, informing them that the scissors that were labeled sterile were subject to a sterility test failure. Upon review of the letter, the accounts were requested to inspect their inventory and immediately segregate any affected product and to telephone Cardinal Customer Sales Operations Group at 1-800-766-0706 for return instructions and to obtain credit/replacement. Any questions were directed to Cardinal's Quality Systems, Professional Services at 800-292-9332. |  | 16 | 16 cases | USA Distribution: Kansas and Missouri | Terminated | October 16, 2008 | 247 | 16
Z-1156-2008 | 45875 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIOPLEX 2200 SYPHILIS IgG Kit: 100 tests - Catalog No. 665-1450 - manufactured by Bio-Rad Laboratories in Redmond, WA 98052. The product is a multiplex flow immunoassay intended for the qualitative detection of Treponema pallidium IgG antibodies in human serum. |  |  | November 13, 2007 | 2007 | Bio-Rad Laboratories | False negative results due to reagent packs exhibiting low signal. | Nonconforming Material/Component | On 11/13/2007, the firm faxed Voluntary Field Correction Letters (which are dated 11/9/2007) to its consignees and also contacted them via telephone. The Voluntary Field Correction Letters: a) informed them of the potential for false negative results; b) recommended that they perform QC testing daily with each reagent pack; c) provided a new insert for the Instructions for Use; and d) provided a Response Sheet for each consignee to complete and return to the firm. |  | 796 | 796 kits | Worldwide Distribution -- USA, Canada, France, United Kingdom, Italy, Germany, South Africa, and Switzerland. | Terminated | May 23, 2011 | 1287 | 11616
Z-1176-2008 | 46745 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | CalMed Pericardial Sump, with 1/4"/0.635 cm Connector, Single Use Only. Do Not Reuse; REF SU-20602; CalMed Laboratories, Costa Mesa, CA 92626 |  |  | January 24, 2008 | 2008 | California Medical Laboratories Inc | Fragment of material: The product may have a loose, stainless steel stringer flash located at the distal tip of the product. | Process control | CalMed emailed letters to consignees on 1/24/2008 notifying customers of the recall. The letter advises that that the following Medical Devices are subject to an immediate Voluntary Recall forCorrection: Pericardial Sump (Catalog #SU-20601), Pericardial Sump (Catalog #SU-20602), and Malleable Pericardial Sump (Catalog #SU-20802). It is asked that further distribution or use of any remaining product should cease immediately. Where appropriate, the firm asks that consignees notify their own customers who received the product of the same.The Product is being recalled due to a potentially loose, stainless steel stringer flash located at the distal tip of the Product. Customers are asked to contact the firm's Quality Assurance Manager immediately for purposes ofissuing a Returned Goods Authorization and coordinating the return of the Subject Medical Devices such that they may be inspected and nonconforming product removed. A response is requested via e-mail. |  | 31124 | 31,124 units of all models | Worldwide Distribution - USA, Turkey, Brazil, Germany, Italy, Spain, Philippines, South Africa, Japan, Israel, Saudi Arabia, United Arab Emirates, India, Ireland, China, Hungary, Poland, Taiwan, Belgium, Pakistan, Lithuania, Greece, Canada, Syria, Israel, Austria, Denmark, Bangladesh, Costa Rica, Mexico, Spain, and Sultanate of Oman. | Terminated | March 05, 2013 | 1867 | 31124
Z-1155-2008 | 44884 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 |  |  | September 04, 2007 | 2007 | GE Healthcare Clinical Systems | Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t | Software change control | GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software. |  | 62 | 62 | Worldwide Distribution - USA (NY &amp; AL), Korea, China, India, Brazil, Mexico, Germany, Greece, Italy, Poland, Russia, Saudi Arabia, Spain, and Czech Republic. | Terminated | May 29, 2012 | 1729 | 62
Z-1173-2008 | 46691 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL Acetabular Shells; Trident PSL HA Solid Back, |  |  | January 21, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals. | Process design | Recall notification letters were sent to Stryker Branches/agencies, OR Supervisors and Chief of Orthopaedics on 1/21/08. A Patient information sheet was sent on 2/4/08 to surgeons and hospitals. Per call with Center 2/4/08, recall letter to be revised and RES updated upon receipt.The letter was revised and sent on February 28, 2008 to include Trident PSL Acetabular shells. This letter was sent to Risk Management at hospitals and included the scope of the recall , the potential hazard and recommendation to physicians to monitor patients consistent with care for those receiving total hip replacement. It was also sent to surgeons and a revised letter to Stryker branches. A product acknowledgement form was included in all letters to indicate receipt of letter and quantity of product on hand, if applicable. |  | 10191 | 10191, both PSL types | Nationwide distribution. | Terminated | September 25, 2012 | 1709 | 10191
Z-1180-2008 | 46716 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System, Model number IS2000 Revision number A51 P5; endoscopic surgical control system, Intuitive Surgical, Inc., Sunnyvale, CA 94086 |  |  | December 27, 2007 | 2007 | Intuitive Surgical, Inc. | Delay in responding: In certain circumstances, the device may not respond immediately to a user's command such as master clutch or camera control. | Software design | Beginning 12/28/2007, Intuitive Surgical Field Service Representatives visited each site to correct the issues. A notification letter was provided to all sites. |  | 9 | 9 units | USA Distribution: TX, NJ, WA, NV, HI, MI, NH and CA | Terminated | August 15, 2008 | 232 | 9
Z-1161-2008 | 46204 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | S/5 iCentral, Cardiovascular Monitoring Devices, Versions L-NET03, L-NETC03, L-NET05 and L-NETC05, GE Healthcare, Wauwatosa, WI 53226. |  |  | December 13, 2007 | 2007 | GE Healthcare | Failure to alarm: An active monitor could become disconnected from the unit without any notification or alarm. | Device Design | Consignees were notified via a Medical Device Correction letter on 12/13/07. The letter explained the circumstances in which the "connection to monitor loss" alarm would not display, if a monitor was disconnected. The letter also included an attachment entitled "How to ensure a 'Connection to monitor lost' alarm on iCentral." For additional information, contact 1-800-558-7044. |  | 2146 | 2146 | Worldwide Distribution including USA states of AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, ME, MI, MN, MO, MS, NC, NH, NJ, NY, PA, SC, TN, TX, VA, VT, WA and WI; and countries of Algeria, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Guans, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libyan, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Yemen. | Terminated | August 13, 2013 | 2070 | 2146
Z-0189-1 | 2263 | [6 empty] | Immunosimplicity brand anti-SSA Test kit, Catalog 720-250, Each Kit |  |  | October 31, 2000 | 2000 | DIAMEDIX CORP. | Unexpected equivocal/weak positive results |  | The firm notifed both clients on 3/21/03 during their service telephone/e-mail communications to correct the problem, and they were given a correction work around. A software patch was installed at both clients on 6/25/2003. |  | 113 | 113 | Product was distributed to 18 domestic (user accounts, e.g. labs) in TN, GA, NY, MS, OH, PA, AL, CA, VA, KY, CA, MO, HI, AND TO 5 international accounts as follows: Delta Biologicals, Cipollia and Zanetti, Fiumicino Airport, Roma, Italy; Movaco, S.A.,Iberconco S.A., Loading terminal office NR-9, Barcelona Airport, Spain; Rochem Biocare Columbia S, Deposito Snider y Cia, S.A., Sant Fe de Bogota, Colombia; | Completed |  |  | 113
Z-1160-2008 | 45955 | No | CFP | Immunoreactive insulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ultrasensitive Insulin Assay Kit, Part Number 33410. The product is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of insulin levels in human serum and plasma (EDTA) using the Access Immunoassay Systems. Beckman Coulter, Inc., Chaska, MN 55318. |  |  | September 19, 2007 | 2007 | Beckman Coulter Inc | False negative results when used to test serum samples (as compared to plasma samples). | Other | On September 19, 2007, Beckman Coulter mailed to its consignees an Urgent Product Corrective Action Letter informing them that the affected lots of the product were giving false negative results when used to test serum samples (as compared to plasma samples). The letter instructed consignees to discontinue using the affected lots of the product, and to re-test using an EDTA plasma sample if there was any discordance between the results of the serum sample (if it was obtained using an affected lot of the product) as compared with the clinical presentaion of the patient and the results of other diagnostic tests. The letter also instructed consignees to return, within 10 days, the Response Sheet . |  | 8537 | 8,205 units in the US; 332 in Canada | Worldwide: including USA and Canada. | Terminated | May 09, 2012 | 1694 | 8537
Z-1162-2008 | 46280 | No | BTC | Reservoir bag. | Anesthesiology | Anesthesiology | 510(K) Exempt | 1 Liter Latex-Free Breathing Bags; Reservoir Bag; Vendor Part Number: 720059B, Medline Item Number: 40056; Available in various Anesthesia Circuit Kits and Complete Delivery System kits including: 1) Item Number CDS980230I, NICU-OR CDS; 2) Item Number DYNJ22691, Basin Supply Pack; 3) Item Number DYNJ24724, Spine-Crani Anesthesia Kit; 4) Item Number DYNJAAHY11, Hyperinflation System-UHS-1L; 5) Item Number DYNJAP0010, Anesthesia Circuit Ped-LF; 6) Item Number DYNJAP0011A, Pediatric Circuit-LF; 7) Item Number DYNJAP0012, Pediatric Anesthesia 40 In. CI; 8) Item Number DYNJAP0100, Pedi Breathing Circuit-LF; 9) Item Number DYNJAP0101, Anesthesia Circuit-Ped; 10) Item Number DYNJAP4012, Anesthesia Circuit Ped; 11) Item Number DYNJAP4013, Anesthesia Circuit; 12) Item Number DYNJAP4702A, Anesth Circuit Ped GS Line; 13) Item Number DYNJAP4707, Anesth Circuit 40in Ped-LF; 14) Item Number DYNJAP4722, Anes Cir 40 Ped GSEPW HYBABYH; 15) Item Number DYNJAP4729A, Circ Ped Anes Super-LF; 16) Item Number DYNJAP4730A, Circuit Pediatric Anesthesia; 17) Item Number DYNJAP4732, Anesth Circ Ped 40in 2B/V; 18) Item Number DYNJAP4740, Anesth Circ 40in Ped; 19) Item Number DYNJAP6101, Anesthesia Circuit; 20) Item Number DYNJAP6105, Anesth Circuit 60in Ped; 21) Item Number DYNJAP6106, Anes Circ Ped 60in BI GSE-LF; 22) Item Number DYNJAP6107, Anesth Circ Ped 60in ANG HME; 23) Item Number DYNJAP6108, Anesth Circuit 60in Ped; 24) Item Number DYNJAP6109, Anesth Circ Ped 60in B/V 1L-LF; 25) Item Number DYNJAP6111, Anes Cir Ped 60 GSEPW PALL FIL; 26) Item Number DYNJAP6112, Anesthesia-Peds Circuit; 27) Item Number DYNJAP6113, Circuit Pedi 60in; 28) Item Number DYNJAP6114, Anesth Circ 60in Ped; 29) Item Number DYNJAP6116, Anesth Circ 60in Ped; 30) Item Number DYNJAP6120A, Pediatric Circuit 40in LF; 31) Item Number DYNJAP6124, Anesth Circ Pedi 60in 1L GSL; 32) Item Number DYNJAP6127, Anes Circ Pedi 60in-LF; 33) Item Number DYNJAP6130, Anesth Circ Ped 60in 1L-LF; 34) Item Number DYNJAP6132A, Circuit Pediatric Anesthesia; 35) Item Number DYNJAP6133, Anesth Circuit Peds 60in-LF; 36) Item Number DYNJAP6136, Circuit Peds Anesthesia; 37) Item Number DYNJAP6137, Circuit Pediatric Anesthesia; 38) Item Number DYNJAP7217A, Anesth Circuit 72in Exp Ped BI; 39) Item Number DYNJAP7225A, Circuit Pediatric Anesthesia; 40) Item Number DYNJAP7227, Anes. 90in Exp Pedi -LF; 41) Item Number DYNJAP7233, Anes. Circuit 90in Exp Pedi-LF; 42) Item Number DYNJAP7234A, Anesthesia 72in Pediatric Exp; 43) Item Number DYNJAP7235A, Anesth Circ 72in Ped; 44) Item Number DYNJAP7236A, Anesthesia Pedi 72in; 45) Item Number DYNJAP7237, Circuit Pedi 90in EXP; 46) Item Number DYNJAP7241A, Circuit Pediatric Anesthesia; 47) Item Number DYNJAP7243B, Circ Anesth Ped 2 B/V LFree; 48) Item Number DYNJAP7246C, Circuit Pediatric Anesthesia; 49) Item Number DYNJAP7247, Anes Circ Peds 90in Exp; 50) Item Number DYNJAP7248, Anesth Circ Ped 72in Exp; 51) Item Number DYNJAP7261, Circuit Pediatric Anesthesia; 52) Item Number DYNJAP7262, Circuit Pediatric Anesthesia; 53) Item Number DYNJAP7263, Pediatric Anesthesia Circuits; 54) Item Number DYNJAP9005, Anesthesia Circuit; 55) Item Number DYNJAP9721A, Anesth Circuit Pedi HMEF LF; 56) Item Number DYNJAP9725, Pedi Expand Circuit; 57) Item Number DYNJAP9726, Anesth Circ Ped 90in; 58) Item Number DYNJAP9729, Anesth Circ 90in Exp Ped; 59) Item Number DYNJAP9730A, Circuit Pediatric 60in CORR-LF; 60) Item Number DYNJAP9732A, Circuit Pediatric Anesthesia; 61) Item Number DYNJAP9735, Peds Circuit; 62) Item Number DYNJAP9739, Circuit Pediatric Anesthesia; 63) Item Number DYNJAP9740, Circuit Pediatric Anesthesia; 64) Item Number DYNJAP9743, Circuit Pediatric Anesthesia; 65) Item Number DYNJAPF6003, Anesth Circ 60in Ped Unilimb; 66) Item Number DYNJAPF6005A, Anesth Circ 60in Uni Peds; 67) Item Number DYNJAPF6006, Anes Circ Peds 60in Uni; 68) Item Number DYNJAPF6009, Peds Anesthesia Circuit-LF; 69) Item Number DYNJAPF6011, Circuit Pediatric Anesthesia; 70) Item Number DYNJAPF601 |  |  | December 07, 2007 | 2007 | Medline Industries Inc | Separation from Bushing: The breathing bag may separate from the bushing during use. | Device Design | Consignees were notified by a Sub-Recall Notice letter sent on 12/5/07. The letter instructed users to check their inventory for any affected product and to return a response card indicating the quantity of affected product. The users were instructed to return the affected products to the recalling firm so that they can be reworked. Users who did not wish to return the affected products could request a replacement product and were instructed to destroy the affected products. The letter also informed users that affected products in the manufacturer's inventory were being replaced and these replaced products would contain the label "INSPECTED." For additional information, contact 1-866-359-1704. |  | 77750 | 77,750 bags | Nationwide Distribution. | Terminated | October 16, 2008 | 314 | 571500
Z-1220-2008 | 46281 | [6 empty] | VECTRA GENISYS, Model No. 2761, Chattanooga Group, Hixson, TN 37343, Made in the USA |  |  | December 17, 2007 | 2007 | Chattanooga Corp Inc | Device malfunction: causing electric shock and burn. | Device Design | Beginning on 12/17/2007, Chattanooga Corp. sent two recall notification letters, via UPS mail, to its consignees informing them of the situation. The letters stated that the recall was to the user level and requested that the user perform the software upgrade, which will eliminate the possibility of shock and burn. |  | 4047 | 4047 units | Worldwide Distribution - including USA | Terminated | February 19, 2010 | 795 | 22000
Z-1236-2008 | 45946 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Solar HA Humeral Stem with Plasma Coating 7mm to 17 mm Product Catalog Number: 5351-4307 to 5351-4317. For use as a Shoulder Replacement. Stryker Orthopaedics Made in USA |  |  | November 15, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Delamination: Loosening of the titanium plasma coating on the product was observed on twelve non-distributed units. | Process control | On 11/28/2007, Stryker Orthopaedics sent, via FedEx, Urgent Product Recall notification letters to distributors, hospitals, surgeons and hospitals' Chiefs of Orthopaedics informing them of the firm's voluntary recall of the affected product. They were instructed to discontinue use of the product and to have the units returned to the firm. Distributors were further instructed to notify their customers using the sub-recall letter provided by the firm. The letter to surgeons also informed them to monitor their patients who had received an implant of the product. |  | 915 | 915 units | Nationwide - USA | Terminated | January 25, 2010 | 802 | 915
Z-1188-2008 | 46397 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Reusable Ohmeda Compatible Finger Clip Sensor, Model Number: E403-02, Product Code: PS2225. To monitor oxygen saturation level of patients under doctor's care. Kentec Medical Inc., Irvine, CA 92614 |  |  | May 18, 2007 | 2007 | Kentec Medical Inc | Incompatible components: The Finger Clip Sensors - with a Serial Number prefix of PA514, PA515, YE508, ZC502 and ZG512 - are missing an additional diode and higher resistor in its existing configuration making them inoperable when used to interface with those Ohmeda Tuffsat monitors previously re-configured by Ohmeda to include a detection circuit for diodes. | Process change control | Consignees were notified via telephone on 5/18/07 and an "Important Voluntary Market Withdrawal Notice" letter on 5/23/07. The letter instructed consignees to segregate and return all affected devices. The letter included an "Acknowledgement of Market Withdrawal Notification" form for customer to complete and return to the recalling firm. For additional information, contact 1-800-825-5996. |  | 41 | 41 | Nationwide Distribution including states of AZ, CA, CO, FL, KY, LA, MA, NC, NJ, TX and UT. | Terminated | February 29, 2012 | 1748 | 41
Z-1241-2008 | 46381 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa HeartLab Cardiovascular DICOMstore Version 2.04.40.00 software, Heartlab, Inc., Westerly, Rhode Island 02891. The software is used in the Agfa HeartLab Cardiovascular System. |  |  | June 21, 2007 | 2007 | AGFA Corp. | MIsidentification: The electrocardiogram (ECG) data of one patient is misidentified as the data of another patient. | Software design | On 6/21-26/2007, AGFA Healthcare notified, via telephone, the affected consignees of the situation. On 7/2/2007, affected consignees were sent, via registered mail, an "Urgent Medical Device Correction" letter instructing them to: a) remove the affected software DICOMstore 2.04.40.00 and replace it with the prior software version, DICOMstore 2.04.34.02; and b) review patient ECG reports to determine if any patients were actually affected and to notify the firm. AGFA Healthcare also provided a customer service contact number (877) 777-2432. |  | 43 | 43 units | Worldwide Distributution - USA, including states of AZ, CA, DC, FL, IL, IN, KY, MA, ME, MN, MS, OH, OK, OR, NH, PA, SC, TN, TX, WI, and country of Canada. | Terminated | August 04, 2008 | 410 | 43
Z-1242-2008 | 46382 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Heartlab Cardiovascular Results Management Product: Versions 2.03 and 2.04. Heartlab, Inc., Westerly, Rhode Island 02891 |  |  | July 27, 2007 | 2007 | AGFA Corp. | Incorrect Results: Erroneous echocardiographic measurement values due to mathematical formula being mis-configured. | Software design | On 7/27/2007 to 8/16/2007, AGFA Healthcare notified, via telephone, consignees of the situation. On 7/27/2007, a formal customer notification letter was sent to affected consignees, providing them with a respective listing of patients that were potentially affected. AGFA informed the consignees that the firm would correct any incorrect patient ECG data. AGFA also informed their consignees that they would provide a configuration update to correct the problem. A notification confirmation form to be return to the firm was also attached to the notification letter, which also provided the firm's customer service number: (877) 777-2432. |  | 6 | 6 units | Nationwide Distribution -- USA including states of GA, NH, NY and TX | Terminated | October 10, 2008 | 441 | 6
Z-1167-2008 | 46408 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series Defibrillator with 12 Lead ECG, Automated external defibrillator, Software Version 3.00 and higher; Zoll Medical Corporation, Chelmsford, MA 01824. |  |  | January 15, 2008 | 2008 | ZOLL Medical Corporation, World Wide Headquarters | Incorrect Patient Records: Patient records that have been stored, then later transmitted or printed may not be the correct record. Records that are transmitted or printed immediately following acquisition are correct. | Software design | Consignees were notified by an Urgent Device Correction letter sent on 1/14/08. The letter instructed users to complete and return a response form. Upon receipt of the response form, the recalling firm will have sent a software update. The letter also provided instructions for users to follow in order to avoid the software error before the software update was received. For additional information, contact 1-800-348-9011. |  | 5018 | 5,018 units | Worldwide Distribution including USA and countries of Canada, AE, Australia, Belgium, Chile, Germany, Spain, France, UK, Hungary, India, Italy, Jordan, Kuwait, Netherlands, Russia, Saudi Arabia, Singapore and ZA. | Completed |  |  | 5018
Z-1606-2008 | 47518 | No | MTF |  |  | Immunology | PMA | Dimension Total Prostate Specific Antigen (TPSA) Flex Reagent Cartridge (RF451). Siemens, Newark, DE 19714 |  |  | March 24, 2008 | 2008 | Siemens Healthcare Diagnostics Inc. | False Results: The product may exhibit falsely elevated results. The affected lots demonstrate sporadic variability in recovery of calibrator and quality control; therefore this issue may not be readily detected upon calibration or during routine processing of Quality Control. | Device Design | The recalling firm issued a letter (Urgent Field Safety Notice) dated 3/25/08 to inform their customers to discontinue use of any remaining inventory, retest using an alternate lot for those samples from post-prostatectomy patients which recovered positive TPSA results less than 1.0 ng/nL. The firm has requested that you call the them for no-charge replacement product at 1-800-441-9250 and/or for technical assistance. |  | 1631 | 1631 cartons | The product was shipped to medical facilities nationwide. The product was also shipped to Germany, Singapore, Canada, Japan and AE | Terminated | October 15, 2008 | 205 | 1631
Z-2196-2008 | 48655 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Update Instructions "TH003/08/S" applies to Siemens brand SIMIVIEW 3000, SIMVIEW NT, and MEV ASIM S systems; |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm sent an Update Instruction, "TH003/08/S" for Motion Enable Switch/holder assembly. The switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result. | Software design | On June 25, 2008 all consignees were notified via written letter and on-site visit by the firm's sales representative, informing them that the switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result and providing instructions on the recall. A new software version will be provided. For assistance contact Siemens at 1-925-246-8407. |  | 300 | 300 units | Worldwide and Nationwide | Terminated | December 22, 2010 | 1017 | 300
Z-1306-2008 | 46812 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | IMPAXÂż 6.2.1 Network Gateway Server, License Key. Catalog # ABC code: EXEGE000. The product is used to transfer medical imaging data. AGFA Corp., Greenville, South Carolina 2960. |  |  | December 11, 2007 | 2007 | AGFA Corp. | Failure of Data Transmission: No imaging data could be transmitted by the device, because a key value needed for data transmission was incorrectly entered in the affected units by AGFA Healthcare. | Release of Material/Component prior to receiving test results | On 12/11/07, AGFA Healthcare sent out an e-mail to all affected consignees informing them of the situation. The firm also arranged to install the corrected License Key on all affected units of the product. |  | 12 | 12 units | Nationwide Distribution - USA, including states of CA, GA, IA and SD. | Terminated | July 21, 2010 | 953 | 12
Z-1248-2008 | 46679 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards LifeStent FlexStar Biliary Stent System; Biliary catheter and accessories; Model Number: EX060803CD; Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 05, 2007 | 2007 | Edwards Lifesciences, Llc | Incorrect Expiration Date: Products were labeled with an incorrect shelf-life expiration date. Therefore, the sterility of the device may be compromised. | Packaging process control | Consignees were notified by an Urgent Product Recall letter sent on 12/5/07. The letter instructed users to return any affected product to the recalling firm for replacement products. The letter also advised users that the likelihood of patient complications is extremely low for products that have already been implanted in patients. Users were requested to return a confirmation form after checking the inventory for any affected products. For additional information, contact 1-949-250-3779. |  | 678 | Lot #FR6G0122 - 3 UNITS, FR6G0123 - 2 UNITS, FR6H0193 - 4 UNITS, FR6H0201 - 6 UNITS | Nationwide Distribution | Terminated | January 11, 2012 | 1498 | 10320
Z-1307-2008 | 46813 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Impax 4.5, Impax 5.0, Impax 5.2 and Impax 5.2x Picture Archiving Systems (PACs) with integrated Voxar 3D software; and Impax 5.0 and Impax 5.2 PACs with integrated MediCAD software. The softwares are used to provide digital X-ray images for pre-surgical orthopedic planning. AGFA Corp., Greenville, South Carolina 29601 |  |  | February 15, 2007 | 2007 | AGFA Corp. | Misidentification: The orthopedic planning X-ray images for one patient are misidentified as the images for another patient. | Software design | On 2/15/07, AGFA notified their consignees of the situation by sending them "Urgent Medical Device Correction" letters, instructing them on how to work around the problem until the firm was able to install the upgraded software that corrected the problem. On 12/11/2007, the firm reported that the upgraded software had been installed at all affected sites. |  | 879 | 879 units | Worldwide Distribution - USA including states of AR, AZ, CA, CO, DC, FL, GA, IL, IN, KY, KS, MA, MD, ME, MI, NC, NJ, NV, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA and WA, and country of Canada. | Terminated | June 30, 2009 | 866 | 879
Z-1314-2008 | 47210 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | GentleTouch System-Colostomy/Ileostomy Postoperative Kit; Ostomy pouch and accessories; UPC Code Number: 30003-020938; 2 3/4 in., 70mm; ConvaTec, Skillman, NJ 08558. |  |  | March 07, 2008 | 2008 | ConvaTec | Mislabeled: The carton label reads "For postsurgical use following a urostomy procedure." It should read "For postsurgical use following a colostomy or ileostomy procedure." | Labeling mix-ups | Consignees were notified by an Urgent Product Recall letter sent on 3/7/08. The letter instructed users to return any unused affected product to the recalling firm for a credit. For additional information, contact 1-800-422-8811. |  | 9299 | 9,299 units | Worldwide Distribution including USA and country of Canada. | Terminated | October 21, 2008 | 228 | 9299
Z-1779-2008 | 47833 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips defibrillator, model M5066A. Sold in the US by Philips as an OTC under the brand name Philips "HeartStart On Site Defibrillator" and outside the U.S. by Philips with the brand name Philips "HeartStart Defibrillator". The HeartStart is an automated external defibrillator for use on site and consists of a computer with digital interface and cartridge containing pads. |  |  | April 30, 2008 | 2008 | Philips Medical Systems | Buttons Sticking - The On/Off and Shock buttons on the front panel of the device may stick in place when pressed and fail to respond to the button press, causing a delay in the delivery of a defibrillation shock when the system advises a shock. | Component design/selection | Three separate Urgent - Device Recall letter were sent out on April 30, 2008 via tractable letter using UPS. 1) The end-customer letter advises of the recall, contact information to arrange a replacement of the device and instructions for checking the equipment to insure it is working correctly. "If your HeartStart Home Defibrillator is one of the affected devices or if you have any questions about this recall, please contact Philips Healthcare at 1.800.263.3342, option 5, or visit www.HeartStartHome.com to arrange for a free replacement defibrillator. If you have given your HeartStart to someone else as a gift, please call us promptly and let us know how to contact that person to arrange for the replacement. While waiting for your replacement unit, Philips recommends that you perform a battery insertion test (BLT) of your HeartStart, as described in the Owner's Manual provided with it. Simply remove the battery, wait 5 seconds, reinstall the battery, then press the shock button when directed. After the BIT has finished, press the green On/Off button twice. If either button does not seem to work, please contact Philips immediate." 2) The dealer/distributor and retail letters advises of the recall; request that the dealer/distributor provide end-customer information. "As part of a coordinated effort involving notification of regulatory authorities, other Philips entities, and support personnel, Philips will be contacting all affected customers directly. It is not necessary for you or your sub-distributors to contact customers." 3) The international customer contacts will be made via established Philips sales and distribution channels. Philips will be replacing international units under a Field Change Order Process (to be provided by the firm). In addition, Philips will provide a website that will provide customers with information related to the recall (www.HeartStartHome). |  | 3959 | 3,959 units | Worldwide including USA and countries of Austria, Australia, Canada, Colombia, German, Denmark, Dominican Republic, Spain, France, England, Ireland, Italy, Japan, Netherlands, Norway, New Zealand, Panama, Sweden, Singapore, and South Africa. | Terminated | October 23, 2012 | 1637 | 4596
Z-2019-2008 | 46287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.1 and 2.2 - CXP analysis SW V2.1 single user upgrade, Automated differential cell counter; Part Number 722395; Beckman Coulter; Miami, FL 33196. |  |  | May 04, 2007 | 2007 | Beckman Coulter Inc | Incorrect Display: In certain modes, the protocol will not initially display the correct data in some plots. In addition, when running EWL files, any edited SampleID1 is not automatically updated in all tubes in the panel. | Software design | Consignees were notified of this safety recall by an Urgent Product Corrective Action letter sent on 5/3/07. The letter provided instructions on how to ensure correct data acquisition. Users were requested to return a response form acknowledging receipt of the letter. For additional information contact 1-800-526-7694. |  | 501 | 445 units in the US; 56 units in Canada | Worldwide Distribution including USA and country of Canada. | Terminated | July 11, 2011 | 1529 | 501
Z-1235-2008 | 46833 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | AMS In-Fast Ultra Kit with Polypropylene Suture, REF 72403885. Sterile Transvaginal Bladder Neck Support System. |  |  | January 28, 2008 | 2008 | American Medical Systems | Missing part: The units in two lots of In-Fast Ultra Transvaginal Bladder Neck Support System may contain only one bone screw, but should contain two as indicated on the labeling. Two screws are required to complete the procedure. The missing screw will result in the inability to complete the procedure as expected if there are no additional screws readily available. | Process control | All customers received phone calls on 1/28/08 explaining the error and associated risks and were followed up with detailed letters. The letter "AMS Important Product Recall" dated 1/28/08 described the product and problem. Customers were advised to return the enclosed "Recall Acknowledgement Form" to AMS along with any affected product that they have in inventory. |  |  |  | Nationwide Distribution. | Terminated | October 13, 2010 | 989
Z-1315-2008 | 46806 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Anesthesia Machine; Gas machine for anesthesia or analgesia; Catalog Number: 8604700; Drager Medical Inc, Telford, PA 18969. |  |  | February 05, 2008 | 2008 | Draeger Medical, Inc. | Detachment of Component: The chassis on some machines were manufactured out of specifications. As a result, the caster may break loose from the chassis. | Nonconforming Material/Component | Consignees were notified by an Urgent Medical Device Recall letter sent on 2/7/08. The letter informed users that the recalling firm will contact them to schedule appointments to inspect any potentially affected machines. The letter instructed users to continue using the machines, before the inspection, making sure to carefully follow the warnings in the instruction manual. If a caster on the machine appears unstable before the inspection is conducted, users were advised to contact the recalling firm. For additional information, contact 1-800-543-5047. |  | 196 | 196 UNITS | Nationwide Distribution. | Terminated | September 09, 2008 | 217 | 196
Z-1379-2008 | 46947 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical FLO2 Emergency Non-rebreather High Flow O2 System; Positive end expiratory pressure breathing attachment; Reference Number: P60000; Smiths Medical, Rockland, MA 02370. |  |  | March 04, 2008 | 2008 | Smiths Medical ASD, Inc. | Incorrect Oxygen Level: Device may not provide expected oxygen concentration or the expected flow rate due to incorrectly molded part. Patient will not receive proper amounts of air/oxygen when in need. | Nonconforming Material/Component | Consignees were notified on 2/26/08 via telephone to place the affected products on hold pending further instruction from the recalling firm. A follow-up letter was issued on 3/4/08 advising consignees to return any product in stock for a replacement or credit. Distributors were requested to provide their customer lists to the recalling firm in order to notify their customers of this recall. |  | 650 | 650 | Worldwide Distribution including USA states of CA, FL, IN, KS, MN, MO, NC, NH, NY, PA and TN, and country of Canada. | Terminated | August 20, 2012 | 1630 | 846
Z-1381-2008 | 47219 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio Total Knee Cruciate Retraining Femoral Component; Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis; Product Number: 70-300-7R; Stryker Corp, Howmedica Osteonics Corp, Mahwah, NJ 07430. |  |  | February 08, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Mislabeled: The labeling indicates a Scorpio CR Femoral Component, however, the device inside the package is actually a Scorpio PS Femoral Component. | Labeling mix-ups | Consignees were notified by letter on 2/8/05. For additional information contact 201-831-5825. |  | 7 | 7 units | Worldwide Distribution including USA and countries of Canada, Mexico, Italy, The Netherlands and the UK. | Terminated | July 16, 2008 | 1254 | 12
Z-1233-2008 | 46800 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Infrared (IR) Hand Control System, an optional wireless control component of the Arjo Maxi Sky 600 and BHM V4 Ceiling Patient Lifts; the hand control system is composed of two components, part 700.13640.33 - IR hand control, and part 492.00038 - IR receiver board |  |  | February 20, 2008 | 2008 | Arjo, Inc. | The ceiling lift may not stop lateral movement after releasing the left (or right) action button on a ceiling lift equipped with an infrared hand control. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 2/15/08 on the BHM Medical letterhead to the end user accounts who received the affected IR remotes, advising them of the potential for the remote to become difficult to operate or to create uncontrolled movement of the ceiling lift. The accounts were instructed to remove the IR remote from service until the field correction is issued by BHM Medical. Instructions for removal of the IR equipment were provided in the Safety Advisory Notice included with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating the number of units requiring correction. |  | 25 | 25 units | Nationwide Distribution. | Terminated | September 11, 2010 | 934 | 25
Z-1917-2008 | 48084 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Celsus Laboratories Heparin Lithium Lyophilized; Lithium Heparin for use in blood collection tube devices and in vitro diagnostic devices. |  |  | March 31, 2008 | 2008 | Celsus Laboratories Inc | Contamination: The firm learned that up to 2.2% oversulfated chondroitin was found in a batch of crude heparin they received from their supplier. | Nonconforming Material/Component | Celsus notified their customers by e-mail on April 24, 2008, explaining that oversulfated chondroitin was found in a batch of crude heparin they received from their supplier. Celsus recommended the device manufacturers quarantine material from the specified lots. |  | 211818 | HL-31707=11,542g, HL-31807=12,836g, HL-32708=8,374g, HL-32808=12,128g &amp; HL-32908=5,000g | Worldwide Distribution. The Lithium Heparin was sold to companies in USA (Nebraska, New Jersey, Ohio), Puerto Rico, Austria, and Denmark. | Terminated | September 21, 2012 | 1635 | 211818
Z-1308-2008 | 46814 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit; Programmable implanted infusion pump; Model Number: 8596SC; Packaged with catheter model number 8731 or 8731SC. |  |  | January 09, 2008 | 2008 | Medtronic Neuromodulation | Incorrect Connector Pin: The Medtronic 8596SC Intrathecal Catheter Pump Segment Revision kits were packaged with the incorrect connector pin. | Packaging process control | Consignees were notified by an Urgent Medical Device Recall letter sent in January 2008. The letter informed users that the recalling firm would contact them regarding removing the affected products from their inventory. Users were requested to return a response form indicating the number of affected product. For additional information, contact 1-800-707-0933. |  | 18 | 18 | Nationwide Distribution including states of MI, OR, IN, NY, FL, PA, MN, TX, CO and UT. | Terminated | October 08, 2010 | 1003 | 18
Z-1479-2008 | 46888 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Sequoia Spinal System Surgical Kits; Pedicle screw spinal system; Non-sterile (kits intended to be sterilized prior to use); Rx Only; |  |  | February 13, 2008 | 2008 | Abbott Spine | Screw Head Dissociation: The polyaxial screw head may dissociate from screw shaft during surgery. Also, the polyaxial screw driver is experiencing difficulty (broken/bent driver shafts, difficulty attaching screws to the driver and jamming of the collet.) | Device Design | The recalling firm notified sales reps regarding the recall by phone on 02/13/08. The sales reps were instructed to discontinue kit use in surgeries and to return inventory. Surgeons who had used the Sequoia System in surgery were notified of the recall via letter on 02/20/08. Surgeons were asked to monitor for post-operative dissociation in implanted patients and to notify Abbott Spine sales reps if any incidences occur. |  | 38 | 38 surgical kits. | Nationwide Distribution including states of AR, AZ, CA, CT, CO, FL, IL, MD, MI, MN, NC, OH, OK, OR, PA, TX, VA and WI. | Terminated | July 28, 2008 | 166 | 38
Z-1474-2008 | 47225 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Medline EndurO2 Oxygen Concentrator without Oxygen Monitor; Portable oxygen generator; Reorder Number: HCS021; Packaged in 1 unit per case; Product is an oxygen generator designed to provide oxygen to patient who has been prescribed oxygen therapy by a physician. |  |  | March 04, 2008 | 2008 | Medline Industries Inc | Overheating of Components: There is the potential for an electrical component in the oxygen concentrator to overheat, causing damage to the unit and making it nonfunctional. | Device Design | Consignees were notified by a letter sent on 3/4/08. The letter informed users that the recalling firm will replace the affected units with new units or issue a credit for the returned recalled units. Distributors were requested to notify their accounts of the recall. The letter requested users to complete and return an enclosed response sheet, indicating the number of units to be replaced or returned for credit. For additional information, contact 866-359-1704. |  |  |  | Nationwide Distribution | Terminated | June 17, 2010 | 835
Z-1902-2008 | 47409 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Levitronix CentriMag Back-Up Console |  |  | March 17, 2008 | 2008 | Levitronix, Llc | Stops pumping: Interruption of CentriMag System support may occur when using a Valleylab Force FX-C Electrocautery Unit | Device Design | Levitronix issued an Urgent - Voluntary Device Correction letter on 3/17/08 mailed to the US distributor and requested to contact their accounts. The letter revises the warning in the CentriMag Console Manual. Levitronix is issuing revised manuals.On March 13, 2008 the firm emailed and mailed (FedEx) to the one Canadian Hospital and Health Canada was informed of the Voluntary Advisory Notice via email on Wednesday, March 19, 2008.Rest of the World: Wednesday, March 12, 2008 throughMonday, March 17, 2008. The Notice was issued as aVoluntary Advisory Notice after review and follow up with theUK MHRA. On 7/24/08 the firm issued a revised Recall Letter to advise users not to use Valley Lab Force FX-C or SSE2L Electrosurgery device with the firm's CentriMag Blood Pumping System and include the significant risk to health. The firm revised their letter and Black Box Warning and reissued the letters to the hospital, surgeons, etc. on 7/24/08.On 2/5/09 the firm issued a "Follow-Up to Urgent Device Correction of 3/17/08 " to advise users of the software upgrade to be installed in the CentrigMag Primary software Version 3.0 and CentriMag Back Up console , software version 2.0. Site visits will be made by Levitronix between Feb 2009 and April 2009 , software upgrade installed and new Operating Manuals provided. |  | 237 | 237 units | Worldwide Distribution: USA, Canada, Switzerland, UK, Italy, France, Greece, Spain, Slovenia, Czech, Turkey, South Africa, Argentina, Chile, Panama, Singapore, Taiwan, Israel, and Saudi Arabia. | Terminated | July 06, 2010 | 841 | 867
Z-1372-2008 | 46212 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vascular Drape, Surgical drape and drape accessories; Reference Number: EF 00-902776-01; Packaged in a box of 10. |  |  | December 20, 2007 | 2007 | OEC Medical Systems, Inc | No Premarket Notification: Product does not have FDA 510(k) approval for use. | PMA | Consignees were notified by an Urgent Recall Notice letter on 12/20/2007. The letter instructed users to stop using and to dispose of the recalled drapes. For additional information, contact 1-800-874-7378. |  | 158 | 158 units | Nationwide Distribution including states of AL, CA, FL, MA, MD, MO, NJ, NY, PA, OR, SC and TX. | Terminated | July 22, 2008 | 215 | 158
Z-1928-2008 | 48100 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Bio-Medicus Femoral Cannula/Introducer, for Cardiopulmonary Bypass. Catalog Numbers: CB96345-023, CB96345-023, and CB96345-029. |  |  | May 09, 2008 | 2008 | Medtronic Perfusion Systems | Medtronic is recalling selected Carmeda-coated cardiopulmonary bypass (CPB) products (e.g., oxygenators, reservoirs, pumps, cannulae and tubing packs) that were manufactured with contaminated heparin. Medtronic's heparin supplier recently reported that it had provided some batches to Medtronic that were contaminated with oversulfated chondroitin sulfate (OSCS). | Nonconforming Material/Component | Medtronic issued a Press Release on 05/07/2008. An Urgent Medical Device Recall letter was sent to customers on May 07, 2008. The letter identifies affected devices, explains the situation and requests that the impacted product be returned to Medtronic for disposal. The end user will be requested to complete a Recall Certificate acknowledging the receipt of this information. |  | 46 | 17 US, 29 OUS | Worldwide Distribution: USA, Australia, Belgium, Canada, Czech Republic, France, Italy, Japan, Norway, and Taiwan. | Terminated | May 08, 2012 | 1460 | 844
Z-1566-2008 | 47203 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AxSYM Ultrasensitive hTSH II Master Calibrators |  |  | March 06, 2008 | 2008 | Abbott Diagnostic International, Ltd. | High Control value out of range after calibration. --AN INCREASED FREQUENCY OF HIGH CONTROL VALUES BEYOND THE UPPER LIMIT OF THE RANGE SPECIFIED BY THE PACKAGE INSERT. | Nonconforming Material/Component | The firm issued Product Recall letters dated February 22, 2008, requesting remaining units of the affected product be discontinue using and destroyed. A Customer Reply Immediate Action Required letter was enclosed to indicate acknowledging receipt of the product recall letter and to provide information on the number of kits destroyed before March 7, 2008 and faxed to 1-800-777-0051. |  | 882 | 882 total units | Worldwide Distribution --- USA including states of AZ, CA, CO, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, WA, WI, WV, and WY, and Puerto Rico, and countries of Canada, Mexico, Brazil, Chile, Germany, Japan, and Hong Kong. | Terminated | January 25, 2012 | 1420 | 882
Z-1571-2008 | 47253 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | EUB-5500 VISION 5500 Diagnostic Ultrasound Scanner w/Step 5 (V05-7) Software. Imaging device intended to provide the physician with physiological and clinical information. |  |  | February 25, 2008 | 2008 | Hitachi Medical Systems America Inc | Miscalculation reading. A software error in the firm's EUB-5500/HI VISION diagnostic scanning system causes a miscalculation of the left ICA/CCA ratio when using the Carotid Calculation package for patient scans. | Software design | A Device Correction Letter was issue by the firm on 02/25/2008. The letter asks the affected customers not to use the value of ICA/CCA on the Measurement Report for diagnosis, but instead to follow the Immediate countermeasure detailed in the 'Note in Use' document which accompanies the Device Correction letter. The customers are asked to follow the Immediate Countermeasure until they receive a Temporary Countermeasure to correct the ICA/CCA ratio calculation. The Temporary Countermeasure will include: 1) a Disk to import a new carotid measure package; 2) an Installation booklet detailing how to import the new carotid measurement package and optional customization tools; and 3) an Operator manual addendum Upon release of the countermeasure, the firm's Applications or Servcie personnel shall perform the corrective action as identified in the Service Pack (A) Installation Instructions. Finally, upon release of the next software revision, service personnel will replace the software at each customer location. |  | 20 | 20 units | Nationwide Distribution --- including states of OH, FL, CA, TX, DE, VA, NY, MO, KS, NJ, MD and ID. | Terminated | February 24, 2012 | 1460 | 20
Z-1376-2008 | 46708 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard DCM CR Tibial Bearing, 14MM x 79/83MM, CR STD bearing, sterile, Biomet Orthopedics, Inc., Warsaw, IN; Part 183464. |  |  | February 12, 2008 | 2008 | Biomet, Inc. | Mislabeled: The packaging of the two products was mixed up, resulting in the products being labeled with an incorrect size. If implanted, the product may not fit properly, which could require repeat surgery. The problem could also result in a delay in the procedure. | Packaging process control | Distributors were notified via recall letter (Urgent Medical Device Removal Notice) dated 2/12/08 to locate the devices in their possession and at hospital sites, discontinue use, and return them to Biomet. Hospitals having received the devices will be provided with a recall letter dated 2/12/08 which describes the nature of the problem and the removal of these instruments. A third letter dated 2/12/08 was provided for implanting physicians. |  | 8 | 8 | Worldwide: USA and Australia | Terminated | December 22, 2009 | 679 | 12
Z-1587-2008 | 47400 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon TS Femoral Trial; |  |  | March 05, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Sharp edges will cause a glove to tear and cut someone.Sharp edges on the slots machines in select lots of the Triathlon TS Femoral Trials, sizes 1-8, have been discovered. | Process control | Stryker Branches/Agencies were notifed by letter (Urgent Product Recall) on March 5, 2008. Reps were asked to examine inventory and visit hospitals that were provided with the device to retrieve product. Reps were instructed to retrieve all of the affected product lots and return them to their branch or agency warehouse for reconciliation. All products should be reconcilied on the Product Accountability Form that is attached to the letter and sent via Inter-Org Transfer back to Stryker Orthopaedics. Questions and re-orders of the product that is being returned should be directed to 201-831-5825. |  | 2 | 2 | Nationwide Distribution --- including states of CT, NJ, PA, VA, NC, GA, FL, AL, MI, MN, IL, TX and AZ. | Terminated | October 25, 2010 | 964 | 104
Z-1373-2008 | 46751 | Yes | MRM |  |  | Cardiovascular | PMA | Lumax Cardiac Resyncronization Therapy Defibrillator (CRT-D); Implantable, dual-chamber defibrillator; Model Number: HF-T; Catalog Number 355263; Software Version 603.U; Catalog Model Number 359251; ICS 3000 units are composed of three Catalog/Model Numbers: 349527, 349530, and 355547. The model 349530 is the component which is loaded/updated with software. |  |  | January 10, 2007 | 2007 | Biotronik Inc | Timing Feature Error: An implementation error allows the usage of the V-V timing feature, which is not approved for use in the U.S. | Software design | Consignees were notified by a letter sent 1/10/07. The letter advised users to stop usage of the affected software. Users were instructed not to implant additional devices until the software program has been corrected. The recalling firm requested that all affected software be scrapped or returned to them. Users were requested update the software with a corrected version as soon as it becomes available. |  | 53 | 53 of the 603.U discs distributed. | Nationwide Distribution including states of AL, AZ, CA, CO, DE, FL, GA, IL, LA, MI, MO, NC, NJ, NV, NY, OH, OR, PA, SC, TX and WA. | Terminated | May 05, 2011 | 1576 | 53
Z-2005-2008 | 48101 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Trillium AFFINITY NT, 511T, Hollow Fiber Oxygenator with Trillium Biopassive Surface, with Plasma Resistant Fiber. MODEL Numbers: 95215, 95216, 95217, 95249, 95250, 95252, 0H46R28, 0M81R41, 0R35R13, 0Y95R22, 1702R51, 1A11R16, 1A19R26, 1B22R5, 1F66R12, 1G47R4, 1K86R3, 1K93R5, 1L71R20, 1M05R12, 1P07R6, 1X24R, 2892R18, 2D64R12, 2E55R9, 2N90R8, 2V85R4, 3B89R12, 3C98R7, 3P11R4, 3R39R2, 3R42R14, 4B46R16, 4E58R16, 4F13R2, 4J01R7, 4Q39R1, 4R68R5, 4R68R6, 4W65R, 4X44R7, 5633R30, 5C89R2, 5D83R, 5R69R1, 5U35R, 9059R20, 9243R19, CB0A12R15, CB2T58R4, CB3L15R5, CB4D99R4, CB4E05R6, CB4J24R2, CB4J67R1, CB4K47R1, CB4K64R5, CB5B53R, CB5C01R2, CB5C82R9, CB5P29R4, CB5Q88R, CB5S97R2, DP38T, HY0W55R17, HY0Y86R17, HY1709R34, HY1C73R16, HY2F13R22, HY2F13R23, HY2J04R2, HY2K12R5, HY2K16R5, HY2K51R14, HY2K60R3, HY2Q31R1, HY2Q32R2, HY2T49R3, HY2U43R5, HY2U98R5, HY2W85R9, HY3B29R13, HY3C53R1, HY3E77R13, HY3H52R13, HY3L42R12, HY3L78R1, HY3M41R4, HY3M43R1, HY3M45R1, HY3Q11R14, HY3T77R2, HY3U59R4, HY4A96R3, HY4B44R, HY4B50R, HY4E05R8, HY4E66R5, HY4V82R2, HY4X31R7, HY4Y15R11, HY4Z59R7, HY4Z62R5, HY5C74R15, HY5D07R, HY5D70R1, HY5L38R1, HY5N41R2, HY5N42R3, HY5P52R, HY5Q00R, HY5S13R, HY5S33R2, JP1Z53R11, M310113Q, M310319A, M310523D, M310525C, M310531G, M310538A, M310543F, M310824V, M310852C, M310854A, M311028K, M311224A, M311818C, M320234E, M320410V, M320608K, M321014K, M321505L, M322318C, M330112G, M330344D, M330607H, M332108C, M332607G, M333405W, M333809B, M334007C, M334007C, M334008D, M334209A, M335615B, M336202V, M337412A, M337414A, M337415A, M337417A, M337418B, M338001P, M340305C, M340615E, M343105G, M380324D, M390206E, M390322C, M390508D, M390824C, M391817L, M391817M, M391908B, M392814A, M392815A, M395216B, M395217B, M395319B, M396010D, M397013G, M397912B, M398309C, M399510A, M399511A, M410811K, M430737A, M443307B, M443822B, M443823B, M447005R, M460524D, M460530B, M460531B, M460532B, M462007B, M470233C, M470237C, M470945S, M470945T, M490750B, M495512K, M580211B, M580416B, M580418E, M720407F, M720409F, M720714C, M720718D, M720720B, M930401B, M930402B, M960116A, M960119A, M973502E, M973901C, M999308D, P391401B, TL0A42R21, TL0W56R7, TL1E95R12, TL1J50R3, TL2C82R4, TL2H89R7, TL2J77R, TL2K45R6, TL2L66R5 TL2L67R, TL2M25R6, TL2P63R1, TL2P93R5, TL2P94R9, TL2R00R3, TL2R66R, TL2S03R4, TL2S21R1, TL2S24R3, TL2S80R24, TL2S80R30, TL2S80R30, TL2T46R7, TL2U42R13, TL2U59R, TL2W21R, TL2W22R5, TL2W90R11, TL2Z00R4, TL2Z34R, TL2Z35R, TL3A08R16, TL3B47R8, TL3C66R4, TL3C69R16, TL3E97R5, TL3E98R7, TL3E99R2, TL3F17R2, TL3F33R1, TL3G03R2, TL3G89R1, TL3H76R12, TL3H85R1, TL3H86R6, |  |  | May 07, 2008 | 2008 | Medtronic International Trade Inc | Medtronic is advising that certain Trillium-coated Medtronic cardiopulmonary bypass (CBP) products (e.g., oxygenators, reservoirs, pumps and tubing packs) were manufactured with heparin batches contaminated with oversulfated chondroitin sulfate (OSCS). Based upon currently available information, Medtronic has determined that Medtronic Trillium-coated products can continue to be used in cardiopulm | Nonconforming Material/Component | Medtronic Cardiovascular Customers were sent the May 05, 2008 Urgent Medical Device Corrective Action letter to advise them that certain Trillium-coated cardiopulmonary bypass products were manufactured with heparin which may be contaminated with oversulfated chondroitin sulfate. The letter discussed the risk evaluation and gave recommendations to healthcare professionals for patient care. Affected devices can continue to be used in cardiopulmonary bypass procedures. |  | 59500 | 59500 | Worldwide Distribution, including USA. | Terminated | December 01, 2008 | 208 | 59500
Z-2037-2008 | 48755 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly QS) System - PFILS (Patient File Server) software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit. |  |  | June 30, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Patient information mix-up: When the PFILS application, configured with a backup server has experienced a network interruption, PFILS verify could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 7/20/07 of the problem with the PFILS application that could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. The letter provided the users with a workaround and provided them with a Frequently Asked Questions document which listed the availability of a software patch to fix the problem. A follow-up letter dated 6/27/08 was sent on 6/30-7/2/08 to those accounts who had not requested a PFILS Hotfix CD-ROM or had not had their software recently upgraded. The letter a copy of the 7/20/07 letter, the most recent FAQ, the PFILS Hotfix CD (PFIL67_66) and Installation Procedure for the PFILS Patch instructions for installation of the software patch. The accounts were advised that they can request an analysis of existing patient records to determine whether they have been affected by the PFILS issue, by submitting a non-critical Technical Support Request on the GE Centricity Perinatal Web site: http://www.gehealthcare.com/usen/security/cis/stupgrades.html, indicating "PFILS ANALYSIS REQUEST" in the Description of Issue field. Any questions were directed to Technical Support at 800-433-2009. |  | 1031 | 1,031 units | Worldwide: USA, Puerto Rico, Canada, and China. | Terminated | September 06, 2010 | 798 | 1031
Z-1557-2008 | 47456 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medi-Lifter III Plus Total Lift, Model 7000-2000 and Summit Total Lift, Model 4800-2000; patient lifts with mast scale (600 lbs. capacity); |  |  | April 10, 2008 | 2008 | Arjo, Inc. | Scale could break and fall potentially causing injuries to someone who is on it. There is a potential for the mast scale load cell ssembly to fracture and allow the mast/boom to fall. Should a patient be in the lift at the time, the patient could fall and sustain injuries. | Device Design | Arjo sent Urgent Device Recall Customer Notification letters dated 4/1/08 on the BHM Medical letterhead to the end user accounts who received the affected patient lifts with mast scales, advising them of the potential of fracture at the load cell assembly which might cause the scale to fall. Should a patient be in the lift at the time, the patient could fall, potentially sustaining injuries. The accounts were instructed to immediately stop using the patient lifts. A BHM representative will contact the accounts within 5 days to arrange for the correction of the lifts. The accounts will be given the option of returning the lifts for repair (with shipping costs covered by BHM), or purchasing a replacement lift at a significantly reduced price. Any questions or concerns were directed to the BHM Service Department at 1-819-868-0441 or service@bhm-medical.com. |  | 37 | 37 lifts | Nationwide Distribution --- including states of California, Florida, Illinois, Michigan, Minnesota, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Washington, and Wyoming. | Terminated | November 18, 2009 | 587 | 37
Z-1549-2008 | 47474 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medex MX821L 72 inch (182.9cm) Administration Sets, Standard Drip, Sterile Single Use, Latex Free device packaged 25 units per case. |  |  | March 24, 2008 | 2008 | Smiths Medical Asd Inc | Misbranded/mispackaged device was distributed. Product labeled to contain 72 inch Standard Drip ( IV) Administration Sets, actually contained 60 inch (IV) administration sets with a' Y' site. | Packaging process control | On 3/24/2008, the firm issued separate URGENT: PRODUCT RECALL NOTIFICATIONs sent via certified mail to their distributor and medical customers. Both recall notifications inform the customers of this labeling error and ask that the customers: 1) check their inventories for the presence of the suspect product; 2) complete and return the attached 'Product Recall Notification Response Form'; and 3) make arrangements to return any unused affected product to the recalling firm in exchange for a credit or replacement product. In addition, the distributors are asked to either provide the recalling firm with a list of their sub-account customers so that Smith Medical can contact the enduser customers and notify them of the recall, or as an alternative measure, the distributor can notify their branches and customers directly, via their own certified letter. |  | 350 | 350 units | Worldwide Distribution --- USA including states of PA, NJ, IL, MN and ND and country of Canada. | Terminated | September 30, 2010 | 920 | 350
Z-1460-2008 | 46919 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Acuity with Conebeam Computed Tomography; Radiation therapy simulation system; Model Number: H77. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | Incorrect Image Orientation: The image orientation tag may not be set correctly if the patient has been scanned with an orientation other than Head First Spine. When such images are imported to a treatment planning system, they may appear with either a mirrored or a rotated view. | Software design | Consignees were notified by an Urgent Medical Device Correction letter issued on 02/05/2008. The letter provided recommended actions for users to follow in order to avoid the software error. The letter informed users that a service pack is being created to correct the issue and that the recalling firm will notify users when it becomes available. For additional information, contact 888-827-4265. |  | 16 | 16 UNITS | Worldwide Distribution--USA including states of CA, MA and PA, and countries of UK, Chile, Belgium, Canada and Germany. | Terminated | September 30, 2009 | 636 | 16
Z-1550-2008 | 47334 | No | FZW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Dermacarriers II skin graft carriers, 1.5 to 1 dermacarrier, Rx, latex free, sterile, Zimmer, Dover, Ohio; REF 00-2195-012-00. |  |  | April 03, 2008 | 2008 | Zimmer Inc. | Sterility compromised. Lack of assurance of sterility, as the packages may not have been sealed on the manufacturer's end. | Packaging process control | Consignees were notified via letter (Urgent: Medical Device Correction) dated 4/3/08 to identify their stocks on hand, verify package sealing, to discard any non-sealed packages, mail the Business Reply Card regardless if there are any remaining affected products. |  | 99345 | 99345 of all units | Worldwide Distribution --- including USA and countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, Ecuador, England, Finland, France, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Iraq, Israel, Italy, Japan, Jordan, Kenya, Largos, Lebanon, Lithuania, Malta, Mexico, Morocco, Netherlands, Nigeria, Panama, Peru, Poland, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunis, Turkey and Venezuela. | Terminated | December 17, 2009 | 623 | 99345
Z-1569-2008 | 47236 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect Free T4 MasterCheck; a single (specified) analyte control (assayed and unassayed); list 6C50-05. |  |  | February 27, 2008 | 2008 | Abbott Laboratories | Testing material becomes unstable at the recommeded storage temperature. Some of the Architect Free T4 MasterCheck and Total T4 MasterCheck lots are not stable at the recommended storage condition of 2 to 8 degrees Celsius through expiration dating. These lots may exceed the suggested tolerance limits listed in the Reportable Range/Calibration Verification section in Appendix B of the Architect | Nonconforming Material/Component | Abbott sent a Product Recall letter dated 2/27/08, to all customers who received the affected product lots, informing them that the storage conditions are being changed to -10 degrees C or colder for all lots of material manufactured after lot numbers 82130M and 84110M until the cause has been determined and a corrective action has been implemented. All new shipments will contain additional instructions indicating that the ARCHITECT Free T4 and Total T4 MasterChecks are to be stored frozen until first use. Customers were requested to check their inventory for lot numbers 82130M and 84110M, discontinue use of the lots and destroy any remaining inventory. For lots 82140M, 82150M and 84120M, if they know the date the lot was received or the date it was thawed, they may continue to use lots if they have been stored at 2 to 8 degrees C unopened for less than 60 days. If the material has been stored at 2 to 8 degrees C for greater than 60 days or they do not know the number of days it has been thawed, they were instructed to discard the material. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the kits. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT in the U.S. Customers outside of the U.S., were requested to contact their local customer support representative. |  | 197 | 197 units | Worldwide Distribution --- USA including state of Puerto Rico and country of Canada. | Terminated | November 05, 2008 | 252 | 325
Z-1556-2008 | 47347 | Yes | NIO |  |  | Cardiovascular | PMA | CORDIS S.M.A.R.T.Âż Control" Nitinol Stent System |  |  | March 10, 2008 | 2008 | Cordis Corporation | Data transcription error. The 10 affected lots of Catalog Number C10040SL were manufactured with a prior version of the guidewire lumen material, which is not validated with the current manufacturing process. This condition may result in a separation of the device. If a separation was to occur, and the separated fragment of lumen material was not retrieved on the guidewire, blockage of the vessel | Mixed-up of materials/components | A letter (RE: Medical Device Recall) dated March 11, 2008, and Acknowledgement Form was sent overnight to the following 6 individuals in each account: Dir. Cardiac Cath Lab, Dir. Radiology Services, Dir. Vascular Services, Risk Manager, Materials Dir., Hospital Administrator; requesting that the affected product be set aside, and that the form be reviewed, signed, and faxed back to Cordis. Affected product is to be either returned or the local sales rep should the contacted to facilitate the return of the affected product. All product will be replaced. Customer Service can be contacted if there are additional questions relative to product return, replacement, or availability. |  | 398 | 398 units | Worldwide Distribution --- including USA and countries of Japan and Canada. | Terminated | October 01, 2009 | 570 | 398
Z-2138-2008 | 48974 | No | MVR | Implantable clip. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | GEM 2753 Microvascular Anastomotic Coupler, 2.5 mm, REF GEM2753/I, P/N 5111-00250-010, Rx only, Irradiation Sterilized |  |  | May 14, 2008 | 2008 | Synovis Micro Companies Alliance Inc | Coupler rings may slip out of the delivery toolcaused by jaws larger than specification. | Device Design | Urgent Recall Notification was sent via Fed Ex on May 15, 2008 to 432 hospitals. The letters were addressed to the attention of the Risk Manager. The recall notification instructed the hospital to immediately discontinue use, examine their inventory and, if found, to remove the product and quarantine it for return. A second notification was sent on July 21, 2008 to non-responding consignees. |  | 3812 | 3,812 units | Worldwide Distribution --- including USA and countries of Canada, Japan, UAE, Korea, South Africa, Australia, Taiwan, Norway, Netherlands, Sweden, Greece, Norway, Turkey, Portugal, Taiwan, Denmark, Italy, Germany, Japan, Spain, Austria, and Great Britain. | Terminated | April 09, 2010 | 695 | 3812
Z-0274-2008 | 47235 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D Classical R&F; System with control room PC; Image-intensified fluoroscopic x-ray system; ; Model Number: 5117866-27. |  |  | August 01, 2007 | 2007 | GE Healthcare | Inaccurate cassette size display reading/ Inaccurate mAs reading: 1. When the device is in PBL mode and a larger film cassette is inserted into the table Bucky (replacing a smaller cassette) after a protocol is selected on the console, the collimator will not adjust to the larger cassette size. The collimator stays at the original smaller cassette size and allows the operator to take an exposure. | Software design | Consignees were notified by two "Product Safety Notification" letters on July 10, 2007. Each letter addressed one of the two safety issues with the affected product. The letters provided safety instructions on how avoid the software error and informed users that their representative would contact them to schedule a visit for a software correction installation. |  | 101 | 101 | Worldwide Distribution --USA including states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, ID, IL ,IN, LA, MA, MD, MI, MO,MS, NC, ND, NE, NH, NJ, NM, NV, OH,OK, OR, PA, SC, SD, TN, TX, UT, VA, VT and WA; USA territory of: PR and countries: EGYPT, SAUDI ARABIA, MACEDONIA, FRANCE and GREAT BRITAIN. | Terminated | September 10, 2008 | 406 | 101
Z-1459-2008 | 46852 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 MultiLeaf Collimator, Medical charged-particle radiation therapy system; Software Version 7.0; Product Code: H56; Accessory to Medical Linear Accelerator. |  |  | January 30, 2008 | 2008 | Varian Medical Systems Oncology Systems | Dose Calculation Error: Software anomaly may result in failure of a MLC leaf to reach planned position, potentially resulting in misadministration of dose to a patient. | Software design | Consignees were notified by a Medical Device Notification letter sent on 01/29/2008. The letter advised users to cease use of the device until a corrected version of the software is available. A representative will contact users to upgrade the application once the revised software has been validated and released for distribution. |  | 3 | 3 units subject to correction | Nationwide Distribution including states of NC, FL and CA. | Terminated | August 18, 2010 | 931 | 3
Z-1190-2008 | 46753 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris C1160 Universal Flexible Processing Tray, for use with STERIS SYSTEM 1 Processor, used for sterilization of flexible endoscopes. |  |  | February 19, 2008 | 2008 | Steris Corp | Failure to effectively sterilize: Under certain conditions, the design/operation of the C1160 Tray may compromise the ability of the SYSTEM 1 processor to correctly assess the state of the high pressure pump during the Sterile Cycle. The L6 sensor fails to alarm when the high pressure pump is not working properly. | Device Design | Steris sent an Urgent Safety Alert/Recall letter dated 2/8/2008, to all customers. The letter stated that any device sterilized in the C1160 Tray should not be used clinically prior to a completion of a successful diagnostic cycle. Until further notice, you must run a diagnostic cycle after each sterile cycle involving a C1160 tray or alternately you must discontinue use of these trays. Steris service technicians will schedule visits to correct the problem. |  | 28766 | 28,766 units | Worldwide Distribution, including USA, Puerto Rico, Canada, and Japan. | Terminated | February 24, 2012 | 1466 | 28766
Z-1191-2008 | 46788 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | INTEGRA NeuroSensor Probe, REF NS-P, component of Cerebral Blood Flow and Intracranial Pressure Monitoring Kit, Integra NeuroSciences |  |  | January 28, 2008 | 2008 | Integra Neuro Sciences | Incorrect pressure reading: Due to a manufacturing error, the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically, the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg, that could result in a maximum Insertion Error of +/- 6mmHg. | Process design | Integra NeuroSciences sent a Recall Notification with a Recall Acknowledgement form to all customers with product currently in their possession through an Integra Field Sales Repesentative on January 28, 2008. The notification advises that the firm has determined that due to a manufacturing error, the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically, the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg that could result in a maximum Insertion Error of +/- 6mmHg. Customers are asked to examine their inventory immediately to determine if they have any NeuroSensor Model NS-P's lot #100063 on hand and if so to not use the product. Customers are advised to contact their local Integra representative for further assistance. The letter advises that the lot/serial number is located on a white 1/2" x 1/2" label on the large beige optical connector, and can readily be seen through the bottom of the product tray packaging. The letter also advises that this recall pertains to NeuroSensor Model NS-P's, lot # 100063 -all serial numbers. A response form is included.In addition, the Field Sales Representatives have been instructed to physically remove all lot 100063 of NeuroSensor Probes from inventory and provide replacement product. |  | 43 | 43 units | USA Distribution: CA, GA, NM, TX, NE, WA, and NY. | Terminated | April 22, 2010 | 815 | 43
Z-1638-2008 | 47428 | No | IMG | Ultrasound and muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | VECTRA GENISYS, electromechanical therapeutic ultrasound and powered muscle stimulation device, Model No. 2761, Chattanooga Group. |  |  | February 28, 2008 | 2008 | Chattanooga Corp Inc | Device may fail to administer therapy during use. | Component design/selection | Chattanooga Group sent an Urgent Medical Device Recall letter dated 02/29/2008 to consignees to advise them of problem and the recall. The first group of letters was sent to the distributors via certified mail. The letter stated that the devices were failing and the recall was to the User level. The firm is recalling the devices for repair and calibration. The firm followed with a user notification, by US Mail not certified beginning on March 11, 2008. |  | 479 | 479 devices | Nationwide Distribution. | Terminated | February 19, 2010 | 722 | 1565
Z-1181-2008 | 46846 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira LifeCare PCA Label Utility, Version 1.0 CD. The installation/user guide is entitled "LifeCare PCA Barcode Label Utility Program." This software program is designed to allow entry of each drug concentration into a formulation library on a pharmacy computer for standardized printing of barcode labels for compounded drug product vials. |  |  | January 03, 2008 | 2008 | Hospira Inc. | Incorrect dosage labels created: Customers have experienced independent changes in drug concentration (mg/mL to mcg/mL and vice versa) and expiration dates, and the addition of trailing zeros to the barcodes. | Software design | Hospira sent letters dated 1/3/08 to the immediate recipients of the product via FedEx overnight letter on the same date. The letters informed the accounts of the reports of independent changes in drug concentration while editing a saved formulation using the Hospira LifeCare PCA Label Utility Software CD that was provided to some customers in a Hospira MedNet Drug Library Creation Team Introductory Kit-LifeCare PCA kit. The accounts were requested to remove the utility from any PC in their facility on which it is currently installed, destroy the CD and sign and return the enclosed removal verification form verifying that they have removed the software from their system. The letter listed commercially available software packages that the account could validate and implement in their facility. Any questions concerning the utility software withdrawal were directed to Hospira at 1-800-241-4002, option 4, then option 1. |  | 127 | 127 CDs | Nationwide Distribution and Canada | Terminated | March 09, 2010 | 796 | 127
Z-1538-2008 | 46916 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Wedge HS Suture Anchor II with Needles, 6.5 mm, Model number 3910-100-145 |  |  | February 07, 2008 | 2008 | Stryker Endoscopy | Product package may crack and cause loss of sterility. | Packaging | Stryker Endoscopy initiated the recall on 02/07/2008 by notifying its consignees via telephone, as well as sending out Urgent: Device Removal notices. Recipients of affected lots are to inspect their inventory for part/lot numbers and quarantine any affected product. Affected products are to be returned to Stryker Endoscopy and Customer Service should be contacted for replacement alternatives. Also, the "Acknowledgment of recipt" card needs to be signed and returned to Stryker Endoscopy. |  | 53 | 53 units | Product was distributed to 8 consignees in CA, OH, PA, CO, MA, and NJ. | Terminated | December 21, 2010 | 1048 | 927
Z-1348-2008 | 46890 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Da Vinci S 8mm Instrument Cannula with outlet for use with the Da Vinci Surgical System; |  |  | January 04, 2008 | 2008 | Intuitive Surgical, Inc. | (1) Incorrect dimension on Luer on smoke evacuation cannulae not allowing for secure attachment function.(2) Incorrect labeling. External labeling lot number differs from lot number etched on the cannula. | Process control | On 1/4/08, the firm initiated the recall and its notification was via Urgent Device Recall letters explaining the reason for the recall, requesting immediate discontinued use, segregate non-conforming product in a secure area, and a Customer Service Rep will make contact in order to arrange for the retrieval of the material. |  | 39 | 39 units | Nationwide | Terminated | March 31, 2011 | 1182 | 39
Z-1620-2008 | 47521 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FLAT DRAIN -10MM, CHARIKER-JETER DRESSING KITS (LARGE AND SMALL) -Product numbers 6680217 and 6680221 - Wound Management Product. |  |  | March 31, 2008 | 2008 | Smith And Nephew, Inc. Wound Management Division | Incorrect labeling. Kit labels may incorrectly state that a flat drain is contained as a kit component while some kits may contain a round drain. | Labeling mix-ups | Consignees were notified of recall by letter (Correction and Removal Notification) sent to customers on 3/31/08. They were asked to check their stocks, identify, and return all affected units to Smith &amp; Nephew according to the instructions. Consignees are to contact the Customer Care Center at 1-800-876-1261, if they have additional questions. |  | 4796 | 4796 | Nationwide Distribution. | Terminated | November 18, 2008 | 232 | 4796
Z-1239-2008 | 46773 | No | JKS | Carbon monoxide test system. | Toxicology | Toxicology | 510(K) Exempt | International Technidyne Corp. (ITC) AVOXimeter 4000 Co-oximeter, Catalog Numbers: AVOX4000-110 and AVOX4000-220 |  |  | February 13, 2008 | 2008 | International Technidyne Corp. | Incorrect reading: ITC discovered that instruments built or repaired since approximately April 2007 were inadvertently left with the diagnostic mode enabled when shipped to customers. | Process design | ITC sent Urgent - Medical Device Field Recall - ITC No. 08-001 letters dated 2/08/2008 to hospitals and/or medical facilities identified as having affected instruments. The letter instructed users not to use the instrument if the diagnostic mode is enabled. ITC will reset affected devices and return them to the customers. |  | 63 | 61 US, 2 International | Worldwide Distribution: USA, United Arab Emirates, Malaysia, Saudi Arabia, India, Germany, United Kingdom, Greece, Romania, Singapore, ROC, Pakistan, Netherlands, Korea, France, Belgium, Switzerland, Canada, Thailand, Slovenia, Turkey and Australia. | Terminated | July 24, 2009 | 527 | 273
Z-1529-2008 | 46370 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with Data Innovations Instrument Manager, Part Numbers 626553 and 6605628 |  |  | August 20, 2007 | 2007 | Beckman Coulter Inc | Labeling provided by two integrated software systems are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. | Software design | A Product Corrective Action (PCA) letter, dated August 16, 2007, informed the customers that the User documentation and Online Help for the FC500 with Data Innovations lnstrument Manager are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. The letter asks the customers to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They are also to review all sample identification, demographics, data and results before reporting results out. A response form is included. Questions and concerns regarding the notification should be directed to Beckman Coulter Customer Service at 1-800-526-7694. |  | 141 | 141 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | October 15, 2012 | 1883 | 141
Z-1530-2008 | 46372 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSDiva Software v6.0 and 6.1 with the HP XW4400 Computer System |  |  | December 21, 2007 | 2007 | BD Biosciences | Values might not update in statistics views after certain functions are performed in a worksheet. | Software design | The firm issued an Important Product Information notification letter on 12/21/2007. The letter instructs users to reload data by closing and reopening an experiment before reporting results in order to determine if their statistics have been impacted. They are asked to fill out and return a form pertaining to this problem. BD Biosciences will release a version of BD FACSDiva software to correct this problem. The letter also stated that BD Biosciences rep will contact the users when the new software is available. User are asked to contact BD Customer Support Center at 1-877-232-8995 if they have any additional questions. |  | 74 | 74 units | Nationwide Distribution. | Terminated | September 24, 2008 | 278 | 74
Z-1561-2008 | 46750 | No | GTJ | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | BD DifcoTM Neisseria Meningitidis Antisera Group Y 241113, 1mL, in vials, which were labeled in part, 8019456 6250890 N. MEN. GP Y UNLABELED LYOPH, Becton, Dickinson and Company, Sparks, MD 21152. The product is packaged in 100 vial units |  |  | September 06, 2007 | 2007 | BD Diagnostic Systems / Lee Laboratories | Exhibits cross reactivity with Neisseria meningitides Group W135 | Mixed-up of materials/components | BD Diagnostics sent an Urgent Product Recall letter, dated September 2007, to all accounts, requesting the discontinue of distribution and that any remaining inventory of the recalled product be discarded. BD will issue credit or send replacement. A receipt of notification needs to be completed and sent back whether or not customers have any remaining inventory. Also, a customer list of end-users who were shipped this product needs to be provided so that the firm may inform them of the issue. |  | 200 | 200 units | Worldwide Distribution --- including USA and country of New Zealand. | Terminated | March 24, 2009 | 565 | 1240
Z-1368-2008 | 46906 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EvoTech Endoscope Cleaner &amp; Reprocessor, Product |  |  | February 13, 2008 | 2008 | Advanced Sterilization Products | Cleaning cycles are being cancelled. Specific issues are:a) Water heater tank is rusting resulting in the in-line 0.2 micron filter prematurely clogging,b) Interface board malfunction resulting in a premature lid opening, andc) Minor fluid leaks within the system (commonly at the 0.2 micron filter attachment points). | Other | Feburary 13, 2008 ASP issued an Urgent Product Recall Letter. ASP recommends that customers use an alternative reprocessing method until they return the system. However, if they do not have an alternate method for processing endoscopes and choose to continue using the EvoTech ECR System during this time, endoscopes processed through completed cycles in accordance with the User's Guide will be appropriately cleaned and highlevel disinfected. ASP is prepared to tailor an alternative solution to minimize the interruption to the customer's processing needs. |  | 64 | 64 units | Worldwide Distribution --- including USA and countries of Canada and Japan. | Terminated | November 29, 2011 | 1385 | 66
Z-1568-2008 | 46950 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 3M Comply EO Chemical Indicators Strips, Catalog #1251 |  |  | March 19, 2008 | 2008 | 3M Company / Medical Division | A production specification discrepancy may cause the indicators to show an inaccurate result, which could incorrectly lead customers to conclude the sterilization cycle was adequate. | Nonconforming Material/Component | Customers and Distributors were each sent a "Voluntary Product Recall" letter dated 3/17/08 on 3/19/08. The letter states that use of the affected product should immediately be discontinued, stocks should be checked for any remaining inventory, and reprocess any unused packs using the product not from the affected lot or from an alternative supplier. The letter requested consignees to notify 3M whether the facility has affected product or not by sending the "Product Return Form". Replacements will be provided to those with affected products remaining in their stock. |  | 51992 | 12,998 cases ( 4 boxes per case &amp; 480 strips per box)) | Nationwide Distribution --- including states of AL, AR, AZ, CA, CO, CT,DC, DE, FL, GA,HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS,MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | March 01, 2012 | 1443 | 51992
Z-1628-2008 | 47355 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilization System, Product Code 10101 |  |  | February 25, 2008 | 2008 | Advanced Sterilization Products | Oil mist filter fails allowing oil mist to be emitted into the vicinity of the sterilizer, which may result in the release of a mist, "haze," or "smoke." | Device Design | The firm sent an Urgent Device Correction letter dated 4/23/08, notifying customers of a situation that may occur with all models of STERRAD Sterilization Systems (100S, 50, 200, NX and 100NX). As a precaution, personnel should leave the room if mist is observed and avoid returning until the mist has cleared. ASP asks that customers circulate this letter to any STERRAD system users within their facility and also to the Chief Administrative Officer, the Director of Biomedical Engineering, Director of Materials Management and the Chief of Nursing. The letter instructs customers on what to do if they experience a filter failure - a) cancel the cycle if the sterilizer is running [and as always, do not use instruments from an incomplete cycle], b) leave the room, c) discontinue use of the STERRAD Sterilizer until the system is repaired. d) Personnel should avoid working in the room until the mist has cleared. The letter advises that ASP service will replace the sterilizer's filter and increase the frequency of filter changes by adjusting the Planned Maintenance (PM) schedule within the software on their STERRAD 100S and STERRAD NX Sterilizer. Per the notice, for each STERRAD Sterilizer model, ASP will do the following: a) STERRAD 200 Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter and gasket assembly at no charge. b) STERRAD NX Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter assembly at no charge and adjust the PM schedule to 650 cycles or 6 months (whichever occurs first). c) STERRAD 100S Sterilizer - beginning April 28, 2008 through October 2008, ASP will adjust the PM schedule to 750 cycles or 6 months (whichever occurs first). d) STERRAD 50 and 100NX Sterilizers - No adjustments to the oil mist filters or PM schedules are required at this time. PMs will be performed as currently scheduled. Information for MedWatch reporting is included in the letter. |  | 9950 | 9,950 | Worldwide Distribution -- including USA and countries of France, Germany, Italy, United Kingdom, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Australia, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Peru, Canada, Czech Republic, Middle East, Portugal, S. Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India, and Slovenia. | Terminated | November 21, 2011 | 1365 | 9950
Z-1463-2008 | 46850 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Viasys Healthcare, Pulmonetic Systems LTV 1000 Ventilator |  |  | February 22, 2008 | 2008 | Cardinal Health 203, Inc | The failure of an electronic component could result in failure of the ventilator to breathe for the patient and/or failure of the ventilator to properly alarm to alert the caregiver. | Device Design | An Urgent Medical Device Recall letter dated February 22, 2008 was sent to ventilator consignees/owners and distributors. The letter describes the issue, identifies the affected devices, states that the devices will be corrected by Cardinal Health and provides direction to verify ventilator operation and audible alarm function daily. Service Centers were provided information to identify affected printed circuit board assembly replacement kits' lot numbers and asked to return them. An "Update" was mailed 03/04/2008 to identify additional ventilator serial numbers that also need correction. |  | 4809 | 4809 (2427 US, 2382 OUS) | Worldwide Distribution --- USA including Puerto Rico and countries of Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Canada, Chile, China Peoples Republic, Colombia, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Norway, Panama, Paraguay, Philippines, Portugal, Saudi Arabia, Serbia &amp; Montenegro, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and West Indies. | Terminated | August 16, 2012 | 1637 | 20603
Z-1633-2008 | 47361 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | 6F Viking Guiding Catheter, Percutaneous catheter, Shape ALR 1.2, Part Number: 1001934-06, Abbott Vascular, Santa Clara, CA 95054. |  |  | March 17, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Mislabeling: The part number on the product label packaging may not match the product. | Labeling mix-ups | Consignees were notified via an Urgent Device Recall letter dated and hand delivered starting 3/18/08. The letter requests users to remove and quarantine the affected lots, which will be picked up and replaced by the recalling firm. Users were also requested to return the enclosed Efficiency Check Reconciliation Form. Abbott Vascular will exchange all returned devices. For additional information contact 1-800-227-9902. |  | 10 | 10 units | Nationwide Distribution. | Terminated | July 07, 2011 | 1207 | 17
Z-0866-2008 | 47371 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular Analytics System, Modular D2400 Module GMMI clinical chemistry analyzer; Catalog number 04998618001. |  |  | March 21, 2008 | 2008 | Roche Diagnostics Corp. | Possible mismatch between patient and result. | Process change control | Consignees were notified via Urgent Medical Device Correction letter dated 3/21/08, which instructed them to not manually order QC test selections while the system is in operation mode and that they can only manually order QC when the system is in Stand-by mode. Consignees are instructed to consult with their physicians or pathologist to determine specific clinical implications for their patients. The firm states that this issue will be resolved in the next version of software. |  | 11 | 11 of all models | Nationwide Distribution --- including states of Florida, Georgia, Indiana, New Jersey and North Carolina. | Terminated | November 18, 2009 | 607 | 11
Z-2141-2008 | 48697 | Yes | NIM |  |  | Cardiovascular | PMA | Boston Scientific NexStent Monorail, NexStent Carotid Stent and Monorail Delivery System. Material Number (UPN) M001553030, Catalog Number 55-303. [OUS Material Number (UPN) M001553000, Catalog Number 55-300] |  |  | June 06, 2008 | 2008 | Boston Scientific Corporation | Tip detachment: Boston Scientific has received complaints involving a detachment of the tip from the NexStent Carotid Stent delivery system. Potential health hazards resulting from this type of failure include increased procedure time, vessel wall injury, stroke and/or emergency surgery to remove the detached tip. Since the potential for the tip to detach occurs during delivery of the sten | Component design/selection | Boston Scientific sent an urgent Voluntary Medical Device Recall letter, dated 06/06/08, to affected customers. The letter described the issue, identified affected product, requested to cease using or distributing the product and provided direction on how to return the recalled devices to Boston Scientific, Quincy, MA. A Reply Verification Tracking Form was asked to be returned to Boston Scientific. Since the issue is with the delivery system, patients already implanted with the stent are not affected. |  | 1946 | 1946 units (1570 U.S., 376 OUS) | Worldwide Distribution: USA, Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, Italy, Kuwait, Luxembourg, Netherlands, Norway, Panama, Poland, Portugal, Romania, Spain, Switzerland, and Ukraine. | Terminated | February 16, 2011 | 985 | 1946
Z-2152-2008 | 48435 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | Vitrex Blood Gas Capillary Tubes, MC0024, REF 189301, 250 pcs, 1,45 x 1,85 x 125 mm app. 200Âżl, Heparinized: 6 IU Na &amp; 9 IU Li/100 Âżl, |  |  | May 08, 2008 | 2008 | Opti Medical Systems Inc | This lot of Heparinized Capillary Tubes contain contaminated lithium heparin. | Nonconforming Material/Component | Opti Medical sent a Recall Notice, dated May 8, 2008 to all consignees. Consignees were notified via email on/about May 5th thru 7th , 2008. They were instructed to discontinue use of the affected lot and to immediately return all remaining product to OPTI Medical for replacement. |  | 31750 | 31750 tubes | Worldwide Distribution: USA, Puerto Rico, Austria, China, Chile, Cyprus, Great Britain, Japan, Netherlands, and Senegal | Terminated | January 22, 2010 | 624 | 31750
Z-1635-2008 | 47373 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Integra LifeSciences Corporation Disposable Convenience Kit, PICC Insertion Tray, CUS791 |  |  | March 18, 2008 | 2008 | Integra Pain Management | Hep-Lock Flush vials included in Integra's PICC Insertion Trays contain a heparin component that has been recalled by Baxter for contaminated heparin. | Nonconforming Material/Component | Integra notified the consignee by letter on 03/18/2008. Instructions were to cease distribution/use of the affected kits and arrange for return to Integra. Also, consignee was instructed to notify all customers who received the kits to cease use and return any affected product for further disposition. |  | 500 | 500 kits | CO | Terminated | November 05, 2008 | 232 | 500
Z-1833-2008 | 48124 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HydraGlide Heparin coated PVC Straight Thoracic Catheters with Tapered Connector Tip, |  |  | May 12, 2008 | 2008 | Atrium Medical Corporation | Heparin-coated Thoracic Catheters were manufactured with heparin allegedly contaminated with oversulfated chondroitin sulfate | Nonconforming Material/Component | Atrium Medical issued a press release on 5/12/2008 and notified consignees by facsimile on 5/12/2008. Customers were instructed to contact the firm for a return goods authorization and to notify the firm upon receipt of the recall notice. |  | 41 | 41 cases (10 individual/case) | Worldwide Distribution: USA, US Virgin Islands, South Africa, Austria, Taiwan, Venezuela, and Saudi Arabia. | Terminated | August 20, 2009 | 465 | 1842
Z-1442-2008 | 46920 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems On-Board Imager, Versions 1.0.15 and 1.2.05 Imaging Accessory to Medical Linear Accelerator. The On-Board Imager Device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields in relation to anatomical and/or fiducial landmarks. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | If used with a third party radiation therapy treatment planning software system, mistreatment may occur because of a misalignment. | Software design | The firm has upgraded all user sites through normal course of service and installation activities. Notification letters were distributed starting January 03, 2008. |  | 291 | 291 units | Worldwide Distribution -- USA including states of GA, MI, NY, MN, TX, TN, CA, MD, KS, AL, KY, CT, WI, LA, NC, NJ, PA, IL, RI, FL, NE, AK, VA, WV, NV, MO, TN, OR, IN, MA, CT, AL, AZ, NH, ID, OH, SC, CO, CT, ND, WA, and UT, and countries of Sweden, Switzerland, Germany, Denmark, Japan, UK, France, Italy, Norway, Taiwan, New Zealand, Canada, Poland, Australia, Colombia, China, Spain, Belgium, Russian Federation, India, Iceland. | Terminated | November 27, 2010 | 1059 | 291
Z-1445-2008 | 47244 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | QWIX Screw 4.3mm diam; length 26 mm; The stabilization screw is indicated for fixation of bone fractures or for bone reconstruction. |  |  | February 27, 2008 | 2008 | Integra LifeSciences Corp. | Certain QWIX Fixation Screws (Reference numbers 111426S and 111428S) have been etched and labelled with an incorrect length. | Labeling mix-ups | Integra Reps were sent recall notification with a recall acknowledgement on 2/27/2008. The customers who were shipped affected product were also notified on that date. |  | 42 | 42 units | Nationwide Distribution --- including states of NC, IL, CA, NC and TX. | Terminated | August 21, 2008 | 176 | 62
Z-1619-2008 | 47495 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI Âż3180 Oximeter - Patient monitoring medical device, Manufactured by Smiths Medical PM, Inc. For continuous patient monitoring of blood oxygen levels, pulse rate and pulse strength measurements. It is only used in hospital or clinical settings or emergency response environments. |  |  | February 27, 2008 | 2008 | Smiths Medical PM, Inc. | The Smith Medical PM, Inc. BCI ÂżÂż3180 Pulse Oximeter tantalum capacitors C38, C72, C74, and C98 were installed backwards during assembly at the board manufacturer. This exposes the components to reverse polarity voltages that may lead to their pre-mature failure. As a result, the identified monitors may exhibit failure modes such as the monitor may shut down and not produce alarms or signals to | Vendor change control | Consignees were contacted by an URGENT: Product Safety Information Voluntary Recall Notification letter dated March 11, 2008 with attached Safety Action Bulletin 08-SAB01 dated February 27, 2008 and email between March 7 and March 13, 2008. The consignees will need to return the device to be evaluated in the Smiths Medical PM, Inc. |  | 900 | 900 | Class II Recall - Worldwide Distribution --- USA including states of GA, AL, IL, TX, MI, VA, IA, TN, PA, NC, HI, NY, FL, AZ, OH, CA, MS, UT and Countries: Argentina, Bangladesh, Iran, Thailand, Mexico, India, Germany, Guatemala, Pakistan, Chile, Colombia, Brazil, Saudi Arabia, Morocco, Kuwait, Trinidad and Tobago, Ecuador, Paraguay, Uruguay, Canada, Poland, Portugal, Peru, Qatar, Cyprus, Greece, Italy, Spain, United Arab Emirates, Israel. | Terminated | December 17, 2011 | 1389 | 900
Z-1622-2008 | 47753 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂż Insulin Pump, Model 1700, Smiths Medical International Ltd WD24 4LG UK, Smiths Medical MD, Inc., St Paul, MN 55112. This is a syringe infusion pump intended for Continuous Subcutaneous Insulin Infusion for the control of diabetes. |  |  | March 10, 2008 | 2008 | Smiths Medical MD, Inc. | Smith Medical discovered an issue with a motor component in specific serial numbers of the Deltec Cozmo ÂżInsulin Pump, Model Numbers 1700 and 1800. This issue affects how the motor operates and may cause an over-delivery of insulin which could result in injury to the user. | Component change control | Consignees were mailed an Urgent Medical Device Recall letter dated March 10, 2008. The Urgent letter addressed to "Distributors, Clinicans, Patients and Other Customers" contains Description of the Problem, Results of Investigation, and Advise on Action to be taken by the User. The patient users will be required to return a Recall Confrirmation Form. The USA distributors, USA prescribing physicians and the USA managed Care Accounts (Insurance Companies) will be mailed also a cover letter. |  | 436 | 436 | Class II Recall - Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME. RO, MN, MO, MS, ME, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and countries of Greece, Netherland, Sweden, France, Germany, Turkey, China, Great Britain, Israel, Italy, Kuwait, New Zealand, Australia. | Terminated | January 13, 2011 | 1039 | 1502
Z-1694-2008 | 47766 | No | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | The OptiVantage DH Injection System is a contrast delivery system and is designed to inject radiopaque contrast media into a patient's vascular system to obtain diagnostic images when used with computed tomography equipment. The J-Bow component is used to connect the powerhead to a suspension system. |  |  | April 07, 2008 | 2008 | Mallinckrodt Inc | Screws can begin to loosen and shear off to the point of the J-Bow falling from the suspension system. This may cause injury to patients or hospital staff. | Component design/selection | On April 18, 2008, Mallincrodt Inc. issued an Urgent Device Correction letter informing customers to check their inventory, and determine if they have any defective system units. If they observe any loose screws, sheared screws and/or elongated opening(s) around the screws, they are to immediately take the system out of service and contact the firm's service representative. Covidien's customers were all asked to inspect their inventories and to complete an 'Effectiveness Check Return Form' whether or not the product is defective. However, only those system units observed to have loose screws, sheared screws, and/or elongated openings around the screws, will be scheduled to receive correction by a Covidien Service Representative. |  | 783 | 783 Units | Product is distributed in the US throughout the following states: AL, AR, AZ, CA, CO ,CT, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WV. | Terminated | May 20, 2011 | 1138 | 783
Z-1621-2008 | 47767 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-MS ÂżÂż3 Ambulatory Infusion Pump, Model 7400 Smiths Medical MD, Inc., St. Paul, MN 55112 USA. Electromechanical pumps used for general drug delivery infusion therapies. Used mainly by home care patients but can also used in health care facilities. |  |  | March 07, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical discovered an issue with a motor component in specific serial numbers of the CADD-MS Âż3 Ambulatory Infusion Pump. This issue affects motor operation and may cause an over-delivery of medication which could result in injury to the user. | Component change control | Consignees were mailed an Urgent Medical Device Recall letter dated March 7, 2008. (Urgent Field Safety Notice dated March 7, 2008 for International Consignees.) The Urgent letter was addressed to "Risk &amp; Safety Mangers, "Distributors", "Clinicians", "Patients" and "other Customers". The letter contains a Description of the Problem, Results of Investigation, and Advice on Action to be taken by the User which includes contacting Smiths Medical MD, Inc. Customer Service Department to make arrangements for a replacement pump. The letter also requests consignees to return the confirmation form. |  | 646 | 646 | Class II Recall - Worldwide Distribution --- USA including states of PA, TX, KY, FL, MN, GA, HI, NM, MD, and countries of CZ, GB, IT. | Terminated | December 17, 2011 | 1380 | 646
Z-1526-2008 | 44927 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Neonatal Enhancement for the Engstrom Carestation, Model # 1505-3272-00, GE Healthcare P.O. Box 7550, Madison, WI 53707-7550, USA |  |  | July 02, 2007 | 2007 | Datex - Ohmeda, Inc | Neonatal Flow Sensors supplied for use with the Neonatal Enhancement for the EngstrÂżm Carestation may result in the sensors providing invalid values and alarm messages when exposed to high flow rates. | Nonconforming Material/Component | Consignees were sent a Medical Device Correction Notification letter on July 2, 2007. The firm's short term recommendation is that Neonatal Flow Sensors may continue to be used with added viligance around potential patient circuit disconnections and the associated alarms. Ther firm recommends that the MV exp high and Circuit leak alarms be set appropriately for each patient in order to detect a patient disconnect. The long term recommendation was to order the replacement part when available and dispose of the old part when the replacement is received. |  | 394 | 394 | Worldwide Distribution --- USA including state of TX and countries of Austria, Bangladesh, Belgium, Costa Rica, Denmark, Finland, France, Germany, India, Italy, Mexico, Saudi Arabia, South Africa, Sweden, Switzerland, United Kingdom, United Arab Emirates. | Terminated | July 03, 2012 | 1828 | 394
Z-1528-2008 | 46323 | No | GNB | Shigella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Lee Laboratories brand Shigella Antiserum Poly Group B, catalog number 210722, 10 mL vials and labeled in part ***Lee Laboratories, Inc Grayson, Georgia 30017*** |  |  | November 30, 2007 | 2007 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent for identification of Shigella bacteria in patient samples may cause false negative results. | Nonconforming Material/Component | The recalling firm notified distributors by faxed letter flagged as "Urgent Product Recall" and e-mail on 11/30/07. Distributors were instructed to cease distribution and discard identified product and return a list of end user consignees for direct notification by Becton Dickinson. End users were notified by letter flagged as an "Urgent Product Recall" on 11/30/06 through 12/06/07. Customers were instructed to discontinue and discard affected product for replacement. The notification included a reply form that BD requested users to fax back with the quantity of replacement vials needed. |  | 643 | 643 (recall total) vials | Worldwide Distribution | Terminated | November 12, 2008 | 348 | 643
Z-1515-2008 | 46941 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER, MOSAIQ versions 1.3 and 1.4 with Elekta and Siemens Machines, Model number 40003-001000-RO, Manufactured by IMPAC Medical Systems, Sunnyvale, CA. |  |  | February 14, 2008 | 2008 | Impac Medical Systems Inc | Software issue may result in change to intended treatment field, potentially resulting in mistreatment. | Software design | Firm issued field advisory to its consignees on 02/14/2008. Phone calls will also be made to affected customers to ensure reception and understanding of the field notification. Firm advises users with affected product to review the jaw positions of IMRT fields to ensure that they match the RT Plan after an Apply Limits or Machine Change occurs, until software at their site can be updated to the corrected version. A software patch is under construction. |  | 36 | 36 units | Worldwide Distribution --- USA including states of OH, OK, NJ, MN, OR, TN, NY, CA, AZ, NV, IA, GA, AL, MD, MI, PA, IN, WA, FL, TX, CO, and countries to Canada, Australia, Germany, Greece and Sweden. | Terminated | June 24, 2009 | 496 | 36
Z-1444-2008 | 46942 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Medical Systems iSite PACS (Picture Archiving and Communications Systems) version 4.1.17.0 - Radiology Image management system. |  |  | January 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | When using the Freehand Region of Interest (ROI) tool, errors may occur in area calculation yielding incorrect results. The firm discovered that a descrepancy exists in the measurement provided by the tool. | Software design | The firm issued product notification letter February 20, 2008, to its consignees. The notification alerts customers to the behavior, requires them to post the notification and notify users not to use the Freehand ROI feature until they are upgraded to a corrected software version. Philips will deliver a software upgrade to correct the problem. |  | 2 | 2 units | Worldwide Distribution --- USA including states of MN, CO, and PA, and country of Germany. | Terminated | September 25, 2008 | 240 | 5
Z-1525-2008 | 38337 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | FACSDiva Software Versions 4.0.1 to 5.0.2, Flow Cytometry Acquisition &amp; Analysis software. Catalog number 641609 |  |  | April 05, 2007 | 2007 | BD Biosciences | When data file containing one or no fluoresence parameters is exported, the software will automatically apply compensation to this file and all subsequently exported files. | Software design | The firm has issued a customer Important Product Information letter via mail on 04/20/2007 to inform them of the software issues. The letter states that to avoid this problem, make sure each tube or well has at least two fluorescence parameters before exporting FCS 3.0 data files. Manually exporting tubes or wells one at a time or restarting the software after the export of an FCS data file containing one or no fluorescence parameters will prevent the problem from occurring. Customers are requested to fill out and return the response form and to contact BD Customer Support Center if there are any additional questions. A software fix is being developed in version 6.0 and a free maintenance release with no product enhancments (version 5.0.3). |  | 1074 | 1074 units | Worldwide Distribution --- USA including states of NY, NM, MI, GA, MD, MA, AL, CA, VA, NC, OH, TX, CO, NJ, PA, MI, CT, IA, MO, NV, KY, IL, TN, FL, MN, FL, UT, WA, OK, CT, WI, SD, AZ, and IN, and countries of Buenos Aires, BE, Brazil, Canada, Chile, China, Hong Kong, Japan, Singapore, Taiwan, Venezuela, Mexico, Trinidad and Tobago, and Columbia. | Terminated | September 30, 2009 | 909 | 1074
Z-1750-2008 | 47764 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Omnifit EON Surgical Protocol; |  |  | May 17, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The Table 2 Broach and Neck Trial Sizing (Literature # LSP48 04/04) on page 3 is incorrect, reading 5mm less than what it should be for certain sizes. | Labeling design | Styrker Branches/Agencies were notified of the incorrect information printed in the protocal on May 17, 2007 via an Important Product Correction letter. The letter requested that all incorrect copies of the protocol be located and destroyed immediately. They requested that the Acknowledgement Form be completed and faxed, and that a Stryker Orthopaedics Customer Service rep be contacted in order to reorder the destroyed product. |  | 2501 | 2501 units | Worldwide Distribution --- including USA and countries of Canada, Australia, Latin America and Greece. | Terminated | August 19, 2008 | 460 | 2501
Z-1764-2008 | 47459 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | HerpeSelect 1 and 2 Immunoblot IgG Kit, IB0900G |  |  | July 06, 2007 | 2007 | Focus Diagnostics Inc | A small percentage of samples tested in specific lots, the HSV common antigen band appears lighter (less reactive) than the reading control band. Reduced reactivity of the common antigen band may lead to inaccurate interpretation of a sample as negative. | Process design | Focus notified all clients which received the recalled lots starting on July 6, 2007 via an Important Medical Device Information letter by fax or email. The letter explains how to identify suspect results and the steps to be taken if suspect results are obtained. If customers find the interpretation of results with the indicated kit lots to be problematic, customers have the option of having their kits replaced at no charge within 90 days of receiving the notification. |  | 190065 | Lot Code 063169 - 97 Kits, Lot Code 063354 - 16 Kits, Lot Code 063355 - 74 Kits | Worldwide Distribution --- USA including states of OH, CA, CT, IL, NY, and countries of Barbados, Canada, Germany, Israel, Slovakia, France, and England. | Terminated | June 25, 2009 | 720 | 190065
Z-1370-2008 | 46771 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Exacta-Mix 2400, Pharmacy Compounding System is used in the pharmacy to compound multiple source ingredients into on final solution based on user-specified sequences and volumes. |  |  | February 11, 2008 | 2008 | Baxa Corporation | Flush solution from TPN compounder may be added to patient's TPN bag. | Software design | Consignees were notified by letter (Exacta-Mix Compounder Software Safety Alert) on 02/11/2008 advising them of the necessity to install a waste bag prior to a machine flush. |  | 1000 | 1000 units | Nationwide, including military facilities in CA, FL, HI, MS, NC, TX, VA and VA facilities in NY, OR, SC, UT. Foreign distribution to Australia, Canada, China and New Zealand. | Terminated | March 16, 2012 | 1495 | 1221
Z-1322-2008 | 46870 | [6 empty] | Implantable Cardioverter Defibrillator |  |  | January 16, 2008 | 2008 | St Jude Medical CRMD | A condition that could lead to a ventricular sensing anomaly in Epic and Atlas family of lmplantable Cardioverter Defibrillators(ICDs) has been identified. A loss of ventricular sensing would prevent the ICD from being able to detect an arrhythmia. If an arrhythmia occurred and the device was unable to detect it, then a patient would not receive potentially life saving therapy. | Software design | A January 16, 2008, letter entitled, "Important Device Upgrade Information" was sent to physicians by FedEx. In the US, sales representatives will follow-up with the physician and will upgrade all programmers. For the rest of the world, it will be monitored by various St. Jude offices and local distributors with this information being fed back to St. Jude Medical CRMD in Sylmar. In addition, St. Jude Medical issued a press release on January 18, 2008 and information concerning this advisory can be found on St. Jude's website, www.sirn.com. |  |  |  | Nationwide and Internationally: Argentina, Australia, Austria, Barbados, Belgium,Bolivia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Denmark, Dominican Republic,Ecuador, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Korea, Malaysia,Mexico, Netherlands, New Zealand, Norway, Peru. Poland, Portugal, Republic of Panama,Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, United Kingdom, Uruguayand Venezuela. | Terminated | April 03, 2012 | 1539
Z-1772-2008 | 47781 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Bariatric Wheelchair, 28" Wide and 20" Deep Seat, 700 lb. Weight Capacity; Manufactured in Taiwan for Medline Industries, Inc., Mundelein, IL 60060 USA; |  |  | March 31, 2008 | 2008 | Medline Industries Inc | Injuries to fingers. There is a potential for users to pinch their fingers between the seat bars when opening the wheelchairs. The injuries can be severe from fracture to severing. | Component design/selection | A Voluntary Field Correction letter dated 3/31/08 was issued to customers. The letter informed them of reports of users catching their fingers under the seat bars, which are identified on the chair with warning stickers. The customers were requested to change the backrest with the enclosed replacement upholstery kit, properly dispose of the old back rest, and place the enclosed cautionary tags on both sides of the chair seat. Instructions for installation of the upholstery and cautionary tags were included, as was a sheet entitled "How to Open a Wheelchair" which gives step by step instructions with illustrations for opening the wheelchair, as well as the warning "CAUTION KEEP YOUR FINGERS ON TOP OF THE UPHOLSTERY AT ALL TIMES. DO NOT PLACE YOUR FINGERS ON THE SIDES OR BOTTOM OF THE CHAIR SEAT." The accounts were requested to complete the enclosed response sheet and fax it back to Medline at 847-643-4482, indicating the serial numbers of chairs corrected. Any questions were directed to 800-633-5463, extension 5349, between 8:00 AM and 5:00 PM CST. Distributors were requested to notify their accounts of the recall. |  | 201 | 201 wheelchairs | Class II Recall - Worldwide Distribution --- including USA and Canada. | Terminated | May 11, 2010 | 771 | 1138
Z-1378-2008 | 46938 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Radius Capspin is a non-sterile instrument for use with other devices in spinal surgery. |  |  | February 19, 2008 | 2008 | Stryker Spine | The Radius Capsin has separation of the tip from the inserter. | Device Design | Recall notification was sent via FedEx on February 19, 2008 to all branches/agencies. All questions are to be directed to Michelle Klein, Regulatory Compliance Specialist - 201-760-8256. |  | 52 | 52 | Product was distributed to a limited population during a product launch for the purposes of gathering feedback on the product. | Terminated | August 19, 2008 | 182 | 117
Z-1183-2008 | 46316 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Power Supply Assembly Sled |  |  | October 17, 2007 | 2007 | Beckman Coulter Inc | Electrical grounding failure: During manufacturing, internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering. | Process control | Beckman Coulter mailed a Product Corrective Action (PCA) letter on October 17, 2007 to all Access and Access 2 lmmunoassay Systems, SYNCHRON LXi 725 System, and UniCel DxC 600i System customers.The customers were informed that during manufacturing, internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering. The letter referenced aproactive initiative to inspect the power supply sled assemblieson all instruments referenced. A response form was included. |  | 3247 | 3055 instruments in the US; 192 instruments in Canada | Nationwide Distribution and Canada | Terminated | April 27, 2012 | 1654 | 3247
Z-1184-2008 | 46336 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel Dxl 800 Access Immunoassay System; |  |  | October 08, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: There is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. | Device Design | Beckman Coulter mailed a Product Corrective Action (PCA) letter the week ofOctober 8, 2007 to the customers who have the affected UniCel Dxl 800Access lmmunoassay systems advising that there is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. The letter further advises that Beckman Coulter will contact the customer to arrange a date to remove the modificatlons but until the modification is completed, the system should be operated in aroom temperature of 20ÂżC to 30ÂżC. This can be acomplished by monitoringand recording room tamperature in the area of the instrument. Customers are also asked to review historical quality control records to ensure assay performance is within their laboratory specification and that patient results should be reviewed in context of other diagnostics tests and the clinical presentations of the patient. A response form is included. |  | 120 | 111 units in the US; 9 units in Canada | Nationwide and Canada | Terminated | January 11, 2012 | 1556 | 120
Z-1189-2008 | 46728 | [6 empty] | Bio-Rad MONOLISA Anti-HBc EIA, Catalog Number 26186. In vitro diagnostic. |  |  | January 07, 2008 | 2008 | Bio-Rad Laboratories | Incorrect result: Elevated reactivity with a kit lot 232BGG. Patient samples that are negative for antibodies to anti-HBc could be assigned a positive result. | Process control | Bio-Rad called each customer and alerted them to the issue on 1/7-8/2008. A letter was faxed and sent via overnight carrier (DHL) to each customer site on 1/8/2008. Customers were told not to use the affected lot of the product and to discard it. The firm will provide another lot of the product to customers. |  | 31 | 31 distributed | Nationwide Distribution. | Terminated | March 07, 2011 | 1155 | 31
Z-1478-2008 | 46781 | No | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allen Medical C-Flex, Polar Head Positioner |  |  | November 27, 2007 | 2007 | Allen Medical Systems, Inc. | Handle jamming in the open position, preventing the device from adequately supporting the head during spinal surgery. Possible delay in surgery, or injury. | Device Design | Allen Medical notified accounts by telephone on 11/08/07 and the Sales Representative. The Hospitals confirmed that the device would be sent back for evaluation and upgraded. |  | 2 | 2 | MA, MS | Terminated | August 26, 2008 | 273 | 2
Z-1447-2008 | 46388 | No | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | GivenÂż Diagnostic System with PillCamÂż SB Capsule - Gastrointestinal capsule imaging system. The device is an ingestible telemetric gastrointestinal capsule imaging system is used for visualization of the small bowel mucosa as an adjunctive tool in the detection of abnormalities of the small bowel. The device captures images of the small bowel with a wireless camera contained in a capsule. This device includes an ingestible capsule (containing a light source, camera, transmitter, and battery), an antenna array, a receiving/recording unit, a data storage device, computer software to process the images, and accessories. |  |  | December 13, 2007 | 2007 | Given Imaging Inc | The cards included in the product's education kit (proposing a clinical algorithm for capsule endoscopy in patients with suspected Crohn's disease), were printed incorrectly. | Employee error | Consignees were notified by letter on 12/13/2007. The letter contains a diagram illustrating the error on the incorrect card and a diagram illustrating the corrected card. Consignees were instructed to discard the incorrect card if they still have it and to complete a response form and return it via fax to Given Imaging. |  | 1392 | 1392 units | Worldwide Distribution --- US including Puerto Rico and country of Canada | Terminated | November 15, 2010 | 1068 | 1392
Z-1695-2008 | 47824 | No | LZS |  |  | Ophthalmic | PMA | Excimer Laser System; WaveScan WaveFront System, Software Version 3.65, for use with the VISX STAR S4 and STAR S4 IR Excimer Laser Systems. System is a diagnostic instrument intended for the automated measurement, analysis and recording of refractive errors and wavefront aberrations to the target shape and then generates the commands for the VISX STAR laser to create the target shape on the cones. |  |  | November 17, 2006 | 2006 | VISX INCORPORATED, A SUBSIDIARY OF AMO INC | Firm became aware of a significant overcorrection in one eye of a patient who had received a wavefront-guided LASIK treatment at a facility in Germany. The probability of a recurrence of this error was determined to be extremely remote. To ensure that this type of incident will not happen again, a software update was developed. | Software design | VISX issued a notification letter on November 27, 2006, to all distributors and direct accounts. The letter informed the recipients that they were performing a field correction in the form of a mandatory upgrade on all VISX WaveScan WaveFront Systems and Offline Programming Modules to correct an issue. The field upgrade should be made by February 16, 2007. A Dear Doctor letter was also issued instructing surgeons to upgrade all Offline Programming Modules using an enclosed CD. |  | 1310 | 1310 units | Class 2 Recall - Worldwide Distribution: USA, Venezuela, Colombia, Italy, Saudi Arabia, Kuwait, Egypt, UAE, Jordan, Lebanon, Australia, Canada, France, Germany, India, Ireland, Japan, Sweden, Vietnam, Czech Republic, Slovak Republic, Bulgaria, Philippines, South Korea, Turkey, UK, Portugal, Spain, Argentina, Israel, Cyprus, India, Mexico, Greece, Singapore, Chile, Poland, Brazil, China, Thailand, Taiwan, Hong Kong and Russia. | Terminated | June 08, 2010 | 1299 | 1310
Z-1769-2008 | 47773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Medical charged-particle radiation therapy system. Siemens Linear Medical Accelerators equipped with COHERENCE Therapist R2.0 and Oncologist R2.0. Product update, TH 030//07/S. The COHERENCE Therapist R2.0 software numbers are: 7358299, 7360857, and 8141835 and applies only to those parts used with the following devices: Coherence AG Therapist; Coherence Therapist System; Coherence Impression Therapist; PRIMEVIEW 3i; AG Therapist 3rd party V&R; Coherence Therapist 2.0 and PRIMEVIEW 3i 2.0. |  |  | October 03, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Software. The firm issued an Advisory Letter Update instruction TH030/07/S as a result of a complaint that images viewed in the device may shift, which could result in mistreatment. The advisory alerts all its affected customers of the image issue to prevent potential mistreatment. | Software design | Customer Safety Advisory Notice was issued on 10/03/2007 to consignees, via certified mail. Consignees are strongly cautioned that they should not further correct the image alignment at this time as doing so will result in incorrect offset values which could lead to incorrect repositioning of the patient. Consignees are instructed to check the "Filter Images on Loading" and "Saving of Filtered Images" config settings, and apply any additional filters to the image BEFORE using the interactive shift or landmark registration tools. Do not correct any perceived issues. A software fix is being validated and will be available soon. |  |  |  | Class II Recall - Worldwide Distribution --- including USA and countries of Angola, Australia, Brazil, Canada, Croatia, Denmark, Egypt, France, Germany, Greece, Hungary, India, Iran, Italy, Japan, Malaysia, Netherlands, New Zealand, P.R. China, Poland, Portugal, Republic Korea, Russian, Fed. Saudi Arabia, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad, Tobago, Turkey, United States, Vietnam. | Terminated | April 29, 2011 | 1304
Z-1773-2008 | 47785 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Infusion System. EKOS EndoWave Infusion System. EKOS EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core. Labeled as Sterile, Single Use Only. Connector clip is CIC Connector Clip. The infusion system is intended for the controlled and selective infusion of physician-specified fluids, including thrombolytes, into the peripheral vasculature. |  |  | April 01, 2008 | 2008 | EKOS Corporation | Catheter disconnection. There is a potential for the catheter to disconnect from the Catheter Interface Cable (CIC) pod which connects to the Control Unit. | Other | Firm representatives will visit and distribute connector clips and instructions for use to consignees beginning 04/01/2008 at no cost to the consignee. |  | 218 | 218 units | Class II Recall - Nationwide | Terminated | May 13, 2011 | 1137 | 218
Z-1186-2008 | 46345 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Impactor, 3/8" - 24 Thread Insert; Model 31-400603. Instrument for hip prosthesis acetabular cup component. |  |  | January 16, 2008 | 2008 | Biomet, Inc. | The weld at the lock location may fracture during impaction. | Device Design | Biomet notified distributors and medical facilities by letter dated 1/16/08. Distributors with distributor-owned instruments at medical facilities were notified by email on 1/17/08, and provided with a letter to present to their associated medical facilities describing the recall. |  | 59 | 59 | Worldwide Distribution: USA, Germany, Mexico, Netherlands, Sweden and United Kingdom. | Terminated | September 22, 2009 | 615 | 1004
Z-1539-2008 | 46960 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKNEE Tibial Insert Trial, Size 2, 13mm Right; Part #801-01-640; manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | February 20, 2008 | 2008 | Encore Medical, Lp | Tibial insert trial incorrectly marked for thickness; device thickness marked correctly as 13mm on the underside of the trial, but is incorrectly marked as 15mm on the side. | Error in labeling | Firm notified consignee by phone and asked them to return any outstanding product on 02/20/08. |  | 6 | 6 units. | Distribution --- country of Germany. | Terminated | September 02, 2008 | 195 | 6
Z-1514-2008 | 46297 | Yes | LWP |  |  | Cardiovascular | PMA | VITALITY EL, models T127, T167, T177 Implantable Cardioverter Defibrillator (ICD), pulse-generator, dual Chamber |  |  | January 07, 2008 | 2008 | Boston Scientific CRM Corp | Repetitive mechanical stress applied to the titanium case can induce component damage and device malfunction in certain ICD and CRT-D devices implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. This update identifies additional VITALITY ICD | Device Design | An Urgent Medical Device Information Update to May 2006 Product Advisory, dated January 04, 2008, was sent to implanting physicians and physicians following implanted devices beginning Monday, January 07, 2008. The letter described a malfunction that can occur when the ICD is implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. The physicians were asked to review records to determine if an affected device was implanted subpectorally. An AP radiograph should be done to determine specific device orientation for subpectoral implants. If a subpectoral implant is in a susceptible orientation, consider repositioning or replacement for physically active patients who regularly need device therapy. Follow patients with susceptible devices at 3-month intervals. |  | 25000 | BSC CRM estimates that the number of devices implanted in a susceptible orientation is likely less than 1% of the total active implants 25,000 (w.w.) | Nationwide and Worldwide: Bolivia, Panama, Iran, Saudi Arabia, Kuwait, United Arab Rep., Austria, Belgium, Cyrus, Czech Republic, Denmark, Finland, France, Germany, Greece, Iceland, Ireland, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, UK, India, Australia, New Zealand, Malaysia, Singapore, Hong Kong, China, Israel, Turkey, Thailand, and Canada. | Terminated | December 30, 2008 | 358 | 25000
Z-2139-2008 | 47801 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA |  |  | April 11, 2008 | 2008 | Abbott Laboratories | Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical. | Software change control | Recall initiated on April 11, 2008. A Product Correction letter and customer reply form were sent to all currently active CELL-DYN Sapphire Customers. The letter requests that users be certain that the transmission of test results receipt by the host interface is mapped according to the Numerical Result Label, when configuring their CELL-DYN Sapphire Hematology Analyzer for results transmission to a Clinical Laboratory Information System or Middleware. Affected product was distributed to a total of 4 distributors. However, notification was made to all 154 customers that received the CELL-DYNE Sapphire product. If the customer or the health care providers that they serve have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT. |  | 15 | 15 copies of the specification were sent worldwide | Nationwide Distribution. | Terminated | March 12, 2010 | 700 | 15
Z-1471-2008 | 46903 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Xmatrx FISH (fluorescent in situ hybridization) Automated Staining System, Catalog number AS4010B |  |  | January 24, 2008 | 2008 | Biogenex Laboratories | Bulk reagent dispensing mechanism may break resulting in reagent spill and possible exposure to chemicals. | Component design/selection | The firm initiated written notification on 01/24/2008. The firm has asked its consignee to inform its consignees of the recall. |  | 10 | 10 units distributed. | Product was distributed to a single distributor, which further distributed the product. | Terminated | January 08, 2011 | 1080 | 10
Z-2150-2008 | 48184 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Equal Diagnostics Lipase Color Reagent Kit |  |  | April 14, 2008 | 2008 | Genzyme Corporation | Calibration and Quality Control outside of established range resulting from an increase in initial OD readings with the reagent contained within the kits. | Nonconforming Material/Component | Genzyme notified accounts by Fed'X on 4/14/08 via an Imprtant Field Correction letter, requesting users to discard kits and follow the "User Update" instructions previously provided to minimize calibration error. Customers are to complere and sign the Confirmation of Notification form and fax it back to the firm regardless if they have any affected product or not. Also, requires that the customer confirms that the REagent OD limits for First and Last H are set at 0.15 for Lipase on the Olympus AU400/400e, AU600, AU640/640e and AU2700/AU5400 Systems. |  | 3443 | 3,443 | Nationwide Distribution | Terminated | September 16, 2008 | 155 | 3443
Z-2140-2008 | 48465 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Osteo-Clage Cable System Package Assembly (01-0001-S). Product labeled in part, "2.0mm x 850 mm CoCr Cable / Sleeve... REF 01-0001-S LOT W59192... Acumed 5885 NW Cornelius Pass Road Hilsboro, OR 97124-9432 ". |  |  | March 28, 2008 | 2008 | Acumed LLC | Cables of incorrect diameter were used in product assembly, resulting in the sleeve not being able to bind to the cable after crimping. | Employee error | On 3/28/08, the firm sent an Urgent Notification letter to their consignees. The letter explained the issue and requested the customers return the product to Acumend. On 7/16/08, the firm sent a 2nd notification, via mail, to thier customers. The Urgent Notice: Device Recall 2nd Notification, advised customers of the recall, the hazard and requested customers to return existing product to Acumed. |  | 50 | 50 units | Nationwide Distribution --- including states of AR, CA, PA, GA, KS, MO, and TN. | Terminated | February 03, 2010 | 677 | 50
Z-1541-2008 | 46961 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio |  |  | March 04, 2008 | 2008 | Philips Medical Systems | Medications prescribed for one patient were printed on the record of another patient. | Software design | Philips Medical issued letter notification, dated 29 Feb 2008 to customers titled; "Urgent Medical Device Correction Notice".Customers were asked to not solely rely on the Worklist to remind them when to administer orders or interventions until the software has been upgraded. Customers should review the reminders and verify orders and interventions on the Medication Administration Record and Administration record. The correction will be in the form of a software upgrade and will be performed remotely by Philips service personnel. |  | 138 | 138 for products | NationwideCanada, Australia, Austria, Belgium, Denmark., France, Finland, Germany, Great Britain, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, and UAE | Open, Classified |  |  | 138
Z-1532-2008 | 46908 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Groove Power Wheelchair with QC (non-expandable) electronics, Model number EIPW13. |  |  | January 11, 2008 | 2008 | Sunrise Medical Inc | Product experiences a longer than expected stopping distance, which may cause injury to the user, objects, or the chair itself. | Software design | Initial phone contact began on January 11, 2008. Medical device correction letters were also sent to affected dealers. |  | 141 | 141 units total | Product was distributed to 126 retailers nationwide throughout the USA, as well as Canada and Australia. | Terminated | November 16, 2009 | 675 | 141
Z-1542-2008 | 47252 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Altaire MRI System Emergency Rundown Switch Unit (ERDU) is intended to provide the physician with physiological and clinical information obtained non-invasive and without the use of ionizing radiation. |  |  | February 11, 2008 | 2008 | Hitachi Medical Systems America Inc | The firm's Altaire MR Scanner's Emergency Rundown Switch Unit (ERDU), that is used to shut down the device in an emergency, was found to be defective. An ERDU failure could delay or prevent patient access leading to possible patient death or serious injury. | Component design/selection | The firm issue a DEVICE CORRECTION letter to their customers on 2/11/2008. The letter informs the customers of the Emergency Rundown Switch Unit (ERDU) defect and provides a warning which lists instructions for the safe use of the ERDU Unit. The letter informs the customer(s) that their local Hitachi Field Service Engineer (FSE) will be scheduling a visit to test the ERDU in order to verify its state of operation. The letter also states that This test requires specific tools in order to verify the operation without actually quenching the magnet and that Hitachi will begin the test ING in March, 2008 and will continue the program until all installed ERDU units are verified.Should an emergency arise requiring you to activate the ERDU, verify the Emergency Run Down light illuminates and that the magnet does quench (a very loud noise will be audible) before entering the scan room without screening personnel and equipment. Notify the Hitachi Customer Service Department at (800) 800-4925 immediately should you initiate an Emergency Run Down. |  | 235 | 235 units | The recalled device units were distributed to customers located in the following states nationwide: CA, OH, IL, NJ, NY, FL, TX, NE, AR, MD, OR, MN, NV, MS, OK, IN, TN, MO, SC, VA, HI, LA, NC, NM, WY, WA, AL, PA, UT, WA, and KY | Terminated | September 24, 2010 | 956 | 235
Z-2156-2008 | 48472 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Legend tapered dissecting tool, Model F2/8TA23 |  |  | May 30, 2008 | 2008 | Medtronic Powered Surgical Solutions | Cutting flute geometry of surgical dissecting tools may be outside of established tolerance. This may lead to reduced cutting performance, inability to perform desired dissection, increased risk of overheating, and potential need for revision surgery. | Process control | Firm notified consignees by Urgent: Medical Device Recall letter sent May 30, 2008. The letter requires consignees to immediately compare their inventory of products against the list provided and quarantine any unused dissecting tools. The letter also states that Medtronic would provide replacement product to match the quantity and type returned by the consignees' institution. |  | 1306 | 1,306 units | Nationwide Distribution. | Terminated | August 22, 2009 | 449 | 1306
Z-2157-2008 | 48596 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stem Trial, Triathlon Revision Instruments; Non-Sterile; 100 mm length |  |  | May 14, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Triathlon TS Stems Trails may crack, potentially causing fragments of the plastic handle to fragment and break off completely. | Process control | Urgent Product Recall notification letters were sent to all Stryker Branch/Agency Manager/Quality contacts on May 14, 2008 via Federal Express. The letter requests that consignees examine their inventory and hospital consignment locations to identify the product; and retrieve all affected product lots and return it to their branch or agency warehouse for reconciliation. The letter also states that a Stryker Orthopaedics Customer Service rep should be contacted for alternative products available. |  | 16862 | Catalog Number 5565 T-009 to 025: 2629 US, 238 international: Catalog 5566 T-009 to 025: 2625 US, 239 international | Worldwide Distribution --- including USA and countries of Australia, Canada, Sweden, Germany, Spain, Portugal, United Kingdom. | Terminated | January 26, 2010 | 622 | 25292
Z-2234-2008 | 48685 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | StatSampler 200 uL, Li Heparin (micro blood collector) |  |  | June 20, 2008 | 2008 | StatSpin, Inc., dba Iris Sample Processing | EDTA inadvertently mixed with lithium heparin solution contained in capillary tubes. | Mixed-up of materials/components | Iris Sample Processing notified customers by telephone followed by an Urgent: Product Recall letter dated 6/19/08. Users are requested to return inventory, to not use product from Lot Number 7296, and to repeat tests where the recalled tubes were used.The firm identified an additional lot : 8031 was effected and recalled the lot to two consignees. |  | 171 | 171 vials | Worldwide Distribution --- USA including states of CA, CO, FL, KY, KS, GA, NC, NJ, OR, PA, and countries of Canada, Hong Kong, UK. | Terminated | May 30, 2012 | 1440 | 171
Z-1234-2008 | 46383 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Nautica (GK) Mattress Stretcher, manufactured by Georg Kramer GES.M.B.H. &amp; CO. KG A-4061 PASCH ING, Neuhauserweg 14 Austria. These mattress stretchers were sold for installation on or as a component of the following patient bath lifts: a) Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5: Part number 88004051, Nautica Bath Lift System; b) Distributed by Whitehall Manufacturing, P.O. Box 3527, City of Industry, CA 91744, Manufactured by BHM Medical Inc., Canada: Part number 88004051.12, Nautica Bath Lift System |  |  | January 23, 2007 | 2007 | Arjo, Inc. | Repeated use of the ratchets that lock the position of the back and leg rest of the stretcher may damage the retention pin that prevents the latch from disengaging from the anchor system. This would allow the backrest to pivot freely down to the floor. | Device Design | BHM Medical initially notified the end users on 1/23/07 of the potential for the back and leg rest latch ratchets to become damaged, allowing the rests to pivot freely. Should a patient be in the stretcher at the time, the patient could fall, potentially sustaining cuts, bruises or even more severe injuries. The letters provided safety instructions to follow until a permanent fix is available, and requested the accounts to complete and return the enclosed response form acknowledging receipt of the letter and dissemination of the safety instructions to all employees operating the lifts with the affected GK stretchers.BHM sent Customer Field Correction Notification letters dated 1/23/08 to the accounts, along with the Technical Advisory Notice - TAN A dated 1/3/08, reiterating the safety instructions from the 1/23/07 letter and requesting the accounts to, if at all possible, remove the equipment from service until an authorized engineer can replace the ratchets in the backrest and leg rest latches.All consignees have been contacted. |  | 7 | 7 units | Nationwide Distribution --- including states of California, Connecticut, Florida, Mississippi and Texas. | Terminated | August 25, 2008 | 580 | 7
Z-1564-2008 | 46807 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision RXi Digital Remote X-Ray Imaging R&F; System. It is a remote R&F; system consisting of a tilting table with integrated spot film device, 65 or 89 kW pulsing generator and high heat capacity X-Ray tube. The system can be configured to meet a wide variety of digital record and fluoroscopy exams for various clinical applications with either 32 cm or 40 cm image intensifier. |  |  | December 18, 2007 | 2007 | GE Healthcare | GE Healthcare Precision RXi tilt joystick became stuck in the Trendelenburg direction causing the patient to slide off the table. Also recall conducted to investigate the possibility of accidentally activation of the remote console joystick where a fatal accident occurred on a non-GE Remote R&F; System. . | Device Design | Two letters "GE Product Safety Notification" dated 12/18/07 and 3/7/08, addressed to Hospital Administrators, Manager of Radiology and Radiologists, were sent to each consignee. Each letter addressed a separate safety issue. Letters described the affected product, safety instructions, corrective action and contact information. The joystick can become stuck causing unintentional motion of the tube.Field Modification Instructions (FMI 10802/10803) have been issued to the field to install additional warning labels on the patient table, install a protector for the console joystick and inspect and correct the clearance between the console joystick and its mounting location.If your facility has a Precision RXi or Precision e system manufactured before March 2007, a GE Healthcare Field Engineer will schedule a field visit to install the added joystick protection. This upgrade will be performed at no charge.If you have any questions/concerns regarding this issue, please contact your local GE Healthcare sales or service representative. |  | 193 | 193 | (USA)AKCOILLAMAMIMSNCNJNYPASCTNUTVAWAWI(OUSA)BR CA MX AU C3 KR PH TW AE AT BE BF BG CM CY CZ DE ES FI FR GB GE GR GY HR HU IT LB LV MT NA NG PL PT RO RU SA SD SE SI SK SY TN TR UA YU ZA | Terminated | August 09, 2012 | 1696 | 265
Z-1567-2008 | 46921 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | FASTPAK batteries labeled Manufactured for Medtronic Physio-Control 11811 Willows Rd, Redmond, WA USA. Batteries are Nickel-Cadmium (NiCd) battery packs used to provide DC power to LIFEPAK 5, 10, 11, or 12 defibrillator/monitors. The FASTPAK battery has a capacity of 1 amp/hour (1.0AH). The FASTPAK battery can be charged in a Battery Support System, Battery Support System 2, Mobile Battery Service Station, or the LIFEPAK 10, 11 or 12 defibrillator/monitor when the device is powered by an external power adapter. The FASTPAK battery does not have a pushbutton fuel gauge like the LIFEPAK NiCd. The fuel gauge provides a visual indication of battery capacity. |  |  | February 27, 2008 | 2008 | Physio Control, Inc. | Batteries mislabeled - batteries labeled as 2.4 amp/hour batteries, however are 1.0 amp/hour batteries. | Labeling mix-ups | On February 27, 2008, the firm sent a "URGENT MEDICAL DEVICE - ACCESSORY RECALL" letter via certified return receipt to their consignees. The letter states the batteries are fully functional 1.0 amp/hour batteries mislabeled as 2.4 amp/hour batteries and that the batteries will provide half the monitoring time and half the number of defibrillation shock as of a 2.4 amp/hour LIFEPAK NiCd battery. Enclosed with the letter is a confirmation sheet for the consignee to indicate the number of recalled batteries they have and fax back to the firm. The firm will send out replacement batteries and the consignee is to then send back the mislabeled batteries.If there are any questions, contact the technical support group at 1-800-442-1142 - option 5 or visit www.physio-control-notices.come/fastpak. |  | 811 | 811 batteries | Batteries were distributed nationwide and to Canada, Netherlands, Asia Pacific and Latin America. | Terminated | March 07, 2011 | 1104 | 811
Z-1615-2008 | 46965 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art System with software versions 3.0 and 3.1 - Radiation Therapy System. The device is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy. |  |  | February 22, 2008 | 2008 | TomoTherapy Incorporated | Registration adjustment field numbers may appear on printed pages of the Register Tab as nonsense characters or values different than the numbers shown on the Operator Station (OS) display. Setup will move the couch to the correct offset position. If the couch is manually adjusted to the incorrect value, the patient may not be the correct position during treatment. | Software design | Consignees were sent a letter "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" dated February 22,2008. The letter described the Issue, Product Affected, Cause, Recommended Customer Action and Resolution.If you have not received this letter or have questions, contact the firm as outlined below:North America: 1 866 368 4807Belgium: 0800 38783France: 0805 631 565Germany: 0800 000 7541Italy: 800 986 399Netherlands: 0800 0201364Spain: 800 300049Switzerland: 0800 001927United Kingdom: 0808 238 6035China/CNC (Northern): 10 800 712 1701China (Southern): 10 800 120 1701Hong Kong: 800 967912Japan: 0044 22 132374Singapore: 800 1204 683South Korea: 0079 81 4800 7204All other locations: +1 608 824 2900 or +32 2 400 44 44 |  | 32 | 32 | Worldwide Distribution --- USA including states of CA ,CT, DC, FL, IL, LA, MD, MI, MO, NE, NJ, NY, SC, TN, VA, WI , and countries of Canada, Germany, Taiwan, and Turkey. | Terminated | December 02, 2008 | 284 | 32
Z-1558-2008 | 47463 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 24mm Rotating C-Mount Coupler, Model Number 1088-020-121S2. The device in indicated for use in general laparoscopy, nasopharyngoscopy, ear endoscopy, sinuscopy, and plastic surgery wherever a laparoscope, endoscope or arthroscope is indicated for use. |  |  | March 31, 2008 | 2008 | Stryker Endoscopy | The coupler on this device was manufactured without set-screws that are used to hold and center the rear assembly and the endobody. The coupler is part of the camera system and affects the ability to focus. | Nonconforming Material/Component | Only one customer had all of the defective couplers. The firm called this customer on March 27, 2008, and asked that the 28 couplers remaining in their possession be returned. |  | 28 | 28 units | Product was delivered to one consignee in Iowa. | Terminated | September 17, 2008 | 170 | 28
Z-1640-2008 | 47441 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Logix TPN Software Solution, catalog 2M8410, consists of the Logix Order Entry (OE) Installation CD, version 1.0.28 and Logix Compounder Manager (CM) Installation CD, version 3.1.8. The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | March 28, 2008 | 2008 | Baxter Healthcare Corp. | Software anomaly results in inaccurate information being printed out on the Delivery Report, even though the compounder performed the compounding correctly. | Software design | Baxter sent the 3/28/08 Urgent Device Correction letter to affected customers, to the attention of the Director of Pharmacy, via first class mail to advise them of the potential inaccuracy of the Delivery Report due to a software anomaly. The accounts were instructed to review each delivery report for accuracy. If the report contains zero actual volumes the user must manually reprint the report. The letter outlined the process for reprinting the Delivery Report. Any questions were directed to the Baxter Nutritional Professional Services at 1-800-422-2751. The accounts were requested complete the enclosed reply form acknowledging receipt and understanding of the letter and dissemination of the information to their staff. A software upgrade will be provided.If you have questions about this letter or you did not receive this letter contact Baxter Nutrition Professional Services at 1-800-422-2751. |  | 5 | 5 units | Arizona, Minnesota, Wisconsin, Puerto Rico and Canada | Terminated | August 25, 2008 | 150 | 5
Z-1641-2008 | 47471 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-MS"3 Ambulatory Infusion Pumps, Model 7400, Rx only, CE 0473, SN 499000. The device is designed for subcutaneous, intravenous, epidural and intrathecal infusion of medication. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has received reports that the device powers down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their Pump, there is a potentia | Labeling design | A "Smith Medical Urgent Medical Device Correction" letter dated February 22, 2008 was distributed to consignees. The letter addressed to Risk/Safety Managers, Distributors, Clinicians, Patients and other Customers, instructs them to fully tighten the battery cap and to routinely monitor the pump for damage that may result in power down situations. Consignees are to return the "Urgent Medical Device Correction" form.If you have questions about the letter or did not received the letter contact Mike Herbert at 651-628-7049 or Jeff Esselman at 651-628-7920. |  | 2303 | 2303 | AZCAFLGAHIILKYMDMIMNNCNENJNMNYPATNTXUTCANADACHILECZECH REPUBLICGERMANYGREAT BRITAINITALYJAPANNETHERLANDSNORWAYSWEDEN | Terminated | January 11, 2011 | 1054 | 2303
Z-1643-2008 | 47524 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syn go Imaging XS is a picture archiving and communication system intended to display, process, read, report, communicate, distribute, store and archive digital radiological images, including digital mammography images. |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm discovered a problem where the selected patient images on the device may display an additional image from another patient or study. | Software design | The recalling firm issued a Customer Safety Advisory via certified mail to affected customers per Update Instructions IM011/08/S. The letter informs customers of the potential issue and provides instructions to detect its occurrence. A software update to correct the issue is already available. A service representative will visit affected sites to install the software update.If you have questions concerning the advisory letter or you did not receive a letter contact Alexander Vogel at +49 (9131) 84-3840 or alexander.vogel@siemens,com and make reference to complaint 08-1KM-0073 date 2008-03-06. |  | 236 | 236 units | The product was shipped to medical facilities in AL, AZ, CA, CO, CT, FL, GA, IA, KS, KY, MD, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, WI, and WV. | Terminated | November 19, 2008 | 254 | 236
Z-1648-2008 | 47758 | [6 empty] | Seprafilm Adhesion Barrier (chemically modified sodium hyaluronate/carboxymethylcellulose absorbable adhesion barrier). Japan Distribution Only |  |  | April 10, 2008 | 2008 | Genzyme Corporation | Sterility may be compromised | Employee error | Genzyme notified Distributors and hospitals by letter dated 4/10/08, requesting ceasing distribution and return recalled product. |  | 517 | 517 (10 packs) | NationwideKorea, Taiwan, Japan | Terminated | September 29, 2015 | 2728 | 3149
Z-1697-2008 | 47871 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Dental Cement-Bisco BISCEMÂż Dual-Cured Self-Adhesive Resin Cement, Opaque; The product is intended to serve as a temporary tooth filling or as a base cement to affix a temporary tooth filling, to affix dental devices such as crowns or bridges, or to be applied to a tooth to protect the pulp. |  |  | April 28, 2008 | 2008 | Bisco Inc | During product stability testing, Bisco determined that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Bisco also determined that product component(s) may thicken slightly over time. | Incorrect or no expiration date | Bisco sent Urgent Medical Device recall letters dated 4/25/08 to the U.S. consignees via 1st class mail on 4/28/08. The letters informed the users that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Future production lots will be labeled with a shelf life of 1 year with a requirement to refrigerate the product when not in use. The accounts were requested to destroy their remaining stocks of the product and record the number of used or partially used syringes destroyed, indicating the lot numbers of the units destroyed on the enclosed response form and fax it back to Bisco at 847-534-66101. Questions were directed to Bisco Dental Products Customer Service at 1-800-247-3368. |  |  |  | Nationwide including Puerto Rico and internationally to Argentina, Australia, Bulgaria, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Iceland, India, Indonesia, Israel, Italy, Jordan, Korea, Lebanon, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Portugal, Romania, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vietnam | Terminated | September 14, 2010 | 869
Z-1703-2008 | 47522 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Circular Mapping Catheter-Biosense Webster LASSO 2515 Variable Circular Mapping Catheter- The catheter is indicated for multiple electrode electrophysiological mapping of the cardiac structures of the heart, i.e., recording and stimulation only. The device is designed to obtain electrograms in the atrial regions of the heart. |  |  | March 25, 2008 | 2008 | Biosense Webster Inc | This recall was initiated after the firm received several complaints reporting issues with deflection. The catheter may become caught in the mitral valve, requiring surgical procedures, valve may be torn while trying to dislodge, unable to retract the catheter through the sheath and require surgical intervention to remove, pulling on catheter to remove may lead to large atrial septal tear, and sep | Device Design | A letter entitled, "Urgent Safety Field Notice" dated March 25, 2008, and certification form were sent to all customers and Johnson &amp; Johnson affiliates who purchased the product in the past year as of March 25, 2008, due to the expiration date. For customers in the US, two letters and forms were sent out - one addressed to the EP/Cardiology Lab and a second to the Risk Management Office.If you have questions or have not received a letter contact Customer Service at 1-800-729-9010. |  | 14295 | 14,295 units for both products | Nationwide and Worldwide to:Argentina,Australia,Austria,Belgium,Brazil,Canada,Chile,China,Columbia,Czechia,Denmark,Egypt,Finland,France,Germany,Greece,Hong Kong,Hungary ,Iran,Ireland,IsraÂżÂżl,Italy,Japan,Lithuania,Luxembourg,Mexico,Middle East,Netherlands,Norway,Panama,Poland,Portugal,Russia,Serbia,Slovenia, South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | May 03, 2012 | 1500 | 14295
Z-1649-2008 | 47800 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | BCI Non Invasive Blood Pressure Monitor |  |  | April 02, 2008 | 2008 | Smiths Medical PM, Inc. | Due to a component that may be installed backwards, the device may turn off or may immediately reset. | Employee error | Consignees were contacted on 4/2/08 with a Safety Action Bulletin, 08-SAB02, and Technical Worksheet 08-TW01 dated April 1, 2008, which included information regarding the potential safety issue and the proper inspection method for the monitor(s).If you have not received the Safety Action Bulletin or the Technical Worksheet contact Smiths Medical PM, Inc. at Telephone: (262) 542-3100 or Toll Free: (USA): (800) 558-2345 or email info.pm@smiths-medical.com. |  | 195 | 195 | USA: IL, FL, TX, IA, CA, ID, MD, AZ, MN, PA, CO, IN, KS, WA, TN, AL, WI, NM, LA, and PR.OUSA: Argentina, Canada, Ecuador, Greece, Hong Kong, Italy, Pakistan, Peru, Russia, saudi Arabia, Slovakia, Thailand, United Kingdom | Terminated | December 17, 2011 | 1354 | 195
Z-1702-2008 | 47908 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Opthalmic microsurgical system-AMO WHITESTAR SIGNATURE Phacoemulsification System, Product Number NGP680300. The device is a modular ophthalmic microsurgical system that facilitates both anterior segment (i.e., cataract) and posterior segment (i.e., vitreoretinal) ophthalmic surgery. |  |  | April 02, 2008 | 2008 | Advanced Medical Optics, Inc. | AMO initiated this field correction event after becoming aware of a trend in complaints associated with AMO Vitrectomy Cutter used in conjunction with the AMO WHITESTAR SIGNATURE Phacoemulsification System. The physician may be unable to make a cut during cataract surgery. | Device Design | On April 2, 2008, AMO began distribution of an Urgent Safety Notice to all customers in the United States via overnight carrier (Federal Express). The U.S safety notice was used as a template for all global AMO communications regarding the Urgent Safety Notice , many of which required translation into local language prior to distribution. Included in this correspondence is a listing of the serial numbers of the affected products.The Notice informed customers of the problem and gave the customers modifications to the Vitrectomy Cutter Priming Instructions. In addition, the notice informed the customers that an AMO Field Service Engineer will be scheduling a site visit to modify a pressure setting on the AMO Whitestar Signature system. Contact the AMO Phaco Technical support at 1-877-AMO-4LIFE for assistance. AMO will provide updates to the FDA, including all information regarding the recall effectiveness as part of the recall progress reports. |  | 191 | 191 units | Nationwide and to: AL, AR, CA, CO, FL, HI, IL, IN, MD, NC, ND, NJ, NY, OH, OR, PA, PR, SC, TX &amp; WA and worldwide to: Azerbaijan, Belgium, Germany, Denmark, Egypt, Spain, France, Great Britain, Croatia, Hungary, Israel, Italy, Jordan, Kuwait, Lebanon, Liechtenstein, Netherlands, Norway, Poland, Portugal, Saudia Arabia, Sweden, Turkey, South Africa, Australia, India, New Zealand and Singapore | Terminated | January 21, 2010 | 659 | 191
Z-2160-2008 | 48602 | No | IZJ | Cine or spot fluorographic x-ray camera. | Radiology | Radiology | 510(k) | General X-Ray System types with Digital Spot Imaging (DSI) software |  |  | May 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | X-ray images may be stored in the wrong patient file or corrupted, images from two different patients may be combined in one patient's folder or may be combined on one image (corrupted). | Software design | On 5/20/08 the firm issued an URGENT - DEVICE CORRECTION letter to the consignees advising of the problem, requesting consignees to check the stored or printed images to ensure that it represents the data and images captured from the examination. The firm will issue a Field Change Order to their service representatives providing new software, when software is available. This should be issued by the end of June. |  | 41 | 41 units | Nationwide Distribution. | Terminated | September 28, 2010 | 858 | 41
Z-2162-2008 | 48604 | No | FZT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ENDOPATHÂżETS Endoscopic Linear Cutter, TSW35. |  |  | April 24, 2008 | 2008 | Ethicon Endo-Surgery Inc | The firm discovered a component production issue which may cause the cartridge to deploy an incomplete staple line. | Other | Firm issued an Urgent: Device Recall notification instructing the user to not use products from these lot numbers. Examine their inventory immediately to determine if they have affected product on hand and remove them from inventory. Complete the Business Reply Form and fax it back. To return affected product, copy the BRF, place in box with your product, and affix the pre-printed UPS shipping label. |  | 10734 | 10,734 pieces | Worldwide Distribution --- including USA and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CAYMAN ISLANDS, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, MEXICO, PAKISTAN, POLAND, PORTUGAL, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM &amp; VENEZULA. | Terminated | April 03, 2012 | 1440 | 42434
Z-2159-2008 | 48601 | formalin, neutral buffered | IFP | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | 10% Buffered Formalin Phosphate; Certified |  |  | May 01, 2008 | 2008 | ThermoFisher Scientific Co., LLC | Assay tests on retention sample of 10% Buffered Formalin Phosphate showed it to be 3.417%, which is below the specification of 4-5%. | Nonconforming Material/Component | Urgent: Product Recall notification letters were sent by US mail on May 1, 2008. The letters asked that those in receipt of the product check their stock for this product, destroy any of the material that remains from the referenced lot, and then contact their local customer service center for a replacement. The letter also states that the appropriate corrective measures have already been instituted to prevent a recurrence. |  | 1640 | 1052 x 4 liters, 564 x 20 liters | Worldwide Distribution --- including USA and Canada. | Terminated | August 13, 2012 | 1565 | 1640
Z-2167-2008 | 48611 | shield, eye, ophthalmic (including sunlamp protective eyewear and post-mydriatic eyewear) | HOY | Ophthalmic eye shield. | Ophthalmic | Ophthalmic | 510(K) Exempt | Radiation Protective Eyewear (only involving multifocals. |  |  | June 05, 2008 | 2008 | Phillips Safety Products | Radiation Protective Eyewear does not meet the radiation protection levels stated on the labeling. | Process control | Customers were all notified by phone, followed by an Urgent Recall letter dated June 3, 2008. The letter asked that customers examine their stock immediately to determine if they had any of these radiation protective bifocals on hand, discontinue use or distribution of the lot, and set the products aside for further instructions. Also, if customers distributed these items, they needed to contact their accounts and advise them of the recall and have them return their outstanding recall stock to them. The firm will send replacement products in the weeks to follow. |  | 845 | 845 pairs of eyewear | Worldwide Distribution --- including USA and countries of Spain, Norway, Great Britain and the Netherlands. | Terminated | December 24, 2008 | 202 | 845
Z-1716-2008 | 47468 | No | BYJ | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | 510(k) | Generation 4 electronic liquid gauges used on the Caire Liberator, Stroller, Sprint, and Hi-Flow Portable oxygen unit. |  |  | October 23, 2007 | 2007 | Caire, Inc. | Moisture ingress into the electronic liquid level indicator may cause inaccurate liquid oxygen level readings. The oxygen could run out without the patient knowledge. | Device Design | A Technical Service Bulletin was sent to consignees on/about October 23, 2007 to notify customers of the potential issue. Technical support and information for ordering replacement Generation 4 liquid gauges were offered.For technical concerns or if you have not received the notification from the firm, contact U.S. Technical Service at 1-800-482-2473 or European Technical Service at +44 (0) 797 1260077. |  | 3470 | 3470 | U.S. Nationwide and the following foreign countries: Argentina, Austria, Brazil, Canada, China, Chile, Central America/Caribbean, Colombia, Costa Rica, Europe, Austria, Belgium, Denmark, France, Germany, Greece, Hungary, Irish Republic, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, United Kingdom, Mexico, South America, Peru and Uruguay. | Terminated | December 17, 2009 | 786 | 3470
Z-1868-2008 | 48427 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encore Modular Femoral Neck, 12 degree/35 mm; Ref #410-35-112; distributed by Encore Medical, Austin, TX 78758. |  |  | May 08, 2008 | 2008 | Encore Medical, Lp | Product not manufactured to specifications; modular femoral necks found to contain 24 scallops, instead of the standard 12. | Nonconforming Material/Component | Firm notified sales agents of recall on 05/08/08 via an Urgent: Medical Device Recall letter. Sales agents asked to immediately examine their inventory and quarantine product subject to recall. Agents are asked to notify their customers if they had further distributed the product. All product must be returned to Encore along with the response form. |  | 20 | 20 units | Worldwide Distribution --- USA including states of IN, KS, GA, WA, SD, and countries of Greece and Australia. | Terminated | September 08, 2009 | 488 | 20
Z-2154-2008 | 48467 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CorometricsÂż MODEL 250 Series Maternal/ Fetal Monitor, Models 256 and 259, part number 2024489-002. |  |  | May 05, 2008 | 2008 | Datex Ohmeda, Inc. dba GE Healthcare | Communication connector within monitor is defective and may cause potential loss of telemetry. | Component design/selection | The recalling firm notified healthcare professionals by an Urgent Medical Device Correction letter dated 07/01/2008 to Healthcare Administrators, Risk Managers, Nurse Managers and Directors of Clinical Engineering. The letter advised of a defect in a communication board that may cause a potential loss of telemetry monitoring. The firm recommended that the device should not be used with the firm's 340 Series telemetry system pending replacement of the defective board by service engineers. The Healthcare Professionals are asked to contact the GE Healthcare Technical Support center at 1-800-558-7044. |  | 133 | 133 units | Nationwide Distribution. | Terminated | August 07, 2012 | 1555 | 133
Z-2165-2008 | 48607 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Scimed Maverick 2 Monorail" PTCA Catheter, 3.0 mm 15 mm |  |  | May 01, 2007 | 2007 | Boston Scientific Corporation | The units within this lot may not meet thickness specifications surrounding the wire exit port wall, which may result in an air embolization if the Port Septum Wall burts. | Nonconforming Material/Component | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required" letter dated May 11, 2007. The letter addressed to Risk mangers, described the issue and product. The letter requested the immediate discontinued use of and segregation of recalled product, A Reply Verification Tracking Form was requested to be completed and returned. |  | 24 | 24 Units | International Distribution --- including countries of Greece, Germany, Great Britain, Spain, Sweden, Belgium, and Algeria. | Terminated | September 10, 2008 | 498 | 24
Z-2170-2008 | 48615 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ev3 Primus GPS Biliary Stent System |  |  | April 22, 2008 | 2008 | Ev3, Inc | A Primus GPS Biliary Stent System was mislabeled. The mislabeling resulted in a 8 mm diameter, 27 mm length stent being labeled as a 8 mm x 37 mm stent. | Employee error | The single affected consignee was visited by an ev3 representative to verify the use of the device or exchange it for an unaffected device. Information recorded on an ev3 Device Reconciliation Form. |  | 1 | 1 Unit | Nationwide Distribution --- state of NJ | Terminated | December 10, 2011 | 1327 | 1
Z-2166-2008 | 48608 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Scimed Maverick 2 Monorail 12mm Balloon Catheter, 2.5 mm x 12 mm |  |  | November 01, 2005 | 2005 | Boston Scientific Corporation | Two lots may not be correctly labeled. Specifically , a 2.5 mm x 12 mm unit was incorrectly labeled as a 3.5 x 20 mm unit on the box and tyvek label. | Packaging process control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required"letter dated November 7, 2005. The letter addressed to Risk mangers, described the issue and product. The letter stated that immediate discontinued use of and segregation of the recalled product was required. A Reply Verification Tracking Form was requested to be completed and returned. |  | 67 | 67 Units | International Distribution --- including countries of Czech Republic, France, Germany, Great Britain, Italy, Spain, Sweden, India, and Iran. | Terminated | September 10, 2008 | 1044 | 67
Z-2169-2008 | 48614 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ev3 Protege GPS Biliary Stent Systems |  |  | April 22, 2008 | 2008 | Ev3, Inc | A Protege PS Biliary stent system was mislabeled. The mislabeling resulted in a 9 mm diameter, 60 mm length stent being labeled as a 6 mm x 60 mm stent. | Employee error | The single affected consignee was visited by an ev3 representative to verify the use of the device or exchange it for an unaffected device. Information recorded on an ev3 Device Reconciliation Form. |  | 1 | 1 | Nationwide Distribution --- state of LA. | Terminated | December 10, 2011 | 1327 | 1
Z-1776-2008 | 47807 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Q-Logic Controller Software used on the following power wheelchairs: Q6000, Q6000XL, Q6000Z, R4000, Q600, Q600XL, Q610, Q1107, Q600E, and Dynamo 1107 |  |  | April 11, 2008 | 2008 | Pride Mobility Products Corp | Software - Watch dog timer feature was disabled. | Software design | The recalling firm issued an Urgent Device Correction letter on 4/11/08. This letter instructs the dealers to contact the end users so that the units can be upgraded to enable the watch dog feature. |  | 4129 | 4129 | Class II Recall - Worldwide Distribution --- including USA and Canada. | Terminated | November 13, 2009 | 581 | 4129
Z-1774-2008 | 47795 | [6 empty] | Hall Easy-FitÂż Prosthetic Heart Valve - Aortic A7700. The device uses a single pyrolytic carbon disc occluder pivoting within a titanium housing. The housing is machined into final configuration from a single piece of solid titanium and has no welds or introduced bends. It is an open-membered design in which all members project centrally without distal attachment. A rotatable fabric sewing ring is affixed to the valve. The valve is provided on a holder (collect) to facilitate insertion and to provide protection during shipping and handling. The valve is sterile and non-pyrogenic if seal and package are not opened, damaged, or broken. |  |  | March 25, 2008 | 2008 | Medtronic Heart Valve Division | Medtronic has identified an issue with 685 Hall Easy-FitÂż valves Models A7700 and M7700 in which the holder of the valve was incorrectly secured to the valve. This could potentially make it difficult to easily disengage the holder from the valve. | Process control | A Medtronic "Urgent Communication" letter dated March 24, 2008 was sent to consignees on 3/25/08. The letter describes the product and problem. The letter requests that they return the impacted devices and to return the reconciliation form "Medtronic Recall Certificate (FCA0807)" confirming whether they have the valves in inventory or if they have been implanted. |  | 69 | 69 (USA), 117 (EUROPE), 5 (CANADA), 179 (HONG KONG) | Class II Recall - Worldwide Distribution --- including USA states of TX, WY, CA, MI, WI, MO, IL, NY, TN, AR, LA, KY, WA, WV, and countries of Canada, Hong Kong, Australia, Albania, Belgium, Croatia, Germany, India, Italy, Netherlands, Poland, Romania, Russian Federation, South Africa, and Turkey. | Terminated | May 15, 2012 | 1512 | 99
Z-1777-2008 | 47819 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Century+ Bed Model P1400 HBSW siderail upgrade kits; accessory model numbers P9964-11, P9964-12, P9964-13, P-9964-14, P9964-15, P9964-16, P9964-41 and/or P9964-42. |  |  | April 28, 2008 | 2008 | Hill-Rom, Inc. | A risk of entrapment will exist between the head and foot rails if installation instructions are followed. | Other | Hill-Rom issued an Urgent Medical Device Correction letter on 4/28/2008 notifying consignees that they will be visited by a qualified Hill-Rom service technician or designee and provide a revised copy of the instruction sheet. The technician will either assist the facility to ensure proper installation or verify by signature that the facility understands the issue and will verify their beds had the upgrade kits installed in such a way that an entrapment gap does not exist. |  | 69 | 69 | Class II Recall - Nationwide Distribution --- including states of Florida, Illinois, New Mexico, Texas and Virginia. | Terminated | September 22, 2009 | 512 | 69
Z-1778-2008 | 47829 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Patella System Reamer Shaft Assembly; |  |  | April 15, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | The Xcelerate Patella Reamer Shaft Assembly and the Reamer Adapter do not mate properly. This could result in the removal of excessive bone from the patella leading to a thinner patella; a patella that does not have enough bone for adequate engagement with pegs or patella fracture. | Process control | An IMPORTANT MARKET WITHDRAWAL letter was issued on April 19, 2005 and sent Federal Express to those consignees who received the device. The letter instructed them to examine their inventory to identify instruments, reconcile these instruments on an attached accountability form; fax a copy of this form within 2 days of receipt of the letter; retrieve and return affected product; contact Stryker Customer Service to re-order. |  | 24 | 24 total | Class II Recall - Worldwide Distribution --- including USA and Canada, United Kingdom. | Terminated | September 24, 2008 | 1258 | 24
Z-2174-2008 | 48616 | No | LKK |  |  | General Hospital | PMA | Model 8731SC Intrathecal Catheter |  |  | April 14, 2008 | 2008 | Medtronic Neuromodulation | The SC catheters cannot completely engage with the Model 700-04M portal connector. The result of this condition is a possible leakage or disconnect of the SC catheter from the portal during the continuous therapy trial period. | Device Design | Medtronic informed all Belgium implanting and managing physicians of the SynchroMed II infusion system. Each customer was mailed a Dear Healthcare Professional letter (Medtronic Urgent: Field Safety Notice) on 4/3/08 providing important safety information and patient management recommendations regarding the device incompatibility. The firm recommends that healthcare professionals do not use SC catheters if a connection with an implantable access port is required. Also, if a continuous infusion trial is required with use of an implantable access port, they should use a model 8711 intrathecal catheter. The firm states that for patients who are currently undergoing a continuous trial with an implantable port, replacing the catheter connector before connecting to the SynchroMedd II using the Medtronic model 8578 revision kit is recommended. Medtronic also recommends discussing the risks of connection failure with patients. Local field reps will personally follow-up with affected customers to confirm and document notification. |  | 232 | 232 total from all models | International Distribution --- country of Belgium. | Terminated | September 10, 2008 | 149 | 232
Z-1309-2008 | 46834 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HOUVA III Phototherapy System with DermaSense Technology comes on one skid within 10 cardboard boxes. |  |  | February 19, 2008 | 2008 | National Biological Corp | Software allows operator to override "low line voltage" error warning and store light intensity value. | Software design | National Biological sent an "Important safety reminders for your Houva DermaSense phototherapy device" letters, dated 2/11/08 to users. It states that the user should collect a "Profile" of the light energy and record it on a daily/weekly flow chart. The chart should be in plain view so each operator can clearly determine the units output. If an error message appears in the display panel, treatment should not be continued unless the phototherapy booth operator was able to identify and correct the cause of the error. Make sure both breakers are in the "ON" position before starting up the unit each morning and before giving treatment. |  | 48 | 48 Units | Worldwide distribution --- USA including states of: CA, DE, FL, GA, ID, IL, IN, MA, MI, MN, NC, NH, NY, OH, OR, PA, TX, VA, WA and WI, and country of Canada. | Terminated | September 20, 2011 | 1309 | 48
Z-1650-2008 | 47885 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Ambulift, standard basic model manually powered patient lift. The Ambulift is used to move and lift patients from their beds, transport them to the bathroom, and lift them in and out of the bath. |  |  | May 05, 2008 | 2008 | Arjo, Inc. | The armrests of the lift chair may push up, allowing the patient to slip down and out of the seat. The resultant fall could result in serious or fatal patient injuries. | Device Design | Arjo sent Customer Field Correction Notification letters dated 5/2/08 on the Arjo letterhead to the end user accounts who received the affected Ambulift Patient Lifts, to the attention of the Administrator/Risk Manager, via UPS 2nd day delivery on 5/5/08. The letters informed the consignees of the incident with the patient slipping from the seat of the lift and sustaining fatal injuries and advised them to always ensure that a clinical assessment to verify that the resident has sufficient upper body strength to maintain a steady position on the chair, and that the resident has sufficient mental capacity to adhere to the caregiver's instructions, is carried out by a qualified nurse or therapist before lifting residents with the Ambulift. To further enhance safety, a seat belt is being made available, which must be fitted to the chair, enabling it to be used when required by clinical assessment. Instructions on how to apply, use, and maintain the seat belt will be provided with the belts. The belts will be provided at no cost. The accounts were requested to review, complete and return the enclosed customer response form, indicating the number seat belts needed. Any questions were directed to Arjo's Quality Department at 1-800-323-1245, ext. 6140 or 6165. |  | 11 | 11 lifts | Worldwide Distribution-USA includes the states of AL, AZ, DE, FL, KY, MO, NY, OK, and VA | Terminated | February 03, 2010 | 639 | 14
Z-2153-2008 | 48466 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Acu-Loc Distal Radius Plate, Standard, Left. Product labeled in part, " Acu-Loc Distal Radius Plate, Standard, Left... PL-DR50L W48003... Manufactured or Distributed by: Acumed LLC Hillsboro, OR ". |  |  | June 29, 2007 | 2007 | Acumed LLC | Plates are incorrectly laser marked. The PL-DR50L plates are incorrectly laser marked with PL-DR70L. | Process control | The firm issued an Urgent Notification: Product Removal letter on 6-29-2007. The notification was reportedly sent by either FAX or e-mail. The notification informed consignees to remove the plates from their inventory and return them to Acumed. The notification did not include any instructions for a sub-recall. On 6/13/08, the firm sent a second Urgent Notification: Product Removal letter. This letter describes the product as PL-DR50L - Lot Code - W48003, explains the problem and asks customers to return product to Acumed. |  | 51 | 51 pieces | Worldwide Distribution --- including USA and countries of South Africa, Italy, Finland, and Brazil. | Terminated | February 03, 2010 | 950 | 51
Z-2155-2008 | 48468 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm X 32.5mm Cortical Screw used with the Polarus Humeral Fixation System,. Product labeled in part, "Polarus Humeral Fixation System&amp; SIZE: 3.5mm X 32.5mm Cortical Screw... REF HCO3325-S... Acumed 5885 NW Cornelius Pass Road Hillsboro, OR 97124-9432 ". |  |  | December 27, 2005 | 2005 | Acumed LLC | Package of 3.5mm X 32.5mm cortical screw may contain 3.5mm X 22.5mm cortical screw. | Process change control | The firm issued a Product Removal notification on 12-27-2005 to their customers. The notification was sent by either FAX or by e-mail. The Product Removal Notification informed their customers of the issue, asked them to remove the screws from inventory and return the screws to Acumed. A reminder notification, dated 1/20/08 was issued to customers identifying the product and the problem and requesting the product be returned. |  | 112 | 112 pieces | Worldwide Distribution --- USA including states of CA, KY, FL, TX, OH, OR, NY, ID NJ, NE, LA, WI, UT, MO, MS, GA, AZ and countries of United Kingdom, Italy, Puerto Rico, Turkey, Spain, Japan, South Africa, and Belgium. | Terminated | February 03, 2010 | 1499 | 112
Z-2201-2008 | 49030 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Pacific Consolidated Industries Mobile Oxygen Storage Tank, MOST 20-2B, Part Number: 793080-002 |  |  | September 08, 2007 | 2007 | Pacific Consolidated Industries LLC | The recall was initiated after the firm received two customer complaints of ruptured bourdon tubes that had resulted in bellowed out face gauges. With one unit the operator was able to shut off the oxygen flow, however one rupture resulted in a fire. No injuries were reported. | Vendor change control | PCI sent a Recall Notification letter to the Spanish Air Force September 8, 2007. The Recall Notification letter was mailed to the US Defense Supply Center on July 24, 2008 after a recall agreement was signed on July 8, 2008 between Pacific Consolidated Industries and the Defense Supply Center. The recall notification to Spanish Air Force informed the consignee of the problem and asked the consignee to remove the affected units from service until a Technical Service Representative could be sent to Spain to inspect the affected units and perform the corrective action. The recall notification to Defense Supply Center instructed the consignee to keep the affected units in quarantine (units placed in quarantine by US Air Force August 3, 2007) until they have been refurbished under recall agreement that was signed on July 8, 2008. |  | 1032 | 1032 units | Worldwide to US Air Force and also to Spain | Terminated | November 01, 2012 | 1881 | 1032
Z-1617-2008 | 47388 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific SteeroCath-T Ablation Catheter. This catheter is indicated for general indications, treatment of supraventricular tachycardia (svt) and ventricular tachycardia (vt). This product in not indicated for use as an ablation catheters indicated for treatment of atrial flutter or for atrial fibrillation. |  |  | February 14, 2008 | 2008 | Boston Scientific Corp | Package sterile barrier may be breached, compromising sterility. | Nonconforming Material/Component | Recall initiated on Feb. 14, 2008, via Urgent Medical Device Recall notification being sent to consignees, along with a reply verification tracking form.If you have questions or need a copy of the recall notification, contact Deborah J. Springer at 1-408-935-4944 or springed@bsci.com. |  | 827 | 827 units total | Worldwide Distribution --- including USA and countries of Saudi Arabia, Spain, Italy, Israel, Denmark, Canada and Australia. | Terminated | January 08, 2011 | 1059 | 827
Z-1705-2008 | 47774 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Image Intensified fluoroscopic x-ray system (includes OEC 9800, OEC FluoroTrak 9800 Plus, OEC 9800 Plus, and OEC 9800MD Motorized C-arm System), GE Healthcare, Surgery, Salt Lake City, UT. |  |  | April 15, 2008 | 2008 | OEC Medical Systems, Inc | Use of existing four-pedal footswitch on a different machine may cause various operational errors. | Device Design | Consignees were notified by letter on 03/21/2007 and advised that GE would be replacing the 4-pedal switches at some time in the future. Replacement began on 04/15/2008. Contact GE Healthcare at 1-800-874-7378 for assistance. |  | 78 | 78 units | Nationwide, including one VA facility in FL. No military or foreign distribution. | Terminated | December 21, 2011 | 1345 | 103
Z-1642-2008 | 47487 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | AD 7 Patient table supplied with Allura Xper FD10 and Allura Xper FD10/10. Product is non-tilting but can be purchased with a tilt option. |  |  | March 14, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for radiologic patient table to become immobile and unable to move again due to force sensor sensitivity to electromagnetic radiation. | Device Design | On 03/22/2004, firm sent letters via certified mail to notify customers of issue. Letter advises customers of potential risk for patients and users, and states to not use system for critical examinations or treatments if table movements suddenly stops or makes a freeing movement. The letter indicates a representative will contact customer for installation of a Hardware filter, free of charge, to prevent Electromagnetic interference (EMI) from affecting the force sensor. Please contact Philips at 1-800-722-9377 #5, #4, #1. |  | 39 | 39 units | Product is distributed to hospitals and medical centers nationwide to the following states: ND, FL, KY, MN, VA, ME, PA, NJ, CT, TX, CA, OH, TN, OR, GA, NY, SC, MO, WA, IL. | Terminated | July 18, 2012 | 1587 | 39
Z-1715-2008 | 47384 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | In vitro Diagnostic-Clinitest hCG Cassette Pregnancy Test, Radio immunoassay for Human Chorionic Gonadotropin (hCG) test system Product Number: 1760 The Clinitest hCG Pregnancy Test is for in vitro diagnostic use as a qualitative method in the rapid detection of human chorionic gonadotropin (hCG) in urine specimens. |  |  | April 04, 2008 | 2008 | Siemens Medical Solutions Diagnostics | False negative hCG results due to decreased sensitivity | Component design/selection | Siemens notified Distributors ( Distributor Bulletin Urgent Field Safety Notice) and end-users (Customer Bulletin Urgent Field Safety Notice) Clinitest hCG Cassette Recall by letter dated 4/3/08. Users are requested to, discontinue use, discard the product, and immediately contact Siemens technical support representative or distributor for replacement cassettes from a different lot.If you have any question or have not received the recall notification, contact your local Siemens Distributor Relations Representative or Siemens Technical Support at 877-229-3711 option 13 then 5. |  | 8180 | 8180 | Nationwide | Terminated | March 15, 2012 | 1441 | 8180
Z-1524-2008 | 47517 | No | MIE |  |  | Cardiovascular | Contact ODE | CryoValve, Pulmonary Valve &amp; Conduit |  |  | March 26, 2008 | 2008 | CryoLife, Inc. | Tissues taken from a donor whose blood culture was found positive for Serratia marcescens, was distributed. | Material/Component Contamination | Hospitals were notified by letter on 3/24/2008. They were advised not to implanted the tissue and if implanted, they do not recommend explant unless medically necessary. A self addressed, stamped post card was included to confirm receipt of the notification. Serial number 8895391 was reported implanted. Serial number 8895131 was returned to CryoLife for destruction. |  | 1 | 1 tissue | Hospitals in DE, TN | Terminated | January 12, 2010 | 657 | 1
Z-1593-2008 | 47424 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinderÂż Spinal Fixation System T-Handle Driver; Part 2151-7 (13" shaft); Non-sterile, Reusable Instrument; manufactured by Abbott Spine, Austin, TX 78727. |  |  | March 24, 2008 | 2008 | Abbott Spine | Reports received regarding disassembly of T-Handle Driver components during surgery dropping the closure top into the surgical site. | Error in labeling | Firm sent notification letters to sales agents via letter on 03/24/08. Enclosed with the notification letter, the firm provided updated product inserts for the InCompassÂż and PathFinderÂż Spinal Fixation System (Cat. #2199-0002-LB Rev. H and #1199-0002-LB Rev. H, respectively), as well as the "Disassembly and Assembly Instructions for PathFinderÂż &amp; InCompassÂż T-Handle Drivers" (Cat #1199-0040-MKC Rev. A).Sales Agents instructed to distribute updated inserts and instructions to all locations that handle the drivers for cleaning and to destroy all copies of old inserts.If you have questions or have not received the recall notice from the firm, contact Kirstin Widding at 512-533-1852. |  | 224 | 224 units. | Nationwide Distribution --- including states of AR, AZ, CA, CT, CO, FL, IL, LA, MD, MI, MO, MN, MT, NC, NM, NJ, NY, OH, OK, OR, PA, SD, TX, VA and WI. | Terminated | December 02, 2009 | 618 | 502
Z-1475-2008 | 47352 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinitest hCG Cassette Pregnancy Test, Radio immunoassay for Human Chorionic Gonadotropin (hCG) test system |  |  | October 17, 2007 | 2007 | Siemens Medical Solutions Diagnostics | False positive hCG results | Employee error | Siemens issued a Service Bulletin to Siemens Healthcare Diagnostics Branches via email on 10/17/007, and a customer bulletin to outside the US . Users are requested to discontinue use and discard the product. A Completion Notification was requested to be completed and returned by the faculties. |  | 3065 | 3065 kits | NationwideCanada, France, Korea, Netherlands, UK | Terminated | March 15, 2012 | 1611 | 3065
Z-1749-2008 | 47536 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UBiTÂż-IR300 Infrared Spectrophotometer, Software Versions 7.0, 7.01 and 7.11. |  |  | January 18, 2008 | 2008 | Otsuka America Pharmaceutical, Inc. | Medical device power supply may overheat and cause smoke to be emitted from the device. | Nonconforming Material/Component | The initial distributor in the US notified consignees 03/14/08 and 03/26/08 by letter. Consignees were provided two different letters depending on the hours of use of the spectrophotometer. The notifications advised of the firm's preventive service and upgrade program for the instrument's power supply, in which consignees would receive a temporary replacement instrument while the service is completed. Consignees with spectrophotometers having &lt;15000 light source hours were notified similarly, and were provided filters and instructions for cleaning and replacing the filters to protect the power supply from unnecessary accumulation of debris. Contact firm at 1-888-MERETEK for assistance. |  | 517 | 517 units | The affected medical devices were distributed to medical facilities and physician offices nationwide and to distributors in Mexico. | Terminated | February 18, 2011 | 1127 | 517
Z-1646-2008 | 47535 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | i-CAT 17-19 Cone Beam 3D Dental Imaging System. This product is also sold under a private label as KaVo 3D eXam |  |  | March 19, 2008 | 2008 | Imaging Sciences Intl Inc | Incorrect scatter information in the manual. | Error in labeling | Firm issued Urgent Medical Device Recall letters dated 3/19/08 to their customers. The distributors were provided background information and specific instruction on the steps to perform to correct the problem.If you have questions or have not received this letter, contact Francine McCoonse at 215-997-5666 ext 1340. |  | 1203 | 1203 units | The products were shipped to distributors and dental offices nationwide. The products were also shipped to Canada, Austria, Belgium, Czech Republic, England, Finland, France, Germany, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Portugal, Lavenia, Spain, Sweden, Switzerland, Turkey, United Kingdom, Australia, New Zealand, Chile, Israel, South Korea, UAE, India, Argentina, Brazil, China, Columbia, Hong Kong, Jamaica, Korea, Lebanon, Malaysia, Morocco, Russia, Scotland, Taiwan, Thailand, and Singapore. | Terminated | May 05, 2009 | 412 | 1203
Z-1748-2008 | 47527 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Cyclone anterior cervical plate, 4 Level, 88MM (Ti-6Al-4V); REF 600-04-188. |  |  | March 26, 2008 | 2008 | Zimmer Inc. | The screw locking cap may fracture when the surgeon rotates it into the locked position. | Device Design | Consignees were notified via voicemail on 3/26/08 and a letter was sent to all distributors on 3/27/08 informing them that the plate should no longer be used and that all lots are being recalled. A letter dated 6/18/08 was issued to implant physicians recommending patient monitoring. Contact Zimmer Spine at 1-952-857-5636 for assistance. |  |  |  | Arizona, California, Florida, Georgia, Louisiana, Michigan, Minnesota, Montana, New Hampshire, North Carolina, Oklahoma, Pennsylvania, Tennessee, Texas, Utah, Washington and Wisconsin. | Terminated | November 12, 2009 | 596
Z-1591-2008 | 47423 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂżÂż Insulin Pump, model 21-1800. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smith Medical has received reports of Deltec CozmoÂżÂż Insulin Pumps powering down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the Pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their blood gl | Device Design | Consignees and the USA demonstration pump physicians were mailed an "Urgent Medical Device Correction" letter on 2/22/08. The letter stated Details on Affected devices, Description of the problem, Advice on action to be taken by the user, Additional Information and New Warning. These consignees will be required to return the Correction Confirmation Form. The 17 countries with international mailing addresses were sent the Urgent Medical Device Correction letter (Urgent Field Safety Notice) via the Smiths Medical International Regulatory department based in Watford UK. The Smiths Medical Distributors and Affiliates will be required to mail the Safety Notice to their direct users. The USA distributors were mailed the Urgent Medical Device correction letter with a cover letter. These consignees will be required to add the letter to in-house inventory. The USA prescribing physician and the the USA Managed Care Accounts (Insurance Companies) were mailed the Urgent Medical Device Correction letter with a cover letter. These consignees will not be required to respond and no follow-up will be done by Smiths Medical.Contact Mike Herbert at 651-628-7049 if you have any questions or did not receive the recall information from the firm. |  |  |  | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and countries of AU, CA, CL, CN, DE, EG, FR, GB, GR, IL, KW, MX, NL, NZ, ON, PL, SA, SE, TR, UK. | Terminated | January 19, 2011 | 1062
Z-1766-2008 | 47529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 ultrasound system at SW revision 200.0.054 or 200.0.059A, with one of the following options installed: 3-Scape Imaging, Factory, PN 05937011 or 3-Scape Real Time Imaging, UPG ACU, PN 10033688. Material number 10032746, 10037592, 10032746, 10037591 or 10038202. |  |  | March 21, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Software issue may result in distorted images and inaccurate measurements, which could lead to a misdiagnosis. | Software design | The firm initiated a field correction on 03/21/2008. Site visits to correct the problem and install the revised software version 200.0.063A are planned for all consignees. No other notification was made. The firm can be contacted at 650-694-5993. |  | 488 | 488 active units total | Product was distributed worldwide to 539 consignees, in the USA to MD, SC, AZ, FL, IL, WI, NJ, NC, CA, OH, VA TX, PA MO, NH, IA, NE, MN, NY, GA, MI, WA, HI, MA, SC, WV, VA, ND, IN, WY, UT, MS, TN, PR, SD, KY, OK, OR, and worldwide to Sweden, Saudi Arabia, Austria, New Zealand, Italy, Switzerland, Mexico, Germany, Australia, Canada, Great Britain, France, China, South Korea, Singapore, Denmark, Brazil, Poland, Norway, Faroe Islands, South Africa, Russian Federation, Romania, Trinidad and Tobago, Greece, India, Spain, Thailand, Ukraine, Nicaragua, Japan, India, Czech Republic, Egypt, Taiwan, Turkey, Malaysia, Qatar, Bosnia an dHerzogovina, Vietnam, Brunei Darussalam, Italy, Kuwait, Turkey, Portugal | Terminated | September 30, 2009 | 558 | 488
Z-1823-2008 | 47834 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Headset Model Number RTC-2K for use with the Serene Sound Digital MRI Compatible High Fidelity Stereo Sound System. |  |  | May 14, 2007 | 2007 | Resonance Technology Inc | Potential safety concerns - Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to the skin. | Device Design | The strategy was notifying customers by written communication. The letter advises Resonance Technology, Inc. is voluntarily recalling the MRI Headset, Model No. RTC-2K and RTC- 551-070 because of potential safety concerns. Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to skin. The letter also advises that Resonance Technology Inc. has designed and developed the significantly improved MRI Headset, Model No. RTC-661-060-09P-001 (upgrade) and began shipping this product in January 2007. Customers who received the Headset Model; RTC-2K and RTC- 551-070, are asked to upgrade to MRI Headset Model No. RTC-661-060-09P-001. Response is requested as to the number of headsets to be replaced. In addition, consignees are advised that if they further distributed this product, to identify their customers and notify them at once of this product recall. If you have questions, contact Olivia Shayesteh at 818-882-1997 or via email at Olivia@mrivideo.com |  | 20 | 20 units | Nationwide and Internationally: Canada, Japan, Taiwan, United Kingdom, Netherlands, Italy, Spain, Germany, Switzerland, France, | Terminated | July 07, 2011 | 1515 | 49
Z-1826-2008 | 48078 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 06, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software anomaly results in incorrect study date and time information being displayed in the report screen and title, which may result in a potential patient misdiagnosis. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 5/5/08 to all customers who have the GE Centricity PACS Software Versions 3.0.0.X, 3.0.1.X, 3.0.2.X and 3.0.3.X informing them of the incorrect study date and time information being displayed in the report screen and title, which may lead to a potential patient misdiagnosis. The users were instructed to refer to the image overlay for the correct date and time until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 177 | 177 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Italy, Japan, Malta, South Korea, Spain, Switzerland and United Kingdom. | Terminated | February 03, 2010 | 638 | 177
Z-1821-2008 | 48123 | No | MVV |  |  | Neurology | 510(k) | Psi Bands acupressure/acustimulation wrist bands in the following colors: Racer Black (UPC 859570001036), Daisy Chain (UPC 859570001029), Color Play (UPC 859570001043) and Cherry Blossom (UPC 859570001012). |  |  | April 26, 2007 | 2007 | Psi Health Solutions | Product was marketed before its 510(k) application cleared. | PMA | Recall imitated 04/26/07. Retailers were informed via phone and email and mail.The product was marketed prior to the firm receiving clearance to market from the FDA. Contact Carla Falcone at 831-373-7712 if you have questions. |  | 21000 | 21000 units | Product was distributed to approximately 21,000 consumers and retailers throughout the US. | Terminated | December 18, 2008 | 602 | 21000
Z-1825-2008 | 47840 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity AW Suite 2.0 Software running on the Centricity RA 1000 or RA600/CA1000 Workstation; for diagnostic image analysis; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 09, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta protocol or customized protocol of the Centricity AW Suite 2.0 running on the Centricity RA 1000 or RA600/CA1000 workstations. | Software design | GE Healthcare IITS sent Urgent Safety Notice letters dated 5/9/08 to all customers who have the AW Suite Software Version 2.0, informing them of the incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta analysis protocol or customized protocols. The users were instructed not to use the aorta analysis protocol or the customized protocols for an aneurysm until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 207 | 207 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Great Britain, Hong Kong, Ireland, Israel, Italy, Japan, Malaysia, Malta, Portugal, Saudi Arabia, South Korea, Spain, Sweden, Switzerland and Turkey. | Terminated | September 24, 2010 | 868 | 207
Z-1752-2008 | 47245 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | PT-3B Control Unit - component of the EndoWave Infusion Catheter System. EndoWave Infusion System consists of three main components, a single use EndoWave Infusion Catheter System which is comprised of a Drug Delivery Catheter (CD) and removable Ultrasound Core (USC) and a reusable EKOS PT-3B Control System. Box is labeled EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core |  |  | March 10, 2008 | 2008 | EKOS Corporation | Potential for unit to burn patient. | Software design | In March 2008, the firm planned to visit each consignee to install the new software version 01.05.02 and hand deliver letters, dated 02/20/2008, notifying consignees of the reason for the action. The old software version is 01.04.02. The software was upgraded because the Ultrasound Core heats very quickly and if removed from the DDC while operating it may result in burns to the patient or user. The software will turn off the Ultrasound Core automatically if removed from the DDC during operation, and will prompt to restart therapy if ultrasound is turned off during treatment.On 04/17/2008, the firm issued a new letter with the date of April 2008 informing customers of the issue. In addition the letter indicates that the corrective action, release of the new software is delayed to incorporate further changes, and that customers will be notified when the software is available for their unit. Customers are advised that the device is safe to use when operating according to the Instructions for Use. EKOS contact number 1-888-356-7435 |  | 218 | 218 units | Products were distributed nationwide to hospitals. | Terminated | May 13, 2011 | 1159 | 218
Z-1820-2008 | 48116 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Allez Spine Del Mar Pedicle Screw System: IT-FN 1007, Del Mar Sterilization Tray (Part Number: 30013) |  |  | April 14, 2008 | 2008 | Allez Spine, LLC | This field action was initiated because the cleaning, decontamination and sterilization procedures specified in the "Instructions for Use" for the impacted systems are incorrect.The use of these parameters in performing cleaning, decontamination and sterilization procedures may not result in clean and sterile product. | Error in labeling | Allez Spine issued a safety notice with attached fax-back attachment on April 14, 2008 via FedEx to all its hospital users, distributors and field representatives informing them of the failure of the sterilization validation and requesting them to use the new recommended sterilization parameters. The notice included copies of the updated Instructions for Use for the Allez Spine Laguna Pedicle Screw System and the Allez Spine Del Mar Pedicle Screw System. The safety notice requested that the revised recommended instructions be used for all future cleaning, decontamination and sterilization. Copies of the revised instructions were also provided with the safety notice along with links to the company website where additional copies may be downloaded.Contact Mark De Baca at 415-357-9908 if you have any questions. |  | 2 | 2 systems | Nationwide to AZ, CA, CO, GA, IL, NV, TN, TX, UT | Terminated | January 25, 2010 | 651 | 66
t | o | o |  | b | i | g
Z-0156-2009 | 49452 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Weck Hemoclip Traditional Ligating Clips, Rx Only, Sterile, Teleflex Medical, RTP, NC 27709 USA. The use of the reusable Teleflex Medical metal ligating clips is transient, invasive and intended for the application of hemostatic clips during general open surgery. Ligation clips are non-absorbable to thermostatically restrict flow of fluids within vessels. Clips are removed from the cartridge and applied in such a manner as to engulf vessels. Teleflex Medical metal ligating clips are intended for use in procedures involving vessels or anatomic structures for which the surgeon determines ligating clips are the best choice. Surgeons should select the size type, and material of the clip based upon their experience, judgment and needs. |  |  | August 22, 2008 | 2008 | Telefelx Medical | A hole in the sterile unit blister pack was detected that would compromise sterility. | Packaging process control | Consignees were notified via UPS next day on/about August 22, 2008. They were instructed to cease use and distribution and quarantine all affected product immediately. If they have any of the affected product they were instructed to return all it to Teleflex Medical via provided UPS-Ground. An Acknowledgement form was provided to be completed and returned via fax to Teleflex Medical to confirm receipt of the notification and amount of product on hand for return. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks from the second mailing. Consignees also may be contacted by phone, e-mail, or fax. A record of these contacts will be maintained. Upon return to Teleflex Medical, product will be quarantined, assessed for recall effectiveness, and either reworked or destroyed. Contact Telefex Medical at 1-800-334-9751 ext. 8066 for assistance. |  | 786036 | 786036 (total pieces) | Nationwide | Terminated | March 15, 2012 | 1301 | 786036
Z-0030-2009 | 49462 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Platinum Plus" LT , REF 46-601, UPN M001466010, LT/0.018/180, Sterilized with ethylene oxide gas, Rx only, Legal Manufacturer Boston Scientific Corporation, One Boston scientific Place, Natick, MA 01760-1537, Made in USA: 8600 NW 41st Street, Miami, FL 33166-6202. The Platinum Plus" Guidewire facilitates placement of a catheter during diagnostic or interventional intravascular procedures. The wire can be torqued to facilitate navigation through tortuous arteries and/or avoid unwanted side branches. |  |  | August 25, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific is initiating a voluntary recall of the Platinum Plus Guidewire (peripheral). BSC has received reports that product may be labeled as a 260 CM long guidewire when the actual packaged device is a 180 CM long guidewire and the product may be labeled as a 180 CM long guidewire when the actual packaged device is a 260 CM long guidewire. There is no expected potential for injury to | Labeling mix-ups | A "Boston Scientific Urgent Voluntary Medical Device Recall or Correction - Immediate Action Required" letter dated 8/25/08 was sent to each affected account (To The Hospital Field Action Contact). The letter includes the customer steps for recall, a complete listing of all product batches/lots being recalled, the Reply Verification Tracking Form(s), and a return shipping label. Contact Boston Scientific Corporation at 1-763-494-1700 for assistance. |  | 47 | 47 | USA: CAFLGAILINLAMNMONJNVNYOHPATNTXUTWAWI andWVOUS: France, Germany, Great Britain, Italy, and the Netherlands. | Terminated | December 10, 2011 | 1202 | 82
Z-0034-2009 | 49494 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 3-0 Black Mono Nylon Suture, 12 units per box. product number 925B. The product intended for use in general soft tissue approximation and/or ligation including use in ophthalmic procedures. |  |  | September 09, 2008 | 2008 | Surgical Specialties Corp | mislabeled product - labeled as 3-0 black mono nylon contains 4-0 black braided silk suture | Employee error | The recalling firm issued a Recall letter dated 9/9/08 to its customers. The letter informed the customers of the problem, the need to conduct a sub-recall, and return the product. Contact Surgical Specialties Corp. at 1-610-404-1000 for assistance. |  | 956 | 956 boxes | USA--AZ, CA, CT, FL, GA, IA, IL, KS, KY, LA, MD, MI, MS, NC, NJ, NY, OH, TN, TX, VA, and WA. OUS: Australia, Canada, and Fiji Islands. | Terminated | March 27, 2009 | 199 | 956
Z-0021-2009 | 49435 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VICRYL RAPIDE (polyglactin 910) Braided Coated Synthetic Absorbable Suture, Undyed, Non-USP. |  |  | August 11, 2008 | 2008 | Ethicon, Inc. | Package defect compromised the integrity of the primary seal, which could lead to premature suture degradation and/or impair the sterile barrier of the product. | Packaging | Letters were sent to distributors, Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Room on August 11, 2008, requesting that all affected lot numbers of Vicryl Rapide be discontinued immediately from sale and/or use. A separate correspondence was sent to hospitals. Letters to Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Rooms were advised of the recall and asked to discontinue use of any of the affected lot codes of sutures. Return and credit instructions were provided to the Director of Material Management. Contact Ethicon, Inc. at 1-908-218-2553 for assistance. |  | 6912 | 6912 | Worldwide: US and Canada OUS: Australia, New Zealand and Indonesia. | Terminated | June 05, 2012 | 1394 | 6912
Z-0014-2009 | 49022 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Mobetron Model 1000, mobile electron linear accelerator, a device system which delivers radiation to an intraoperative site. |  |  | June 20, 2008 | 2008 | Intraop Medical Corporation | Docking screw/nut failure, which may result in a treatment head dropping, potentially impacting the patient. | Component design/selection | Recall initiated on 06/20/2008, advising users via Advisory Notice that a safety upgrade would be available. Two options are given to either self inspect the product or schedule an Intraop Authorized Field Service Engineer to visit their facility to perform the inspection. The letter notes that if no nut failure was evident, the machine could still be used. The firm continues to send retrofit kits. |  | 24 | 24 units | Worldwide Distribution | Terminated | April 05, 2011 | 1019 | 24
Z-0013-2009 | 49003 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences***Fogarty Fortis***Arterial Embolectomy Catheter |  |  | June 23, 2008 | 2008 | Edwards Lifesciences, LLC | Potential for tubing fracture near the tip of the catheter. | Process control | The firm sent Urgent Field Safety Notices dated June 23, 2008 to customers. The letter gave details on affected products, a description of the problem, and advise on action to be taken. Edwards has decided to recall these lots of Fogarty Fortis catheters which were manufactured between 9 June 2007 and 20 June 2008. The expiration date for these lots is between 11 September 2009 and 26 September 2009. The attached sheet lists product by model and lot number which you have ordered during that time frame. Please stop using these lots of catheter products immediately. Edwards requests that you return all Fogarty Fortis catheters from these lots that are still in your inventory. Please call Edwards Customer Service to arrange for return of any unused product and obtain information about replacement product. After you have verified your inventory, please complete the attached Confirmation Form, even if you no longer have any affected units at your facility. This will allow us to verify the completion of this recall action. Please return the Confirmation Form by FAX to: +31.412.47.66.99 , Attention: Mark vanSchijndel Transmission of this Field Safety Notice: This notice needs to be passed on all those who need to be aware within your organization. Please transfer this notice to other organizations on which this action has an impact and where or to any organization where the potentially affected devices have been transferred. Please maintain awareness on this notice and resulting action for an appreciate period to ensure effectiveness of the corrective action. We sincerely regret the inconvenience caused by this action and greatly appreciate your immediate attention to this matter. If you have specific questions about this notification, please contact Customer Service |  | 2975 | 2,975 units | International Distribution --- including countries of Austria, Belgium, Bulgaria, Switzerland, Czech republic, Germany, Denmark, United Kingdom, Italy, Luxembourg, Netherlands, Sweden, Slovak Republic, South Africa, and Taiwan. | Terminated | September 19, 2012 | 1549 | 2975
Z-0032-2009 | 49465 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet brand Discovery Elbow, 3 mm x 75 mm right ulna with bearing/bond coat, TI 6AL 4V Alloy/ARCOM UHMWPE, For cemented use only, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 114813. The product intended use is for Orthopedic implant. |  |  | May 23, 2008 | 2008 | Biomet, Inc. | The component in the package is not the correct size. | Packaging process control | Consignees were notified by letter dated May 28, 2008 to locate and return the recalled lots. Contact Biomet Orthopedics at 1-800-348-9500, ext. 3983 for assistance. |  | 6 | 6 | USA--California, Illinois, Indiana, Massachusetts, Ohio Puerto Rico, and Canada, OUS: Germany and Great Britain. | Terminated | September 22, 2009 | 487 | 11
Z-0022-2009 | 49436 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec High Viscosity Injector, 4mm |  |  | September 15, 2008 | 2008 | Becton Dickinson and Company | Product labeled as 4 mm High Viscosity Injector tip contains a 6 mm tip. | Employee error | BD contacted accounts by telephone on 9/15/08 requesting users to discard units and complete the response form. Contact Becton Dickinson and Co. at 1-781-906-7950 for assistance. |  | 1485 | 1485 units | AZ, CA, CO, IL, OH, FL, PA , TN, TX,and VAForeign: Australia and Japan. | Terminated | February 29, 2012 | 1262 | 1485
Z-0020-2009 | 49379 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Aisys, Datex-Ohmeda, Inc., a General Electric Company going to market as GE Healthcare, GE Healthcare PO Box 7550, Madison, WI. 53707-7550 |  |  | August 28, 2008 | 2008 | Datex - Ohmeda, Inc | The electronic vaporization system in the GE Aisys Anesthesia machine contains the following components: Backpressure Valve, Inflow Check Valve and Cassette Interface Board, that can contribute to independent failures. | Manufacturing material removal | GE Healthcare "Urgent Medical Device Correction" letters dated 8/12/08 were sent on 8/20/08 individually to Health Administrator/Risk Manger, Head of Anesthesia and Director of Clinical Engineering. The letters described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The Safety Instructions state that the machines can be safely used as long as the following guidelines are followed:1. Carry out the daily system checkout for any system in use.2. The Aisys Anesthesia will alert the user both visually and audibly if it detects a fault condition.3. Have an alternative method of anesthesia available at all times.Once each machine has been corrected an "Update" will be provided by the service technician responsible for the correction. Each customer/machine will receive a "GE Healthcare Urgent Medical Device Correction Update" letter which informs the customer that this correction is a follow-up activity to the letter that was sent earlier. |  | 1571 | 1571 | Worldwide Distribution --- USA including states of AL, AK, CA, CO, FL, GA, IL, IN, KS, KY, ME , MD, MA, MI, MO, MT, NE, NV, NH, NY, NC, OH, OK, SC, SD, TX, VA, WV, WI, and DC, and countries of SOUTH AFRICA, TURKEY, THAILAND, SINGAPORE, SWEDEN, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MEXICO, LIBYA, KUWAIT, KOREA, JAPAN, JORDAN, ITALY, UNITED KINGDOM, FRANCE FINLAND, SPAIN, ESTONIA, DENMARK, GERMANY, COLOMBIA, CHINA, SWITZERLAND, CANADA, BRAZIL, BAHRAIN, BELGIUM, AUSTRALIA, AUSTRIA, and UNITED ARAB EMIRATES. | Terminated | March 08, 2016 | 2749 | 1571
Z-0028-2009 | 49445 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical, Lap Banding Kit, REF: K0406 Contains: (1) C0658, 11x100mm Threaded Shielded Trocar w/Universal Seal (1) C0522, 5x100mm Threaded Premium Flat Blade (1) C2202, 150mm Insufflation needle (1) C0605, 15x100mm Threaded Separator Access System (1) CB030, 5mm x 35cm Direct Drive Disposable Scissors (2) C4120, 5mm x 38cm Direct Drive Grasper Reposable Cartridge, for laparoscopic procedures. |  |  | June 25, 2008 | 2008 | Applied Medical Resources Corp | Potential inability to insufflate through the stopcock. Under certain circumstances, a flexible elastomeric component inside the 15mm trocar canstretch into a configuration that blocks the flow of insufflation gas (C02). | Process control | Customers were notified via FedEx or UPS on June 25,2008 with an overnight letter. The letter advises that Applied Medical is conducting a voluntary recall of the 15mm Separator Access Systems, Models C0604, C0605, C0606 and C0607 due to the potential inability to insufflate through the stopcock. Per the letter, the firm is asking that all 15mm product listed in the letter be returned. A response form is provided. |  | 14113 | 14,113 total | Nationwide and Canada distribution. | Terminated | March 26, 2012 | 1370 | 14113
Z-0088-2009 | 49490 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC. |  |  | June 19, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. | Software change control | An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008. |  | 779 | 779 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 21, 2009 | 336 | 779
Z-0087-2009 | 49442 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Medical Apollo Anesthesia Machine, catalog number 8606500 |  |  | August 22, 2008 | 2008 | Draeger Medical, Inc. | Sporadic errors in various device functions, including low readings of the single gas flows for oxygen, nitrous oxide, or air delivery, or mechanical ventilation failure. Alarms and error messages function as designed. | Nonconforming Material/Component | Draeger Medical sent a Medical Device Recall letter to their customers via certified mail on 8/22/08 informing them of the problem. Until the exchange of affected printed circuit boards has been implemented, the firm is recommending that users closely monitor all device displays and error messages and if applicable act in accordance with the advisory statements and remedial information provided in the Instructions for Use as well as making provisions for alternative patient care measures. |  | 214 | 214 units | Nationwide Distribution. | Terminated | September 01, 2009 | 375 | 214
Z-0086-2009 | 49411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Ultrasound, Acuson/Sonovista X300 ultrasound systems, software revisions Sonovista X300 software revision 1.0.06 and 1.0.07, Acuson X300 sofware revision 1.0.08 ultrasound system with onscreen display. |  |  | August 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Incorrect value calculations by the device may result in inaccurate aortic stenosis estimates. | Software design | Siemens initiated a Field Correction 08/19/2008. Service personnel will conduct a site visit and distribute a Customer Safety Advisory Notice, correct the problem and install a revised software version. |  | 169 | 169 units total | Worldwide Distribution: USA, Singapore, Germany, Denmark, Romania, Sweden, Poland, Portugal, Australia, Chile, India, Spain, Thailand, Malaysia, South Africa, Hungary, Saudi Arabia, Qatar, Peru, Italy, Azerbaijan, Greenland, Austria, Venezuela, Puerto Rico, Brazil and Japan. | Terminated | April 21, 2011 | 975 | 169
Z-0012-2009 | 48638 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Diagnostics NAPA2, N-Acetyl-Procainamide, COBAS INTEGRA, Roche Diagnostics, Indianapolis, IN.; 20737852322. |  |  | June 11, 2008 | 2008 | Roche Diagnostics Corp. | Some of the labels are glossy, resulting in the bar code being difficult to read by the bar code reader. | Labeling Change Control | Consignees were notified via an Urgent Medical Device correction letter dated 6/11/08 informing users of the problem and providing them with work-around instructions. The instructions suggests 1) inserting the cassette rack slowly; 2) using mainly reagent slots A, B, or C; 3) tilting the rack slightly clockwise so the barcode scanner does not read the barcode at an exact 90 degree angle; and 4) confirming the rack is properly seated on the track and displayed on the Cassette Status screen. If the cassette barcode does not register after following the Work-Around instructions, then customers are required to contact Diagnostics Technical Support. |  | 2 | 2 | Nationwide Distribution | Terminated | September 22, 2009 | 468 | 3624
Z-0015-2009 | 49107 | No | OHM | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Medline Latex Free C-Section CDS-LF and C-Section Supplemental CDS-LF; a single patient prescription procedure pack and disposables, including a Gerber Nuk Pacifier; Made in USA by Medline Industries, Inc., Mundelein, IL 60060 |  |  | August 04, 2008 | 2008 | Medline Industries Inc | The latex free labor &amp; delivery packs contain a latex Nuk pacifier. The pacifier is packaged separately and offers no risk of exposure of latex particles to other components of the pack. | Employee error | Recall letters dated 8/4/08 were sent to the direct accounts on the same date, informing them that the labor and delivery packs were labeled Latex Free, but contained a Latex Nuk pacifier. The letters were customized for each account, listing the specific affected reorder numbers and lot numbers shipped to each account. The accounts were requested to immediately check their stocks for the affected lot numbers and quarantine the affected product found. The accounts were given three options: A) Attach the enclosed stickers to the packs so that the user is aware of the latex Nuk pacifier, B) Call Medline to receive replacement non-latex pacifiers, but also place the stickers on the packs, C) Return the packs to Medline for replacements. The accounts were requested to complete the enclosed response form, indicating which option the account selected and whether they needed any additional stickers, and fax the form back to Medline at 847-643-4482 by 8/18/08. Any questions were directed to Kathy Dunne at 1-866-359-1704. |  | 579 | 579 cases | Nationwide Distribution --- including states of Alabama, Arizona, Georgia, Missouri, Nevada, New Jersey, and Washington. | Terminated | October 20, 2008 | 77 | 2115
Z-0101-2009 | 49417 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos TF, model number 10093902 |  |  | August 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Liquids may enter the system and cause potential malfunction and possible hazard to patients, user or other persons. | Device Design | The recalling firm was made aware of a potential problem with the device from the manufacturer in Germany. The recalling firm sent a copy of the Customer Safety Advisory Notice, dated June 18, 2008, received from Germany to their customers. The recalling firm has representatives visiting each location to modify the units. |  | 211 | 211 units | Nationwide Distribution | Terminated | October 05, 2009 | 419 | 211
Z-0036-2009 | 49405 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences LifeStent FlexStar Self-Expanding Biliary Stent System, Edwards Lifesciences LLC Irvine, CA 92614, Model EX060203CD. The LifeStent FlexStar self-expanding biliary stent system is intended for use in the palliation of malignant strictures (neoplasms) in the biliary tree. |  |  | August 29, 2008 | 2008 | Bard Peripheral Vascular Inc | Some LifeStent FlexStar Systems may exhibit a gap between the tip of the delivery system and the primary sheath such that the guidewire lumen is visible. | Process control | The customer letters for the tip to sheath gap were sent on 8/29/2008 with return receipt. The letter advises that C.R. Bard has determined that some LifeStent FlexStar Systems may exhibit a gap between the tip ofthe delivery system and the primary sheath such that the guidewire lumen is visible. Per the letter, the firm's analysis has determined that a gap between the tip of the delivery system and the delivery system sheath may be associated with difficult advancement to the treatment site. Users are instructed to inspect all devices to verify that there is not a gap between the tip and sheath (orange colored guidewire lumen is visible), prior to use and to not use any devices that exhibit a tip to sheath gap condition. If devices that exhibit a tip to sheath gap are identified, customers are advised to call Customer Service at 1-800-321-4254 to return the device and a replacement will be provided at no cost. The letter also advises that the Instructions for Use for the LifeStent FlexStar System will be updated to include the inspection step which is provided in the letter. |  | 92 | 92 units | Nationwide | Terminated | April 06, 2012 | 1316 | 9889
Z-0111-2009 | 49412 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Ultrasound, Acuson/Sonovista X300 ultrasound systems, software revisions 2.0.1 to 2.0.05, 3.0.01 and 3.0.02, manufactured by Siemens Ultrasound |  |  | August 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Thermal Index cranial (TIC) is not displayed for the Neo-Head exam type with the C8-5 transducer. This is a required display to support the alara principle when imaging cranial structure. | Employee error | Field correction was initiated on 8/19/2008. Under Siemens update program US020/08/S &amp; US021/08/S Siemens service personnel make a site visit, deliver a customer safety advisory notice to the customer and correct the problem by installing the revised Software versions. After the planned field correction is performed at the customer site the Service personnel log that information through Siemens internal web or reporting tools into the Siemens field update database. |  | 72 | 72 units | Worldwide Distribution --- including USA and countries of France, Colombia, Poland, Taiwan, Saudi Arabia, Austria, Malaysia, India, Czech Republic, Croatia, Argentina, Venezuela, Italy, Philippines, Mauritius, Spain, Bolivia, Germany, Ukrain, Russia, United Kingdom, Brazil, South Africa and Canada. | Terminated | December 20, 2010 | 853 | 72
Z-0100-2009 | 49402 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | M3811B Philips Telemonitoring Clinical Review Software part of Philips Telemonitoring System, M3810A, software revision identified as Build 1.1.2.11d, B.02.07. Prescriptive medical device used to automatically collect and transmit medical information (weight, blood pressure, non-diagnostic ECG) over phone lines between provider and patient. |  |  | September 11, 2008 | 2008 | Philips Medical Systems | Multiprint report may contain incorrect vital data for patients | Software design | Philips contacted customers by sending an Urgent Medical Device Correction Notification on 9/11/08. Letters to customers sent via certified mail and receipt of delivery will be tracked.Customers are asked to follow the Action To Be Taken By Customer/User section of the Urgent Medical Device Correction Notification. This section instructs users to allow the upgrade of the software upon logging in to the system. Contact Philips at 1-866-246-7316 for assistance. |  | 52 | 52 units | Nationwide | Completed |  |  | 52
Z-0105-2009 | 49061 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific FilterWire EZ" Embolic Protection System, Model FW EZ 190cm, 3.5-5.5 mm, US, Material H749201001900, Catalog # 20100-190, Sterilized with irradiation, Made in USA: 2011 Stierlin Court, Mountain View, CA 94043-4655. Filter wire EZ" Embolic Protection System is indicated for use as a guide wire and embolic protection system to contain and remove embolic material (thrombus/debris) while performing angioplasty and stenting procedures in coronary saphenous vein bypass grafts and carotid arteries. The diameter of the vessel at the site of filter loop placement should be between 2.25 mm and 5.5 mm for coronary saphenous vein bypass graft procedures and between 3.5 mm and 5.5 mm for carotid procedures. |  |  | May 20, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific is initiating a field correction for 48 lots/batches of the Filter Wire EZ" Embolic Protection System. Boston Scientific has determined that the Directions for Use (DFU), which accompanies each packaged device, may be missing. If the product does not include a DFU and the user is unable to find a copy, the user will have to seek another product box containing the DFU. This may | Employee error | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required" letter dated May 20, 2008. The letter addressed to Risk Mangers, described the issue and product. A Reply Verification Tracking Form was requested to be completed and returned. Contact Boston Scientific Corporation at 1-800-272-3737 for assistance. |  | 4338 | A total of 4338 units distributed. | US: AL AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD. TN, TX. UT, VA, VT, WV, WA, WI, WY, and PR.European Countries: Austria, Belgium, Cyprus, Czech Republic, Denmark, Finland, France, Great Britain, Germany, Greece, Hungary, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, Spain, Sweden, and Switzerland, Intercontinental: Taiwan, Korea, Australia, China, Singapore, Canada, Columbia, and Mexico. | Terminated | December 10, 2011 | 1299 | 4338
Z-0114-2009 | 49416 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety Lancet for capillary blood sampling, 1.6 mm x 28G, Mini, Sarstedt. (Blood lancet) |  |  | August 27, 2008 | 2008 | Sarstedt Inc | The product from this lot may not retract the lancet into the safe position inside the lancet body after the trigger is actuated. | Process control | The consignee was notified via letter that was hand delivered on 08/27/2008. All of the distributed product was picked up from the consignee and returned to Sarstedt, Inc. None of the product had been used. All action was completed on 8/27/2008. Contact Sarstedt at 1-828-465-4000 for assistance. |  | 1000 | 1000 pieces | NC | Terminated | April 24, 2012 | 1336 | 1000
Z-0113-2009 | 49518 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Zilver Expandable Metal Biliary Stent System, ZILBS-10-6, Stent Diameter: 10mm, Stent Length: 6cm, Introduction System: 7 FR., Disposable-Single Use Only, Rx Only, Cook GI Endoscopy, Winston-Salem, NC. The device is used in palliation neoplasms in the biliary tree. |  |  | September 11, 2008 | 2008 | Cook Endoscopy | A section of the introduction system may detach after the stent has deployed. | Process control | The consignee was notified by letter on/about 09/11/2008. The distributor was instructed to contact their customers and coordinate the return of all affected product to Cook Endoscopy. A responds form was attached to the letter to be completed and returned to Cook Medical to confirm receipt of notification and indicate quantity to be returned. Contact Cook Endoscopy at 1-336-744-0157, ext. 6325 for assistance. |  | 12 | 12 | USA: IL and OUS: China | Terminated | April 01, 2009 | 202 | 12
Z-0102-2009 | 49496 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging with versions VB20B, VB20D, VB20F, VB20G, VB25B, VB30A, VB30A_SP1, and VB30A_SP2. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Image may not be visible after merge. | Software design | The recalling firm issued Customer Safety Advisory Notice dated 8/22/08 to inform the consignees of the potential issue and provide instruction to avoid its occurrence. To avoid potential risk of this issue, users are asked to check whether the extended HL7 messages as stated in the letter are mandatory for their RIS-PACS workflow, and disable these messages if they are not. Siemens is preparing a modification under the UI IM086/08/S for the affected versions of syngo Imaging that will resolve the potential malfunction. The field modification will be immediately distributed and installed by the firm's Service as soon as it is available. |  | 59 | 59 units | Nationwide Distribution --- including states of AL, AZ, CA, CO, CT, FL, IA, KS, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, TN, TX, and WI. | Terminated | November 19, 2008 | 89 | 59
Z-0112-2009 | 49459 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal (formerly Quantitative Sentinel) System - Fluid Total Precision software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | On the I&O; chart, the IN, OUT and NET fluid totals values will not honor the numeric precision configuration, always displaying a whole number, losing decimal point accuracy. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 4/20/07 of the problem with the Fluid I &amp; O Module when fluid precision is required. The letter instructed those accounts that require fluid precision to contact GE Technical Support at 1-800-433-2009 to obtain a software patch to fix the problem. A follow-up letter dated 8/29/08 was sent on the same date to the accounts informing them that GE had re-classified the issue as a class II safety issue and that all customers need to have their software updated. The accounts were requested to contact GE Tech Support at 800-433-2009 to schedule a time to upgrade their system to 6.70.6 P02 as soon as possible. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 847-277-5240. |  | 470 | 470 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 17, 2010 | 599 | 470
Z-0092-2009 | 48432 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMEVIEW 3i System equipped with a Coherence Therapist R2.0 or R2.1, and PRIMEVIEW 3iR2.0 or R2.1, part number 7341428 |  |  | November 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Flat panel positioning calibration could be off by as much as 4 mm without the machine discovering detail. | Software design | Service calibration inspection/correction Update instructions TH038/07/S was issued on 11/06/2007 to all customers. The instructions identified the affected products, stated the reason for the update, and listed the steps to update the software. Customer notification was also issued. Software patch planned for release in June 2008. |  | 13 | 13 units | Worldwide Distribution | Terminated | December 18, 2010 | 1138 | 251
Z-0161-2009 | 49545 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Accessory Head Rest for Alphastar Operating Tables |  |  | September 26, 2008 | 2008 | MAQUET Inc. | During installation or cleaning there is a risk of crushing the finger by handling the head rest in the area of the gas strut while simultaneously actuating the release lever. | Other | Urgent Device Correction letters were sent beginning September 26, 2008 via Fed Ex. Letters identified the potential for finger crush in conjunction with the head rest. Proper operating instructions are provided as well as a warning sticker to be applied to the head rest indicating proper positioning. All letters included a customer response form to be completed by each institution. Information requested is: that device correction letter has been read and instructions followed; that the corrected operating instructions have been received; that the warning sticker has been properly placed on the headrest, or that they do not have the head rest identified in the device correction letter. Questions should be addressed to USA Maquet Inc at 1-888-627-8383 #1. |  | 1972 | 88 units in US; 1884 units worldwide | Nationwide Distribution | Terminated | April 16, 2009 | 202 | 1972
Z-0116-2009 | 49024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. |  |  | July 04, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images. | Software change control | A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance. |  | 16 | 16 | Nationwide in LA, WI, DE, PA and OH, and internationally in Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium and France | Terminated | December 22, 2010 | 901 | 16
Z-0115-2009 | 48995 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Series 7000 Tibial Impactor/Extractor, Catalog number 3770-000 |  |  | July 02, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is a potential inability to release the Scorpio Series 7000 Tibial Impactor/Extractor from a tibial baseplate during use. | Component design/selection | US recall notification (product correction) letters were send out by Federal Express with return receipt on July 2 and 9, 2008.Advisory notices were sent to foreign customers on July 2, 2008.Both letters include a Product Correction Bulletin that instructs all used to lubricate the divide with an approved medical lubricant, especially the extractor shaft and threads. It is recommended that all moving parts be evaluated for functionality and useful service prior to each surgery and instructions on how to perform this process. Contact Stryker Howmedica Osteonics Corp. at 1-201-831-5918 for assistance. |  | 5126 | 5126 | Worldwide including the US and Asia Pacific, Canada, Europe, Middle East, Africa, Japan and Latin America. | Terminated | August 17, 2012 | 1507 | 5126
Z-0117-2009 | 49399 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Exact Couch with Exact Couch Top, a component of the C-Series Clinacs, Acuity and Ximatron treatment systems. Model numbers H14, H18, H27, H72, H77. A powered radiation therapy patient support assembly is an electrically powered adjustable couch intended to support a patient during radiation therapy. |  |  | August 11, 2008 | 2008 | Varian Medical Systems Oncology Systems | During couch movement, a patient's fingers may get pinched at certain points which could result in broken, disengaged or pinched fingers. | Device Design | Recall initiated on 08/11/2008. Notification will be sent to Varian Sales, Marketing and Service organizations as well as to all affected users. Contact Varian Medical Systems Oncology Systems, at 1-650-424-5731 for assistance. |  | 3296 | 3296 units | Worldwide including to the US and to Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania, Luxembourg, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, The Netherlands, United Kingdom, Algeria, Australia, Bahrain, Bangladesh, Belorussia, Bolivia, Brazil, Canada, Chile, China, Colombia, Ecuador, Egypt, Guadeloupe, Hong Kong, India, Indonesia, Iran, Israel, Jamaica, Japan, Kazakhstan, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mexico, Morocco, Nepal, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Russian Fed., Saudi Arabia, Serbia, Montenegro, Singapore, South Africa, South Korea, Sudan, Syria, Taiwan, R.O.C., Thailand, Turkey, U.A.E, Ukraine, Uruguay, Venezuela, and Vietnam. | Terminated | November 27, 2010 | 838 | 3296
Z-0151-2009 | 49439 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ray TFC Device with End Caps: |  |  | August 18, 2008 | 2008 | Stryker Spine | The label on one side on the Ray TFC Device (Ray Cage) is incorrect. It reads 20 mm x 26 mm instead of 20 mm x 21 mm. | Labeling mix-ups | Urgent Product Recall notification letters were sent via Fed Express on August 18, 2008 to Stryker Branches and Hospital Risk Managers. The firm asked customers to examine their inventory and hospital locations to identify the product; reconcile all products on the Product Accountability Form and attached distribution sheet; and fax a copy of the Accountability Form and distribution sheet to Sarah Thompson within 5 days of receipt of the notice. All questions are to be directed to Susan Krasny at 201-760-8150. |  | 69 | 69 units | Nationwide Distribution | Terminated | November 04, 2008 | 78 | 69
Z-0119-2009 | 49407 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Sonesta 6210 Fluoroscopy Procedure Table; a motorized procedure table with adjustments (via a remote control) for height, tilt, back and seat cushion position; Stille Surgical AB, Sundbybergsvagen 1A, SE-17173 Solna, Sweden |  |  | August 27, 2008 | 2008 | Stille AB | The table contains a battery backup, not mentioned in the User Manual, which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However, if a short circuit occurs, it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply, as the battery automatically w | Component design/selection | Stille sent Field Safety Notices dated 8/27/08 to the direct accounts, informing them that the table contains a battery backup, not mentioned in the User Manual, which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However, if a short circuit occurs, it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply, as the battery automatically will supply continuous power to the control box. The accounts were requested to complete and return the enclosed acknowledgement card to Stille by fax or e-mail as soon as possible and no later than 9/26/08. The accounts were advised to not further use the table in conjunction with a C-arm or any other heavy object until a service technician can permanently remove the battery from the control box according to the enclosed Service Manual Amendment No. 1. Stille will make a permanent correction to the table once it has been validated and approved. Contact Stille Surgical Inc. at 1-800-655-1614 for assistance. |  | 114 | 114 units | Nationwide and Canada | Terminated | December 10, 2010 | 835 | 114
Z-0163-2009 | 49567 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX Multipurpose Radiography System, stationary X-Ray system, Model number: 7414803. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Tube support arm or the detecting mounting hardware may become loose at the joint with the telescope. | Equipment maintenance | Updated Instructions were issued in August 2008 (AX049/08/S) to Siemens Service Engineers to have them inspect and tighten any loose hardware on the affected systems. |  | 7 | 7 units | Nationwide distribution including states of CA, IN, KY, UT, VA and WI. | Terminated | October 20, 2008 | 56 | 7
Z-0152-2009 | 49443 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Auto Syringe AS50 Infusion Pump, product code 1M8550; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A. |  |  | September 05, 2008 | 2008 | Baxter Healthcare Corp. | Non-conforming Electrostatic Discharge (ESD) grounding squares may cause the pump to be susceptible to shorting-out of circuitry, resulting in a loss of audio, and/or interruption of therapy. | Nonconforming Material/Component | Urgent Product Recall letters dated 9/5/08 were sent on the same date via first class mail to the direct accounts, to the attention of the Director of Biomedical Engineering and Facility Risk Manager. If the account had device(s) that had been serviced by Baxter with a part affected by this action, a serial number list of the affected device(s) was provided with the notification. Customers were instructed to remove the affected device(s)/spare parts from service and contact Baxter to arrange for their return and repair. Alternatively, if removal of the device is not a viable alternative for the customer and they wish to perform self-service, the customers will also be provided with a Service Bulletin to provide instruction on replacement of the rear case assembly. |  | 509 | 509 units | Worldwide Distribution --- USA including Puerto Rico, and countries of Canada and Korea. | Terminated | February 03, 2010 | 516 | 509
Z-0103-2009 | 47839 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Zyoptix XP Microkeratome maxon motor, Ref #ZXP09183. The label shows the product is manufactured by Bausch &amp; Lomb Incorporated, Rochester, NY. |  |  | April 23, 2008 | 2008 | Bausch &amp; Lomb Inc | The plastic tip on the drive shaft may come off the shaft and cause the blade to stop oscillating. | Device Design | Phone calls were initiated on 4/28/08 and an Urgent Recall letter dated 4/23/08 was issued 4/28/08 via overnight mail to U.S. customers explaining the reason for recall. The letter also requested that customers stop using the motors identified and return of the motors for evaluation. A return form was enclosed. |  | 69 | 69 motors | Worldwide Distribution --- including USA and countries of Brazil, The Netherlands, Canada, India, Malaysia, Mexico, and Singapore. | Terminated | April 02, 2012 | 1440 | 113
Z-0164-2009 | 48661 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian RV Software |  |  | June 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | A malfunction within the Varis software might result in a misadministration (underdose). The software does not correctly interpret the number, causing field parameters having decimal values to be incorrect. | Software design | Recall initiated on June 2, 2008. A product notification letter was distributed by certified mail to all affected consignees. |  | 18 | 18 | Worldwide Distribution --- including USA and countries of Canada, Austria, France, Germany, South Korea, Sweden, and the UK. | Terminated | May 03, 2011 | 1065 | 18
Z-0165-2009 | 48711 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular E Module immunoassay analyzer, Roche Diagnostics, Ind., IN; GMMI Nos. 04998642001 and 03617505001. |  |  | June 24, 2008 | 2008 | Roche Diagnostics Corp. | A software bug may result in pipetting from an incorrect reagent pack and/or assigning calibration curve parameters incorrectly. | Other | The consignees were notified of the problem via letter dated 6/24/08 and were provided instructions for a workaround. |  |  |  | Nationwide Distribution --- including states of Florida, New Jersey and North Carolina. | Terminated | October 22, 2009 | 485
Z-0170-2009 | 49128 | No | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Alpha-Omega Services, Inc. Product No:NFW0001-010 |  |  | November 15, 2007 | 2007 | Alpha Omega Services Inc | There is a possibility that the Flexiguide Needle tip is defective. There is the potential for the needles tip to leak, thereby breaking the sterile boundary and possibly contaminating other devices with body fluids. | Nonconforming Material/Component | Notification of customers initially began November 15, 2007 based on a cursory visual inspection of the defective needle tips. The initial letter advised that there is a possibility that the Flexiguide Needle tip is defective in that there is the potential for the needle's stylet or an HDR source wire/cable to puncture through the defective needle tip during implant or treatment and penetrate the surrounding tissue thereby injuring the patient and/or breaking the sterile boundary. Further distribution and/or use of these products should not take place. The recall notification was updated on November 27,2007 once results from actual testing were available. The letter advises that there is a potential for the needle's tip to leak, thereby breaking the sterile boundary and possibly contaminating other devices with body fluids. Further distribution and/or use of these products should not take place. Customers are to return the products for credit. A response form is provided. Contact Alpha-Omega Services, Inc. at 1-800-346-7894 for assistance. |  | 40 | 40 pkgs (10 each pkg) | NV, AZ, CA, MI, UT, MO, CO, and OH | Terminated | December 09, 2008 | 390 | 198
Z-0171-2009 | 49131 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid AST (SGOT) Reagent Set, Manufactured by Pointe Scientific, Inc., Canton, MI in 150 ml (Product no. A7561-150) and 450 ml (Product no A7561-450). For the quantitative determination of aspartate amoinotransferase in human serum |  |  | August 25, 2008 | 2008 | Pointe Scientific, Inc. | The R1 component may be contaminated with Serratia liquefaciens, resulting decreased absorbances and failure of the reagent to produce test results. | Material/Component Contamination | Consignees were notified via recall letter dated 8/22/08 and mailed 8/25/08 to examine their inventory, dispose of any recalled product found, and to notify the firm for replacements. Contact Pointe Scientific Inc., at 1-800-757-5313 for assistance. |  | 109 | 109 kits | Nationwide, Bolivia, Ecuador, Greece, Indonesia, Mexico, Nepal, Philippines, Romania and Serbia. | Terminated | November 27, 2009 | 459 | 134
Z-0183-2009 | 49540 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Vacora Biopsy Vacuum Assisted Biopsy Probe; Catalog numbers VB10118 and VB10140. |  |  | September 12, 2008 | 2008 | Bard Peripheral Vascular Inc | Some of the Vacora Biopsy Vacuum Assisted Biopsy Probe thumb wheels from these lots may fracture when fired using the prime/pierce option. | Device Design | Product Recall Notification Letter sent dated September 11, 2008 Attention: Risk Managers "All VB10118 and VB10140 catalog numbers with expiry dates prior to September 2008 are now beyond their expiration date. In accordance with accepted medical practice, they should have been discarded and not used. Do not use or further distribute any affected products. Once the product affected by the recall has been removed from your inventory; fill out the recall &amp; effectiveness check form, call Silvia Carrillo at 1-800-321-4254 ext 2727 to issue your with either a Return Authorization Number or Consignment Recall Number. fax the form to BVP at 1-800-440-5376 |  | 225 | 225 | Nationwide Distribution --- including states of AL, AR, CA, FL, KY, MA, NY, OH, SC, TN, TX, VA, and WA. | Terminated | September 19, 2012 | 1468 | 225
Z-0182-2009 | 49514 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Circlip component of the Stryker Flat Panel and Navigation Arm System. Stryker Flat Panel and Navigation Arm System are intended for use as a ceiling-mounted device to support a medical grade flat panel monitor or camera in the patient area. The Circlip within the flat panel and navigation arms is part of the mechanical stop designed to prevent the arm from rotating infinitely. It also prevents the cables routed through the arm from becoming twisted during routine movement and rotation of the arm. |  |  | September 05, 2008 | 2008 | Stryker Communications Corp | Circlip component used to suspend flat panel and navigation arm system may become dislodged. | Device Design | Notification letter was sent to consignees September 5, 2008. Field service will be performed at all affected sites. Old Circlips will be replaced with new clamps that do not allow the arm to fall. Contact Stryker Communications Corp. at 1-972-410-7310 for assistance. |  | 705 | 705 units | Nationwide and International Distribution | Terminated | October 02, 2009 | 392 | 705
Z-0181-2009 | 49464 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Inserter, 3/8" thread, Biomet Orthopedics, Inc., Warsaw, IN; Model 31-400600. Non-powered, hand-held instrument intended for use as an impactor during orthopedic surgery. |  |  | July 28, 2008 | 2008 | Biomet, Inc. | The pin and clip may fracture during use. | Device Design | Consignees were notified by recall letter dated 7/28/08 to remove the recalled lots from use and to return them. Contact Biomet, Inc. at 1-800-348-9500 or 1- 574-372-3983 for assistance. |  | 85 | 85 | Nationwide, Australia, Belgium, Canada, Denmark, New Zealand and Sweden. | Terminated | September 22, 2009 | 421 | 85
Z-0173-2009 | 49393 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IDS 25-Hydroxy Vitamin D EIA Enzymeimmunoassay for the quantitative determination of 25-hydroxyvitamin D and other hydroxylated metabolites in serum or plasma. Ref: AC-57F1 Lot 60897 Exp 2009-08; Immunodiagnostic Systems Ltd, Unit 10, Boldon Business Park, Boldon, Tyne &amp; Wear, NE35. 9PD, UK. Enzymeimmunoassay for the quantitative determination of 25-hydroxyvitamin D and other hydroxylated metabolites in serum or plasma |  |  | August 02, 2008 | 2008 | Immunodiagnostics Systems Inc | A small number of customers have experienced low absorbance values when using kits from this batch. | Nonconforming Material/Component | All customers who received 25-Hydroxy Vitamin D EIA lot # 60897 were contacted by mail. Customers were instructed to sign a confirming document for having received the notification letter and return via fax. Customers are requested to contact Immunodiagnostic System Inc's office if they have any kits of 60897 remaining in stock. Rather than replacing only the biotin buffer as stated in the letter dated August 2, 2008, Immunodiagnostic Systems Inc. will replace the kit per the letter dated August 8, 2008. Contact Immunodiagnostic System Inc. at 1-480-836-7435 for assistance. |  | 691 | 691 kits | Nationwide | Terminated | August 13, 2012 | 1472 | 691
Z-0196-2009 | 49549 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 26" Vision Elect HDTV Surgical Viewing Monitor, model number 0240030960, Manufactured by Stryker Endoscopy San Jose |  |  | September 22, 2008 | 2008 | Stryker Endoscopy | Knob parts stick together, resulting in poor functionality of the knob. | Device Design | Urgent: Device Correction letters dated 9/22/2008 were sent to Sales representatives and all accounts. Sales reps were also notified via email. Each Sales Rep was provided rework instructions, a list of affected accounts and serial numbers, and new knobs to install on affected units. Customers were asked to inspect their inventory for the part number, they were informed that their sales rep would be responsible for replacing the rotary knobs on all affected units, and that they should contact their sales rep to schedule the knob replacement. |  | 2265 | 2265 devices | Worldwide Distribution --- USA including states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Montana, North Carolina, North Dakota, Nebraska, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and countries of Australia, Canada, China, England, France, Germany, Greece, Iberia, India, Israel, Japan, Korea, Mexico, Netherlands, Poland, Portugal, Scandinavia, Singapore, South Africa, Switzerland, and Taiwan. | Terminated | December 21, 2010 | 820 | 2265
Z-0199-2009 | 49565 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smallbore T-Port Extension Set with Female Luer Lock Connector and Male Luer Lock T-Port with Ultrasite Injection Site 5 in. The product is shipped 100 units per carton. |  |  | September 17, 2008 | 2008 | B. Braun Medical, Inc. | Product was assembled incorrectly and connected to the wrong part. | Employee error | The recalling firm issued an Urgent - Medical Device Recall letter dated 9/17/08 to the distributors informing them of the problem and the need to return the product. The letter states that further distribution of the product should be immediately discontinued. Utilizing the "Product Removal Acknowledgement" form, customers should combine their inventory with the inventories reported by their customers and record them on the form. If affected products and associated lot numbers are currently in their possession, customers should call Customer Support Department to arrange for shipping and replacement product. A Customer Support Rep will provide instructions for handling affected product. |  | 3900 | 3900 units | Nationwide Distribution --- including states of CA, FL, NC, NJ, and TX. | Terminated | May 05, 2009 | 230 | 3900
Z-0186-2009 | 49446 | No | MHY |  |  | Neurology | PMA | Medtronic DBS" 3387 Lead Kit for Deep brain stimulation, DBS Leads: Proximal 40 mm, Distal 12.0 mm, Medtronic, Inc. |  |  | June 14, 2008 | 2008 | Medtronic Neuromodulation | Medtronic has received reports of DBS" leads being damaged at the proximal connector end of the lead when the lead cap is used in the implant procedure. The lead damage is the result of excessive torsion and / or tensile lading of the proximal connector end of the lead. Data also suggest that excessive torque while tightening or loosening the setscrews may twist the setscrew connector block and d | Process design | Consignees were sent a "Medtronic Urgent: Medical Device Correction" letter 8/2008. The letter addressed to Healthcare Professionals, described the product involved, explanation of the issue, recommendations for Implant and included an attachment "Updated Instructions for Using the Lead Cap". Contact Medtronic Neuromodulation at 1-800-707-0933 for assistance. |  | 3939 | 3939 | AL, AZ, AR, CA, CO, CT, FL, DC, GA, HI, ID, IL , IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE , NV, NH, NJ, NM, NY, NC, ND, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT , VT, VA, WA, WV, and WI.ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROATIA, CYPRESS, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, LEBANON, LUXEMBOURG, MEXICO, MOROCCO, NETHERLANDS, NORWAY, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | December 31, 2011 | 1295 | 26313
Z-0174-2009 | 49360 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Premier Value Lancing Device. UPC Code: 840986016707, The Package Includes: 1 Lancing Device, 1 Clear Cap and 5 Premier Value Ultra Thin Lancets, Distributed by: Chain Drug Consortium, LLC, Boca Raton, FL 33431. |  |  | July 18, 2008 | 2008 | Facet Technologies | The Alternate Site Cap was not included in the package as indicated on the labeling. | Error in labeling | The consignees were notified by letter on/about 7/18/2008. The Return Material Authorization number was communicated to the customer. They were asked to place the items on hold and Can Am will obtain any quantities. The Returned Product Mechanics Process will: Inventory will be sent back to Facet in two stages, Inventory on hand at Cimetra will be sent back to Facet, Inventory received back from customers will be sent back to Facet in one shipment 30-60 days after notice. Inventory received back after this second shipment will be destroyed. |  | 2976 | 2976 units | Nationwide--GA and FL. | Terminated | August 24, 2009 | 402 | 2976
Z-0189-2009 | 49456 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T6 System; a dual-detector variable angle gamma camera with a six slice CT scanner; Siemens Medical Solutions USA, Inc., Molecular imaging, Hoffman Estates, IL 60192. The system is used to perform CT scans and nuclear imaging studies with the same instrument, obtaining attenuation corrected images and providing registration of anatomical and physiological images within the patient's anatomy. The Symbia T6 System is used for advanced oncology, neurology and cardiology applications. |  |  | September 29, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | The protective plastic cap over the CT gantry power switch on the Line Connection Box may come loose, exposing the energized electrical contacts within the switch, thereby causing an electric hazard. | Nonconforming Material/Component | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 9/29/08 to all users of the Symbia T6 and Symbia T16 Systems. The accounts were informed of the potential risk that the protective plastic cap over the CT gantry power switch could come loose, causing an electrical hazard. The customers were advised to avoid contact with the gantry power switch until the protective plastic cap has been replaced, and to leave the power switch on at the line connection box and cycle the CT gantry on and off by toggling the CT gantry circuit breaker within the circuit box. They also were advised that a Siemens Customer Service Representative would contact them within 90 days to schedule a corrective update to their Symbia system free of charge. The accounts were requested to complete and return the enclosed mandatory fax-back form to 1-865-218-3019, acknowledging the receipt and understanding of the 9/29/08 letter and providing their Symbia System model number and serial number. Contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 for assistance. |  | 11 | 11 units | Nationwide to the District of Columbia, Florida, Georgia, Illinois, Indiana, Missouri, Ohio, and Texas, and internationally to Australia, Brazil, Canada, Great Britain, India, Italy, Japan, Norway and Switzerland. | Terminated | December 09, 2009 | 436 | 21
Z-0162-2009 | 49566 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GlobTek, Inc. Power Supply for Medical Use Part No: TR9CE4000LCP-Y-MED MODEL: GTM21097-5012, Made in China. |  |  | August 07, 2008 | 2008 | Smith And Nephew, Inc. Wound Management Division | External battery chargers used with the Version 29 VISTA Negative Pressure Wound Therapy pumps are failing to properly charge the pump's battery. | Nonconforming Material/Component | Each consignee was notified by recall letter (Field Correction Notification) dated 8/7/08 via FEDEX overnight. Within five (5) working days of the date of the customer notification a Smith &amp; Nephew representative will contact the customers to assess their inventory and schedule delivery of replacement battery chargers. Contact Smith &amp; Nephew, Inc. (Customer Care Center) at 1-800-876-1261 for for assistance. |  | 2550 | 2,550 | Nationwide within U.S. and internationally to United Kingdom, New Zealand, Germany, Canada and Australia. | Terminated | October 01, 2009 | 420 | 2550
Z-0200-2009 | 49581 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Monitor suspensions used with Advantx Legacy, RFX Classical, SFX-90, Model # 46-240838G2, manufactured after March 8, 1985 for Advantx Legacy, RFX and SFX systems. |  |  | March 09, 2007 | 2007 | GE Healthcare | GE Healthcare has recently become aware of a potential risk of monitor dropping associated with the monitor suspension of the Advantx system that may impact patient safety. It has been reported that suspensions manufactured after March 8, 1995 did not have the required thread locking agent applied to the setscrew, allowing the setscrew to back out over time, which can lead to the monitor dropping | Component change control | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated September 13, 2007. The letter was addressed to Hospital Administrators/Managers of Radiology/Radiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information. The letter recommends that customers move the monitor suspension to an open area in the room where viewing is still possible and make sure to minimize the movement of the monitor until a Qualified GE Healthcare Field Engineer has inspected the suspension. The letter also states that a Qualified GE Healthcare Field Engineer will schedule a field visit to inspect the monitor suspension. Customers should contact their sales or service rep is they have any questions. |  | 21 | 21 | Worldwide Distribution --- USA including states of AZ, CA, FL, GA, IA, MD, MO, NC, NY, OH, OK, TN, TX,WI, and LA, and country of Slovak Republic. | Terminated | August 24, 2012 | 1995 | 21
Z-0202-2009 | 49578 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000 Emergency and Transport Ventilator Continuous Respirator with software version 1.10 pre-installed; Catalog number 2M86300. |  |  | October 03, 2008 | 2008 | Draeger Medical, Inc. | May experience an interruption of ventilation for approximately 5 seconds | Software design | The firm issued an Urgent - Medical Device Recall letter dated September 2008. Software version 1.10 may under specific conditions experience an interruption of ventilation for approximately 5 seconds. During that time period, the patient airway pressure may drop to zero. This phenomenon was identified by Draeger engineers and to date has not been reported by any user. The firm's investigations revealed that when ventilating in BIPAP mode with the pressure rise time (Ramp) set to the highest setting and the expiration time set to exactly 0.5 seconds, the inspiration phase can coincide with the internal auto zero function of the device. Under these circumstances a pressure spike will occur upon reaching Pmax +5mbar (cmH2O) which with software version 1.10 will lead to the inspiratory phase stopping. Ventilation will resume after approximately 5 seconds. The device will generate a Paw Measurement INOP alarm. Until the new software is available, customers are being instructed to avoid the use of the following combination of settings: 1) BIPAP mode; 2) Ramp = highest, and 3) Texp = 0.5 seconds. |  | 50 | 50 units | Nationwide Distribution --- including states of AZ, CA, FL, IL, IN, KY, LA, ME, MI, MO, NC, NE, NJ, NY, OH, PA, TX, UT, VT, WA, WV, and WY. | Terminated | October 16, 2009 | 378 | 50
Z-0153-2009 | 49444 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Allez Spine Laguna Pedicle Screw System; Catalog Number: IT-FN1003 |  |  | August 21, 2008 | 2008 | Allez Spine, LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food &amp; Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA. | Pending | Allez Spine issued a field notification beginning August 21, 2008 via FedEx to all its hospital users, distributors and field representatives informing them of their decision to recall the Allez Spine Laguna Pedicle Screw System, Allez Spine Del Mar Pedicle Screw System and the Allez Spine Cross Connectors as a result of a Warning Letter issued to AlIez Spine LLC by the Food &amp; Drug Administration.The recall letter requested their customers to return their entire inventory of AlIez Spine manufactured products as soon as possible, gave instruction for product return and provided a contact person and telephone number for information regarding the recall. |  | 6472 | 6472 units for all systems | Nationwide Distribution --- including states of AZ, CA, CO, GA, IL, NV, TN, TX, &amp; UT. | Terminated | January 19, 2010 | 516 | 6472
Z-0118-2009 | 49403 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communication System), software versions 3.6.28.x and 4.1.x. |  |  | August 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Two defects have been identified: 1) potential to display a patient on the canvas page that is different from the patient whose images are displayed on diagnostic monitors when using Conference Presentation States and when opening two or more studies from a folder or an exam worklist and 2) potential to miscalculate measurements when Pixel Spacing and Imager Pixel Spacing DICOM tag values are bot | Software design | A notification letter (Urgent Product Correction) dated 8/29/08 was sent via certified mail to all customers alerting them of the issue, requiring posting of notification. Philips Healthcare Informatics, Inc. strongly recommend all customers advise their iSite Radiology users not to open multiple patient exams simultaneously from a folder or worklist using Conference Presentation States. A software correction is being worked on, and should be available by end of Q3 2008 for version 4.1.45.2. A mandatory upgrade is required for all customers who are currently using iSite version 4.1.45.2 and former versions. For specific questions regarding these issues, contact the Customer Advocate (CARE) or Customer Care at Philips Healthcare Informatics, Inc. at 1-877-328-2808. In your communication with Customer Care at Philips Healthcare Informatics, Inc. , please refer to this notification and associated FCO 22945704-08292008-004. |  | 27 | 27 units. | Worldwide Distribution including USA (MN, CO and PA) and countries of Italy, Germany, Netherlands, Sweden, France, Portugal, Switzerland, Thailand, Japan and Australia. | Terminated | November 30, 2010 | 823 | 27
Z-0132-2009 | 49520 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Respiratory Gating System which includes: Pulmonary Digital Assembly Kit and Pulmonary Toolkit Oncology Pro, Model #455011203131 and Model #455011203191. |  |  | September 22, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | Images not correlated: A leak between the Tube Interface and the Outlet Tube of the transducer may result in a failure to produce respiratory correlated images. | Process design | On 9/8/2008 Philips Medical Systems sent recall notification letters to their customers requesting they inventory their stock. They sent instructions to verify that the product is working properly. They are also sending Field Service Engineers to each customer to conduct field test. |  | 507 | 507 units | Worldwide distribution: USA, Australia, Austria, Bahrain, Belgium, Canada, China, Czech Republic, Denmark, England, Finland, France, Germany, Israel, Italy, Japan, Jordan, Korea, Lithuania, Malaysia, Netherlands, Norway, New Zealand, Puerto Rico, South Africa, Saudi Arabia, Spain, Switzerland, Sweden, and Thailand. | Terminated | March 01, 2012 | 1256 | 507
Z-0180-2009 | 49406 | Yes | LFK | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | VCARE (Vaginal-Cervical-Ahluwalia's-Retractor-Elevator) LARGE, Uterine Manipulator, REF 60-6085-102. Made in USA, CONMED CORPORATION, UTICA, NEW YORK USA 13502. The ConMed VCAREÂż Retractor/Elevator is indicated for manipulation of the uterus and injection of fluids or gases during laparoscopic procedures such as Laparoscopic Assisted Vaginal Hysterectomy (LAVH), Total Laparoscopic Hysterectomy (TLH), minilap, laparoscopic tubal occlusion, or diagnostic laparoscopy and also maintains pneumoperitoneum by sealing the vagina once colpotomy is performed. |  |  | September 08, 2008 | 2008 | ConMed Corporation | The firm received several complaints sighting detachment of the balloon at the distal end of the shaft. It was also noted that the incidence of the forward cup slipping off the shaft and being retained in the vaginal canal had increased. | Process change control | The domestic recall notices were mailed on 9/08/08 by USP Priority Mail pouches. The international notices were sent via FedEx International Priority on 9/08/08. The letter described the problem, the health risk and requested that customers inventory the product and return them back to the firm. Customers are requested to fax the reply form to the firm. Contact ConMed Corporation at 1-800-765-8375, ext 3237 for assistance. |  | 32453 | 32,453 (total devices) | Nationwide, Australia, Belgium, Brazil, Chile, Ecuador, France, and Switzerland | Terminated | August 29, 2012 | 1451 | 32453
Z-0191-2009 | 49497 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Osteosynthesis Reduction Spoon; |  |  | August 14, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker has become aware of the potential for breakage of the Reduction Spoon designed prior to 2003. | Device Design | Recall notification letters were sent on August 14, 2008 to all branches and hospitals via Federal Express. Questions are to be directed to 201-831-5718 or a Trauma team member listed in the letter. |  | 1070 | 1070 units | Nationwide | Terminated | May 29, 2012 | 1384 | 1070
Z-0130-2009 | 49558 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Oscor Adelante Luer-Lock Peel Away Introducer Set, for introduction of diagnostic or therapeutic devices into the body. |  |  | July 03, 2008 | 2008 | Oscor, Inc. | Difficulty breaking the sheath hub and subsequently to peel the sheath off for Introducer Set, Adelante size 7F. | Process design | Oscor notified the Direct Account by letter on 07/03/2008. Direct account was asked to return all unused Introducers back to Oscor. All returned products should be labeled with the Return Goods Product number. |  | 396 | 396 | Product was distributed to one direct account in MN. | Terminated | December 09, 2008 | 159 | 396
Z-0150-2009 | 49056 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK Express Automated External Defibrillator (AED). Labeling on the device states LIFEPAK CR Plus/LIFEPAK EXPRESS Defibrillators with ADAPTIV Biphasic Technology; Medtronic Emergency Response Systems, Inc. 11811 Willows Road NE Redmond, WA. The battery operated, small portable device is primarily deployed in Public Access Defibrillator (PAD) environments where the expected use for any one device is once every 20 months. Product is capable of 30 full discharges. |  |  | July 24, 2008 | 2008 | Physio Control, Inc. | The device may not power on although it indicates it is ready for use and would not be able to provide defibrillation therapy. | Nonconforming Material/Component | On 07/24/08, the firm sent an "URGENT MEDICAL DEVICE RECALL" letter with a confirmation sheet via certified return receipt to customers. The letter notified the customer of the recall and the hazard associated with the recall and provided instructions to customers: 1) Immediately verify the AED powers on and voice prompts begin. After verification, press and hold the On/Off button for approximately two seconds to turn the AED off. 2) Keep the AED in use and perform monthly inspections that include the additional AED power on test. Perform a monthly inspection which is consistent with the device labeling. 3) If "OK" is visible on the Readiness Display and AED powers on, it is ready for use. 4) If, at any time, the AED does not power on or if any other indicator displays (i.e. battery or CHARGE-PAK symbol, attention symbol or wrench symbol) immediately call Physio-Control Technical Support (1-800-442-1142). The letter advises customers a Physio-Control representative will contact them within 60 days to arrange a replacement. A confirmation sheet was sent with the letter; to be returned to Physio-Control, Inc. For specific questions or additional information, contact Physio-Control, Inc. at 1-800-442-1142 weekdays between 6:00 a.m. and 4:00 p.m. (PST) or at www.physio-control.com. |  | 2230 | 2,230 Units | Worldwide including USA (and Puerto Rico) and countries of Argentina, Australia, Bahamas, Barbados, Brazil, Canada , Chile, Colombia, Dominica, Germany, Hong Kong, Hungary, Jamaica, Japan, Mexico, Netherlands, New Zealand, Singapore, Uruguay, Venezuela, and Virgin Islands. | Terminated | August 08, 2013 | 1841 | 39020
Z-0205-2009 | 44852 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 base, 110/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion, regional perfusion and cardiopulmonary bypass procedures. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | Local pump controls including start/stop button, select button and manual button may not function. | Component design/selection | Consignees were notified by recall letter dated 7/18/08 and were provided with revised operating instructions telling them to revert to the central control module should these problems occur. |  | 626 | 626 | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 926
Z-0207-2009 | 48174 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory versions 5.3, 5.4, 6.1, 6.2 and 6.3. |  |  | April 17, 2008 | 2008 | Sunquest Information Systems, Inc. | Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong pat | Component design/selection | Customers were sent a preliminary notice on 4/17/2008 followed by customer notice PSN-08-04 on 05/07/2008 via email. The notice advises that this notice is follow-up to the Preliminary notice distributed on 4/17/08 and applies to all clients who have an Instrument Generated Order (IGO) interface.Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong patient. When this occurs, Sunquest generated reports and inquiry contains the incorrect results.The letter provides the various steps that cause the issue to occur. Sites Potentially Affected by this Problem:Sunquest Laboratory using Instrument Generated Orders (IGO) Interface, Version 5.3 and laterAction Sunquest Information Systems is taking to eliminate this Problem:Change Request 1-62VWLJ has been created to address this issue. This will be available as part of Sunquest Laboratory Instrument Build version 1.6.0023 and later.To request the software correction, customers can create a Service request via the Client support Webpage at www.sunquestinfo.com/supportweb. The notice also provides temporary procedural instructions:Effective immediately, do NOT perform Online File Cleanup (OFC) on any IGO instrument interface methods.-Manual OFC should not be performed using functions OFC, or Instrument Maintenance (IX or IXM), option 11 for IGO interfaces only.-If using function Instrument Maintenance (IX or IXM), option 1, 2, 2, Method, 2 (Automatic File Cleanup). This should be set to "N". Then stop and re-start the Instrument Results Processor.Note: Online file cleanup for non-IGO interfaces needs to be performed as usual.-Pending and Overdue Logs need to be monit |  | 419 | 419 clients had the affected code | Worldwide Distribution --- including USA and countries of Canada, Bermuda, Ireland, United Kingdom, United Arab Emirates, and Saudi Arabia. | Terminated | December 05, 2011 | 1327 | 419
Z-0223-2009 | 49466 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose (HEX) Hitachi; 3-HG920-L. (Clinical chemistry) |  |  | September 10, 2008 | 2008 | Pointe Scientific, Inc. | The inability of the product to maintain stated performance specifications through the stated shelf life. | Process control | Consignees are being notified by letter dated 9/10/08 which will be sent out via fax, email or certified mail to all direct account and drop-ship customers. Customers are advised to dispose of any recalled product found and to consider test results obtained with these lots questionable. Contact Pointe Scientific, Inc. at 1-800-757-5313 or 1-734-487-8300 for assistance. |  | 63 | 63 | Nationwide, Canada, China, Romania and Trinidad. | Terminated | November 27, 2009 | 443 | 131
Z-0226-2009 | 49470 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD GeneOhm MRSA 48 ct, Catalog #441244. IDI-MRSA assay is a qualitative in vitro diagnostic test for the direct detection of nasal colonization by methicillin-resistant Staphylococcus aureus (MRSA) to aid in the prevention and control of MRSA infections in healthcare settings. The test performed on the Smart Cycler instrument with a nasal swab specimen from patients at risk for colonization, utilizes polymerase chain reaction (PCR) for the amplification of MRSA DNA and fluorgenic target-specific hybridization probes for the detection of amplified DNA. IDI-MRSA assay is not intended to diagnose MRSA infections nor to guide or monitor treatment for MRSA infections. Concomitant cultures are necessary only to recover organisms for epidemiological typing or for further susceptibility testing. |  |  | July 21, 2008 | 2008 | BD Diagnostics (GeneOhm Sciences, Inc) | BD Diagnostics GeneOhm (BDDG) initiated the recall after they identified that the product has the potential to identify a patient as falsely positive for colonization with methicillin resistant Staphylococcus aureus.The potential health hazard related to a false positive result from this assay would be that a patient may be misclassified as having MRSA colonization when in fact they do not have | Nonconforming Material/Component | BD Diagnostics GeneOhm (BDDG) initial recall notification was made on 07/21/2008 by telephone and followed up with written notification on 07/22/2008. Domestic customers were advised to check their stock to determine if they have inventory remaining of the recalled lot. If so, BD GeneOhm Technical Service will arrange for return of unused product from the customer. The were provided with a Recall Verification form to complete and return. The customers were instructed if they have remaining tests from the affected referenced lot, please discontinue use immediately. They were also told that BD GeneOhm recommend that they review the results of samples tested using this lot of product. Look for any trends outside of the normal expected results, including any changes in MRSA positivity rates compared to your normal prevalence. If they have any questions or concerns following the review, please contact BDDG Technical Service to arrange for consultation with our Medical Director.An telephone number 1-858-334-6356 was given for customers to contact BDDG Technical Service. |  | 60 | 60 kits (48 test/kit) =2880 test | Nationwide to FL, NJ, NY, OH &amp; PA and Internationally to Belgium | Terminated | March 04, 2010 | 591 | 3460
Z-0279-2009 | 49392 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Outback LTD Re-Entry Catheter, Sterile, Single Use. The Cordis LuMend OUTBACK LTD Re-Entry Catheter is indicated for placement and positioning of guidewires within the peripheral vasculature. The cannula is a nitinol needle that is used to re-enter the true lumen of the vessel. |  |  | August 20, 2008 | 2008 | Cordis Corporation | The Cordis LuMend Outback LTD Re-Entry Catheter - separation rate of the cannula to the deployment slide of the handle assembly is higher than anticipated. The cannula is unable to be retracted into the device after deployment due to a separation of the inner key from the cannula. | Other | U.S. Customers - A Notification of Recall Letter with an Acknowledgement Form was sent overnight to 5 individuals in each account on 08/18/2008. For non-U.S. customers, email notification with return acknowledgement is used to notify the affiliate distributors, who then notify and reconcile the customers in those countries (08/18/2008). Contact Cordis special Customer Service at (800) 551-7683 for assistance. |  | 11893 | 11,893 | Worldwide | Terminated | December 01, 2008 | 103 | 11893
Z-0192-2009 | 49515 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hydrophobic HALO360 Filter that is |  |  | September 18, 2008 | 2008 | Barrx Medical, Inc. | Some units may contain the wrong filter, which does not have the proper lock and may result in a leak. | Mixed-up of materials/components | The recall was initiated on September 18, 2008. All users and international distributors were informed by letter. The replacement filters were sent on 9/19/2008. Contact Barrx Medical, Inc. at 1-408-745-8000 for assistance. |  | 650 | 650 units | Worldwide Distribution, including USA, Austria, Australia, Belgium, Canada, France, Netherlands, Switzerland, Sweden, and UK . | Terminated | April 26, 2011 | 950 | 650
Z-0227-2009 | 49484 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar V40 Femoral Stem |  |  | August 20, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | The warning label "do not use with heads more than +10, offset" does not appear on 135 degree stem boxes. This warning is contained in the instructions for use. | Labeling mix-ups | Recall notification letters were sent by Fed Ex on August 20, 2008 to Stryker branches/agencies and Hospital Risk Management departments. Letters to Branches stated the reason for the recall and the potential hazard. Both letters state that it is recommended that physicians follow patients in a manner that is consistent with a physicians routine standard of care for total hip arthroplasty. Contact Stryker Howmedica Osteonics Corp. at 1-201-831-5000 for assistance. |  | 77 | 77 units | Nationwide | Terminated | June 01, 2009 | 285 | 77
Z-0228-2009 | 49492 | No | CAX | Compensated thorpe tube flowmeter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Bag Tee Assembly Model P1407A . An accessory to the dental flowmeter which has a connection function for complete delivery of mixed gases to the breathing circuit portion of the gas scavenging apparatus. The Bag Tee assembly features a non-breathing valve and an Emergency Air Intake located on the bag tee which comply with the American Dental Association guidelines. |  |  | August 26, 2008 | 2008 | Porter Instrument Co Inc | inverted check value on the bag tee assembly which can cause leaking of mixed N2O/O2 gas | Employee error | On 9/2/08, the firm issued an Important Field Correction Notice to their customers informing them of the problem and how to determine if their product was defective and needed to be returned. Contact Porter at 1-800-457-2001 and ask for Dental Customer and refer to the Bag Tee Field Correction Notice. |  | 220 | 220 units | Nationwide | Terminated | September 24, 2010 | 759 | 220
Z-0271-2009 | 49525 | No | LGW |  |  | Neurology | PMA | Precision Charger 1.0 (Model Number SC-5300) for the Precision Spinal Cord Stimulator System labeled as BIONICS SCS-CHARGING KIT CONTENTS Charger (1) Base Station (1) Belt Power Supply (1) Adhesive Kit ADVANCED BIONICS CORP 12740 San Fernando Rd. Sylmar, CA 91342 USA. The Precision Spinal Cord Stimulator System (Precision System) is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs, including unilateral or bilateral pain associated with following: failed back surgery syndrome, intractable low back pain and leg pain. The Precision System includes an implantable 16-output, multi-channel stimulator (IPG - Implantable Pulse Generator) with a rechargeable battery power source. The IPG, commonly implanted in the abdomen or buttock area (less frequently in the subclavicular area), can be configured to accept one or two leads. The leads are implanted in the epidural space, adjacent to the spinal column. The stimulation delivered at the distal end of the lead is intended to mask the pain signals. An externally used Charger is utilized to transcutaneously recharge the IPG (Implantable Pulse Generator) battery. Currently, Boston Scientific distributes Charger 2.0 (Model Number SC-5312). For improved heat management, Charger 2.0 is equipped with thermistor (temperature sensing component), designed to activate/deactivate charging at a set temperature. |  |  | September 22, 2008 | 2008 | Advanced Bionics Corp | Since April 2004, 27 of 8,769 (0.31%) patients have reported receiving second degree burns in the area of charging and 3 of 8,769 (0.03%) patients have reported receiving third degree burns in the area of charging while using Charger 1 .0. Some of these events were the direct result of the patient disregarding the instructions for use, such as sleeping with the Charger on or placing it directly on | Other | In the United States, notification will be sent to patients with the affected devices. Additionally, their physicians on record will be provided with the notification. The Company is currently planning to initiate this field action on September 22, 2008. The U.S. notification will be sent via Fed Ex or U.S. Postal Service certified mail. The letter advises that Boston Scientific Neuromodulation is voluntarily recalling its first generation charging device, Precision Charger 1.0 (Model Number SC-5300) for the Precision Spinal Cord Stimulator System. The Charger 1.0 devices are being recalled and will be replaced with the new second generation Precision Charger 2.0 (SC-5312, currently in distribution) that features a temperature limiter for improved temperature management. The Charger 1.0 can be used until replacement with Charger 2.0 by following the instructions on Page 2 of the letter. Customers are asked to complete the acknowledgement form at the end of the letter and return it to Boston Scientific in the return envelope provided. A representative from Boston Scientific will contact them upon receiving the completed form to schedule an in-person visit at their physician's office to exchange the device. Contact Boston Scientific Technical Serives at (866) 566-8913 for assistance. |  | 6549 | 6549 units | Nationwide and Internationally - Ireland, Korea, Mexico, The Netherlands, United Kingdom, Canada, Australia and Israel | Terminated | January 24, 2012 | 1219 | 6549
Z-0212-2009 | 49385 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BioStar OIA FLU AB, Catalog 90007. Packaged 30 devices per tray. BioStar, Louisville, CO 80027 |  |  | July 02, 2008 | 2008 | Inverness Medical BioStar Inc. | Diagnostic kits for flu were distributed with incorrect components. | Mixed-up of materials/components | Consignees were notified by a Important Customer Notification letter on 7/2/2008 and told to re-examine patient records that indicated a negative result and to destroy any remaining recalled product. The accompanying record should be filled out and faxed to Technical Service. If consignees further distributed any kits, they were to forward the notification to these accounts to assure that patient results were reviewed, and destruction of remaining kits could be verified. For assistance, customers can call 800-637-3717. |  | 509 | 509 kits | Nationwide Distribution --- including state of IL. | Terminated | November 04, 2008 | 125 | 509
Z-0201-2009 | 49590 | [6 empty] | Vibe Technologies, VIBE machine, Vibrational Integrated Biophotonic Energizer. |  |  | October 01, 2008 | 2008 | VIBE Technologies | Medical device marketed without marketing approval for claims that include cures cancer, infections, depression. | PMA | Vibe Technologies notified their customers by letter on October 1, 2008 and were told to discontinue use as a medical device. This letter notification included: 1. A warning label to be placed on the Vibe machine stating that it is not a medical device and should not be used as such.2. An updated operation manual/users guide containing no medical conditions or human body claims.3. A certification to be signed by the user and returned acknowledging that they: received the letter; affixed the warning label to the Vibe machine; understand that the Vibe machine does not affect the structure/function of the human or animal body; will not promote the Vibe machine as a medical device; will remove any medical claims from their website; destroy any Vibe literature making medical claims. Approximately three weeks after the certified mailing, the firm will begin performing recall effectiveness checks. |  | 851 | 851 units | Worldwide | Terminated | May 21, 2012 | 1328 | 851
Z-0131-2009 | 49531 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 18 gage XTW needles, sold as a component in SafeSheath Hemostatic Tear-away Introducer System with Infusion Side Port kits or sold Bulk Non Sterile (BNS). |  |  | September 12, 2008 | 2008 | Thomas Medical Products Inc | Difficulty inserting guidewire through introducer needle. Needle hub is not tapered down to the needle cannula. | Process control | On 9/11/08, Thomas Medical Products issued letters to the distributors informing them of the problem, to cease distribution and await further instructions. Additional instructions to the distributors were to notify all their account via an advisory notice and request that the user not use the needle supplied with the kit and to replace with another commercially available sterile 18 gage introducer needle. |  | 246585 | 246,585 needles | Nationwide distribution via distributors in CA, MA, and MN. | Terminated | May 06, 2009 | 236 | 246585
Z-0278-2009 | 48739 | No | MRV | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | S/PÂż Orange 50g Glucose Tolerance Beverage, 10 oz glass bottles, 12 bottles/case, labeled in part ***Cardinal Health McGaw Park, IL 60085-6787 USA*** For prescription use only. Product is used in support of diagnostics testing for hyperglycemia and hypoglycemia. |  |  | April 23, 2008 | 2008 | Nerl Diagnostics LLC | Beverage may contain glass particles. | Other | The recalling firm notified distribution centers by letter dated 04/23/08 and advised of a single report of glass found in a bottle. The notifications further advised distributors to quarantine referenced product, notify their customers and return the enclosed reply form. Distributors were requested to return or destroy product for replacement. Please contact NERL Diagnostics, LLC at 1-401-824-2046 for assistance. |  | 25908 | 25908 bottles | Nationwide. | Terminated | November 26, 2012 | 1678 | 25908
Z-0137-2009 | 49426 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Bicor X-Ray System, Model number 9023136. |  |  | August 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Unintended movement of the system table and/or C arm. | Software design | Siemens Medical Solutions issued a letter to consignees on 8/11/2008 to inform them of the problem and a workaround should the issue occur prior to software revision. |  | 20 | 20 units | Nationwide. | Terminated | December 21, 2012 | 1593 | 975
Z-0193-2009 | 49529 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis. Angiographic X-ray System. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Patients could possibly fall off the table if not properly secured when: 1) moving them onto the patient table, 2) moving them on the patient table or 3) removing them from the patient table. | Labeling design | The recalling firm issued a Customer Safety Advisory Notice dated 7/04/08 (including an Addendum to the Operators Manual) which alerts the operator of possible injury to patients if they are not properly secured to the table during examinations and describes the recommended fastening procedures required to properly secure the patient to the table.For questions or additional information, contact Siemens Medical Solutions, Inc. at 610-448-4634. |  | 1530 | 1,530 units | Nationwide Distribution. | Terminated | September 23, 2009 | 397 | 1638
Z-0282-2009 | 50113 | No | JEY | Bone plate. | Dental | Dental | 510(k) | HT X-DRIVE SCREW, Part # 91-2416, W.Lorenz Surgical. Screws for use in the stabilization and fixation of mandibular fractures and mandibular reconstructive surgical procedures. |  |  | July 11, 2008 | 2008 | Biomet Microfixation, Inc. | The packaging for the 91-2416, 2.4X16MM HT X-DRIVE SCREW, Lot 493200 actually contained the 99-6577, 2.0X7MM FOSSA X-DRIVE SCREW. | Labeling mix-ups | An "Urgent Medical Device Recall Notice" was sent to customers on 07/11/2008 by certified letter. The letter described the problem, the risk and asked that customers discontinue use of the product and return them to Biomet. Also, customers were provided with a Inventory Reconciliation sheet and asked to fax it back to Biomet. Please contact Biomet Microfixation at 1-800-874-771 or 904-741-4400, ext. 448 for assistance. |  | 10 | 10 | Nationwide and Canada | Terminated | September 11, 2009 | 427 | 10
Z-0213-2009 | 49440 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Mobile Care Server used in conjunction with Web Viewer, Pocket Viewer, and Cellular Viewer WIT with L-WEB05 software. |  |  | July 16, 2008 | 2008 | GE Healthcare | Issue 1 -Possible failure of CIC Pro audible alarms: Requests for telemetry defaults from Mobile Care Server could cause excessive CIC Pro memory consumption. This excessive memory consumption may lead to potential loss of audio at the CIC or a system reset or reboot. Issue 2 - Mobile Care software/Mobile Viewers don not recognize ApexPro FH Telemetry transmitter: Mobile Care server does not ac | Software design | Consignees contacted via a GE Healthcare Urgent Medical device Correction letter dated 8/26/08 and supplemental instruction for installation of software" Mobile care server with software, L-WEB05 and iCentral with software, L-NET05 supplemental Instruction for installation". The letter was addressed to Healthcare Administrator/Risk manger, chief of Nursing and director of Biomedical engineering / IT. The letter also described the problem, affected Products,and Short Term Recommendations . |  | 292 | 292 | Worldwide Distribution --- USA including states of CALIFORNIA, IOWA, INDIANA, VERMONT, and WASHINGTON and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CROATIA, CYPRUS, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY, and UNITED KINGDOM. | Terminated | November 24, 2015 | 2687 | 292
Z-0273-2009 | 49554 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI, Inc. dissector PlasmaKnife (DPK) |  |  | October 01, 2008 | 2008 | Gyrus ACMI Corporation | Sterility may be compromised | Process control | Gyrus ACMI initiated a voluntary recall by an Urgent: Medical Device Recall letter on 10/01/08 informing accounts of the problem, cease any further use of the affected product, remove any of the affected product service representative to obtain a Return Goods Authorization, return the product, and complete the reply form and fax it to the Regulatory Affairs Dept. so that the response can be tracked. |  | 1340 | 1340 units | Worldwide Distribution --- including USA and countries of Canada, UK, Italy, Spain, Norway, Australia, and Taiwan. | Terminated | May 07, 2012 | 1314 | 1340
Z-0284-2009 | 49606 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | ReliOn Insulin syringes - 1cc, 31 G |  |  | October 09, 2008 | 2008 | Covidien LP | Mislabeled: Package labeled as an insulin syringe for use with U-100 insulin contains an insulin syringe for use with U-40 insulin. Risk of overdose of insulin. | Process design | Tyco Healthcare Group LP (Covidien) notified Can-Am Care, Corp. by letter sent by e-mail and by Federal Express on 10/9/08. The letter requests that they immediately inventory, quarantine any remaining stock they may have of the product, return, and notify their customers. Wal-Mart conducted a consumer mailing on October 14, 2008, and is posting placards in their stores. Can-Am Care, whose name appears on the label of the product has concurrently posted a recall notice on their website. |  | 471000 | 471,000 (4710 boxes of 100) | Nationwide | Terminated | July 06, 2010 | 635 | 471000
Z-0188-2009 | 49448 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Alphastar 1132.01AX/BX, 1132.02AX/BX, and 1132.03AX/BX Universal Operating Table |  |  | September 16, 2008 | 2008 | MAQUET Inc. | Table tilts: The Alphastar Surgical Table piston rod head may break and cause a full inclination of the table top during Trendelenberg positioning if there is insufficient lubrication. | Equipment maintenance | Maquet sent Recall Notification letters to all customers on September 16, 2008 by Federal Express. Letters inform customers of the potential problem with the Alphastar tables, a description of the failure and hazards involved. Hospitals are asked to confirm whether yearly maintenance by an authorized technician have been performed. If preventive maintenance has not been performed, customers are requested to contact their Maquet Service rep. A Customer Service Response form is included and customers are to complete if to acknowledge whether of not the facility is still in possession of an affected table. If so, the model number, serial number and facility name &amp; address is to be completed so that Maquet reps can visit the facility and perform the corrective action. |  | 141 | 141 in US | Nationwide | Terminated | January 12, 2010 | 483 | 141
Z-0203-2009 | 44851 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Base, 100/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion, regional perfusion and cardiopulmonary bypass procedures. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The pumps may exhibit blank or distorted displays, loss of local control and/or may have local control knobs that are difficult to use. | Other | Consignees were notified via letter dated 7/18/08 and instructed to use the central control module should the problem occur. Contact Terumo Cardiovascular Systems Corp at 1-800-521-2818 for assistance. |  | 626 | 626 | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 926
Z-0274-2009 | 49600 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PurePoint System Operator's Manual, Catalog Number 8065751131, Rev. B for use with the Alcon PurePoint Laser, Catalog Number 8065750597 |  |  | September 26, 2008 | 2008 | Alcon Laboratories, Inc. | Alcon had identified that indications for use unapproved by the Food and Drug Administration are included in the PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Revision B. | Employee error | An Alcon representative has been provided with a package of information for each customer site. All domestic customers are being provided an Urgent: Medical Device Correction letter describing the change to the PurePoint Operator's Manual and are being issued a revised PurePoint Operator's Manual (Revision C). The letter advises that Alcon Research, Ltd, has initiated a voluntary correction based upon unapproved indications for use identified within the Alcon PurePoint Laser System Operator's Manual, Revision B. The letter advises that the following indications for use have been updated or removed on page 1.14, Professional Operator's Information, in Revision C of the PurePoint Laser System Operator's Manual:Updated:"AMD; Wet or Dry to include macular degeneration" has been updated to read "Choroidal neovascularization secondary to age-related macular degeneration (AMD)"."Trabeculoplasty for treatment of Chronic/Primary Open Angle Glaucoma (COAG, POAG), Acute Angle ClosureGlaucoma (AACG), and refractory glaucoma" has been updated to read "Trabeculoplasty for treatment of Chronic/Primary Open Angle Glaucoma (COAG, POAG) and refractory glaucoma"Removed:Macular photocoagulation; including grid, focal, laser drusen scatter (panretinal)Transcleral cyclophotocoagulationIntra-ocular tumors; to include choroidal hemangioma, choroidal melanoma, retinoblastomaOtosclerotic hearing loss. No other changes to the manual have been made in this revision and no other Alcon products are affected by this correction. Per the letter, an Alcon Representative will provide customers with a new PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Rev. C. and customers are return their existing PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Rev.B. A signed confirmation of receipt of the letter is required from all customers. |  | 41 | 41 units | Worldwide Distribution --- USA including states of IL, WI, TX, MI, PA, NY, CT, SC, NC, LA, FL, GA, CA, OR, WA, CO, and NV, and countries of India, Bulgaria, Barbados, and Dominican Republic. | Terminated | March 07, 2012 | 1258 | 41
Z-0209-2009 | 49092 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | Iris Sample Processing SafeCrit 40mm Tubes. Catalog number 59-000869-002 H. Iris Sample Processing, 60 Glacier Drive, Westwood, MA 02090. |  |  | August 01, 2008 | 2008 | Drummond Scientific Co | presence of OSCS contaminant | Material/Component Contamination | The recalling firm issued recall letters dated 8/1/08 via certified mail to the label owners informing them of the problem and the need to notified their customers. |  | 4000 | 4000 vials | Nationwide Distribution --- including states of FL and MA. | Terminated | December 18, 2008 | 139 | 19000
Z-0281-2009 | 49451 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ASP Automatic Endoscope Reprocessor, AER Plus Product Code 20300 and AER with printer Product Code 20301 |  |  | September 11, 2008 | 2008 | Advanced Sterilization Products | Residual high-level disinfectant solution remaining in endoscopes that have been reprocessed in the ASP Automatic Endoscope Reprocessor. If high-level disinfectant, or other fluids, remains in the endoscope after reprocessing, contact with the mucous membranes may occur and result in chemical burns, irritation (chemical colitis), or other symptoms. | Component design/selection | ASP is mailing customer notification CL-101367 and AD-52976-001 in an Urgent Product Correction envelope starting the week of September 8, 2008. The letter advises that the firm has received customer reports of residual high-level disinfectant solution remaining in endoscopes that have been reprocessed in the ASP Automatic Endoscope Reprocessor (AER). If high-level disinfectant, or other fluids, remain in the endoscope after reprocessing, contact with patient mucous membranes may occur and result in chemical burns, irritation (chemical colitis), or other symptoms. ASP has tested the PENTAX 70 series endoscopes containing FWJ (Forward Water Jet) and has confirmed that residual fluids may remain in the PENTAX 70 series endoscopes with a FWJ after reprocessing. To eliminate any residual fluids, ASP is recommending customers follow the revised instructions when reprocessing PENTAX 70 series endoscopes with a FWJ. A copy of the revised instructions is provided with the notification letter. |  | 1945 | 1,578 units installed in the US and 367 units installed internationally | Worldwide Distribution --- including USA and countries of Italy, UK, Uruguay, Argentina, Chile, Colombia, Ecuador, Puerto Rico, Mexico, Australia, China, Singapore, Taiwan, Middle East, Canada, Czech Republic, Portugal, South Africa, Spain, Hungary, Poland, Brazil, Russia, Egypt, and Slovenia. | Terminated | December 27, 2011 | 1202 | 1945
Z-0283-2009 | 49475 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | One Touch Ping Glucose Monitoring System |  |  | August 04, 2008 | 2008 | Animas Corporation | display old, inaccurate values for bolus amount delivered and amount of planned bolus totals | Software design | On 8/6/08, the recalling firm telephoned all users to inform them of problem and the need to return the system. |  | 251 | 251 units | Nationwide Distribution | Terminated | November 12, 2008 | 100 | 251
Z-0215-2009 | 49371 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | VACUETTEÂż, 2 ml Lithium Heparin Venous Blood Collection Tubes. Reference #454237, 13x75 green cap-white ring, 24 racks of 50 pcs (1200 pcs in total, non ridged, Sterile). Manufactured by: Greiner bio-one NA, Inc., 4238 Capital Drive, Monroe, NC 28110. |  |  | July 29, 2008 | 2008 | Greiner Bio-One North America, Inc. | Labeling Error: Incorrect tube label reads "Z Serum Clot Activator" instead of "LH Lithium Heparin" tube.All other information on the tube label including color coding of tube was correct. Case and rack label information showed the correct information with 2 ml Lithium Heparin. | Labeling mix-ups | Consignees were notified by letter on 07/29/2008. Notification letters to distributors instructed them to stop distribution of the product and destroy all affected lots of the product. In addition, notification letters addressed to users instructed them to destroy and discard all products from the item and affected lots. The "Product Disposition Form" must be completed by the consignee and returned to Greiner Bio-One North America, Inc.For questions or additional information, contact Greiner Bio-One North America, Inc. at 704-261-7800. |  | 33 | 33 cases (39,600 pieces) | Worldwide including USA (states of AK, AL, AR, CA, CO, DE, GA, ID, IL, IN, KS, LA, MA, MI, MO, MS, NC, ND, NE, NH, NY, OH, OK, OR, PA, RI, SC, SD, TX, VA, VT, WA, and WI) and country of Canada. | Terminated | August 26, 2009 | 393 | 33
Z-0295-2009 | 49604 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Sarns Level Sensor II (yellow) (Part # 195215) alert level sensor transducer for use with Sarns Advanced Perfusion System I, Perfusion System 800 and Perfusion System 900; Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; REF 195215. Level sensors are distributed as a finished device, as part of a kit, with a perfusion system 8000 safety monitor or with an Advanced Perfusion System 1 base. |  |  | May 23, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The level sensor may not properly couple to the reservoir, resulting in a sensor "not attached" message, an alert or alarm condition, or a failure to detect a low level condition. | Nonconforming Material/Component | U.S. consignees were notified by (Urgent Medical Device Removal) letter dated 5/23/08 via FedEx and international affiliates were notified via e-mail on 5/27/08. Consignees were advised of the problem, and to continue using the affected sensors until replacements are available in late 2008.For questions or additional information, contact the Terumo CVS Customer Service at 1-800-521-2818. |  | 574 | 574 | Worldwide including USA and countries of Australia, Belgium, Bolivia, Canada, Chile, Ecuador, Egypt, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Latin America, Malaysia, Panama, Philippines, Saudi Arabia, Singapore, South Africa, South Korea, Sri Lanka, Taiwan, Thailand, United Arab Emirates, Venezuela and Vietnam. | Terminated | August 04, 2009 | 438 | 1121
Z-0197-2009 | 49553 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DeRoyal Lap Appy TraceCart, REF53-1776, Rx only, non-sterile. (custom surgical kit) |  |  | October 01, 2008 | 2008 | DeRoyal Industries Inc | Surgical kits contained a recalled Endopath ETX 35 mm Endoscopic Linear Cutter White Reload/Cartridges. | Nonconforming Material/Component | The firm sent out recall notices to their sales representatives, the distributors, and the hospitals addressed to the Recall Coordinator/OR Supervisor on 10/2/08. These notices stated that the recalling firm's sales representatives would perform the field correction in lieu of returning the kits. Contact DeRoyal Industries, Inc. at 1-865-362-6157 for assistance. |  | 4 | 4 kits | LA &amp; KY | Terminated | December 02, 2009 | 427 | 34
Z-0280-2009 | 49575 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 7F Bard SSV 13 cm Split Sheath with Valve and sideport |  |  | October 01, 2008 | 2008 | C. R.Bard, Inc./BardElectrophysiology Division | Failure to insert the guidewire through the introducer needle. | Process control | Bard issued a recall notice on 1 October 2008 [note that the letter is dated 29 September 2008] to all recipients of affected product via certified mail. Users are requested to discontinue use and return inventory. Contact Bard Customer Service at 1-800-824-8724 for assistance. |  | 3105 | 3,105 kits | Nationwide | Terminated | March 09, 2010 | 524 | 3105
Z-0275-2009 | 49602 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ATEC Breast Biopsy and Excision System 12 gauge, 20 centimeter handpiece, sterile, Hologic, Inc., Indianapolis, IN; Model number ATEC 1212-20. |  |  | September 22, 2008 | 2008 | Hologic, Inc. | The distal tip of the needle may become detached and remain in the patient, requiring surgical removal. | Process control | Consignees were notified via an Urgent Medical Device Recall letter dated 9/17/08 asking them to immediately check their inventories and quarantine all affected product from the listed lots. Also, if the product had been further distributed, users were to be identified and notified of the recall. Products are to be held in quarantine until they are returned to Hologic. The response form must be completed and returned to fulfill regulatory requirements and document actions taken at your facility. Handpieces whose pouches had been opened were to be destroyed, as they could potentially shift during transport and expose dangerous sharp points to personnel handling the boxes. Upon receipt of the attached response form, Hologic will promptly send an equivalent amount of replacement product free of charge. |  | 5120 | 5,120 (1,024 cases of 5) | Worldwide Distribution --- including USA and countries of Australia, Belgium, Brazil and Singapore. | Terminated | November 18, 2009 | 422 | 5120
Z-0276-2009 | 50068 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Barco Surgical Display, Model MDSC-2124 |  |  | September 09, 2008 | 2008 | Barcoview | The front protective cover may loosen and completely fall off. | Other | Consignees were notified via a Field Safety Notice letter on 09/09/2008. They were instructed to inspect the units to check if the front panel is tightly secured to the display housing. Any front panel found to be lose or showing a gap between the front plate and the back shell of the monitor is to be reported the BARCO Service &amp; Support organization and is to be secured using adhesive tape. Follow up Field Safety Notice was sent to consignees on 09/26/2008 indicating the root cause and the firm's corrective field actions implemented October 15, 2008. |  | 25 | 25 units | Nationwide Distribution --- including states of AK, GA, IL, and MA. | Terminated | June 02, 2009 | 266 | 25
Z-0287-2009 | 49571 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | NAMIC Custom Angiographic Kit, Right Heart Kit, REF/Catalog # 60190254, Sterile, and bulk Non-Sterile Manifolds Assemblies, Product #s 40338200 and 40336200 (foreign product for further processing, e.g., packaging, labeling, sterilization, and distribution under BSC). Manufactured by: Boston Scientific Corporation/NAMIC Technology Center/Glens Falls, New York 12801. |  |  | August 06, 2008 | 2008 | Navilyst Medical | Firm became aware that a specific batch of the product may contain loose plastic particulates in the fluid pathway. | Process design | On 8/12/08, Navilyst Medical sent a fax and phoned the sole domestic hospital account that had received the affected units. Navilyst Medical instructed the consignee to immediately locate, segregate, and quarantine the units. The hospital consignee confirmed that all units were still in their control and that they will maintain them in quarantine awaiting return to Navilyst Medical. The sole international account, Boston Scientific in Ireland was notified of the problem on 8/06/08. This was followed up with an Urgent Medical Device Recall Letter and Verification Form sent via FedEx on 8/20/08. |  | 16220 | Domestic - 20 units; Foreign - 16,200 bulk units | Worldwide Distribution --- USA including state of MO and country of Ireland. | Terminated | June 16, 2009 | 314 | 16220
Z-0288-2009 | 49709 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Biomet 2.0 MM Stainless Steel (SS) Crimp Sleeve, sterile, use w/2.0 MM S.S. cable # 350800, Biomet Orthopedics, Inc., Warsaw, IN; REF. 350805. |  |  | September 24, 2008 | 2008 | Biomet, Inc. | The outer package is properly labeled, but the inner package may be labeled as a femoral component. | Labeling mix-ups | The firm's sales force was notified of the problem via an Urgent Medical Device Removal Notice letter dated 9/24/08, and provided instructions to retrieve the product from customer accounts and to provide them with a copy of the removal letter. Distributors were to immediately locate and remove the identified implants from circulation, carefully follow the instructions, and fax a copy of the Response Form prior to the return of the product. If the product was further distributed, hospital personnel MUST be notified via the enclosed "Dear Biomet Customer" notice. The letter MUST be given to hospital personnel responsible for receiving recall notices but distributors are charged with the location and return of those products. |  | 48 | 48 | Worldwide Distribution --- USA including states of Arkansas, Florida, Kansas, New Jersey, and Ohio, and countries of Argentina, Canada, Chile Italy, Mexico and United Kingdom. | Terminated | June 29, 2009 | 278 | 48
Z-0218-2009 | 49463 | No | MLY |  |  | Physical Medicine | 510(k) | Gebauer's Fluro-EthylÂż Nonflammable Topical Anesthetic Skin Refrigerant (Aerosol Can) P/N 0386-0020-20. The product consists of a can, a valve and an actuator. |  |  | September 03, 2008 | 2008 | Gebauer Company | Some units of Fluro-Ethyl containing the defective valve which could malfunction and spray refrigerant out from the side of the valve in addition to or instead of spraying in the normal inverted position from the product's actuator. The product can also leak coolant from the side of the valve onto the fingers of the user. | Component design/selection | A recall notification letter (Recall Notice) (dated 9/3/2008) was issued on 9/16/2008 to customers. The notice instructs customers to check their inventory for any remaining recall product and if found, return the product to the Gebauer Company. The letter also advises distributors to notify their customers and follow instructions as noted above, i.e. customers inventory their stock for recalled product and if found, return to the Gauber Company. For questions or additional information, contact Gebauer Company at 216-581-3030 (ext. 120) or 800-321-9348. |  | 12717 | 12,717 Units | Nationwide Distribution. | Terminated | October 13, 2010 | 770 | 12717
Z-0214-2009 | 49136 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory System. Software Version: Sunquest Laboratory v5.4.2, v6.1, v6.2 and v6.3. |  |  | August 01, 2008 | 2008 | Sunquest Information Systems, Inc. | While performing in-house testing for the Sunquest Laboratory System and under certain conditions, there have been instances in which an error that occurred at the database level or originated in the operating system (Cache)was not communicated or displayed in the Graphic UserInterface (GUI) application so the activity continued. While this error was registered in the error log in the SystemU | Software design | A notice (PSN-08-07) was issued to all customers on 08/01/2008 via e-mail. The notice advises customers of the problem and actions Sunquest Information Systems has taken to eliminate the Problem. Sunquest Client Support will contact their facility to schedule correcting the error trap setting. The notice also provides temporary procedural instructions to include monitoring error log using System Utilities (Function UTL) and suboption 2, System Error/Message Inquiry. Contact Client Support at 520-570-2252 for questions or additional information. |  | 139 | 139 | Worldwide Distribution including USA and countries of Canada, Denmark, Saudi Arabia and United Arab Emirates. | Terminated | November 28, 2011 | 1214 | 139
Z-0195-2009 | 49546 | No | EWO | Audiometer. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | GSI Audera (version 2.6 software) Viasys Healthcare, Inc. |  |  | September 16, 2008 | 2008 | Cardinal Health, Inc | GSI Audera systems requires a system software update due to issues with the Split-Screen and the Vestibular Evoked Myogenic Potential (VEMP) functions. A. Split Screen Field Correction:GSI Audera units with version 2.6 software may mislabel Auditory Evoked Potential (AEP) waveform responses with respect to the identification of the stimulus ear. When using 2.6 software in split-screen mod | Software design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated September 16, 2008 on 9/22/08 and 9/26/08. The letter was addressed to the Customer/Distributor. The letter instructed consignees to perform installations if they have the following software: 1) For Version 2.6 software, you must install Version 2.6.5 software to correct a Split- Screen performance issue and to remove the VEMP function. 2) For Version 2.0 or 2.1 software, you must install the Protocol Update Software to remove the VEMP function. Versions 2.0 and 2.1 do not contain Split-Screen capability. "Attachment A-Verification Form" must be completed by the recipient/distributor.If you need assistance in making the software revision determination or have further questions, please contact the Cardinal Health (formerly VIASYS) Neurocare/GSI Technical Support Group at 1-800-356-0007, 1-800-700-2282, 608-273-5000, or by e-mail at gsiservice@viasyshc.com or gsiservice@cardinalhealth.com. |  | 297 | 297 | Nationwide including states of AL AZ CA CO CT DC FL GA HI IA IA IL IN KS LA MA MD ME MN MO MS NC NC NH NJ NM NY OH OK PA RI SC TN TX UT VA WA and WI. | Terminated | May 01, 2012 | 1323 | 297
Z-0216-2009 | 49427 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect LH MasterCheck (List 6C25-05). An in-vitro diagnostic, consisting of 4 - 2 mL vials of LH MasterCheck, Levels 0, 1, 2 and 3. Manufactured for Abbott Laboratories, Abbott Park, IL 60064. |  |  | September 05, 2008 | 2008 | Abbott Laboratories | The values listed in the Architect LH MasterCheck, Lot 82520 data sheet are incorrect. When MasterChecks do not perform as intended, the integrity of the system cannot be verified and patient results would not be generated. Other commercially available materials are available to customers in order to meet CLIA requirements. | Error in labeling | A Product Correction, Immediate Action Required letter and Customer Reply Form dated 9/5/08 were sent to all Architect LH MasterChecks (List Number 6C25-05) customers who received Lot 82520M. Customers were informed that the values listed in the Architect LH MasterCheck Lot 82520 data sheet are incorrect. Customers were requested to check their inventory for the affected lot, and if any kits were found, to substitute the values printed on the data sheet with the values supplied in the letter. In addition, Customers were requested to keep a copy of the letter on file and to complete the Customer Reply Form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for the Architect LH MasterCheck testing, and fax it to Abbott by 9/19/08 at 1-800-777-0051. Follow-up telephone calls will be placed to customers who do not provide a written reply via the Customer Reply Form. U.S. Customers please contact Abbott Laboratories Customer Service at 1-877-422-2688 for questions regarding this information. |  | 30 | 30 kits | Nationwide Distribution including Arizona, California, Florida, Illinois, Kansas, Louisiana, Massachusetts, Missouri, Montana, New York, Ohio, Pennsylvania, South Carolina, Texas, Utah and Virginia. | Terminated | March 18, 2010 | 559 | 30
Z-0272-2009 | 49544 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bone Marrow Aspiration Kit, consists of 3 main components packaged in a box. 1. Marrow Loc (Bone Marrow Specimen Acquisition System), 2. "J" Style Bone Marrow Biopsy Needle (11g (3.0mm)x10cm (4.0)) 3. BD 20ml Syringe. |  |  | August 27, 2008 | 2008 | Theken Spine LLC | The kits were manufactured in 2005 with syringes that expire on 10/2009 and labeled with an overall expiration on 10/2010. | Other | On 8/27/2008 the firm sent recall notification letters to each of their customers notifying them of the problem with the product. They were requested to return any remaining product back to the recalling firm and to call Customer Service at (866) 942-8698 for an RGA and to arrange shipment. On 8/28/2008 the firm conducted follow up contact each distributor via telephone. |  | 115 | 115 total | Nationwide Distribution --- including states of CA, NC, NJ, NM, NY, TX, and UT. | Terminated | April 13, 2012 | 1325 | 115
Z-0217-2009 | 49441 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Reflex Hybrid 4.0 x 14 mm Variable, Self-Tapping Bone Screw (Non- sterile). |  |  | August 18, 2008 | 2008 | Stryker Spine | The Reflex Hybrid 4.0 x 14 mm Variable, Self Tapping Bone Screw, Lot NSZ, was determined to have been anodized with the incorrect color for that screw. | Nonconforming Material/Component | An URGENT Product Recall Letter and a Stryker Spine Accountability Form were sent via FedEx with return receipt on August 18, 2008 to Stryker Branches and Hospital OR Supervisors. Letters notified branches of the recall and requested that they immediately identify and quarantine affected product and notify hospital locations that received shipment of product within 72 hours prior to identification of the problem. Letters sent to Hospital OR Supervisors described the problem and requested that all inventory locations be checked for affected product. If recalled product is found, Stryker Spine Representatives must be contacted to retrieve the recalled product from the facility. The Stryker Spine Accountability Form must be completed and returned via fax, 201-760-8269 to Stryker Spine. Contact Stryker Spine Clinical and Regulatory Affairs at 201-760-8150 or 201-760-8069 for questions or additional information. |  | 183 | 183 units | Nationwide Distribution. | Terminated | June 19, 2009 | 305 | 183
Z-0301-2009 | 48712 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 28 mm diameter, neck length +3.5 mm, sterile, Zimmer, Warsaw, IN; REF 00-6418-028-03 and 00-6428-028-03. |  |  | June 11, 2008 | 2008 | Zimmer Inc. | When used with cobalt/chromium hip stems, the average test values for ceramic head bursting falls below the guidance document limits if the taper is made by casting instead of by forging. | Device Design | The firm issued Product Compatibility Update # 13, dated 6/11/08, to consignees informing them which products are not compatible, and thus should not be used in combination. |  | 49474 | 49,474 of all products | Worldwide Distribution --- including USA and countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, England, Finland, France, Germany, Greece, Hungary, India, Israel, Italy, Korea, Lebanon, Mexico, Netherlands, New Zealand, Poland, Romania, Singapore, Slovakia, Spain, Switzerland and Ukraine. | Terminated | November 10, 2009 | 517 | 49474
Z-0004-2009 | 49413 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Test Strips, part number 020-994-02, manufactured by LifeScan Inc, Milpitas, CA, 25 count. |  |  | September 19, 2008 | 2008 | Lifescan Inc | Products exceed inaccuracy threshold , with many complaints resulting from inaccurate low results. | Process control | Urgent: Medical Device Recall letters, dated 09/19/2008, were sent to consumer end users, direct account distributors and wholesalers, health care providers, and pharmacists notifying them of the recall. The letter states that if customers have strips from Lot 2829235, they should call LifeScan at 1-866-247-1281 and tell customer service that they have strips that need to be replaced. Customer service and QA scripts were provided to LifeScan staff to provide to any callers who inquire about the recall. The distributors and wholesalers were to locate and sequester any inventory from the affected Lot 2829235; request all recalled inventory from customers; call Carolina Logistics for an RGA; and when only the affected returned product is received, a credit will be issued. |  | 35308 | 35308 vials | Nationwide Distribution --- including Puerto Rico. | Terminated | December 27, 2010 | 829 | 35308
Z-0306-2009 | 49397 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List numbers 08H00-01, 08H00-03 |  |  | August 25, 2008 | 2008 | Abbott Laboratories | Exposed wire resulted in minor electrical shock and burn to the service technician | Device Design | Recall issued August 22, 2008. Technical service bulletin was issued for all field personnel via internal database. No customer communication is planned as customers are instructed *not* to remove the cover, which is necessary to access the area. |  | 958 | 958 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Brazil, Canada, Chile, China WFOTC, Colombia, Germany, Hong Kong, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | January 06, 2011 | 864 | 958
Z-0297-2009 | 48628 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02) |  |  | May 22, 2008 | 2008 | Abbott Diagnostic International, Ltd. | AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file.The instrument might require installation of the AxSYM Digoxin II assay file. | Software change control | A Product Information letter, dated May 22, 2008, was issued to customers. It requested that one of the following two actions be taken if the AxSYM Digoxin II assay file is not installed on your AxSYM instrument and you want to run the AxSYM Digoxin II assay. If you have one of the previously manufactured software disks listed, [AxSYM Digoxin II Assay Disk, LN 2C43-01, or higher; or AxSYM TDM/Transplant Assay Disk Version 1.0 (3D53-01) or Version 2.0 (3D53-02)], install the AxSYM Digoxin II assay file. Do not install any other assay file. If you do not have one of the assay disks listed, contact your local Customer Service representative to have this assay file installed or install the file from a previous version if you have it on hand. |  | 9069 | 9069 kits | Worldwide Distribution --- USA including states of AL, AZ, AR, CA, CT, FL, GA, IL, KS, MD, MA, MI, MN, MS, MO, NV, NJ, NY, ND, OK, OR, PA, RI, TX, VA, WA, Hawaii, and Puerto Rico, and countries of Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Ecuador, Peru, Chile, Uruguay, Argentina, Puerto Rico, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand, Curacao, and Panama. | Terminated | September 09, 2011 | 1205 | 9069
Z-0305-2009 | 49097 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂż, labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. |  |  | June 20, 2008 | 2008 | Datex Ohmeda, Inc. dba GE Healthcare | Manufacturer of neonatal incubators changed the device labeling to account for new use and care instructions. | Equipment maintenance | A notification letter (Urgent Medical Device Correction) was sent to consignees on 06/20/08. The notification letter addressed healthcare professionals, including neonatal, infection control and nursing departments at medical facilities. In add tion, the letter advised of revisions to the current Giraffe OmniBedÂżÂż cleaning instructions contained in the product Operators Manual. Instructions were enclosed with the notification letter and are intended to supplement those already present in the manual.Contact the GE Healthcare Technical Support center at 1-800-558-7044 for questions or additional information. If outside of the USA, please contact your local GE Healthcare Service Representative. |  | 7878 | 7,878 beds | Worldwide Distribution: The devices were widely distributed to national and international medical facilities. | Terminated | March 24, 2010 | 642 | 7878
Z-0289-2009 | 49867 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard(TM) DCM PS Plus 16mm x 63/67mm Tibial Bearing (REF. 183726), ARCOM UHMWPE, sterile, Biomet Orthopedics, Inc., Warsaw, IN. |  |  | September 22, 2008 | 2008 | Biomet, Inc. | The package is properly labeled, but the laser etch for size incorrectly reads "16 X 63/63" instead of the correct size, "16 X 63/67". | Process control | A notification letter (Urgent Medical Device Removal Notice) dated 9/22/08 was sent to the firm's sales staff with instructions to locate and retrieve the product and to notify customers at affected hospitals of the problem.Questions related to this notice should be directed to the Biomet Orthopedics, LLC Corrective Actions Manager at 800-348-9500 or 574-372-3983, Monday through Friday, 8 A.M. to 5 P.M. (EST). |  | 11 | 11 | Worldwide Distribution: USA including states of Florida, New Jersey, Maryland, Pennsylvania, South Carolina, South Dakota, and Virginia, and country of Finland. | Terminated | September 22, 2009 | 365 | 11
Z-0340-2009 | 50073 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilizer (Product Code 10101). The sterilizer includes software and hardware components. |  |  | October 17, 2008 | 2008 | Advanced Sterilization Products | 1) Inability of the sterilizer to detect when an injectiontakes place without hydrogen peroxide being transferred to the vaporizer bowl. This situation can result in a cycle completion with the injection of insufficient hydrogen peroxide sterilant and 2) Inability of the sterilizer to detect an obstructionin the door travel path while the door is closing. This situation can result in a hand b | Component design/selection | Customer notifications will be sent the week of November 24, 2008. The notification letter (URGENT: Product Correction) will advise of two situations that could cause the STERRAD 100S Sterilizer to malfunction. The first situation involves the failure of the STERRAD 100S System to detect when an injection takes place without hydrogen peroxide having been transferred to the vaporizer bowl. This can occur under combinations of the following conditions:- the pneumatic pressure is low on the machine, or- the injector valve has a small leak,AND- a cassette fails to advance, or- a used cassette is re-inserted into the sterilizer.This situation can result in a cycle completion with insufficient hydrogen peroxide. The second situation involves the failure of the STERRAD 100S System to detect an obstruction in the door travel path while the door is closing. Under some conditions, the door sensors may become mechanically unsecured/disconnected and will not detect an obstruction. This situationcan result in an operator's hand becoming pinched and bruised while the door is closing. To correct both situations above, Advanced Sterilization Products(ASP) has revised the software and is making physical hardware additions to the STERRAD 100S System.Contact ASP Customer Care Center at 888-783-7723, option 2 for questions or assistance. |  | 6324 | In the U.S., there are 2,623 units installed. There are also 3,701 units internationally. | Worldwide Distribution including USA and countries of France, Germany, Italy, UK, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Australia, Austria, China, Greece, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Middle East, Peru, Canada, Czech Republic, Portugal, South Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India and Slovenia. | Terminated | November 29, 2011 | 1138 | 6324
Z-0323-2009 | 50079 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Arjo Maxi 500 Patient Lift with 4 Point DPS (Dynamic Positioning System) and Scale; a Non-AC-Powered Patient Lift; Models KM560181; BHM Medical Inc., Magog, QC, Canada |  |  | October 20, 2008 | 2008 | Arjo, Inc. | The pivot bolt that attaches the hanger bar to the scale of the patient lift can break, resulting in the hanger bar falling. | Nonconforming Material/Component | Arjo sent an Urgent Device Recall Customer Notification letter dated 10/20/08 to all affected end users to the attention of the Administrator/Risk Manager via UPS 2nd day on the 10/20/08. The letters informed the accounts of the potential for the hanger bar to detach from the scale due to shearing of the pivot bolt. The accounts were advised to review the enclosed Safety Advisory Notice, note the affected model and serial numbers, and avoid using the lift under load on a high friction floor like carpet or soft tile while using the hanger bar to move the lift. The accounts were advised that an Arjo Field Service Technician would contact them within two weeks to schedule a visit to the account to perform repairs on their Maxi 500 lifts. Any questions were directed to Arjo Quality Department at 800-323-1245, ext. 6118. The accounts were also instructed to complete and fax back to Arjo the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, the numbers of affected units, and acknowledgment of receipt and understanding of the Urgent Device Correction. |  | 336 | 336 units | Nationwide Distribution. | Terminated | August 19, 2010 | 668 | 374
Z-0310-2009 | 46312 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 6 inch diameter Roller Pump; Model 801041. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The pumps may fail to power up or experience unplanned pump stops. | Process control | Consignees were advised of the problem via an Urgent Medical Device Correction: Safety Advisory letter dated 7/18/08 and each consignee was contacted by phone to assure receipt. The letter noted that the firm was not recommending that customers stop using the product. If customers found the risk unacceptable they were to contact Technical Service, 1-800-521-2818. Customers were to review the notice, assure all users were aware of the notice, and confirm receipt by completing the attached Customer Response Form and faxing it to the number indicated on the form. |  | 1429 | 1429 | Worldwide Distribution --- including USA and countries of Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Israel, Japan, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Vietnam. | Terminated | July 07, 2009 | 354 | 2548
Z-0122-2009 | 49584 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis dTA, Axiom Artis Modular Angiography X-Ray System, Model number 7008605. |  |  | September 02, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Incorrect contour finding: Image may be calibrated to the wrong size catheter. | Software design | Siemens issued a Customer Safety Advisory Notice dated 8/29/08 to inform customers of the potential issue. The recalling firm issued Update Instructions AX040/08/S to the service engineers which provides information to fix the problem. |  | 133 | 133 units | Nationwide. | Terminated | October 05, 2009 | 398 | 577
Z-0259-2009 | 49506 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | PREMISE, PART NO. 32813, 10 PACK UNIDOSE PREMISA A3 (INTERNATIONAL ONLY), dental composite |  |  | September 05, 2008 | 2008 | Kerr Corp | Material stiff: Material in several lots of Premise Unidose appears to stiffen and become difficult to extrude over time. | Component design/selection | Kerr sent an Urgent Medical Device Recall letter to consignees via USPS first class mail the week of September 8, 2008, requesting they return the product for replacement or credit.The recall notification requested the return of the affected products because the material appears to stiffen and become difficult to extrude over time. Customers were instructed if they have any of the affected lots of product in their inventory to immediately return it to Kerr. Any product returned will be replaced at no charge or credited to your account. In addition, customers were instructed to fax the attached Recall Return Form to 714-516-7739.If the customers were an authorized Kerr distributor, they were requested to identify and recover the affected product lots listed on the attached Lot Number Table that may have been shipped to their customers.Also, customers were instructed to contact Kerr Customer Care at 1 -800-537-7123 directly to handle the arrangements of a quick return and replacement |  | 22478 | 22,478 packs for all affected lots | Worldwide: USA, Australia, Canada, Hong Kong, Malaysia, Mexico, and Singapore. | Terminated | March 07, 2012 | 1279 | 22478
Z-0308-2009 | 49564 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Proven Cemented Semi-Constrained Total Knee. Part number SC1591. The size is 38 and the thickness is 10 mm. |  |  | September 03, 2008 | 2008 | Stelkast Co | Five (5) Three Peg Patella SC1591-29 were packaged in containers labeled for SC1591-38 sharing the same lot number. | Employee error | On 9/4/08, the recalling firm telephoned their customers to inform them of the problem and asked them to immediately recover all SC1591-38 with lot number 19906-052008. Once the product is in their possession, customers should contact Stelkast Customer Service for a Return Authorization (RA) Number prior to shipment to Stelkast. |  | 5 | 5 units | Nationwide Distribution --- including states of IA and WV. | Terminated | December 04, 2008 | 92 | 5
Z-0294-2009 | 49476 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Battery Caps used with Animas IR1250 Insulin Pump. |  |  | August 04, 2008 | 2008 | Animas Corporation | Unintentional rebooting: Pump products exhibit an intermittent loss of power due intermittent loss of contact between battery cap and battery canister resulting in the device resetting. The failure of the battery cap may result in failure of the device to administer insulin therapy which may result in hyperglycemia. | Nonconforming Material/Component | The recalling firm issued a notification letter to the patients with insulin pumps on 8/5/08 to inform them of the problem and that they needed to replace the battery. The recalling firm contacted OneTouch Ping users by telephone on 8/5/08 for this issue and another issue. The recalling firm issued letters to distributors informing them of the problem and notified customers to replace the batteries. In addition, a letter was issued to physicians to alert them of the situation with regard to current patients. Consumers can contact 1-800-423-4087 for additional information or questions. |  | 1624 | 1,624 units | Worldwide Distribution including USA (AR, CA, FL, KY, LA, MA, NE, OH, PA, SC, SD, TN, and UT) and countries of Australia, Austria, Canada, Czech Republic, Finland, France, Israel, Italy, New Zealand, Spain, Switzerland, and Sweden. | Terminated | February 04, 2010 | 549 | 7152
Z-0327-2009 | 50123 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Luminos dRF without ST filter. . Solid state x-ray imager (flat panel/digital imager). Model number 10252048. |  |  | September 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Issue may occur where an intended movement of the joystick may unintentionally initiate an x-ray exposure. | Nonconforming Material/Component | The recalling firm issued Update Instruction AX042/08/S dated August 2008, to affected customers. This Update instruction provided Siemens service engineers with instructions for the replacement of the joystick module at each facility. |  | 1 | 1 unit | Nationwide Distribution --- including states of KY, NJ and PA. | Terminated | September 23, 2009 | 369 | 4
Z-0415-2009 | 50149 | No | FDG | Gastroenterology-urology fiberoptic retractor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Fredrick's Converse Retractor #5 F/O, Model Number: 01-0485. |  |  | October 24, 2008 | 2008 | Biomet Microfixation, Inc. | Fiber optic cable is missing from the Frederick's Converse Retractor Lot # 042408D08. | Nonconforming Material/Component | Consignee was notified by certified mail via an Urgent Medical Device Recall Notice dated 10/24/08, with return receipt. The letter identifies the product and states the issue with the product. The letter required the customer to discontinue use of the item and return it to Biomet Microfixation so that they may rework the instrument. It also asked the customer to notify any hospital personnel responsible for receiving recall notices of this action if the product had been further distributed. The RGA number should be utilized to return the products. The Inventory Reconciliation sheet should be filled out and faxed to 904-741-9425. |  | 3 | 3 | Nationwide Distribution --- including states of Chicago, IL and La Crosse, WI. | Terminated | September 11, 2009 | 322 | 3
Z-0331-2009 | 50096 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Easy Core" Biopsy System, 18 ga. x 15 cm, REF 43-453 , M001434531, Sterilized with ethylene oxide gas. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific Corporation initiated a recall of its easy Core" BIOPSY SYSTEM due to difficulty cocking or arming the cannula latch on the device. This difficulty may result in an inability to use the device. | Device Design | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated October 21, 2008. The letter was addressed to The Hospital Field Action Contact. The letter described the problem and products being recalled. The letter stated that it was very important that the attached product table be used carefully and consider both inner and outer packaging product codes when searching for affected/recalled product as the UPN numbers on the inner and outer labeling are different. Requested to discontinue use of product, identify, and segregate and return affected products to BSC. Letter also requested to complete and return the Reply Verification Tracking Form. |  | 3808 | 3808 total for all products | Nationwide Distribution --- including states of AL, AZ, AR, CA, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT, NE , NV, NH, NH, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and PR. | Terminated | December 10, 2011 | 1159 | 3808
Z-0298-2009 | 48645 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | "LOCATOR Abutment for 5.7 Screw-Vent &amp; Compatibles." Zest Anchors, Inc. Product is sold for use with any Zimmer Dental 5.7 mm diameter Tapered Screw-Vent Implant. |  |  | April 09, 2008 | 2008 | Zest Anchors Inc | The recall was initiated after the firm discovered through a customer complaint that the Locator Abutments for 5.7 Screw-vent &amp; Compatibles do not fit properly in all sizes of the Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implants (TSV6B8, TSV6B10, TSV6B11, TSV6B13, TSV6B16, TSV6H8, TSV6H10, TSV6H11, TSV6H13 and TSV6H16). The Locator Abutment bottoms out in the pilot hole of these implants c | Nonconforming Material/Component | The recall notification letter (Recall Notice) was sent to consignees on April 23,2008. The letter includes the following information on how to handle the recall.It is very important that you determine the use of the Locator Abutment(s) as follows, mark the appropriate box, ,and return this letter to Zest Anchors, Inc. in the stamped, self-addressed envelope:1. The Locator Abutment(s) listed above WERE NOT used on any Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implants. THE FIT IS CORRECT AND NO ACTION IS NECESSARY.2. The Locator Abutment(s) listed above WERE used on a Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implant. PLEASE CONTACT OUR DENTAL OFFICE TO ARRANGE THE R PLACEMENT OF THE ABUTMENT WITH A NEWLY DESIGNED LOCATOR ABUTMENT. and3. We have current STOCK of the Locator Abutment(s) listed above with planned use for a Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implant. PLEASE CONTACT OUR DENTAL OFFICE TO ARRANGE THE REPLACEMENT OF THE ABUTMENT WITH A NEWLY DESIGNED LOCATOR ABUTMENT.For questions or additional information, contact Zest Anchors, Inc. at 800-262-2310 or 760-743-7744. |  | 996 | 996 units | Worldwide Distribution including USA and countries of Spain, Canada, Australia and South Africa. | Terminated | December 02, 2008 | 237 | 996
Z-0307-2009 | 49561 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Uni-CP Compression Forceps (spreader); |  |  | September 30, 2008 | 2008 | Integra LifeSciences Corp | Reports of the Uni-CP Compression Forceps breaking during the Compression of the Uni-CP plate | Nonconforming Material/Component | Recall Notice letters were sent to all Integra Sales Specialists on September 30, 2008 by first class mail. The letter informed them of the product and the problem concerning the product. The Sales Specialists were asked to return all inventories under their control with part number 339001ND from lots E6WS, E7UY, &amp; E86M. They are required to fill out the attached form identifying the lot number for all Newdeal Compression Forceps (Spreader) in their inventory along with the serial number(s) of their Uni-CP set(s). The form should be filled out even if the sales specialist does not have a Uni-CP. The form should be faxed or mailed back to Walt Ryan in Customer Service, 609-799-3297. A RMA number will be generated for the return of any affected inventory. A replacement will be sent and upon receipt, the forceps (spreader) should be replaced and returned to the address indicated on the letter. |  | 90 | 90 | Nationwide Distribution | Terminated | May 11, 2009 | 223 | 90
Z-0358-2009 | 46313 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Terumo Sarns Sternal Saw II System Power Unit, 115 V, Terumo Cardiovascular Systems Corp., Ann Arbor, MI; Catalog # 15670. |  |  | February 14, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The drive power cable cannot be fully inserted into the power unit or can only be inserted with difficulty, resulting in the device not being operational. | Nonconforming Material/Component | Consignees were notified via Urgent Medical Device Removal letter dated 2/14/08 to not use the device if it exhibited the problem, provided a response card, and stated that replacement motors will be shipped to the person listed on the returned response card. |  | 69 | 69 total | Worldwide Distribution --- including USA and countries of Belgium, China, Columbia, Malaysia, Philippines, United Arab Emirates, and Venezuela. | Terminated | July 07, 2009 | 509 | 69
Z-0441-2009 | 50078 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 5 F BardSelect 11 cm Percutaneous Catheter Introducer Set |  |  | October 07, 2008 | 2008 | Medtronic, Inc. | Failure to insert the guidewire through the 18 gauge introducer needle | Nonconforming Material/Component | Medtronic notified Medtronic Sales representatives via UPS for hand delivery of the Urgent Notification of Product Recall letter to the affected customers on 10/7/08 advising users to segregate the product so that it would not be used. The letter stated that a Medtronic Sales Rep would visit the site to account for all units that were received and would initiate a credit, as appropriate. Customers were requested to cooperate and assist with completing the necessary documentation. |  | 1830 | 1,830 | Worldwide Distribution --- USA including states of IL and FL, and countries of Belgium, Honduras, and Germany. | Terminated | June 07, 2011 | 973 | 6555
Z-0286-2009 | 49498 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Synthetic Leather and Biothane Waist and Wrist Restraints, Catalog Numbers: 2217SL and 2217B |  |  | July 25, 2008 | 2008 | J T Posey Company | Failure to restrain: This recall was initiated after a complaint was received by the firm regarding a Posey 2217SL unit indicating that a patient was able to pull the rivet or washer through the synthetic leather allowing the patient to have freedom of arm movement. A second complaint was received indicating that a patient was able to tear the strap at the adjustment notch on a Posey 221 7SL unit | Device Design | On July 29, 2008 Posey sent a Recall Letter to all affected customers. The recall letters were sent by certified mail with return receipt signature. The recall letter instructed the customers not to distribute or use the affected products. Any existing inventory subject to this recall was requested to be returned to Posey Company. Also, if they were a distributor, they were asked to please additionally forward a copy of the recall notice and reply form to each of their customers to whom you have further distributed the subject product. A Recall Fax Back form was provided with the recall letter for the affected facilities to fax back to inform Posey of their existing inventory.An RMA number was issured to each facility and a return UPS label was provided for any product that will be returned to Posey. Additionally, the sales representatives for each of the customers will be contacted by Posey Customer Service. |  | 145 | 145 units | Worldwide, USA, Australia, Canada, France &amp; The Netherlands. | Terminated | March 30, 2012 | 1344 | 145
Z-0484-2009 | 50071 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer, Product Code 10033 |  |  | October 07, 2008 | 2008 | Advanced Sterilization Products | ASP has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD NX Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. | Nonconforming Material/Component | The Urgent: Product Correction customer notifications, dated October 6, 2008, were hand delivered by field representatives starting October 7, 2008. The letter advises that Advanced Sterilization Products has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. A list of the affected sterilizer serial numbers is included with the letter. The letter advises that the Field Service Engineer has come to their facility to inspect and, if necessary, replace this power supply. Customers are asked to sign the letter acknowledging receipt of this letter upon completion of the work by the Field Service Engineer. |  | 79 | 32 potentially affected systems in the US - 20 installed, 5 uninstalled, 7 potentially affected. There are also 15 units internationally. | Worldwide Distribution --- USA including states of WV, CA, TN, AZ, WI, NC, MO, NJ, DE, GA, SC, WA, FL, MA, NM, MI, and AL and countries of Ecuador, China, Israel, Japan, Korea, Middle East, Australia, Czech Republic, Spain, Brazil, and Russia. | Terminated | November 07, 2011 | 1126 | 79
Z-0456-2009 | 49460 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly Quantitative Sentinel) System - Alert and Reminder software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies in the Alert and Reminder feature could result in a delay of treatment. When attempting to select the last visible alert or reminder choice, the next choice on the list below the desired choice is selected, and an inconsistent color may be displayed for the same clinical element across a set of work stations. | Software design | The software was updated to version 6.80.1 on 4/22/08, however, the consignee was notified of the software defects relating to the last visible Alerts and Reminders choice and the inconsistent display for the same clinical element across a set of work stations, via Product Safety Notification letter dated 8/29/08, sent via first class mail on 9/24/08. If there were any questions, the customer was to contact Tech Support at 1-800-433-2009. |  | 1 | 1 unit | Nationwide Distribution --- state of Ohio. | Terminated | December 03, 2008 | 69 | 1
Z-0457-2009 | 49468 | No | JFP | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL7XX Calibration 1 Solution S1720. P/N S1720, 944-024 |  |  | September 05, 2008 | 2008 | Radiometer America Inc | The firm found that for this particular lot that the barcode, used for entering the characteristics of CAL1 Solution into the analyzer, does not reflect the actual values of the solution. As a consequence, the calibration curves for these parameters and thereby the measured results for patient samples and quality controls will be biased. | Process control | On, 9/2/2008 the firm sent out recall notification letters requesting their customers discard any Calibration 1 Solution (CAL1) S1720 (part number 944-024), lot WA-04, complete the attached return form and fax it to 800-736-0601. The firm will send out replacements for all affected product that the firm shows was shipped to their location. Further questions could be directed to Technical Support at 800-736-0600, option 2. |  | 3200 | 3,200/200ml bottles | Worldwide Distribution --- USA including states of AZ, CA, FL, IA, ID, IL, IN, KY, LA, MD, NC, ND, NE, NY, OH, OK, PA, SD, TX, VA, WA, and WI, and country of Canada. | Terminated | April 25, 2012 | 1328 | 3200
Z-0324-2009 | 50085 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo MultiModality WorkPlace (MM WP). Software product(Model Number 10140720). |  |  | September 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Values derived from Dynamic CT data sets may be incorrect. The corresponding images are not affected. | Software design | The recalling firm has issued a software update to affected customers via Update Instructions CT034/08/S. This software update corrects the values from the volume perfusion post processing application. A firm representative will visit each consignee to install the software update beginning 9/12/08. |  | 130 | 130 units | Nationwide Distribution. | Terminated | September 23, 2009 | 376 | 130
Z-0408-2009 | 50081 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PALINDROME EMERALD SPORT PACK Cuffed Dual Lumen Catheter with Pre-curved Shaft and Heparin Coating, REF 8888123411, Size 14.5Fr/Ch (4.85mm) x 23cm. TYCO HEALTHCARE GROUP LP, Mansfield, MA 02048. --- NOTE: Both the "sport pack" and "kit" package configurations include the catheter, Venetrac insertion stylets, pull apart sheath, and other accessories necessary for placement. The Venetrac insertion stylet is an optional accessory. The 510k for the Venetrac accessory is K051584. The 510k for the catheter is K060509. |  |  | September 19, 2008 | 2008 | Covidien | Incorrect size of the Venetrac stylets (an optional accessory). They are too short and cannot be used. | Mixed-up of materials/components | Starting on 9/19/08, Covidien Sales Representatives visited each of the hospitals that received the catheters, informed the staff that the Venetrac insertion stylets were too short, and placed pink corrective sticker labels on any unused product. Should you require additional information, please contact jim.welsh@covidien.com or 508-216-8532. |  | 10 | 10 units | Nationwide distribution including states of OH, WI, LA, TX, UT, CA, OR, and WA. | Terminated | December 04, 2008 | 76 | 117
Z-0319-2009 | 49713 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | AS3000 Anesthesia Delivery System. Datascope Patient Monitoring; |  |  | October 07, 2008 | 2008 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Two issues have been identified with the AS3000 unit.1. The threadlocker of the caster (wheels) may not have been utilized, which may allow the caster to loosen and possibly separate from the unit.2. The use of select brands of pre-pack absorber in the absorber canister of the AS3000 has been associated, in some cases, with gas leakage around the pre-pack, rather than through the absorber mate | Component design/selection | Notification (Urgent Product Field Correction) letters were sent to the one consignee on October 7, 2008 by Federal Express overnight delivery. |  | 2 | 2 units | Only 2 units are in distribution, both in the same facility in Redondo Beach, CA. | Terminated | December 08, 2008 | 62 | 2
Z-0318-2009 | 49572 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Toddler Chair Hip Strap (Belt). The hip belt is made with nylon webbing and a plastic buckle. 510(k) exempt, Medical Device Listing # R061877. |  |  | September 12, 2008 | 2008 | Community Products, LLC | Some hip straps were assembled incorrectly by threading the buckle on the strap upside down. If a child were left unattended in spite of the warnings on the product and in the manual, a defective hip strap could loosen, allowing the child to slip down in the chair and this could create a choking hazard. | Process control | On 9/12/08, Rifton Equipment began notifying the customers by phone using a phone script. This phone call notified the customers of the recall and also served as an effectiveness check that they were notified. The customers were instructed to remove the product from use until the replacement kit is received and the old strap is returned. The replacement kit contains instructions with ID codes and a place for the customer to verify that they have made the replacement as instructed. It also contains the required number of new hip straps. A verification check was sent with the replacement straps, which the customer will sign and return to the firm verifying that they have replaced the straps for the affected serial numbers. Rifton Equipment is providing the customer with return shipping labels which were sent by UPS 2nd day air. The replacement straps use black thread instead of tan thread for the sewing on the ends of the straps so that the firm can be sure that the returned straps are not the new replacement straps.For questions or additional information, contact Rifton Equipment at 1-800-571-8198. |  | 147 | Potentially 147 units may be affected. | Worldwide Distribution including USA and countries of Spain and Great Britain. | Terminated | March 15, 2010 | 549 | 147
Z-0312-2009 | 49569 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis System dFC using software version VB31D with 30x40 cm2 flat detectors. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | System may switch to emergency fluoroscopy mode. | Software design | A Customer Safety Advisory Letter dated 8/20/08 was issued with Update Instructions AX045/08/S. The letter informed the customer of the potential issue and provided instruction to avoid its occurrence. A software update (VB31E) was released to correct the issue. |  | 4 | 4 units | Nationwide Distribution including states of AR, CA, CO, FL, IA, IL, IN, KS, KY, MI, NY, TX, VA, and WY. | Terminated | January 26, 2010 | 519 | 21
Z-0414-2009 | 50105 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel Healthcare RT240 Adult Breathing Circuit Kit. |  |  | October 22, 2008 | 2008 | Fisher &amp; Paykel Healthcare Inc | The recall was initiated because certain lot dates of the RT240 Adult Breathing Circuit Kits manufactured on or before June 27, 2008 include a heated breathing circuit that may be more susceptible to damage when used in excess of the specified seven-day maximum duration of use, which may increase the risk of malfunction or fire. | Process change control | Fisher &amp; Paykel Healthcare, Inc. (FPH) issued an URGENT-PRODUCT RECALL letter (dated October 8, 2008) including a Customer recall letter for consignees/customers to Cardinal Health. Cardinal Health will, in turn, send the Customer recall letter dated October 10, 2008 to its consignees/customers. Fisher &amp; Paykel Healthcare (FPH) also issued a similar set of recall notification letters to Norco on October 15, 2008.Recipients of the recall letter must perform the following:1) Check all Cardinal Health stock of the product for the affected lots, 2) Immediately quarantine the affected product, 3) Complete a Device Recall Response Sheet, and 4) Fax or e-mail the completed Device Recall Response Sheet. Fisher &amp; Paykel Healthcare, Inc.'s "Dear Customer Letter" for Cardinal Health and Norco instructs consignees to complete a customer response (i.e. fax enclosed Customer Acknowledgement Form) to FPH. The customer letter also provides instructions for product return.Fisher &amp; Paykel Healthcare, Inc. recall communication instructions to Cardinal Health and Norco was for all affected product to be returned and quarantined by the distributors. The distributors will return any affected stock to Fisher &amp; Paykel Healthcare, Inc. for documented destruction. |  | 229320 | 229,320 units | Worldwide Distribution including US and country of Canada. | Terminated | March 26, 2012 | 1251 | 229320
Z-0343-2009 | 50179 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | AMO Healon D Ophthalmic Viscosurgical Device, Product Code: 10204011, Lot Number: UD30654, 30 mg/mL fill size |  |  | October 29, 2008 | 2008 | Advanced Medical Optics, Inc. | Endotoxin levels above specifications have been noted in some syringes tested for lot number UD30654 of the AMO Healon D Ophthalmic Viscosurgical Device. Endotoxin levels above the maximum USP level may be potential causes of an inflammatory response and/or TASS in patients following surgery. | Process design | On Thursday, October 30, 2008, AMO began distribution of an Urgent Device Recall Notice via Federal Express overnight delivery to the twenty-six (26) AMO Healon D OVD, Lot UD30654, customer accounts in the United States. The Urgent Device Recall Notice include&amp; a facsimile reply form.The Recall Notice informed the customers that AMO has received reports that some physicians using lot number UD30654 of the AMO Healon D OVD have observed an inflammatory reaction in their patients at the one-day post-operative visit. The customers were asked to undertake the following actions:(1) STOP using and remove from your inventory all units of AMO Healon D OVD, Lot UD30654. (2) Complete the included Facsimile Form making sure to note on the Facsimile Form the Quantities of AMO Healon D OVD present in your facility to be returned to AMO and fax to AMO Customer Service at 714-247-8722 within the next three (3) business days.The customers were also informed that an AMO Customer Service Representative will be contacting them to assist with the return of product affected by this action. If they have AMO Healon DC OVD units to be returned as part of this recall event and have not been contacted by an AMO Customer Service Representative, please call our toll free number at 1-877-AMO-4LIFE (1-877-266-4543) so that we can arrange the return of the product. |  | 1450 | 1450 | Nationwide to: MA, NJ, NY, WV, NC, SC, KY, OH, IN, MI, MN, OK, TX, WY OR, WA and Puerto Rico | Terminated | February 21, 2012 | 1210 | 1450
Z-0311-2009 | 46314 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Sarns Level Sensor II Pads. For use with Sarns Advanced Perfusion System I, Perfusion System 800 and Perfusion System 900. Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; REF 195240. |  |  | January 13, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The level sensor holder may detach from the adhesive pad causing the sensor to lose contact with the venous reservoir, which may result in a detached sensor indication or false alarm. | Nonconforming Material/Component | Consignees were notified via recall letter (Urgent Medical Device Removal letter) dated 2/13/08 and requested to report stocks on hand so that replacements could be sent. |  | 31980 | 31,980 pads (25,620 to U.S. and 6,360 international) | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Japan, Saudi Arabia, Singapore and Yemen. | Terminated | July 07, 2009 | 541 | 31980
Z-0354-2009 | 44839 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 (4 inch diameter Roller Pump; Model 801040). |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | Overspeed and underspeed errors, erratic pump behavior, jerky operation at low RPMs, pump instability, pump slowdowns and pump stops due to overspeed max events may occur due to a grease leak onto the motor speed encoder disk. | Component design/selection | Consignees were notified of the problem via recall letter dated 7/18/08. The firm began conducting field upgrades on the roller pumps in 07/08. For questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 1834 | 1,834 | Worldwide Distribution including USA and countries of Australia, Belgium. Canada, Columbia, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 5116
Z-0447-2009 | 48470 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Shoulder Body Assembly, Left (SH-1540L-S) and Modular Shoulder Body Assembly, Right (SH-1540R-S). Product labeled in part, "Acumed 5885 NW Cornelius Pass Road Hillsboro, OR 97124-9432 " |  |  | May 13, 2008 | 2008 | Acumed LLC | The packages for Modular Shoulder Body Assembly Left (SH-1540L-S) and Modular Shoulder Body Assembly Right (SH-1540R-S) may contain assemblies for the opposite side as indicated. | Labeling mix-ups | The firm issued a Urgent Notice: Device Recall on 5/13/08. The firm notified consignees of the issue and asked them to return the product to Acumed with the enclosed product recall effectiveness form for customers. For further information contact Acumed customer service at 1-888-627-9957 |  | 20 | 20 units | World-wide distribution including states of GA, MD, CA, MO, and AL, countries of United Kingdom, South Korea, South Africa, and Spain. | Terminated | May 13, 2010 | 730 | 20
Z-0407-2009 | 49522 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Mentor Aseptic Transfer Set, Cat. #350-8400 (contains BD 60mL Luer-Lok Syringe). Aseptic transfer set manufactured by Mentor Texas, LP, Irving, TX 75038 (BD 60 mL Luer-Lok syringe manufactured by Becton Dickinson, Franklin Lakes, NJ 07417). |  |  | August 22, 2008 | 2008 | Mentor Texas, Inc | The Mentor Aseptic Transfer Set contains a component, the BD 60mL Luer-Lok Syringe, which is under recall by Becton Dickinson due to a package integrity issue. | Package design/selection | Firm sent notification, along with a copy of the Becton Dickinson recall letter, to consignees on 08/18/08. Consignees asked to inspect inventory of Mentor Aseptic Transfer Sets for affected BD 60mL syringes and to complete and return the included Device Accountability Worksheet. For additional information, please contact Mentor Customer Service at 800-235-5731. |  | 32480 | 32,480 Aseptic Transfer Sets (4 syringes/set) | Worldwide Distribution; product sent to all states (except Alaska), Washington DC and Puerto Rico. Product also sent to the following countries: Australia, Canada, Costa Rica, Korea, New Zealand and South Africa. | Terminated | August 26, 2010 | 734 | 32480
Z-0458-2009 | 50065 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Gravity Compensating Accessory; |  |  | October 03, 2008 | 2008 | Integra LifeSciences Corp. | Integra NeuroSciences has become aware that certain Gravity Compensating Accessory implants have the potential for CSF leakage under certain conditions. | Process control | Urgent: Product Recall Notification letters, dated 10/3/08, were sent out on October 9, 2008. Letters explain the reason for the recall and the catalog and lot numbers affected. If the product has been implanted, it is recommended that the patient be monitored periodically using the standard of care determined appropriate by the patient's attending physician. If the product is found in inventory, customers are asked to quarantine the product and contact their Integra Sales Rep for assistance and for arrangements for product return. A Recall Acknowledgement and Return Form is to be completed and returned to Integra for reconciliation purposes. |  | 49 | 49 units | Worldwide Distribution - MI, IL, NY, IN, CA, IN and NC and Carnegie, Victoria, Melbourne, Australia. | Terminated | October 06, 2009 | 368 | 49
Z-0418-2009 | 49523 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dual Port Wizard Low Profile Replacement Gastrostomy Device, Code 00202W, 16F, 2.4cm, 10cc, Sterile EO, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | September 23, 2008 | 2008 | C R Bard Inc | Gastrostomy device anti-reflux valve may allow leakage from the stomach. | Component design/selection | Customers notified by letter (Product Field Action letter) on September 23, 2008 and instructed to remove and return any affected product to Bard Access Systems, Inc (BAS). Customers were instructed in the recall letter to respond by completing and faxing the enclosed Reply Form and Inventory Reconciliation Form. For assistance, contact BAS customer service at 800-290-1689. |  | 1399 | 1399 units | Worldwide Distribution including US, Latin America and countries of Canada, Japan, Finland, Norway, Sweden, Austria, Belgium, Holland, Great Britain, Italy, France, Switzerland, Spain, Germany. | Terminated | November 08, 2011 | 1141 | 20263
Z-0460-2009 | 49458 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies result in patient safety issues involving patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the curr | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation customers via letter (Product Safety Notification) dated 9/25/08 of the software anomalies resulting in patient safety issues involving the patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the current exam as a historical study and vice versa. The letter provided the users with safety instructions to follow until the software is updated. Direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-277-5240. |  | 546 | 546 units | Worldwide Distribution including USA and countries of Australia, Belgium, Brazil, Canada, Cayman Islands, Chile, Denmark, Ecuador, France, Germany, Hong Kong, Italy, Japan, Korea, Malaysia, Malta, Mexico, New Zealand, Singapore, Spain, Switzerland, Taiwan, United Kingdom and Venezuela. | Terminated | September 06, 2010 | 711 | 546
Z-0485-2009 | 50075 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Enzymatic Creatinine Flex reagent cartridge, DF270. |  |  | September 29, 2008 | 2008 | Dade Behring, Inc. | Reagent may exhibit unflagged inaccurate patient sample results. | Component design/selection | The recalling firm telephoned all customers on 9/29/08 to inform them of the problem. Customers were issued a letter dated 9/29/08 (and a response form) with instructions. Customers were instructed to complete the response form and fax (302-631-7242) it to Siemens Healthcare Diagnostics. For questions or additional information, contact the Siemens Technical Solutions Center at 800-441-9250. |  | 12773 | 12,773 cartons | Nationwide Distribution. | Terminated | March 27, 2009 | 179 | 12773
Z-0483-2009 | 49586 | No | LTD |  |  | Toxicology | Contact ODE | RNAgentsÂż Total RNA Isolation System (Cat. # Z5110). Kit Box label/Box drawer label. For Laboratory Use. Promega Corporation, Madison, WI. USA. |  |  | June 04, 2008 | 2008 | Promega Corpopration | Promega Corporation received reports of leaking bottles of Phenol:Chloroform:Isoamyl Alcohol (Part # Z5112), a component of the RNAgentsÂż Total RNA Isolation System (Catalog # Z5110). Promega product Z5112 is a mixture of phenol, chloroform and isoamyl alcohol. If phenol contacts skin, it can cause burns and poisoning. The fumes of this product are also harmful. | Packaging | Consignees were sent a "Promega Urgent Product Alert-Risk of Injury" letter dated June 9, 2008. The letter described the product, problem and provided instructions on handling, inspection and disposition of product. In addition, E-mail notifications were sent to distributors and branch offices on May 29,2008. For questions and additional information, contact your local Promega representative or Promega Technical Service at 800-356-9526 and press "4", or e-mail at techserv@promega.com. |  | 836 | 836 | Worldwide Distribution: USA including states of IN, ME, NE, NY, WI, PA, NC, FL, OR, UT, MI, IL, AL, OR, GA, CA, MD, TX, MA, UT, MO, OH, and LA, and countries of Canada, Mexico, Brasil, China, Japan, Germany, Singapore, Great Britain, Australia and France. | Terminated | June 26, 2012 | 1483 | 836
Z-0454-2009 | 48660 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Vidiera NsP Nucleic Sample Preparation, Part Number: A22421, Version 1.0.41 |  |  | August 30, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: under certain circumstances, when transfer volume is not within the set range, the Vidiera NsP software does not report sample exclusion and does not flag the excluded sample in the "Run Results" report. The software error associated with the excluded sample appears as "Liquid was not successfully transferred to well#..........the sample will be excluded from further proce | Software design | The customer notifications were initiated by Beckman Coulter on September 11, 2007 when all affected accounts were contacted via telephone, followed by a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers of the issue and instructed them to:(1) Review the "run Summary" event log for your samples. (2) Review the attached instructions describing the steps for identifying the samples affected by this error. (3) Complete and return the enclosed response form within 10 days so that we may maintain our records.The letter informed the customers that Beckman Coulter would provide a software update within 60 days. The letter also asked their customers to share the PCA with their staff and provide a copy of the PCA to any customers that they may have forwarded the affected products to.For questions regarding the PCA, please contact firm at 714-961-3634. |  | 6 | 6 | Nationwide distribution including states of PA, MI, UT, WA, IL, and CA. | Terminated | January 13, 2009 | 502 | 6
Z-0338-2009 | 49420 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Boston Scientific iLab Ultrasound Imaging System, models 120INS and 240INS. Its intended use is for ultrasound examination of intravascular pathology. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | September 05, 2008 | 2008 | Boston Scientific Corporation | Computer malware: Products may be infected with a worm (computer malware) which could infect a computer network to which it may be connected. | Process control | Customer letters sent on Friday, Sept 5 via FedEx, 2008 to all customers. This letter will be followed up by a customer service visit. Package delivery will be tracked, as well as Reply Verification tracking forms.For further information, please contact firm at 508-652-5594. |  | 357 | 357 units | World-wide distribution including USA and countries of Europe, Middle East, Africa, and Japan. | Terminated | May 18, 2009 | 255 | 357
Z-0504-2009 | 50157 | [6 empty] | Single-use sterile puncture attachment UA1256-U used with the BK Medical Ultrasounds: |  |  | September 08, 2008 | 2008 | Bk Medical | No 510k approved | Other | BK Medical notified accounts via a Recall Notice letter dated 9/03/08 to inform customers of the unapproved status and to remove the device from use and it will be replaced with B-K Medical's reusable metal guide UA1256. Customers were to fill out the attached "Reply Form" and return it via fax or the postage paid envelope. Any questions should be directed to B-K Medical Systems, Inc. at 1-800-876-7226. |  | 664 | 664 pieces | Nationwide Distribution --- including states of AR, CA, CO, CT, IL,IN, OK. OH, MA, MS, MD, NC, and NY. | Terminated | January 29, 2016 | 2699 | 664
Z-0503-2009 | 50151 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Move, Model NDA0200-20 Batteries Pack; BHM Medical Inc., 2001 Tanguay, Magog, QC, Canada J1X 5Y5 |  |  | November 05, 2008 | 2008 | Arjo, Inc. | The battery pack NDA0200-20, lot 070826D, sold with the Maxi Move patient lift or as replacement parts, has a defective connector that could lead to the inability of the battery pack to recharge and/or short circuit, resulting in smoke emission. | Component design/selection | Arjo sent the Urgent Device Recall Customer Notification dated 11/5/08 and Field Safety Notice (FSN) dated 10/31/08 to all affected Maxi Move customers on 11/5/08. The accounts were informed of the problem with the connector of the battery pack NDA0200-20, lot 070826D, sold with the Maxi Move patient lift (model numbers starting with KMC)or as replacement parts. They instructed the accounts to discontinue use of the battery pack if it does not recharge or has an intermittent electrical contact. The accounts were requested to complete and return to Arjo Inc., the enclosed customer response form indicating that they have received the notification package. The accounts will be contacted by an Arjo representative within a week to arrange the replacement of the battery pack and corresponding connectors (male part/plug) on the lift and battery charger. |  | 273 | 273 battery packs | Nationwide Distribution | Terminated | January 05, 2010 | 426 | 273
Z-0482-2009 | 49563 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ACE and ACE Alera Hemoglobin A1C : EZA1C Reagent, EZA1C Controls, and EZA1C Calibrators. |  |  | August 22, 2008 | 2008 | Alfa Wassermann, Inc. | Internal studies observed occasional unexpected outlier results in whole blood samples tested for Hemoglobin A1c using the EZA1c reagent on the ACE or ACE Alera clinical chemistry systems. | Component design/selection | An initial URGENT: CUSTOMER ADVISORY NOTICE dated August 22, 2008 was sent to all customers and distributors. The letter stated the problem and asked customers to cease using the EZA1C reagent, controls and calibrators. Customers were asked to hold any affected inventory while additional testing was conducted by Alfa Wassermann Diagnostic Technologies, LLC. (AWDT). A second communication, URGENT: CUSTOMER ADVISORY NOTICE (and Product Response Form) dated September 18, 2008 was sent instructing customers to destroy all product and package inserts in stock for the EZA1C products listed in the letter. Information was provided on receiving credit and an alternative product for use. Customers are requested to complete and return the Product Response Form by fax (888-646-2535) to AWDT indicating their receipt of the letter and the amount of product destroyed. Distributors' letters requested their customers be notified of the recall and that the Product Response Form be sent directly to AWDT. Direct questions to AWDT Customer Solution Center at 866-419-2532. |  | 69 | Reagent Lot F2481 = 264 kits and Reagent Lot F2603 = 69 kits. (Control Lot HBC00507-1=322 kits and Calibrator Lot HBS00907-1=262 kits). | Nationwide Distribution. | Terminated | April 15, 2009 | 236 | 69
Z-0412-2009 | 50098 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Arjo Maxi 500 Patient Lift with 2-Point Spreader Bar and Scale. A Non-AC-Powered Patient Lift. BHM Medical Inc., Magog, QC, Canada. Model KM560101, and Model KM560001 - Maxi 500 with 2-Point Spreader Bar, if it was equipped with the optional scale kit 700.05505. |  |  | October 30, 2008 | 2008 | Arjo, Inc. | If the spring pin is not properly reinstalled after maintenance, the pivot bolt could unscrew by itself within a limited period of time, resulting in a hanger bar detachment. | Employee error | Arjo, Inc. sent an Urgent Device Recall Customer Notification letter dated 10/29/08 to all affected end users to the attention of the Administrator/Risk Manager via UPS 2nd day on the 10/30/08. The letters informed the accounts of the potential for the hanger bar to detach from the scale if the spring pin is not replaced into the pivot bolt after performing maintenance on the lift. The accounts were advised to review the enclosed Urgent Field Safety Notice, note the affected model and serial numbers, and inspect their lifts for the presence of the spring pin in the pivot bolt. If the user finds that the spring pin is missing, they are instructed to immediately stop using the lift and contact their Arjo representative. The accounts were also instructed to remove previous versions of the Operating and Product Care instructions, and replace them with the current version 001.20815.EN Rev3 dated October 2008 provided with the letter, ensuring that the proper personnel are made aware of the changes outlined in the Field Safety Notice. Direct questions to the Arjo, Inc. Quality Department at 800-323-1245, ext. 6118. The accounts were also instructed to complete and fax (888-594-2756) to Arjo, Inc. the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, the numbers of affected units, and acknowledgment of receipt and understanding of the Field Safety Notice. |  | 46 | 44 lifts, 2 scale kits | Nationwide Distribution. | Terminated | August 19, 2010 | 658 | 427
Z-0481-2009 | 49562 | No | LDJ | Cannabinoid test system. | Toxicology | Toxicology | 510(k) | Cannabinoids Intercept Micro Plate EIA 100 plate kit |  |  | September 08, 2008 | 2008 | OraSure Technologies, Inc. | Readings, high: intermittent high absorbance readings. For further information, please contact the firm at 610-882-1820. | Other | On 9/8/08, the recalling firm telephoned the customer to inform them of the problem. As a follow up to the telephone call, the recalling firm issued a recall letter dated 9/11/08 to the customer. |  | 1 | 1 kit | The product was shipped to one customer in New York. | Terminated | December 18, 2008 | 101 | 1
Z-0498-2009 | 50103 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | FC4000 MicroTargeting Electrode Kit 5x |  |  | October 20, 2008 | 2008 | FHC, Inc. | Mislabeled: May contain an electrode longer than labeled size | Pending | FHC notified customers via an Urgent: Medical Device Recall letter dated 10/20/08 advising users to remove the recalled electrodes from inventory and set the product aside for inspection by your FHC representative. Customers were asked to document their actions by completing the enclosed form and faxing it back to FHC at 207-666-8539. |  | 3 | 3 Kits (30 electrodes-10 electrodes/kit) | Nationwide Distribution --- including states of CA, CO, FL, LA, NM, PA, TN, and PR. | Terminated | May 02, 2012 | 1290 | 62
Z-0502-2009 | 50125 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA Neoplastine CI Plus 10; Product Catalogue Number: 0667 |  |  | October 20, 2008 | 2008 | Diagnostica Stago, Inc. | Internal investigation has determined the potential lack of homogeneity between product vials. | Process control | Product Recall Notice letters were sent to all customers on October 20, 2008 vis US mail. The letter stated the reason for the notice, the corrective action, and replacement process. Customers are to confirm that they received the Recall Notice by completing the attached "Product Recall Response Form" and return by fax to the Regulatory Affairs Department at 973-867-6028 or 973-695-0095. The firm is prepared to immediately implement the following corrective action: 1) perform QC testing in accordance to directions from the package insert on each vial and destroy vial and replace inventory if determined product was unsatisfactory OR 2) return inventory for replacement with new lots of inventory. If customers have any questions or need any technical assistance, they are to call the DSI Hotline at 1-800-222-2624, option 1. |  | 31207 | 31,207 units | Nationwide Distribution. | Terminated | June 24, 2010 | 612 | 32487
Z-0509-2009 | 50172 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Custom Cranial Implant Kit, small. Distributed by Stryker Craniomaxillofacial, Kalamazoo, MI; Catalog no. 54-00101. |  |  | October 24, 2008 | 2008 | Stryker Craniomaxillofacial Division | Lack of assurance of sterility. | Process control | The firm's sales representatives and consignee risk managers, as well as implanting surgeons, were notified individually of the problem via recall letter dated 10/24/08, advised of the risk for serious infections, and instructed to follow-up with patients for infections for at least six months after surgery. The recall letter also included details for identifying and returning any remaining implant kits. Direct questions to Stryker's RA/QA Representative 269-324-5346 ext. 4258. The firm issued a press release concerning this recall on 12/23/08. |  | 322 | 322 of all products. | Nationwide Distribution. | Terminated | December 31, 2009 | 433 | 322
Z-0496-2009 | 50091 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Heartmate II Left Ventricular Assist System (LVAS). Model/Catalog Number 1355 and Model/Catalog Number 102139. |  |  | October 24, 2008 | 2008 | Thoratec Corp | The firm has become aware that, over time, wear and fatigue of the percutaneous lead connecting the HeartMate II Left Ventricular Assist System (HM II LVAS) blood pump with external system controller may result in damage that has the potential to interrupt pump function and may require a reoperation to replace the pump. | Device Design | An initial "Urgent: Medical Device Correction" letter dated October 21, 2008 and a follow-up letter dated October 24, 2008 were sent to affected HeartMate II Left Ventricular Assist System (HM II LVAS) customers. The letters identified the problem, immediate action and preventive action to be taken by customers and recommended that the pump be replaced as soon as possible if damage to the percutaneous lead is confirmed. If customers suspect that a HM II LVAS patient may have a damaged percutaneous lead, please contact Thoratec Technical Services (800-456-1477) for assistance. A return response (Acknowledgement Form) was included with the letter for consignees to complete and fax it to Thoratec Regulatory Affairs at 925-847-8628. Direct questions to the Thoratec Regulatory Affairs Department at 925-730-4135 (Monday-Friday, 8 a.m.-5 p.m. PST). Clinicians and patients with questions may contact the firm at 800-528-2577, or if calling outside the USA, 925-847-8600 (7 days a week, 8 a.m.-5 p.m. PST). Copies of the press release may be found on the firm's website, www.thoratec.com, under Investor Relations/Press Releases. |  | 2335 | 2,335 units | Worldwide Distribution. | Terminated | January 26, 2011 | 824 | 2335
Z-0500-2009 | 50115 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | FastPack(R) Total PSA Immunoassay; Catalog #: 25000001 |  |  | September 29, 2008 | 2008 | Qualigen Inc | The recall was initiated after the firm investigation of customer complaints of falsely elevated results determined the product was not meeting product performance expectations. | Other | The recall was initiated with Qualigen faxing a copy of the Product Notification letter, dated 9/29/2008, to all customers who have received any sublot of the affected product . Faxes were sent beginning 9/29/08. The recall letter informed the customers that the affected product was not meeting the firm's expectations for product performance. Customers were instructed to please stop using the affected lots immediately as they have been noted to produce falsely elevated results and they should confirm any result that appears to be inconsistent with "a patient's medical history, clinical examination, and other findings". In addition, they may wish to retest any patients that were run with a different FastPack lot.Beginning on 9/30/08, all customers were contacted by phone to confirm that they received the recall letter.The customers were shipped replacement products at no cost and as part of the product replacement process, they were instructed to return any unused Total PSA FastPacks/Kits of lot 0807002 to Qualigen. |  | 720 | 720 kits | Worldwide Distribution --- including USA and countries of Canada, UK , Italy, Switzerland, Sweden, &amp; South Korea. | Terminated | June 02, 2011 | 976 | 720
Z-0593-2009 | 50108 | No | MQE | Assisted reproduction needles. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Oocyte Recovery Needles,Oocyte Recovery sets, 750 mm tube with flushing connector, Product Number: KÂżLN-750, and KÂżLN-750. |  |  | October 23, 2008 | 2008 | Irvine Scientific Sales Co Inc | Smiths Medical International became aware that certain lots of Wallace oocyte Recovery Sets have a damaged needle tip. | Packaging | The recall was initiated by Smiths Medical on October 22, 2008 with the forwarding a recall notification to Irvine Scientific. Smiths Medical International sells the products to Irvine Scientific for distribution in the U.S. Irvine Scientific is the U.S. Agent. Irvine Scientific issued an Urgent Field Safety Notice Letter, dated 10/22/2008, with attached fax-back Customer Response form to their customers via US Postal service beginning on October 24, 2008. Irvine Scientific requested their customers to complete an inventory summary on the Customer Response form and contact the firm to arrange for the return of any remaining product in their inventory to the firm. |  | 71 | 71 units | Nationwide Distribution --- including states of CA, FL, PA, IL, OR &amp; UT. | Terminated | September 02, 2011 | 1044 | 71
Z-0486-2009 | 50086 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | FHC microTargeting Platform DBS Measuring Fixture, |  |  | October 08, 2008 | 2008 | FHC, Inc. | Measuring fixture is incorrectly graduated. | Process control | FHC provided letter notification to FHC Field Representatives on 10/08/08 to hand deliver to customers. FHC Field Representatives will remove incorrectly marked product and provide replacement product. Direct questions to an FHC representative or FHC Quality and Regulatory Affairs at 207-666-5651, ext. 2210. |  | 5 | 5 units | Nationwide Distribution including states of CO, FL, NC and NM. | Terminated | May 02, 2012 | 1302 | 5
Z-0524-2009 | 49077 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Progesterone (LKPG 1,5) |  |  | July 21, 2008 | 2008 | Siemens Medical Solutions Diagnostics | Siemens Healthcare Diagnostics has confirmed a long-standing high bias with LKPG 1, 5. (The bias is approximately 25% at 10 - 20 ng/mL as opposed to approximately 3% at 1-3 ng/mL - for more detail see the bias plot that is included with this notice.). | Other | Siemens initiated the recall on 07/22/2008 by forwarding A Customer Bulletin by fax or mail to all affected Siemens Healthcare Diagnostic Branches for communication with the affected customers.The Customer Bulletin informed their customers that the firm is conducting a field correction and requested that they discontinue using any IMMULITE/IMMULITE 1000 kit lot prior to LKPG 1,5 330. They were also instructed to complete attached Completion Notification Form and return to their local Siemens Healthcare office or distributor as soon as possible. The were informed that replacement kits will be provided once the completed forms have been received.The Customer Bulletin recommend their customer discuss the content of the letter with their laboratory director regarding the need to review previous test results, conduct patient follow up, and/or repeat testing. The bias should not impact the in-vitro fertilization clinical application because the average change for samples within the range of 1-3 ng/mL is less than 5%. Although there is a greater degree of shift at higher concentrations, the clinical impact is considered to be relatively low when using the test for the diagnosis of ectopic pregnancy and luteal phase defects because in these situations, the progesterone result should always be considered in conjunction with other tests. PROGESTERONE DETERMINATIONS USED IN THE DIAGNOSIS OF PLACENTAL INSUFFICIENCIES OR FOR MONITORING PROGESTERONE THERAPY MAY HAVE BEEN AFFECTED BY THE HIGH BIAS.The letter also indicated SIEMENS IS INTRODUCING A RE-CALIBRATED IMMULITE PROGESTERONE KIT BEGINNING WITH KIT LOT LKPG1, 5 300, WHICH WILL BEGIN ROUTINELY SHIPPING AUGUST 15 TH. AT THAT TIME, KITS WITH THE PREVIOUS CALIBRATION WILL ONLY BE AVAILABLE UPON REQUEST AND AVAILABILITY OF INVENTORY FOR CUSTOMERS WHO STILL NEED TO PERFORM BASELINING STUDIES. An Important Notice, which is reproduced at the end of this notice below, will be included with the recalibrated kits |  | 12908 | 12,908 | Worldwide Distribution --- including USA and countries of Argentina, Aruba, Australia, Austria, Bangladesh, Bolivia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Mexico, Netherlands, Nicaragua, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Republic of Panama, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Arab Emirates (UAE), United Kingdom, Uruguay, and Venezuela. | Terminated | July 25, 2012 | 1465 | 12908
Z-0479-2009 | 49557 | No | LDJ | Cannabinoid test system. | Toxicology | Toxicology | 510(k) | Micro Plate EIA Oral Fluid Cutoff Calibrator which is a component of the Intercept Micro Plate EIA kit. |  |  | September 08, 2008 | 2008 | OraSure Technologies, Inc. | Results, false-positive test: low % displacement result of 36.0% (spec 39-60%). For further information, please contact the firm at 610-882-1820. | Nonconforming Material/Component | The recalling firm issued a letter entitled "URGENT PRODUCT RECALL NOTICE" dated 9/8/08 to their customers to inform them of the problem and the need to return the product. |  | 2198 | 2198 units | Worldwide Distribution --- USA including states of AZ, CA, CT, DE, FL, GA, ID, KS, KY, MA, MD, MI, NC, NH, NV, NY, PA, VA, and WY, and countries of United Kingdom, Italy, Germany, and Canada. | Terminated | May 05, 2009 | 239 | 2198
Z-0476-2009 | 49550 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VASOVIEW 4 Endoscopic Vessel Harvesting System, with VASOVIEW Flexible Endoscopic Tools in the following configurations: VH-10242: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic scissors; and VH-10243: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic bisector. |  |  | September 09, 2008 | 2008 | CTS, Inc. dba Guidant Cardiac Surgery | Packaging unsealed: Sterility of product may be compromised due to packaging issue. For further information, please contact the firm at 408-635-6874. | Packaging process control | The firm issued a letter entitled "URGENT FIELD ACTION - IMMEDIATE ACTION REQUIRED *** VASOVIEW(TM) 4 Endoscopic Vessel Harvesting System (Kit Only)" via federal express to its consignees on September 9, 2008. The firm has asked its consignees to return the product. Follow-up calls were made to ensure recall effectiveness. |  | 7935 | 7,935 units | Worldwide Distribution --- including USA and country of Australia. | Terminated | December 08, 2010 | 820 | 7935
Z-0459-2009 | 50092 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Sun Med GreenLine/D Laryngoscope Blades Compatible with Fiber Optic Green System Handles, Sterile Disposable. Contents 20. |  |  | April 27, 2007 | 2007 | Sun Med, Inc. | A breakage problem of the acrylic bundles at both distal and proximal ends of the light tube for the MD Laryngoscope Blade GreenLine D Mac 3 has been detected. If breakage occurs during use, there is the potential for leaving approximately a one inch piece of acrylic tube in the patient's mouth. | Device Design | SunMed sent initial recall letter on April 27, 2007 to all distributors requesting that kits containing the MacIntosh size 3, consisting of lots number Gj and Ha be returned to facility. Replacement information was included in letter. On December 12, 2008, the firm issued a Device Recall Press Release with additional information on the potential for the tube to break during use leaving an approximate 1 inch piece of the acrylic tube in the patients mouth and subsequently requiring medical or surgical intervention to prevent serous injury to the patient. SunMed has notified its distributors and customers by certified mail and has arranged for return and replacement of all recalled products. Consumers with question may contact their distributor of the GreenLine/D or SunMed at 1-800-433-2797. |  | 2682 | 2,682 boxes (20 per box) | Nationwide Distribution | Terminated | July 06, 2010 | 1166 | 2682
Z-0621-2009 | 50487 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Flextome Cutting Balloon Device Over-the-Wire Delivery System, Coronary Atherotomy System, Over-the-Wire Delivery System, Catalog No. CB0340006, 4.00mm, Sterile. Made in Ireland: Interventional Technologies Europe Ltd., A subsidiary of Boston Scientific Letterkenny, County Donegal, Republic of Ireland. US Customer Service 888-272-1001. |  |  | November 21, 2008 | 2008 | Boston Scientific Corporation | For one batch/lot of flextome Cutting Balloon (OTW), an incorrect compliance chart was packaged inside the sterile pouch. For the affected batch/lot, the compliance chart packed inside the pouch of the 4.0mm diameter Flextome Cutting Balloon device is for a 3.5mm diameter Flextome Cutting Balloon device. This labeling discrepancy could create a procedural delay as users seek to explain the discr | Mixed-up of materials/components | An Urgent medical Device Field Correction, dated November 21, 2008, was sent to affected customers/hospitals for delivery on November 24, 2008. The letter described the issue; identified product affected; use or further distribution of the affected product should be ceased immediately; and a correct compliance chart included with the notification is to be secured to the outer sterile pouch. Affected product is not asked to be returned as the correct compliance chart is enclosed and is to be secured to the outer sterile pouch. A Reply Verification Tracking Form is to be completed and faxed to Customer Service Call Center. |  | 75 | 75 | Nationwide Distribution --- including states of AL, AR, AZ, CA, CO, GA, HI, ID, IL KS, MA, MD, MI, MN, MO, MS, NV, NY, OH, PA, TN, TX, UY, VA, and WI. | Terminated | December 10, 2011 | 1114 | 75
Z-0622-2009 | 50489 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Genicon 5mm Pyramidal Trocar, Sterile, Genion 6869 Stapoint Court, Suite 114 Winter Park, FL U.S.A. 32792, Rx Only. Catalog Numbers: 100-005-001 &amp; 100-005-002. |  |  | September 25, 2008 | 2008 | Genico, Inc., dba Genicon | The stainless steel tips on the Genicon 5mm Pyramidal Trocar exhibited contamination that resembled rust and there was also pitting on the tip of the trocar. | Manufacturing material removal | Consignees were notified with a letter, dated 9/19/2008, by e-mail and fax on 09/25/2008. The letter stated a description of the issue, analysis and reliability data, indications of Device Malfunction, and recommendations. There was also a request for all defective devices to be returned to the distributor or Genicon for replacement. Genicon Customer service was to be contacted at 407-657-4851. |  | 2965 | 2,965 units (1340 OUS and 1625 US) | Worldwide Distribution --- including USA and countries of Africa, United Kingdom, Denmark, Colombia, Belgium, Italy, Finland, Hong Kong, and Switzerland. | Terminated | January 09, 2009 | 106 | 2965
Z-0445-2009 | 44838 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | April 17, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The power supply may fail to charge the batteries due to various hardware malfunctions. | Nonconforming Material/Component | Terumo notified U.S. consignees via Urgent Medical Device Correction: Safety Advisory recall letter dated 4/17/08 and international consignees were notified via email on 4/24/08 informing them of the potential problem. |  | 256 | 256 | Worldwide distribution: USA, Australia, Belgium, Canada, Colombia, Dominican Republic, Honduras, Hong Kong, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Philippines, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, and Viet Nam. | Terminated | July 07, 2009 | 446 | 417
Z-0449-2009 | 49712 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | Hologic, Inc., 10 Year Fracture Risk Questionnaire Option for QDR X-Ray Bone Densitometers Software |  |  | September 24, 2008 | 2008 | Hologic, Inc. | Software error may lead to a high estimate of major fracture probability. | Software design | Hologic notified customers by telephone on 9/23 and 24/08 and instructed them to not use the software. A customer notification letter was sent via FedEx to customers on September 25, 2008.The letter requested de-installation of the software and return of the software to Hologic in a pre-paid envelope. |  | 15 | 15 software discs | Nationwide distribution in: CO, MA, MN, NM, NY, NC, OR, OH, PA, and TX. Also distributed in the UK. | Terminated | March 07, 2016 | 2721 | 15
Z-0467-2009 | 49495 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | CASMED 500cc Unifusor Classic with piston gauge and stopcock valve |  |  | September 20, 2008 | 2008 | CAS Medical Systems, Inc. | Leaks: Infusion Cuffs may exhibit air leakage at the union of the tube and bag, and not maintain pressure. For further information, please contact the firm at 203-488-6056. | Equipment maintenance | CAS Medical Systems notified accounts by letter dated 9/19/08 requesting return and sub-recall be conducted to their accounts. The letter was entitled "UNIFUSOR(R) INFUSION CUFFS AND PRIVATE LABEL (CASMED-MANUFACTURED) INFUSION CUFFS *** URGENT PRODUCT CORRECTION NOTICE." For further information, please contact the firm at 203-488-6056. |  | 6384 | 6,384 pc | NationwideForeign: Malaysia, Austria, Belgium, Czech Rep., France, Germany, Greece, Italy, Norway, Spain, Sweden, Switzerland and the UK. | Terminated | August 05, 2010 | 684 | 80864
Z-0480-2009 | 49560 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | INSURE(R) QUIK FIT(TM) DEVELOPER KIT, Lot number B1478, Product Numbers: 30025 (InSure Quik FIT Developer Kit) and 50025 (INSURE(R) QUIK FIT(TM) containing Developer Kit and Patient Collection Kits. |  |  | October 03, 2008 | 2008 | Enterix, Inc. | False readings: Reduced sensitivity could affect the test line area and/or control line on the test strips, which may lead to inaccurate results. For further information, please contact Enterix, Inc. at 1-800-531--3681. | Other | Medical Device Recall Letters were sent on October 3, 2008 to all customers by certified mail. The letters were entitled "URGENT *** MEDICAL DEVICE RECALL (U.S. ONLY). Questions are to be directed to Enterix at 1-800-531-3681, |  | 1125 | 45 kits (includes 25 patient collection kits and 1 developer kit) | Product was distributed to medical facilities and laboratories throughout the US. | Terminated | March 26, 2009 | 174 | 1125
Z-0617-2009 | 50220 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model 515A Oxygen Concentrator |  |  | October 30, 2008 | 2008 | Sunrise Medical | potential for sieve beds to have been incorrectly assembled | Device Design | On 10/30/08, the recalling firm began to telephone customers of the problem. The recalling firm followed up with Urgent: Medical Device Recall letters, dated 11/7/2008. The letters requested that users cease the use of the affected device and immediately contact their provider to make arrangements to have the Concentrator replaced. The firm has set up a toll free line to assist customers with identification of affected units and the return process. |  | 428 | 428 units | Worldwide Distribution --- USA including states of AL, CA, FL ,GA, IN, IL, MA, MD, MI, MS, NC, NE, NH, NM, NY, PA, SD, TN, TX, VA, and WI, and countries of Canada, Columbia, Germany, Mexico, Netherlands, Singapore, and United Kingdom. | Terminated | October 27, 2012 | 1458 | 428
Z-0604-2009 | 50154 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | MassTRANSIT Infusion Catheter Kit, 605-731K &amp; 605-780K, Cordis Neurovascular |  |  | June 30, 2008 | 2008 | Cordis Corporation | Two lots of Cordis Neurovascular Masstransit Infusion Catheter Kits has the potential for sterile PTFE contamination within the inner lumen of the catheter. | Pending | A Medical Device Recall letter, dated 6/30/2008, and Acknowledgement Form were sent overnight. The firm asked customers to set aside any affected product in a manner that ensures it will not be used. Customers are asked to review, sign and return the enclosed Acknowledgement Form directly to Cordis Neurovascular. Customers can either return any product with reference to the RA# on the form, or contact their local sales reps to facilitate return of the affected product. Customers will receive a credit for the units they return. If any affected product has been forwarded to another facility, the facility should be contacted to arrange return of the product. If there are any additional questions, Customer Service may be also be contacted at 800-551-7683. |  | 27 | 27 | Nationwide Distribution --- including states of Colorado, Connecticut, Florida, Minnesota, Nebraska, Nevada, Ohio, Oregon, and Virginia. | Terminated | January 02, 2009 | 186 | 27
Z-0618-2009 | 50225 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flexi-Seal Fecal Management System Advanced Odor Control Kit. |  |  | November 20, 2008 | 2008 | ConvaTec | It was discovered that a portion of the inner nylon sleeve of the catheter was visible at the balloon end of a sample device. | Nonconforming Material/Component | Urgent Product Correction Letters were sent by first class mail on November 18, 2008 to Healthcare Professionals and ConvaTec Reps. The reason for the recall was identified along with possible risk to patient. Healthcare professionals were asked to use sufficient lubricating jelly during the insertion process with special care given to those who have a tendency to bleed or are on anti-coagulant medication. Physician should be notified if rectal bleeding or abdominal symptoms develop. ConvaTec reps were asked to contact all health care professionals to whom product or samples were given and reinforce the importance of proper lubrication and following the directions. A return form is to be completed by ConvaTec reps for each visit indicating the name of the Health Care Provider and that the they personally received notice of the product correction. All remaining sample products are to be returned to Besta, Franklin, WI. |  | 795 | 795 | Worldwide Distribution --- including USA and country of Canada. | Terminated | March 09, 2009 | 109 | 795
Z-0610-2009 | 50180 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Stryker Leibinger New Generation/Universal Mandibular System, Mand. Recon Plate, Straight, w. Template, Stryker Leibinger GmbH &amp; Co., Freiburg, Germany; REF # 55-28911. |  |  | September 19, 2008 | 2008 | Stryker Leibinger USA | A template is supposed to accompany the product, but it was not included with this lot. | Process control | Consignees were notified via an Urgent Medical Device Recall letter dated September 2008 and requested to examine their inventory for the affected lot and return the recalled product. Customers are to fill out the attached forms to confirm the return-shipment and send it back. The firm will inform customers immediately regarding replacement deliveries. |  | 23 | 23 | Nationwide Distribution | Terminated | December 15, 2009 | 452 | 23
Z-0508-2009 | 48177 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Novel PEEK VBR XS 16-22 MM, S-5, 26MM, Part Numbers : 64753-126. |  |  | August 16, 2007 | 2007 | Alphatec Spine, Inc. | Firm confirmed that two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number &amp; lot number. They were incorrectly laser marked with the wrong part numbers and sizes. Two lots were mixed from the following lot numbers: Part Number 64733-120 Lot Number 611884, and Part Number 64753-126 and Lot Number 611881. | Process control | The recall was initiated on August 16, 2007 with Alphatec Spine contacting each of the four direct distributors by telephone. This was followed with a Product Recall Notice, dated August 27, 2007, that was mailed to all affected customers. The distributors were informed that the firm was conducting a voluntary recall because two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number &amp; lot number. The firm asked the customers to return the product, and were informed that replacement product will be sent back. |  | 4 | 4 units | Nationwide Distribution --- including states of AZ, FL, NY &amp; TX. | Terminated | January 12, 2009 | 515 | 7
Z-0522-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HD3. Product labeled HD3 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 4279 | 4279 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181
Z-0472-2009 | 49517 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Ultra Laboratory System Software; for recording, reporting and distribution of lab results. GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 22, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software computer, error: The abnormal result flag is not being sent in the HL7 messages on results that are lower than the normal range. For further information, please telephone GE Healthcare Integrated IT Solutions at 800-437-1171. | Software design | GE Healthcare notified their Centricity Ultra Laboratory Information System customers via letter dated 9/19/08. The letter was entitled "PRODUCT SAFETY NOTIFICATION." The letter states that the software correction was installed or will be installed to correct the abnormal results flag not being sent in the HL7 on results that are lower than the normal range. The letter summarized the safety concern and the action taken. Any questions were directed to GE Technical Support at 888-778-3373 or 905-479-1249. The accounts were requested to complete the enclosed customer reply form, supplying the software version number, and fax it back to 847-277-5240. |  | 3 | 3 units | Globally - United States including the states of New York and Texas; and the country of Canada. | Terminated | May 18, 2010 | 603 | 3