## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0523-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HDI4000. Product labeled HDI4000 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. Product also marketed as Philips 4000 New Standard, Philips 4000 System and Multibeam System. According to the firm, these units have the same user manual as the HDI4000 system. The different names are used internally as a description for the different configurations of model HDI 4000. Over the product lifecycle, the product configuration changed as follows: MULTIBEAM SYSTEM : This was the original HDI 4000 system branded with the former company name 'ATL' and labeled, "Distributed by ATL Ultrasound Bothell, WA. USA 98041"; PHILIPS 4000 SYSTEM :This configuration changed the legacy 'ATL' brand / colors to the 'Philips' brand / colors; PHILIPS 4000 NEW STANDARD: This Philips brand configuration eliminated hardware that supported the 'CV'. The CV option was initially standard, but later became optional depending on customer preference. The product ID label is the same for both PHILIPS 4000 SYSTEM and PHILIPS 4000 NEW STANDARD. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 1902 | 1902 Units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181
Z-0544-2009 | 49457 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Dynamics Workplace, version 5.0, Siemens Medical Solutions, Ann Arbor, MI; 10090310. |  |  | July 23, 2008 | 2008 | Siemens Medical Solutions, USA, Inc | If the text is manually modified and additional phrases added as a result of worksheet interactions, some of the phrases may not populate the report. | Software design | Consignees were notified of the problem via a Customer Safety Advisory Notice dated 7/23/08, which provided guidance on how to minimize the risk of the problem occurring, and told customers that a service patch will be installed when it becomes available. |  | 306 | 306 of all products. | Nationwide Distribution. | Terminated | October 18, 2010 | 817 | 306
Z-0595-2009 | 50111 | No | LYT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ossur 3" Ceramic Tip Pin, 4 pack, Model Number: 517DS |  |  | September 26, 2008 | 2008 | Ossur | Ossur has initiated this voluntary recall as a precautionary measure because the firm is unable to guarantee the sterility of the affected products. | Packaging process control | Ossur initiated contact with the customers to locate the pins on Sept. 26, 2008with the Product Manager for this device contacting each consignee by phone.The customers were informed the recall was initiated because the firm was unable to guarantee the sterility of the affected products. Customers were asked to check their inventory for the affected parts. If products were located, they were informed that an Ossur customer service representative will be emailing you with return information any they are to return these pins to the firm as soon as possible. They were also told Ossur does not currently have replacement product available, but they will contact them immediately when they are available.If the products were not located, they were asked to determine where these pins are currently located, if they have been applied to a patient? If so, are they able to track down that patient and bring them in for an evaluation? |  | 12 | 3 pouches (4 pins/pouch) = 12 pins | Nationwide Distribution --- including states of FL, KS, MI, MO, NY, &amp; WA. | Terminated | April 27, 2010 | 578 | 43
Z-0285-2009 | 49135 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M System, a multiparameter physiological patient monitoring system intended for use on adult, pediatric and neonatal patients, within a hospital or facility environment |  |  | April 02, 2008 | 2008 | GE Healthcare | A software error has been discovered when using Combination Monitoring with Solar 8000M/i - PDM - Telemetry in combo mode. This error has only been noted with this particular configuration and mode of monitoring. It has not been observed in other monitoring applications. The software error will result in a reset of the Solar 8000M/i, causing a loss of monitoring for approximately 1-2 minutes a | Software design | A GE Healthcare consignee letter was sent out 9/9/08. The letter dated August 26, 2008 was addressed to Healthcare Administrator/Risk Manger. The letter described the problem, product and provided recommendation. |  | 2782 | 2782 | Nationwide to: AL, AR, AZ, CA, FL, GA, IA, ID, IL, IN, MA, MI, MN, MO, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, and WI. OUS to include: Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, France, Germany, Indonesia, Ireland, Italy, Jordan, Korea, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Pakistan, Saudi Arabia, Spain, Sweden, United Kingdom, and Venezuela. | Terminated | May 17, 2012 | 1506 | 2782
Z-0651-2009 | 50206 | support, head and trunk, wheelchair | IMS | Wheelchair accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | High-low:x frame for the x:panda mechanical chair. |  |  | November 06, 2008 | 2008 | Snug Seat Inc | Breakage in the welds was discovered in a section of the chair frame. | Process control | Consignees were notified by letter dated 11/6/2008. Snug Seat sent a repair kit to their retail dealers on 11/18/2008. The kit includes a safety plate, written and video replacement instructions as well as the tools (allen wrench, crescent wrench and two nuts) necessary to make the safety update. Customers were identified using purchase order numbers, and product serial numbers to enable the dealers to identify the purchaser of the product. Due to HIPPA regulations, purchaser information was not available to Snug Seat, Inc. to contact directly. Written notification will be followed up with a telephone call to each dealer to ensure they have received the retrofit kit. Snug Seat will also insure they understand the importance of the kit and how to properly install and document receipt of the update for their customers. |  | 67 | 67 units | Nationwide. | Terminated | February 17, 2011 | 833 | 67
Z-0439-2009 | 49580 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SteriLyteÂżÂż Liquid Bicarbinate, Formula CL-02; 1 gallon (3.78 L); manufactured by Rockwell Medical Technologies, Inc., Grapevine, TX 76051. This product is used in conjunction with C-Series RenalPureÂżÂż Liquid Acid Concentrate or DC-Series Dri-SateÂżÂż Dry Acid Concentrate in artificial kidney (hemodialysis) machines. |  |  | September 29, 2008 | 2008 | Rockwell Medical Technologies, Inc | Microbial growth was observed in the retain samples of 4 lots of Liquid Bicarbonate after products had been distributed. | Environmental control | Recall notification was faxed and/or e-mailed to consignees on 09/29/08 and then sent via mail on 09/30/08. A follow-up response form was faxed to consignees starting 10/01/08.Consignees asked to return product for a credit. For additional information, contact Rockwell Medical Technologies, Inc., Customer Service at 1-800-449-3353. |  | 2281 | 2,281 cases (4 units per case) | Nationwide. Product distributed to AL, AR, FL, GA, NC, SC, TN, TX, and VA. | Terminated | July 29, 2010 | 668 | 2976
Z-0339-2009 | 49607 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. The product is intended for use in radiation therapy. |  |  | September 18, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Inc. discovered an issue on the Hi-Art system during the course of ongoing testing. It has been identified that treatment plans involving very small structures (volume &lt;2cc), reading in a high dose gradient region, may exhibit inaccuracies with the DVH curve and the Dose Statistics for the small structures. This error can only happen if planning was performed using very small voxels | Software design | Consignees were sent on 9/18/08 a "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" letter dated September 17, 2008. The letter described the issue, product affected, recommended actions and resolution. Questions may be directed to TomoTherapy Inc. Customer Service Center at 1-868-368-4807. |  | 165 | 165 | AKALARAZCACOCTDCFLHIIAILINKYLAMAMDMIMNMONENJNMNVNYOHOKORPAPRSCSDTNTXVAWAWI andWYBelgiumCanadaSwitzerlandGermanySpainFranceGreat BritainIndiaItalyJapanKoreaNetherlandsTurkey andTaiwan | Terminated | April 19, 2011 | 943 | 165
Z-0342-2009 | 50083 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 9900 Elite GSP fluoroscopic x-ray system GE Healthcare, Surgery, Salt Lake City, UT. The system is intended to provide an x-ray display of human anatomy. |  |  | August 29, 2008 | 2008 | OEC Medical Systems, Inc | Under certain conditions, x-ray system may experience operational impairment. | Device Design | Consignees were notified by letter on 08/29/2008. Interim solutions were provided for proper start-up procedures and for avoidance of mechanical impact to the C-arm. A start Up caution label was included with the notification. Further information available at 800-874-7378. |  | 448 | 448 units | Nationwide, including AK, KY, NC, VA WA. and WI | Terminated | April 24, 2012 | 1334 | 1054
Z-0321-2009 | 50077 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Sensation 64 X-Ray System, Model number 8377520. This is a computed tomography x-ray system which is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. |  |  | September 03, 2008 | 2008 | Siemens Medical Solutions USA, Inc | software error - does not allow protocol to properly synchronize with ECG triggers | Software design | The recalling firm issued a Customer Safety Advisory Notice dated 9/3/08 to affected customers via Update Instructions CT045/08/S. This letter informs customers of the potential issued and provides instruction to avoid its occurrence. For additional information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-4500. |  | 9 | 9 units | Nationwide | Terminated | September 23, 2009 | 385 | 177
Z-0325-2009 | 50118 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FS-OMNI, Fusion Omni-Tome, Sterile, Rx only, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105, EC Representative: Cook Ireland Ltd, O'Haioran Road, National Technology Park, Limerick, Ireland. The Fusion Omni-Tome is supplied sterile and is a single use device. This device is used in cannulation of the ductal system and for sphincterotomy. |  |  | October 30, 2008 | 2008 | Cook Endoscopy | The wire control port may have been omitted during manufacture. | Process control | Cook Ireland (distributor) was contacted via email on 10/30/08 and was instructed to contact their customers and coordinate the return of all recall product to Cook Endoscopy. A response sheet was attached. Recall status reports will be issued monthly. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed. Should you have any questions, please contact Cook Endoscopy Customer Quality Assurance Department at 1-336-744-0157. |  | 11 | 11 | Ireland | Terminated | April 07, 2009 | 159 | 11
Z-0645-2009 | 50502 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | CardioVations Ethicon, a Johnson and Johnson company, EndoClamp aortic catheter 100 cm, REF EC1001, Sterile R, individually packaged, Heartport, Inc., Somerville, NJ 08876. This product is indicated for use in patients undergoing cardiopulmonary bypass. It occludes the ascending aorta, vents the aortic root and allows delivery of cardioplegia solution to arrest the heart. The pressure lumen allows monitoring of the aortic root pressure |  |  | November 26, 2008 | 2008 | Edwards Lifesciences Research Medical, Inc. | Balloon rupture: Catheter balloon may rupture during cardiopulmonary bypass procedure. For further information, please telephone Edwards Lifesciences Research Medical, Inc. at 801-565-6173. | Nonconforming Material/Component | All consignees were notified by letter on 11/26/2008 and requested to return any unused product to Edwards. The letter was entitled "PRODUCT RECALL." Questions could be referred to the Edwards representative or call 800-424-3278. |  | 963 | 963 units | Globally - USA including the states of AK, AL, CA, CO, CT, DC, DE, FL, GA, IA, IL, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI, and WV. VA facilities in VA and WI. No military distribution. Foreign distribution to Australia, Canada, Switzerland and UK. | Terminated | August 10, 2011 | 987 | 2230
Z-0448-2009 | 48734 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homocysteine Reagent Test Kit. The Homocysteine test kits were manufactured under 2 brand names. The first brand name Catch, the label reads "CATCH Catch Inc., Bothell, WA 98011 HOMOCYSTEINE REAGENT KIT, Catalog Number A11167, 100 Test Kit, Manufactured at Intersect Systems, Inc., Longview, WA 98632". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1 HCY A (28mL), Homocysteine Reagent R2 HCY R2 (13.5 mL), Homocysteine HCY C (2.8 mL), Homocysteine 0.0 umol/L (2mL) and Homocysteine 27.0 umol/L (2mL). The second brand name Equal the label reads "EQual DIAGNOSTICS FOR IN VITRO DIAGNOSTIC USE ONLY HOMOCYSTEINE 3-Part Kit w/calibrators, Catalog Number E02-058-03, Equal Diagnostic, Exton, PA 19341". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1, Homocysteine Reagent R2, Homocysteine Reagent R3, Homocysteine Calibrator 0.0 umol/L and Homocysteine Calibrator 27.0 umol/L. All individual kits are packaged into individual boxes and then packaged into 1 box. Homocysteine Reagent is intended for in vitro quantitative determination of total homocysteine in serum and plasma. |  |  | January 18, 2007 | 2007 | Catch Incorporated | A high homocysteine level can be reported as falsely low. | Device Design | On 01/18/07 an "URGENT: PRODUCT CORRECTIVE ACTION" letter was sent to the consignees. The letter states the circumstances in which a high homocysteine rate can result in a false low homocysteine rate, a work around for the problem which a flag must be programmed into the user defined homocysteine test parameters and instructions to advise their customers about the situation was addressed in the letter so the consignees can take the recommended corrective action. A confirmation sheet was included with the letter to be signed by a representative of each distributor and returned to Catch Inc. Contact Catch Inc. at 1.425.402.8960 for assistance. |  | 2699 | 2,699 kits | CA and PA. | Terminated | September 28, 2010 | 1349 | 2699
Z-0542-2009 | 49382 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche ISE Internal Standard Gen.2 for use on Roche cobas c systems, Roche Diagnostics, Inc., Indpls., IN; Catalog no. 04522320190. Product is an in vitro diagnostic. |  |  | September 15, 2008 | 2008 | Roche Diagnostics Corp. | Test results, low: Sodium, potassium and chloride results on the cobas 6000 ISE unit may exhibit a downward drift in controls and patient samples within the recommended calibration interval of 24 hours. For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. | Package design/selection | Consignees were notified by letter dated 9/15/08, which provided a work-around for them to use until the problem can be permanently resolved. The letter was entitled "URGENT MEDICAL DEVICE CORRECTION." For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. |  | 7909 | 7909 | Nationwide Distribution. | Terminated | December 01, 2009 | 442 | 7909
Z-0453-2009 | 50156 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Angiotech, Bone Marrow Harvest Needle, Medical Device Technologies, Inc. The device is intended for use Bone Marrow Biopsy. |  |  | July 16, 2008 | 2008 | Medical Device Technologies, Inc. | Product packaging for Bone Biopsy trays/blisters may have a hole in the formed tray. | Other | Each consignee was notified of recall by letter on 07/16/2008 . The letter described the problem and indicated the corrective action to be taken. A Customer Acknowledgment Form with request to fax form to firm's QA Dept was provided. Contact Medical Device Technologies, Inc., Quality Assurance Department at 1.352.338.0440, ext 350 and 353 for assistance. |  | 130 | 130 units | Nationwide - FL, TX, WI, PA, MN, SC, NC, KY, TN, CA, AR, OR, NJ, and IA. | Terminated | April 21, 2009 | 279 | 430
Z-0455-2009 | 49010 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software, Part Numbers: 626553, Version 2.2 |  |  | August 01, 2007 | 2007 | Beckman Coulter Inc | Software error: The recall was initiated after Beckman Coulter identified that the CXP User documentation does not adequately characterize the use of the Live Gate feature.The primary use of a Live Gate is for the international removal of unwanted events, such as debris. The data outside of the Live Gate will not be stored to the final listmode file, this reducing the data set size.The softwa | Software design | The customer notifications were initiated by Beckman Coulter on July 23, 2007 when a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers that the user documentation does not adequately characterize the use of the Live Gate feature. The letter provides the customer with explanation of Live Gate and actions to be taken. For further information, please contact Beckman Coulter, Inc. by telephone at 714-993-5321. |  | 565 | 565 | Globally to the United States and Canada. | Terminated | January 13, 2009 | 531 | 565
Z-0607-2009 | 50162 | No | LJB | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Measles IgG Enzyme Immunoassay Test Kit, DIAMEDIX. The assay is intended for the semi-quantitation of human IgG antibodies to measles virus in human serum by indirect immunoassay to aid in the assessment of the patient's immunological response to measles virus, to determine the immune status of individuals and, when evaluating paired sera, as an aid in the diagnosis of measles infection. |  |  | October 10, 2008 | 2008 | Diamedix Corporation | Possible false-negative test results: The Sample A Diluent Lot # 50768 assigned to Is-Measles IgG Lot #'s 41108 and 41108B has higher conductivity than normal, which can sporadically interfere with the MAGO Plus/Aptus sensing of and picking up of samples. As a result, if the sample is not picked up/processed, it could lead to the possibility of obtaining false negative results in some cases. | Other | Notification of all domestic customers via Certified Mail. Notification of the one international customer via Federal Express. |  | 974 | 974Test Kits | Worldwide Distribution --- including USA and country of England. | Terminated | May 12, 2010 | 579 | 974
Z-0601-2009 | 50144 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Synergy Dynamic, Powered Patient Rotation Bed, Model number 5607. |  |  | October 09, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | These devices have the potential to overheat. | Device Design | All end users (those who either purchased the device and any customer that has rented) were notified by an Urgent Field Safety Notice on/about 10/31/2008. The letter informed the customers of the safety risk and list steps they can take to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that were sold, the consignee of each device was notified by letter for the safety risk and what they can do to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that are rented, an internal notification was sent via e-mail to Service Centers (Because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the Consignee). Each Service Center will perform the correction on the units affected. |  | 2083 | 2,083 units (Total all four units) | Nationwide Distribution. | Terminated | December 10, 2009 | 427 | 2083
Z-0519-2009 | 49014 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Compress/Finn Instrument Set Compress Anchor Plug Assembly, stainless steel, Biomet Orthopedics, Inc., Warsaw, IN; REF RD425032. |  |  | July 18, 2008 | 2008 | Biomet, Inc. | The set screw may not fully seat, which will prevent the handle from mating to the instrument during use. | Process control | Distributors were notified via an Urgent Medical Device Removal Notice, dated 7/18/08, which instructed them to locate and remove the instruments, carefully follow the instructions on the enclosed "FAX Back Response Form," fax a copy of the Response Form to 574-372-1683 prior to return of product, and to notify the hospital personnel via a copy of a recall letter. The letter also stated that the instruments would be reworked and returned to the distributors. Distributors are to confirm receipt of the notice by calling, 800-348-9500, ext 3983. Questions related to the notice should be directed to Mary Hardesty at 800-348-9500 or 574-372-3983. |  | 41 | 41 of all products. | Worldwide Distribution --- including USA and countries of Australia, Belgium and Finland. | Terminated | July 07, 2009 | 354 | 41
Z-0608-2009 | 50167 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi CXR4 Computed Tomography Scanner. This product is an x-ray imaging device that produces cross-sectional images of he body at different angles. The system reconstructs, processes, displays, and stores the collected images. The device output can provide an aid to diagnosis when used by a qualified physician and is intended for general purpose CT applications. |  |  | November 05, 2008 | 2008 | Hitachi Medical Systems America Inc | Software error: The device has a software error that can result in a condition where the displayed image orientation does not match the display orientation markers (right-left reversal). | Software design | On 11/5/2008 Hitachi Medical Systems sent recall notification letters entitled "URGENT - DEVICE RECALL CORRECTION" to its customers informing them of the problem with the software. They stated that a Hitachi service representative will contact each customer to schedule a software upgrade. For further information, please contact Hitachi Medical Systems America Inc. by telephone at 330-425-1313. |  | 42 | 42 | Nationwide Distribution --- including states of AZ, CA, DE, FL, ID, IL, IN, KS, KY, LA, MA, MO, NJ, NY, OH, OK, PA, PR, SC, TX, VA, and WI. | Terminated | September 21, 2012 | 1416 | 42
Z-0579-2009 | 49573 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Knee system Primary Tibial Baseplate #5 Beaded wPA; Catalog Number 5526-B-500, Cementless use only. Sterile. |  |  | September 21, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that one lot of the Triathlon Primary Beaded PA Baseplate Size 5 may not have the Peri-Apatite coating. | Process change control | Important Market Withdrawal notification letters were sent via Fed Ex on 9/20/06 to the direct accounts. Notification letters informed of the reason for the market withdrawal and requested that all branches examine their inventory and hospital consignment locations to identify the product. Product was to be reconciled on the Product Accountability Form and returned to Stryker as per instructions. Contact Stryker Howmedica Osteonics Corp at 1-201-831-5718 for assistance. |  | 10 | 10 | PA, GA, AL, WI, UT, and OR. | Terminated | January 09, 2009 | 841 | 10
Z-0647-2009 | 50500 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CT syngo Dual Energy Viewer on the syngo Multimodality Work Place (MMWP) and on the CT workplace, Model number 10094833. |  |  | November 13, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Indicated orientation does not match actual orientation of the patient. | Software design | The recalling firm issued a Customer Safety Advisory dated 10/8/08 to affected customers via Update Instructions CT052/08/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. Customers should consider "no saving of calculated Dual Energy data sets from non-supine and mirror-option scans via "Save Sets CT"-button." The firm also highly recommends for customers "to check the correct display of "left", "anterior" and "posterior"markings, e.g. by checking the position of the spine with respect to the posterior marking in the image." The error should be fixed within certain upgrade and installation activities on the affected software. |  | 374 | 374 units | Nationwide Distribution | Terminated | March 19, 2013 | 1587 | 374
Z-0643-2009 | 50501 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST CVi Angiographic Injection System with Interface Cable for Toshiba Infinix-i cardiovascular X-ray system, Part Number 700318-002. The ACIST injection system is an angiographic injection system used in interventional cardiology, radiology, and vascular surgical procedures. The ACIST injection system supplies radiopaque contrast media to a catheter at a user-determined variable flow rate and volume which can be instantaneously and continuously varied. Manufactured by: ACIST Medical Systems, Inc., 7905 Fuller Road, Eden Prairie, MN 55344, USA. The ACIST CVi injection system is able to synchronize with certain X-ray imaging systems from Siemens, Toshiba, GE and Philips. |  |  | November 10, 2008 | 2008 | Acist Medical Systems | Multiple Injections may result. It has been determined that when using the ACIST CVi Angiographic Injection system to synchronize with the Toshiba Infinix -i cardiovascular x-ray system there is an injection scenario that may result in multiple injections of contract media when one injection was intended. | Software design | ACIST is distributing an Urgent Medical Device Correction letter, dated November 10, 2008, to affected customers. The letter describes the issue and the scenario when the issue to occur, instructs them to discontinue use of the synchronized mode between the ACIST Angiographic Injection System model CVi and the Toshiba Infinix-i X-ray fluoroscopic system by disconnecting the interface cable immediately. The letter states that ACIST is working with Toshiba to resolve this issue. A response form is asked to be returned either by fax, e-mail or mail. If there are any questions, they should be directed to either ACIST at 952-995-9300 or local Sales Reps or Clinical Applications Specialist. |  | 18 | 18 cables | Worldwide Distribution --- USA including states of AR, AZ, FL, and NY and countries of France, Germany, Italy, Japan, and Spain. | Terminated | May 07, 2011 | 908 | 18
Z-0505-2009 | 50168 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Sinu-Lift(TM) System, LLC., Model SLS-1, Sterile, Innovative Implant Technology, |  |  | September 30, 2008 | 2008 | Innovative Implant Technology | Possible device failure: The cutter's shaft of the Sinu-Drill of the Sinu-Lift System could potentially get stuck inside the tap, thereby not allowing the cutter to spring out of the tap upon encountering the sinus membrane, increasing the possibility of perforating the sinus membrane. | Nonconforming Material/Component | All consignees will be notified by letter entitled "URGENT: MEDICAL DEVICE RECALL" dated September 30, 2008, using priority mail. Consignees will also be contacted by telephone. Consignees will be asked to return the product using the prepaid FedEx label included with the letter. The returned product will be quarantined. For further information, please telephone Innovative Implant Technology at 305-975-4101 |  | 299 | 299 | Worldwide Distribution --- including USA and countries of Canada, Mexico, Japan, Australia, Japan, Egypt, and Turkey. | Terminated | October 01, 2009 | 366 | 299
Z-0871-2009 | 50546 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural-Knee II System Modular Cemented Tibial Baseplate, Right, Size 3, For cemented use only, sterile, Zimmer, Warsaw, IN; REF 6420-01-230. |  |  | November 25, 2008 | 2008 | Zimmer Inc. | Sterility compromised: Lack of assurance of sterility, as a corner flap of the sterile barrier packaging on the inner cavity may be trapped in the seal of the barrier packaging for the outer cavity and the seal integrity of the outer sterile barrier cannot be assured. | Packaging process control | Zimmer distributors were notified by email of letter dated 11/25/08 on 11/25/08. This letter is entitled "URGENT: DEVICE RECALL." Secondary consignees/user accounts were notified by a copy of the 11/25/08 letter, which was sent on or before 12/05/08. Consignees were instructed to stop using the product and to return it to Zimmer. For further information, please contact Zimmer, Inc. by telephone by dialing 800-613-6131. |  | 10 | 10 | Nationwide Distribution --- including states of Arkansas, Florida, Michigan, North Carolina, New York and Wisconsin. | Terminated | November 12, 2009 | 352 | 10
Z-0487-2009 | 50087 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Easy Core Biopsy Device, 18 ga. x 21 cm, sterile, latex free. Boston Scientific, Spencer, IN; UPN 500-104, REF M0065001040 and M0065001041. |  |  | October 17, 2008 | 2008 | Boston Scientific Corp | These devices may experience difficulty cocking and arming, resulting in an inability to use the devices. | Nonconforming Material/Component | U.S. consignees were notified via letter dated 10/17/08, which informs them of the problem, instructs them to remove from use and quarantine the recalled lots and requests completion and submission of the Reply Verification Tracking form via fax at 508-683-5578. For questions about this recall, contact your local Boston Scientific representative. |  | 14254 | 14,254 of all devices. | Worldwide Distribution including USA and countries of Algeria, Austria, Australia, Belgium, Canada, Denmark, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Ireland, Italy, Japan, Mexico, Netherlands, Norway, Philippines, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Japan and United Arab Emirates. | Terminated | December 15, 2009 | 424 | 15114
Z-0870-2009 | 50540 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Web 3.0 software; the software allows physicians convenient, quick access to medical images and related data they need via a web browser, password and a standard PC; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software error: There are two potential safety situations with Centricity Enterprise Web software - a forced log off may occur while using the Centricity Enterprise Web during an open session, and Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity Enterprise Web software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomalies resulting in patient safety issues involving a forced log off during an open session while using the Centricity Enterprise Web and measurements displayed in Centricity Enterprise Web will use the calibration provided by the original scanning modality. Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. For further information, please contact GE Customer Care Center by telephone at 800-437-1171. |  | 742 | 742 units | Worldwide Distribution including USA and countries of Australia, Austria, Bahrain, Belgium, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Malaysia, Malta, Martinique, Netherlands, New Zealand, Nigeria, Norway, Portugal, Reunion, Russia, Saudi Arabia, Slovenia, Singapore, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | May 18, 2010 | 522 | 742
Z-0580-2009 | 49597 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABBOTT Clinical Chemistry Bilirubin Calibrator, List Number 1E66-04. |  |  | September 12, 2008 | 2008 | Abbott Laboratories Inc. | The field action was initiated after an investigation for the Total Bilirubin reagents (List Number 8G62 and List Number 6L45) determined that the matrix of the secondary standard used in the value assignment of the calibrator is sensitive to the Diazo method. The matrix of this secondary standard caused a positive bias. Because of these findings, a new Total Bilirubin calibrator value assignm | Component design/selection | On September 12, 2008, a Product Correction letter with an attached Customer reply form was mailed via US Postal Service to all affected customers that have received the affected lots. The letter instructed customers to take the following actions:1. Identify the lot number(s) of Clinical Chemistry Bilirubin Calibrator currently used in your laboratory;2. Replace the calibrator value sheet with the appropriate attachment provided with this letter. Discard the previous calibrator value sheet;3. Configure the revised Total Bilirubin calibrator values provided in the Attachments, for the appropriate calibrator lot number and reagent list number.ARCHITECT cSystems: Refer to System Configuration in Section 2, Installation Procedures and Special Requirements in the Architect System Operations Manual.AEROSET: Refer to Calibrator/Control Configuration in Section 2, Installation Procedures and Special Requirements of the AEROSET System Operations Manual;4. Calibrate the Total Bilirubin assay. Verify the calibration with at least two levels of controls according to the established quality control requirement for your laboratory;5. Evaluate your quality control. Shifts in quality control and proficiency testing samples may vary among laboratories and should be evaluated according to your laboratory procedures. NOTE: Refer to the Attachments (E and F) for the representative shift with the revised calibrator values (Patient sample and Control/QC recovery); and6. Evaluate your reference range and any shifts in patient results. It is recommended that each laboratory established its own expected ranges, which may be unique to the population it serves.Please retain this communication for future reference. If you have forwarded any Clinical Chemistry Bilirubin Calibrator to another laboratory, please provide a copy of this letter to them.For questions, please contact ABBOTT Customer Service at 877-422-2688 (or 877-4ABBOTT). Customers outside of the |  | 7101 | 7,101 units worldwide (2,267 units in US). | Worldwide Distribution including United States, Mexico, Columbia, Uruguay, Puerto Rico, Canada, Hong Kong, New Zealand, Bahamas, Guatemala, Argentina, Germany, Thailand, Trinidad/Tobago, Venezuela, Chile, Brazil, El Salvador, Singapore, Australia, Curacao and Cayman Island. | Terminated | April 24, 2012 | 1320 | 7101
Z-0869-2009 | 50535 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon PKR Insert X3 #1 LM/RL - 8mm, 5630-G-108, Lot code NXDMEE |  |  | November 25, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Label mix-up: Stryker Orthopaedics became aware of a lot for lot mix up between Triathlon PKR Insert X3 #1 LM/RL 8mm, 5630-G-108 lot code NXDMEE and Triathlon PKR Insert X3 #1 RM/LL- 12 mm 5630-G-122 lot code M8JMEE. | Labeling mix-ups | Recall notifications entitled "URGENT PRODUCT RECALL" were sent via Federal Express on 11/25/2008 to Stryker branches, hospital risk management, Hospital Chief of Orthopaedics. For further information, please contact Stryker Howmedica Osteonics Corporation by telephone at 201-831-5718. |  | 29 | 8 mm - 1 unit; 12 mm - 8 units. | Nationwide distribution: Product was distributed to 7 Stryker branches and one hospital. | Terminated | August 30, 2010 | 643 | 29
Z-0129-2009 | 48188 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.2, Part Number: 626553. This product is used for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. These properties are measured when the cells pass through one or two laser beams in single-file. |  |  | October 11, 2007 | 2007 | Beckman Coulter Inc | Software error: This field action was initiated after Beckman Coulter determined that the CXP software does not refresh the dataset automatically under certain circumstances. The User documentation does not clearly state that the user should refresh the dataset after changes are made. | Software design | Software error: A Product corrective Action (PCA) letter was mailed via US Postal Service on October 17, 2007 to all Cytomics FC 500 Flow Cytometry System with CXP Software accounts. This letter was entitled "URGENT: PRODUCT CORRECTIVE ACTION." Customers were informed that the user documentation does not clearly state that the user should refresh the dataset after changes are made. The letters provide the customers with an explanation and action to be taken whenever a protocol containing data is modified by making changes to cytosettings, regions or gates.Customer were instructed to review all data plots and results before reporting results out. any uncharacteristic patterns in the data or unexpected results must be verified by running Quality Control and if necessary, preparing new samples.In addition, customer were instructed to please share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Also they were asked to complete and return the an attached Customer Response Form within ten days.Customers with questions or concerns regarding the notification, could call Beckman Coulter Customer Service at 1-800-526-7694 in the United States or Canada. Outside the US and Canada, they were to contact their local Beckman Coulter Representative. |  | 1188 | 594 Worldwide: 518 in the United States and 76 in Canada | Globally - United States and Canada | Terminated | January 13, 2009 | 460 | 1188
Z-0578-2009 | 49526 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Custom Waste Management Kit, K10-01756/B, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095 |  |  | September 12, 2008 | 2008 | Merit Medical Systems, Inc. | Leak, disposal bag: A Merit Disposal Depot (MDD) bag contained in certain custom waste management kits may leak. For further information, please contact Merit Medical Systems, Inc. by telephone at 801-208-4748. | Process control | All consignees were notified by sales reps via phone, fax or visit between 09/15/2008 and 10/09/2008. Sales reps were to instruct them to immediately discontinue use and quarantine all affected products. A Product Retrieval Form was to be completed and signed by the customer's site representative. More information available at 801-316-4998. |  | 399 | 399 kits | United States including the states of AR, AZ, CA, CO, FL, KY, IL, IN, MA, MI, MN, MO, NJ, NY, PA, SC, TX, UT, VA, WI, and WV. No military or foreign distribution. VA facilities in CA and FL. | Terminated | January 13, 2009 | 123 | 399
Z-0581-2009 | 49608 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension EXL Clinical Chemistry System used with software versions 8.5.1 and 8.5.1SP3. |  |  | September 30, 2008 | 2008 | Dade Behring, Inc. | Error messages, "Wash Aspirator Failure" and "Wash station unable to recover" are not generated when one or both HM wash probes are totally occluded. | Software design | The recalling firm notified consignees by telephone to inform them of the problem. A follow-up letter, "Urgent Field Safety Notice" dated 9/30/08 was issued with information describing the problem, the possible impact of an occluded wash probe on assay results, and emphasizing regular performance of routine wash probe maintenance as detailed in the operator guide to prevent wash probe occlusion. For questions, please contact the Siemens Technical Solution Center at 800-441-9250. |  | 57 | 57 units | Nationwide Distribution including states of AR, CA, CO, IA, ID, IL, IN, LA, MD, MI, MO, MS, NC, NE, OH, OK, PA, SD, TN, TX, VA, and WI. | Terminated | May 05, 2009 | 217 | 57
Z-0817-2009 | 50088 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips M8000-64100 Table Mount |  |  | October 22, 2008 | 2008 | Philips Medical Systems | Table Mounts assembled with incorrect length hardware may fall | Process control | Philips issued an "Urgent Medical Device Correction" notification to all potentially affected customers on 10/22/08 via UPS for the US customers. The letter mailed to customers outside the US by Philips' regulatory contacts in each country via a tracked method.The letter informs customers of the problem and gives instructions on how to identify affected table mounts. Field Service Engineers will be dispatched to each customer site to inspect each Table Mount. If customers detect any sign of a loose connection between the monitor and the mounting plate while they are awaiting the Field Service Engineer, they are instructed to make sure that the monitor is not in a position to fall. They are asked to disconnect the monitor from the table mount. For additional information, contact Philips Medical Systems at 1-978-687-1501. |  | 14292 | 14,292 units | NationwideCanadaForeign: Argentina, Aruba, Australia, Austria, Bahain, Bangladesh, Belgium, Bosnia, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Estonia, Egypt, Finland, France, Gabon, Germany, Great Britain, Greece, Hong Kong, Iceland, India, Iraq, Ireland , Italy, Japan, Korea, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Namibia, Netherlands, Norway, Peru, Philippines, Poland,Portugal, Russia, Saudia Arabia, Singapore, SloveniaSouth Africa, Spain.,. Srilinka, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, UAE, Uzbekistan, Venezuela, Vietnam, and Zambia | Terminated | March 24, 2017 | 3075 | 14292
Z-0550-2009 | 49510 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Burrs |  |  | August 08, 2008 | 2008 | Sterilmed Inc | SteriMed, Inc is conducting a product removal of a small group of devices individually packaged and sealed within a pouch. The removal is being conducted because the possibility exists that some of the pouches were not properly sealed on one end.A breach in packaging seal or a failure in packaging integrity has the possibility of risk to the patient in terms of transmitting organisms capable | Packaging process control | Consignees were notified of the recall via verbal communication in person by on-site technician and or their local sales representative on 8/8/08. Others were contacted via fax or e-mail followed by on-site visits 8/11/08. SteriMed also provided consignees (Materials Manager/Risk Manager) with a "Urgent Product Removal" recall letter (August 14, 2008) in person informing them of the potential of receiving affected product. Included in the Product Removal Notification was a Customer Reply Form to inventory quantity of product received and quantity returned. Contact SterilMed, Inc. at 1-888-541-0078 for assistance. |  | 125 | 125 | AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, MA, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WI, and WV OUS: CANADA | Terminated | April 26, 2011 | 991 | 5070
Z-0619-2009 | 50235 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 1 ml Integra TB Syringe with Retracting Precision Glide Needle; 1 ml 27 g 1/2 (0.4mm x 13 mm); Sterile; REF # 305298 |  |  | November 24, 2008 | 2008 | Becton Dickinson &amp; Company | Expiry date on the shelf carton is incorrect. It reads 2013-13, it should read 2013-03. | Incorrect or no expiration date | Recall letters were sent on November 24, 2008 to all direct customers and potential end users of the product by UPS second day air. BD requests that inventory be examined and affected product be quarantined and returned as per recall letter instructions. Questions should be directed to BD customer service at 888-237-2762, M-F 8.OO AM to 5:00 PM EST. |  | 334400 | 334,400 units | GA, KS, OH, MO, FL, PA, KY and NY. | Terminated | March 27, 2009 | 123 | 334400
Z-0646-2009 | 50480 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Drive Belt for Radiation Therapy Simulation systems SIMVIEW 3000 and SIMVIEW NT with CT option or IDI. |  |  | November 06, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Excessive wear on drive belts used for Image Intensifier and/or X-ray head drive mechanisms may result in uncontrolled movement, which may cause injury. | Device Design | Firm has reissued the Urgent Medical Device Customer Letter regarding the belt replacement and a reminder message to replace the belts yearly during the annual Preventative Maintenance. The letter requests that customers follow the guidelines, which consist of making any required SAD adjustments utilizing the Image Intensifier drive, the X-Ray Head Drive, or both; being vigilant of any change in the type or level of noise from either of the drive belt assemblies while they are in use as this may be an indication of wear; discontinuing clinical operation if there is any doubt, and contacting their local Siemens Service organization to inspect the belts. Currently, the firm is distributing the current belt kit to replace the entire installed base of active SIMVIEW Simulators (3000 and NT models). |  | 223 | 223 units | Worldwide Distribution | Terminated | December 09, 2010 | 763 | 223
Z-0879-2009 | 50698 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biomet Microplasty Elite Instruments Vanguard Slidex Femoral Block, 70 mm, Biomet Orthopedics, Inc., Warsaw, IN; REF 32-485156. A manual surgical instrument for general use; a non-powered, hand-held, or hand-manipulated device; instrument guide. |  |  | December 18, 2008 | 2008 | Biomet, Inc. | Falling component: The universal pin, which holds the sliding instrument saw guide in its slot, may loosen during surgical procedures and fall into the wound. | Other | Consignees were notified by letter entitled "URGENT MEDICAL DEVICE REMOVAL NOTICE" dated 12/18/08 advising them of the problem and that the instruments should be removed from use and returned. For further information, please contact Biomet, Inc., by telephone on 800-348-9500. |  | 444 | 444 of all products | Worldwide Distribution --- including USA and countries of Australia, Canada, Japan and the United Kingdom. | Terminated | January 20, 2010 | 398 | 444
Z-0838-2009 | 50136 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | NeuroTherm Simplicity III RF Electrode |  |  | July 28, 2008 | 2008 | Neurotherm, Inc. | Distal tip may detach from the probe | Labeling design | NeuroTherrn sales representatives were notifiedtelephonically or through e-mail. End users were notified through visits by NeuroTherm sales representatives on 7/28/08 and 8/15/08. For questions, please contact NeuroTherm, Inc. at 1-978-777-3916. |  | 71 | 71 units | CO, FL, IL, OH, TN, TX , VA, and WYForeign: Australia, China, and UK, | Terminated | September 24, 2009 | 423 | 71
Z-0611-2009 | 50193 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ReSolve Hydrophilic Coated Locking Biliary Catheter, REF RBC-8-038, 8.5 F, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | November 07, 2008 | 2008 | Merit Medical Systems, Inc. | Biliary drainage catheter tips may separate from the catheter shaft after placement. | Other | Customers were notified by Merit representatives by phone or visit beginning 11/07/2008. They were to identify and quarantine any unused inventory and complete a Product Retrieval Form. For further information, call 800-356-3748 x 4929. |  | 284 | 284 units | Nationwide. OUS distribution to Canada, Colombia, France, Germany, Mexico, Ireland, Netherlands, South Korea, and UK. | Terminated | August 05, 2009 | 271 | 1164
Z-0856-2009 | 50208 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, a component of the CyberKnife Robotic Radiosurgery System. Model number 025007, manufactured by Accuray Inc., Sunnyvale, CA |  |  | October 30, 2008 | 2008 | Accuray Inc | Product may not be tensioned properly, potentially causing unexpected rotation or descent. | Process control | All affected sites were notified of the potential problem via Customer Advisory Notice dated 30 October 2008. The letter states required actions that the customer should follow at the beginning of each treatment day. The tension of the belt drive and the support bolts in question will be inspected by Accuray Field Service personnel at the affected facilities. In addition, the preventive maintenance procedure will be updated to include regular inspection and replacement of the belt-drive. Any questions or concerns regarding this issue should be directed to Accuray Customer Support at 1-877-668-8667 or 1-408-716-4700 or customersupport@accuray.com. |  | 13 | 13 units | Worldwide Distribution --- including USA and countries of France and Turkey. | Terminated | September 10, 2009 | 315 | 13
Z-0582-2009 | 49609 | [6 empty] | QuantumPulse machine. Vibe Technologies, Greeley, CO. |  |  | October 01, 2008 | 2008 | VIBE Technologies | Medical device marketed without marketing approval for claims that include strengthening, recharging and removing toxins from cells. | PMA | Customers were notified by letter on September 16, 2008 and were told to discontinue use as a medical device. This letter notification included: 1. A warning label to be placed on the Vibe machine stating that it is not a medical device and should not be used as such;2. An updated operation manual/users guide containing no medical conditions or human body claims; and3. A certification to be signed by the user and returned acknowledging receipt of the letter; affixed the warning label; will not promote the Quantum Pulse as a medical device; will remove any medical claims from their website; destroy any Quantum Pulse literature making medical claims and destroy any previous operation manual/user guides.Please contact VIBE Technologies at 800-356-9594 for questions or need assistance in complying with these actions. |  | 58 | 58 | Worldwide Distribution. | Terminated | May 03, 2012 | 1310 | 58
Z-0858-2009 | 50483 | No | LPB |  |  | Cardiovascular | PMA | RF Ablation System Foot Switch. (Can be used with Boston Scientific Maestro 3000 and EPT-1000XP Cardiac Ablation Cardiac Ablation Systems). Manufactured by Altech Corporation, Flemmington NJ and distributed by Boston Scientific Corporation, San Jose, CA, catalog numbers/UPN numbers 21840/M004 218400. |  |  | November 25, 2008 | 2008 | Boston Scientific Corp | Product malfunction may result in delivery of energy without depressing the foot switch that controls delivery. | Process control | Recall initiated on November 25, 2008. Urgent Company Initiated Medical Device Recall - Immediate Action Required letters, dated 11/26/2008, were sent to consignees via Federal Express and included return and reply forms for consignees to respond. The letters asked that further distribution or use of any foot switch affected by this recall be ceased immediately. Attached to the letter was a document entitled "Instructions for Determining Affected Product," which provided instructions and serial numbers for determining if customers have product affected by this recall. In addition, a document entitled "Recall Instructions" was also attached. Affected product should be segregated and returned to Boston Scientific in accordance to the "Recall Instructions," and the Reply and Verification Tracking Form should be completed and returned. Customers will receive replacement product for all recalled product returned to Boston Scientific. Local Sales Reps should be contacted in regards to any questions customers may have pertaining to this recall. |  | 322 | 322 units | Worldwide Distribution --- including USA and countries of France, Great Britain, Saudi Arabia, Spain, Canada, Portugal, Germany, Italy, Australia, Russia, Netherlands, Ireland, Libya. | Terminated | April 26, 2011 | 882 | 322
Z-0857-2009 | 50477 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator, when used in combination with Syngo RT Therapist Version 4.1 build 67 and automatic Motion of Flat Panel, part number 8139789, manufactured by Siemens Medical Solutions, Oncology Care System. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Under certain conditions, the flat panel may move unexpectedly and collide with the patient. | Labeling design | The firm issued an Urgent Medical Device Customer Safety Letter, dated 10/23/2008. With the letter included a detailed description of the issue, the condition under which this issue may occur, and recommendation to avoid them. Customers are to read the information carefully and distribute to the responsible persons. The Customer Safety Letter should be included their syngo RT Therapist Express manual and remain there until the next release of software is installed. In the interest of safety, the preventive measures should be performed and all affected personnel should be informed immediately. The firm is tracking responses via return cards and completion protocol. |  | 31 | 31 consignees have the configuration noted above. | Worldwide Distribution --- including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Belgium, France, New Zealand, Malaysia, Spain and Canada. | Terminated | December 09, 2010 | 777 | 31
Z-0396-2009 | 47848 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ConMed Linvatec Hall Surgical Blades, Sterile ST SAG BLDE CRS, Catalog # 00502313300, 00502313700, 00502313000, 00502313100, 00502313500, 00502313200, and 00502314300. |  |  | February 21, 2008 | 2008 | Linvatec Corp. | Sterility compromised. There is a possibility that certain catalog number and lots of ConMed Linvatec Hall Surgical Blades, may have a hole(s) in the corner(s) of the blister package(tray), thereby potentially compromising the sterility of the devices. | Process control | An Urgent Medical Device Recall letter was issued. Customers must check their facility's inventory immediately for the catalog and lot numbers in the attached list. They also need to complete and sign the enclosed Reply Form and mail or fax the Reply Form to ConMed Linvatec using the enclosed self addressed postage paid envelope or fax 727-319-5701. Customers have an option to either re-sterilize the products using the re-sterilization instructions in the letter or they can return the devices to ConMed Linvatec for a credit using the enclosed preprinted FedEx Air Bill with ConMed Linvatec's Bill Number and the return goods number. Recall is being conducted to the distributor and end user level. |  | 13388 | 13388 | Worldwide Distribution | Terminated | January 15, 2009 | 329 | 199977
Z-0531-2009 | 49100 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Mainline Confirms hCG urine pregnancy test, Mainline Technologies, Inc., Ann Arbor, MI; Catalog # 6008C. |  |  | August 18, 2008 | 2008 | Mainline Technology, Inc. | Lack of assurance of safety and efficacy: Unapproved for marketing in the U.S., non-compliance with cGMP regulations and no stability data to support labeled expiration dates. | PMA | Consignees were notified via a Medical Device Removal Notice dated 8/18/08 to locate, discontinue use, and return product along with a copy of the letter. The recall was extended to an additional product via letter dated 9/5/08. |  | 33919 | 33919 kits | Worldwide Distribution --- including USA and countries of Belgium and Switzerland. | Terminated | February 07, 2011 | 903 | 45330
Z-0538-2009 | 49112 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Depuy DELTA Xtend, Standard Humeral PE Cup, Dia 42/+9 STD, sterile, Depuy France; REF 130742209. |  |  | July 10, 2008 | 2008 | Depuy Orthopaedics, Inc. | Compressed air may become trapped between the cup and the stem, preventing the tapered walls from engaging and causing the cup to feel unstable. | Device Design | A letter, "Urgent Device Correction" dated July 8, 2008 was sent via certified mail to implanting surgeons advising them of the issues and guidance on how to mitigate the problem. A communication was also sent to all sales representatives and territory offices on July 15, 2008. For questions, please contact Depuy Orthopaedics, Inc. Customer Service at 1-800-473-3789. |  | 1521 | 1,521 of all devices. | Nationwide Distribution. | Terminated | October 28, 2009 | 475 | 1521
Z-0620-2009 | 50472 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Foot switch, Part Number 1009.81. Maquet GmbH &amp; Co. KG, 76437 Rastatt, Germany. |  |  | December 08, 2008 | 2008 | MAQUET Inc. | A bearing point of the pedal of the foot switch may come loose. With continued use of the foot switch, one or both springs beneath the pedal can fall out of their mounting, and it is possible the pedal comes to rest on the switch causing an unintended motion of the operating table during patient use. | Device Design | An Urgent Device Correction letter (with an enclosed Customer Response form) was sent via Federal Express on Monday, December 8, 2008 to all customers who had received tables. Customers were notified in the letter of the problem and actions to be taken including not to use any affected foot pedal for safety reasons. In addition, customers were instructed to complete and return the Customer Response Form via fax (732-667-1903) to Maquet Inc. Please call USA Maquet, Inc. at toll free number 1-888-627-8383 #1 (then select the option for Surgical Workplace products) for questions. |  | 587 | 587 units total (54 units in US) | Worldwide Distribution. | Terminated | June 18, 2012 | 1288 | 587
Z-0855-2009 | 50143 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse 8.5 (build 8.2.2X) systems with distributed Calculation Framework (DCF) 8.2.22, model number H48, manufactured by Varian Medical Systems, Inc., Palo Alto, CA. |  |  | October 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | When using the multiple static segment option, the device may alter dose delivery. | Software design | An Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps on October 2, 2008. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm recommends that customers DO NOT USE the Multiple Static Segment (MSS) option in Varian Leaf Motion Calculator version 8.2.22. The firm also recommends setting the default delivery method for Varian Leaf Motion Calculator to Sliding Window. When using Eclipse 8.5, always follow the Safety Instructions found in chapter 2 of "Eclipse External Beam Planning Instruction for Use", P/N B501357R01A, March 2008. If customers believe a patient treatment has been affected by this issue, they should notify a Varian Service or Support person as soon as possible. |  | 64 | 64 units | Worldwide Distribution --- including USA and countries of Switzerland and the Netherlands. | Terminated | November 27, 2010 | 786 | 64
Z-0597-2009 | 50120 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical drive brake (on patient supports). Used on certain Philips Medical Systems MX series and Brilliance series CT scanners. |  |  | September 23, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | The product is being recalled because the patient support may travel downwards without being commanded to move. | Process control | A On 11/15/2007 &amp; 6/6/2008 the firm sent Urgent Device Notification letters to their customers. They notified their customers that a Field Service Engineer will inspect and repair the brake system. |  | 559 | 559 | Worldwide Distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and countries of Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Canada, China, Colombia, Costa Rica, Czech Republic, Denmark, Egypt, England, France, Germany, Greece, China, India, Indonesia, Ireland, Iraq, Italy, Japan, Jordan, Korea, Lebanon, Libya, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, South Africa, South Korea, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 01, 2012 | 1255 | 559
Z-0506-2009 | 50181 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 70 Semi-Electric Bed, Hill-Rom Company, Batesville, IN.; Model HS-968. The bed is a general purpose bed for use with low to medium acuity patients in the home care environment. |  |  | November 11, 2008 | 2008 | Hill-Rom, Inc. | Failure, mechanical: If the bed mechanism is cranked downward when the bed is already in its lowest position, the springs may cause the bed to rise suddenly if the bed is empty, more slowly if occupied by a patient, or when the patient exits the bed. For further information, please contact Hill-Rom, Inc. at 800-445-3720. | Other | Consignees were notified of the problem by letter dated 11/11/08 and instructed to position the bed at least 3 inches above the lowest bed position in order to prevent the problem, and then to only use the pendant electronic controls to raise and lower the head and foot sections of the bed, until Hill-Rom can repair the bed. The letter was entitled "Urgent Medical Device Correction." |  | 567 | 567 | Nationwide. | Terminated | December 31, 2009 | 415 | 567
Z-0625-2009 | 49568 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Vision 3 Flat Panel Shroud, Part #0682-000-414 (Individual Flat Panel Shroud, Ver. 2, Reworked), 0682-000-415 (SV3 Shroud Retro Fit Kit) and 0682-000-416 (Flat Panel Retro Fit Kit); distributed by Stryker Communications, Flower Mound, TX 75028. The 0682-000-414 (FP Shroud version 2 reworked) is a shroud designed to cover the cabling necessary on a Stryker Version 3 Monitor when mounted on the Flat Panel Arm. |  |  | September 02, 2008 | 2008 | Stryker Communications Corp | Paint chips: Shroud used to cover cables on the Stryker Vision 3 flat panel monitor (monitor used in surgery rooms during procedures) was not manufactured/painted to specification, resulting in the possibility that paint chips could fall off and fall into non-sterile and/or sterile fields during surgery. For further information, please contact Stryker Communications Corporation by telephone at 97 | Nonconforming Material/Component | Firm began notifying consignees of recall via letter on 09/02/08. Consignees advised that a representative would shortly visit the facility to replace the shroud with one that complies with specifications. The recall letter is entitled "URGENT: Device Recall." |  | 116 | 116 units. | Worldwide distribution; Product distributed nationwide to consignees in the states of CA, FL, IL, IN, MA, MN, NC, NM, OH and TX and to the following countries: Canada, Spain and the United Kingdom, | Terminated | January 08, 2010 | 493 | 116
Z-0616-2009 | 50198 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Esmark Elastic Bandage, 4" x 3 yards, sterile, for single use only, latex free. Distributed by Cardinal Health, McGaw Park, IL 60085-6787. |  |  | November 18, 2008 | 2008 | Cardinal Health | The natural rubber latex bandages were mislabeled as latex free. | Labeling Change Control | A recall letter dated 11/17/08 was issued to the Directors of Materials Management and sent via DHL 2-day service on 11/18/08. Accounts were advised that the elastic bandages containing latex were mislabeled as latex free. Customers were requested to return their inventories of the affected product to Cardinal Health for credit or replacement by calling the Presource Customer Sales Operations Group at 800-766-0706 for return instructions. Customers were also requested to complete the enclosed Cardinal Customer Acknowledgement form by indicating the amount of affected product being returned and fax it to 847-689-9101. Please direct questions to Quality Systems, Professional Services at 800-292-9332. |  | 9521 | 9,521 units (of all devices) | Worldwide Distribution including USA, Guam and Korea. | Terminated | November 04, 2009 | 351 | 9521
Z-0868-2009 | 50525 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dri-Lok Disposable Cannula, Part numbers 3910-075-500, 3910-075-501, 3910-075-502, 3910-075-650, 3910-075-651, 3910-075-652, 3910-075-800, 3910-075-801, 3910-075-802, 3910-090-500, 3910-090-502, 3910-090-650, 3910-090-652, 3910-090-800, and 3910-090-802. |  |  | December 01, 2008 | 2008 | Stryker Endoscopy | Leak: Leak-resistant barrier may not properly seal and leak when an instrument is inserted, and keep leaking after instrument is removed. | Component change control | Recall initiated 12/01/2008, Sales representatives were notified via telephone and e-mail. Product advisories entitled ""Urgent: Device Removal" and dated December 1, 2008, were sent to all consignees. All unused product is to be returned to the firm. For further information, please contact Stryker Endoscopy by telephone at 408-754-2124. |  | 4194 | 4,194 boxes (24,750 units) | Worldwide Distribution --- including USA and countries of Australia, Brazil, Canada, Chile, China, France, Germany, Greece, India, Italy, New Zealand, Poland, Portugal, Sweden, Spain, Singapore, Switzerland, and UK. | Terminated | January 10, 2011 | 770 | 4194
Z-0514-2009 | 48190 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Quantimetrix Dip and Spin Control Urine Dipstick/Microscopics Control. |  |  | June 01, 2007 | 2007 | Quantimetrix Corporation | This recall was initiated due to efficacy concerns with the stabilized erythrocytes, leukocytes and lymphocytes contained in this product group. | Employee error | The recall communication was initiated when customer both direct and distributors who have purchased the lots were mailed a recall notification letter dated June 8, 2008 with instructions which was sent via first class mail.Customers were requested to assist in the voluntary recall of the listed product and if any of the affected lots have been consumed please review their records to ensure that the obtained values conform to the stated specifications found in the lot specific product insert. |  | 24937 | 24,937 sets for all products. | Nationwide Distribution including states of AK, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MN, MO, NC, ND, NH, NM, NY, OR, PA, SC, SD, TN, TX, VA, WA and WY. | Terminated | February 18, 2009 | 628 | 24937
Z-0629-2009 | 50207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Cyberknife Robotic Radiosurgery System. A radiation therapy device, MultiPlan (MP) Treatment Planning Software and Iris Variable Aperture Collimator, Software version 3.0. |  |  | November 14, 2008 | 2008 | Accuray Inc | System may use random incorrect data to calculate dose. Resulting dose calculation error can exceed 100% of correct dose which may lead to serious patient injury. | Software design | An Urgent Advisory Notification letter dated November 14, 2008 was issued to all sites via Fed Ex. The letter identifies the issue and required actions. The user was instructed not to modify current extrapolation depth settings and to contact Accuray Incorporated Customer Support prior to making any changes.Please direct any questions or concerns regarding this issue to Accuray Incorporated Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (outside of USA) or customersupport@accuray.com. |  | 4 | 4 units | Nationwide Distribution. Product was distributed to 4 medical facilities in the US. | Terminated | December 02, 2010 | 748 | 4
Z-0473-2009 | 49539 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally, rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. |  |  | January 30, 2008 | 2008 | GE Medical Systems, LLC | Software error, computer: GE Healthcare has became aware of an issue with Error Management affecting the Innova 2100, 3100/3100 IQ &amp; 4100/4100IQ. It was reported that when the foot switch was pressed, the live monitor screen was blank and it looked like there were no X-rays being emitted. But the message "Acquisition in progress" was displayed on the DL monitor, even after the foot switch was r | Software design | Consignees sent a "GE Urgent Medical Device Correction" letter dated February 28, 2008. The letter was addressed to Hospital Administrators/Hospital Risk Management Department/Managers of Radiology-Cardiology/ Radiologists-Cardiologists. The letter described the affected products, safety Issue, short term recommendations, long term solution and contact Information. The first batch of letters was received by customers on 5/28/08. For further information, please contact GE Healthcare at 262-544-3894 |  | 427 | 427 | USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL IN, KS, KY, LA, ME, MD, MA, MI,MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY,NC, ND, OH, OK OR, PA, RI, SC, SD, TN, TX UT,VT, VA, WA,WV, and WI; the United States territory of Puerto Rico; and countries of ALGERIA, AUSTRALIA, BELARUS, BELGIUM, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ,FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN,KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND,NORWAY, PANAMA, POLAND, PORTUGAL RUSSIAN FEDERATION SAUDI AR ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SUDAN, SWEDEN SWITZERLAND, SYRIA, TAIWAN, TURKEY UNITED KINGDOM, and VENEZUELA. | Terminated | March 07, 2016 | 2959 | 1127
Z-0652-2009 | 50141 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker Endoscopy FloControl Arthroscopy Pump Integrated Tubing, Model number 0350600006, for use with the FloControl Arthroscopy Pump. |  |  | October 01, 2008 | 2008 | Stryker Endoscopy | Unable to pump fluids. Silicone tubing is stiffer than usual, resulting in motor defect message, leading to inability to pump saline through tubes for surgery. | Process control | Stryker Endoscopy initiated the recall 10/01/2008. Information gathered and recall packets sent via FedEx to all affected accounts. Packets included advisory notices and acknowledgement card, as well as shipping labels to return product. Each account was instructed to locate and quarantine product for return to Stryker endoscopy. Sales representatives received a voicemail/email notification that included a list of all affected accounts. |  | 1440 | 1440 devices | Nationwide. | Terminated | December 16, 2010 | 806 | 1440
Z-0598-2009 | 50127 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Alimed Gait Belt (Transfer Aid), Antimicrobial Plastic buckle, 54" |  |  | October 30, 2008 | 2008 | AliMed Corporation | Misassembled gait (transfer aid) belt. | Employee error | Alimed notified accounts via U.S. Postal Service (Return Receipt Requested) with a Recall Notice dated 10/30/2008 and instructs users to review their records and notify all users to whom the product may have been issued. If any of the product has been issued outside of their facility, customers are to forward the correspondence to them in it's entirety. Customers are to complete the enclosed form for return of their product (new or used) and the issuance of replacement product or credit. |  | 993 | 993 units total | Nationwide Distribution | Terminated | September 09, 2011 | 1044 | 993
Z-0628-2009 | 50147 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with versions 2.2.4, 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. |  |  | October 14, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Incorporated identified a potential issue with the Hi-Art system during the course of ongoing testing. The Operator Station Calibration panel provides access to view and modify machine specific configuration settings. Access to these settings has always been restricted to individuals with appropriate security rights, being limited to only the "Superuser" and "Field Service engine | Software design | A TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated October 17, 2008 was sent to all consignees. The letter addressed the Issue, Product Affected, Recommended Actions and Resolution. Please direct questions to the TomoTherapy, Inc. Customer Care Center by email (www.support@tomotherapy.com) or by telephone, North America: 1-866-368- 4807 and all other locations: 1-608-824-2900 or 32-2-400-44-44. |  | 215 | 215 | Worldwide Distribution --- USA including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NM, NY, OH,OK , OR, PR, PA, SC, SD, TN, TX, VA, WA, WI, WY, and DC, and countries of BELGIUM, CANADA, SWITZERLAND, CHINA, GERMANY, SPAIN, FRANCE, GREAT BRITAIN, HONG KONG, ITALY, JAPAN, KOREA, NETHERLANDS, SINGAPORE, TAIWAN and TURKEY. | Terminated | February 10, 2011 | 849 | 215
Z-0907-2009 | 50750 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Expressaire, Disposable Tourniquet Cuff 34", 1 line (18), sterile, Manufactured for Sammons Preston, A Patterson Medical Company, Cedarburg, WI.; Cuff type A70107010. Each cuff is marked ExpressAire by Sammons Preston and A70107010. |  |  | November 24, 2008 | 2008 | Virtus, Inc. | The inflation tube may come loose at the connection to the cuff during use. | Employee error | Sammons Preston was notified of the recall by letter dated 11/24/08, which instructed them to examine their inventory immediately for the specific lot numbers. The quantity on hand should be indicated on the table attached tothe letter and the info should be faxed to Virtus at 812-933-0749. Customers that may have tourniquet cuff inventory of identified lot numbers should be contacted. Once recall items have been located in quarantine, Tara Collins, Quality Technician, should be contacted at 812-933-1121 for Return Material Authorization number. |  | 3654 | 3654 | Nationwide Distribution --- including state of Wisconsin. | Terminated | October 22, 2010 | 697 | 5958
Z-0626-2009 | 49591 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pre Isotonic Personal Lubricant labeled PRE' Patient Lubricant. |  |  | September 22, 2008 | 2008 | INGfertility, LLC | Product contaminated with bacteria (Brachybacterium rhamnosum, a non pathogen gram positive cocci). | Pending | Firm telephoned all wholesale, clinic, and retail customers on 9/22/08 requesting that product be returned. |  |  | Unknown*** | Worldwide Distribution including USA and country of Ireland. | Terminated | May 02, 2012 | 1318
Z-0885-2009 | 50544 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer NEXGEN Complete Knee Solution Cruciate Retaining (CR) Provisional - Articular surface, Size Blue, 14 MM Height, nonsterile, Zimmer, Warsaw, IN; REF 5971-50-14. |  |  | December 12, 2008 | 2008 | Zimmer Inc. | Some 12 mm thick devices may have been etched and packaged as 14 mm thick, and vice versa. | Employee error | Zimmer sales force and distribution sites were notified of the recall via email on 12/15/08 and instructed to inventory, return the affected products and to provide a copy of the recall letter to those customers to whom they had further distributed the products. The 'Urgent: Device Recall" letter dated 12/12/08, described the problem and the potential risks. For additional information, contact Zimmer, Inc., at 1-800-613-6131. |  | 131 | 131 of both products. | Nationwide, Canada, Denmark, Finland, Germany, Poland, Spain, Sweden and United Kingdom. | Terminated | November 12, 2009 | 335 | 131
Z-0854-2009 | 50112 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Ruggles Surgical Instrument ACF Distraction Screws |  |  | October 22, 2008 | 2008 | Integra LifeSciences Corp. | A packaging error had occurred where 16 mm screws were packaged and labeled as 12 mm screws and vice versa. | Packaging | Urgent: Product Recall Notification letters were sent via Fed Ex on October 22, 2008. The letter asked customers to quarantine any of the affected stock on hand and contact their local Integra Sales Representative for further assistance and for arrangement of product returns. Customers were to fill out the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified in the letter. Additional questions can be directed to Christin Roach at 609-936-5413.A Urgent: Product Recall Second Notification dated 1/5/2009 was sent to customers providing them with additional information on the packaging error and included immediate and/or product used during surgery. This notification also requested that all surgeons be notified who were or may have been provided with any of the screws affected by this recall. |  | 351 | 351 units | Nationwide Distribution | Terminated | July 15, 2009 | 266 | 351
Z-0818-2009 | 50232 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier AFFINITI Humeral Stem Standard 10mm, Catalog No. 0020010. Sterile. Manufactured for Tornier, 10750 Cash Rd., Stafford, TX 77477. Tornier's Affiniti Shoulder System is a non-constrained glenohumeral prosthesis intended for use as a total or hemi-shoulder replacement system. The total shoulder consists of a metal humeral stem, a metal humeral head and an ultrahigh molecular weight polyethylene glenoid. The hemi-shoulder consists of a metal humeral stem and a metal humeral head. |  |  | October 20, 2008 | 2008 | Tournier, Inc. | Tornier, Medical has received a report of a mislabeled sterile AFFINITI head product. As a result, Tornier has decided to voluntarily recall the potentially affected lots, including two (2) lots of AFFINITI Stems. | Labeling Change Control | An Urgent Product Recall letter, dated 10/17/08, was hand delivered to end users by distributors and sales representatives beginning 10/20/08. The recall letter identified the issue and requested end users to immediately discontinue use of the affected products and return any unused affected products to Tornier. Distributors were to notify any and all accounts that may be end users or to provide this information to Tornier, whom will contact the end users. A questionnaire and return instructions were provided with the letter. The questionnaire was asked to be completed and returned to Tornier as soon as possible. For additional information, contact Tornier Customer Service at 1-888-494-7950. |  | 401 | 401 total | AR, CA, FL, NC, NE, OH, PA, TN, TX, WA, and WI | Terminated | May 07, 2011 | 929 | 401
Z-0860-2009 | 50491 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSet True Flow Series, Hemodialysis Blood Tubing with Priming Set and Transducer Protectors, with Access Flow Reversing Connector (Twister) |  |  | November 25, 2008 | 2008 | Fresenius Medical Care North America | Port may crack and separate resulting in potential blood exposure/blood loss for the patient | Process control | Fresenius Medical contacted customers by e-mail,telephone and hard copy letter advising users to immediately discontinue use and return inventory. All internal customers were emailed on 11/25/08. All external customers were US Certified Mailed on 11/26/08. All internal customers were US Certified Mailed on 12/2/08 &amp; 12/3/08. |  | 351743 | 351,743 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | September 24, 2009 | 303 | 351800
Z-0836-2009 | 50538 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomaly resulting in a patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 138 | 138 units | Worldwide Distribution including USA and countries of Canada, China, Denmark, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Lebanon, Mexico, Netherlands, Saudi Arabia, Spain, Taiwan and United Kingdom | Terminated | May 18, 2010 | 522 | 138
Z-0929-2009 | 50770 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition AS, Model number 8098027. |  |  | December 09, 2008 | 2008 | Siemens Medical Solutions USA, Inc | possible improper installation of tilt transport lock/tilt limiter | Labeling Change Control | The recalling firm received a letter dated 11/25/08 from the manufacturer informing them of the problem. As a result, the recalling firm has issued Update Instruction CT065/08/S to have the transport lock/tilt limiter on the device visually inspected. The tilted scan modes have been disabled until the modified lock/tilt limiter has been installed. |  | 42 | 42 units | Nationwide Distribution --- including states of AZ, CA, CO, FL, IL, KS, KY, MA, MI, MN, MO, NE, NY, OH, PA, SC, VA, and WY. | Terminated | September 23, 2009 | 288 | 42
Z-0842-2009 | 48760 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Endoureterotomy Stent Set, 10Fr/13 cm, 6.0Fr/24 cm, sterile, Cook Urological, Inc., Spencer, IN; Ref EUSCFS-060024-10-13; GPN REF G18132. Ureteral stent. |  |  | July 31, 2008 | 2008 | Vance Products Incorporated | The bond may break where the two materials meet, resulting in separation of the stent. | Process control | Consignees were notified via "Urgent: Recall Notice" letter dated 7/31/08. The letter described the failure, instructed that the products be isolated and returned to the firm. For additional information, contact Cook Urological at 1-812-339-2235, ext. 7241. |  | 1739 | 1739 of all products. | Nationwide, Australia, Canada, Colombia, Hong Kong, India, Ireland, Israel, Japan, Malaysia, Mexico, South Korea, Taiwan, Venezuela and Viet Nam | Terminated | November 27, 2009 | 484 | 1739
Z-0587-2009 | 49603 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8627-10. 10 mL Reservoir. The contents [STERILE] of the inner package have been sterilized by ethylene oxide gas. The implantable Medtronic SynchroMed EL programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. |  |  | August 25, 2008 | 2008 | Medtronic Neuromodulation | Medtronic issued a letter to healthcare professionals providing important safety information regarding MRI (magnetic resonance imaging) effects on SynchroMed EL and SynchroMed II implantable infusion pumps. As stated in product labeling, the magnetic field of an MRI will temporarily stop the rotor of the pump motor and suspend drug infusion for the duration of MRI exposure for all SynchroMed pump | Labeling Change Control | An Urgent Medical Device Correction, Important Information on Potential MRI Effects letter, dated August 2008, was sent to health care professionals beginning August 25, 2008. The letter described the issue, provided patient management recommendations and included enclosures: Roller Study Procedure to confirm pump function, Post MRI Pump Interrogation Guidelines and New Information Regarding Potential MRI Effects. |  | 76253 | Approx 76,253 of all devices. | Worldwide Distribution --- USA including Washington D.C., and Puerto Rico, and countries of Algeria, Aruba, Austria, Belarus, Belgium, Croatia, Cyprus Czech Republic, Denmark, Egypt, Finland, France, French Polynesia, Germany, Greece, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy , Jordan Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Slovakia, Slovenia, Sough Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State, New Caledonia, Brazil, Dominican Republic, Mexico, Hong Kong, Singapore, Canada, Australia and Japan. | Terminated | May 17, 2012 | 1361 | 76253
Z-0001-2009 | 50138 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual Multi-Slice Computed Tomography (CT) Scanner System (software). |  |  | April 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | An issue associated with the "Timed Scan" application was found. The system does not wait for the programmed time interval and starts the x-ray after selection of the scan button when certain criteria is met. | Software design | A letter was issued to customers prior to the release of the field correction during June 2008. The system software will be updated to correct the non-compliances. For information or support concerning this issue, please contact the Neusoft Medical Systems Service Support Department (helpdesk@pnms.neusoft.com) or contact Philips Medical Systems North America Co. Phillips at 425-487-7602. |  | 19 | 19 units in the US. | Nationwide Distribution --- USA including states of CO, FL, GA, IL, KY, NC, TN and TX. | Terminated | October 29, 2010 | 932 | 19
Z-0609-2009 | 50173 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | DLP(R) Pericardial/Intracardiac Sump Cannulae, 20 Fr., P/N: 12112. This product is intended for draining the pericardial sac or the cardiac chambers. |  |  | September 29, 2008 | 2008 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Separates: Six complaints were received from two locations. In all cases the sump tip assembly separated from the tubing during device removal and had to be retrieved from the patient. There were no adverse events affecting these patients. | Nonconforming Material/Component | Consignees were sent a "Medtronic Urgent Medical Device Recall Notice" dated October 1, 2008. The letter described the product and issue. It also requested its customers to return the product and to complete, sign and fax the Recall Certificate "Medtronic Recall Certificate (FCA 0902)". For further information, please contact Medtronic Cardiovascular Revascularization &amp; Surgical Therap by telephone at 763-391-9000. |  | 560 | 380 US ; 180 OUS | Worldwide Distribution --- including states of NC, PA, IA, WI, IN, VA, IL, KY, and TX, and countries of Australia, France, Hong Kong, Netherlands, and Poland. | Terminated | May 07, 2011 | 950 | 560
Z-0887-2009 | 50230 | No | ILX | Shortwave diathermy. | Physical Medicine | Physical Medicine | 510(k) | Regenesis Biomedical PROVANT Wound |  |  | September 29, 2008 | 2008 | Regenesis Biomedical Inc | In the course of device testing, the firm determined that the affected devices may, under certain unusualclinical circumstance, be unable to generate a therapeutic dose of radio frequency energy. As such,for these few devices, the effectiveness of Provant therapy may be impaired. | Finished device change control | Written notification dated September 29, 2008 was sent to consignees by overnight mail notifying them of the nonconformance and providing instructions for return of the affected PROVANT Model 4201A units to Regenesis. Receipt of the notification and return shipping of the units will be verified. Contact Regenesis Biomedical Inc. for additional information at 1-480-970-4970. |  | 7 | 7 units | WY, VA, and NC | Terminated | July 08, 2009 | 282 | 7
Z-0888-2009 | 50485 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Solar Shoulder Humeral Stem |  |  | November 13, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a visual defect in certain blister packages of certain devices. | Packaging | Recall notification letters were sent via Federal Express on November 21, 2008 to Stryker Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeon. The letters explained the problem and indicated that there are no risk factors. Contact Stryker Orthopaedics for additional information at 1-201-831-5718. |  | 322 | 322 total units - all types | Nationwide. OUS: Australia, Canada, Japan, Argentina, Brazil, Germany, Spain, Columbia, Guatemala and Sweden. | Terminated | January 20, 2010 | 433 | 322
Z-0884-2009 | 50498 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration Modular Conical Distal Stem; |  |  | November 21, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker became aware of a lot for lot mix up. | Labeling mix-ups | Recall notification letters sent via Federal Express on Nov. 21, 2008 to Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons. Letters explain the labeling mix up and the potential risks to patients if the wrong size stem is used. Stryker branches are asked to examine inventory and that of hospital locations to identify affected product. All affected product is to be reconciled on the Product Accountability form and the form faxed back to Stryker. Contact Stryker Howmedical Osteonics Corporation for additional information at 1-201-831-5718. |  | 4 | 4 | Class II Recall - Nationwide Distribution | Terminated | June 25, 2012 | 1312 | 4
Z-0904-2009 | 50694 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer TIV Self Tap Screw, 40 mm. Manufactured by Zimmer Manufacturing B.V., A Subsidiary of Zimmer, Inc., Warsaw, IN; Item No. 48231904000. NOTE: This product is not sold in the United States. |  |  | December 12, 2008 | 2008 | Zimmer Inc. | The package may contain a 40 mm long screw instead of the labeled 42 mm long screw, or vice versa. | Process control | The firm's facility in Japan was notified via letter, "Urgent: Device Recall" (dated 12/11/08) and requested to complete an enclosed Inventory Return Certification form. The letter instructed distributors to notify accounts to the user level to stop using the affected implants and return them to the Zimmer Distribution Center in the U.S. and fax (574-372-4265) a copy of the completed Inventory Return Certification form to Zimmer, Inc. For questions about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 100 | 100 | International Distribution including country of Japan. | Terminated | October 30, 2009 | 322 | 508
Z-0926-2009 | 50189 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump 2 Blockaid, 400 mL PainPump for Continuous Nerve Block, sterile, 5 per box. Stryker, Kalamazoo, MI; REF 575-100, Product Number 0575-100-000. |  |  | November 12, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | The pump may display an E-4 error message, which causes the pump to stop dispensing the pain medication. | Process design | Consignees were notified via letter dated 11/12/08 and instructed to locate and destroy the affected lots. Customers were also instructed to complete and return the enclosed Business Reply form via fax (336-499-8603) to Stryker Instruments Regulatory Department. For questions regarding the recalled product, please contact Stryker Instruments at 1-800-800-4236, extension 3808 or extension 4056. |  | 2009 | 2,009 boxes | Worldwide Distribution including USA and country of Canada. | Terminated | November 30, 2009 | 383 | 2009
Z-0930-2009 | 50169 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX SP Power Wheelchair with Stability Lock, |  |  | November 27, 2007 | 2007 | Invacare Corporation | Invacare has become aware that on some chairs, the stability lock feature may not be engaging properly or consistently. If the stability lock feature does not engage properly, the wheelchair may veer to one side or it may rock forward onto its front riggings. | Pending | On 11/27/2007 the firm sent Urgent Field Correction Needed recall notification letters and response forms to their customers informing them of the problem. They instructed their customers to immediately inspect and adjust the Stability Lock feature, distribute the enclosed update to the Owner's Manual to all end-users, and return the enclosed response form notifying them that the inspection and/or adjustments have been completed. |  | 4171 | 4171 Wheelchairs | Worldwide Distribution --- including USA and countries of Canada, Australia, Germany, New Zealand and the United Kingdom. | Terminated | August 01, 2013 | 2074 | 4171
Z-0606-2009 | 50155 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko AB Uno Lifts, Powered (non-AC) Patient Lift: |  |  | November 11, 2008 | 2008 | Liko North America Inc | Actuator of the lift has the potential to separate when the upper arm is fully extended. | Process control | Hill Rom notified customers with a Field Safety Notice, dated December 5, 2008 to notify customers of the issue, potential risks, and to inspect units per their instructions and remove units from service if the unit fails inspection. Hill-Rom will provide a permanent fix, using an "OuterTube" to be placed on all units. |  | 24070 | 24070 total units | Worldwide Distribution --- including USA and country of Canada. | Terminated | February 04, 2016 | 2641 | 24070
Z-0654-2009 | 50503 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | CAD SCIENCES LLC 3TP PrecisionPointTM Software, for biopsy guidance. |  |  | December 09, 2008 | 2008 | iCAD, Inc. | Software modules not approved for this indication. | PMA | iCAD initiated a voluntary field correction by telephone and letter on 12/5/08 by disabling these features at all affected sites. Since Precision Point and TheraMap are controlled by a software license key, iCAD will disable these features by updating the license at the affected sites. |  | 35 | 35 | Nationwide distribution to: CA, FL, HI, KS, NH, NJ. NY, MA, MO, MI ,NC, OH, SC, and WI. | Terminated | January 28, 2016 | 2606 | 35
Z-0928-2009 | 50539 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 21, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 11/21/08 of the software anomaly resulting in a patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. The letter provided the users with safety instructions to follow until the software is updated. Please direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 476 | 476 units | Worldwide Distribution including USA and country of Canada. | Terminated | September 06, 2010 | 654 | 476
Z-0663-2009 | 50510 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Ethicon Endo-Surgery ENDOPATHÂż Electrosurgery Probe Plus II 29 cm Needle Probe (Product Code EPS09) |  |  | December 09, 2008 | 2008 | Ethicon Endo-Surgery Inc | Nickel exposure: The stainless steel tips on the affected product were degrading, causing a small amount of nickel to be released, which may lead to health risks in some patients, particularly those with a known or suspected nickel allergy or sensitivity. Ethicon Endo-Surgery determined the presence of free nickel in the affected product could elicit a reaction in those patients with nickel sens | Process design | On 12/9/2008 Ethicon Endo-Surgery sent recall notification letters to their customers instructing them to stop using the product, examine all inventory to determine if they have the affected product on hand and remove the affected products from inventory. They requested that each customer complete the Business Reply Form and fax it back to Stericycle and also return any affected product to Stericycle. |  | 2461 | 13 lots, 2,448 pieces | Worldwide distribution: USA, PUERTO RICO, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, DOMINICAN REPUBLIC, EGYPT, ENGLAND, EQUADOR, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, MEXICO, PAKISTAN, PANAMA, POLAND, REPUBLIC OF YEMEN, RUSSIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD, TURKEY, UNITED ARAB EMIRATES, URUGUAY, and VENEZUELA. | Terminated | October 20, 2010 | 680 | 48372
Z-0932-2009 | 50680 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS InteXen LP Collagen Dermal Matrix, REF 72404066. Product Size: 8 x 12 cm. AMS InteXen LP is a sterile non-perforated processed porcine collagen dermal matrix. Sterile EO. American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA. |  |  | December 05, 2008 | 2008 | American Medical Systems | The product size shown on the carton labeling does not match the size on the product inside. | Labeling mix-ups | Customers were first notified by phone on 12/05/08. They were told of the situation and were requested to return the devices. Customers were also provided an Important Product Recall letter, dated 12/05/08. This letter identified affected product, told customers to discontinue use of product and to return it. An acknowledgement form was also asked to be returned. AMS Customer Service should be contacted at 1-800-328-3881 to receive a Return Authorization number and arrange for shipment of replacement product. |  | 5 | 5 | Nationwide Distribution --- including states of TN and TX. | Terminated | June 05, 2012 | 1278 | 5
Z-0883-2009 | 50119 | Yes | MVK |  |  | Cardiovascular | PMA | Monitor portion of the LifeVest device - specifically the Defibrillator printed circuit assembly (PCA) within the monitor. |  |  | October 21, 2008 | 2008 | Zoll Lifecor Corporation | The charging circuit will not shut down properly. | Component design/selection | Physicians with active patients were notified of the product removal by letter via first-class mail dated 10/21/2008. The letter describes the problem and the potential risk. For additional information, contact ZOLL Lifecor Corporation at 1.800.543.3267. |  | 1685 | 1685 units | Nationwide. OUS: Germany, United Kingdom, Italy, and France. | Terminated | November 04, 2009 | 379 | 1685
Z-0664-2009 | 50128 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Dual Taper Wedge Pressed Fit Femoral Component. Hip Prosthesis, Part number SC1439 |  |  | October 07, 2008 | 2008 | Stelkast Co | Sterility may be compromised due to breach in packaging. | Package design/selection | On 10/7/08, Stelkast telephoned the two customers of the problem and the need to return the products. |  | 6 | 6 units | USA distribution to dealers in TX and WV. | Terminated | April 01, 2009 | 176 | 6
Z-0669-2009 | 50129 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HARDY DIAGNOSTIC McFarland Standard 0.5; in vitro diagnostic, glass vial, one vial per box |  |  | September 22, 2008 | 2008 | Trek Diagnostic Systems | Incorrect result: The absorbance has changed as the standards have aged; the standards' turbidity has become denser. The recalled lot is now reading 0.15 ABS @ 625 nm. TREK'S internal specification and the manufacturer's (Hardy Diagnostics) specification is 0.08 - 0.10 ABS @625 nm. CLSI specifies 0.08 - 0.13 ABS @ 625 nm. | Nonconforming Material/Component | On 9/22/2008 TREK Diagnostic sent a recall notification letter to their customers advising them of the problem with the product. They instructed their customers to inventory their stock for any remaining recall lot and immediately destroy. |  | 4 | 4 vials | USA distribution in OH and IL. | Terminated | July 06, 2010 | 652 | 4
Z-0668-2009 | 49711 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus CS Acute Care Troponin I TestPak, Catalog number CCTNI, in vitro diagnostic. |  |  | June 23, 2008 | 2008 | Dade Behring, Inc. | Incorrect result: Potential for false elevation of CTnI results | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice dated 6/23/08 with instructions to immediately discontinue use and discard the affected lots. Customers were also instructed to contact Siemens Customer Service Center at 800-241-0420 to request replacement product at no charge. |  | 5868 | 5868 cartons | Worldwide distribution: USA, Singapore, Saudi Arabia, Japan, Germany, Canada, Bahrain, Australia, Korea, and India. | Terminated | January 29, 2009 | 220 | 5868
Z-0837-2009 | 50140 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV Peripheral Imaging Catheter, model number H749364560, catalog number 36456, manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 09, 2008 | 2008 | Boston Scientific Corporation | Product is mislabeled as having a shelf life of 3 years, where the actual shelf life is 1 year. | Labeling mix-ups | Recall initiated on October 6, 2008, and notification letters sent on October 8, 2008. Letters were sent via FedEx. Responses are being tracked along with delivery verification. For additional information contact Boston Scientific Corporation at 1-508-652-5594. |  | 1033 | 1033 units | Worldwide Distribution --- USA including Puerto Rico and country of Panama | Terminated | July 27, 2009 | 291 | 1033
Z-0666-2009 | 50137 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Maximo II VR D284VRC, for ventricular pacing and defibrillation. |  |  | July 25, 2008 | 2008 | Medtronic Inc. Cardiac Rhythm Managment | Maximo II VR ICD, DR ICD and CRT-D inadvertently include the Capture Management algorithm (Atrial/Right Ventricular/Left Ventricular Capture management). The Capture management algorithm monitors daily pacing thresholds and may increase or decrease pacing amplitudes in response to patients' needs. Although the algorithm is operating correctly, it is not accessible to the clinician for reprogramm | Process design | A Medtronic Inc. letter dated July 30, 2008 was sent to Physicians. The letter described the product and problem. The letter advised to contact Technical Services for specific questions regarding the care of patients or how to program the Maximo II devices. The letter instructed to collect and hold all non-implanted Maximo II devices and that a Medtronic Representative will contact the doctor to make arrangements to update the programmer 2090. |  | 376 | 376 | No US distribution. OUS to include: AUSTRIA, DENMARK, FINLAND, GERMANY, ITALY, SPAIN, SWEDEN, UK, GREECE, HONG KONG, INDIA, JORDAN, and SAUDI ARABIA. | Terminated | May 07, 2011 | 1016 | 1108
Z-0935-2009 | 50479 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard 3D Max Mesh, Left large , Sterile |  |  | December 08, 2008 | 2008 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Product labeled as Left large, may contain a Right Large | Mixed-up of materials/components | Davol issued a letter notification dated 12/25/08 via Federal Express requesting accounts discontinue use, remove product from inventory and return product.For additional information, contact Davol at 1-401-463-7000. |  | 232 | 232 units | Nationwide | Terminated | May 02, 2012 | 1241 | 232
Z-0913-2009 | 50749 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 20Fr (6.7 mm) Safety PEG Kit, Push Method, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6653, inner pouch UPN M00566530 and outer carton UPN M00566531. |  |  | December 30, 2008 | 2008 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Other | Consignees were notified by "Urgent Medical Device Recall" letter dated 12/30/08. The letter described the problem with the product, instructed customers to determine whether they store these products in the outer carton or the inner pouch at their facility, to locate all affected products, and to return them to the firm. For additional information, contact Boston Scientific Corporation at 1-508-683-4678. |  | 2460 | 2460 of all products | Nationwide. | Terminated | October 30, 2009 | 304 | 2460
Z-0670-2009 | 44850 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Air Bubble Detection Module, Terumo Cardiovascular Systems, Ann Arbor, MI; Catalog number 802110. |  |  | November 20, 2007 | 2007 | Terumo Cardiovascular Systems Corporation | The system may malfunction due to t-filter corrosion on an internal circuit board. | Nonconforming Material/Component | U.S. consignees were notified of the problem by letter on 11/16/07. International affiliates were informed via email on 11/21/07. Consignees were informed that Terumo would schedule appointments to install new pressure modules and reconfigure the systems beginning in Feb. 2008 or that parts would be shipped to consignees wishing to perform the work themselves. For additional information, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 1219 | 1219 (total) | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | August 06, 2009 | 625 | 1219
Z-0867-2009 | 50516 | [6 empty] | REF 01-6851 Office System 1.5/IMF Small Office Set, 156MMX125MM, (6.4"X4.9") Polypropylene, Ultem, Radel, AL, SS. Non Sterile Product. W.LORENZ SURGICAL, Walter Lorenz Surgical, Inc. 1520 Tradeport Drive |  |  | August 06, 2008 | 2008 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling the Lactosorb Endobrow Lift Complete, Office Fixation Kit, 1.5/IMF Small Office Set and IMF Set trays containing instruments associated with the Lactosorb and IMF systems. The finished trays were packaged and shipped to facilities without the instructions for use and without the single use symbol on the packaged tray label. | Other | An IFU and correction letter will be issued to the field. Each distributor will receive instructions for delivering the correction letter to all customers. The customer letter will contain a reconciliation form to be returned once the customer receives the IFU for the twist drill. Contact Biomet Microfixation for additional information at 1-800-874-7711. |  | 33 | 33 | Distributed domestically throughout the U.S. Distributed internationally to Argentina, Australia, Canada, Columbia, Hong Kong, Japan, Mexico, New Zealand, Puerto Rico, Taiwan, Thailand, Denmark, France, Netherlands, Greece, Italy, Norway, Finland, Spain, Czech Republic, and Venezuela. | Terminated | September 11, 2009 | 401 | 420
Z-0927-2009 | 50515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV SA Linear Accelerator with syngo RT Therapist Version 4.1 build 67 and build 122. Manufactured by Siemens Medical Solutions Oncology Care Systems, Concord, CA. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | If a plan contains a sequence of beams, including at least 1 ARC treatment beam and if during the setup stage of a beam following an ARC beam in the delivery sequence a termination occurs, and the user selects the immediate resumption feature, syngo RT Therapist Express systems will display and record a second delivery of a prior beam and not the delivery of the current beam in certain circumstanc | Software design | An Urgent Medical Device Customer Safety Letter dated October 23, 2008 was issued to consignees via certified mail. The letter identified the issue and preventive measures for consignees. Monitoring will be done via return cards. Follow-up to be conducted via telephone. |  | 36 | 36 units | Worldwide Distribution including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium, New Zealand, France, Malaysia, Spain, and Canada. | Terminated | December 09, 2010 | 777 | 36
Z-0903-2009 | 50512 | No | OEX | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Spectranetics SLS II 14Fr. Laser Sheath Kit, REF 500-012, Sterile EO, individually packaged, Lot number 080314B, Spectranetics, 96 Talamine Court, Colorado Springs, Colorado 80907. |  |  | October 15, 2008 | 2008 | Spectranetics Corporation | Catheter pouch was mislabeled with the wrong size. | Error in labeling | Consignees were notified by sales representatives who visually inspected their inventory. The sales representatives were given a form to fill out to confirm whether the consignee had the catheters with the serial number still in inventory. The forms were faxed back to Spectranetics indicating if the catheters were used or not and if they were not used, were they labeled correctly. For assistance, contact Spectranetics Corporation at 1-719-447-2319. |  | 24 | 24 units | CA, FL, MI, NC, NY, OH, PA, TX, and VA. | Terminated | February 26, 2009 | 134 | 24
Z-0946-2009 | 50521 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Epicor LP Connecting Cable, CC-6-LP-CE, Catalog Number 12263, an accessory to the Epicor Cardiac Ablation Control System, manufactured by St.Jude Medical, Sunnyvale, CA |  |  | December 09, 2008 | 2008 | St. Jude Medical Inc | Firm has received multiple reports of difficulties connecting the product and/or the presence of error code messages occurring with the use of the connecting cable and the Ablation Control System unit. | Other | The firm issued notification by "Urgent Product Recall" letter dated 12/4/2008, to all consignees. The letter describes the issue, potential risk and states that the product should be returned. For additional information, contact St Jude Medical at 800-722-3423. |  | 158 | 158 units distributed. | Worldwide - NY, IN, NC, UT, OR, CA, IL, MI, TX, TN, and FL and country of Belgium. | Terminated | December 28, 2010 | 749 | 158
Z-0931-2009 | 50523 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Low Profile U-Joint Driver for the SynFix Mini-Open System, Catalog number 03.802.231. |  |  | November 20, 2008 | 2008 | Synthes Spine | screwdriver tip may break off and lodge in the screw head | Device Design | The recalling firm emailed all Spine Sales Consultants on 11/20/08. They were instructed to immediately examine their stock and discuss appropriate technique with surgeon users. In addition, a Medical Device Correction letter dated 11/19/08 was issued via USPS certified mail to the medical facilities. This initial letter stated the issue and requested the users take the following actions: discard all SynFix Technique Guides and replace with the updated literature contained in the mailing packet with the letter, and is identified by code J7022-E (Rev. E); to ensure receipt and understanding of the information, the Verification Section of the letter needs to be completed and returned to Synthes by either fax: 610-719-5120 or scan/e-mail: lewis.lynne@synthes.com. An Urgent: Medical Device Recall letter, dated 1/16/2009, was also sent out. This latter letter requested that users examine their stock immediately and remove any of the affected devices from their inventory. Customers are also to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the Verification Section; and return the letter with the recalled product to the address listed. There are also actions for those who do not have the affected product. Questions should be directed to 800-620-7025 x 5453 or 610-719-5453. |  | 381 | 381 units | Class II Recall - Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA ,SC, TN, TX, UT, VA, WA, and WI, and countries of Canada and Switzerland. | Terminated | October 27, 2009 | 341 | 381
Z-0949-2009 | 50787 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator System Equipped with 160 MCL Accessory Holder, Manufactured by Siemens Medical Solutions, Oncology Care System, Concord, CA |  |  | December 20, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The range of tolerance is to high for the accessory holder, therefore, Wedges and other accessories do not latch to the locking mechanism correctly. | Device Design | On 12/20/08, the firm issued letters via certified mail to direct consignees, informing them of the affected product and providing instructions on the recall.Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for additional information. |  | 43 | 43 units | Nationwide and Worldwide | Terminated | December 09, 2010 | 719 | 43
Z-0933-2009 | 50756 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Ergolift 400 Lbs. Floor Lift, Manufactured by BHM Medical Inc., 2001 Tanguay, Magog, QC, Canada, J1X 5Y5; model/part numbers ERGOLIFT and ERGOLIFT-2 |  |  | December 08, 2008 | 2008 | Arjo, Inc. | The manufacturer, BHM Medical, has received reports of structural failure of the mast during use. | Process control | Arjo sent the Urgent Device Recall letter dated 12/22/08 and Field Safety Notice (FSN) dated 12/8/08 to their Ergolift customers on 12/22/08. The accounts were informed of the potential structural failure of the mast on the BHM Medical Ergolift. The letter instructed the accounts to discontinue use of the lifts immediately. Those lifts with serial numbers ERLI-0001 and ERLI-0899 will be taken out of service permanently, and will be replaced with a new device at a significant discount. Those lifts with serial numbers ERLI-0900 and ERLI-1717 will be corrected by a reinforcement kit installed by a qualified technician. The accounts were requested to complete and return to Arjo Inc. within 10 days the enclosed customer response form indicating that they have received the notification, indicating the serial numbers on hand needing correction, and whether they need any replacement lifts. The accounts will be contacted by an Arjo technician within weeks to arrange for the correction of the lifts and/or an Arjo representative for the replacement lifts. |  | 220 | 220 lifts | Nationwide Distribution. | Terminated | January 21, 2011 | 774 | 220
Z-0948-2009 | 50751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communications System) Versions 4.1.x up to and including 4.1.51.3, manufactured by Philips Healthcare Informatics, Foster City, CA. |  |  | December 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Product may display clinical reports with the study date and time in GMT format, which may not be obvious to the end user, who may expect to see the local time in the report. | Software design | Product notification was issued on 12/29/2008, instructing users to identify affected products, and instructing users to refer to image overlays for local study date and time information. The firm is planning a field correction expected to be available in "Q1 2009". For additional information, contact Philips at 650-293-2624. |  | 40 | 40 units | International - Italy, Denmark, Germany, Netherlands, Sweden, Spain, Portugal, Switzerland, Spain, Turkey, France, Saudi Arabia, Australia, Thailand, New Zealand, Japan, and Brazil. | Terminated | November 29, 2010 | 700 | 40
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4
Z-0923-2009 | 50184 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump, Two Site Infusion Kit with ExFen 2.5", 120 mL Infusion Kit, sterile, 5 per box. Manufactured for Stryker Instruments, Kalamazoo, MI. Product Number 0500-166-025, REF 500-166-25. |  |  | November 07, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | There is a potential for the pump to overinfuse. | Process control | A "Medical Device Recall Notification" letter dated November 11, 2008 was issued to consignees. The letter identified the problem and instructed consignees to locate and destroy all affected products. Customers were also instructed to complete the enclosed Business Reply form and fax it to Stryker Instruments Regulatory Department at 866-521-2760. For questions regarding the recalled product, contact Stryker Instruments at 1-800-800-4236, extension 4056 or extension 3808. |  | 118 | 118 boxes of all products. | Nationwide Distribution. | Terminated | December 01, 2009 | 389 | 118
Z-0794-2009 | 48991 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | King Systems Universal Flex2 Breathing Circuit, Latex Free, 50 per case, King Systems Corporation, Noblesville, IN; REF # DF477S-6121Z. |  |  | August 01, 2008 | 2008 | King Systems Corp. | The inner line may be severed or broken into two pieces, thus hindering the delivery of oxygen and anesthetics to the patient and possibly elevated patient carbon dioxide levels. | Pending | Consignees were notified by an Urgent Product Recall letter on or about 7/29/2008 to cease sales, notify their customers to the hospital level, and to return product on hand. |  | 500000 | Approx. 500,000 units of all products. | Worldwide Distribution - including USA and countries of Australia, Colombia, Germany, Netherlands, and Peru. | Terminated | November 12, 2009 | 468 | 500000
Z-0939-2009 | 50192 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | VOYAGER NC Coronary Dilatation Catheter, 4.5 X 12.0 mm. |  |  | November 05, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The recall was initiated after Abbott Vascular received reports that, in a small percentage of cases, the protective, yellow balloon sheath of the 4.5mm diameter VOYAGER NC Coronary Dilatation Catheter has been tight. If unnoticed by the user, there is a potential to cause damage to the balloon dilatation catheter during removal of the sheath. The balloon sheath covers the balloon, maintaining th | Package design/selection | An Urgent Device Recall letter (dated November 6, 2008) with an Efficacy Check Reconciliation form was issued and hand-carried to affected customers by their respective Sales Representative who assisted in locating and returning the recalled units. The letter described the problem and informed customers that Abbott Vascular will replace returned recalled product with like product or with VOYAGER NC sizes when they become available. Additionally, customers were instructed to contact their Abbott Vascular local account representative to review their inventory, complete and return the attached Efficacy Check Reconciliation Form and return affected products to Abbott Vascular. |  | 126 | 126 units | Worldwide Distribution. | Terminated | August 13, 2012 | 1377 | 224
Z-0956-2009 | 50204 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | HR90K Platinum Series Headpiece Model Number AB-7300-001 ADVANCED BIONICS LLC 12740 San Fernando Rd. Sylmar, CA 91342 USA. The HiRes90K headpiece is used in conjunction with speech processors manufactured by Advanced Bionics. |  |  | July 24, 2008 | 2008 | Advanced Bionics LLC | Mistuned headpiece - Advanced Bionics recently became aware of a tuning issue with some Platinum Series and S-Seriesheadpieces shipped from May 2007 through April 2008. Tuning refers to the process through which the headpiece's communicationfrequency is adjusted to match that of the internal device. Potential clinicalsymptoms include:-No lock-Intermittent lock-Implant locks, but cann | Employee error | Advanced Bionics is notifying clinicians of this situation so that they can provideappropriate clinical management. The notification letter sent to clinicians in theUnited States was sent via FedEx so that delivery information could be confirmed. The letter advises that Advanced Bionics recently became aware of a tuning issue with some Platinum Series and S-Series headpieces shipped from May 2007 through April 2008. Auria and Harmony headpieces are notaffected by this issue. Tuning refers to the process through which the headpiece's communication frequency is adjusted to match that of the internal device.Since discovery of this issue firm has conducted an extensive technical analysis showing that the use of these headpieces does not pose any safety risk to the patient. While there has been no increase in thenumber of headpiece-related complaints, firm advises of the potential clinicalsymptoms:-No lock-Intermittent lock-Implant locks, but cannot measure impedances-Reduction in battery life, primarily for C1 patients with thick flaps.Per the letter, as long as the system maintains lock, there is no degradation in sound quality orstimulation. Further, the impact on battery life for CII and HiRes 90K patients is minimal.Medical professionals are advised that If they have any patients who display any of these clinical symptoms that they attribute to an affectedheadpiece, to advise the patient to contact Advanced Bionics Customer Service at 877-244-9541 and they will send them a replacement promptly. |  | 2049 | 2,049 units | Nationwide | Terminated | March 23, 2012 | 1338 | 4617
Z-0969-2009 | 50758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | January 12, 2009 | 2009 | GE Healthcare Integrated IT Solutions | If a user site created a custom default display protocol (DDP) to display the current exam versus the historical exam and the user-defined DDP fails, the system displays the exams using a GE default hanging protocol. This may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 1/12/09 of the software anomaly resulting in a patient safety issue involving Default Display Protocols. The issue may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 778 | 778 units | Nationwide, Canada, Bermuda, Mexico, Ecuador, Chile, and Brazil | Terminated | September 06, 2010 | 602 | 778
Z-0934-2009 | 50134 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Connecting Straps &amp; Belts/Locking Cuffs. |  |  | September 08, 2008 | 2008 | J T Posey Company | The recall was initiated after an investigation conducted by the Posey Company determined that a patient may be able to catch the belt on a stationary object (wall, pipe, etc.) and use their body force to tear or break the strap, creating freedom of arm movement. In some patient populations, such freedom of movement could pose an unwanted hazard of serious injury to a patient or caregiver. | Other | An "Urgent Medical Device Recall (Correction)" letter dated August 15, 2008 was mailed via US Postal Service to affected customers. The letter described the issue, instructed customers to secure a replacement label with the correct application technique over the picture in the 2008 Posey Catalog and provided a corrected Posey Locking Cuffs Instruction Sheet (19273). The Posey Company also provided updated literature for the Posey Locking and Non-Locking Connecting Straps 19270 and 19254.The Posey Company asked customers to complete and return the "Reply Fax-Back Acknowledgment" Form via fax (626-443-5418) or e-mail (vcimmarusti@posey.com) within 7 days. Contact your Sales Representative with any questions or concerns. |  | 51200 | 51,200 (Approximately) | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, France, Germany, Ireland, Italy, Korea, Lithuania and Netherlands. | Terminated | April 16, 2012 | 1316 | 51200
Z-0936-2009 | 50774 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker AVS AL Spacer Trial Inserter; |  |  | September 08, 2006 | 2006 | Stryker Spine | DESIGN FLAW: Issues were reported concerning fracture of handle, difficulty in removing inserter and inserter difficult to lock. | Device Design | Market Withdrawal letters were sent out on September 8, 2006 to all branches who received device. The letter asked customers to examine their inventory and hospital consignment locations to identify the product. Also, customers needed to reconcile all products on the attached Product Accountability Form and fax a copy back. The affected products were to be retrieved and returned using the "Product Returns Procedure". |  | 37 | 37 | Nationwide Distribution --- including state of NJ. | Terminated | February 05, 2009 | 881 | 37
Z-0961-2009 | 50545 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROMOS Reverse, Glenosphere Loosening Adapter, REF 24 000 047, S&N; 75005271, Rx only, QTY(1), non-sterile |  |  | November 21, 2008 | 2008 | Smith &amp; Nephew Inc | The product was breaking during use. | Other | All affected Smith &amp; Nephew International Distributors were notified via email on 11/21/2008. They were instructed to immediately notify their accounts concerning the recall. Contact Smith &amp; Nephew Inc at 901-399-6771 for assistance. |  | 51 | 51 units | Germany, France, Italy Austria, Switzerland, Belgium, United Kingdom, Greece, Australia, South Africa, Spain, and Cypress | Terminated | November 27, 2012 | 1467 | 210
Z-0962-2009 | 50699 | No | NRA |  |  | Orthopedic | PMA | Biomet Oxford Unicompartmental Knee, Phase 3, Right Medical Tibial Tray, Standard Size B, CoCr, sterile, Biomet UK Ltd, Bridgend, UK; . |  |  | December 19, 2008 | 2008 | Biomet, Inc. | Packages labeled as "right" actually contain a "left" component. | Packaging process control | Biomet salesmen were notified by phone and instructed to retrieve the products. One customer was notified by letter dated 12/19/08 that use of the product should cease and that a salesman would pick up the product. A second customer had implanted the product and was hand delivered a letter dated 12/19/08 instructing that the patient should be monitored and symptoms treated as necessary. Contact Biomet at 574-362-3983 for assistance. |  | 2 | 2 | Massachusetts and West Virginia. | Terminated | July 07, 2009 | 200 | 2
Z-0968-2009 | 50677 | No | NPQ | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | NFlex Stabilization System 150 mm, Catalog number NF60150-1. |  |  | November 07, 2008 | 2008 | Synthes Spine | the dimensional specifications are incorrect | Nonconforming Material/Component | The recalling firm issued an e-mail on 11/18/08 to their sales force to inform them of the problem and the need to remove the product.The recalling firm issued Medical Device recall letters dated 11/19/08 to their customers to inform them of the problem and the need to return the product. Contact Synthes at 800-620-7025 ext. 7025 or 610-719-5750 for additional information. |  | 1508 | 1508 (total) | Worldwide: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV. OUS: Germany, Denmark, Switzerland, and Korea. | Terminated | October 05, 2009 | 332 | 1508
Z-0872-2009 | 50676 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SERVO Ventilator 300A (SV300A); |  |  | December 19, 2008 | 2008 | MAQUET Inc. | SERVO Ventilator 300A could cause hypoventilation if the use of the PC or PRVC mode with the Automode function set to "On" in Neonatal and Pediatric patient ranges with small tidal volumes and short set inspiratory rise time. | Device Design | URGENT: Device Correction letters were sent out on December 22, 2008. Letters alert users to the potential problem and the situation in which it may occur. Letter indicate the set inspiratory rise time should be set in the range of 7-10% and that the humidifier chamber water level be maintained according to manufacturers instructions. The letter is enclosed as a supplement to the SV300A Operating Manual. Consignees were directed to complete a "Device Correction-Customer Fax Response" to indicate number of units of the affected product and that they followed the instructions for including the attached supplemental instructions to the operating manual. For questions about the recall, contact your local Maquet, Inc. representative. |  | 3688 | 3,688 units in US | Nationwide | Terminated | July 08, 2009 | 201 | 3688
Z-0945-2009 | 50755 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Polling Procedure version 1.00.09, used in IMPAX 6.2.1 and 6.3.1.; Catalog numbers: E18R4000, L5NQJ000, L57OP000, LBCDB000, and E2ZCK000. |  |  | December 24, 2008 | 2008 | AGFA Corp. | Polling procedure may fail when multiple merge/split is carried out. | Software design | All sites potentially affected by this product were notified of the issue directly via telephone. Each site was directed to observe a ten minute polling window as an interim mitigation until the normal mitigation by software Service Update is performed at the site by Agfa Support. An Urgent Safety Notice was sent via FED-EX on/about 12/24/2008 to all sites. The letter describes the potential issue and both interim and formal mitigation. Acknowledgment, via FAX-Back, that the information was received and understood was requested from the sites. |  | 5 | 5 units | Nationwide Distribution --- including states of FL, IL, OR, SC, and SD. | Terminated | April 01, 2010 | 463 | 5
Z-0970-2009 | 50759 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 20, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Unread exams may inadvertently get marked as dictated in PACS workstation when the user signs off the report in Talkstation and opens the next exam. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation with Talk Extend software customers via letter (Urgent Medical Device Correction) dated 11/10/08 of the software anomaly resulting in a patient safety issue involving Talk Extend, which could contribute to a delay of care or misdiagnosis due to delay in communicating pertinent exam results or failure to treat due to lack of an attached dictated report by the radiologist for physician review. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 22 | 22 units | California, Florida, Illinois, Iowa, Louisiana, Oregon, Pennsylvania, Texas, Utah, Virginia and Wisconsin. | Terminated | February 03, 2010 | 440 | 22
Z-0624-2009 | 45967 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-Arm Intraoperative Imaging System |  |  | November 27, 2007 | 2007 | Medtronic Navigation, Inc | Navigation accuracy problem: Accuracy may fall outside of specification when used in combination with the Medtronic Synergy Experience StealthStation System Spine Software with the O-Arm Intraoperative Imaging System and tracker. | Software design | Medtronic Navigation notified Field representatives on 11/27/07 to perform site visits to implement FCA 2007-005 - O-Arm Intraoperative Imaging System Accuracy, and deliver the Customer Letter notification. Confirmation of the visit was documented. |  | 54 | 54 units | Worldwide Distribution: USA, Canada, Czech Republic, Denmark, Germany, Netherlands, S. Korea, Poland, South Africa, and Singapore | Terminated | January 22, 2016 | 2978 | 54
Z-0540-2009 | 49114 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | MicroScan Synergies Plus Dried Overnight Gram Negative MIC/Combo Panels, 20 panels per box. Medical device designed for use in determining quantitative and/or qualitative antimicrobial agent susceptibility and/or identification to the species level of colonies, grown on solid media, of rapidly growing aerobic and facultative anaerobic gram-negative bacteria. |  |  | July 31, 2008 | 2008 | Siemens Microscan | Falsely lowered imipenem MICs may result, leading to major errors in results. | Other | Consignees were notified on July 31, 2008 by letter (Urgent Device Correction) sent via FedEx overnight. International consignees were notified by Siemens Healthcare Diagnostics associates in affected countries. The letter described the problem and provided users with safety instructions to follow until an update is provided. Contact Siemens Healthcare Diagnostics Inc. at 916-374-3075 for additional assistance. |  | 569375 | Total of both Negative and Synergy: 569,375 boxes | Worldwide | Terminated | November 29, 2010 | 851 | 569375
Z-0650-2009 | 50178 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical Stair-PRO Evacuation Chair civilian use stair chair, patient transport device, Model 6253. |  |  | November 18, 2008 | 2008 | Stryker Medical Div. of Stryker Corporation | Handle could come off. The bolts may back out over time, resulting in a risk of dropping the patient and/or injury to transport personnel, a delay in medical treatment, or a delay in evacuation. | Process control | Consignees were notified by Urgent Medical Device Correction letter dated 11/18/08 and given the serial numbers of their units, an upgrade kit and instructions for performing replacement of the bolts. Consignees were informed Stryker will perform the upgrade if requested. |  | 102 | 102 | Worldwide Distribution: USA, India, Netherlands and United Kingdom. | Terminated | October 27, 2010 | 708 | 213
Z-0951-2009 | 49474 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS Centrifuge Module (Hettich Robotic centrifuge); List Number: 07L02-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | September 09, 2008 | 2008 | Abbott Laboratories, Inc | ACCELERATOR APS Centrifuge Module being recalled due to potential for micro-cracks to form over time in the metal centrifuge rotor buckets that may cause the centrifuge buckets to disconnect from the rotor in the centrifuge during centrifugation. | Device Design | Notification letters sent to consignees on 09/09/08. Consignees provided instructions for inspecting the buckets for cracks and were asked to contact customer service reps for replacement buckets. Consignees also asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. |  | 6 | 6 | Nationwide Distribution --- including states of TX, UT, and VA. | Terminated | October 04, 2010 | 755 | 6
Z-0950-2009 | 48172 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect / Aeroset Clinical Chemistry Iron/Magnesium Calibrator, List Number: 1E69-03, Lot Numbers: 54187M200 |  |  | April 24, 2008 | 2008 | Abbott Laboratories Inc. | An incorrect Magnesium calibrator 2 value (CAL 2) was provided for Clinical Chemistry Iron/Magnesium Calibrator lot number 54187M200. Use of the incorrect Magnesium CAL 2 value causes quality control and patient magnesium results to be reduced by as much as 17.24% above 0.8mEq/L (1.0 mg/dL or 0.4 mmol/L). | Error in labeling | On April 24, 2008, a Product Correction letter with attached Customer reply form was mailed via US Postal Service to all affected customers that have received the affected lots. Customer reply form were included in all US customers letter. The letter instructed customers to:1. Identify the lot number of Clinical Chemistry Iron/Magnesium Calibrator currently used in your laboratory2. If you have lot number 54187M200, cross out the incorrect Magnesium CAL 2 value in the value sheet, commodity 30-3535/R1, June 2005. Place a copy of this letter in all calibrator kits with this lot number.3. Enter the revised Magnesium CAL 2 value provided in this letter in your system. NOTE: Ensure that the value used corresponds to the units of concentration used in your laboratory. *ARCHITECT cSystems - Refer to Enter a calibrator concentration (cSystem) in Section 2, Installation procedures and special requirements of the ARCHITECT System Operations Manual. *AEROSET - Refer to Calibrator/Control Configuration in Section 2, Installation Procedures and Special Requirements of the AEROSET System Operations Manual.4. Calibrate the Magnesium assay. Verify the calibration with at least two levels of controls according to the established quality control requirement for your laboratory.Customers were also instructed to please retain this communication for future reference and if they have forwarded any Clinical Chemistry Iron/Magnesium Calibrator to another laboratory, please provide a copy of this letter to them. An 800 number was provided in the recall letter for any questions the affected customers or any of the health care providers they serve may have. |  | 1287 | 1287 | Worldwide Distribution --- USA including states of AL, AZ, CA, CT, FL, GA, IA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, and WI, and countries of Mexico, Chile, Puerto Rico, Canada, Thailand, Curacao, Guatemala, Argentina, Dominican Republic, Germany, New Zealand, Jamaica, Ecuador, Brazil, El Salvador, Hong Kong, and Bermuda. | Terminated | July 09, 2009 | 441 | 1287
Z-0971-2009 | 50790 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | 3.7 V Li-Polymer Battery, used in Precess MRI Compatible Patient Monitoring System and Achieva, Intera and Panorama 1.0T. REF 9065 or REF 9067. Invivo Corporation, Orlando, FL 32826 |  |  | August 13, 2008 | 2008 | Invivo Corporation | Battery Failure: The Invivo Lithium Polymer Wireless Module Battery Part Number 9065 and the Lithium Polymer Fiber Optic Module Battery Part Number 9067 is being recalled because the battery, itself, when dropped from 3 feet/1 meter onto a hard surface can produce a flame and then emit smoke. The firms investigation shows that the battery failure does not occur when a module is dropped with a ba | Process design | Urgent Medical Device Recall Notice letters, dated August 2008, and response forms were mailed to each consignee. Distributors were instructed to forward copies of the recall letter and response forms to their customers. The letters identified the affected products, gave the reason for the recall, the hazard involved, how to identify the product, actions to prevent the failure, and procedures to mitigate risk if failure occurs. The letter stated that Invivo would provide replacement batteries but until replacements were provided, customers should handle batteries over an elevated area and to make sure batteries are handled in the absence of paper or other flammable materials. If customers wish to remove the battery from the module in order to conserve battery life, they are to disengage the battery from the module and place the battery on a table or in the battery charger. Do not allow the battery to be disengaged and remain partially inserted in the module as this may cause the battery to slip from the module and fall to the floor. |  | 8476 | 8,476 | Worldwide Distribution | Terminated | November 19, 2010 | 828 | 8476
Z-0952-2009 | 50209 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | EPOC BGE Test Card |  |  | October 16, 2008 | 2008 | Epocal | elevated range of potassium levels | Other | The recalling firm visited the hospital on 10/18/08 to conduct an investigation. At the end of the visit, the firm representatives removed the remaining cards from the hospital and returned to the manufacturing location. |  | 1000 | 1000 test cards | Nationwide Distribution --- including state of CA. | Terminated | February 06, 2009 | 113 | 1000
Z-0973-2009 | 50848 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Sterile Field Saline Flush; |  |  | January 13, 2009 | 2009 | Excelsior Medical Corp | Sterility Issue: There is the possibility that the affected lots have pouches that were not sealed properly and could potentially compromise the external sterility of the syringes. | Packaging process control | An Urgent Medical Device Recall letter (dated January 13, 2009) was sent to consignees via certified mail, with return receipt requested. The letter provided the reason for the recall and instructions to discontinue use and return (or destroy) the recalled product. The letter also instructed customers to complete the enclosed Customer Reply Form and fax (866-360-6211) to Excelsior Medical Corporation. If the affected product is destroyed on site, fax a copy of the certificate of destruction with the completed Customer Reply Form. Arrangements will be made by Excelsior Medical Corporation for the return of recalled product. Please direct questions to Excelsior Medical Customer Service at 800-487-4276. |  | 341580 | 341,580 units | Nationwide Distribution. | Terminated | August 06, 2009 | 205 | 341580
Z-0635-2009 | 50227 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo TenderFlow Pediatric Arterial Cannulae, 3.3mm (10 Fr) OD, with 1/4 inch non-vented connector, 9 inch (22.9 cm) long; Catalog no. 813572. |  |  | October 16, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The introducer or stylet is difficult to remove from the pediatric arterial cannula body, which may result in the need to remove the cannula and replace it. | Device Design | An Urgent Medical Device Recall letter dated 10/16/08 was sent to consignees advising them of the problem; requesting that they immediately verify whether or not they had unused cannulae with the affected lot numbers; requesting that they isolate any affected product to prevent use; complete the Customer Response Form, indicate on form if it is medically necessary to use product without introducer; fax Customer Response Form to Terumo CVS at the number listed; and Terumo CVS would contact customers to arrange for the return of unused product. The letter stated that credit would be issued for any unused TenderFlow cannulae that were returned. The quantities to be returned should be indicated on the Customer Response Form. Consignees were advised that, if it was medically necessary to use the affected cannulae, they should be used without the introducer and in accordance with instructions for such use. |  | 288 | 288 (170 plus 78 samples to US and 40 international) | Worldwide Distribution --- including USA and countries of Australia, Belgium and Canada. | Terminated | August 04, 2009 | 292 | 2892
Z-0974-2009 | 50887 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Zimmer Dental Calcitite 4060-2 Dense Nonresorbable Hydroxylapatite Bone Grafting Material, 4 ampules, 0.5g each, 40-60 mesh size, sterile. Zimmer Dental, Carlsbad, CA; Item # 0003. |  |  | January 22, 2009 | 2009 | Zimmer Inc. | Fragments of the glass vial may be present in some of the packages from these lots. | Packaging process control | The consignee in Japan was notified by e-mail and "Urgent: Device Recall" letter dated January 22, 2009 to stop using the product and to return it to the firm. |  | 225 | 225 kits | International Distribution including the country of Japan. | Terminated | October 23, 2009 | 274 | 225
Z-0972-2009 | 50825 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Endoscopy brand 3.5mm Super 90-S SERFAS Energy Suction Probe; |  |  | June 20, 2008 | 2008 | Stryker Endoscopy | The probes may have inadequate suction due to obstruction, which can result in unintended ablation due to reduced visibility within the surgical area. | Process control | On 6/20/08, the firm initiated the recall and its notification was via Urgent: Device Removal letters explaining the reason for the recall. The letter recommended that recipients of the letter inspect their inventory for the part number/lot numbers listed and quarantine any of the affected product. Affected products should be returned to Stryker Endoscopy using the Pre-paid shipping label provided. Customers should call Stryker Endoscopy's Customer Service Department at 1-800-624-4422 for replacement product. Importantly, customers should sign and complete the enclosed "acknowledgment of receipt" card and return it to Stryker Endoscopy or fax it to 408-754-2521 with the customer's account info included. If there are any questions regarding the notice, a Regulatory Affairs Representative should be contacted at 408-754-2161 or 408-754-2000. |  | 2463 | 2463 units | Worldwide Distribution. | Terminated | June 22, 2009 | 367 | 2463
Z-0975-2009 | 50161 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Surgical Membrane Oxygenator. |  |  | October 17, 2008 | 2008 | Medtronic Inc | Medtronic has identified an issue with the labeling of the I-4500 Silicone Oxygenators (P/N 61399402655). Lot numbers of this oxygenator manufactured between August 2007 and September 2008 have an incorrect "Recommended maximum blood flow range" on the label. The label on the device states a range of 5.0 - 8.5 L/min but should state a range of 5.0 - 6.5 L/min. These oxygenators are acceptable t | Error in labeling | A consignee letter " Urgent Medical Deive Notice" dated October 17, 2008 was sent to the impacted customers on October 17, 2008. The letter described the issue and product and confirmed the correct maximum internal flow of 6.5 L/min for the I-4500 Silicone Oxygenators. The letter states the product will not be removed from the field. Customers were instructed to complete the" Medtronic Product Notification Certificate (FCA 0903)" acknowledging receipt of the information and fax the completed certificate to 763-391-9603. Please direct questions about the recall to your Medtronic Sales Representative or call 763-391-9124. |  | 408 | 408 | Worldwide Distribution USA including states of CA, CO, FL, IL, IN, KY, LA, MD, MI, MN, MO, NV, NJ, NY, OH, OK PA, TN, TX, UT, VA, and WI, and countries of LATVIA, ITALY, ARMENIA, RUSSIAN FEDERATION, POLAND, CHILE and CHINA. | Terminated | April 19, 2011 | 914 | 408
Z-0980-2009 | 50223 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical Universa Soft Ureteral Stent and Positioner with Hydrophilic Coating, 5.0Fr/26cm, sterile, Cook Urological Inc., Spencer, IN; REF USH-526-R, GPN REF G49966. The product was placed in a shipping box which is preprinted with the word "firm" instead of "soft", but a product specific label stating "soft", as above, was also placed on the box. |  |  | November 10, 2008 | 2008 | Vance Products Incorporated | "Soft" stents were placed in "firm" stent outer boxes. The soft stent is only to be left in the patient for 6 months, but the firm stent can be left in the patient for 12 months, so this error could result in the stent being left in the patient for longer than indicated for the device. | Labeling mix-ups | Consignees were notified via recall letter, "Urgent: Recall Notice" dated November 10, 2008, and by follow-up phone call. Customers were instructed to isolate the recalled product, complete the enclosed form and fax to Cook Urological at 812-829-2353. After receipt of the completed form, Cook Urological will issue a Return Goods Authorization number for each customer to return the affected product to Cook Urological via FedEx or UPS. |  | 23 | 23 | Nationwide Distribution including states of Georgia, Oregon and Texas. | Terminated | November 25, 2009 | 380 | 23
Z-0863-2009 | 50522 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Model 7310PR-D VacuAide Compact Suction Unit. A portable, AC-powered or compressed air-powered suction pump. The device may include a microbial filter. |  |  | October 27, 2008 | 2008 | Sunrise Medical | Improper suction: Regulator knob does not turn properly which can affect the suction level. | Device Design | DeVilbiss notified dealers and users (through the dealers) to recall the units, by a telephone call October 27, 2008 followed by an Urgent Medical Device Recall mailing November 3, 2008. The letter advised customers what models are affected and how to contact the firm of dealer to have the affected unit replaced. Corrective and preventative action on current manufacturing has been implemented. |  | 78 | 78 units | Worldwide Distribution: USA, Mexico, Guatemala, Colombia, and Costa Rica. | Terminated | September 25, 2012 | 1429 | 78
Z-0915-2009 | 48598 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Tibial Array and Femoral Array. |  |  | May 02, 2008 | 2008 | Mako Surgical Corporation | The product in its current state may fail during use, which would cause the instrument to be in-accurate. | Component design/selection | A MAKO employee will visually verify product at each account for any product affected by this recall. In addition, a recall letter dated July 1, 2008 was mailed to users (surgeon and owner of system, e.g. hospital). Customers were instructed to complete the enclosed Acknowledgement form and return it to your Clinical Technical Specialist (CTS) or in the enclosed pre-paid addressed envelope to the Mako Surgical Corporation. Please direct questions to your CTS or call 954-927-2044, extension 604. |  | 44 | 44 | Nationwide Distribution USA including states of FL, NY, CA, TX, GA, TN, NC and PA. | Terminated | August 21, 2009 | 476 | 44
Z-0976-2009 | 50548 | [6 empty] | Turbo Elite 0.9mm Turbo Elite (TE) RX, Model #410-154, tip diameter 0.9mm/.038 in, 4 Fr, sterile EO. Manufactured by Spectranetics, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 04, 2008 | 2008 | Spectranetics Corporation | The catheter pouch label was incorrectly labeled with the wrong size. | Error in labeling | The mislabeled catheter was returned to Spectranetics. The second possible mislabeled catheter was visually inspected by a Spectranetics representative and confirmed the catheter was labeled correctly. |  | 2 | 2 units | Nationwide Distribution including states of TX and FL. | Terminated | February 13, 2009 | 224 | 2
Z-0982-2009 | 50700 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy DNP Anatomic Fracture Repair System, JIG DNP Anatomic, Depuy, Warsaw, IN; REF DNPAJIG. |  |  | December 23, 2008 | 2008 | Depuy Orthopaedics, Inc. | The DNP nail plate may not attach to the alignment guide (jig) instrument, making it difficult to insert the nail plate into position or to drill needed holes accurately. | Device Design | The firm's sales force was instructed via "Urgent Device Correction" letter issued December 29, 2008 to inspect all of these instruments in their possession, and those instruments at their hospital customers, and to leave a copy of the letter with each hospital customer who has the instruments on site.The letter describes the issue and corrective actions for customers. Please direct questions about this recall to Depuy Orthopaedics, Inc. at 888-783-7156. |  | 3250 | Approx. 3,250 of all products | Nationwide Distribution. | Terminated | October 18, 2010 | 664 | 3250
Z-0979-2009 | 50697 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gold Probe 7FR Working Length: 350 cm, Probe Tip O.D. 2.3 mm (7Fr), Rx Only. Electrical Connector: Standard. Single Use Only. Legal Manufacturer, Boston Scientific Corporation, One Boston Scientific Place Natick, MA 01760-1537 USA. Made in USA: 8600 N.W. 41st Street Miami, FL 33166-6202 USA. |  |  | December 24, 2008 | 2008 | Boston Scientific Corporation | Packaging tray for the Gold Probe Single-Use BiPolar Electrohemostasis Catheters may have a crack. Cracks in the tray or damage to the tray of the Gold Probe device may lead to a compromised sterile barrier. | Packaging process control | A notification letter (dated December 24, 2008) was issued to customers. The letter identified the issue. Consignees are instructed to: 1) Discontinue use and segregate affected product, 2) Complete a Reply Verification Tracking Form and fax to the Boston Scientific Corporation (BSC) at 508-683-5578 and 3) Package and ship recalled product as instructed in the letter. If affected product is found, customers are instructed to call BSC Customer Service at 1-866-868-4004 between 8:30 a.m. through 6:00 p.m. (Eastern Standard Time).Please direct questions about the recall to your local Sales Representative. |  | 10841 | 10,841 | Worldwide Distribution. | Terminated | May 11, 2010 | 503 | 109610
Z-0862-2009 | 50499 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stainless Steel Surgical Blades Part no 4-312B. The product is supplied sterile individually packaged in boxes containing 100 blades. |  |  | November 17, 2008 | 2008 | Miltex, Inc. | Blade may break: Over milling of the blade resulted in tip more narrow than specified. | Process control | An Urgent Medical Device Recall letter dated 11/17/08 with an accompanying fax back form was sent to each customer. Each distributor was instructed to review their inventory, complete the form, and return any affected product. The distributors were also instructed to conduct a sub recall notification. |  | 200 | 200 boxes | Nationwide Distribution and Canada. | Terminated | February 23, 2009 | 98 | 200
Z-0861-2009 | 50496 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Eclipse treatment planning system, Model Number H48, versions 6.5/7.x and Inspiration 8.0. |  |  | November 21, 2008 | 2008 | Varian Medical Systems Oncology Systems | Possible incorrect wedge orientation in treatment: Software labeling was issued without warning message relating to preservation of user selected values for Siemens wedge orientation, wedge limits and SAD values for Elekta machines, and collimator operating limits for Saturne Machines when the software version is upgraded. | Software design | An Urgent Device Correction product notification letter/field safetly letter dated November 21, 2008, was distributed to all affected users, with a description of the problem and user corrective action steps. The letter was also distributed to Varian Sales, Marketing and Service organizations, informing them of the issue. |  | 516 | 516 units | Worldwide Distribution: USA, Austria, Belgium, Canada, Czech Republic, Finland, France, Germany, Greece, Hungary, Italy, Ireland, Lithuania, Norway, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, Netherlands, Turkey and the United Kingdom. | Terminated | March 10, 2011 | 839 | 516
Z-0910-2009 | 50554 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base, Terumo Cardiovascular Systems, Ann Arbor, MI. Model 801764. (Not distributed within the United States). |  |  | November 07, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The power manager board may cause a short circuit, resulting in loss of system power, loss of central control monitor or electrical arcing. | Nonconforming Material/Component | An "Urgent Medical Device Correction" letter was sent to consignees on November 7, 2008. The letter described the reason for correction and provided instructions to customers. Terumo Cardiovascular Systems Corporation will replace the power manager boards on the affected devices. Customers were instructed to complete and fax the Customer Response Form to CVS Customer Service at 1-800-292-6551 (US only) or respond by telephone at 1-800-521-2818. Please direct questions to CVS Customers Service at 1-800-521-2818. |  | 515 | 515 | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 242 | 1080
Z-0940-2009 | 49610 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker InTouch 1.0 Critical Care Bed; Model 2130. |  |  | January 05, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The touch screen on the footboard may display a "tilt over range" error, resulting in a precautionary stop of the bed motions which result in the bed appearing to have locked out some of its bed positioning features. | Device Design | Consignees were notified of the problems via letter dated January 5, 2009, advised to check their beds and to take them out of service if the foot board touchscreen displayed a "tilt over range" error, and that a Stryker Medical field service representative would contact them to perform software upgrades on the affected beds. Please direct questions to Stryker Medical at 1-269-324-6644. |  | 1288 | 1,288 of all models | Worldwide Distribution including USA and countries of Canada, Japan and Switzerland. | Terminated | November 05, 2010 | 669 | 1288
Z-0919-2009 | 50064 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science Powerheart Automated External Defibrillator G3 9390A |  |  | October 13, 2008 | 2008 | Cardiac Science Corporation | Potential for AED to not deliver defibrillation therapy. | Software design | On October 13, 2008 Cardiac Science began sending via Fed-Ex letters and software upgrade kits. Mailing of letters and software upgrade kits is anticipated to be completed by November 21, 2008. The letter instructs consignees to install the software update, assure that the AED information test file generated by the update software is returned to the firm, and complete the postage paid response card and return it to the firm. If the affected devices are no longer in the possession of the consignee, the consignee is to contact the firm and forward the update kit and letter to the person who has the AED in their possession. A separate letter is being sent to distributors. The letter advised to install the software upgrades for any units in stock or to forward the software upgrade along with the end user letter enclosed and that the AED information text file generated by the update software must be returned to the firm. The firm will visit the locations of individual national account customers and install the update. OEM partners will be notified and provided the information and update kits to provide to their customers. |  | 2572 | 2,572 devices | Worldwide Distribution: USA, Puerto Rico, Virgin Islands, Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czechoslovakia, Denmark, Dominican Republic, France, Germany, Greece, Guam, Hong Kong, Hungary, Ireland, Israel, Italy, Jamaica, Japan, Kenya, Kuwait, Libya, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, UK, and Ukraine. | Terminated | December 08, 2011 | 1151 | 24183
Z-0944-2009 | 50541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One-Channel list number 16026-04 and Two-Channel list number 16027-04 Infusion System, a general purpose infuser, Distributed and Manufactured by Hospira Inc, Morgan Hill, CA. |  |  | December 03, 2008 | 2008 | Hospira Inc | Hypersensitivity to normal motor friction/drag may cause malfunction errors and alarms that are designed to stop pump operations. | Software design | Recall initiated 12/03/2008. Consignees were notified of the issue via an Urgent: Device Field Correction letter, dated 12/3/2008, and were requested to respond appropriately to any alarm or warnings displayed by the device. The firm is making arrangements to screen each Symbiq pump in customer facilities and implement the appropriate enhancements. |  | 11840 | combined 11,840 units | Worldwide Distribution --- including USA and countries of Canada and Australia. | Terminated | April 03, 2012 | 1217 | 11840
Z-0990-2009 | 50552 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Minimally Invasive Solutions Technology Offset Rasp Handle, Left, 30 degrees, Zimmer, Warsaw, IN; REF 00-7806-035-01. |  |  | December 19, 2008 | 2008 | Zimmer Inc. | The c-clip located near the connection end may fracture or fall off during use in surgical procedures. | Labeling design | U.S. consignees were notified by "Urgent: Device Recall" letter issued via e-mail on December 19, 2008. The letter described the problem and instructed customers to immediately quarantine and return the recalled product along with the completed Inventory Return Certification form to the Zimmer, Inc. Distribution Center. Consignees were also instructed to fax a copy of the completed Inventory Return Certification Form at 1-574-371-8603. International consignees were also notified, but Zimmer, Inc. will gather the instruments from those customers. Direct questions about the recall to Zimmer, Inc. at 1-800-613-6131. |  | 1461 | 1,461 of both products. | Worldwide Distribution --- including USA, Belgium, Czech Republic, Finland, Germany, Greece, Italy, Poland, Romania, Spain and United Kingdom. | Terminated | December 22, 2009 | 368 | 1461
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423
Z-0985-2009 | 50746 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Orthosoft Navitrack, OS Unicondylar Knee 1.0 - Universal; software application for orthopedic surgical stereotaxic instrument. |  |  | August 19, 2008 | 2008 | Zimmer Inc. | Incorrect parameters: If the user switches from one surgeon profile to another surgeon profile after completing anatomical landmark acquisition tasks, the application will continue to use the initial surgeon profile parameter values in the calculation of the predicted hip-knee-ankle alignment estimates, which may have different resection or component thickness parameters than are appropriate for | Software design | Zimmer notified consignees via letter dated 8/19/08, which advised them not to change the selected surgeon profile after the acquisition of landmarks has begun, and the issued will be fixed with the next software release. |  | 75 | 75 (34 U.S. and 41 international) | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Singapore, South Africa, Switzerland and United Kingdom. | Terminated | April 28, 2010 | 617 | 75
Z-1038-2009 | 50839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore. Product codes: 2M8153, 2M8163, 2M9163, DNM8153, DNM9163, BRM8153, BRM8153T, BRM8153RT, PNM8153, PNM8153T, PNM8153RT, CNM8153, GNM8153, HNM8153, TRM8153 and WNM8153. |  |  | January 23, 2009 | 2009 | Baxter Healthcare Corp. | Baxter has identified failure codes that lead to a delay or an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages for the Colleague Infusion Pumps. | Device Design | Baxter Healthcare Corporation sent an "Urgent Device Correction" letter dated January 23, 2009 to all Colleague Infusion Pump customers, to the attention of the Vice President of Nursing with copies to the Chief Operating Officer, Director of Biomedical Engineering and Director of Risk Management to advise them of failure codes that lead to an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages associated with the Colleague pumps. The letters listed the failure codes involved that lead to an interruption of therapy which could lead to serious injury and/or death; advised institutions to have contingency plans to verify that back-up pumps are available; and provided steps for addressing an interruption of therapy with any failure code. The letters also provided information about the hazards associated with improper cleaning of the pumps and instructions for addressing damaged battery messages. Customers were instructed to complete the attached Customer Reply Form acknowledging receipt of the letter and fax it to 888-569-1585. Direct questions to your local Baxter Representative or to Baxter Global Technical Services by calling 1-800-843-7867, Prompt 2, between 7AM and 5 PM Central Standard Time (CST). |  | 81775 | 81,775 pumps | Worldwide including USA, Puerto Rico, Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Canada, Cayman, Chile, China, Colombia, Curacao, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Honduras, India, Ireland, Israel, Italy, Jamaica, Korea, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Qatar, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | October 04, 2011 | 984 | 302872
Z-0955-2009 | 50796 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Meter Remote for OneTouch Ping Glucose Management System |  |  | October 27, 2008 | 2008 | Animas Corporation | incorrect dosing recommendation | Software design | The recalling firm issued Urgent: Medical Device Correction letters dated 10/27/08 to their customers/user to inform them of the problem. The letter stated that the firm was in the process of correcting the issue and that a new OneTouch Ping Meter-Remote would be provided as soon as one became available. In the interim, customers could still safely use their OneTouch Ping System provided that they use the ezCarb calculator on the meter-remote when it is paired and communicating with the pump. The letter listed guidelines to ensure that the device is properly communicating with the pump. The letter stated that customers should follow the directions in the Owner's Booklet to troubleshoot any communication issues. Questions regarding this letter were to be directed to Animas at 866-423-4087. A patient acknowledgment letter dated November 18, 2008 was also issued. Another Urgent: Medical Device Correction letter, dated November 2008, was issued along with the replacement device. This letter asked customers to download their current meter-remote data, set up their replacement meter-remote, and then return their current meter-remote. |  | 5093 | 5093 units | Nationwide Distribution --- including states of KY, UT, FL, OH, MA, and TX. | Terminated | December 15, 2009 | 414 | 5093
Z-0947-2009 | 50736 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | VARI-STIM III Nerve Locator/Stimulator, Sterile, Medtronic Xomed, Inc. |  |  | November 21, 2008 | 2008 | Medtronic Xomed, Inc. | Device may fail to operate or operate intermittently during use and such operation could result in a lack of nerve stimulation when nerve tissue is actually contacted by the stimulator. | Employee error | All consignees were contacted by telephone (initiated on November 10, 2008). After telephone contact, each consignee was sent a follow-up fax/letter, "Urgent Product Recall Notification". The letter described the issue and actions for customers to take in response to the recall. Customers were instructed to check their inventories and return affected product to Medtronic Xomed, Inc. Customers were also instructed to fax their completed "Customer Contact and Product Return Checklist" to Medtronic Xomed, Inc. at 1-904-296-2386. Direct questions to Medtronic ENT Customer Service at 1-800-874-5797 (select Option 1 for Customer Service) and ask for the Vari-Stim Field Action Contact Person. |  | 22 | 22 boxes (220 units) | Worldwide Distribution --- including USA and country of Canada. (22 boxes were distributed within U.S. Three boxes were distributed to Medtronic Canada but returned) | Terminated | October 01, 2009 | 314 | 22
Z-1013-2009 | 50738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 Posterior Stabilized Tibial Insert |  |  | December 18, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a lot for lot mix-up between part numbers 5532-G-713, lot code LBD528 (Triathlon X3 Posterior Stabilized Tibial Insert, size 7-13 mm) and part 5532-G-813, lot code LBD512 (Triathlon X3 Posterior Stabilized Tibial Insert Size 8-13 mm) | Labeling mix-ups | Urgent Product Recall letters were sent via Federal Express on December 18, 2008 with return receipt to Stryker branches, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letter stated the issue, the potential hazards, and the risk mitigation factors. The firm asked customers to examine their inventory and hospital locations to identify affected product; retrieve all affected product and return it to their branch or agency warehouse for reconciliation; reconcile all products on the attached Product Accountability Form; and return affected product. Questions should be directed to Karen Ariemma at 201-831-5718. |  | 126 | 9 of lot code LBD528; 14 of lot code LBD512 | Nationwide Distribution: FL and IL. | Terminated | September 11, 2012 | 1363 | 126
Z-0992-2009 | 50132 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Bone Wax. |  |  | October 28, 2008 | 2008 | Ethicon, Inc. | A hospital reported to their Ethicon sales representative that they had an "unmarked" box of bone was on their shipping dock that they could not receive due to lack of appropriate documentation. Upon investigation, this box was determined to be missing from a shipment between two steps of the manufacturing process. Further checks determined that two other boxes were missing. | Process design | Letters were sent via UPS next day air on October 28, 2008 to all customers who purchased finished product codes W31G and DW31B (on inner box labels) from January 1, 2008 through October 22, 2008. |  | 2200 | 2,200 foil pouches | Nationwide Distribution. | Terminated | April 29, 2009 | 183 | 2200
Z-0997-2009 | 51160 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley PED3.0 Cuffless Pediatric Tracheostomy Tube, REF 3.0 PED, Sterile. Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Assembled in Mexico. The product is provided as one each shipped in an inner carton to end customers. |  |  | January 14, 2009 | 2009 | Covidien Limited | Difficulty inserting an obturator and/or suction catheter into the tracheostomy tube. | Finished device change control | Consignees were sent an "Urgent Device Alert" letter dated 01/14/2009, by mail via Federal Express overnight delivery and instructed to evaluate any of the affected lot numbers of tracheostomy tubes prior to placement into a patient's stoma. If either the obturator or the desired sterile suction catheter does not fit properly into the tube, the tube must not be inserted into the patient. Further instructions are available at Nellcor Technical Services by calling 1-800-635-5267, option 3, then option 1. New product was expected by mid February, 2009, to address the issue. |  | 8931 | 8,931 units | Nationwide Distribution, including Puerto Rico. | Terminated | May 21, 2012 | 1223 | 8931
Z-1031-2009 | 50176 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Silhouette Axial Torque Handle, Zimmer Spine, Inc., Minneapolis, MN; REF 7010-0151-01. |  |  | November 07, 2008 | 2008 | Zimmer Inc. | Some of the torque wrenches apply less than the minimum torque to the screw locking cap when the torque indicators are aligned as described in the surgical technique. | Process control | The firm's sales force was notified by an Urgent: Device Recall letter dated 10/7/08, which instructed them to stop using the instrument and quarantine it immediately; physically count all affected product in their territory and record data on the Inventory Return Certification Form; fax a copy of the completed form to Zimmer Spine; and return the affected devices with the completed Inventory Return Certification Form. User accounts/implanting physicians were informed by an Urgent: Device Recall letter dated 12/10/08 and instructed to monitor their patients. |  | 76 | 76 | Worldwide Distribution --- USA including states of California, Minnesota, Ohio, Pennsylvania, and Tennessee, and countries of Denmark, Germany, Iran, Italy, Russia, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | December 22, 2009 | 410 | 76
Z-1033-2009 | 49372 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InteliFUSE, Drill Bit, 2.00 mm, J-Latch, QTY: 1, REF: DB.200.JL, non-sterile, 1515 Poydras st., Ste 1490, New Orleans, LA 70112, MDCI Arundel House, 1 Liverpool Gardens, Worthing, West Sussex BN11 1SL, United Kingdom |  |  | July 30, 2008 | 2008 | Intelifuse Inc | Manufacturing specification for drill bit was incorrect which, when used in conjunction with the drill guide, could cause the production of metal shavings. | Process control | The firm initiated their recall on July 31, 2008 via telephone and Urgent Recall Notice to their current distributor, Integra. The letter asked them to complete, sign, and return the enclosed Recall Acknowledgment and Return Form. The letter also stated that InteliFUSE would travel to their designated Distribution Center(s) to ensure testing and replacement of all defective units once the form has been received. The firm sent follow-up Urgent Recall Notice letters on August 14, 2008 directing the distributor to recall from the hospitals to whom they distributed the product and to two hospitals which were shipped the products directly. The letter asked that use of the affected products be immediately ceased and that the any stock on hand be quarantined. The Recall Acknowledgment and Return Form should be completed, signed, and returned and the products should be returned to InteliFUSE for immediate replacement. |  | 2149 | 2149 units | Nationwide Distribution - NY and NJ. | Terminated | June 22, 2011 | 1057 | 4282
Z-1034-2009 | 50899 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra Mosaik Putty, 15cc, Catalog # PTY10256 |  |  | January 16, 2009 | 2009 | Integra LifeSciences Corp | Potential for open seals in the inner foil pouch of certain products. The outer seal remains intact and therefore the sterility of the inner pouch is maintained. | Pending | Urgent: Product Recall Notification letters were sent on January 21, 1009 via Federal Express. The letter asked customers to immediately quarantine any stock on hand and contact their local Integra Sales Representative for further assistance and arrangement of product returns. The enclosed "Recall Acknowledgement and Return Form" indicating whether or not customers possess any of the affected product should also be completed and returned. Company is conducting recall to the hospital level. Questions should be directed to Judy O'Grady at 609-936-2317. |  | 88 | 88 units | Nationwide Distribution -- including states of FL, CA, GA, FL, CT, MT, IL, TX, WA, TN, NC, IN, AZ, MD, CO, MO, and TN. | Terminated | June 25, 2009 | 160 | 119
Z-0953-2009 | 50511 | No | GNW | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | LIGHT DIAGNOSTICS SimulFluor Flu A/Flu B DFA Kit |  |  | November 24, 2008 | 2008 | Millipore Corporation | Certain lots of Simulfluor Flu A / Flu B Reagent may cross-react with Respiratory Synctial Virus. The continued use of Simulfluor Flu A / Flu B Reagent Catalog #5250 may lead to a potential misdiagnosis of influenza B. | Pending | Millipore is performing a voluntary recall by notifying customers and distributors by telephone on November 24, 2008. In addition, an Urgent: Medical Device Recall notification letter with the same date was sent via certified mail.Customers were advised to examine their stock immediately and discontinue the use of the specified lots. Customers were instructed to complete the enclosed Effective Check Survey confirming disposal of defective lots and return this form to Millipore's Regulatory Department. Customer service is to be contacted at 1-800-437-7500 for replacement product. Millipore Corporation will issue credit or replace the recalled lots that customers currently have in inventory. |  | 404 | 404 units | Worldwide Distribution -- USA including states of AL, CA, CO, CT, FL, ID, IL, KS, MA, MI, MS, NM, NY, OH, OK, OR, PA, SD, TX, UT, VA, VT, WA, WV, and WI and countries of Canada, Mexico, Italy, Belgium,Germany, France, Spain, andUnited Kingdom. | Terminated | August 10, 2012 | 1355 | 662
Z-1006-2009 | 50146 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Reliability issues - potential to shock a non shockable rhythm or not shocking a shockable rhythm; | Component design/selection | Welch Allyn sent a Customer Recall Notice dated February 26, 2009, to all customers (customer list supplied by distributors). Welch Allyn will upgrade all units with software version 2.02 and lower to version 2.06.02 by Proc. AED10 Prod Service Min HW/SW Update Rev. E 820-0178-00. |  | 10735 | 10, 735 | Worldwide Distribution. | Terminated | March 19, 2012 | 1190 | 10735
Z-1018-2009 | 50711 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical Medfusion Model 3500 Syringe Infusion Pump, Software version V3.0.6. A small, lightweight and portable syringe infusion pump. |  |  | December 18, 2008 | 2008 | Smiths Medical MD, Inc. | Delay in therapy. Keypad issues with Medfusion Syringe Infusion Pump, model numbers 3500-306 and 3500-415 ("Pump") have occurred during implementation or shortly after first being placed into use. These issues can be intermittent, and may render the Pump inoperable or the Pump may not respond to certain keypad button presses. | Nonconforming Material/Component | Consignees were initially notified of the recall via a telephone conversation beginning December 18, 2008 and continued through January 2009. After the initial telephone contact, an "Urgent Medical Device Recall" letter dated December 22, 2008 was sent via electronic mail. The letter was addressed to "Affiliates, Distributors, Clinicians, and other Customers". The letter gave a Description of the Problem and Advise on Action to be Taken by the User. Direct questions to Smiths Medical Customer Service department at 1-800-426-2448 (option 1, then option 3).A Smiths Medical Customer Service Representative will have contacted customers. These pumps will be fitted with new keypads with gold connectors. A second Smiths Medical "Urgent Medical Device Recall Additional Affected Units"recall letter was sent after the recall was expanded. The second letter dated March 6, 2009 was addressed to Safety/Risk Mangers, Healthcare Providers, Biomedical Engineers, Distributors, Clinicians and other Customers. The letter listed the Details of Affected Devices, Description of the Problem, Advise on Action to be Taken by the User, Results of Investigation and Contact Details. |  | 957 | 957 | Worldwide, including USA, AMERICAN SAMOA, AUSTRALIA, CANADA, CHILE, GERMANY, UNITED KINGDOM, HONG KONG, IRELAND, ISRAEL, SOUTH KOREA, NETHERLANDS, SAUDI ARABIA, SWEDEN | Terminated | July 06, 2012 | 1296 | 3509
Z-0996-2009 | 51039 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Most Hinged Tibia Baseplate Spacer, size 1, 6 mm, sterile, Zimmer, Warsaw, IN; REF 5020-06-001. |  |  | January 29, 2009 | 2009 | Zimmer Inc. | The package contains 11 mm screws instead of the specified 6 mm screws. | Packaging process control | The consignee in Germany was notified by an Urgent: Device Recall letter dated 1/29/09, and delivered by email on 2/19/09. The letter stated the reason for the recall and clinical implications. The firm asked the consignee to stop using the provisional and quarantine it immediately; physically count all affected product and record data on the Inventory Return Certification Form included with letter; fax a copy of the completed form to Zimmer, Inc at 574-371-8603; request an IRA from Lindsey Arnett; and return recalled product along with completed Inventory Return Certification Form and IRA to Simmer Distribution Center. The letter states that credit will be issued upon receipt of the returned product. Also, if product has been further distributed, a copy of the recall letter should be provided to the customers. Questions or assistance should be directed to Zimmer, Inc at 1-800-613-6131. |  | 3 | 3 | International Distribution only -- Country of Germany. | Terminated | November 12, 2009 | 287 | 3
Z-1005-2009 | 50142 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 Automated External Defibrillator |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Potential for device to shut down prematurely under certain circumstances. | Device Design | On 4/13/07 the firm released Service Bulletin part no. 810-2614-00, SB - AED 10 S/W V020600 DOWNGRADE to V020500. That mandatory bulletin, for units with software version 2.06.00 and up indicates that all returning AED 10 units shall have their software version downgraded to 2.05.00. |  | 89 | 89 units | Worldwide distribution: USA, Belgium, Canada, France, Ireland, Israel, Mexico, and United Kingdom. | Terminated | February 15, 2012 | 1157 | 89
Z-1022-2009 | 50084 | No | KIJ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Rotary Handpiece, Single Trigger, For use with Stryker System 6. Stryker Instruments, Kalamazoo, MI; REF 6203. |  |  | October 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | The handpieces may continue to run even when the trigger is not depressed, which could result in the unintended cutting of bone or soft tissue damage. | Component design/selection | Consignees were notified of the problem via letter dated October 10, 2008, provided interim instructions to reduce the risk, and informed that Stryker instruments will replace the trigger on all handpieces. Consignees were also instructed on the following: 1) Review the letter and distribute to all affected departments, 2) Complete and sign the enclosed Business Reply Form and fax to 336-499-8603 and 3) Call Stryker Instruments Service Department at 888-308-1983 to schedule service for the handpieces. |  | 25312 | 25,312 of all units | Worldwide Distribution including USA and countries of Australia, Brazil, Canada, Chile, Colombia, Costa Rica, England, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Pakistan, Panama, Peru, Philippines, Poland, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan and Turkey. | Terminated | October 27, 2010 | 747 | 25312
Z-1007-2009 | 50532 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Potential for device to deliver less than the programmed energy. | Process control | Welch Allyn notified distributors by letter on 12/15/2008 requesting that they notify their customers. Welch Allyn sent an Urgent: Medical Device Recall letter to distributors and customers on February 26, 2009. |  | 14012 | 14,012 devices | Worldwide distribution. | Terminated | March 19, 2012 | 1190 | 14012
Z-1060-2009 | 50856 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PROSTIVA RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. |  |  | June 05, 2008 | 2008 | Medtronic Neuromodulation | Reported incidents of needle retraction failure when using the PROSTIVA RF Model 8929 Hand Piece. The Medtronic Prostiva RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. | Device Design | A Medtronic "Urgent Medical Device Correction" letter addressed to "Healthcare Provider" was sent August 6, 2008. The letter described the issue, product involved and the Prostiva hand piece Disassembly Procedure that should be followed in the event the user is unable to retract the needles during a procedure. The letter also requested the return of the "Medtronic Reply Form/receipt Confirmation" form. |  | 40168 | 40,168 | Worldwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, RI, SC, SD, IN, TX, UT, VA, VT, WA, WI, WY, and WV and countries of Australia, New Zealand, China, Austria, Denmark, Finland, France, Germany, Italy, Kuwait, Lebanon, Norway, Qatar, Russian Federation, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 17, 2012 | 1442 | 40168
Z-1059-2009 | 51176 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AcuSnare Polypectomy Soft Snare, Order number SASM-1, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105, packed 5 boxes, 10 units per box. |  |  | January 21, 2009 | 2009 | Cook Endoscopy | The snare head is hexagonal but the handle and packaging indicate an oval reorder number. | Process control | Consignees were notified via USPC, Certified Mail and email on/about 1/21/2009 with an Urgent - Product Recall Notice. The distributor (Cook Australia) was instructed to return quarantined product to Cook Endoscopy. All medical users were instructed to review the attached list and quarantine any affected product remaining at their facility. They were to complete the form at the bottom of the page and send/fax it. Users are to return any affected product to Cook Endoscopy. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed. |  | 10 | 10 boxes (100 snares) | Worldwide Distribution -- USA including states of MI, NY, and WI and country of Australia. | Terminated | April 06, 2010 | 440 | 10
Z-1010-2009 | 50861 | No | LQR | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Stonetome Stone Removal Device. Triple Lumen Sphincterotome, 5.5 Fr, 20 mm, 11.5 mm, sterile, Boston Scientific Corp., Natick, MA; REF 3511. |  |  | January 08, 2009 | 2009 | Boston Scientific Corp | The ultraviolet light barrier on the pouch may be discolored, which will allow ultraviolet light into the package and may cause the latex retrieval balloon to become brittle and to leak or burst when inflated. | Packaging | Consignees were notified via letter dated January 8, 2009 to discontinue use of the affected product and to return it to Boston Scientific. Customers were instructed to complete and fax the Reply Verification Tracking Form to Boston Scientific at 508-683-5578. Direct questions about this recall to your local Sales Representative. |  | 294 | 294 of all products. | Nationwide. | Terminated | October 22, 2009 | 287 | 294
Z-1091-2009 | 50131 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose Oxidase Reagent as follows: Pointe Scientific Glucose Oxidase Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog # G7519-1L, G7519-500 and G7519-SAM. 2MM Diagnostics Glucose Oxidase Reagent Set, 2MM Corporation, Miami, FL; Reorder # RT-160-500. Pointe Scientific Glucose Oxidase Reagent, Pointe Scientific, Inc., Canton, MI; Catalog, MI; Catalog # 2-G7519-L, 7-G7519-100, 7-G7519-200, 7-G7519-1000, 8-G7519-100, and 12-RT-160-500. |  |  | November 03, 2008 | 2008 | Pointe Scientific, Inc. | The product is unable to maintain the specification for linearity through the shelf life of the product. | Process control | Consignees were notified via an Urgent: Medical Device Recall letter dated 11/3/08 that all test results obtained with these lots are questionable and that attending physicians should be informed to determine whether additional testing is required. Also, consignees should examine their inventories for any of the lot numbers listed and dispose of them. If the product has been further distributed, identify customers and notify them of the product recall by including a copy of the recall letter. The Recall Return Response Form should be completed and returned and upon receipt Pointe Scientific will replace any disposed of reagent. Questions should be directed to the Technical Support Department at 1-800-757-5313. |  | 1500 | 1500 liters | Worldwide Distribution -- USA, Hong Kong, Indonesia, Israel, Jordan, Mexico, Nigeria, Poland, Russia, Serbia, Sri Lanka, Trinidad, and Venezuela. | Terminated | December 03, 2009 | 395 | 1500
Z-1014-2009 | 50185 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Medtronic TSRH 3DX Implant Set; SPS002058. Sets are unlabeled, but contain individually labeled 3Dx Splined Connectors in sizes small, medium and large. |  |  | October 31, 2008 | 2008 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | The thickness of the connector may be under specification, which could cause the implant to break during surgery or after implantation. | Process control | Hospital risk managers and implanting physicians were notified separately by letters dated 1/9/09. The letters advised of the recall, reported that a firm representative would visit the facility to pick up the recalled products. Physicians were instructed to closely monitor their patients during postoperative visits, to notify any other physician who may be monitoring the patient in lieu of the implanting physician, and to advise their patients to contact them immediately if they have any sudden and significant symptoms (increased discomfort, sudden change in their clinical symptoms, or new symptoms such as leg pain, back pain or irritation of local tissue). Customers were also instructed to complete and return the attached questionnaire to Medtronic SOFAMOR DANEK via mail (enclosed self-addressed stamped envelope) or fax at 901-399-2047. Direct questions to your Medtronic SOFAMOR DANEK Sales Representative or the firm's Global Quality Department at 800-876-3133, extension 6333. |  | 43 | 43 sets | Nationwide Distribution including states of Arizona, Connecticut, Georgia, Indiana, Minnesota, New Jersey, New York, South Carolina, Texas and Washington. | Terminated | November 18, 2009 | 383 | 4153
Z-1092-2009 | 51159 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 3100 / 3100 IQ, Cardiovascular Imaging System, GE Healthcare, Waukesha, WI 53188 (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | March 06, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare reported an issue with the Exposure Switch Redundancy system with Innova Systems 3100 &amp; 4100, in which this error stays on, even after the pedal is released, inhibiting further x-rays, and a system reset is required to recover. | Software design | On site visits were performed from 3/6/08 - 12/19/08 to update software on all 5 units. |  | 4 | 4 | Worldwide Distribution -- USA including states of MD, TX, and VA and country of China. | Terminated | September 04, 2014 | 2373 | 5
Z-1074-2009 | 51205 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 (Intermittent Catheter), Sterile, Uncoated 6 French Pediatric Length REF # 52306 |  |  | January 14, 2009 | 2009 | Rochester Medical Corp | Rochester Medical observed some Magic3 and HydroSil(R) catheter packages in which the catheter was caught in the package seal during the packaging process. | Package design/selection | Consignees were contacted by telephone or in-person visit. The communication provided the reason for the recall and a list of affected products and their lot numbers, and instruction for the removal. Distributors were given instructions to notify their customers in turn. For the customers that were not reached via telephone call or in-person, a letter, "Voluntary recall of magic product samples," was used for the recipients of product samples and sent via e-mail. The letter was dated January 14, 2009 and asked customers to destroy all of their remaining sample inventory of the affected lot numbers. |  | 12 | 12 | Worldwide Distribution -- USA including states of CA, FL, GA, ID, IL, LA, MD, MA, MI, MN, NJ, NM, NY, OH, PA, and WA and countries of United Kingdom. | Terminated | April 20, 2011 | 826 | 1960
Z-1070-2009 | 50471 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch SureStep Test Strips, Part numbers 020-052-01 and 010-359-04, for use with SureStep brand Blood Glucose Meters, manufactured by LifeScan Inc., Milpitas, CA |  |  | November 25, 2008 | 2008 | Lifescan Inc | Deformity in the test strip may result in insufficient blood transfer to reaction area, resulting in inaccurate test results. | Process design | Urgent: Medical Device Recall letters, dated 11/24/2008 were issued to all consignees on November 25, 2008, requesting return of affected lots to LifeScan and instructing how to obtain replacement strips free of charge. A request for sub-recall communications were made as well, and the firm posted notification on its website. Customer service and Sales account reps were provided information to address questions regarding the recall. |  | 23288 | 23,288 vials | Worldwide Distribution -- USA and the Caribbean. | Terminated | December 06, 2010 | 741 | 23288
Z-1088-2009 | 50689 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Panel, Model Number: 95000. |  |  | November 06, 2008 | 2008 | Biosite Inc | The recall was initiated after the firm's investigations confirmed that, while lot #186625 passed all internal specifications and quality requirements, the incidence of false positive PCP results has exceeded their post market surveillance thresholds for customer complaints in the field. In addition, the firm has been able to reproduce false positive results reported by customers and as a result | Other | Consignee notification was initiated on 11/6/2008 with Urgent Medical Device Recall notifications, dated 11/5/2008 and mailed to domestic customers sent via US Postal Service Priority Mail and to foreign customers via FedEx. The notice informed customers of the reason for the recall and requested the customers immediately discontinue all use of lot #186625 and discard all affected product in accordance with your local regulations. The customers were also instructed to complete and fax an attached Customer/Distributor Verification form within 10 days. |  | 2634 | 2634 kits | Worldwide Distribution -- USA and country of Canada. | Terminated | October 13, 2011 | 1071 | 2634
Z-0642-2009 | 50476 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medtronic COMPLETE SE Self Expanding Biliary Stent System, Model number SC760L, includes a pre-loaded self-expanding Nickel-Titanium (Nitinol) stent with 8 tantalum radiopaque markers, 4 on each end, and an over the wire retractable sheath delivery system. |  |  | November 24, 2008 | 2008 | Medtronic Cardiovascular | Mislabeled: Product may contain wrong instructions for use-- may contain International IFU instead of the correct US (for Biliary indication use only) IFU. | Process control | Medtronic sent a notification entitled "URGENT - VOLUNTARY MEDICAL DEVICE RECALL" to consignees on 11/21/2008. All products are to be returned to the firm's distribution center. For further information, please contact Medtronic Cardiovascular by telephone at 707-566-0111. |  | 23 | 23 units | Nationally in the states of NY, TX, CA, IA, AZ, PA, IN, and KS. | Terminated | March 27, 2009 | 123 | 23
Z-1090-2009 | 49548 | No | MKQ |  |  | Pathology | PMA | BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R. |  |  | September 26, 2008 | 2008 | Bd Diagnostic Systems Tripath | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. | Software change control | BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008. |  | 662 | 662 units | Worldwide Distribution | Terminated | January 15, 2013 | 1572 | 662
Z-1008-2009 | 50224 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Dryspot (R) Staphytect Plus, Catalog #DR0100M, 120 tests/kit. |  |  | October 07, 2008 | 2008 | Remel, Inc | Control and test wells are reversed on some of the reaction cards. | Pending | The recalling firm issued Important Medical Device Customer Advisory Notice letters dated 10/7/08 via regular mail to their customers explaining the reason for recall. The letter requested that they review their inventory for the product lots indicated in the letter; discard all of the lot remaining in their inventory; complete and return the Product Inventory Checklist. |  | 32 | 32 kits dist. by recalling firm | Nationwide Distribution -- including MN, MI, IL, and NC. | Terminated | July 30, 2009 | 296 | 32
Z-1009-2009 | 50170 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS c 111 Analyzer with ISE module option, GMMI No. 04777433001; Roche Diagnostics, Indianapolis, IN. |  |  | November 17, 2008 | 2008 | Roche Diagnostics Corp. | Due to a manufacturing error at the supplier, not all cobas c 111 pinch valve caps are properly glued to the correct position on the pinch valve shaft. Without the glue, the cap can become loose or misaligned. The effect of these faulty valves is that they may not open or close properly which can disturb the fluid transportation within the ISE module resulting in problems with ISE calibration, con | Nonconforming Material/Component | Consignees were notified via an Urgent Medical Device Correction letter dated 11/17/08 and instructed Roche will contact them to make an appointment to inspect their analyzers. A work around was provided, for interim use, in the letter. Customers were instructed to complete and fax back the "ISE Controls Request - Purchased from a Distributor." A no charge replacement product will be shipped to customers. |  | 66 | 66 | Nationwide Distribution | Terminated | December 08, 2009 | 386 | 66
Z-1041-2009 | 50174 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II TURBO Hemoglobin Testing System with CDM v 4.0, model number 270-2600, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | October 17, 2008 | 2008 | Bio-Rad Laboratories Inc | Sample results may be assigned to the wrong barcode number and tube position numbers, leading to discrepant sample results. | Pending | An Urgent Medical Device Correction letter was sent out on October 17, 2008 via fax and phone call. The letter recommended that customers review the Summary Report for duplicate tube position numbers or barcode IDs within the same rack. If there are questions or a need for assistance, customers are to contact their regional Bio-Rad office. Follow-ups will be made once root cause is identified. |  | 92 | 92 units | Worldwide Distribution | Terminated | November 12, 2010 | 756 | 92
Z-1040-2009 | 50163 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035. |  |  | November 06, 2008 | 2008 | Leica Microsystems, Inc. | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. | Software change control | Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made. |  | 5 | 5 units | Nationwide Distribution -- including states of CA, IL, KS, and NY. | Terminated | March 16, 2011 | 860 | 5
Z-1160-2009 | 50867 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Tri-State Centurion PICC Insertion Tray, sterile, 15 per case, Tri-State Hospital Supply Corp., Howell, MI.; Reorder CVI2415A. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Lack of assurance of sterility, as the package seals may be inadequate. | Packaging process control | Tri-State Hospital Supply Corporation notified all consignees via Urgent Product Recall Notice letter dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand, and to hold the product for retrieval. |  | 60 | 60 | Nationwide. | Terminated | December 31, 2009 | 351 | 1186
Z-1133-2009 | 51287 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782B. Sizes: 5 CM in length, REF (240) 27782B. Each catheter is provided in a sealed Tyvek plastic tray placed within a sealed Tyvek/Mylar pouch. The sealed pouch is then placed within a chipboard carton along with product IFU and patient wallet cards. The product is ETO sterilized. |  |  | February 12, 2009 | 2009 | Medtronic Neurosurgery | The ventricular catheters may become detached from the Snap Base assembly after implantation and this condition may increase the need for emergency revision surgery. | Device Design | An Urgent - Medical Device Recall notification was sent on February 13,2009 to the attention of risk managers, advising them to cease use of the Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782 ( all versions and all lot numbers), the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27802 ( all versions and all lot numbers), and the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27708 ( all version and all lot numbers). Each customer is to cease the use of the listed product and any unused unit needs to be returned to Medtronic. Further, each account is required to account for product used/implanted and to advise surgeons of the recall information. Each account will recieved an individualized Customer Product Reconciliation Form provided by their Medtronic Sales Representative for determining the status of any product shipped to them. |  | 3048 | 3048 of all units | Worldwide Distribution -- USA, Ireland, and Australia. | Terminated | November 08, 2010 | 634 | 3048
Z-1052-2009 | 50896 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Impaction Bur Guard REF 2296-301, 2296-300, 0260-901-300, &amp; 5400-300; Stryker Instruments, Kalamazoo, MI. |  |  | January 27, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Bur guards were distributed without an expiration date, and if they are used after they have expired it could result in the patient or operating room staff experiencing a burn or laceration. | Process control | Customers were notified via a Medical Device Recall Notification letter dated 1/27/09 and instructed to immediately check all stock areas or operating room storage and inspect their bur guards for an expiration date. They were also instructed to quarantine any affected product found, to indicate on the Business Reply Form the quantity of affected product on hand, to destroy any affected product found without an expiration date by breaking off at least 3 of the tabs on the bur guard, then dispose of the affected product in normal trash, and complete and return the business reply form. Customers are to fax the Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2760. Upon receipt of the Business Reply Form, a replacement bur guard(s) will be sent for any affected product destroyed. |  | 70 | 70 | Nationwide Distribution | Terminated | December 17, 2009 | 324 | 70
Z-1042-2009 | 50704 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Diasorin LIAISON (R) Chemiluminescence Analyzer Model # 15970. |  |  | December 05, 2008 | 2008 | Diasorin Inc. | DiaSorin determined that there was a potential for a short to occur in a a residual starter pump electrical cable in the LIAISON (R) Analyzer. | Device Design | DiaSorin Inc authorized Field Service Engineer has visited the sites and performed the necessary field correction. Upon completion of their work, the field service engineer generated a Field Service Report and a copy was left with the customer. |  | 9 | 9 | Worldwide Distribution -- USA including states of NC, NJ, PA, CA, WA, MO, UT, and OR and country of Canada. | Terminated | May 22, 2012 | 1264 | 9
Z-1089-2009 | 49438 | No | LJM | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Sm ELISA Test System |  |  | August 19, 2008 | 2008 | Zeus Scientific Inc | The Calibrator Value is printed incorrectly on the Sm ELISA Test System, lot # 08041819. The Calibrator Value (CV) printed on the label is 102; the correct CV should read 266. | Employee error | Urgent Recall Notice letters were sent out by Federal Express night mail on August 19, 2008 to all direct account. Wampole (Inverness) will notify their customers in a separate letter. Recall letters explained the reason for the recall and that the product does meet the company's criteria when it is used qualitatively to interpret index values or OD rations as outlined in the kit package insert. Customers are advised in the letter that the Calibrator Value of 102 is incorrect and that it should be 266. Instructions to correctly express the specimen results using specimen and calibrator OD values semi-qualitatively as Unit Values and the Calibrator Value (CV) or 266 are included. Letters also state that results for all patients using the incorrect CV of 102 should be recalculated using the correct CV of 266. All letters included a fax back form indicating that they have received the recall notification, have notified customers to correct the CV value for the Semi-Quantitative Conversion and have notified customers that have received the kit to correct the Semi-Quantitative Conversion of Optical Density to AAU/mL for the affected lot number of the Sm ELISA Test System. |  | 265380027 | 33 of lot 08041819, 27 of lot 08041811. | Nationwide Distribution | Terminated | April 30, 2009 | 254 | 265380027
Z-1073-2009 | 50520 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dual Taper Wedge Pressed Fit Femoral Component. |  |  | October 31, 2008 | 2008 | Stelkast Co | A possible breach of the sterile barrier packaging (Tyvek pouch). | Packaging | The recalling firm telephoned the distributors to inform them of the problem and the need to return the product. |  | 187 | 187 units | Nationwide Distribution -- states of NC, TX, and WV. | Terminated | April 14, 2009 | 165 | 187
Z-1020-2009 | 51177 | No | NQO |  |  | Orthopedic | PMA | Spacers of varying lengths used with the InSTANT Stranding System, Part numbers: TG3138, TG3139, TG3140, TG3141 and TG3142. Theragenics Corporation, 5203 Bristol Industrial Way, Buford GA 30518. |  |  | December 08, 2008 | 2008 | Theragenics Corporation | Some pouches were found to have weak seals potentially compromising sterility. | Packaging process control | Consignees were notified first by telephone on December 8, 2008 and sent a follow up letter dated January 13, 2009. They were instructed to return any of the affected lots remaining in their inventory. |  | 62 | 62 units | Nationwide Distribution -- states of VA, WI, and CO. | Terminated | April 06, 2009 | 119 | 62
Z-1043-2009 | 50821 | No | LRF | Cryptococcus neoformans serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BEC Growth Chek Microbial Suspension - Candida albicans CA1 (100 dose), Product #10231-100 |  |  | January 13, 2009 | 2009 | Lexamed | Candida albicans microbial suspension, certified to deliver less than 100 organisms per dose, was found to have a population count which exceeded 100 organisms per dose, following distribution. | Process control | The recalling firm contacted the customer and notified them of this recall via telephone on 1/13/2009, followed by an Urgent Product Recall Notice which was sent to the customer via Certified mail on 1/14/2009. The letter stated the problem with the affected lot and asked that the customer examine their inventory, remove the affected lot, and return the used and unused vial to LexaMed. The customer is to complete and return the enclosed acknowledgement card by mail or fax. Questions should be directed towards Anne Schuler, QA Manager at 419-693-5307 or by e-mail aschuler@lexamed.net. |  | 1 | 1 vial | USA: Utah | Terminated | December 09, 2009 | 330 | 1
Z-1204-2009 | 50482 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus Inc. PLAC Test Reagent Kit, Turbidimetric Immunoassay for the Quantitative Determination of Lp-PLA2 in Human Plasma or Serum, liquid ready-to-use two-reagent kit assembled in a kit box, manufactured by diaDexus Inc, South San Francisco, CA 94080, Catalog number 90110 |  |  | November 04, 2008 | 2008 | DiaDexus, Inc | Storage equipment failure at distributor site due to Hurricane Ike caused product to experience a temperature variation and go out of range. The integrity of the product could not be assured. | Other | The firm and its distributor contacted the consignees via phone on 11/04/08. The firm informed them of the issue and inquired about the products' status. An Urgent: Device Recall letter was issued on 11/04/08 via FedEx. The letter requested that use of the kit be immediately discontinued and returned promptly. A phone call was made to follow up requesting return of the product. |  | 27 | 27 units | Nationwide Distribution | Terminated | December 21, 2009 | 412 | 27
Z-1122-2009 | 51236 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ILED Surgical Light, Model numbers: ILED3, ILED5 and ILED5K |  |  | February 27, 2009 | 2009 | Trumpf Medical Systems, Inc. | Cracks may form around the screw connections of the cover plates causing brittleness, which may lead to fragments breaking off and falling into the operating area. | Other | Consignees were notified via an Urgent Safety Information letter starting 02/27/2009. The Recall Notice asks customers to visually inspect the ILED covers for cracks and in particular to focus on the corners. If cracking is visible they are asked to remedy by taping with a clear tape over the cracks and report it to TRUMPF Medical Systems, Inc. immediately. TRUMPF will be replacing all covers with a newly designed and tested cover. Original priority in the United States will be set to start with the oldest units affected. If TRUMPF receives notification that cracking does exist then that facility will rise to the highest priority. Additionally, TRUMPF will issue a new approved cleaning and disinfection instruction, since investigations have shown that the cracking is significantly influenced by the use of certain disinfectants containing alcohol. The operator manual previously stated to avoid using alcohol-based cleaners on some parts of the ILED, but after reviewing this it was determined that it was not clearly stated to avoid use on the covers. |  | 4480 | 4480 units (540 in U.S.) | Worldwide Distribution. | Terminated | October 22, 2009 | 237 | 4480
Z-1206-2009 | 51053 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll AED Plus Defibrillator |  |  | February 12, 2009 | 2009 | ZOLL Medical Corporation, World Wide Headquarters | Device fails to discharge the defibrillation energy. | Process design | On 3/31/09, the firm revised their recall strategy and required all users to a perform a "Mandatory Upgrade" software The software upgrade allows all potentially affected devices to monitor battery charging performance through periodic self-testing. If defective batteries are detected at any time prior to the recommended maximum of five years, users are prompted by the device to install fresh batteries. Frim issued Press Release on April 2, 2009.In the first notification Zoll notified customers by letter on 2/12/09 via certified mail. Customers instructed to replace device batteries every three years. In lieu of replacing batteries every three years, customers will also have the choice to update theirdevices with software that will monitor batteries for the identified defect. ZOLL will be send email notification to all consignees who have provided an email address as part of their contact information. ZOLL will also be publishing information of the notification on industry magazines. http://www.zollaedplusbatteryhelp.com/ |  | 183535 | 183,535 units | Worldwide distribution | Open, Classified |  |  | 183535
Z-1053-2009 | 50914 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Humeral Assembly, extra small, 4 inch length, sterile, Zimmer Inc., Warsaw, IN; REF 32-8105-027-04. |  |  | January 30, 2009 | 2009 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone. | Labeling design | Consignees were notified by "Urgent Device Correction" letter dated January 30, 2009 and provided revised surgical techniques for these extra-small implants. Direct questions about this recall to Zimmer, Inc. at 574-267-6131. |  | 14613 | 14,613 of all products | Worldwide Distribution, including USA, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Jordan, Latin America, Netherlands, Russia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey and Venezuela. | Terminated | December 22, 2009 | 326 | 14613
Z-1231-2009 | 51223 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Electric Stretcher with the Quick Drop CPR option, Model 1550. |  |  | February 27, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The head section backrest (fowler) may not raise up, or, if up, may drop suddenly due to wearing out of the actuator threads on the fowler. | Device Design | Consignees were notified by an Urgent - Medical Device Correction letter dated 2/27/09, advised of the problem and instructed to locate stretchers and verify that there is no grinding noise when activating the fowler electrically. Consignees were also instructed to remove any stretchers with a fowler that was non-operational, or which makes a grinding sound, from service until the firm makes repairs and contact the firm. Stretchers will be corrected within the next four months by a Stryker field service representative. The enclosed post card should be returned to confirm receipt of the notification. The notice should be forwarded to any users that have borrowed or purchased the stretchers and the firm should be notified of their new location. Also, if any of the affected stretchers have been disposed of or are no longer in use, the firm needs to be provided with their serial numbers. Questions or concerns should be directed to Andrea Kuthe at 269-324-6581. |  | 3607 | 3607 | Worldwide Distribution | Terminated | April 25, 2011 | 787 | 3645
Z-1044-2009 | 50518 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | ANA Fluorescent Test System and ANA Colorzyme Test System, Sub assembly catalog number SA 2013, Kit Catalog Numbers, 2000, 2002, 2003, 2004, 2005, 2100, 2200, 4000, 4200, 8024.1, and 8060, manufactured by Immuno Concepts N.A. Ltd, Sacramento, CA. |  |  | December 04, 2008 | 2008 | Immuno Concepts Inc | Samples can occasionally run out of the wells and cross contaminate adjacent wells. | Nonconforming Material/Component | The firm notified customers via a "Notice of Market Withdrawal" letter, e-mail and telephone. The letter described the problem and instructed customers to removed affected product from their inventory, complete the attached form and return it via fax to Immuno Concepts, Inc. at 916-363-2843. The firm will issue replacements as requested. Direct questions to Immuno Concepts, Inc. at 916-363-2649. |  | 378 | 378 kits | Worldwide Distribution, USA (NC, WI, MS, MD, IL, GA, CA, MI and Puerto Rico) and countries of Sweden, Mexico, Spain and Denmark. | Terminated | November 10, 2009 | 341 | 378
Z-1039-2009 | 51065 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, S/5 iCentral and iCentral Client, |  |  | January 26, 2009 | 2009 | GE Healthcare | GE Healthcare is providing information about two issues associated with iCentral L-NETO5 and L-NETC05 v4.3, 4.4, 4.5 and 5.0 when used in combination with Dash 2000, 3000, 4000 and 5000; Solar 7000, 8000, 8000M and 8000i; Eagle 3000 and 4000 and ApexPro. The iCentral does not provide audio or visual alarming for the following: A FIB, VT&gt; and PVC(1) alarms when used with ApexPro and VT&gt;2 and PVC | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated January 26, 2009 addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the safety issue, affected product details, safety instructions, product correction and contact information. Direct questions to GE Healthcare at 1-800-558-7044, Option 2 and ask for Technical Support or contact your local GE Healthcare Distributor. |  | 2146 | 2,146 | Worldwide Distribution, USA (and PR) including states of AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY,LA, ME, MI, MN, MS, MO, NH, NJ, NY, NC, OH, PA, SC, TX, VT, VA, WA, and WI and countries of AUSTRALIA, AUSTRIA, BARBADOS, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GAYANA, HONG KONG, HUNGARY, ICELAND, INDIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, PAKISTAN, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SLOVEN IA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, SYRIA ARAB REPUBLIC, THAILAND, TUNAS IA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA and YEMEN. | Terminated | August 13, 2013 | 1660 | 2146
Z-1045-2009 | 50519 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | RELISA ENA Multiparameter Antibody Screening test, enzyme immunoassay test system, manufactured by Immuno Concepts N.A. Ltd., Sacramento, CA. |  |  | November 25, 2008 | 2008 | Immuno Concepts Inc | In certain patient samples, it is difficult to properly discriminate between anti-Sm/RMP positive samples and samples that are only anti-Sm positive. | Process control | The firm has issued telephone calls, e-mails and notification letters to its consignees. |  | 25 | 25 kits | Worldwide Distribution including USA, Australia, England, Sweden and Austria. | Terminated | November 17, 2009 | 357 | 25
Z-0999-2009 | 50875 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Viasys Healthcare, Sonora Transcranial Doppler (TCD) System. Model Sonara. Part Number of Units: 982A0470 Sonara Bilateral. |  |  | January 06, 2009 | 2009 | Cardinal Health, Inc | Internal testing has identified several several scenarios where unexpected System lockups, shutdowns or system errors would render the Sonara unavailable for use until the Sonara can be restarted. These issues can be corrected by installing the Version 6.0 b045 software updating the Audio Driver and updating to XP Service Pack 3. Additionally, improper wiring could cause the cooling fan to not op | Device Design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated January 12, 2009 beginning the week of February 2, 2009 and continuing into week of February 9, 2009. The letter was addressed to customers describing the problems and systems affected and solutions. Cardinal Health requested customers to complete the actions listed in the letter to install the Field Correction and return the completed "Verification Form" via mail, e-mail or fax to Cardinal Health-NeuroCare. Further information can be obtained by calling 1-800-356-0007 ext 2191 or fax to 1-877-244-3805. |  | 64 | 64 systems | Worldwide Distribution -- USA including states of FL, DC, LA, NJ, NY, PA, TN, TX, WI, CA, MO, and LA and countries including THAILAND, SINGAPORE, SERBIA-MONTENEGRO, PORTUGAL, PHILIPPINES, POLAND, NORWAY, NETHERLANDS, KOREA, ITALY, ISRAEL, HUNGARY, HONG KONG, GREAT BRITAIN, SPAIN, GERMANY, CZECH REPUBLIC, SWITZERLAND, CANADA, BULGARIA, and BOSNIA-HERZEGOVINA. | Terminated | July 03, 2012 | 1274 | 100
Z-1061-2009 | 51136 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray Adaptor, 5.0 MM, sterile, Biomet Orthopedics, Inc., Warsaw, IN, REF 141491. |  |  | February 13, 2009 | 2009 | Biomet, Inc. | A 2.5mm locking insert was packaged with the 5.0mm Offset Tibial Tray Adaptor instead of the 5.0mm locking insert. | Packaging process control | Consignees were notified by "Field Safety Corrective Action" letter dated February 13, 2009 to conduct a recall to the hospital level and to locate and return the recalled product to Biomet Orthopedics. Customers were also instructed to complete the enclosed "Fax Back Response Form" and return it via fax to Biomet Orthopedics at 1-574-372-1683. Direct questions about this recall to 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8 a.m. to 5 p.m. |  | 17 | 17 | Worldwide Distribution including Germany, Italy, Netherlands, Spain and the United Kingdom. | Terminated | August 11, 2009 | 179 | 17
Z-1171-2009 | 51173 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Internal Sterilizable Paddles (electrodes) and Handles without discharge control for use with LIFEPAK 9 and LIFEPAK 9P defibrillators/monitors. |  |  | February 16, 2009 | 2009 | Physio Control, Inc. | Sterilization efficacy can not be confirmed for the current methods of sterilization listed in the product labeling. | Other | Recalling firm sent Urgent Medical Device Recall letters, dated 2/2009, advising customers to discontinue use of the listed sterilization methods and to remove them from service since there is no sterilization method that can be supported by current data. They are to use the enclosed form to identify and confirm if they have the listed products and follow the instructions on the form for returning information to the firm. |  | 3218 | TOTAL: 3218 units; 2052 in the US, 1166 internationally | Worldwide Distribution | Terminated | August 17, 2011 | 912 | 3218
Z-1174-2009 | 51402 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Integra Bilayer Matrix Wound Dressing, 8 in X 10 in (20 cm x 25 cm); Non-pyrogenic, Do not Re sterilize, Store flat at room temperature, Avoid Excessive heat, Avoid freezing. Sterile. Catalog number BMW810. |  |  | March 10, 2009 | 2009 | Integra LifeSciences Corp. | Package Seal Integrity: Integra became aware of the potential of open seals in the foil pouch of the Integra Bilayer Matrix Wound Dressings. | Nonconforming Material/Component | Urgent: Product Recall Notification letters, dated March 10, 2009, were sent out by first class mail to all consignees. The letters asked consignees to quarantine any stock on hand and contact their local Integra Sales Representative for further assistance and for arrangement of product returns. Consignees also need to complete and return the "Recall Acknowledgement and Return Form" indicating whether or not they possess any affected product. Additional questions should be directed to Judith O'Grady at 609-936-2311. |  | 125 | 125 units | Nationwide Distribution -- including states of CA, FL, MN, TX, NC, GA, NY, OH, MI, NE, and PA. | Terminated | June 26, 2009 | 108 | 125
Z-1170-2009 | 50795 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ReadyMed Elastomeric Ambulatory Pump, Model Numbers: R050050, R100100, R100200, R250050, R250100, and R250167. |  |  | December 19, 2008 | 2008 | Cardinal Health 303 dba Cardinal Health | The recall was initiated as a result of a potential for out of specification flow rates. The product may not consistently meet the labeled flow rate range of +/- 15%. | Process design | Cardinal sent an Urgent: Medical Device Recall Notification letter, dated 12/19/2008, to consignees on December 22, 2008 stating the Public Reason for Recall, and informing them of the Short Term and Long Term Actions. In addition a modification to the "Dear Pharmacist" letter, FAQ and Recall Response Card was mailed beginning January 26, 2009 to clarify that the pharmacist's ReadyMed products should be sent directly to their distributor and the credit for their product will also be provided through their distributor. Cardinal Health will work directly with the distributor and not the pharmacist as they are not Cardinal Health's direct consignee.Short Term: All ReadyMed elastomeric ambulatory infusion system models are subject to this recall. Cardinal Health is requesting that the use of the ReadyMed elastomeric ambulatory infusion system be discontinued and that the product be returned to Cardinal Health for credit.-Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card. Note on the card the ReadyMed elastomeric ambulatory infusion system models, lot numbers and quantities that will be returned. Cardinal will provide a courier mailing package for that return and will process consignees' credit.-Customers who purchased through a distributor will receive their credit through that distributor and not directly from Cardinal Health.-Consignee distributor has already been notified of this recall. If you have any questions, please contact your distributor directly, or call the Cardinal Health ReadyMed Recall Call Center at the number below.-Enclosed is a Frequently Asked Questions (FAQ) Letter for Pharmacists to address any questions or concerns related to the use of the ReadyMed elastomeric ambulatory infusion system.Long Term Actions: cardinal Health is continuing to investigate the out of specification condition of the ReadyMed elastomeric ambulatory infusion system and reviewing possible corrective measu |  | 2600000 | Approximately 2.6 million units | Nationwide Distribution | Terminated | September 13, 2011 | 998 | 2600000
Z-1096-2009 | 50696 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dual Port Wizard Low Profile Replacement Gastrostomy Device, REF 00200W, 16F x 1.2cm, 10cc, Sterile EO, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | December 22, 2008 | 2008 | C R Bard Inc | Out of specification replacement gastrostomy device was distributed. | Process control | Consignees were notified by "Product Recall" letter distributed on January 7, 2009 via Federal Express overnight mail and asked to check all inventory locations, remove unused affected product and return to Bard Access Systems, Inc (BAS). Customers were also instructed to complete the enclosed Reply Form and Inventory Reconciliation Form and fax it to BAS at 1-801-522-5674. Additional information is available by calling BAS Customer Service at 1-800-290-1689. |  | 450 | 450 units | Worldwide Distribution, USA (states of AL, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, NH, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV) and countries of Canada, Japan, Finland, Norway, Sweden, Austria, Belgium, Holland, United Kingdom, Italy, France, Switzerland, Spain and Germany, Australia, and Latin America. | Terminated | May 18, 2010 | 512 | 29939
Z-1172-2009 | 51135 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic RV Lead Integrity Alert , Cat. # SW012, 1.0 software in EnTrust(R) (D153ATG, D153VRC, D154ATG, D154VRC) and EnTrust(R) Escudo (D144DRG, D144VRC) defibrillators. |  |  | November 03, 2008 | 2008 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic has identified an issue related to installation or removal of the Lead Integrity Alert software in EnTrust(R) and EnTrust(R) Escudo" defibrillators. In those devices only, installation or removal of LIA will inadvertently turn off two audible patient alerts. | Software design | Between November 24 and December 9, 2008 US CRDM are responsible to hand deliver the consignee letter, a Medtronic "Urgent Patient Management Information," dated November 2008, to impacted physicians to make them aware of the issue and continue to promote and install LIA (by using the new LIA tipcard). For any physician listed that a confirmation sheet was not received by end of day, December 9, 2008, the letter was mailed to the physician on Dec 10, 2008. The letter described the issue and the product involved. The letter also gave recommendations and informed consignees in regard to a future update to the LIA software for EnTrust devices. |  | 55755 | 55,755 | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Netherlands, Netherlands Antilles, Norway, Oman, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom Vatican City State, Venezuela, and Virgin Islands. | Terminated | May 10, 2012 | 1284 | 55755
Z-1072-2009 | 50771 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Soft Tip Sheath Multipurpose Introducing Sheath Kit, Model Numbers: |  |  | December 24, 2008 | 2008 | Boston Scientific Corp | The marker band may detach from the sheath and a clinical failure could result in an embolism. | Process design | The firm initiated the recall on December 24, 2009 and notified consignees via "Urgent Company Initiated Medical Device Recall - Immediate Action Required" letters which explain the reason for the recall and request the affected product be returned to the Boston Scientific Corporation Distribution Center. An amendment letter with two additional Convoy Advanced Delivery Sheath Kit models (M0045748ST0 and M0045749ST0) was sent to affected customers on January 27, 2009. Direct questions about this recall to your local Sales Representative. |  | 31340 | 31,340 of all units | Worldwide Distribution. | Terminated | April 06, 2010 | 468 | 31340
Z-1181-2009 | 51300 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pilling(R) Coronary Scissors, 120Deg w/bar, Catalog numbers: and 352178, Teleflex Medical, Research Triangle Park, NC 27709 |  |  | March 03, 2009 | 2009 | Teleflex Medical | The product has a rough square finish instead of a smooth rounded finish which can cause damage to arteries. | Process control | Consignees were notified by an Urgent Medical Device Recall letter, dated 03/03/2009, via UPS Next Day. Hospitals were instructed to immediately discontinue use of the products, inspect the scissors for the date codes and to return the product to Teleflex Medical for replacement. Distributors were instructed to forward the letter to their consignees to retrieve relevant product from those locations. A Recall Acknowledgement Form was attached to be completed and faxed back to Teleflex Medical. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or they may be contacted by phone, e-mail or fax. A record of these contacts will be maintained. Affected product retrieved by Teleflex Medical will be destroyed and replacement product provided to the consignee. Records of destruction will be maintained. |  | 497 | 497 of all units | Worldwide Distribution -- USA including states of AL, AZ, CA, CO, DC, DE, FL, IA, IN, KS, MA, MD, MI, MO, MS, NH, NJ, NY, NV, OH, OK, OR, PA, RI, TN, TX, UT, and WA, and countries of Australia, Canada, France, Germany, India, Japan, and Singapore. | Terminated | March 31, 2011 | 758 | 497
Z-1065-2009 | 50171 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001. |  |  | November 12, 2008 | 2008 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Software change control | Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. |  | 22 | 22 systems in all | Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. | Terminated | September 22, 2009 | 314 | 22
Z-1064-2009 | 50478 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-Dyn Ruby Hematology Analyzer, List Number 08H67-01 |  |  | November 17, 2008 | 2008 | Abbott Laboratories | Carryover failures in software revisions 1.0ML and 2.0ML result in elevated PLT background count. | Other | Recall initiated on 11/17/2008. A Product Correction letter and customer reply form is sent to all currently active CELL-DYN Ruby customers. The customers are asked to keep the communication with their CELL-DYN Ruby System Operator's Manual and to note that the information listed in the table replaces the carryover specification provided in all applications of the manual. Questions from customers in the US should be directed to Customer Support at 1-877-4ABBOTT. Customers outside the US should contact their local hematology customer support representative. |  | 899 | 899 units | Worldwide Distribution | Terminated | November 27, 2010 | 740 | 899
Z-0986-2009 | 50753 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Trio Small Offset Connector |  |  | September 05, 2006 | 2006 | Stryker Spine | Reports were received of the ball ring ejecting from the connector's body when other connectors have already been locked down. | Device Design | A notification letter dated September 12, 2006 and a Product Accountability Form were sent via Federal Express (with return receipt) to all affected branches/agencies. Further information can be found by calling 1-201-760-8256. |  | 2417 | 2,417 units | Nationwide Distribution including OR, CA, MA, VA, GA, FL, NY, NM, NJ, UT, KS, OH, MI, MN, IN, MT, TX, WI, AL, CO, and TN. | Terminated | April 14, 2009 | 952 | 2417
Z-1121-2009 | 50927 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare, Infinia HawkeyeÂż 4, Model Numbers: 5178728, ASM500060, |  |  | June 02, 2008 | 2008 | GE Medical Systems, LLC | The metal pin in the adjustable head holder, which is used to acquire Brain SPECT data on Infinia systems (Infinia, Infinia HawkeyeÂż, and Infinia HawkeyeÂż 4), creates an artifact on the patient's cerebellum, which may lead to the potential for misdiagnosis. | Other | Consignees were sent a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. Contact your local sales or service representative for further information regarding this notification by calling 1-800-437-1171 (within the US or Canada) and 1-262-896-2890 (outside the US). |  | 288 | 122 US and 166 OUS | Worldwide Distribution -- USA (DC and Puerto Rico) including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and countries ofAustralia,Austria,Belgium,Brazil,Canada,China,Colombia,Czech,Denmark,Ecuador,Egypt,Finland,Finland,Egypt,Finland,France,Germany,Greece,Hungry,India,Indonesia,Ireland,Israel,Italy,Japan,Korea,Malaysia,Martinique,Mexico,Netherlands,New Zealand,Norway,Pakistan,Panama,Philippines,Poland,Portugal,Republic of Georgia,Russia,South Africa,Spain,Sweden,Switzerland,Taiwan,Thailand,Turkey,Ukraine, and United Kingdom. | Terminated | May 24, 2012 | 1452 | 1256
Z-1094-2009 | 50990 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia Hawkeye(R), Model numbers: |  |  | September 08, 2008 | 2008 | GE Medical Systems, LLC | There is a possibility for misregistration when brain scans are performed with Infinia with the Hawkeye option. This misregistration could present the potential for misdiagnosis when the patient headrest sags under the weight of a head. | Software design | Consignees were sent from September 2008 - Feb 2009 a GE "Product Safety Notification" letter dated 4/2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. It instructed consignees to continue to use the Infinia Hawkeye system normally for all types of Hybrid scans and stated that as a general practice tomograms should be reviewed carefully using the Xeleris Volumetrix processing protocol to ensure proper registration. Local sales or service reps should be contacted for further information regarding the notification. |  | 417 | 153 US, 264 OUS | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, DC, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK ,OR, PA, PR, SD, TN, TX, UT, VT, VA, WA, WV, and WI and countries of AUSTRALIA, BELGIUM , BRASIL, CANADA, CHINA, COLUMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, FINLAND, FRANCE, GEORGIA, GERMANY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, NEW Zealand, NORWAY, POLAND, PUERTO RICO, RUSSIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | May 24, 2012 | 1354 | 639
Z-1156-2009 | 51195 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | NurseMate with Physio Module; |  |  | February 13, 2009 | 2009 | St. Jude Medical | Installation problem: there is a risk that the monitor may not be adequately secured to the stand, resulting in the monitor falling off which may cause injury. | Nonconforming Material/Component | St. Jude Medical received a recall notice dated January 16, 2009 from its supplier, Smiths Medical PM, Inc. St. Jude Medical prepared and sent their own recall letter dated February 13, 2009 to affected customers. The recall letter described the issue and possible associated risks and provided instructions to customers. A St. Jude Medical representative will contact customers to facilitate replacement of the fasteners. Questions should be directed to your St. Jude Medical sales representative. |  | 4 | 4 units | Nationwide Distribution -- CA, NV, and FL. | Terminated | May 04, 2009 | 80 | 4
Z-1154-2009 | 51297 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Alm Prismatic Surgical Light 5004 SF; |  |  | February 23, 2009 | 2009 | MAQUET Inc. | The screws holding the Prismatic triple surgical lights may break due to dynamic multidirectional bending stress. The firm has received one complaint report. | Device Design | An "Urgent Device Correction" letter dated February 23, 2009 was mailed by FedEx Tracking to all affected consignees. The letter described the issue, potential hazard and recommended precaution. Customers were instructed to complete the enclosed "Customer Response Fax Form" and return it via fax (1-732-667-1903) to Maquet, Inc. Maquet, Inc. will schedule repair service of the affected units once the "Customer Response Fax Form" is received. Direct questions to your local Maquet, Inc. Representative. |  | 54 | 54 of all units, both article numbers | Nationwide Distribution. | Terminated | September 29, 2009 | 218 | 54
Z-1063-2009 | 50850 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | AxSOS Torque Limiting Screwdriver, T8 1.5 Nm 3.0mm Locking Set. REF 702749. This product is non sterile. |  |  | November 24, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Potential Breakage: Stryker Orthopaedics became aware that there is the potential for the tip of the AxSOS Screwdriver and the AxSOS Torque Limiter to break. | Other | "Urgent Product Recall" letters were sent to Stryker Orthopaedics Branches/Agencies, Chief of Orthopaedics, Hospital Risk Management and Surgeons on November 24, 2008 via Federal Express. The letters described the issue and provided the following instructions:1. Examine your inventory and determine if you have any of the affected product,2. Remove and quarantine affected product from your inventory,3. Complete the attached Product Recall Acknowledgement Form (and fax to Stryker Orthopaedics at 201-831-6069), and4. Return the affected product to your Stryker Orthopaedics Sales representative. Direct questions to your Stryker Orthopaedics Sales Representative or call Stryker Orthopaedics at 201-831-5000. |  | 415 | 415 units | Nationwide Distribution. | Terminated | April 06, 2010 | 498 | 882
Z-1173-2009 | 51292 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Medtronic STREAMLINE Bipolar Temporary Myocardial Pacing Lead, 6495, Sterile EO, REF 6495F, Medtronic, INC., Minneapolis, MN 55432 |  |  | February 06, 2009 | 2009 | Medtronic Inc | Medtronic is recalling one lot of STREAMLINE Bipolar Myocardial Temporary Pacing Leads, Model 6495, Lot JGD880205F. This lot of product contains incorrect inner pouch labeling, instead of 6495 it is labeled 6500. The outer carton, shelf box labeling, and instructions for use are correct, and the actual product shipped was consistent with the product ordered. No adverse patient effects have b | Process control | Consignees were sent on 2/6/09 a Medtronic " Urgent Medical Device Recall Notice" dated February 6, 2009. The letter was sent to Risk Mangers. The letter described the product, problem, and requested the removal and quarantine of the affected, unused product. They asked consignees to sign and return the attached recall certificate regarding the status of affected product remaining at their facility. Their Medtronic Sales Representative will contact them to arrange the return of any affected inventory or customers may follow the instructions provided on the attached certificate. |  | 612 | 612 | Nationwide Distribution -- including states of AK, AZ, CA, CO, CT, ID, IL, IN, KS, MI, MS, NE, NV, NH, NM, NY, NC, OH, OR, PA, TN, TX, VA, WA, and WI. | Terminated | May 07, 2011 | 820 | 612
Z-1046-2009 | 50816 | No | LOM |  |  | Microbiology | PMA | AxSYM HBsAg Reagent Pack; Abbott Laboratories, Abbott Park, IL 60064; AxSYM HBsAg is a microparticle enzyme immunoassay (MEIA) intended for the qualitative detection of hepatitis B surface antigen (HBsAg) in neonatal serum, and adult and pediatric serum (including serum collected in serum separator tubes) or plasma (collected in potassium EDTA, sodium citrate, sodium heparin, lithium heparin, or plasma separator tubes); list 9B01-20; |  |  | January 19, 2009 | 2009 | Abbott Laboratories | Defective caps: Customers have observed reagent bottles not opening appropriately, which could lead to a probe crash on the instrument. | Employee error | Consignees were sent a Product Correction-Immediate Action Required letter on 1/19/09. The letter informed users of the reported increased frequency of probe crashes when using the affected reagent pack lots and provided instructions on what actions the laboratory can take if they have a defective reagent pack or experience a probe crash with their system. Users need to inspect each reagent pack for broken or improperly sealed caps; manually open each bottle to ensure the flipper bar raises the reagent cap correctly; close each flipper bar prior to loading reagent pack on instrument; and discard any reagent packs that do not open properly. If a probe crash is observed prior to testing samples again, a probe crash recovery must be performed following the procedure in the Operations Manual. A copy of the communication needs to be provided to any laboratory that you may have forwarded affected lots to. Users were requested to return a customer reply letter acknowledging receipt and understanding of the letter. For additional information, contact Abbott Labs at 1-877-4ABBOTT. |  | 1540 | 1,540 kits | Worldwide Distribution -- USA including Puerto Rico and countries of Germany, Mexico, Hong Kong, and Japan. | Terminated | September 11, 2010 | 600 | 10123
Z-1198-2009 | 51178 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Low-Level Output Cable. |  |  | February 04, 2009 | 2009 | Datascope Corporation | Non-functioning cable: the Low-Level Output Cable when connected to a Datascope Intra-aortic Ballon Pump (IABP) and patient monitor does not provide a signal to the patient monitor. | Process design | An "Urgent Medical Device Recall" letter dated February 4, 2009 was delivered to affected consignees by Datascope Field Representatives. Direct questions to your local Sales/Service Representative or call the Datascope Corporation at 1-973-244-6314. |  | 20 | 20 | Nationwide Distribution -- states of IL and FL. | Terminated | August 11, 2009 | 188 | 20
Z-0989-2009 | 50872 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂż Insulin Pump, Model 1700 containing 1800 software. |  |  | November 20, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has become aware of a display irregularity with the Deltec CozmoÂż Insulin Pump, Model 1800 and Model 21-1700 containing Model 1800 software. There have been adverse events reported that the amount of Extended Bolus delivered was not accurately displayed by the Pump. The amount of extended Bolus displayed on the Pump Home Screen 2 and in the Bolus Summary Report is less than what w | Software design | Consignees were sent a Smiths Medical "Urgent Device Recall" letter dated November 20, 2008. The letter was addressed to Distributors, Clinicians, Patients and other Customers. The letter included a Description of the Problem, Advice on Action to be Taken by the User and requested the return of the Confirmation Form.In particular, users were given the option to check the "Replacement Pump" option on the Confirmation Form initiating the process for pump replacement with a recertified pump followed by a return of the defective product. The Confirmation Form can be faxed to 1-800-628-6322 in the U.S. or 1-651-628-7485 outside the U.S., completed at www.cozmore.com/notice or calling 1-800-501-5748 in the U.S. |  | 586 | 586 | Worldwide Distribution --- Including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MD, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and DC and countries of AUSTRALIA, NEW ZEALAND, UNITED KINGDOM, IRELAND, CANADA, CHINA, GREECE, POLAND, SLOVENIA, SWEDEN, ISRAEL and BRAZIL. | Terminated | June 05, 2012 | 1293 | 17047
Z-1212-2009 | 51080 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | 8 Fr 30cc Ultra 8 IAB, |  |  | February 02, 2009 | 2009 | Arrow International, Inc. | A fault in the connector of the pump tubing assembly may result in the volume setting on the pump defaulting to 2.5 ccs or 5 ccs (depending on IABP model) rather than the appropriate 30, 40, or 50 cc volume. | Process design | An "Urgent Medical Device Recall" letter dated February 2, 2009 was issued to consignees via US Certified Mail to notify users of the problem. Users were instructed to check their inventory for the affected products, quarantine, and return the them to Arrow International, Inc. Customers were also instructed to follow the "Mitigation Directions" in circumstances where alternate Intra-Aortic Balloon (IAB) catheters/adaptors or other forms of therapy are not available (until receipt of available product). Canadian accounts were notified on February 9, 2009 by courier and International accounts notified on 6 February 2009 and 10 February 2009 via Registered Post/UPS.Direct questions to your local Sales Representative or the Arrow International Customer Service at 1-800-523-8446 (US) or 1-919-361-4062 (outside US).Press Release issued April 13, 2009 is posted to firm's webiste:http://www.arrowintl.com/iab_recall.asp. |  | 4258 | 4,258 units | Worldwide Distribution: USA, Argentina, Australia, Taiwan, Singapore, New Zealand, India, Brazil, Belgium, Korea, Peru, Finland, and Denmark. | Terminated | July 06, 2010 | 519 | 61778
Z-1203-2009 | 50529 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | 300 microliter CO-RE tips for use with Microlab STAR series instruments, Part Number 235902, 5,760 tips per case, manufactured by Hamilton Company, Reno, NV. |  |  | December 01, 2008 | 2008 | Hamilton Co | Some tips may be occluded and are unable to aspirate or dispense liquid. | Process control | An "Urgent: Product Recall Letter" dated December 1, 2008 was issued to consignees via certified mail and e-mail. The letter described the problem and potential effects of use. Consignees were instructed to complete the enclosed "Hamilton Company Product Recall Questionnaire" and return it via e-mail or fax (775-856-7259) to the Hamilton Company. After receipt of the completed questionnaire, the firm's Customer Service Department will contact consignees to provide Returned Material Authorization information for returning affected product. Direct questions to the Hamilton Company at 775-858-3000, ext. 236. |  | 28 | 28 cases | Nationwide Distribution -- MA, CA, MN, CT, and NC. | Terminated | July 21, 2009 | 232 | 28
Z-1251-2009 | 51212 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Replacement System (MRS) Cemented Stems, Sterile, Catalog Numbers: 6485-3-000, 6485-3-008, 6485-3-009, 6485-3-010, 6485-3-018, 6485-3-019, 6485-3-300, 6485-3-308, 6485-3-309, 6485-3-310, 6485-3-318, and 6485-3-319. |  |  | January 30, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Instruction for use incorrect: There are no statements or warning on labels or Instructions for Use that 8, 9 and 10 mm diameter MRS Cemented Stems are indicated only for use in the GMRS Proximal Tibial and Distal Femoral Components. | Labeling False and Misleading | Urgent Product Recall letters, dated 1/30/2009, were sent to all Branches/agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons on January 30, 2009 via Federal Express. The letter asked consignees to examine their inventory and determine if they had any affected products, remove and quarantine affected product from their inventory, and complete the attached Product Recall Acknowledgment Form. The letter also stated that a Stryker Orthopaedics Sales representative would contact the consignee to reconcile and return the product . Consignees should contact their Stryker sales rep if they weren't not contacted by them. Questions should be directed to Colleen O'Meara at 201-831-5970. |  | 1161 | 1161 units | Nationwide Distribution | Terminated | April 08, 2010 | 433 | 1161
Z-1202-2009 | 51042 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Direct Bilirubin Reagent Set, Pointe Scientific Inc., Canton, MI. |  |  | February 02, 2009 | 2009 | Pointe Scientific, Inc. | The product performance fails due to precipitation of the R2 component. | Process control | Consignees were notified by "Urgent: Medical Device Recall" letter dated February 2, 2009 and mailed on February 6, 2009, and instructed to conduct the recall to the retail level if they had further distributed the product, to destroy any stocks on hand, and to complete and return the "Recall Return Response Form" via mail, e-mail, or fax (734-483-1592) to Point Scientific, Inc.Direct questions to the firm's Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 24 | 24 kits | Worldwide Distribution -- USA, Chile and Mexico. | Terminated | October 13, 2010 | 618 | 24
Z-1239-2009 | 51492 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Intellicath CCO/VIP Thermodilution Catheter with AMC Thromboshield (an Antimicrobial *Heparin Coating) For use with Vigilance Monitor. REF: 139HF75 |  |  | February 09, 2009 | 2009 | Edwards Lifesciences, Llc | Cardiac output measurements were unavailable and the following message is displayed in the Vigilance monitor: "Fault. CCO Catheter Verification. Use Bolus Mode". | Software design | Edwards initiated the worldwide recall via Urgent Customer Notification letters dated February 9, 2009 and are being sent by Federal Express starting on February 10, 2009. These letters are requesting that customers cease using Swan-Ganz" Continuous Cardiac Output (CCO) Thermodilution Catheter products from the indicated lots and that they return all unused product. Edwards Lifesciences requested that the customer provide written verification of stock and identify any unused product to be returned. Questions should be directed to Technical Service at 800-822-9837. |  | 107 | 107 units | Worldwide Distribution -- AL, AR, AZ, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, Hawaii, and Puerto Rico and countries of Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, UK, Greece, Ireland, Israel, Italy, Kuwait, Netherlands, Norway, Portugal, Sweden, Colombia, Australia, and Singapore. | Terminated | April 11, 2011 | 791 | 10754
Z-1250-2009 | 51499 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360, CSI ViperWire Orbital Atherectomy Guide Wire, REF: GW14335FRM. Sterilized with Ethylene Oxide. Manufactured for Cardiovascular Systems, Inc., St. Paul, MN 55112, USA. The ViperWire atherectomy guide wire has a shaft diameter of 0.014" and is constructed of stainless steel with a smooth finish and is silicone coated. The spring tip configuration is atraumatic, radiopaque, and can be bent to form a steerable system. |  |  | March 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is recalling 10 devices due to a mislabeling of the shelf life as shown on the device labeling. The labeling shown on the affected devices represents a 24 month shelf life rather than a 6 month shelf life that should have been indicated. | Incorrect or no expiration date | Customer letters were FedEx'd to CSI Sales Professionals on 03/16/2009, who delivered the letters and helped coordinate the removals of affected product. The letter identified the product, provided a description of the issue and provided instructions to immediately remove from use and return the affected guidewires to CSI. A Customer Acknowledgement Form was asked to be completed and returned with the devices. |  | 10 | 10 devices | Nationwide Distribution | Terminated | April 20, 2012 | 1131 | 10
Z-1207-2009 | 51316 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Biomet Marrowstim Concentration Kit w/single Disposable Separator, Bone Marrow Needle, and 30 ml ACD-A. For use with BBI centrifuges only, sterile, Biomet Biologics, Warsaw, IN; REF 800-0611A and 800-0613A. |  |  | March 05, 2009 | 2009 | Biomet, Inc. | A 30 ml tube is packaged with the kit instead of the labeled 60 ml tube. | Process control | Consignees were notified by "URGENT MEDICAL DEVICE REMOVAL NOTICE " dated March 5, 2009 and instructed to locate and return the recalled products. Direct questions about this recall to Biomet Orthopedics at 1-800-348-9500 or 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 27 | 27 total kits of both products | Worldwide Distribution -- USA including states of Massachusetts and Minnesota and countries of Germany and Italy. | Terminated | July 07, 2009 | 124 | 27
Z-0987-2009 | 50829 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD Medication Cassette Reservoir 100 ml with Clamp and Female Leur, Nonvented stopper included REF 21-7002-24, Sterile EO. |  |  | December 09, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical noticed an increased trend in leakage associated with 100 ml CADD Medication Cassette Reservoirs Reorder Number 21-7002-24 Lot Numbers 213X18 and 214X18. If a leak occurs in a CADD medication Cassette Reservoir, it may allow spillage of medication, contamination of the fluid path, and/or air to enter the Reservoir. Smiths Medical has not received any report of patient injuries due | Process control | An "Urgent Medical Device Recall" letter was sent via electronic mail to consignees on December 8, 2008. The letter was addressed to Distributors, Clinicians and other Customers. The letter described the problem and affected product and listed the "Advise on Action to be Taken by the User". Smiths Medical MD, Inc. requested the user to complete and return the "Urgent Medical Device Recall Confirmation Fax Return Form" and fax to the firm at 1-651-628-7485 or e-mail to info.md@smiths-medical.com (in US) or international.customer.service@smiths-medical.com (outside US). Direct questions to the Smiths Medical MD, Inc. Customer Service Department at 1-800-426-2448, Option 1. |  | 214416 | 17,868 cassettes (1,489 boxes 12 cassettes/box) | Worldwide Distribution USA (state of PA) and countries of El Salvador, Australia, France, Finland, Sweden, and Germany. | Terminated | June 07, 2012 | 1276 | 214416
Z-1004-2009 | 50543 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reflection, Microstable Acetabular Liner, REF 71740250, 32mm ID, 50-52 OD, 1 each, sterile EO. Component is a liner used to hold the femoral head used in total hip replacement arthroplasty procedures. |  |  | November 21, 2008 | 2008 | Smith &amp; Nephew Inc | Liners measuring 28 mm ID and were mismarked as 32 mm. | Process control | All affected Smith &amp; Nephew International Distributors were notified via e-mail dated November 24, 2008. In a letter titled "Urgent: Product Recall Notice - Immediate Action Required", distributors were instructed to immediately notify their accounts concerning the recall. All affected product is to be located and returned to the firm immediately.Direct questions to Smith &amp; Nephew, Inc. by calling 1-901-399-5441. |  | 19 | 19 units | International Distribution to Canada, France, Germany and Switzerland. | Terminated | November 21, 2011 | 1095 | 19
Z-1003-2009 | 50918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia HawkeyeÂż 4, Model Numbers: 5178728, ASM500060, H3000WC and H3000WD. |  |  | June 02, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently became aware of excessive pressure applied by the collimator pressure sensitive cover of Infinia (Infinia, Infinia HawkeyeÂż, Infinia HawkeyeÂż 4) that may impact patient safety. It was reported at a customer site that the pressure used to activate the Pressure Sensitive Device (PSD) should be lowered and that the body contouring software requires correction to prevent po | Device Design | Consignees were sent from June 2008 to February 2009 a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.A GE Healthcare service representative will contact each consignee to schedule the installation of the Pressure Sensitive Device (PSD). Please contact your local sales or service representative for further information regarding this recall. Further information can be found by calling 1-800-437-1171 in the US or Canada or 1-262-896-2890 if outside the US. |  | 288 | 122 US, 166 OUS | Worldwide Distribution US (and Puerto Rico) including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV and WI and countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Hungry, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Georgia, Russia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine and the United Kingdom. | Terminated | May 24, 2012 | 1452 | 1256
Z-0995-2009 | 47442 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Demipulse Duo Generator, Model 104; an implantable, multiprogrammable, vagus nerve stimulator. |  |  | February 07, 2008 | 2008 | Cyberonics, Inc | Reset/disabling of the VNS Therapy Demipulse Generator and Demipulse Duo Generator due to magnet interference, resulting in the loss of stimulation. | Software design | Cyberonics notified consignees of recall via letter on 02/07/08. Consignees were notified of reset issue in which during a strong magnetic field, the software may report an error even though no error occurred. This error causes the generator to reset, which disables the generator (i.e., no stimulation delivered).Physicians should inform any potential and currently implanted patients of the possibility of this event. Patients should continue with their regularly scheduled visits as determined by their physician.If this event is suspected, physicians should interrogate the device as described in the Physicians Manual to confirm the generator has been reset and Cyberonics should be contacted for any additional instruction. |  | 108 | 108 units. | Worldwide Distribution - USA, Belgium, Germany, Norway, Spain, Sweden, the Netherlands, and United Kingdom. | Terminated | February 28, 2011 | 1117 | 795
Z-1205-2009 | 51215 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | PLAC Test Calibrator Kit, Five-level liquid calibrator set for use with the diaDexus PLAC Test Reagent Kit, assembled in a kit box with an outer label, Catalog numbers 90108 and 10-0108, manufactured by diaDexus inc., South San Francisco, CA. |  |  | February 09, 2009 | 2009 | DiaDexus, Inc | Premature Expiration-- the product is projected to fall short of expected shelf life, and may result in inaccurate results. | Process control | Distributors and Direct accounts were notified by phone followed by Notification letter and certificate of destruction via FedEx. International customers were notified via e-mail with PDF copies of notices, to be followed by hard copies via FedEx. Customers are requested to record the number of kits (full and partial ) they had on the Certificate of Destruction and to dispose of that number of kits with two people signing the form, and then to fax that form to diaDexus. |  | 105 | 105 kits | Worldwide Distribution | Terminated | January 05, 2010 | 330 | 105
Z-1069-2009 | 50530 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Hematology Reagent Syringe, List number 08H49-02, used on the CELL-DYN Sapphire Automated Hematology Analyzer, List Number 08H00-01, manufactured by Abbott Diagnostics Division, Santa Clara, CA. |  |  | December 10, 2008 | 2008 | Abbott Laboratories | Some syringes were improperly assembled and are therefore not usable, due to an incorrect nut on the syringe shell. | Process control | Recall initiated on December 10, 2008. A product recall letter and reply form were cent to CELL-DYN Sapphire customers who received syringes between September 16, 2008 and October 28, 2008 to US customers, and to November 5, 2008 for Outside-of-US customers. |  | 30 | 30 syringes | Worldwide Distribution -- USA, Australia, Canada, and Germany. | Terminated | December 27, 2010 | 747 | 30
Z-1200-2009 | 51044 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid Auto HDL Cholesterol Reagent Set, Pointe Scientific, Inc., Canton, MI, Catalog Numbers: 10-H7545-80, 12-H7545-162, 3-H7545-L, 7-H7545-1000, 7-H7545-320, 7-H7545-40, 7-H7545-80, H7545-1000, H7545-320, H7545-40, H7545-80, HH945-240, and HH945-480. |  |  | January 29, 2009 | 2009 | Pointe Scientific, Inc. | The R1 reagent has microbial contamination. | Nonconforming Material/Component | Consignees were notified by "Urgent: Medical Device Recall" letter dated January 29, 2009, to conduct the recall to the retail level, to examine inventory and dispose of affected stocks on hand, and to complete the "Recall Return Response Form". Consignees were instructed to return their completed forms via mail, fax (1-734-483-1592) or e-mail to Pointe Scientific, Inc. (PS). Direct questions about the recall to PS Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 590 | 497 kits, 93L | Worldwide Distribution -- US, Chile, Dominican Republic, Egypt, Ghana, Greece, India, Israel, Italy, Jamaica-West Indies, Korea, Malaysia, Mexico, Philippines, Poland, Romania, Serbia, Trinidad-West Indies, Turkey and Venezuela. | Terminated | October 20, 2010 | 629 | 1471
Z-1123-2009 | 51276 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Operon Electrical Operating Room Table. |  |  | February 02, 2009 | 2009 | Berchtold Corp. | Difficulty in pressing the button needed to release a bed part resulted in an injury to a caregiver's finger which became stuck in metal latch under the gas cylinder of the bed piece. | Device Design | An official notification letter dated February 3, 2009 was hand-carried by a local Berchtold Corporation Technical Service representative to each account. The letter indicated that the Berchtold Corporation Service Technician will complete recommended upgrades. Direct questions concerning this recall to your BERCHTOLD Technician Service at 1-800-243-5135 and press the Number 2. |  | 453 | 453 units | Worldwide Distribution -- US and Canada. | Terminated | June 04, 2009 | 122 | 453
Z-1197-2009 | 49019 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Immunoassay System Hybritech PSA Calibrators, Product Number: 37205. |  |  | January 08, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter incorrectly shipped Access Hybritech PSA calibrators containing dual calibration cards (Hybritech and WHO) to some Access Hybritech PSA Customers. The WHO calibration for PSA is not currently approved in the United States. | Packaging process control | An "Urgent: Product Corrective Action" letter dated January 10, 2008 was issued via fax or mail to customers. The letter described the issue and instructed consignees to discontinue use and discard the affected product. Consignees were also asked to complete and return a "Fax-Back Form" via fax (1-952-368-7610) or mail to Beckman Coulter, Inc. within 10 days of receiving the notification letter.Further questions should be addressed to Beckman Coulter Customer Technical Support at 1-800-854-3633. |  | 123 | 123 Units | Nationwide Distribution -- Including states of AL, AZ, CA, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OK, OR, PA, SC, TN, TX, VA, WI and WV. | Terminated | March 15, 2012 | 1528 | 123
Z-1266-2009 | 51563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging model number 10014063. Affected systems are V30A or earlier with an NFS connected archive and V30A systems where image data was stored to LTS using software version older than V30A. |  |  | February 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | An issue concerning archived image data that is sporadically saved with corrupted data may occur. This issue may occur when syngo Imaging is configured to archive image data on a network file system (NFS) connected via a connected archive (Mount Point Solution). | Software design | A Customer Safety Advisory Notice, dated 02/09/2009 was sent to affected customers via Update Instruction IM013/09S. The letter informs customers of when the malfunction occurs and the potential issue, provides instructions to avoid its occurrence, and how the issue will be resolved. Customers are to ensure that the Safety Advisory is placed in the system's instructions for use and contact their SIEMENS Uptime Service Senter if they have any questions. |  | 20 | 20 units | Nationwide Distribution -- including states of AL, CA, CT, FL, LA, MI, MO, NC, NJ, OH, PA, TN, and WI. | Terminated | September 23, 2009 | 226 | 20
Z-1234-2009 | 51469 | No | ELR | Root canal post. | Dental | Dental | 510(K) Exempt | Dentsply, MAILLEFER, CYTCO-K TITANIUM POST SYSTEM, dental implant components, Model numbers (SKU): C106KL0050A20 (12 posts) and C117K00000020 (4 posts), Swiss Made, Manufactured for DENTSPLY Maillefer, Johnson City, TN 37604 |  |  | March 20, 2009 | 2009 | Dentsply Worldwide Headqrt | The product is manufactured with a left-hand thread instead of a right hand thread. This error could cause tooth breakage if the posts were removed without knowing that the threads were reversed. | Process control | Dentsply notified their distributors and dentists of the recall by email sent on 03/20/2009 by Urgent Recall Notice letter explaining the problem and requesting return of the product from their inventory and from their consignees. The letter requests faxback response (enclosed with the recall notification) within five (5) working days. The letter included a response page to ensure the receipt of the notice. A second follow up letter will be sent to all that do not respond to the first letter.The firm identified another distributor within the company, DENTSPLY Tulsa Specialty Products, who also sent out identical recall letters to those they distributed the product on 03/25/09. |  | 1588 | 1588 posts | Worldwide distribution: USA, Canada, Germany, France, Holland, Switzerland, The Netherlands, Italy, Kuwait, Poland, Argentina, Greece, Russia, Czech Republic, Belgium, Turkey, Hungary, India, and Colombia. | Terminated | February 12, 2010 | 329 | 1588
Z-1208-2009 | 51431 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Screwdriver for 3.5mm Screws, DARCO system, instrument for use in orthopaedic implant surgery. REF: DC4261, Non-sterile, Rx only, 1 each, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002 |  |  | February 19, 2009 | 2009 | Wright Medical Technology Inc | The package insert lists 3 sterilization procedures, 2 of which are not validated. Pre-vacuum steam sterilization is validated. | Process design | Wright Medical notified their distributors of the recall on 03/11/2009 by letter explaining the problem and requesting return of the product from their inventory and from their consignees. The letter included a "fax back" response page to ensure the receipt of the notice. The hospitals and surgeons were notified of the recall also on 03/11/2009. All letters were delivered by Fed Ex. |  | 682 | 682 devices | Worldwide distribution: USA, Canada, Germany, France, Belgium, the Netherlands, Italy, Finland, Austria, Denmark, Norway, Turkey, South Africa, United Kingdom, and Australia. | Terminated | February 03, 2010 | 349 | 682
Z-1235-2009 | 51423 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging, picture archiving and communications system (PACS), Model number 10 014 063 |  |  | January 12, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incorrect and/or incomplete data: Several potential safety issues while working with the syngo Image product including, but not limited to, data loss, errors in patient data due to different patient IDs, delays in image availability, inaccurate results during image measurements, incorrect review or diagnosis of displayed or printed images, measurement inaccuracies from the host modality and image | Software design | syngo Imaging has issued an Addendum to the User Manual, which was distributed to affected customers via Update Instructions IM118/08/S in November, 2008. Two safety related hotfixes (VB20G_HF4 and VB30A_HF26) will be released to those affected only after Update Instructions IM118/08/S has been performed. The addendum informs the user about safety relevant notes which pertain to syngo Imaging versions VB20G and VB30A. Update is a prerequisite for the following safety related hotfixes: VB20G_HF4 and VB30A_HF26. |  | 55 | 55 units | Nationwide Distribution. | Terminated | September 23, 2009 | 254 | 55
Z-1187-2009 | 51239 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2, 400 mL PainPump with 5.0 inch ExFen 2 Site Catheter Set, 5 per box, sterile, Stryker, Kalamazoo, MI; REF 540-50. |  |  | February 11, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | An electronic failure during preparation of the unit, or during patient use, may cause the pump to shut down. | Process design | A "Medical Device Recall Notification" letter dated February 11, 2009 was issued to consignees via phone, mail and/or email. The letter described the issue and instructed customers to discontinue use of the recalled product, destroy it according to instructions provided with the notification and return the "Business Reply Form" by fax to Stryker Instruments Regulatory Department, 1-866-521-2762.Further questions should be addressed to Stryker Instruments at 1-800-800-4236 extension 4086. |  | 24675 | 24,675 pumps | Worldwide Distribution -- USA including states of AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, WA, WI, WV and WY and countries of Argentina, Australia, Brazil, Canada, England, Singapore and South Africa. | Terminated | December 01, 2009 | 293 | 24675
Z-1319-2009 | 51439 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tri-State Centurion Endoscopy Tray, sterile, 10 per case, Tri-State Hospital Supply Corp., Howell, MI; Reorder MNS4635. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Compromised sterility, as the package seals may be inadequate. | Packaging process control | A letter titled "Urgent Product Recall Notice" was sent by Tri-State Hospital Supply Corporation to all consignees via certified mail titled dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand by sending the Accountability Information Form to the firm via fax (1-517-546-3356) or email and to hold the product for retrieval. Further questions should be addressed to Tri-State Hospital Supply Corp. at 1-517-546-5400 extension 1135. |  | 10 | 10 | Nationwide including states of NC, FL, TX, NJ, IL, MA, GA and IA. | Terminated | December 24, 2009 | 344 | 550
Z-1308-2009 | 51708 | No | OHS | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lumi 8 Intense LED Light Facial System, Sybaritic, Inc.,9220 James Ave South Minneapolis, Minnesota 55431., Mode of Operation: Operation with Intermittent/continuous load, Visible and Invisible LED Radiation, Class 3R LED Product, Refer to IEC 60825-2001 for further information., Input: 100 -230 V AC 50/60 HZ Power Input: 1509 V A....Protective Eyewear required Wavelength 590 &amp; 630 nm. |  |  | June 18, 2008 | 2008 | Sybaritic, Inc | Sybaritic stated that no 510(k) has been filed for the device (Lumi 8 Intense LED Light Facial System).No injuries or illness reported. | PMA | Letter sent to consignee on June 18, 2008 confirming a phone conversation on June 18 regarding the required Field Correction of the device. The letter also described the changes to be performed by the local technician to be in compliance with the FDA. The customer was asked to allow the technician access to their LUMI 8 system; allow technician to replace the computer chips and owner's manual; give current LUMI 8 owner's manual to technician; and sign the work order provided by the technician. Questions should be directed to Fred Berg at 800-445-8418 or fberg@sybaritic.com. |  | 1 | 1 | USA: MI | Terminated | July 05, 2012 | 1478 | 1
Z-1236-2009 | 51319 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cerner Bridge Medical Transfusion Administration and Specimen Collections, Versions 3.4.0, 3.4.1 and 3.4.2. |  |  | January 19, 2009 | 2009 | Cerner Corp | Specimen labels printing incorrect patient demographics. | Software design | Cerner Corporation issued a "Priority Review Flash" dated January 15, 2009 via e-mail to consignees using the software. The letter described the violative nature of the software and explained the proper navigation controls needed to prevent further issues.Further questions should be addressed to Bridge Client Support at 1-800-786-0239. |  | 4 | 4 | Nationwide distribution to states of OH, NC, MO and MI. | Terminated | March 22, 2012 | 1158 | 4
Z-1237-2009 | 51432 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Patient-Controlled Analgesia (PCA) Module (formerly Medley PCA Module) Model 8120 used with the Alaris PC Unit with software versions 8 through 9.1. |  |  | March 04, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | The Alaris PCA module can potentially infuse above or below the intended infusion dose. | Software design | The field action was initiated on March 11, 2009 with Cardinal Health sending by registered return receipt mail a Safety Alert to all customers using the Alaris PC Unit with software versions 8 through 9.1, which encompasses the subset of those customers that use the Alaris PCA Module. The Safety Alert was addressed to the Directors of Nursing, Pharmacy, Risk Management, and Biomedical Engineering at each hospital facility. If a customer did not respond to the Safety Alert, then Cardinal Heath notified that customer by telephone to verify receipt of the Safety Alert.This Safety Alert Informed the customers about a potential risk, and required them to: (1) Determine whether you use version 8 through 9.1 of the Alaris PC Unit. (2)Forward this Safety Alert to all Data Set owners (3) To immediately identify the potential risks review enclosed FAQs and Instructions for Data Set Review. (4)Continue to follow the Institute for Safe Medication Practices (ISMP) and the Joint Commission recommendation for safe medication administration - two nurses checking to ensure that the correct prescription is being entered into the PCA pump when first ordered and when there are any changes in the prescription (paying close attention to dosing units and infusion modes). The customers were also instructed to call our Support Center at 1-888-562-6018 with any questions. An updated Customer Recall Notification was forwarded by registered return receipt mail by Cardinal Health on April 15,2009. This notice reinforced the information in the first notice in addition to informing the customers of the hazard of an over-infusion or under-infusion of medication. Also, customers were given instruction for reporting any adverse reactions experienced with the use of the affected product and/or quality problems via the FDA's MedWatch Program. |  | 164000 | Approximately 164,000 units | Worldwide Distribution -- USA, Canada, Germany, Malaysia, and Singapore. | Terminated | June 26, 2012 | 1210 | 164000
Z-1317-2009 | 51562 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bed, Model P1900. |  |  | March 09, 2009 | 2009 | Hill-Rom, Inc. | An e-ring may not have been installed during production, which could result in an inability to latch the siderail. | Process control | Consignees were notified by an Urgent Medical Device Correction letter dated 3/9/09 and were instructed that a Hill-Rom technician will visit each consignee to inspect and to correct the beds, if necessary. The firm asks that consignees locate and test their beds to make sure the siderails are latching correctly; forward a copy of the letter to users that were loaned or purchased the device and notify Hill-Rom of their new locations or owners; and if affected beds were disposed of or are no longer in use, to inform the firm of their serial numbers. If there are questions customers should contact their Hill-Rom Technical Support for their geographic region. |  | 53 | 53 beds | Worldwide Distribution -- USA and Canada. | Terminated | October 13, 2010 | 583 | 53
Z-1232-2009 | 51253 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro" Software Versions 5.0.3, 5.0.6 ,5.0.7, 5.0.8 and 4.1. |  |  | February 27, 2009 | 2009 | GE Medical Systems Information Tecnology | GE Healthcare is aware of potential network time issues associated with the time change aspects of certain CARESCAPE CIC Pro" software versions. These issues may impact patient safety. No reported patient injuries or illnesses for the use of this device in conjunction with this device correction. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated February 27, 2009. The letter was addressed to the Chief of Nursing/Chief of Biomedical Engineering and Risk Manager/Healthcare Administrator. The letter describes the Safety Issue, Affected Product Details, Safety instructions, Product Correction, and Contact Information.Direct questions to your local GE Healthcare Field Service Representative or the GE Healthcare Technologies Support Center by calling 1-800-558-7044 (Option 2, Technical Support). |  | 12737 | 12,737 | Worldwide Distribution including US (states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, PR, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WY, WI, and WV) and countries of UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA, BOSNIA AND HERZEGOVINA, BANGLADESH, BELGIUM, BRAZIL, BAHAMAS, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, CUBA, CYPRUS, CZECH REPUBLIC, GERMANY, ESTONIA, EGYPT, SPAIN, FINLAND, FRANCE, GHANA, GREECE, HONG KONG, HUNGARY, HONDURAS, INDONESIA, IRELAND, ISRAEL, INDIA, IRAQ, BRITISH INDIAN OCEAN TERRITORY, ITALY, JORDAN, JAPAN, SOUTH KOREA, KUWAIT, LEBANON, LITHUANIA, MALTA, MEXICO, MALAYSIA, NIGERIA, NETHERLANDS, NORWAY, NEPAL, NEW ZEALAND, PANAMA, PERU, PHILIPPINES, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SW DEN, SINGAPORE, SLOVENIA, SLOVAK REPUBLIC, SYRIA, THAILAND, TURKEY, TRINIAD, TOBAGO, TAIWAN, UKRAINE, UNITED KINGDOM, VENEZUELA, YEMEN, and SOUTH AFRICA. | Terminated | September 15, 2014 | 2026 | 12737
Z-1244-2009 | 51421 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | TEVADAPTOR Closed Drug Reconstitution and Transfer System, a closed system used to prepare drugs, including cytotoxic drugs, for intravenous infusion or injection. Model number is MG412110. The device is a single use device. |  |  | December 18, 2008 | 2008 | Migada Inc. | Syringe adaptor is disintegrating when attached to the end of the syringe. | Component design/selection | A Field Safety Notice, dated 12/8/2008, was sent to all TEVADAPTOR customers advising them not to use drugs containing N,N-Dimethylacetamide. Teva Medical is presently revising the instruction leaflet supplied with each TEVADAPTOR to include a warning against the use of drugs containing N,N-Dimethylacetamide. Teva Medical will introduce this revised leaflet to current production as soon as practicable. In the meantime, Teva Medical will advise all users of the potential problem by means of the Urgent Field Safety Notice. Customers are to ensure that the safety notice is passed to all those who need to be aware within their organizations. Questions should be directed to their local contacts at B Braun, USA. |  | 900000 | 900,000 units | Nationwide Distribution | Terminated | July 27, 2009 | 221 | 900000
Z-1254-2009 | 51560 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Sireskop SD Model numbers 8890415 and 8890407. The product is intended for use as radiographic and fluoroscopic imaging device. |  |  | January 08, 2009 | 2009 | Siemens Medical Solutions USA, Inc | During interventional use and applications, the possibility exists for liquids to enter the systems. | Equipment maintenance | A Customer Safety Advisory Notice dated January 8, 2009 was issued to affected customers via Update Instructions AX069/08/S. This letter informed customers of the potential issues and provided instructions to avoid its occurrence. The firm has released an Addendum to the User Manual warning of the potential hazards if liquids penetrate the systems during interventional use and applications.Direct questions about this recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 240 | 240 units | Nationwide Distribution. | Terminated | September 01, 2009 | 236 | 909
Z-1328-2009 | 49127 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stat Profile Critical Care Xpress (CCX and CCX+), CCX Creatinine Membrane Kit for CCX analyzer |  |  | August 11, 2008 | 2008 | Nova Biomedical Corporation | Erroneous, low creatinine results on CCX analyzer | Process control | Initial:USA and Canada - Customer Advisory Notice notification by phone contact with key customer contact facilitated using a Field Correction Script. This contact will be documented within the Nova Biomedical Technical Support database (Pivotal) and on a Field Correction Checklist. All affected inventory will be replaced at customer sites.International - Customer Advisory Notice and Faxback Form being communicated and sent to Nova Biomedical Subsidiaries and Distributors by three methods of delivery (Email, Fax and Phone Call).Follow-up:USA and Canada - Sending the Customer Advisory Notice by email and/or fax to the key customer contact from the initial phone contact.International - Sending the Customer Advisory Notice by email and/or fax to the key contact. |  | 1230 | 1230 | Worldwide Distribution -- USA, Canada, Germany, France, Czech Republic, Taiwan, Japan, Thailand, Italy, Hungary, India, UK, Ireland, Venezuela, Brazil, and Hong Kong. | Terminated | July 22, 2011 | 1075 | 1230
Z-1233-2009 | 51494 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Novation DR, full field digital mammography system. Model Number 06646900. |  |  | January 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Image may appear dark, requiring repeat image acquisition. | Software design | A Customer Safety Advisory Notice dated December 16, 2008 was issued to affected customers via Update Instructions SP040/08/S. The letter described the potential issue and provided instructions to avoid its occurrence. A software update to correct the issue is currently being developed and will be installed on affected systems when it becomes available. |  | 497 | 497 units | Nationwide Distribution. | Terminated | March 19, 2013 | 1516 | 497
Z-1391-2009 | 51925 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Triage Cardiac Panel, Biosite Catalog Number 97000HS, 25 individually pouched devices in a labeled kit box. |  |  | April 27, 2009 | 2009 | Biosite Inc | The recall was initiated because the use of the affected lot may lead to reporting falsely low values for troponin I and myoglobin, possibly resulting in missed or incorrect diagnosis. | Process control | The recall was initiated with "Urgent Medical Device Recall" notices, dated 4/28/2009, with attached "Customer/Distributor Verification Form" being sent to the affected customers via fax, e-mail or direct mail on 04/28/2009. Customers were informed of the reason for the recall and instructed to immediately discontinue all use of the listed products and discard all affected product in accordance with your their regulations. The customers were give two options; "Option 1: Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice, to indicate the number of kits discarded from your inventory and product replacement will be sent to you. Option 2: Contact inverness Medical Technical Service at 1-877-441-7440 and select option 2 from the phone prompt. A representative will confirm the number of kits discarded from your inventory and arrange for replacement product to be sent to you."Customers were also provided telephone, fax and e-mail contact information for Biosite should they have any questions about the information contained in the recall notification. |  | 401 | 401 kits | Nationwide Distribution | Terminated | July 06, 2010 | 435 | 401
Z-1209-2009 | 51206 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | KyphX Xpander Inflatable Bone Tamp, size 15/3, Catalog Number K09A, Manufactured by Kyphon Sarl, Switzerland. For use in spinal surgery. |  |  | February 04, 2009 | 2009 | Medtronic Spine LLC, formerly Kyphon Inc | Product not sterile: 10 units were released for distribution without undergoing sterilization process. | Process change control | Sales representatives were contacted via e-mail on February 4, 2009, and instructed to contact hospitals and inform them of the issue and sequester affected product. They were further instructed to immediately travel to the hospitals and obtain affected product for return to Medtronic LLC. A Customer Communication Notice URGENT NOTICE: DEVICE REMOVAL dated February 9, 2009, was sent to the two consignees. |  | 10 | 10 units affected. | Product was distributed to two hospitals in OH and FL. | Terminated | July 14, 2009 | 160 | 10
Z-1252-2009 | 51473 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo Ventilator 300 (Model SV300) and |  |  | March 23, 2009 | 2009 | MAQUET Inc. | The expiratory valve in the SV 300 ventilator may stay in the open position if the printed circuit board fails. | Device Design | An "Urgent: Device Correction" letter dated March 23, 2009 was sent by Federal Express to customers. The letter described the affected product, problem, indications, potential hazard and actions for users. Customers were instructed to complete the "Device Correction -- Customer Fax Response" form and return via fax (732-667-1903) to Maquet Inc. RA/QA. Direct questions to your local Maquet Inc. representative. |  | 7847 | 7,847 units | Nationwide Distribution. | Terminated | November 16, 2009 | 238 | 7847
Z-1309-2009 | 51049 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL (Polyglactin 910) Suture, Ethicon Inc., A Johnson &amp; Johnson Company, Somerville, NJ 08876 |  |  | February 02, 2009 | 2009 | Ethicon, Inc. | Sterility Issue: The firm decided to remove the product because it cannot assure the sterility of the products due to a failure in the sterilization process. | Process control | Urgent Voluntary Product Recall notification letters were sent via UPS Next Day Air with receipt on February 2, 2009. Letters were sent to customers, Directors of Surgical Services, Director of Risk Management, Materials Manager and Distributors. The letters requested that customers immediately discontinue use of the products with lot numbers listed. Customers are to identify and return only the boxes identified in the attached tables. The affected product should be returned to Ethicon via the pre-paid authorized shipping label. Customers that are not returning affected product should mail the enclosed business reply card to Ethicon. Customers should direct questions regarding this recall to 1-888-202-3694 or their Ethicon sales representative.5/13/09: It was discovered that a lot of PDS II sutures that had been previously identified as affected but not distributed, were mistakenly distributed. Revised letters including this number were sent to 25 customers on 5/13/09. |  | 1758360 | 1,758,360 total | Worldwide Distribution -- USA, Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Greece, Hong Kong, Israel, Japan, Korea, Luxemburg, Mexico, New Zealand, Norway, Panama, Puerto Rico, Sweden, Taiwan, United Kingdom, Uruguay, and Venezuela. | Terminated | October 21, 2009 | 261 | 1758360
Z-1262-2009 | 51320 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Nexgen Complete Knee Solution Fluted Stemmed Tibial Broach, Size 7, 8, nonsterile; REF 5997-61-07. |  |  | March 13, 2009 | 2009 | Zimmer Inc. | The surface finish inside of the broach is not to specification and the instrument may corrode if cleaned and sterilized according to instructions. | Finished device change control | Distributors were notified by an "Urgent: Device Recall" letter dated March 13, 2009. The letter described the affected product, reason for recall, health risks and actions for distributors and user facilities. Customers were instructed to quarantine affected instruments and return them to the Zimmer Product Service Department. Distributors were also instructed to provide the name and address of hospitals where the instruments were used so that a copy of the recall letter can be forwarded to the user facilities.Direct questions about this recall to Zimmer, Inc. by calling 800-613-6131. |  | 37 | 37 | Worldwide Distribution -- US, Australia, and Germany. | Terminated | December 22, 2009 | 284 | 37
Z-1327-2009 | 47475 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Diagnostics Chemical Limited Iron Assay |  |  | March 14, 2008 | 2008 | Diagnostic Chemicals, Ltd. | Precipitate in product may cause elevated background absorbance interference. | Process control | Genzyme Diagnostics P.E.I. Inc. notified the 1 account by a Removal Notification letter on 3/14/08 requesting that they stop using and discard any remaining stock. The customer is to request replacement through their normal channels. Questions or concerns should be directed to Penny White of Genzyme. |  | 153 | 153 kits | Nationwide Distribution | Terminated | May 16, 2012 | 1524 | 153
Z-1263-2009 | 51447 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Carina Home Ventilator. |  |  | January 21, 2009 | 2009 | Draeger Medical, Inc. | The device may not generate a pilot line disconnect alarm if the pilot line becomes disconnected, which may potentially result in a substantial reduction in tidal volume delivery. | Software design | An "Urgent - Medical Device Recall" Letter dated January 2009 was issued to consignees via certified mail. The letter described the issue, specific conditions, background information, and precautions for customers. An additional letter dated January 22, 2009 was included with the "Urgent - Medical Device Recall" letter which instructed consignees to inform an enduser of the contents of the notification letter because some of the consignees are home care device/service providers. When the new hardware/software solution is available, affected devices will be updated free of charge. Direct questions about the notification letter to Drager Medical, Inc. by calling 1-800-543-5047 (press "1" at the prompt and then press "2349#"). Direct questions regarding updating your device when the new software is available, contact Drager Service Technical Support at 1-800-543-5047 (press "4" at the prompt). |  | 40 | 40 units | Nationwide Distribution -- states of CA, FL, GA, HI, IL, LA, MA, MN, OH, and TX. | Terminated | January 05, 2010 | 349 | 40
Z-1326-2009 | 50740 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical BCI(R) Advisor(R) Vital Signs Monitor, model 9200, catalog numbers: 920654225, 920674225, 920674235, 925454220, 925454235, 925454320, 925454325, 925454330, 925454335, 925457325, 925474325, 925654225, 925654325, 925754325, 925754335, and 92M774325. |  |  | December 22, 2008 | 2008 | Smiths Medical PM, Inc. | Recent improvements were made to the Advisor case plastic resulting in the inserts being deeply embedded into the plastic. This affects the ability of the monitor to be adequately secured on the wall mount bracket or roll stand bracket, resulting in the monitor falling off. There have been no adverse events or injuries reported. | Device Design | An Urgent Product Safety Information, Voluntary Recall Notification, dated 12/22/08, was sent to affected consignees. The notification contained a Safety Action Bulletin which provided actions to be taken. The Safety Action Bulletin described the issue, requested customer to contact Smiths Medical to ensure proper mounting hardware is utilized for those consignees using a roll stand or wall mount. If the consignee further distributed the product, they were asked to forward on the Safety Action Bulletin to the end user. Page 3 of the Safety Action Bulletin was asked to be completed and returned to Smiths Medical PM, Inc.A second "Urgent Product safety Information Voluntary Recall Notification" letter dated January 16, 2009 was sent to consignees. The second letter included a risk statement associated with the recalled product. |  | 143 | 143 monitors (17 US, 126 OUS) | Worldwide Distribution -- USA including states of FL, GA, NJ, NY, and TN and countries of Argentina, Brazil, Ecuador, Greece, Peru, Romania, Russia, Singapore, Slovenia, Trinidad and Tobago, United Kingdom, and Uruguay. | Terminated | December 17, 2011 | 1090 | 143
Z-1397-2009 | 50201 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity TriWin Laboratory Information System; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | November 17, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software defect in the Centricity Lab (TriWin) product may incorrectly link patient information to a wrong report. Also, when viewing results in Patient Viewer, and a coagulation result is higher or lower than the linearity limit, then all of the results following the higher or lower result will show the same result as the first flagged result. | Software design | GE Healthcare notified their Centricity TriWin Laboratory customer via letter dated 11/17/08 of the software anomalies resulting in patient safety issues involving the patient information incorrectly being linked to a wrong report, and when viewing results in Patient Viewer, and a coagulation result is outside of a linearity limit, then all of the results following the out of linearity result will show the same result as the first flagged result. The user was informed that a software patch has been prepared to resolve the issues and will be installed on their system. Any questions were directed to GE Technical Support at 888-778-3375 or 905-479-1249, select option 1. The accounts was requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 905-305-0046 to schedule the software upgrade. |  | 26 | 26 units | Worldwide Distribution -- USA and country of Canada. | Terminated | September 11, 2010 | 663 | 26
Z-1255-2009 | 51071 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche MagNA Pure LC Workstation; Product Numbers: 12236931001, 12236931444, 12236931692 and 03670325001. Robotic workstation for automated isolation of nucleic acids from different kinds of crude sample material such as whole blood, serum, blood cells, culture cells, tissue, bacteria and fungi. |  |  | February 09, 2009 | 2009 | Roche Diagnostics Corp. | Protruding screw threads may cut the user when cleaning the interior of the instrument. | Device Design | In a letter titled "Urgent Medical Device Correction" dated February 9, 2009, consignees were notified of the protruding screws and instructed to use caution when cleaning the instrument until a protective plate is installed by the firm's field representative.Further questions should be addressed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 371 | 371 | Nationwide Distribution including Puerto Rico and states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 07, 2009 | 301 | 371
Z-1261-2009 | 51687 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Sirus Intramedullary nail for tibia, cannulated, diameter 12 L, 330 mm, sterile, Zimmer GmbH, Winterthur, Switzerland; REF 02.02631.233. |  |  | March 23, 2009 | 2009 | Zimmer Inc. | The package is labeled as containing 330 mm length nails, but actually contains 300 mm length nails. | Other | Distributors were initially notified by phone on March 23, 2009. Distributors and user accounts were notified by an "Urgent: Device Recall" letter dated March 31, 2009. The letter described the affected product, reason for recall, risks to health, other information and actions for customers. Customers were instructed to quarantine recalled product and to return it to the Zimmer Distribution Center. For questions about the recall, contact Zimmer, Inc. at 1-800-613-6131. |  | 32 | 32 | Worldwide Distribution -- US (states of Florida, Missouri, New Jersey, New York, Ohio, Oklahoma and Texas) and countries of Austria, Japan, Korea and Germany. | Terminated | October 23, 2009 | 214 | 32
Z-1195-2009 | 51193 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. The Innova Biplane systems are the first and only digital flat panel biplane systems with a full sized lateral plane to cover lateral anatomy without requiring multiple contrast injections and radiation exposures. The product is indicated for use in generating fluoroscopic images of human anatomy for diagnostic and interventional cardiac angiography procedures. These systems can be operated in a mobile or fixed site environment. |  |  | November 21, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has become aware of a potential image storage failure on Innova 2121 IQ/ 3131 IQ cardiovascular imaging systems that may impact patient safety. This intermittent communication failure on the image storage chain may result in the following situations: Record abort, Record not available, Fluorostore function not available, Review not available, Boot/Reset failure. Fluoroscopy imag | Device Design | A GE Healthcare "Urgent Medical Device Correction" letter dated February 2, 2009 was sent to consignees via mail. The letter was directed to Hospital Administrator/Risk Manger, Managers of Radiology/Cardiology and Radiologists/Cardiologists addressing the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.Further questions should be addressed to GE Healthcare at 1-800-437-1171. |  | 85 | 85 | Worldwide Distribution -- USA including states of CA, CO, FL, HI, IA, IL, IN, KS, MD, MA, MN, MO, NY, OH, OR, PA, TN, TX, WA, WV, WI and VA and countries of the United Kingdom, Romania, Japan, Taiwan, Bahrain, France, Malaysia, India, Canada, Chile, Indonesia, Ireland and Italy. | Terminated | November 21, 2013 | 1826 | 85
Z-1395-2009 | 50531 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II Hemoglobin Testing System with CDM v 4.0, model number 270-2000, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | December 09, 2008 | 2008 | Bio-Rad Laboratories Inc | The wrong barcode number and tube position number may be assigned to samples leading to discrepant sample results. | Pending | Recall initiated on December 10, 2008. Notification conducted by phone call and fax for domestic consignees. Each subsidiary will receive the subsidiary Urgent: Medical Device Correction notice, customer medical device correction response form and number of customers distributed to by e-mail. Each subsidiary must contact customers and provide the customer notice and response form as needed. The notice stated that Bio-Rad Laboratories would release a follow up communication providing a permanent solution upon completion of an investigation. In the interim, it was recommended that customers carefully review the Summary Report for duplicate barcode IDs or or tube position numbers within the same rack. If there are any questions or assistance needed, the local regional Bio-Rad office should be contacted. |  | 42 | 42 units | Worldwide Distribution -- USA, Argentina, Australia, Brazil, Canada, China, Czech Republic, France, Germany, Hungary, India, Italy, Mexico, New Zealand, Singapore, South Korea, Sweden, Thailand, and the UK. | Terminated | November 12, 2010 | 703 | 42
Z-1199-2009 | 51455 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL Model M4735A Defibrillator, Automatic, External |  |  | March 30, 2009 | 2009 | Philips Healthcare Inc. | Rotary energy select switch in affected M4735A devices may fail and prevent the user from turningthe device on, rendering the device unusable for monitoring and defibrillation therapy. | Nonconforming Material/Component | An Urgent - Medical Device Recall Letter and Notice dated March 2009 were issued to affected customers. Mailings of the Urgent - Medical Device Recall Letter and Notice were handled by Philips Healthcare representatives in each affected geography. Customers were asked to follow the "Actions to be taken by customer/user" section of the Urgent - Medical Device Recall Notice. In this section, customers were instructed to remove the device from service if possible, while they await the correction of their device. If the customer is unable to remove the device from service, they are asked to identify a readily available backup device to use in the event that their device becomes unusable.All customers will be contacted by a Philips Healthcare representative to arrange for replacement of the energy select switch in the affected devices (Mandatory Field Change Order FCO86100076). Contact your local Philips Healthcare representative for additional information or support concerning this issue. |  | 26992 | 26,992 | Worldwide Distribution -- UNITED STATES, CANADA, ALBANIA, ARGENTINA, AUSTRIA,BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, DOMINICAN REPUBLIC, EGYPT,EL SALVADOR, FINLAND, FRANCE, GERMANY,GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND,ISRAEL, ITALY, JAPAN, JORDAN,KAZAKHSTAN, KENYA,KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACAU, MACEDONIA, MALAYSIA,MALI, MARTINIQUE, MAURITIUS, MEXICO, MONGOLIA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU,PHILIPPINES, POLAND,PORTUGAL, PUERTO RICO,ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA,SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN,TANZANIA,THAILAND, TRINIDAD AND TOBAGO, TURKEY,TUNESIA, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA,VIETNAM, VIRGIN ISLANDS and YEMEN. | Terminated | September 04, 2012 | 1254 | 26992
Z-1400-2009 | 51131 | No | GTQ | Streptococcus spp. exoenzyme reagents. | Microbiology | Microbiology | 510(K) Exempt | Remel Staphaurex Plus Kit, Catalog No. R30950102 (150 tests) and Catalog No. R30950201 (450 tests), rapid latex agglutination test for the identification of Staphylococcus aureus The responsible firm listed as Remel Europe Ltd, Dartford, England. |  |  | November 18, 2008 | 2008 | Remel, Inc | These lots may exhibit false positive reactions with strains of Staphylococcus epidermidis that may result in the mis-identification of Staphylococcus aureus. | Nonconforming Material/Component | Consignees were notified of the recall via letter dated 11/20/08 explaining the reason for recall and requesting any remaining units from the lots be discarded. A Product Inventory Checklist was enclosed to report the amount of product remaining in their inventory. Any non-responding consignees were contacted via telephone during the period of 12/30/08 to 1/11/09 to ensure the effectiveness of the recall. When the recall was expanded to include additional lots, the recalling firm issued an Important Medical Device Customer Recall Notice letters dated 3/16/09 again explaining the reason for recall. The letter also asked consignees to review their inventories for the product lots indicated and requested that remaining units from the lots be discarded. This letter also had a Product Inventory Checklist enclosed for completion. |  | 446 | 446 kits | Worldwide Distribution -- USA and Canada. | Terminated | December 03, 2009 | 380 | 446
Z-1411-2009 | 51886 | Yes | FLS | Apnea monitor. | Anesthesiology | Anesthesiology | 510(k) | SmartMonitor 2 Infant Apnea Monitor. |  |  | April 23, 2009 | 2009 | Respironics, Inc. | Audible alarm failure. | Process design | A notification letter dated April 23, 2009 was issued to consignees informing them of the issue. Consignees were instructed to locate the affected units in their inventory and return them to Philips Children's Medical Ventures-Youngwood Service. Any units that were shipped by consignees to their customers are to be retrieved and returned to Respironics. Consignees were also instructed to complete the enclosed Business Reply Form and Serial Number Reconciliation List and fax to Stericycle at 888-345-5369 or email to respironics1985@stericycle.com. For questions and additional information about the recall, contact Stericycle Customer Service by calling at 888-345-4630. |  | 4992 | 4,992 units | Nationwide and Canada. | Terminated | October 19, 2011 | 909 | 4992
Z-1196-2009 | 50202 | No | MDS |  |  | Clinical Chemistry | PMA | Abbott brand FreeStyle Navigator, Receiver Warranty Kit, Blood Glucose Monitoring System, Part Number: 70789-01. |  |  | November 01, 2008 | 2008 | Abbott Diabetes Care, Inc. | The product has the Units of Measure set to mmol/L rather than mg/dL. The blood glucose results could be misinterpreted. This may lead to under treatment and potential for hyperglycemia. | Packaging process control | The Consignees of the affected FreeStyle Navigator Receiver Kit was contacted via telephone, Federal Express or an Abbott representative beginning on November 1, 2008. Contact Customer Service at 1-866-597-5520 with further questions. |  | 2 | 2 units | US Distribution to states of NY and AZ. | Terminated | June 03, 2009 | 214 | 2
Z-1394-2009 | 51561 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas 6000 analyzer series (including cobas c 501 and cobas e 601 modules) Middleware Solutions Instrument Manager Server as follows: v.8.04.02.10, catalog no. 04785568001; v.8.04.03.01, catalog no. 04785568002; and v.8.06.02, catalog no. 05109019001. |  |  | April 01, 2009 | 2009 | Roche Diagnostics Corp. | Some data alarm flags do not transfer from the analyzer to the laboratory information system, since not all data flags are recognized by the instrument manager/middleware. | Component design/selection | Consignees were notified by "Urgent Medical Device Correction" letter dated April 1, 2009 and instructed to add error codes to the Instrument Manager driver and to contact the manufacturer of their LIS for instructions on adding data alarm codes to the LIS. Consignees were also instructed to complete the attached fax form after adding the error codes to their system and fax to 1-888-345-4690. Contact Roche Diagnostics Technical Support at 1-800-428-2336 for questions about information contained in the notification letter. |  | 498 | 498 of all products. | Nationwide Distribution. | Terminated | December 07, 2009 | 250 | 498
Z-1344-2009 | 51334 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific, FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style) |  |  | March 18, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific Corporation is initiating an urgent Field Correction related to its FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style). During the attempted removal of the device, sutures were separated from the catheter and remained in the patient. | Labeling False and Misleading | Consignees were sent a Boston Scientific " Urgent Medical Device Field Correction Notification" letter dated March 18, 2009. The letter was addressed to "The Hospital Field Action Contact". The letter described the problem, product affected and the revised catheter removal instructions. They also requested the return of the Acknowledgement Form. |  | 586276 | 586,276 US; 276,013 OUS total for all products | Worldwide Distribution -- USA including DC, Puerto Rico and Guam and countries of United Arab Emirates Austria, Bosnia and Herzegovina, Bahrain, Switzerland, Germany, Egypt, Spain, Finland, France, Great Britain, Ireland, Italy, Kuwait, Netherland, Pakistan, Portugal, Saudi Arabia, Sweden, Slovak Republic, Tunisia, Turkey, South Africa, Netherlands, Costa Rica, Japan, Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Ecuador, El Salvador, Guatemala, Honduras, Hong Kong, Macau, Malaysia, Mexico New Zealand, Panama, Peru, Philippines, Puerto Rico, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | February 23, 2012 | 1072 | 586276
Z-1640-2009 | 51702 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | In Situ Bender Cutter Kit Wide Sterile, Catalog number 530.521S |  |  | April 03, 2009 | 2009 | Synthes Spine | Sterility may be compromised due to a ineffective package design. | Package design/selection | All affected Synthes DMF Sales Consultants were notified of the Medical Device Recall by an Urgent: Medical Device Recall letter notification which was sent via USPS certified mail on 4/6/09. The letter stated the problem and risk. The firm recommended that patients who have had a surgical procedure in which these products were utilized be monitored for signs/symptoms of infection for a period of six months following surgery. If customers have any of the identified devices, they are to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the form; and return the letter with the recalled product to the address provided. Customers that do not have the identified product, are to complete the attached form and return the document to the address provided acknowledging their receipt of the recall information. Questions should be directed to their Synthes CMF Sales Consultant. |  | 2706 | 2706 units total | Worldwide Distribution -- USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the country of Canada. | Terminated | September 15, 2010 | 530 | 2706
Z-1642-2009 | 51868 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration Modular Cone Bodies; titanium Alloy and Hydroxyapatite, Sterile. Catalog number 6276-1-021, Catalog number 6276-1-025, Catalog number 6276-1-219, and Catalog number 6276-1-325. |  |  | April 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Packaging burst test values for the Restoration Modular Cone Body family were found to be incorrectly transferred from the finalized validation to the work specification. Testing showed that burst test results for recalled products fell below the validated parameters. | Process control | Urgent Product Recall letters were sent via Federal Express on April 2, 2009 to Stryker branches, Chief of Orthopaedics, Hospital Rick Management and Surgeons. The letter stated the issue, the potential hazards, and the risk mitigation. Customers were asked to examine their inventory and hospital locations to identify the affected products; retrieve all effected products and return it to their branch or agency warehouse for reconciliation; reconcile all products on the attached Product Accountability Form and fax a copy to the number listed; return the affected product via the Inter-Org Transfer and ship it to the attention of the individual listed; attach the fluorescent orange PRODUCT REMEDIATION sticker to their return, indicating the Product Remediation #, and mark the box with the words "PRODUCT RECALL". Questions should be directed to Colleen O'Meara at 201-831-5970. |  | 17 | 17 units | Nationwide Distribution | Terminated | August 27, 2012 | 1243 | 17
Z-1396-2009 | 48692 | No | DDR | Myoglobin immunological test system. | Immunology | Immunology | 510(k) | Access Myoglobin Reagents Kits, Part Number: 973243. |  |  | September 27, 2007 | 2007 | Beckman Coulter Inc | Beckman Coulter initiated the field action as a result of recent internal testing following the CLSI EP-5A - Evaluation of Precision Performance of Clinical Chemistry Devices protocol, it was determined that the total imprecision for the Access Myoglobin Reagent is greater than or equal to 10% across the expected physiological range rather than the stated claim less than 8%. Based on internal tes | Labeling False and Misleading | An Urgent: Product Corrective Action (PCA) letter was mailed via US Postal Service on October 08, 2007 to the customers that received Access Myoglobin Reagent kits informing them about the issue. The letter also informed customers that the product insert in the Access Assay Manual will be revised with the new range to reflect the new imprecision claim. Customers were instructed to:(1) Share this information with their laboratory staff and retain the letter for theirQuality System documentation,(2) Complete and return the enclosed fax-back response form within 10 days so that Beckman Coulter, Inc. can be assured that the customer received the notification, and(3) Contact Customer Technical Support at 1-800-854-3633 or your local Beckman Coulter, Inc. representative for questions or additional information regarding the notification. |  | 20727 | 20,727 | Nationwide and Canada. | Terminated | May 27, 2009 | 608 | 20727
Z-1398-2009 | 51207 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cell-Dyn Ruby Hematology Analyzer with System Software Version 2.0ML, List Number 08H67-01, Manufactured by Abbott Diagnostics Division, Santa Clara, CA. |  |  | February 12, 2009 | 2009 | Abbott Laboratories | Burn Risk-- when power is cycled to the instrument, there is a possibility that the Mixer Motor Assembly may reach 100 degrees Celsius (212 degrees Fahrenheit ). If the customer touches the motor, burns may occur. | Device Design | A "Product Correction" letter dated February 12, 2009 and "Customer Reply" form were sent to all currently active CELL-DYN Ruby customers using Software Version 2.0 ML. The letter described the affected product, necessary actions and contact information for customers. Customers were instructed to acknowledge receipt of the "Product Correction" letter by completing the enclosed Customer Reply form and return via fax to 1-800-777-0051. A product information letter is being added to the accessory kits shipped with new instruments and version 2.0 ML System Software upgrade kits. For questions regarding this information, contact Abbott Laboratories Customer Support by calling 1-877-422-2688. Customers outside the US, please contact your local hematology customer support representative. |  | 324 | 324 units distributed | Worldwide Distribution -- US, Argentina, Canada, Columbia, Germany, Jamaica, Japan, Mexico, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | December 27, 2010 | 683 | 324
Z-1256-2009 | 50695 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | HemoSense INRatio2 Prothrombin Time Monitor, manufactured by HemoSense Incorporated. |  |  | December 23, 2008 | 2008 | HemoSense Inc | INRatio Monitors that were manufactured with a specific EEPROM memory chip fail during use. The monitors report an error code and do not report a test result. | Nonconforming Material/Component | Letters for each product titled "Urgent: Medical Device Recall Notice" dated December 19, 2008 were issued explaining the problem and instructing users to discontinue use, complete the provided reply form and return the device by mail to Inverness Medical. Distributors of the violative device were instructed to contact their customers and provide the same recall information.Further questions should be addressed to HemoSense Technical Service at 1-877-441-7440. |  | 921 | 921 units | Worldwide Distribution in the US, Germany, Spain and Australia. | Terminated | March 25, 2011 | 822 | 921
Z-1265-2009 | 51424 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex CB-500, 3D Dental Imaging Systems, Model Number G1-15-1-0. This product is manufactured by Imaging Sciences International for Gendex Corporation. |  |  | January 08, 2009 | 2009 | Imaging Sciences International, LLC | faulty calibration | Nonconforming Material/Component | Imaging Sciences International issued a letter to Customers dated January 7, 2009 informing them of the affected device. Users were advised to discontinue use of the device until they are contacted by a service representative and a certified re-calibration of the device is performed. Users were also asked to return a completed "Acknowledgement Form" by fax to the firm at 1-215-997-5665.Additional questions can be made to Imaging Sciences International by calling 1-800-205-3570 Extension 299. |  | 24 | 24 units | Worldwide Distribution -- US to states of FL, OH and PA and countries of Canada, Germany, England, Italy and Switzerland. | Terminated | June 04, 2009 | 147 | 24
Z-1638-2009 | 51898 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgidyne S-Vac round fluted channel silicone drain, 1/8" trocar, 10 FR, round, full fluted, 28 cm length, 4 channel, contains 10, sterile, Sterion Inc., Ham Lake, MN; REF 332187. |  |  | April 13, 2009 | 2009 | Aspen Surgical Products, Inc. | An adaptor, which may be required when connecting the drain to an external device, was not included in the package. | Process change control | An "Urgent Medical Device recall" letter dated April 13, 2009 was issued to consignees via certified mail. The letter described the affected product. and actions for customers. Consignees were instructed to carry out the recall to the retail level and to quarantine and return the product to Aspen Surgical Products, Inc. Consignees were also instructed to complete the enclosed response form and return it via mail, fax or e-mail. Direct questions about the recall by calling Aspen Surgical Products, Inc. at 1-616-698-7100, extension 142 or extension 141. |  | 9 | 9 cases of both products | Nationwide Distribution -- California, Florida, Georgia, Massachusetts and Missouri. | Terminated | November 23, 2009 | 224 | 9
Z-1399-2009 | 50849 | No | DGC | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Immulite 2000 and Immulite 2500 digital Fluidics PC Board, |  |  | January 05, 2009 | 2009 | Siemens Medical Solutions Diagnostics | Design Flaw: The redesign of the Digital Fluidics PCB used in IMMULITE 2000 and IMMULITE 2500 is missing a jumper. | Process design | Siemens Healthcare Diagnostics Inc. issued an "Urgent: Device Recall" notice dated January 2009 mailed to all direct customers of the affected product. The letter described the problem, provided "work-around" instructions and also provided a Recall Response Form to be returned to the firm via fax. For further information, please contact your local Siemens Healthcare support representative or Siemens Healthcare Diagnostics Inc. directly at 1-973-927-2828. |  | 295 | 295 Digital Fludics PCB (88 IMMULITE 2000 and 2500 and 207 as spare parts). | Worldwide Distribution to states including MA, NY, PA, MD, GA, FL, OH, IN, MI, IL, MO, LA, UT, AZ and CA and countries of Canada, Czech Republic, Germany, Poland, Hungary, Belgium, Russia, Republic of South Africa, Greece, Israel, Turkey, Spain, Portugal, Denmark, France, Great Britain, India, China, South Korea and Mexico. | Terminated | April 19, 2010 | 469 | 295
Z-1267-2009 | 51284 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Argon Medical Devices, ArgoGuide Hydrophilic NiTi Guidewire, Catalog Number: AG087301, .035/150cm, regular/straight/standard, sold as a single hydrophilic guidewire in bulk non-sterile packaging. Argon Medical Devices, Inc. Athens, TX 75751. |  |  | February 26, 2009 | 2009 | Argon Medical Devices, Inc | Hydrophilic guidewires have the potential to demonstrate degradation of the Pebax cladding. | Other | Argon Medical Devices Inc. issued an "Urgent Field Safety Notice" letter dated February 26, 2009 informing all consignees of the affected product. Consumers were asked to immediately quarantine and return any product remaining followed by completion and return of the provided response form via email or fax (1-903-677-9396).For further questions or additional information call Argon Medical Devices Inc. at 1-903-677-9319. |  | 3300 | 3,300 units | Worldwide Distribution -- US including Puerto Rico and states of AL, CA, CO, FL, GA, IL, KS, LA, MA, MI, MN, MO, MT, NC, NJ, NV, NY, PA, SC, TN, TX, UT, VA and WI and countries of Germany, Thailand, Turkey and the Netherlands. | Terminated | January 07, 2010 | 315 | 43940
Z-1643-2009 | 51964 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon 25 GA Total Plus Pak, Catalog/Item #8065750220; for use with ACCURUS Vitrectomy System; distributed by Alcon Laboratories, Inc., Fort Worth, TX 76134. |  |  | April 29, 2009 | 2009 | Alcon Laboratories, Inc | Specific lots of Alcon 25 GA Trocar Blades do not conform to finished product specifications in that oxidation can occur on some of the blades. | Process control | Firm began notifying consignees of recall by Urgent: Medical Device Recall letter on 04/29/09 (letters sent via FedEx). Consignees instructed to check inventory for affected lots; to complete and return the Medical Device Recall Response Form; to contact Alcon to arrange for return of the devices; that credit will be issued for the recalled product when it is returned; and in the case that there is no inventory remaining, to indicate that on the Response Form and return it as well. Consignees should contact Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8 for technical assistance. |  | 3523 | 3,523 surgical paks. | Worldwide Distribution -- USA and Canada. | Terminated | October 14, 2009 | 168 | 3582
Z-1646-2009 | 51695 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Roche ACCU-CHEK Spirit insulin pump. Part numbers 4540492001, 4759702001 and 4759729001. |  |  | April 30, 2009 | 2009 | Disetronic Medical Systems, Inc. | The "up" and "down" buttons, which are used for changing the program in the menu or to administer additional insulin through a bolus delivery, may experience intermittent failure or completely cease to operate. | Device Design | The firm issued a press release on 4/30/09 and sent letters with attached Urgent Medical Device Recall (Field Corrective Action) notification dated 5/4/09 to distributors, health care providers and users. Disetronic will replace pumps within the affected serial numbers (SN02119552 to SN10006093) that have intermittent or complete loss of function of the up and/or down button promptly upon contact from the customer with a new pump that is not affected by the potential button failure. For pumps within the affected serial numbers (SN02119552 to SN10006093) which are functioning properly, no replacement is available/possible at this time. Disetronic will gradually phase in replacement pumps for the remaining pumps within the affected range, as they become available. Customers are to complete the enclosed Customer Reply Card to acknowledge receipt of the notification.On 5/17/11, after being notified of the recall's ineffectiveness (in that only 40% of the pumps had been replaced), the firm re-issued recall notification letters including the aforementioned information to their consignees. |  | 14929 | 14,929 | Worldwide Distribution -- USA and Canada. | Terminated | June 05, 2012 | 1132 | 14929
Z-1418-2009 | 51994 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Table Headrest, Catalog/Part #0788-300-012, component of VERTIER Surgical Table. Distributed as part of the VERTIER Surgical Tables or individually by Stryker Communications, Flower Mound, TX 75028. |  |  | May 08, 2009 | 2009 | Stryker Communications Corp | Potential for health care provider/hospital staff member to sustain injury to fingers or hands when removing/handling the headrest component of the VERTIER Surgical Table. | Device Design | An "URGENT: Device Recall" letter dated May 11, 2009 was issued to consignees. Consignees were informed that Stryker service representatives will visit each location to affix a caution label identifying the pinch point on each headrest. Additionally, consignees will be provided with an updated surgical table user manual. Direct questions to Stryker Communications Corporation by calling 972-410-7100. |  | 230 | 230 headrests. | Nationwide Distribution and Canada. | Terminated | April 26, 2010 | 353 | 230
Z-1645-2009 | 51959 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Kimberly-Clark, Synthetic Powder Free Vinyl Exam Gloves, Small, Medium and Large, 100 gloves per box, Non-Sterile, For Medical Use, Catalogue Numbers: 55031, 55032, and 55033. Distributed by Kimberly Clark LLC, Roswell, GA |  |  | April 30, 2009 | 2009 | Kimberly-Clark Corporation | The product may contain defects/holes reducing the product's effectiveness as a barrier. | Release of Material/Component prior to receiving test results | Consignees were contacted by an Urgent: Medical Recall letter on/about 4/30/2009. The letter instructs the Distributor to extend the recall to the end-user level by notifying each of their customers of the recall via a Customer Recall Letter. Both the Distributor Letter and Customer Letter include Response Sheets that allows for both Distributor and Customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. Distributors were instructed to cease further shipment of the product and to physically quarantine the product and fax the completed Response Sheet within 5 business days of receipt of the letter. The customer was instructed to immediately cease further use of the gloves and physically quarantine the product and fax the completed Response Sheet within 5 business days of receipt of the letter. Upon receipt of the Response Sheets, KC will arrange pick-up of the affected products and deliver it to a quarantine location at the contracted warehouse. |  | 9200 | 920 cases/10 boxes per case | Nationwide Distribution -- USA including states of CT, MA, MD, NJ, and NY. | Terminated | January 07, 2010 | 252 | 9200
Z-1417-2009 | 51962 | No | GDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CHIBA NEEDLE MARKED, Model CNM-1806, Size: 18 GA x 15 CM, Qty: 10 each, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | April 06, 2009 | 2009 | Remington Medical Inc. | Incorrect needle (Franseen Needle, FNM-1806) was packaged as a Chiba Needle Marked. | Other | The distributor was notified on Monday, April 6, 2009 via telephone by the QA Manager and a return authorization was issued. A notification letter dated April 13, 2009 was issued via certified mail to the affected customers. The letter described the affected product, reason for recall and actions for customers. Customers were also instructed to conduct recall to user level. For questions or additional information, contact Remington Medical, Inc. by calling 800-989-0057, extension 123. |  | 60 | 60 units | Nationwide Distribution -- KY and NC. | Terminated | May 27, 2009 | 51 | 60
Z-1392-2009 | 51254 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Reaction Vessel, 2000 units per carton; for in vitro diagnostic use with the ARCHITECT i System; Abbott Laboratories, Abbott Park, IL 60064; list 7C15-01. |  |  | February 26, 2009 | 2009 | Abbott Laboratories | Abbott has had an increase in complaints due to the receipt of error code 1006 (unable to process test, background read failure), error code 1007 (unable to process test, activated read failure) and incorrect results when using these lots of ARCHITECT Reaction Vessels. | Nonconforming Material/Component | Consignees were contacted by phone and sent an "Urgent Field Safety Notice Product Recall" dated February 26, 2009. The letter informed users of the problem and detailed how to identify and discard the affected product. In addition, Abbott Laboratories asked the consignees to send a copy of the letter to any laboratory to which they may have forwarded any of affected lots. Abbott sent a second letter dated May 12, 2009 informing consignees of a recall expansion of additional affected lots. The recall was again expanded on 12/17/09 for an addtional 40 lots of reaction vessels.For additional information, contact Abbott Labs at 1-877-4ABBOTT. |  | 27163 | 27,163 cartons | Nationwide including Puerto Rico, and internationally to Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Chile, Argentina, Dominican Republic, Honduras, Nicaragua, Germany, Japan, Hong Kong, Taiwan, Thailand, South Korea, Uruguay, Austria, Bahrain, Belarus, Belgium, Bosnia &amp; Herzegovia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Greece, Hungary, India, Indonesia, Ireland, Italy, Jamaica, Kuwait, Luxembourg, Malaysia, Netherlands, Norway,Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Zimbabwe and Australia. | Terminated | November 04, 2010 | 616 | 27163
Z-1323-2009 | 51668 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Lactosorb Pectus Stabilizer, model number: 01-3805. Distributed and Manufactured by : Biomet Microfixation, Jacksonville, FL. |  |  | March 20, 2009 | 2009 | Biomet Microfixation, Inc. | This recall is based on the occurrence of known adverse effects identified in the device instructions for use. Bending, fracture, loosening, rubbing and migration of the devices can occur as a result of excessive activity, trauma or load bearing.Implantation of foreign materials can result in an inflammatory response or allergic reaction. | Device Design | An "Urgent Medical Device Recall Notice" was issued via certified letter or email on March 20, 2009 to consignees. Customers were instructed to immediately remove and return the affected in-stock product to Biomet Microfixation, Inc. Customers were also asked to complete and return the enclosed Inventory Reconciliation sheet via fax (904-741-9425) to the firm, and to notify recipients of any recalled products that had been further distributed.For questions about the recall, contact Biomet Microfixation at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm EST. |  | 2117 | 2,117 | Worldwide Distribution -- US (including DC and states of AL, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, LA, MI, MO, MN, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA and WI) and countries of Argentina, Austria, Belgium, Chile, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guatemala, Israel, Italy, Lativa, Netherlands, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | September 11, 2009 | 175 | 2117
Z-1421-2009 | 51703 | No | OBP | Repetitive transcranial magnetic stimulation system. | Neurology | Physical Medicine | 510(k) | NeuroStar TMS Therapy System/SenStar Treatment Link. |  |  | March 16, 2009 | 2009 | Neuronetics, Inc. | SenStar Treatment Link may have an open circuit that during use with the NeuroStar TMS Therapy system may cause arcing and heat damage to the SenStar and/or treatment coil. Heating of the SenStar during use has a potential risk of minor scalp burn to patients. | Process control | Neuronetics Inc. notified consignees by telephone and sent a follow-up fax/letter dated March 19, 2009 titled "Urgent Product Recall Notification." The letter describes the issue and actions for consumers to take in response to the recall. Neuronetics will visit each customer in person to complete a "Returned Goods: SenStar Recall Inventory Form." Customer questions should be directed to Neuronetics Customer Service at 1-877-600-7555. |  | 3123 | 3123 Units | Nationwide Distribution to AZ, CA, DC, GA, FL, IL, MD, MI, MN, NH, NY, PA, RI, SC, and WA. | Terminated | May 29, 2009 | 74 | 3123
Z-1401-2009 | 51863 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network with ambulatory telepacks, and software version 8.5.5, 8.5.6, 8.6.1 or 8.7. |  |  | April 10, 2009 | 2009 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Software Anomaly: Software anomaly can occur when the Panorama is attempting to play an audio file, either a physiological or technical alarm sound, at which time the system software may cause the system to freeze due to an interaction in the alarm system. | Software design | Urgent Product Field Correction letters, dated March 27, 2009, were sent by certified mail, return requested, and signature required. The letters identified the affected product and the problem. It also stated the adverse effect on patients and that customers' Panorama software will be upgraded to a current software revision. A Datascope Service Representative will contact customers to arrange for the software upgrade of the Panorama(s) in their facility. This will be done at no cost to the customer. |  | 294 | 294 sites | Worldwide Distribution -- USA and Australia. | Terminated | August 03, 2010 | 480 | 294
Z-1436-2009 | 51982 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | CareVue Chart Release C.O |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Software: A software defect under certain circumstances could lead to the unintended administration of a duplicate dose of a scheduled medication within a single 24 hour period | Software design | Philips Medical notified accounts by letter on 5/6/09 the Urgent - Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within CareVue Chart to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. Philips is currently developing a correction that will be available at no additional charge. A Philips representative will notify customers when the software correction becomes available. If further assistance is needed, customers are to contact the Philips Customer Care Center at 1-800-722-9377. |  | 103 | 103 units | Worldwide Distribution -- USA including states of CA, FL, MA, OH, OK, RI, VT, and WV and countries of Canada, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | March 06, 2012 | 1036 | 103
Z-1224-2009 | 44848 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Central Control Monitor. Terumo Cardiovascular Systems, Ann Arbor, MI; Catalog number 802100. |  |  | January 12, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The monitor may exhibit start-up failure, video/display issues, calibration issues, and partial or complete loss of functionality during use. | Device Design | Consignees were notified via an Urgent Medical Device Correction: Safety Advisory letter dated 1/12/09 advised of the reason for the correction, potential problems, and provided clinical information should the problem(s) arise. Customers are to review the notice, assure that all users are aware of the notice, and confirm receipt of the communication by completing the attached Customer Response Form and faxing it to the number indicated on the form. Questions or concerns should be directed to Terumo CVS Customer Service at 1-800-521-2818. |  | 951 | 951 | Worldwide Distribution -- USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates, and Viet Nam. | Terminated | July 14, 2009 | 183 | 951
Z-1225-2009 | 51220 | No | DYE | Replacement heart valve. | Cardiovascular | Cardiovascular | PMA | Medtronic Mosaic Porcine Bioprostheses with Cinch, Heart Valve; Model 305; Catalog #s 305C21AA, 305C2701, 305C2301, 305C2501, and 305C2901; Sterile LC; Rx only; Medtronic Inc 710 Medtronic Parkway, Minneapolis, MN 55432-5604, Manufactured at: Santa Ana, Ca. |  |  | January 26, 2009 | 2009 | Medtronic Inc | Medtronic identified several heart valves (Mosaic Porcine Bioprosthesis with Cinch &amp; Hancock II Bioprosthesis with Cinch) that were inadvertently distributed after exposure to temperatures exceeding the 77 F (25 C) maximum storage temperature published on the labeling. The product is being retrieved as they do not meet the handling requirements. | Storage | A Medtronic "Urgent Medical Device Recall Notice" dated January 26, 2009, was sent to Risk Managers of each affected account and to each Implanting Physician. The letter described the problem, product and informed consignees that they will be retrieving the non-implanted vales. If additional information is needed, customers are to contact their local Medtronic sales representative or technical services at 1-877-526-7890. |  | 5 | 5 | Nationwide Distribution -- Including states of CT, FL, MD, MA, MI, NC, OH, OR, PA, SD, TN, TX, WI, and WA. | Terminated | May 22, 2012 | 1212 | 15
Z-1324-2009 | 51777 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Square Nut MR Safe, Catalog Number 03.311.060. |  |  | February 09, 2009 | 2009 | Synthes Spine | Numbers are etched incorrectly on the product. | Process control | An "Urgent: Medical Device Recall" letter dated February 6, 2009 was sent to all affected Synthes Trauma Sales Consultants via USPS certified mail on February 09, 2009 and affected In House employees were notified via return receipt email on February 09, 2009. All affected hospitals were notified of the recall through a notification letter mailed via USPS certified mail on February 09, 2009. The notification letter described the affected product, issue and actions for recipients. Direct questions about the recall to Synthes Spine by calling 800-620-7025, Extension 5455 or 610-719-5455. |  | 1137 | 1,137 units | Worldwide Distribution -- US (states of AR, CA, CO, CT, FL, IA, IL, LA, MD, ME, MI, MN, MS, NJ, NY, OK, PA, SC, TX, and UT) and country of Switzerland. | Terminated | December 15, 2009 | 309 | 1137
Z-1249-2009 | 51422 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes InSitu Bender Cutter Kit Narrow Sterile Catalog number 530.522S, used with Synthes Rapid Resorbable Fixation System for craniofacial skeleton fracture repair and reconstructive procedures. |  |  | December 18, 2008 | 2008 | Synthes Spine | The bender is heating prior to activation of the power button, requiring disassembly to stop heating. | Process control | All affected Synthes CMF Sales Consultants were notified of the Medical Device Recall by phone on 12/19/08 and by letter sent via return receipt email on 12/22/08. All affected hospitals were notified by letter mailed on 12/22/08 via USPS certified mail. |  | 40 | 40 units | Nationwide distribution: AL, CA, KS, LA, MD, NJ, NV, NY, OR, PA, TX, and UT | Terminated | December 15, 2009 | 362 | 132
Z-1227-2009 | 51166 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare ApexPro Telemetry System for mobile patient physiological monitoring. |  |  | February 24, 2009 | 2009 | GE Healthcare | Intermittent or continuous loss of patient monitoring. | Nonconforming Material/Component | Consignees were sent on 2/20/09 a GE Healthcare "Urgent Medical Device Correction" letter dated 2/12/09. The letter was addressed to Risk Manger/Chief of Nursing/Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 291 | 291 | Nationwide distribution and Canada. | Terminated | May 29, 2012 | 1190 | 291
Z-1318-2009 | 51449 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Mammo Test Model number 10144185, x-ray guided stereotactic biopsy system |  |  | January 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Table may unintentionally lift during procedure | Nonconforming Material/Component | Siemens Medical Solutions issued a Safety Advisory Notice dated 1/22/09 which informs the customers of the potential issues and recommends that they stop using their MammoTest System until the prone table has been repaired. Linear Medical visited the customers to repair the weld joint on the prone table by January 23, 2009. |  | 4 | 4 units | USA distribution to IL, NY, SC, and WA | Terminated | June 01, 2009 | 131 | 4
Z-1352-2009 | 51693 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Dual Offset Broach Handle, Clamping Style, Left, REF 31-555403. For use in primary hip arthroplasty. |  |  | March 31, 2009 | 2009 | Biomet, Inc. | The handle body may fracture or crack at the slot where the locking lever is inserted to clamp down onto the broach. | Device Design | Biomet notified their distributors by recall letter dated 3/31/09, instructed them to pick up the insturments at the hospitals and to notify the hospital personnel of this recall action in writing. |  | 178 | 178 of all products. | Nationwide, Canada and Finland. | Terminated | December 23, 2009 | 267 | 178
Z-1228-2009 | 51495 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit, catalog number DSK 130-Q, Angiographic Syringe for delivery of contrast media. |  |  | January 29, 2009 | 2009 | Medrad Inc | Possible compromised sterility due to holes in styrene tray | Packaging process control | On 1/29/09, Medrad issued an Urgent Medical Device Recall letter. Customers were requested to check their inventory for any of the affected batch. If the customer has affected product, they are instructed to contact the firm to make arrangement for the return of the product. Non responders were sent additional notifications via FedEx on 3/2/09 and 4/1/09. |  | 2958 | 2958 kits | Nationwide: AL, AZ, CA, FL, IL, LA, MA, MT, NC, NY, OH, TX, and WI. | Terminated | June 04, 2009 | 126 | 2958
Z-1264-2009 | 51438 | No | OEX | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Spectranetics SLS II Laser Sheath Kit, 14F, REF 500-012, Individually packaged catheter introducer. |  |  | March 16, 2009 | 2009 | Spectranetics Corporation | Catheter outer jacket may split during use. | Process control | Spectranetics notified the 8 recipients of the catheters by letter on 3/17/2009. They were told that a Spectranetics sales rep would contact them to coordinate the return of the suspect devices. Questions were referred to 719-447-2462 or 719-447-2539. |  | 17 | 17 units | Nationwide distribution: AL, KY, LA, MI, NC, NY, and PA. | Terminated | November 28, 2011 | 987 | 17
Z-1351-2009 | 51409 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker, Gamma3, The Compact Version of the Gamma Nail System, Operative Technique for Hip Fracture System, Trochanteric and Long Nails. |  |  | March 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails. | Labeling design | Stryker sent Urgent Product Recall letters by Federal Express on March 2, 2009 to Stryker Branches, Chief of Orthopaedics, Hospital Risk Management and Surgeons. Letters were sent to all who have used the Gamma3 products in the past. |  | 18232 | 18232 | Nationwide. | Terminated | August 17, 2012 | 1264 | 38964
Z-1390-2009 | 51673 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor, M3535A/M3536A |  |  | April 10, 2009 | 2009 | Philips Healthcare Inc. | Spontaneous turn-on which could deplete the battery, rendering the device unusable until power is restored. There is also the possibility of a failure mode in which the device fails to respond to user initiated turn-on, rendering it unusable for monitoring and therapy. | Nonconforming Material/Component | Philips notified customers via Urgent Medical Device Recall Letter &amp; Notice dated March 2009. Customers are asked to follow the "Actions to be taken by customer/user section of theUrgent Medical Device Recall Notice. In this section, customers are given the following instructions:"During the interim period, as you await the upgrade for your device, you may continue to use the Heartstart MRx. However, to reduce the risk that a spontaneous turn-on could deplete the battery, it is advised that you keep:a charged spare battery with the Heartstart MRx AC power applied, if possible, when the device is not in useYou should also consider identifying a readily available backup defibrillator to use in the unlikely event that the device fails to turn on. Devices that exhibit a spontaneous turn-on or failure to respond to user initiated turn on should beremoved from service as soon as possible." Philips will replace the therapy switch in affected units free of charge per Mandatory Field Change OrderFC086100077. All customers will be contacted by a Philips Healthcare representative to arrange for replacement of the therapy switch in the affected devices. |  | 22176 | 12,385 units US; 9,791 units OUS | Worldwide Distribution -- USA, Canada, Afghanistan, Argentina, Aruba, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bosnia &amp; Herzegovina, Brazil, Bulgaria, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Oman, Panama, Papua, New Guinea, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudia Arabia, Serbia &amp; Montenegro, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 22176
Z-1435-2009 | 52049 | No | GZI | External functional neuromuscular stimulator. | Neurology | Neurology | 510(k) | WalkAide Stimulator System Control Module, Part Number 20-0100; component of the WalkAide Stimulator System; distributed by Innovative Neurotronics, Austin, TX 78746. |  |  | August 06, 2008 | 2008 | Innovative Neurotronics, Inc. | Knob responsible for regulating the amount of stimulation administered by the WalkAide System control unit may fall off while device is in use. | Nonconforming Material/Component | Innovative Neurotronics Inc notified the consignee via e-mail on August 6, 2008 of the affected product. The Consignee was asked to place a stop shipment on the WalkAides and to return inventory for replacement.Further questions can be directed to Innovative Neurotronics Inc at 1-512-721-1906. |  | 137 | 137 units | Nationwide Distribution to the state of GA. | Terminated | July 28, 2009 | 356 | 137
Z-1434-2009 | 51869 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ADDitIV Primary IV Set with Universal Spike, Backcheck Valve, two Injection Sites and Spin-Lock Connector, Catalog # V1443. |  |  | March 27, 2009 | 2009 | B. Braun Medical, Inc. | Device defect is occlusion/no flow. | Mixed-up of materials/components | An Urgent - Medical Device Recall notice was issued via US Postal Service Certified Mail with registered return receipt mail or FedEx to all customer/distributors in receipt of affected product on 3/27/09. Direct customers received letters addressed to Materials Management that informed them of this recall and provided instructions for disposition of affected product in their inventory. Distributors received the direct customer recall letter along with a cover letter outlining their responsibilities to execute the recall and notify their customers as requested. Distributors are to immediately discontinue further distribution of the affected product, determine their customers' inventory, and have their customers return the product to them. The attached "Product Removal Acknowledgement" should be used to indicate the specifics of the combined distributor and customer inventories. The Customer Support Department should be contacted to arrange for return. A Customer Support Representative will provide instructions for handling the affected product and assist with replacement product. |  | 15450 | 15,450 units | Worldwide Distribution -- USA and Canada. | Terminated | December 15, 2009 | 263 | 397191
Z-1325-2009 | 51701 | No | NAW |  |  | Radiology | PMA | TheraSphere Administration Set (Part Number K125914-034). |  |  | February 24, 2009 | 2009 | MDS Canada Inc. DBA MDS Nordion | Instructions for the use of the pinch clamp that is an integral part of the TheraSphere Administration Set were not included with the distribution of the device. | Error in labeling | A User Bulletin (TS-2009-1) dated February 23, 2009 was issued to all TheraSphere user sites. The bulletin described the issue, affected units, and recommended actions for users. The bulletin instructed users on the proper use of the pinch clamp. MDS Nordion also updated the package insert with these instructions.For questions about the recall, customers in North America contact MDS Nordion Customer Service at 1-866-363-3330 and customers outside North America call 1-877-474-5999 and request Customer Service. Customers may also e-mail MDS Nordion at TheraSphere@mdsinc.com. |  | 26 | 26 | Worldwide Distribution -- US (states of AZ, CA, CO, FL, IL, IN, KS, MD, MI, NC, RI, OH, OR, PA, WI) and country of Canada. | Terminated | June 08, 2009 | 104 | 26
Z-1348-2009 | 51781 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P StarDrive Screwdriver Shaft, T8, self-retaining, quick coupling, for use in spinal fusion surgery. Catalog # 03.617.902 Technique Guide. |  |  | February 10, 2009 | 2009 | Synthes Spine | Tip breakage, due to inadequate instructions for use. The Technique Guide provided incorrect information related to use of the Torque Limiting Attachment during screw insertion and final tightening. | Labeling design | All affected Synthes Spine Sales Consultants were notified of the Labeling Correction via email on February 10, 2009. They were instructed to review changes with their surgeons. They were supplied with an electronic version of the corrected Technique Guide on March 11, 2009. All affected Synthes Sales Consultants and Customers were notified of this labeling change via distribution of the updated Technique Guide and a notification letter. This packet was sent via USPS certified mail on March 26, 2009. To ensure a broad and accurate distribution of notification, affected consultants and customers are defined as having been shipped Zero-P 3.0 mm Titanium Cervical Spine Locking Screws. These devices are used only with the Zero-P system. |  | 1766 | 1766 units | Distributed Nationwide and Canada. | Terminated | October 05, 2009 | 237 | 1766
Z-1414-2009 | 51877 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Prelude Short Sheath Introducer, 4CM in Length with Marker Tip and 0.038" X 50CM, REF No.: PSS-8F-4-038MT, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | April 14, 2009 | 2009 | Merit Medical Systems, Inc. | External tubing may separate during use, allowing blood loss. | Process control | All of Merit's US sales representatives and distributors were notified by Prelude Short Sheath Introducer - Product Recall letter on 4/14/2009. They were instructed to promptly contact their customers. Customers were instructed to isolate, remove from inventory, and return all affected inventory. A Product Retrieval Form was to be returned to Merit. Further information is available at 801-316-4998 or 801-208-4344. |  | 188 | 188 units | Worldwide Distribution -- USA, including states of CA, CO, GA, IL, MO, OK, and OR and countries of Austria, Belgium, Canada, Denmark, France, Germany, Hong Kong, Japan, Latvia, Lithuania, Netherlands, New Zealand, Spain, Sweden, and UK. | Terminated | February 09, 2010 | 301 | 18972
Z-1510-2009 | 52076 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Medtronic SIGMA 100/200/300 Series IPGs, Single Chamber Pacemaker; models SS103, SS106, SS203, and SS303. |  |  | May 18, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Disease Managment | One or more bond wire pairs will lift or separate from the bonding terminals on the device hybrid. This may present clinically as loss of rate response, premature battery depletion, loss of telemetry, or no output. | Device Design | Medtronic Sales representative hand delivered an Important Patient Safety Information letter, dated May 2009, to physicians beginning 05/18/09. The letter describes the issues, provides a predicted failure rate for the 3 populations of devices, and provides Patient Management Recommendations. The letter recommended that physicians should advise their patients to seek medical attention immediately if they experience symptoms (e.g., fainting or lightheadedness). It was also recommended that physicians should consider device replacement for patients who are both pacemaker dependent and who have been implanted with a device in the affected subsets. Medtronic will offer supplemental device warranty if the device is not already at elective replacement time. The last recommendation was that physicians should continue routine follow up in accordance with standard practice for those patients who are not pacemaker dependent. The letter also provides Physician and Patient Support. Letters will be mailed to patients with registered devices beginning 05/27/09. Questions should be directed to a local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 9434 | 9434 | Worldwide Distribution -- USA, including Washington D.C., Guam, Puerto Rico, and Virgin Islands; Albania, Algeria, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latin America, Latvia, Lebanon, Luxembourg, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, Norway, Oman, Pakistan, Palestinian Territory Occupied, Peru, Poland, Portugal, Qatar, Reunion, Romania, Saudi Arabia, Senegal, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sri Lanka, Sudan, Suriname, Sweden, Switzerland, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vatican City State (Holy See), and Venezuela. | Terminated | March 28, 2012 | 1045 | 85378
Z-1422-2009 | 51681 | Yes | JOQ | External programmable pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | EP-4 Computerized Cardiac Stimulator; SJM Part Numbers: 09-1527-0000 / 09-1527-0002; |  |  | March 23, 2009 | 2009 | St. Jude Medical | Non-compliance with IEC Standard 60601-1 Clause 57.6 relative to fusing of AC input line. The neutral line is not fused. The hot line in fused. Clause 57.6 calls for both lines to be fused. | Device Design | All foreign customers were sent Urgent Product Safety Information letters on March 23, 2009. Letters described units affected, the reason for action, and the actions to take to eliminate risk. It is recommended that the information in the letter be circulated to all users, risk managers and the biomedical engineering department responsible for maintaining and inspecting electrical equipment located within their institution. The letter also stated that a SJM representative will be in touch to perform the EP-4 Stimulator fuse modification. Customers are to complete and return the attachment. |  | 694 | 694 total (US, Canada, and OUS) | Worldwide distribution. | Terminated | January 20, 2010 | 303 | 694
Z-1419-2009 | 51413 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 50 Sterilization System Product Code 10050 |  |  | March 09, 2009 | 2009 | Advanced Sterilization Products | ASP has identified a secondary, lower-occurring cause of oil mist within certain STERRAD NX Systems and STERRAD 50 Systems due to potentially defective oil fill plugs. | Component design/selection | Firm mailed out Urgent Device Correction letters dated March 9, 2009, advising customers/users that on April 23, 2008, Advanced Sterilization Products (ASP) initiated a device correction notification related to certain models of STERRAD Sterilization Systems that have the potential to emit oil mist due to a premature mechanical failure in the oil mist filter and/or filter assemblies. All affected STERRAD Systems have been corrected since the notification was issued.Recently, a secondary, lower-occurring cause of oil mist has been identified within the STERRAD NX System and STERRAD 50 System oil fill plugs. STERRAD NX System and/or STERRAD 50 System may potentially have defective oil fill plugs. Less than 1% of the installed STERRAD NX Systems have been affected. ASP has received 18 complaints in 10 months related to this failure for the STERRAD NX System. No complaints related to this issue for the STERRAD 50 System have been received. As of December 22, 2008, ASP has been actively replacing the plugs in these systems with stainless steel plugs during service calls and will continue to replace the oil fill plugs untilall affected units are corrected. The letter advises that there is a risk of symptoms, such as nausea, shortness of breath, dry throat, burning, watery and blurry eyes, and headache occurring if customers inhale oil mist. If you see mist, haze or smoke in the room or detect an oil odor, personnel should complete the following: - Cancel the cycle if the sterilizer is running and as always, do not use instruments from an incomplete cycle- Leave the room- Avoid working in the room until the mist has cleared- Discontinue use of the STERRAD System until the system is repaired.Any additional questions can be directed to the Clinical and Technical Support staff at 1-888-4613. |  | 4222 | 2,209 units installed in US; 2013 units OUS | Worldwide Distribution -- USA, Germany, Italy, France, UK, Switzerland, Benelex, Uruguay, Venezuela, Argentina, Chile, Columbia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Thailand, Taiwan, Turkey, Middle East, Peru, Canada, Czech Republic, Portugal, South Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India, and Slovenia. | Terminated | March 14, 2012 | 1101 | 4222
Z-1441-2009 | 51947 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mavig Portegra 2 which is an accessory to the Siemens angiographic, fluoroscopic and radiographic imaging systems. Component material number 10281151 Surgery lamp, 115V. |  |  | May 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Improper installation of accessory support arm, which could potentially result in the loosening of the arm over time causing it to fall. | Other | The recalling firm has decided to inspect all of the installed units for this condition and where necessary replace the missing clip. The recalling firm has issued an Update Instructions AX011/09/S dated 03.09, which provides instructions on how to perform a visual inspection of the Mavig Portegra 2 arm for the presence of the safety clip and the installation instructions on how to properly install the clip if it is missing. |  | 427 | 427 units | Nationwide Distribution. | Terminated | September 01, 2009 | 123 | 427
Z-1416-2009 | 52008 | No | HFE | Endometrial brush. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cook Medical Tao Brush I.U.M.C. Endometrial Cytology/Histology Sampler (Tao Brush), 9.0Fr/26 cm, 3.5 cm, sterile, Cook Ob/Gyn, Spencer, IN; REF J-ES-090500-BL. |  |  | April 02, 2009 | 2009 | Cook OB/Gyn, Inc. | There is an ink on the brush which has not been tested for use with the Tao Brush. | Process control | The consignee was notified by an Urgent: Recall Notice letter dated 4/2/09. The customer is to forward the notice to all personnel that need to be aware of it; pull from their stock and isolate any affected product; and complete the form. Upon receipt of the form, Cook Women's Health will issue a Return Goods Authorization number and the customer should return the affected products via FedEx using the Return Goods Authorization number. |  | 15396 | 15,396 | Nationwide Distribution. | Terminated | December 03, 2009 | 245 | 15396
Z-1469-2009 | 50745 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard E LUMINEXX Biliary Stent, Size: 14mm x 80mm, Product Number: ZBM14080. |  |  | December 23, 2008 | 2008 | Bard Peripheral Vascular Inc | This action is being taken because a label for an indicated use was applied to product sizes that are not approved for that use. The incorrect indication for use label stated "Now Approved for Vascular Use". | Employee error | Bard Peripheral Vascular, Inc. (BVP) sent an "Urgent - Product Recall Notice" letter to consignees dated December 22, 2008 outlining the reason for the recall. In the letter, customers were informed to check all inventory locations, stop any further distribution of and return affected product to BPV. Consignees were to complete the enclosed Recall and Effectiveness Check Form and fax to 1-800-440-5376 or send information by phone at 1-800-321-4254 extension 2727 or e-mail at silvia.carillo@crbard.com. |  | 94 | 94 units total for all product codes | Nationwide Distribution including states of AL, AR, CA, CT, DE, FL, GA, IL, IN, KY, MA, MD, MO, MS, NC, ND, NJ, NM, NY, OH, PA, TN, TX, VA and WI. | Terminated | April 02, 2012 | 1196 | 94
Z-1260-2009 | 51699 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Remstar Pro M-Series Heated Humidifier System. |  |  | February 05, 2009 | 2009 | Respironics | Malfunction of the J3 connection between the Heater Plate Main printed circuit assembly (PCA) and the heater plate may have an intermittent connection which may lead to excessive heat build-up on the humidifier PCA. | Nonconforming Material/Component | A Service Communication dated February 5, 2009 was provided to all Respironics customers that are approved to service Respironics devices. The communication described the affected device and actions required.Further questions may be addressed by phone at 1-800-345-6443, and by email at service@respironics.com or clinical@respironics.com. |  | 449585 | 449,585 units | Worldwide Distribution -- US including states of OR, OH, KY, TX, MO, AZ, IN and NV and countries of United Arab Emirates, Bosnia, Belgium, Bulgaria, Bahrain, Canada, Czech Republic, Denmark, Egypt, Spain, Finland, Greece, Croatia, Hungary, Ireland, Israel, Jordan, Kuwait, Lebanon, Latvia, Macedonia, Netherlands, Norway, Oman, Poland, Portugal, Romania, Qatar, Russia, Slovenia, Turkey, Serbia and South Africa. | Terminated | July 05, 2012 | 1246 | 449585
Z-1540-2009 | 51211 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT EG7+ Cartridge; pH, PCO, PO, Na , K, iCa, Hct. |  |  | February 23, 2009 | 2009 | Abbott Point of Care Inc. | Equipment problem: i-STAT cartridges were not properly sealed as the packaging system jammed, causing compromised seals. | Packaging process control | Urgent Recall Notice letters pertaining to the EG7+ Cartridges were sent on February 23, 2009 by first class mail to all customers. The firm found one other affected lot of G3+ cartridges and expanded the recall to include one other lot. Letters for the expansion are dated May 2009. Customers were asked to use the attached form to record all of the box numbers if they have any cartridges from the affected lot. Cartridges from the box numbers listed will be replaced an any remaining unused cartridges from those specific boxes are to be returned as instructed on the form. To receive replacement cartridges, customers are to follow the instructions on the enclosed customer reply form. If any cartridges from the affected lot were forwarded to another facility, a copy of the recall letter needs to be provided to them and the information at the bottom of the customer reply form should be completed. Questions should be directed to 1-800-366-8020, Option 1. |  | 26200 | 1048 boxes of 25 cartridges each. | Worldwide Distribution -- USA, Austria, China, India, Saudi Arabia, South Africa, and Taiwan. | Terminated | July 29, 2009 | 156 | 54600
Z-1452-2009 | 52027 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Esteem Blue with Neu Thera Surgeons Glove; a sterile, single use powder-free synthetic polyisoprene surgical glove, latex-free, used as an underglove when double-gloving; 1 pair of gloves per sterile pouch, 200 pairs per case; REF2D73PB90 - size 9; Cardinal Health, McGaw Park, IL 60085 USA, Made in Thailand. |  |  | April 24, 2009 | 2009 | Cardinal Health | Increase in the number of customer complaints due to cuff tears. | Device Design | Urgent: Product Recall letters dated 4/24/09 were sent to the direct accounts on 4/24/09 via UPS 2-day mail, to the attention of the Director of Materials Management. The accounts were informed of the increased complaints of cuff tears, and were instructed to stop using them and destroy them. The letter included a list of gloves that could be substituted for the recalled gloves. The accounts were requested to call the appropriate customer service group to arrange for credit for the destroyed product, and to complete and return by fax the enclosed acknowledgement form, indicating the number of gloves destroyed. |  | 29697 | 29,697 cases total | Worldwide Distribution -- USA, Australia, Austria, Belgium, France, Germany, Japan, New Zealand, Switzerland, Spain, UK, Netherlands, Norway, Sweden, and Turkey. | Terminated | May 11, 2010 | 382 | 29697
Z-1470-2009 | 51865 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Wiseguide Guide Catheter, Sterilized. |  |  | March 11, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific has initiated a recall of the Wiseguide Guide Catheter. They have received reports of difficulties in connecting the product hubs with the Y-Adaptors and toughy borst connectors. If there is difficulty making such connection, there is a potential for prolongation or delay of the procedure in order to exchange the catheter or connector. There have been 41 complaints in the US as | Process change control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated March 11, 2009. The letter was addressed to Risk Manager/Field Action Contacts describing the problem including instructions for affected device handling and a request to return the enclosed Reply Verification Tracking Form.Further questions can be addressed to Boston Scientific at 763-494-1634. |  | 11308 | U.S. and U.S. territories: 22,728; OUS: 11,308 | Worldwide Distribution -- USA, United Arab Emirates, Dutch Antilles, Argentina, Austria, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Czech Republic, Germany, Algeria, Egypt, Spain, Finland, France, Great Britain, India, Indonesia, Ireland, Iraq, Italy, Jamaica, Japan, Kuwait, Sri Lanka, Libya, Netherlands, Norway, Poland, Puerto Rico, Qatar, Saudi Arabia, Sweden, Singapore, Syria, Turkey and South Africa. | Terminated | January 07, 2012 | 1032 | 11308
Z-1474-2009 | 49524 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100 IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188. |  |  | June 17, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has become aware that a few customers have experienced an unusual occurrence of system lockups on Innova single plane cardiovascular X-ray imaging systems while using the image auto-archiving function. when occurring during X-Ray imaging, the acquisition fails and the user has to rest the system. To date, one patient death has been reported related to this issue. | Process control | GE Healthcare sent a "GE Urgent Medical Device Correction" letter dated May 2008. The letter was addressed to Health Care Professionals describing the problem with the affected devices including instructions to avoid further problems until a GE Healthcare service representative is able to make contact and upgrade the system. Further questions may be addressed by calling 1-800-437-1171. |  | 514 | 514 | Worldwide Distribution - US including Puerto Rico and states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY,NC, ND, OH, OK OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV and WI and countries of ALGERIA, AUSTRALIA, BELARUS, BELGIUM, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, FINLAND, FRANCE, GERMANY, GEORGIA, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL RUSSIAN FEDERATION SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SPAIN, SUDAN, SOUTH AFRICA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, TURKEY UNITED KINGDOM, VENEZUELA and SOUTH AFRICA. | Terminated | October 22, 2014 | 2318 | 15816
Z-1507-2009 | 50706 | No | HAX | Skull tongs for traction. | Neurology | Neurology | 510(k) | ReSolve Halo - Open Back Halo Ring, Model Number: 505300D |  |  | November 09, 2008 | 2008 | Ossur Engineering, Inc | Ossur initiated the recall after their investigation of some complaints found that in some instances the ceramic pins are difficult to screw through the halo ring. | Process change control | Ossur contacted direct consignees and sub-accounts beginning on November 10 , 2008 via a phone call informing them of the recall, and the reason as stated in the public reason for recall. Consignees were asked to check their inventory for any of the identified products and contact them with this information. The firm developed a telephone script for the customer service representatives to use.Ossur then forwarded a notification letter to the affected customers beginning on November 12, 2008 via US Postal Service. In the event that any of these parts have already been applied to a patient, Ossur advised customers to follow their normal follow-up protocols, ensuring that the ceramic pins are properly fixed in the skull.At the time of the phone calls Ossur did not have sterilized open back replacement rings available; however, they offered their ReSolve Closed Back Halo ring as an alternative. The firm estimated that the normal product offering should be available once again in approximately four to six weeks from the date of the phone call. |  | 1 | 1 piece | Worldwide Distribution -- USA including states of CA, GA, FL, MD, MI, MO, &amp; WA and countries of Australia, Cyprus, Germany, Italy, Kuwait, Saudi Arabia, Singapore, South Africa, Spain, United Arab Emirates, and the United Kingdom. | Terminated | April 22, 2010 | 529 | 92
Z-1425-2009 | 51202 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Aisys, Anesthesia Gas-Machine. |  |  | February 05, 2009 | 2009 | GE Healthcare | Unintended shut-down: The anesthesia machine may start shutting down without human intervention due to a faulty On/Standby Switch. Prior to shutdown, the system will issue an alarm and play a message that shutdown will occur in 8 seconds. The failure only affects the electrical portion of the switch. If the clinician does not intervene, the patient could experience hypoventilation leading to h | Component design/selection | GE Healthcare contacted consignees via two "Urgent Medical device Correction" letters dated February 5, 2009 and March 10, 2009. The letters are addressed to 3 titles within the affected accounts; Healthcare Administrator/Risk Manger, Chief of Intensive Care and Director of Biomedical Engineering. The letters describes the Safety Issue, Affected Product details, Safety Instructions, Product Correction and Contact Information. The second letter dated March 10, 2009 includes 700 serial numbers that were missing from the first letter. The second letter was only mailed to the consignees of the additional 700 units. |  | 3082 | 3082 | Worldwide distribution: USA, Canada, Mexico, YEMEN, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TUrKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIA, QATAR, PORTUGAL, POLAND, PERU, PANAMA, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LITHUANIA, LIBYA, LEBANON, LATVIA, KUWAIT, KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLoMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA, ARGENTINA. | Terminated | December 10, 2011 | 1038 | 9325
Z-1389-2009 | 51496 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | February 26, 2009 | 2009 | Welch Allyn Protocol, Inc | Issue #1: AED 10 displayed a "Do Not Use" symbol as a result of a blown F1 fuse.Issue #2: AED 10 models that incorporate a spring connector with the speaker experienced intermittent electrical connection to the circuit board causing a loss of voice prompts. | Component design/selection | Welch Allyn sent an Urgent Medical Device Recall letter dated February 26, 2009 to all distributors and customers, describing 5 failure modes. Customers were instructed to respond to the letter by calling 888-345-5356, and decide on replacement or discounted new purchase. |  | 9754 | Issue #1: 4568 units; Issue#2: 5186 units | Worldwide distribution. | Terminated | March 07, 2012 | 1105 | 9754
Z-1423-2009 | 51716 | No | ISA | Therapeutic massager. | Physical Medicine | Physical Medicine | 510(K) Exempt | Sybaritic, Inc., SlimLine Health Environment Capsule, |  |  | June 18, 2008 | 2008 | Sybaritic, Inc | Misbranded and sale without pre-market notification. | PMA | Sybaritic sent a letter to the consignee on June 18, 2008. The letter confirmed their phone conversation on June 16 regarding the required Field Correction of the device. The letter also described the changes to be performed by the local technician to be in compliance with the FDA. |  | 1 | 1 | UT | Terminated | July 06, 2012 | 1479 | 1
Z-1424-2009 | 51952 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Transfer Spike, Blood Processing Accessory, REF BT945, Sequestration Kits |  |  | February 25, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Sterility may be compromised. Non-vented caps that connect to the 3-way stopcock attached to the transfer spike could prevent Ethylene Oxide from contacting the inside surfaces of the stopcock during sterilization, leading to a potentially non-sterile product. Use of the product may cause serious health consequences from potential sepsis/infection and the associated remote risk of death. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Recall Notice" letter dated March 5, 2009. The letter was addressed to Risk Manager. The letter described the product and problem and requested the consignee to complete, sign and return the certificate acknowledgement. |  | 20 | 20 units | Distribution to one medical facility in MN. | Terminated | May 07, 2011 | 801 | 20
Z-1500-2009 | 51887 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Arrow Balloon and stylet temporary pacing catheter |  |  | February 11, 2009 | 2009 | Arrow International Inc | Leakage: Material not lipid resistant, may crack and/or leak. Use of this device with lipid containing solutions or 70% isopropyl alcohol may cause leakage or air embolism. | Device Design | Arrow International issued a letter dated 2/18/09 to all customers informing them of the problem and if a lipid resistant stopcock is required to contact the firm for replacements. New labels state: Warning - do not use with lipid containing solutions or 70% isopropyl alcohol, as these solutions may cause leakage or air embolism. |  | 900000 | 900,000 total, all products | Worldwide distribution: USA, Argentina, Australia, Bahamas, Belarus, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Guatemala, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Singapore, Slovakia, South Africa, South Korea, Spain, Sri lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UK, Ukraine, and Venezuela. | Terminated | July 22, 2010 | 526 | 900000
Z-1486-2009 | 52098 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific FortÂżÂż Guide Wire, Floppy Marker Wire, 300cm x 2 cm, REF 34949-02, for use in intravascular interventional procedures. |  |  | May 12, 2009 | 2009 | Boston Scientific Corporation | Product packaged incorrectly. One batch/lot of wires, that should contain radiopaque marker bands, is missing the marker bands; and a second batch/lot of wires, which should not contain radiopaque marker bands, has the marker bands. Since the marker bands issue would not be visually apparent to the physician there is a potential for prolongation or delay of the procedure in order to exchange the | Process control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated may 19, 2009. The letter was addressed to "Dear Risk Manger/ Field Action Contact". The letter described the problem, customer steps for recall and listing of product batches / lots. The letter requested consignees to return the Reply Verification Tracking Form. |  | 5 | Total of 5 boxes (5 units /box) , | USA distribution: AR, MO, CA, FL, and NV. | Terminated | December 11, 2011 | 943 | 9
Z-1504-2009 | 51817 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100 31 Gauge, 1cc Insulin Syringe, 5/16" (8mm) Needles, Short Needle, packaged 10 syringes per package, 10 packages per carton, NDC 0603-7002-2 |  |  | April 09, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles separate from the barrel of the syringes. | Nonconforming Material/Component | On 04/09/2009, Qualitest issued a press release and sent out recall notifications to their consignees by US Mail. The recall was extended to the User level. Qualitest will send a second letter to all non-responders after thirty days. Thirty days after the second letter, Qualitest will perform a recall effectiveness check by telephone on a percentage of non-responders (in accordance with FDA request). |  | 559500 | 559,500 syringes | Nationwide. | Terminated | November 01, 2010 | 571 | 559500
Z-1505-2009 | 50773 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions E.CAM Patient Handling System (PHS) Field of View (FOV) indicator bars; a component of the E.CAM Emission Computed Tomography System; part number 43 66 618 (right) and 43 81 088 (left). |  |  | December 18, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Patient fingers can get pinched if they attempt to pull themselves up from the patient handling system without any assistance and grabbing onto the flip-up FOV indicator bars. | Device Design | Siemens sent Urgent Field Correction Recall letters dated 12/18/08 to all users of the E.cam Systems. The accounts were informed of the potential risk to patients when they attempt to pull themselves up from the patient handling system (PHS) without any assistance and grabbing the field of view (FOV) indicator bar (aka flip-up armrests). The customers were advised to always assist the patient when getting on/off of the PHS. They also were advised that a Siemens Customer Service Representative would contact them within 90 days to schedule a corrective upgrade to the PHS with a modified FOV guide to mitigate the risk for patient injury free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. |  | 1107 | 1,107 units | Worldwide distribution: USA, Guam, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Mauritius, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, Turkey and United Arab Emirates. | Terminated | May 03, 2011 | 866 | 1107
Z-1506-2009 | 51871 | No | GTZ | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Inverness Medical BioStar Strep A OIA MAX, 100 test kit size, |  |  | April 06, 2009 | 2009 | Binax, Inc. dba IMPD | Potential for false positive Strep A results. | Other | Inverness Medical notified Customers by letter sent via Fedex April 6, 2009 . The letter included a request for a response from the customer. The product is discontinued. |  | 1729 | 1,729 kits | Nationwide. | Terminated | January 06, 2010 | 275 | 1729
Z-1488-2009 | 52023 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 20 ml Syringe |  |  | April 30, 2009 | 2009 | Becton Dickinson &amp; Company | Potential for compromised sterility: Insufficient packaging seals. | Nonconforming Material/Component | Becton Dickinson sent Urgent Product Recall letters on April 30, 2009 to all direct customers. Questions are to be directed to BD Customer Service at 1-888-237-2762.Letters to customers of distributors went out on May 12, 2009. |  | 201600 | 201,600 units | Nationwide Distribution. | Terminated | September 28, 2009 | 151 | 201600
Z-1523-2009 | 52003 | No | NJL |  |  | Orthopedic | PMA | Zimmer NexGen Complete Knee Solution Legacy Knee- Posterior Stabilized LPS Femoral Component Option Size D Left, sterile, Zimmer, Warsaw, IN; REF 00-5996-014-51. |  |  | April 17, 2009 | 2009 | Zimmer Inc. | The implant surface may not have been polished adequately, which could result in increased wear and polyethylene debris generation. | Employee error | Consignees were notified by phone beginning on 4/17/09 and email and letter notifications were sent on 4/22/09. A dear surgeon letter entitled Urgent: Medical Device Recall, dated April 28, 2009, is being sent to implanting surgeons recommending patient monitoring. Consignees are to return the response form indicating their receipt of the notification. Questions or concerns should be directed to (800) 846-4637 or www.zimmer.com. |  | 10 | 10 | Worldwide Distribution -- USA including states of Arizona, California, Colorado, Hawaii, Illinois, Indiana, Michigan, Minnesota, Mississippi, New Jersey, New Mexico, New York, Ohio, Oregon, and Pennsylvania and country of Canada. | Terminated | November 12, 2009 | 209 | 37
Z-1473-2009 | 51960 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Embolectomy Catheters: 4 French, 2 Lumen x 40cm, Product Number: W01798, Single use, Latex Free, Rx only, Arrow International, 2400 Bernville Road, Reading, PA 19605. |  |  | March 25, 2009 | 2009 | Arrow International Inc | The inflation arm tubing may separate from the stopcock adapter. | Process control | Arrow International issued an "Urgent Medical Device Recall" letter dated March 25, 2009 describing proper handling of the affected device. Consignees were also instructed to complete a Recall Acknowledgement and Stock Status Form and fax to Arrow International at 1-800-343-2935.Further questions should be directed to Arrow International, Customer Service Department at 1-800-523-8446. |  | 6910 | 6910 units (total all three products) | Worldwide Distribution - USA and Panama. | Terminated | February 01, 2010 | 313 | 6910
Z-1477-2009 | 52004 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Natural Knee Flex System, Gender Solutions Male, Femoral Component, nonporous, size 4, left, sterile, Zimmer, Warsaw, IN; Product Number: 00-5410-017-01. |  |  | April 29, 2009 | 2009 | Zimmer Inc. | The tyvek lid, which is intended to assure sterility until the device is transported to the operating room, may be missing from the inner tray. | Employee error | Zimmer, Inc. issued an "Urgent: Device Recall" letter dated April 29, 2009 informing all consignees of the affected product. Users were instructed to locate and quarantine all unused devices and ship them to the firm with a completed Inventory Return Certification form. Return Certification forms are also to be sent via fax to Zimmer, Inc. at 1-574-372-4265. For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 14 | 14 | Worldwide Distribution -- US including states of FL, MN, MS, NY, OH, TN, TX and WI and Switzerland. | Terminated | October 21, 2009 | 175 | 14
Z-1524-2009 | 52031 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Normed TX 8 Driver, Non-Cannulated (used with the Ascension Total Foot System); Catalog #503004268; distributed by Ascension Orthopedics, Inc., Austin, TX 78754. |  |  | April 14, 2009 | 2009 | Ascension Orthopedics, Inc | Screwdriver used with the Ascension Total Foot System not manufactured to specification, resulting in the driver not properly engaging and/or stripping the screws making it difficult to properly seat the screws during surgery. | Pending | Firm notified their distributors of recall via phone starting 04/14/09. Distributors replaced the recalled screwdrivers with ones that were manufactured to specification. |  | 9 | 9 units. | Nationwide Distribution -- including states of CO, OR, and TX. | Terminated | September 24, 2010 | 528 | 9
Z-1478-2009 | 52026 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Fetal Monitor FM20; Model Number: M2702A. |  |  | May 12, 2009 | 2009 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Software change control | Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.Questions can be made to Philips response center at 1-800-722-9377. |  | 5663 | 5,663 units | Worldwide Distribution | Open, Classified |  |  | 10451
Z-1525-2009 | 52093 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia Dual Anterior Staple; |  |  | April 12, 2007 | 2007 | Stryker Spine | Labeling Issue: The large Xia Anterior Staple was laser marked with an "S", indicating small. | Error in labeling | Market Withdrawal letters were sent to Stryker branches who received product on April 13, 2007 by Federal Express. The letters asked that they examine their inventory and hospital locations to identify the implant; reconcile these implants on the attached Product Accountability Form and fax a copy to (201) 760-8456; retrieve and return the affected implants; and contact their Stryker Customer Service rep to re-order the product that they are returning as part of the regulatory action. |  | 38 | 38 in US | Nationwide Distribution. | Terminated | June 29, 2009 | 809 | 38
Z-1481-2009 | 52005 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Steinmann Pins, 7/64 in. (2.8 mm) diameter, 9 in. (229mm) length, style 6, nonsterile, qty 6, Zimmer, Warsaw, IN; Model Number: 00-0187-03-69. |  |  | May 01, 2009 | 2009 | Zimmer Inc. | The package contains pins with an incorrect diameter. | Packaging process control | Zimmer, Inc. issued an "Urgent: Device Recall" letter informing users of the affected product. Users were instructed to locate and quarantine all unused devices and ship them to the firm with a completed Inventory Return Certification form. Return Certification forms are also to be sent via fax to Zimmer, Inc. at 1-574-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 37 | 37 | Worldwide Distribution -- US including states of CA, GA, IL, MA, MO, OR, TX and WI and Japan. | Terminated | November 12, 2009 | 195 | 37
Z-1559-2009 | 52041 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Novametrix Neonatal/Pediatric SpO2 Sensor, Foam Wrap Style |  |  | May 13, 2009 | 2009 | Dixtal Medical, Inc. | Neonatal/Pediatric SPU Sensors may be incorrectly packaged in pouches labeled for Adult/Pediatric SPU Sensors. | Labeling mix-ups | Urgent Product Recall Notification was issued to affected accounts by email/facsmile on May 8, 2009 requesting to cease distribution, quarantine and return recalled product. Contact Dixtal Medical with questions; to inquire about your on-hand inventory; or to arrange to receive replacement product at Customer Service: sales@dixtal.com, Phone Number 1-203-269-1112, Fax Number 1-203-269-1760. |  | 675 | 27 Boxes (Packaged 25 sensors/box) | Worldwide Distribution -- US (TX and CT) and countries of Italy, Japan, France and Turkey. | Terminated | March 23, 2012 | 1045 | 675
Z-1539-2009 | 52046 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Bard StatLock Intravenous (IV) Catheter Stabilization Kit, Catalog number: IVK0124, C.R. Bard, Inc., Covington, GA 30014, USA. |  |  | April 02, 2009 | 2009 | C.R. Bard, Inc., Urological Division | The affected units may contain extension sets with the microbore tubing instead of extension sets with the macrobore tubing. | Process control | All consignees were notified by an Urgent: Product Recall Notification letter via Fed-Ex carrier on 04/02/2009. They were instructed not to use or further distribute any of the affected products. They were advised to check their inventory and to remove and return affected product to Bard Medical Division. A Recall &amp; Effectiveness Check Form was included to be filled out and returned via fax, to Bard. |  | 16100 | 16,100 units | Nationwide Distribution - including states of AZ, CA, CO, GA, IA, MN, NE, and TX. | Terminated | December 04, 2009 | 246 | 16100
Z-1582-2009 | 52214 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT software v2.5.0, Model 728306. The Brilliance iCT is packaged INA carton and the software is contained within the Brilliance CT scanner. |  |  | December 22, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | Software anomaly. An artifact may appear in brain scans and brain perfusion studies, when using 80kV and a UA, UB, or UC filter with this CT scanning device. | Software design | An "URGENT - Field Safety Notice" dated December 22, 2009 was issued via certified mail to affected customers. The letter described the affected product, problem, hazards involved, and actions for customers to take. Philips will provide customers with a software update for all affected systems within next 6 months. The firm also sent a second notification in the form of an "URGENT -Medical Device Correction" report dated January 29, 2009 which contained information noted in the "URGENT - Field Safety Notice". Direct questions to your local Philips representative or Philips Customer Care Center (1-800-722-9377), option 5: Other Imaging Systems, option 1: CT or to your local Philips Healthcare office. |  | 44 | 44 | Worldwide Distribution -- US including DC (states of CA, GA, IL, IN, KY, NY, MD, MN, OH, OR, and MD), Germany, Greece, Israel, Italy, Japan, Netherlands, Norway, Russia, South Korea, Switzerland, Taiwan, Canada, and Thailand. | Terminated | September 24, 2010 | 641 | 44
Z-1531-2009 | 51980 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet OSS Elliptical Proximal Femoral Trial 7 cm left provisional, stainless steel, Biomet Orthopedics, Inc., Warsaw, IN; REF 32-472125. |  |  | May 19, 2009 | 2009 | Biomet, Inc. | The retaining ring may fall out of the instrument during surgery. | Process control | Distributors and consignees were notified by an Urgent Medical Device Recall Notice dated 5/19/09. Distributors and consignees were instructed to locate and remove the devices from circulation, carefully follow the instructions on the "FAX Back Response Form," and fax a copy of the Response Form to 574-372-1683 prior to return of the product. If the product has been further distributed, consignees must notify hospital personnel responsible for receiving recall notices of the action, via the enclosed "Dear Biomet Customer" notice. Consignees are responsible for the location and return of products to the firm. The letter stated that the instruments would be reworked and returned. |  | 387 | 387 of all products | Worldwide Distribution -- USA, Australia, Belgium, Canada, Costa Rica, Finland, France, Germany, Mexico, New Zealand, Poland, South Wales, and Spain. | Terminated | October 08, 2010 | 507 | 387
Z-1247-2009 | 50841 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | Ceiling Tube Mount Ceiling Tube Mount (CTM)/Digital Ceiling Tube Mount (DCTM) with 5 inch or 10 inch Column Extensions; a radiographic x-ray tube suspension system designed to facilitate a wide range of procedures; Del Medical Systems Group, 11550 West King Street, Franklin Park, IL 60131-1330. |  |  | January 12, 2009 | 2009 | Del Medical Systems Group | There is the potential for the supporting cable to become detached from the column causing the column to fall and possibly contact the patient or operator. | Component design/selection | Del Medical sent a "Field Correction Recall" letter with an acknowledgement form to the direct account dealers/assemblers on January 12, 2009 - February 4, 2009, informing them that required maintenance inspection of a retaining ring inside the column assembly may not be performed as required. If the retaining ring is not regularly inspected and replaced when worn or loose, there is the potential for the supporting cable to become detached from the column causing the column to fall and possibly contact the patient or operator. The accounts were requested to locate the affected systems and immediately take them out of service until correction by installing a Cable Mount Retention Retrofit Kit to eliminate the potential for the cable to become detached. Affected accounts were instructed to complete the enclosed acknowledgement form and return it via fax (847-288-7911) or mail to Del Medical. If there are any questions, please contact Del Medical Technical Support at 800-800-6006 or 847-288-7000. |  | 213 | 213 devices | Worldwide Distribution -- US, Canada, Scotland and Pakistan. | Terminated | June 17, 2010 | 521 | 213
Z-1485-2009 | 51406 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural-Knee II System Modular Cemented Tibial Baseplate, Right, Size 0, TIB B/PL, for cemented use only, sterile, Zimmer, Warsaw, IN; Model Number: 6420-01-200. |  |  | March 18, 2009 | 2009 | Zimmer Inc. | The tibial spacer will not affix to the distal surface of the baseplate, as intended. | Nonconforming Material/Component | Zimmer, Inc. issued an "Urgent-Device Recall" letter dated March 18, 2009 instructing users to cease using and return the affected product accompanied with a completed Inventory Return Certification form. The inventory Return Certification form should also be faxed to Zimmer, Inc. at 1-574-372-4265.Further questions should be addressed to Zimmer, Inc. at 1-800-613-6131. |  | 834 | 834 of all products | Worldwide Distribution - USA including states of CA, NY, MI, OH, WI, IA, MN, NJ, IL, MO, TX, LA, NC, SC, PA, AR, MS, FL, GA, AL, OK, KS, TN, CT, NM, UT, NV, ID and CO and Switzerland. | Terminated | October 23, 2009 | 219 | 834
Z-1533-2009 | 51924 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TransFx External Fixation System Drill Bit, Quick Connect, 2.5 mm diameter, 180 mm length, nonsterile, Zimmer, Warsaw, IN; REF 4450-56. TransFx drill bits can be found in the following kits: TransFx Intermediate External Fixation System, REF 4450-05-10; TransFx Large External Fixation System, REF 4450-010-10 and TransFx Small External Fixation System, REF 4450-00050-15. |  |  | May 01, 2009 | 2009 | Zimmer Inc. | Drill tip is out of specifications, reducing the efficiency of drilling action, requiring more force and potentially heating the bone surface during operation. | Process control | Distributors and consignees were notified by an Urgent: Device Recall letter dated 5/1/09. Consignees were advised to stop using the device and quarantine it immediately. Distributors were instructed to carry out a physical count of all affected product and record the data on the Inventory Return Certification Form; fax a copy of the completed form to Zimmer, Inc.; and return recalled product along with the Inventory Return Certification Form to the firm. |  | 1391 | 1391 | Worldwide Distribution -- USA, Australia, Canada, El Salvador, England, Germany, Latin America, and Singapore. | Terminated | November 03, 2010 | 551 | 6085
Z-1556-2009 | 52015 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Signa Ovation 0.35 T MR System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has identified a potential safety issue with the Signa OpenSpeed and Signa Ovation MR Scanner tables that may impact patient safety. The table may begin lowering on its own while a patient is on the table. If this occurs while the table is in the magnet bore, the table may tilt and could result in patient injury. The patient's head or feet depending on patient orientation, may hit | Device Design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 16, 2009 addressed to Hospital Administrators, Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter describes the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A GE Healthcare Field Engineer will schedule a field visit to correct the device.Further questions or concerns may be addressed to GE Healthcare at 1-262-521-6681. |  | 203 | 203 | Worldwide Distribution -- US including PR and states of AL, AZ, AR, CA, CO,CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MN, MS, MO,MT, NV, NH, NE, NJ, NY, NC, OH, OK,OR, PA, SC, SD, TN, TX, VA, WA, WV, WI and countries of BRAZIL CHINA, DENMARK, EGYPT, ENGLAND, GERMANY, GUATEMALA, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, OMAN, MEXICO, OMAN, PHILIPPINES, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES and VENEZUELA. | Terminated | April 23, 2013 | 1468 | 374
Z-1480-2009 | 52009 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Optivac Total Hip Kit, Maximum Capacity Cartridge, 1 kit, sterile, latex free, Biomet Orthopedics, Inc., Warsaw, IN; Product Number: 418000. |  |  | April 09, 2009 | 2009 | Biomet, Inc. | The package may contain two short mixing rods instead of one long and one short rod. | Packaging process control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated April 9, 2009 instructing distributors to visit each hospital consignee, deliver a copy of the recall notification and replacement product and to pick up, and return the affected product.Further questions may be addressed to the Biomet Orthopedics at 1-800-348-9500. |  | 250 | 250 | Nationwide Distribution - Including AL, AR, FL, IL, LA, MN, NE, NC, NV, OH, OK, SC, SD, WA and WI. | Terminated | October 06, 2010 | 545 | 250
Z-1570-2009 | 52124 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Sterrad 50 Sterilizer, Product Code 10050. |  |  | January 12, 2009 | 2009 | Advanced Sterilization Products | The capacitor in the vacuum pump component within the Sterrad 50 System could overheat and potentially create a hazardous condition inside the system, resulting in smoke and/or fire. | Other | Advanced Sterilization Products (ASP) sent an "Urgent: Product Correction" letter dated January 12, 2009 to each consignee informing them of the issue and instructed them to report smoke or a burning odors in or around their STERRAD 50 System. ASP strongly recommended for consignees to follow their facility's safety protocols and if possible, immediately disconnect the STERRAD 50 System from its electrical source. Customers were instructed to contact ASP to schedule an onsite visit to investigate the issue. If the issue is caused by a malfunctioning capacitor in the vacuum pump, the firm will replace the malfunctioning component right away. To correct this situation, an ASP Field Service Engineer (FSE) will visit all customers facilities within thenext sixty days to replace the capacitor in the vacuum pump on the affected STERRAD 50 System. |  | 1151 | 1,151 units. | Worldwide Distribution -- US, Germany, Italy, France, United Kingdom, Switzerland, Benelex, Uruguay, Venezuela, Argentina, Chile, Columbia, Belgium, Puerto Rico, Mexico, Philipines, Australia, Japan, Korea, Malaysia, Taiwan, Middle East, Peru, Canada, Czech Republic, Portugal, Spain, Sweden, Hungary, Poland, Brazil, Russia, Finland, Norway, Panama and Switzerland. | Terminated | December 28, 2011 | 1080 | 1151
Z-1549-2009 | 51501 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | April 03, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There is a potential safety issue associated with the use of GE Centricity PACS software related to configuration of DICOM query retrieval with rejected images. | Software design | GE Healthcare Integrated IT Solutions sent an "Urgent Medical Device Correction" letter dated April 2, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers with attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the patient safety issue associated with the affected device and provided safety instructions to follow until the software is updated. User were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issue and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported.Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution - United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 1148 | 1188
Z-1956-2009 | 50739 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Genicon, RX Only, Sterile EO. Genicon 6869 Stapoint Court, Suite 114 Winter Park, FL U.S.A. 32792. EC REP Genicon Europe, Ltd. 44 Airedale Avenue, Chiswick London U.K. W4 2NW. Catalog Number 700-005-010, 5MM. |  |  | May 08, 2008 | 2008 | Genico, Inc., dba Genicon | Genicon is recalling the Genicon Disposable Suction Irrigation 5mm 700-005-010 and 10mm 700-005-005 Tube Sets. The caps on the trumpet valves of the devices could fall off, resulting in the malfunction of the Suction Irrigation Device. | Nonconforming Material/Component | All four consignees were contacted by email and fax on May 8, 2008. Wholesalers were requested to forward copies of Genicon's recall letter dated May 8, 2009 and response form to their customers. Instructions for product return was provided. Direct questions about the recall to Genicon by calling 407-657-4851 (in United States) or 0118 930 5333 (in United Kingdom). |  | 2790 | 2,790 | Worldwide Distribution -- US (MO and IA), United Kingdom, Netherlands and United Arab Republic. | Terminated | October 01, 2009 | 511 | 3260
Z-1544-2009 | 51484 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrainLab Micro-Multileaf-Collimator, model 41600; BrainLab AG, Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  |  | March 19, 2009 | 2009 | Brainlab AG | The Integrated Multileaf-Collimator (MLC) of the linear accelerator could interfere with the intended conformal fields of the BrainLab m3 micro-Multileaf-Collimator (mMLC) during patient treatment. | Other | BrainLAB AG issued an "Urgent Field Safety Notice/Product Notification" letter dated March 19, 2009 informing all customers of the affected device. Consignees were instructed on proper steps to take to avoid further issues. In addition, the most current m3 mMLC user guide, version 7.4, was scheduled to be sent to users in April 2009.For further questions, contact BrainLAB AG by phone at 1-800-597-5911). |  | 13 | 13 units | Nationwide Distribution - AL, CA, CO, FL, GA, IL, KS, KY, LA, MD, MI, MN, MS, ME, NJ, NY, OR, TN, TX, UT, VA, WV, WA and WI. | Terminated | August 19, 2010 | 518 | 50
Z-1580-2009 | 52176 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TumorLoc Software Application for the EBW (Extended Brilliance Workstation, Model # 455011002031, sold as an accessory for the GEMINI PET/CT Systems running CT software versions 3.5.x. , CT host 2.2. |  |  | February 19, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software anomaly in the TumorLoc software for the firm's Brilliance CT scanners was discovered. The same software anomaly was subsequently determined to be present in the firm's GEMINI PET/CT units. The software defect results in a the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | An "URGENT - Field Safety Notice" dated February 12, 2009 was sent via FedEx to domestic customers. The notice described the affected products, issue, involved hazard should problem occur and corrective actions for the customer/user. The letter also notified users/customers that a Philips Healthcare Field Service Engineer will contact and visit their site within 6 months to install a TumorLoc upgrade. Notification will also be sent to the international consignees by Philips personnel. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 11 | 11 units | Worldwide Distribution -- US, Columbia, Germany, France, Netherlands, Canada, Belgium and Slovenia. | Terminated | September 21, 2012 | 1310 | 27
Z-1568-2009 | 52047 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1. Model #455011202691 for Computed Tomography (CT); Model #455011002031 for Nuclear Medicine (NM). |  |  | April 27, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1 due to the fact that when images are saved or filmed that contain the Curved Multi-planar Reformation Section (cMPRS) views that were created with a short curve the Right/Left annotations will be reversed. | Software design | An "URGENT - Field Safety Notice" dated April 27, 2009 was sent to all consignees via certified mail. The letter described the affected products, issue, and actions for the customer/user. The notification letter provides a work around until the new software is available. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 417 | 417 units | Worldwide Distribution -- US including DC (states of AL, AR, AZ, CO, DC, FL, IL, KY, LA, MA. ME, MI, MN, MO, NJ, NY, OH, OR, PA, SC, TN, TX, VT, and WI) and countries of Africa, Albania, Austria, Australia, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Holland, India, Iraq, Israel, Italy, Japan, Korea, Lithuania, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Russia, Singapore, Spain, South Africa, Sweden, Switzerland, Syria, Tahiti, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Venezuela. | Terminated | September 21, 2012 | 1243 | 417
Z-1615-2009 | 52128 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer femoral stem, revision taper, 15 mm diameter, 185 mm stem length, w/compression nut, sterile, Zimmer, Warsaw, IN; Model Number: 9982-15-18. |  |  | May 15, 2009 | 2009 | Zimmer Inc. | The threads on the screw are incomplete/non-functional. | Process control | Zimmer, Inc. began Consignee notification of the affected device via telephone on May 15, 2009. A certified "Urgent: Device Recall" notification dated May 21, 2009 was sent to all consignees instructing them of the affected device, to stop use of and return it with a Return Certification form. The Return Verification form should also be faxed to Zimmer, Inc. at 1-574-372-4265.For shipping assistance, questions or other concerns, contact Zimmer, Inc. at 1-800-613-6131. |  | 10 | 10 | Worldwide Distribution -- US states of IA, NE and CA and countries of Canada, Switzerland, United Kingdom, Hungary, Netherlands and Slovakia. | Terminated | November 12, 2009 | 181 | 10
Z-1657-2009 | 52077 | No | MGB |  |  | Cardiovascular | PMA | Abbott Vascular, StarClose Vascular Closure System Part Number 14677-01, for closure of femoral artery access sites after cardiovascular catheterization. |  |  | May 06, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Potential sterility issue. Internal component quarantined for possible contamination, was released for sale after going through sterilization procedure. | Process design | Between May 6 and May 8, 2009, Sales Representatives visited accounts to recover the product. Recall notifications (Urgent Product Recall letters, dated May 6, 2009) were hand delivered to customers advising them Abbott Vascular is voluntarily recalling one lot of the StarClose Vascular Closure System for a potential stenlity issue discovered internally. Customers are asked to cease use of devices from this lot and work with their local account representative to review their inventory. complete the attached Recall Effectiveness Check Form and rerturn Identified products to Abbott Vascular. |  | 240 | 240 units | Nationwide distribution. | Terminated | December 21, 2011 | 959 | 240
Z-1656-2009 | 51923 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automatic external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | April 24, 2009 | 2009 | Welch Allyn Protocol, Inc | Possible for users to misunderstand signals. The original product labeling told users that a flashing low battery indicator signals that the battery is low and requires changing soon. If a user experiences a flashing low battery status indicator and does not understand that the device can continue to be used, the user might choose not to continue operating the device, thereby causing therapy to | Labeling design | The firm has developed expanded labeling for users and will revise the user manuals for software versins 2.03, 2.04, 2.05, and 2.06. Customers receiving the updated manuals will be sent a cover letter directing them to discard the original user manual and replace it with the updated manual. The April 24, 2009, cover letter will highlight the clarification of directions for use regarding what to do when a low battery indicator is displayed. |  | 46740 | 46740 units | Worldwide distribution: USA, Guam, Puerto Rico, Canada, Mexico, Andorra, UAE, Afghanistan, Australia, Azerbaijan, Belgium, Belarus, Bermuda, Bahrain, Brazil, British West Indies, Switzerland, China, Cyprus, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, France, French Guiana, French Lesser Antilles, Greece, Haiti, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Kenya, Kyrgyz, Lebanon, Lithuania, Luxembourg, Latvia, Monaco, Malaysia, Netherlands, New Zealand, Phillipines, Pakistan, Poland, Portugal, Reunion Island, Russia, Saudi Arabia, Singapore, South Korea, Taiwan, United Kingdom, Uraguay, Yemen, and South America. | Terminated | August 11, 2011 | 839 | 46740
Z-1655-2009 | 52051 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Blackstone Medical: |  |  | May 21, 2009 | 2009 | Blackstone Medical, Inc. | Product marketed without an approved 510(k) | Finished device change control | Blackstone issued the recall notice to distributors by FedEx and email on May 21, 2009. This notification provides instructions to the field representatives how to return the product to Blackstone Medical. The Dear Doctor (Surgeon) letters are being sent Tues, May 26, 2009. |  | 143 | 63 implant kits, 80 instrument kits | Nationwide | Terminated | April 07, 2015 | 2147 | 143
Z-1601-2009 | 50093 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee, A-P Lipped Tibial Insert; 11MM |  |  | May 05, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A box packaged and labeled as a Duracon Total Knee Duration A-P Lipped Tibial Insert 11MM may actually contain a 9MM insert. | Labeling mix-ups | A notification letter and product accountability form were sent via Federal Express on May 5, 2006 with return receipt to branch accounts that had received the affected product. For questions about the recall contact Stryker Howmedica Osteonics Corp. by calling 1-201-831-5000. |  | 8 | 8 | Nationwide Distribution (AL, OH, MN, NM, and CA). | Terminated | July 07, 2009 | 1159 | 8
Z-1649-2009 | 51483 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System; |  |  | March 11, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Patient injury can occur when the patient is lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. Detector motions required for this change can cause collision with the patient or pallet if it is in brain scan position, i.e. the patient or pallet is into the field of view. | Software design | Siemens sent Urgent Field Correction Recall letters dated 1/20/09 to all users of the Symbia S and Symbia T Systems. The accounts were informed of the potential injury to patients while lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. The customers were advised to perform all necessary collimator changes before the patient is positioned on the patient bed, and to examine each of their workflows that include a Tomo Acquisition activity and make sure the Auto Collimator Change check box on the Camera Parameters Tab is not checked. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a corrective software upgrade free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter. |  | 148 | 148 units | Worldwide distribution, including USA, Australia, Austria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Italy, Japan, Korea, Kuwait, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden and Switzerland. | Terminated | September 11, 2010 | 549 | 300
Z-1597-2009 | 51953 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ASC2000 Ambulatory Surgical Table. Model #MTS110003. |  |  | April 17, 2009 | 2009 | Steris Corp | Upon investigation, the firm determined that repetitive rubbing of the hydraulic hoses against the interior of the stainless steel shroud may result in the rupture of the hydraulic hoses causing loss of control of the table surface. | Device Design | A Field Correction Notice dated April 17, 2009 was issued to customers. The letter described the potential problem of the unit. The letter also notified customers of the corrective actions to be taken by the firm. A Steris technician will visit affected accounts to inspect the hydraulic hoses and replace any that show abrasion and will make other preventive corrections. For further information or if you have questions regarding the servicing of your unit, please contact the STERIS Field Service Dispatch at 1-800-288-2805. |  | 412 | 412 units | Worldwide Distribution -- US (states of AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MI, MN, MO, NC, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, VA and WV), AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, ECUADOR, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, KOREA, MEXICO, PHILIPPINES, PORTUGAL, SINGAPORE, SPAIN, SRI LANKA and THAILAND. | Terminated | March 01, 2012 | 1049 | 412
Z-1246-2009 | 51675 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer TM-400 Broach Handle; Part No. 96-210-10001. The product is a surgical instrument for use in conjunction with Zimmer TM-400 Trabecular Metal implants during spinal surgery. |  |  | April 29, 2009 | 2009 | Zimmer Inc. | The broach handle may fracture at the threaded tip during use, which would result in surgical intervention to retrieve the broach or the trial as the threaded tip fragment prevents attachment of another handle to retrieve the device. | Device Design | Distributors and user accounts were notified by "Urgent: Medical Device Recall" letter dated April 29, 2009. The letter described the affected product, reason for recall, health risks, and action for consignees. Distributors were instructed to retrieve the devices and to return them to Zimmer no later than May 13, 2009. This instrument recall will result in an inability to implant the associated devices until a new instrument becomes available.For questions or assistance about the recall please contact Zimmer Spine by calling 800-777-7505. |  | 435 | 435 | Nationwide, Austria, Australia, Belgium, Canada, Finland, Germany, Greece, Korea, South Africa, Sweden, Switzerland, Taiwan and the United Kingdom. | Terminated | October 08, 2010 | 527 | 435
Z-1611-2009 | 52090 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter; Non Sterile, Manufactured by Stryker Spine SAS, Cestas, France. Distributed by Howmedica Osteonics Corp; 325 Corporate Drive, Mahwah, NJ 07430. |  |  | April 30, 2009 | 2009 | Stryker Spine | The Mantis Rod Inserter Shaft was found to have deformation of the stainless stell split sphere when subjected to high forces, prohibiting the user from being able to grip the rod implantation and the split sphere may separate and fall into the patient. | Device Design | Stryker Spine issued an "Urgent Product Recall" notification dated April 29, 2009 via Federal Express to all Consignees. Users were instructed to quarantine all affected product and fax a copy of the Product Accountability Form to 201-760-8406. A Stryker sales representative will retrieve the quarantined product. For further information, contact Stryker Spine at 201-760-8206. |  | 175 | 175 in US | Nationwide Distribution. | Terminated | January 23, 2013 | 1364 | 562
Z-1654-2009 | 51979 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Error in therapy scheduling: If scheduled orders are documented and edited prior to the orders' original schedule time, ICIP can create a second instance of the same order marked as pending. A second instance of an order may result in a patient receiving duplicate therapy or patients may not receive their scheduled therapies. | Software design | Philips notified customers by Urgent Medical Device Correction Notification letter on May 6, 2009. The Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography.Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software corrections. The corrections will be in the form of software upgrades and will be performed by Philips service personnel.On May 27, 2009, Philips issued the second Urgent Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within lCIP to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. |  | 105 | 105 units | Worldwide distribution, including USA, Australia, Belgium, China, Denmark, France, Germany, Ireland, Italy, Kenya, Malaysia, Netherlands, Singapore, Spain, Switzerland, Taiwan, UAE, and UK. | Completed |  |  | 105
Z-1582-2009 | 52223 | No | DRN | External counter-pulsating device. | Cardiovascular | Cardiovascular | 510(k) | Cardiomedics CardiAssist External Counterpulsation System, circulatory assist device, Series 4000. |  |  | April 09, 2009 | 2009 | Cardiomedics Incorporated | Cuff may underinflate, or device may stop working during use. Cardioassist software version 4.3.1 utilized a signal from the plethysmograph that was found not to be sensitive enough due to the shape of the waveform in certain circumstances where there was excessive motion artifact. | Software design | Service visit to customer in May 2009 to upgrade and validate CardioAssist 4000 software to version 4.3.5. |  | 1 | 1 | Florida | Terminated | August 27, 2012 | 1236 | 1
Z-1553-2009 | 51951 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | Cyberkinetics Inserter Wand &amp; V2 Trigger Assembly, Expiration: July 14, 2009; Catalog Number: 4468. |  |  | February 04, 2009 | 2009 | Cyberkinetics Neurotechnology Systems, Inc. | Sterility of product cannot be assured | Packaging process control | Cyberkinetics Neurology Systems (CKI) notified the Distributor, Blackrock Microsystem, of the affected device. Blackrock notified the final user by email on February 4, 2009. |  | 2 | 2 units | Nationwide Distribution - State of MA. | Terminated | August 20, 2009 | 197 | 2
Z-1606-2009 | 52016 | No | MNS | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics BiPAP Focus Noninvasive Ventilator System. |  |  | September 22, 2008 | 2008 | Respironics California Inc | Power Supply failures have occurred on some ventilators. A discrepancy between the power supply strain relief clip and the power cord connector may allow unwanted movement of the cord and the unit's power supply. A disconnection or intermittent electrical connection between the power cord and the unit's power supply during operation or charging can result in transient power surges, which may cau | Component design/selection | Respironics California, Inc. issued an "Attention - Medical Device Field Correction" letter addressed to the Customer/Distributor dated September 22, 2008. The letter informed consignees of the affected device and instructions for removal, device replacement and return of the Business Response form by fax to 1-760-918-0169.Further question may be addressed to Respironics US Customer Service at 1-9800-345-6443 extension 5 or Respironics International Customer Service at 1-724-387-4000. |  | 2270 | 2,270 Power Cords | Worldwide to Argentina, Australia, Belgium, Brazil, Cambodia, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Libya, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Portugal, Puerto Rico, Russian Fed., Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, United States and Venezuela. | Terminated | July 06, 2010 | 652 | 2270
Z-1571-2009 | 52130 | No | OKE | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | Enforcement Discretion | Tri-State Hospital Supply Corp., Tri-State Centurion Venous Access Port Kits, Qty/case 20, sterile, Code No./Reorder KNV80. Each kit contains 1 sterile heparin flush injection solution as follows: Tyco Healthcare Monoject Prefill 100U/mL Heparin Lock Flush Syringe, 10mL, REF # 8881590121. |  |  | April 08, 2008 | 2008 | Tri-State Hospital Supply Corporation | The kits contain Monoject prefill heparin lock flush syringes, which were recalled by the manufacturer due to contamination. | Material/Component Contamination | Tri State Hospital notified their consignees by letter dated 4/8/08 and instructed to discontinue use of the product and to either discard or return it. |  | 100 | 100 kits | Indiana, Mississippi, Ohio and Tennessee. | Terminated | November 19, 2009 | 590 | 810
Z-1653-2009 | 52388 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Pediatric Two-Lumen Central Venous Catheterization Set with Blue FlexTipÂż Catheter, 5 Fr., 2 Lumen, 13cm catheter length, Catalog number: CS-14502. |  |  | June 17, 2009 | 2009 | Arrow International Inc | The spring wire guide is too large in diameter to pass through the needle, dilator and catheter. | Process design | Arrow notified consignees by letter on/about 06/17/2009. They were instructed to cease use and distribution and quarantine all affected product. They were to contact Arrow's Customer Service Department for a Return Authorization Number. Enclosed was a Recall Acknowledgement &amp; Stock Status Form to be completed and faxed to Arrow International to verify receipt of notification and acknowledge the amount of return. |  | 3019 | 3019 units | Worldwide distribution, including USA, Argentina, Brazil, India, Taiwan and Thailand. | Terminated | December 03, 2009 | 169 | 3019
Z-1652-2009 | 52430 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi ECLOS Computed Tomography (CT) Scanners, Software Version: 2.11 or less. |  |  | April 14, 2009 | 2009 | Hitachi Medical Systems America Inc | Incorrect scale on image: A software error can occur if two (2) different Field-of-View (FOV) settings are registered in a protocol using Multi-Reconstruction. This could cause the image generated to be displayed with the wrong scale factor rendering anatomical measurements incorrect. | Software design | HMSA issued an URGENT - DEVICE CORRECTION notification to their customers on 4/14/2009. The notification describes the potential software error and the condition under which these image distortions can occur. The notification goes on to provide the customer with a work-around procedure which they can use to avoid encountering this issue with their current software configuration(s). The letter states that the firm will be installing updated software on all affected systems beginning in June, 2009, and that a Hitachi Service Rep. will be contacting each customer and making arrangements to come out and install a software upgrade on the affected system(s) at no cost to the customer. Lastly, the letter provides a telephone # to the firm's Applications Helpline: 1-877-228-2777, which the customer can use to obtain additional information, if desired. |  | 4 | 4 units | Nationwide distribution, including Puerto Rico (PR). | Terminated | September 21, 2012 | 1256 | 51
Z-1650-2009 | 52436 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging Version V31 Picture Archiving and Communication System, model number 10014063. |  |  | June 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Risk of misinterpretation of patient response to treatment: Date displayed may be incorrect. Filtering based on examination date may lead to an incorrect list of examinations. | Software design | syngo Imaging has issued a Customer Safety Advisory Notice, dated May 28, 2009, to affected customers via Update Instruction IM019/09/S. The notice informs the customers of the potential issue and provides instructions to avoid its occurrence. A software update to correct this issue is currently being developed and will be installed on affected systems by sygo's service representatives when it becomes available. |  | 16 | 16 units | Nationwide. | Terminated | March 19, 2013 | 1387 | 16
Z-1577-2009 | 52171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XoranConnect, software version 4.1.0.6, Xoran Technologies, Ann Arbor, MI. |  |  | February 03, 2009 | 2009 | Xoran Technologies, Inc. | An incorrect patient name label may appear on the screen displaying the selected patient's scan. | Software design | The firm released software version 4.1.0.7 by uploading it through a remote software upgrade on February 3, 2009. User accounts were notified by letter dated February 9, 2009. Direct questions to Xoran Technologies at 1-800-709-6726. |  | 14 | 14 | Nationwide Distribution (Alabama, Arizona, Colorado, Georgia, Illinois, Louisiana, New York, North Carolina, Oklahoma, Rhode Island, Tennessee, Texas and Utah). | Terminated | November 19, 2009 | 289 | 14
Z-1542-2009 | 50824 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TumorLoc software for the Brilliance Big Bore (BB) and the CT Extended Brilliance Workstation (EBW), Model #4550 116 02431 (TumorLoc). |  |  | January 28, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software error in the TumorLoc software was discovered upon receipt of a customer complaint. There is the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | Philips Medical issued an 'URGENT -Medical Device Correction' document dated January 28, 2009. This notification informs consignees of the TumorLoc imaging issue, the inherent hazards involved should this problem occur, the circumstances under which this problem may occur and corrective actions which the user/customer should perform in the event that this problem is encountered. It also informs that user/customer that Philips will have their Service Engineers contact the customers within 6 months to install a TumorLoc upgrade. For further questions, contact the Phillips Customer Care Center at 1-800-722-9377. |  | 356 | 356 units | Worldwide Distribution - US, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Jordan, Libyan Arab Jamahiriya, Lithuania, Malaysia, Netherlands Antilles, Netherlands, New Zealand, Norway, Philippines, Portugal, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, United Kingdom and Venezuela. | Terminated | September 19, 2012 | 1330 | 356
Z-1543-2009 | 51252 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation software (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | March 24, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software anomalies may allow the image orientation to be misinterpreted, posing a potential patient safety issue. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation Software Versions 2.1.x and 3.0.x customers via an Urgent Medical Device Correction notice dated March 2, 2009 of the affected device resulting in patient safety issues. The letter provided users with safety instructions until the software is updated. Consignees are requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 1-847-939-1479.Centricity PACS RA1000 Workstation Software Versions PathSpeed 7.x, 8.x, Centricity 1.X and 2.0.X customers were sent separate letters explaining the recall including instructions. There will be no software updates since those software versions are no longer being supported.For questions, call the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 984 | 984 units | Worldwide Distribution - US including Puerto Rico and countries of Brazil, Canada, Chile, Ecuador, Mexico, Australia, China, Hong Kong, Japan, Korea, Malaysia, New Zealand, Singapore, Taiwan, Austria, Belgium, Switzerland, Germany, Denmark, Egypt, Spain, France, Great Britain, Ireland, Israel, Italy, Kuwait, Lebanon, Malta, Netherlands, Portugal, Saudi Arabia, Sweden and Turkey. | Terminated | September 06, 2010 | 531 | 984
Z-1576-2009 | 52170 | No | MSH | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(k) | Inovel N95 HealthCare Particulate Respirator, Model/Catalog Number: FRN95-SEZ, Expiration Date: 07/13/2010. |  |  | July 23, 2008 | 2008 | Inovel LLC | Display boxes labeled as FRN95-MLEZ, Medium/Large Size Respirators were inadvertently mixed with Display Boxes labeled for FRN95-SEZ Small Size Respirators. | Packaging process control | The recalling process was initiated on July 23, 2007. The affected distributor was notified of the recall through a certified recall letter with an attached Customer Response card. The distributor was instructed to notify and obtain recall product from their customers and return all affected product to Inovel. Direct questions about the recall to Inovel at 1-866-546-6835. |  | 80 | 80 cases | US (Illinois only). | Terminated | July 17, 2009 | 359 | 80
Z-1552-2009 | 51785 | No | OEU | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | Enforcement Discretion | Medtronic Ideal MAXPAC Total System, Catalog Numbers: TS1262R6, TS1262R7, TS1420R, TS1420R1, TS1442R, TS1446R, TS1464T, TS1470R, TS1470T1, TS1480R3, TS1480R5, TS1480R6, TS1487R, TS1487R2, TS8040R28, TS8108R6, TS8163R10, TS8387R5, TS9412R3, TS9453R28, TS9764R9, TS9800R28, TS9800R29, TS9800R31, TS9987R33 and TS9987R34. |  |  | March 23, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Medtronic has identified a small number of One Source Kits and MAXPAC/Total System lots with incorrect Use Before Dates on the outer kit package label. The Use Before Date on the outer label of the kit/pack does not accurately reflect the Use Before Date of the components within the kit/pack. For affected product, the outer label lists a date which exceeds the earliest Use Before Date of the com | Packaging process control | An "Urgent Medical Device Recall Notice" dated March 23, 2009 was sent to customers describing the affected device. Consignees were asked to quarantine product, complete and sign a certificate indicating acknowledgement of the notification and the amount of affected product at the facility, and to return the affected product.For further questions, contact Medtronic at 1-763-514-9605. |  | 1363 | 1363 kits total (1215 US, 148 OUS) | Worldwide Distribution -- US including states of AZ, CA, CO, CT, DE, FL, HI, IL, IN, LA, MI, MN, NC, NJ, NM, NV, NY,OH, OK, PA, TN, TX, and WI and countries of Australia, Japan and Venezuela. | Terminated | May 22, 2012 | 1156 | 1363
Z-1605-2009 | 52036 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | DERMALIFE Takara. Part #E30006-02 - Takara; Renew 2G, w/Pedestal. |  |  | June 22, 2007 | 2007 | Sybaritic, Inc | A correction or removal, conducted to reduce a risk to health posed by a device, was not reported in writing to FDA.Sybaritic, Inc. conducted a recall on DERMALIFE 2G equipment (SpaJet, Spa Oceana, Spa Feng Shui) because of IR Fan failure which resulted in the overheating and failure of the IR emitter and surrounding assembly. | Device Design | Sybaritic, Inc. issued a "Safety Alert" dated April 15, 2008 informing "All Owners/Users of DermaLife 2G Equipment" of the affected device. The letter advised consignees on device use until the mandatory Filed Correction Kit is delivered. For further questions, contact Sybaritic, Inc. at 1-800-445-8418. |  | 10 | 10 | Worldwide Distribution -- US including states of CA, IA, MN, NE, NJ, NY, TX and WA and countries of Macedonia, Hungary, Greece, South Africa, Italy, Hong Kong, United Arab Emirates, Greece, Croatia, Australia, Kuwait, Turkey, Mexico, Belgium, England, Barbados, France, Japan, Lebanon, Moldova, Poland, Portugal, Romania, Russia, Spain and Thailand. | Terminated | July 05, 2012 | 1840 | 363
Z-1569-2009 | 52099 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Amico E-Cylinder Canister Holder, distributed by Stryker Communications, Flower Mound, TX 75028. |  |  | May 19, 2009 | 2009 | Stryker Communications Corp | Design flaw on the E-Cylinder canister holder may result in early fatigue and failure of the bracket screws, which could cause the cage to become detached from its mount and fall. | Device Design | An "Urgent: Device Recall" letter dated May 19, 2009 was issued to customers. Customers were advised to discontinue use of all affected E-cylinder holders. Consignees were asked to return the "Product Accountability Form" included with the recall letter, after which consignees will be provided with an improved cylinder holder at no cost. Direct questions to Stryker Communications at 1-972-410-7100. |  | 66 | 66 units. | Nationwide and Canada. | Terminated | April 26, 2010 | 342 | 66
Z-1567-2009 | 52042 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew Endoscopic Disposable Blades:TURBOWHISKER BLD,EP-1,4.5MM(6) |  |  | May 07, 2009 | 2009 | Smith &amp; Nephew, Inc. Endoscopy Division | Product sterility is compromised due to breach in the sterile packaging. | Pending | Smith &amp; Nephew, Inc. issued a notification letter dated May 5, 2009 via Federal Express to User Facility and/or Sales Representatives on May 7, 2009 and an Email was sent to Sales Representatives. Accounts were requested to 1)complete the form and fax to Smith &amp; Nephew, Inc. at 1-508-261-3636 and 2) return (using the firm's return authorization number) affected product to Smith &amp; Nephew, Inc. Direct questions about this recall to Smith &amp; Nephew, Inc. by calling 1-508-261-3655. |  | 948 | 948 units | Worldwide Distribution -- US, Switzerland, Canada, United Arab Emerites, Belgium, Mexico, South Africa, Great Britain, Australia, Korea, Japan,Germany, Israel, India, Taiwan, Greece, Italy, Turkey and Chile. | Terminated | May 08, 2012 | 1097 | 13638
Z-1613-2009 | 52092 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Biolox-Delta Modular Ceramic Head, 28 mm head diameter, plus 5 neck, type 1, taper, sterile, Biomet Orthopedics, Inc., Warsaw, IN; Catalog Number 12-115112. |  |  | March 16, 2009 | 2009 | Biomet, Inc. | The firm has received reports of fractured heads, requiring revision surgery. | Process control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated March 16, 2009 to users describing the affected device. Consignees were instructed to locate and discontinue use of the product and quarantine until a sales representative can arrange for removal. Consignees were also asked to complete and return the attached "Fax-Back Response Form" to Biomet at 1-574-372-1683.Further questions can be directed to Biomet Orthopedics at 1-800-348-9500. |  | 174 | 174 | Worldwide Distribution -- US including PR and states of AZ, AL, AR, CA, CO, FL, IL, IN, KY, KS, MA, MI, LA, NY, SD, GA, MN, PA, NJ, MD, MI, OH, ME, TX, WA, WI and VA and countries of Australia, Belgium, Canada, Costa Rica, Finland, Netherlands and South Wales. | Terminated | October 06, 2010 | 569 | 335
Z-1679-2009 | 51713 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models F102, F110. |  |  | March 23, 2009 | 2009 | Boston Scientific CRM Corp | Lead connection issues or fracture increases probability of inappropriate therapy: Inappropriate shocks, inappropriate anti-tachycardia pacing therapy, or inappropriate inhibition of needed bradycardia pacing therapy. | Device Design | Boston Scientific sent an Urgent Medical Device Information letter, dated March 23, 2009, to physicians following implanted devices. The letter described the issue, the rate of occurance and provided Patient Management Recommendations. |  | 49148 | 49,148 total | Worldwide distribution: USA, Guam, Puerto Rico, Aruba, Austria, Belgium, Brazil, Canada, Cayman Islands, Colombia, Cyprus, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Kuwait, Latvia, Lebanon, Luxembourg, Mexico, Netherlands, New Caledonia, Norway, Philippines, Poland, Polynesia, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | October 18, 2012 | 1305 | 49148
Z-1673-2009 | 52461 | No | KYB | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(K) Exempt | Optiflex Lens Positioning Unit (LPU) Product Code: 10970, for use in vitreoretinal ophthalmic surgical procedures. |  |  | January 23, 2009 | 2009 | Volk Optical Inc | Lens may fall onto eye surface during procedure. Ophthalmologic device had a loose set screw on the Ball-Seal bushing assembly. As a result, the set screw becomes disengaged from the set screw groove in the bushing assembly and the Lens Positioning Unit slides downward. | Process design | Volk Optical sent a 'MEDICAL DEVICE RECALL' letter to their sole customer on 1/23/2009. The letter requests the return of the remaining two (2) units in the customer's possession (this firm experienced the device failure and reported the consumer complaint which prompted the recall; and had already returned the affected unit to the Volk for refund or replacement). The letter provides Volk's Fedex and UPS numbers for the customer to use to return the recalled units and also provides the E-mail address for the firm's contact person concerning this recall should additional assistance be required. |  | 3 | 3 units | USA distribution to a single customer located in MO. | Terminated | April 12, 2012 | 1175 | 3
Z-1618-2009 | 52191 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Revision Instruments; 1,2,7,8 TGC Upper Tray; Non Sterile, |  |  | May 15, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The screen print for the Valgus Adapter catalog numbers on the Triathlon Trial Cutting Guide Instrument Trays is incorrect. | Process change control | Stryker Orthopaedics issued an "Urgent Product Correction" notice dated May 15, 2009 via Federal Express addressed to Branch/Agency Manager/Quality Contact informing them of the affected device. The letter outlined instructions for correction and requested consignees to complete and return a Product Correction Acknowledgment Form by fax to 1-201-831-6069.Questions can be directed to Colleen O'Meara, Manager, Divisional Regulatory Reporting at 201-831-5970. |  | 27 | 27 | Worldwide Distribution -- US and Canada. | Terminated | April 05, 2010 | 325 | 53
Z-1245-2009 | 51934 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | Hydration Station with Pedestal (Part # E30003-01), Hydration Station with Pedestal &amp; DVD (Part # E30003-01D), Sybaritic, Inc., 9229 James Avenue South, Minneapolis, MN 55431. |  |  | May 15, 2007 | 2007 | Sybaritic, Inc | Sybaritic conducted a recall on the Hydration Station 806 because of the potential for hot water to spurt through steam ports into unit and cause burns. Also, potential for customers to inadvertently place their legs over the side wall of the unit directly onto the steam ports and cause burns.There were 18 injuries and 8 malfunctions reported. | Device Design | Sybaritic, Inc. issued a letter dated May 15, 2007 addressed to "Hydration Station Customer". The letter informed the customer of the new 2007 Hydration Station Upgrade Improvement Kits including instructions and advised the customer to coordinate the installation of the components with their Technical Service Department.Further questions can be addressed to Sybaritic, Inc. Technical Administration at 1-800-445-8418 extension 373. |  | 370 | 370 | Worldwide Distribution - US including states of AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV and WI and CANADA. | Terminated | July 05, 2012 | 1878 | 370
Z-1624-2009 | 51903 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConforMIS iUni ( Unicondylar Knee System iUni Right Lateral Implant Kit Product Code: M57220600040. |  |  | April 23, 2009 | 2009 | Conformis Inc | Microscopic cracks potential for pre-mature revisions in patients implanted with the system. | Process control | ConforMIS contacted accounts and presented the surgeon with the customer communication letter dated April 23, 2009, customized for the surgeon by serial number and patient identification. Product not yet implanted will be recovered. A patient evaluation and monitoring protocol will be recommended for potentially affected product already implanted.Contact ConforMIS for further questions at 1-781-345-9164. |  | 1 | 1 unit | Worldwide Distribution - US including states of CA, LA, MI and TX and Germany. | Terminated | October 14, 2009 | 174 | 17
Z-1554-2009 | 51986 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Waste bottles Part Numbers/Article Numbers: 905-590/D512 and 905-802/D513. |  |  | April 16, 2009 | 2009 | Radiometer America Inc | When removing the waste bottle from the analyzer waste material may leak both from the analyzer and from the waste bottle. | Process design | Radiometer America Inc. issued a recall notification letter dated May 5, 2009 to all customers containing information on replacement of all potentially defective waste containers and instruction on how to handle any mounted waste containers.For further questions, contact Radiometer America Inc. at 1-800-736-0600. |  | 905590 | 905-590 (D512): 35,552 bottles; 905-802 (D513): 45,204 bottles | Worldwide Distribution to US and Canada. | Terminated | May 09, 2012 | 1119 | 905590
Z-1705-2009 | 52521 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Paradigm Quick-set Infusion Sets, Models MMT-396, MMT-397, MMT-398 and MMT-399. |  |  | June 29, 2009 | 2009 | Medtronic MiniMed | Approximately 2% of the infusion sets in the affected lots may not allow the insulin pump to vent properly. Venting is necessary to equalize the pressure in the syringe compartment with the surrounding atmosphere. If the vent does not work properly, this could potentially result in too much or too little insulin being delivered. | Process control | Medtronic issued an "URGENT MEDICAL DEVICE RECALL" notification letter dated July 6, 2009 to health care professionals, followed by notification letters dated July 7, 2009 to distributors, inactive and active customers. Letters to healthcare professionals and distributors were sent via UPS 2-Day Air with signature confirmation of delivery required. On July 8, 2009 and July 9, 2009, a notification letter, along with replacements sets, were mailed via UPS to active patients or customers (who have ordered since January 2009). The notification letter to patients (inactive customers) who last ordered prior to January 1, 2009 were issued via First Class U.S. Mail. All letters to patients include a preaddressed, postage paid card to confirm that the patient has received the recall notice and to identify whether or not they have unused sets to return. As described in the patient letters, patients have the option to confirm receipt of the notification and to order replacement infusion sets online or via an automated telephone voice response system.Distributors who purchased affected Paradigm Quick-set infusion sets must either notify all patients to whom they have provided infusion sets or provide patient names and addresses to Medtronic MiniMed so that the firm can carry out the notification.The firm's website, www.medtronicdiabetes.com/lot8, is available for customer concerns and questions not addressed in the recall letter. The website is also the most efficient way to exchange Lot 8 Quick-set infusion sets and/or place new supply orders.In the event the above website is not accessible or for additional questions, please call 1-800-345-8139 at Medtronic Diabetes 24- hours a day. Doctors who would like to speak with a Medtronic Diabetes medical officer may call 1-818-576-4211, 24-hours a day. Report product problems by calling 1-800-345-8139 at any time.Medtronic issued a press release on July 10, 2009. |  | 28007256 | 28,007,256 units | Worldwide Distribution -- US, Canada, Mexico, Bermuda, Brazil, United Kingdom, Ecuador, Paraguay, Kuwait, Germany, El Salvador and Turkey. | Terminated | November 08, 2010 | 497 | 28007256
Z-1620-2009 | 51345 | Yes | LWP |  |  | Cardiovascular | PMA | Guidant ContakÂż RenewalÂż 3 CRT-D Model H170; 3CRT-D, Model H175; 3HE CRT-D, Model H177; 3HE CRT-D, Model H179; 4CRT-D REF H190; 4CRT-D REF H195; 4HE CRT-D REF H197; 4HE CRT-D REF H199; 4AVT CRT D, REF M177; 4AVT HE CRT-D REF 179. Cardiac Resynchronization Therapy Defibrillator, Sterile EO, Guidant Corporation, 4100 Hamline Avenue North, St Paul, MN 55112-5798. |  |  | March 04, 2009 | 2009 | Boston Scientific CRM Corp | In April 2007, Boston Scientific CRM communicated with physicians regarding the potential for reduced Elective Replacement Indicator (ERI) to Battery End of Life (EOL) time in a subset of implantable defibrillators due to degradation of a low-voltage capacitor. Since that time, the April 2007 advisory population has not experienced any clinically significant changes to either the rate of occurrenc | Nonconforming Material/Component | Boston Scientific CRM notified Doctors via sales representatives of an "Urgent Medical Device Information" letter dated March 4, 2009. The letter described the problem and included Description of Issue, Rate of Occurrence, Patient Management Recommendations, Devices Affected, Warranty Program &amp; Further Information. A Boston Scientific " Product Advisory" letter dated March 4, 2009, was also made available to physicians to distribute to patients. The letter informed the patient that their defibrillator was included in a group of devices that might experience a performance issue. For further questions, Contact Boston Scientific CRM at 1-800-227-3422. |  | 112 | 112 | Worldwide Distribution - US including PR and states of AL, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, IRELAND, ITALY, JAPAN, LEBANON, NETHERLANDS, NEW ZEALAND, POLYNESIA, PORTUGAL, SPAIN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | October 04, 2012 | 1310 | 1383
Z-1548-2009 | 51485 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex VixWin Platinum Imaging Software, Versions 1.0, 1.1 and 1.2; Gendex Dental Systems, 901 West Oakton Street, Des Plaines, IL 60018. |  |  | March 06, 2009 | 2009 | Gendex Dental Systems | There is a compatibility issue when the VixWin Platinum versions 1.0/1.1/1.2 software is used in conjunction with Easy Dental, PracticeWorks, EagleSoft and SoftDent practice management software, resulting in patient files that may contain images that are mixed between patients. | Software design | Gendex issued an "Urgent Medical Device Report" dated March 6, 2009 to the direct account dealers of the affected device via first class mail. The noticed described the issue and provided a software update directed to fixing the problem. In addition, customers were asked to complete and return an Acknowledgement Form by fax to 1-847-718-0601.Further questions should be directed to the Gendex Technical Support team at 1-888-275-5286. |  | 1256 | 1,256 units | Nationwide Distribution. | Terminated | January 07, 2011 | 672 | 1256
Z-1674-2009 | 52129 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Essential Medical Supply brand Essential Rollator with Loop Hand Brakes, models W1700B (Blue) and W1700R (Red), Four Wheel Aluminum Walker/Rollator |  |  | May 08, 2009 | 2009 | Essential Medical Supply, Inc. | The caster may separate from the frame completely causing injury to the user. | Device Design | Essential Medical Supply Inc issued a Notice of Quarantine Letter to all customers on 04/30/09, identifying affected products as W1700B &amp; W1700R Essential Rollators, Serial #'s KLD-E07090001 through KLD-E08020570, received after 11/05/2007. Firm issued a Device Recall Letter to all direct customers on 05/08/09 instructing them to do the following: (1) segregate and quarantine any affected product; (2) identify and notify any customers that may received the affected products, instructing them to return the product to their dealer; and (3) either return the affected products to Essential Medical Supply or destroy the products and send proof of destruction. |  | 2265 | 2265 | Nationwide distribution. | Terminated | September 15, 2010 | 495 | 2265
Z-1599-2009 | 52172 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax 110v Surgical Table Hand Control. |  |  | May 12, 2009 | 2009 | Steris Corp | A limited quantity of the Cmax Hand Controls, P150832-500, were manufactured without Loctite thread locker on the screws that attach the mounting clip to the hand control case. This clip provides a means of storing the hand control on the table side rails. It is possible that screws without Loctite can work loose and the clip can become separated from the hand control. If both screws back out c | Process control | Steris Corp. issued an "Urgent Product Correction" notice dated May 12, 2009 informing consignees of the affected device and that a Steris Service Representative with be in contact to make the necessary inspection and repairs to the device. Steris Field Service Dispatch can be reached at 1-800-288-2805.For further questions, contact Steris at 1-800-548-4873 extension 27318. |  | 217 | 217 Units | Worldwide Distribution -- U.S. states of AL, AR, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WI and WV and countries of Canada and Taiwan. | Terminated | February 26, 2013 | 1386 | 217
Z-1711-2009 | 52431 | No | EXQ | Urodynamics measurement system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Laborie Medical Technologies Urodynamic Catheters, Twin Lumen 6 Fr , 40cm For Cystometry. Catalog Number: CAT-206. |  |  | April 09, 2009 | 2009 | Laborie Medical Technologies Cor | Laborie Medical Technologies reports that there was a mix up in the labeling and packaging of their Urodynamic Catheters. There were 8Fr catheters mixed inside the 6Fr catheter packages. | Packaging process control | Laborie Medical Technologies notified customers by telephone/email beginning April 10, 2009 advising them of the mislabelled product.For further questions, contact Laborie Medical Technologies at 1-800-522-6743. |  | 28 | 28 boxes (10's) | Nationwide Distribution -- CA, GA, IL, KS, MD, PA, OH, OK, UT and WA. | Terminated | August 20, 2009 | 133 | 28
Z-1710-2009 | 52453 | No | HKS | Gonioscopic prism. | Ophthalmic | Ophthalmic | 510(K) Exempt | G-4 HighMag Flange, Model #VG4HM. Standard Gonio Lens for High Magnification Static and Dynamic Gonioscopy. |  |  | June 30, 2008 | 2008 | Volk Optical Inc | The ring of the Gonio lenses distributed was engraved as, 'G4 High Mag Gonio LNF'; however, the engraving should have read simply: 'G4 High Mag Gonio.' | Process design | Volk Optical Inc. contacted their customers and notified them of affected device via E-mail, letter, and Fax transmittals starting June 30, 2008. The recall notification(s) describes the engraving error and instructs the users on steps to replace the affected product.For further questions, contact Volk Optical Inc. at 1-800-345-8655 extension 443 . |  | 15 | 15 | Worldwide Distribution -- Great Britain, China and the U.S. | Terminated | March 09, 2012 | 1348 | 15
Z-1598-2009 | 52283 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fusion Omni-Tome, Rx Only, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105. |  |  | June 12, 2009 | 2009 | Cook Endoscopy | The wire control port of these devices may be missing due to omission of a manufacturing activity. | Process control | The distributor (Cook Ireland) was instructed to contact their customers and coordinate the return of all recall product to Cook Endoscopy. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed.Contact the Customer Quality Assurance Department at Cook Endoscopy for more information at cws.complaints@cookmedical.com or 1-336-744-0157. |  | 20 | 20 units | International Distribution - Italy, Netherlands, Germany and Spain. | Terminated | April 06, 2010 | 298 | 20
Z-1680-2009 | 50901 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Super Secur Fit Plus. |  |  | January 06, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Product deviation : Specific lots of sprayed HA hip stems did not meet Stryker's Internal Material Specification for tensile bond strength and crystallinity. | Process control | An "Urgent Product Recall" letter was sent via Federal Express on January 6, 2009 with return receipt to Stryker Branches/Agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letter described the affected product, issue, potential hazards and risks. Stryker Branches and Hospital Risk Management were asked to examine their inventory for affected product. Stryker Representatives were instructed to contact customers to reconcile and return recalled product to Stryker Orthopaedics. All letters include a response form to be completed and returned indicating that they have been informed of the recall and the amount of product on hand, if any. Direct questions about the recall to Stryker Orthopaedics at 1-201-831-5718. |  | 1628 | 1,628 total stems, all varieties. | Nationwide Distribution. | Terminated | June 25, 2012 | 1266 | 1628
Z-1607-2009 | 52072 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Offset Adaptor Trial - 2MM, Triathlon Revision Instruments. Non-Sterile; Howmedica Osteonics Corp: 325 Corporate Drive, Mahwah, NJ 07430. |  |  | April 16, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Triathlon Offset Adaptors may seize during surgery and the OR staff may be unable to disassemble the instrument using the removal tool. | Process control | "Urgent Product Correction" letters dated April 16, 2009 were sent to all Stryker Representatives, Hospital Rick Management Departments, Hospital Chief of Orthopaedics, Surgeon and countries by Federal Express. The notice explains the Issue, Potential Hazards and Risk Mitigation Factors of the affected product(s) including a request to complete and return the Product Correction Form.For further questions, contact Stryker Orthopedics, Division of Regulatory Reporting at 201-831-5970. |  | 2498 | 2498 units total, all sizes. | Worldwide Distribution. | Terminated | April 05, 2010 | 354 | 2498
Z-1699-2009 | 52227 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Allen medullary cement plugs, 1 - 20 dia. flange/10 mm dia. core, 1 - 24 mm dia. flange/12 mm dia. core (polyethylene with barium sulfate) with inserter, sterile, Zimmer, Warsaw, Indiana; Catatog Number: 00801102001. |  |  | June 08, 2009 | 2009 | Zimmer Inc. | The seal on the sterile barrier pouch may be inadequate, resulting in lack of assurance of sterility. | Packaging process control | Zimmer Orthopedic Surgical Products issued an "Urgent: Medical Device Recall" notice dated June 8, 2009 informing users of the affected devices and actions to be taken including the return of product and acknowledgement of receipt of notice by returning the Recall Certification Form by fax to 1-330-364-0974. |  | 133 | 133 | Worldwide Distribution -- United States, Canada and Japan. | Terminated | December 15, 2009 | 190 | 258
Z-1596-2009 | 52159 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Stat (TM) Super-Fine (TM) Pen Needles and Allison Medical, Inc. Sure Comfort Pen Needles with Silicone Tip, for human use with insulin pen injector device to administer insulin subcutaneously. |  |  | May 29, 2009 | 2009 | Stat Medical Devices, Inc. | Stat Pen Needles with silicone coating have not been cleared by FDA for use with silicone coating which could significantly affect the safety and effectiveness of the device. | PMA | An "Urgent: Medical Device Recall Notification" letter dated June 18, 2009 was issued to direct customers requesting return of affected product and requesting they send notification to any customers which may have received the affected product from them. Direct questions about the recall to Stat Medical Devices by calling 1-978-774-1469. |  | 6424100 | 6,424,100 total pen needles | Worldwide Distribution -- US (CO and Puerto Rico), Canada, and Saudi Arabia. | Terminated | September 07, 2010 | 466 | 6424100
Z-1636-2009 | 51273 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Passive Clip Slings; washable sling; Manufactured in Belgium for Medibo NV, Heikant 5, 3930 Hamont Achel, Belgium; these slings are intended for use with ARJO and BHM Passive Patient Lifts and Ceiling Lifts with four-point hanger bars intended for use with clip slings. |  |  | February 26, 2009 | 2009 | ArjoHuntleigh | If the sling clips are not applied or are applied incorrectly, there is the potential for the patient to fall, resulting in injury to both the patient and/or the caregiver. | Other | An "Urgent Field Safety Notification" letter dated January 30, 2009 to all customers with ARJO/BHM Passive Patient Lifts and Ceiling Lifts with four-point hanger bars via UPS 2nd day on the February 26, 2009. The notification letter informed the accounts of the potential for clip detachments if the sling clips are not applied to the hanger bar attachment lugs. The accounts were instructed to do the following: 1) Apply the stickers provided to the passive patient lift and/or ceiling lift hanger bars as shown in the enclosed Sticker Application Instructions sheet; 2) Retrain lift operators in accordance with the enclosed instructional DVD, the revised Passive Clip Sling Operating and Product Care Instructions manual and the revised Flites Clip Operating and Product Care Instructions manual; and 3) Make lift operators aware of the Clip Sling/ Flites Clip Application Instructional poster and its location within their facility. The accounts were also instructed to complete and fax back to Arjo the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, acknowledging receipt and understanding of the Field Safety Notice.If you have any questions or require any additional information, contact the ARJO, Inc. Quality, Customer Support, or Technical Support Departments via e-mail at quality@arjousa.com, via phone at 1-800-323-1245, or via fax at 1-888-594-2756. |  | 103108 | 103,108 total slings | Nationwide. | Terminated | February 25, 2011 | 729 | 103108
Z-1802-2009 | 52125 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Medtronic INTREPID Spinal System Coverplate, size: Small, Part Number: REF 7961330; medium, REF 7962330; large, REF 7963330; STERILE R. |  |  | February 06, 2009 | 2009 | Medtronic Sofamor Danek USA Inc | The firm received complaints regarding the use of the product in patients with poor bone quality, translational instability, difficulty implanting the cover plate and case of cover plate detachment. | Device Design | Medtronic Sofamor Danek USA, Inc. instructed Sales Representative on starting February 6, 2009 by phone to immediately return product from the field.The firm further issued a "Subject: Medtronic Intrepid Intervertebral Body Fusion Device Recall" notice dated April 2009. The recall notices were mailed directly to Risk Managers and Surgeons, with Recall Questionnaires and Surgeon Confirmation forms included respectively.For further questions, contact your Medtronic Sales Representative or the Medtronic Global Quality Department at 1-800-876-3133 extension 6333. |  | 1059 | 1059 units | Worldwide Distribution -- United States, Australia, Germany, The Netherlands, and South Africa. | Terminated | April 19, 2010 | 437 | 9087
Z-1588-2009 | 52038 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ansell, Derma Prene Ultra, Powder-Free, Smooth Surgical Gloves, REF 8517, Size 8 1/2, 4 boxes of 50, Sterile, Distributed by Ansell Healthcare, 1635 Industrial Road, Dothan, AL, 36303, USA, Made in Sri Lanka. |  |  | May 01, 2009 | 2009 | Ansell Healthcare Products LLC | Premature degradation which could result in tearing during use. | Process control | The firm initiated the recall on May 1, 2009 by letter to Health Canada and by letter dated May 4, 2009 to its US and Canadian consignees. The letters described the affected product, issue and actions for customers. The recall was extended to the hospital and clinic level requesting that the product be returned. Direct questions to your Ansell Healthcare Products LLC representative. |  | 560000 | 560,000 total units--all lots. | Nationwide Distribution and Canada. | Terminated | April 12, 2010 | 346 | 560000
Z-1795-2009 | 52575 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx/MRxE Defibrillator/Monitor Models: M3536A (HeartStart MRx) and M3536J (HeartStart MRxE). Note: M3536JIMRxE was sold to Japan only. |  |  | July 07, 2009 | 2009 | Philips Healthcare Inc. | External transport use environments, the mechanical/electrical connection between the padstherapy cable including pads/CPR cable and the MRx/MRxE may wear at an increased rate ultimately prevent the device from sensing that the padstherapy cable is connected, or, cause the device to inappropriately identify external paddles or internalpaddles when the pads therapy cable is connected | Device Design | Philips Healthcare issued an "Urgent - Voluntary Medical Device Correction" notice dated July 2009 providing information on the affected device, namely, how to identify the affected devices and instructions on actions to be taken. Consignees are asked to contact their local Philips representative to confirm receipt of the recall letter and to arrange for service.Contact Philips Healthcare for further questions at 1-800-722-9377. |  | 11579 | 11,579 units | Worldwide Distribution -- United States, Afghanistan, Albania, Australia, Austria, Bahrain, Brazil, China, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Latvia, Lebanon, Libyan Arab Jamaica, Lithuania, Netherlands, New Caldonia, New Zealand, Norway, Panama, Peru, Poland, Portugal, Reunion, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom and Venezuela. | Terminated | February 19, 2013 | 1323 | 11579
Z-1616-2009 | 52169 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 32 mm diameter, neck length 0 mm, sterile, Zimmer, Warsaw, IN; Catalog Number: 6428-32-02. |  |  | May 29, 2009 | 2009 | Zimmer Inc. | The corner tab of the inner compartment lid was not folded under before the outer lid was sealed, resulting in lack of assurance of sterility. | Packaging process control | Zimmer, Inc. issued "Urgent: Device Recall" notices dated May 29, 2009 informing Consignees of the affected devices. Users were instructed to locate and return affected product and complete and return an Inventory Return Certification form via fax to the firm at 1-547-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 147 | 147 of both products | Worldwide Distribution -- United States, Brazil, Canada, Chile, Czech Republic, Finland, France, Germany, Italy, Korea, Lebanon, Switzerland and United Kingdom. | Terminated | December 22, 2009 | 207 | 147
Z-1692-2009 | 52715 | No | MEQ |  |  | Gastroenterology/Urology | PMA | Prolieve Thermodilatation Catheter Kit, a Transurethral Microwave Therapy (TUMT) System, for BPH (Benign Prostatic Hyperplasia) Thermotherapy, Material Numbers M0068808022 and M0068808170. |  |  | July 20, 2009 | 2009 | Boston Scientific Corporation | Balloon leaks: Boston Scientific has received complaints from the field involving product leaks associated with the anchor balloons and compression balloons. Correspondence between BSC and FDA has led to BSC's decision to remove the Prolieve Catheter Kits from the field. | Process design | Boston Scientific sent Customer notification letters via Federal Express Priority mail on July 20, 2009. A total of 812 US consignees were notified and a total of 26 consignees in Puerto Rico. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in their inventory. Follow-up communications will be made to the accounts that do not respond to this inquiry. |  | 26935 | 26,935 kits | Nationwide distribution, including Puerto Rico. | Terminated | March 04, 2010 | 227 | 26935
Z-1806-2009 | 52705 | No | MAF |  |  | Cardiovascular | PMA | POWERSAIL Coronary Dilatation Catheter: 2.75x18mm (US) |  |  | June 17, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Distal shaft of the POWERSAIL coronary Dilatation Catheters exhibits damage that may result in a leak of contrast material and functional failures. | Process control | Abbott Vascular sent an Urgent Device Recall letter (w/return form) to customers on June 17, 2009 requesting return of identified products to Abbott Vascular Territory Manager. Abbott Vascular issued a press release dated July 28, 2009. |  | 557 | 557 total devices | Worldwide: United States, Australia, Belguim, Canada, China, Czech Republic, France, Germany, Greece, India, Italy, Japan, Jordan, Mexico, Netherlands, Norway, Poland, Puerto Rico, Spain, Sri Lanka, Switzerland, Thailand, and United Kingdom. | Terminated | July 06, 2010 | 384 | 557
Z-1635-2009 | 50800 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Datex-Ohmeda S/5 Anesthesia Delivery Unit |  |  | October 30, 2008 | 2008 | GE Healthcare | GE Healthcare has received reports of breathing circuit or tubing misconnections of the Anesthesia Delivery Unit (ADU), that may impact patient safety. | Other | A GE Healthcare "Urgent Medical Device Correction" letter was sent to consignees on 1/16/09. The letter was addressed to Health Care Administrator, Chief of Anesthesia and Director of Clinical Engineering. The letter described the Safety Issue; Affected Product Details; Safety Instructions, which include clinical personnel being properly trained on the correct connections of breathing circuit tubing and proper execution of System Checkout according to the ADU URMs; Product Correction and Contact Information. |  | 6621 | 6,621 | USA including states of AL, AK, AZ, AR, CA, CO, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, and WI and countries of Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Bosnia &amp; Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Egypt, England, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Iraq, Ireland, Isle of Man (UK), Israel, Italy, Japan, Jordan, Kazahkstan, Kuwait, Latin America, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Yemen. | Terminated | May 17, 2012 | 1295 | 9633
Z-1873-2009 | 52721 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn Vital Sign Monitor 5300 Series/Class II. (VSM 300). |  |  | July 28, 2009 | 2009 | Welch Allyn Protocol, Inc | The Welch Allyn VSM300 Vital Signs Monitor may experience a loss of audio which includes a loss of audio alarms. | Device Design | An "Urgent: Correction" letter dated July 28, 2009 was sent to customers describing the risk associated with loss of audible alarms and recommend that the users mitigate the risk by checking for proper audible alarm function each time they power on the device. (Current labeling describes the variety of beeps and tones that should be present). Users were instructed: 1) if the audio fails, do not use the device, 2) Remove it from service and 3) contact Welch Allyn Protocol, Inc. at the following telephone number (1-800-289-2501) to arrange the return and repair of your device to Technical Services. Direct questions about this recall to Welch Allyn Protocol, Inc. by calling 1-503-530-7935. |  | 68123 | 68,123 | Worldwide Distribution -- US, United Arab Emirates, Afghanistan, Austria, Australia, Belgium, Bulgaria, Bahrain, Botswana, Canada, Switzerland, China, Cyprus, Germany, Denmark, Egypt, Spain, Figi, France, Great Britain, Greece, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Iceland, Italy, Jordan, Kenya, Kyrgyzstan, South Korea, Kuwait, Lebanon, Sri Lanka, Malaysia, Nigeria, The Netherlands, Norway, New Zealand, Oman, Phillipines, Pakistan, Poland, Puerto Rico, Panama, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Thailand, Taiwan and South Africa. | Terminated | August 10, 2011 | 743 | 68123
Z-1675-2009 | 52435 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition Model AS/AS+, Model number 8098027. |  |  | May 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Support mechanism may not engage properly due to improper stowage after opening which may result in internal component damage or potential injury. | Equipment maintenance | The firm has issued a modification of the telescope support arm parking bracket via Update Instruction CT028/09/S. A firm service representative will install the modified support arm parking brackets on the affected systems. |  | 93 | 93 units | Nationwide Distribution -- including states of AR, AZ, CA, CO, CT, FL, IA, IL, IN, KS, KY, MA, MD, MN, MO, MT, NC, ND, NE, NH, NY, OH, PA, SC, SD, TN, VA, WA, WI, and WY. | Terminated | September 01, 2009 | 95 | 93
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415
Z-1626-2009 | 52189 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 24F (8 mm) Safety PEG Kit, Push Method for use with sterile Xylocaine ampule, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6649, inner pouch UPN M00566490 and outer carton UPN M00566491. |  |  | May 05, 2009 | 2009 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Process control | Boston Scientific Corporation issued an "Urgent Medical Device Recall - Immediate Action Required" notice dated May 5, 2009 as follow-up to their "Recall of the Safety and Standard Push PEG Kits" dated December 30, 2008. The notice outlined the issue and Consignees were instructed to locate and return all affected product to the firm. In addition, the firm requests that users complete and return a Reply Verification Tracking Form via fax to 1-508-683-5578.For further questions, contact the Boston Scientific Corporation Field Action Team by phone at 1-508-683-4678. |  | 140 | 140 | Worldwide Distribution -- United States, India, Japan and Slovenia. | Terminated | October 13, 2010 | 526 | 10259
Z-1790-2009 | 52119 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Alaris PC Unit, Model Numbers 8000 and 8015 (formerly Medley PC Unit) with software version 4 and above. |  |  | June 12, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | There is a potential risk for medication not being delivered when IV tubing sets are occluded in the Alaris Pump module. This can occur if IV sets are set up in the infusion pumps but not used for an extended period of time. If users fail to clear the occlusion, then the occluded tubing may cause medication not to be delivered. A delay of medication may result in serious injury and/or death, an | Device Design | The recall was initiated on June 12, 2009 with the firm forwarding via UPS an Urgent: Medical Device Recall Notification with attached Customer Response Card to the Chief Administrative Officer, Director of Pharmacy, Director of Biomedical Engineering, Director of Nursing &amp; Director of Risk Management of all of the affected consignees. Customers were also asked to please promptly complete and return the enclosed Customer Response Card to expedite the correction process.The Recall Notification informed the consignees of:(A) PROBLEM #1: Occlusion Warning Message for Alaris Pump Module; (B) PROBLEM #2 Syringe volume Warning Message for Alaris PCA module; (C) PROBLEM #3 Electrostatic Discharge (ESD) protection circuitry; (D) PROBLEM #4 Fluid Ingress and (E) PROBLEM #5 Alaris System Inter-Unit Interface (IUI) Connectors Labeling. The recall notice included "Required Action for Users &amp; Follow-up actions by Cardinal Health for all five problems. Customers were also asked to report any adverse reactions experienced with the use of this product, and/or quality problems to the FDA's MedWatch Program. In addition, customer were provided the following contact information for the recall: (1) Cardinal Health Recall Center for general questions: By phone at 1-888-562-6018 7am to 5pm (Pacific) ; (2) Customer Advocacy for adverse reports: By phone at 1-800-854-7120, Option 1, Option 1, Option 3 or by email at customerfeedback@cardinalhealth.com 24 hours a day, Sunday-Saturday; (3) Technical Support for Technical Questions regarding the Alaris System: By phone at 1-888-562-6018, 6am-5pm (Pacific).The firm followed-up with the release of a nationwide press statement regarding this issue on July 29, 2009. The press release indicated serial numbers of affected devices, as well as CareFusions short term instructions to customers, and the firms strategy to fix the affected devices can be found at: www.cardinalhealth.com/alaris/medical-device-recall/ (This informatio |  | 186000 | 186,000 units | Worldwide Distribution -- US (including Puerto Rico and Guam), Germany, England, Canada, Netherlands, Malaysia, Singapore, Australia, and New Zealand. | Terminated | March 19, 2012 | 1011 | 1016000
Z-1875-2009 | 52032 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂż MC3 (model M4000) Low Airloss Therapy Units, for the prevention of pressure sores. |  |  | May 12, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | The devices may emit smoke from the blower box assembly. | Equipment maintenance | Urgent Field Safety Notice letter dated May 11, 2009. Hill-Rom service centers were notified by letter on 05/13/2009. Action to be taken are as follows: Rental Customers; In the next 60 days, Hill Rom will be removing the FlexicairÂż II and FlexicairÂż MC3 low Airloss Therapy Units from service to correct the devices. During that time, Hill-Rom may provide an equivalent substitute therapy to accommodate the patients who continue to need a pressure reducing therapy device. If t his malfunction occurs before a substitution or a corrected device is provided, unplug the device, remove the patient from the surface and notify Hill-Rom immediately. Facility-Owned Customers: In the next 60 days, for all facility owned Flexicair II and FlexicairÂż MC3 Low Airloss Therapy Units, a Hill-Rom service technician will either replace the affected components in the blower assembly or deliver an equivalent substitute unit from Hill-Rom's rental fleet. If this malfunction occurs before the device is corrected or a substitute provided, do as follows: unplug the device, remove the patient from the surface, and immediately notify Hill-Rom Technical Support using the attached listing their geographical region. |  | 889 | 889 units | Nationwide and Canada | Terminated | April 03, 2012 | 1057 | 946
Z-2046-2009 | 46722 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns cardiovascular disposable products, Weighted Flexible Sucker, sterile, Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; Catalog No. 804113. |  |  | February 05, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | A loose strand of stainless steel left on the device during manufacturing may result in an unretrieved metal fragment remaining in the patient. | Process control | Consignees were notified via Urgent - Medical Device Removal letters dated 2/5/08 to cease use of the product. Terumo CVS will contact customers to arrange replacement for the affected product upon receipt of the attached response form. Questions are to be directed to Terumo CVS Customer Service at 1-800-521-2818. |  | 7079 | 7079 | Worldwide Distribution --- including USA and countries of Belgium, Brazil, Dominican Republic, Germany, Jordan and Latin America. | Terminated | November 01, 2009 | 635 | 7079
Z-1676-2009 | 52091 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Siemens Medical Solutions USA, Inc., Hoffman Estates, IL 60195-2061 |  |  | April 22, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | The pinion gear could crack and cause the gantry to spin freely thereby damaging the Symbia S or T System and potentially injuring a patient during a scan. | Finished device change control | Siemens sent Urgent Field Correction Recall letters dated 4/22/09 to the affected users of the Symbia S and Symbia T Systems. The accounts were informed of the potential faulty pinion gear which could crack and cause the gantry to spin freely, damaging the system and potentially injuring a patient during a scan. The customers were advised to stop using the system if they hear any abnormal grinding noise from the gantry. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a replacement of the pinion gear free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter. |  | 37 | 37 units | Worldwide Distribution -- USA, Australia, Austria, Brazil, Bulgaria, Canada, China, Finland, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Philippines, Russia, Spain, Taiwan, and Turkey. | Terminated | June 21, 2010 | 425 | 75
Z-1702-2009 | 52226 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer ITST Intertrochanteric/subtrochanteric fixation taper reamer - Asia, nonsterile; REF 2258-50-01. (Note: This product is not distributed in the United States). |  |  | June 05, 2009 | 2009 | Zimmer Inc. | Reamers labeled as having a standard diameter may actually be the dimensions of an Asia reamer, and vice versa. | Process control | Zimmer sent an Urgent Device Recall letter dated 6/5/09 to all consignees and instructed to locate and return the products. |  | 120 | 120 of both products | Worldwide distribution: USA, China, Germany, Japan, Korea and Thailand. | Terminated | December 22, 2009 | 200 | 120
Z-1704-2009 | 52177 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | IMHS CP, 3.2mm, Tip Threaded Guide Pin, REF71687000, for use with IntraMedullary Hip Screw in orthopedic surgery. |  |  | January 22, 2009 | 2009 | Smith &amp; Nephew Inc | Incorrect depth gauge mark etched onto guide wire pin. Guide pin is too long. | Process design | Urgent Medical Device Recall letter dated January 23, 2009. All affected Smith &amp; Nephew Sales Representatives, Distributors and Hospitals were notified of problem and the recall via Fed Ex and e-mail on 01/22/2009. They were instructed to immediately check their inventory and return the product. |  | 99 | 99 units | Distribution to USA, Switzerland and Australia. | Terminated | August 18, 2011 | 938 | 99
Z-1695-2009 | 52228 | No | LPQ |  |  | Orthopedic | PMA | Exogen 4000+ Ultrasound Bone Healing System, SAFHS, Part Numbers 71034100 and 71034101, Rx only, Smith &amp; Nephew, 1450 Brooks Road, Memphis TX 38116. |  |  | June 05, 2009 | 2009 | Smith &amp; Nephew Inc | Problems with transducer may result in a reduced ultrasound output. | Other | All affected Smith &amp; Nephew consignees and distributors were notified of problem and the recall by letters (dated June 5, 2009 and June 8, 2009) via US mail. Patients were directed to return the instruments; physicians were informed that their patients were being contacted; and distributors were requested to return their stock and to notify anyone to whom the product was distributed to also return the product. Direct questions to Smith &amp; Nephew, Inc. by calling at 1-800-836-4080. |  | 4139 | 4139 units | Worldwide Distribution -- US, Switzerland, Canada, Malaysia, Japan, and UK. | Terminated | November 18, 2010 | 531 | 5030
Z-1703-2009 | 52289 | No | MRA |  |  | Orthopedic | PMA | Trident Acetabular System Surgical Protocol. Instructions for use of hip prosthesis. Literature Number LSP55. |  |  | June 09, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Surgical Protocol has been modified to create separate and distinct surgical protocols, one for the TridentÂż PSL Shell (LSP68) and one for the TridentÂż Hemispherical Shell (LSP69) in order to clarify and highlight the differences in reaming technique. | Labeling design | Stryker sent Urgent - Product Correction - Surgical Protocol letters by Federal Express on June 9, 2009 to all Stryker branches/agencies, hospital Risk Management, hospital Chief of Orthopaedics, surgeons and all foreign consignees. |  | 3702 | 2376 in USA, 1326 internationally | Worldwide distribution. | Terminated | May 08, 2012 | 1064 | 3702
Z-1697-2009 | 49148 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Foot Choice Therapeutic Foot Massager, Model Number YS-322. |  |  | August 07, 2006 | 2006 | Suarez Corporation Industries | Complaints of product causing burns to consumers. | Device Design | On August 9- August 14, 2006, the firm sent a letter to all customers which included an "Important Customer Reminder." The Reminder now listed additional warnings about serious burns occurring regardless of the control settings. Direct questions to Suarez Corporation Industries by calling 1-330-494-5504. |  | 71512 | 71,512 units | Nationwide and Canada. | Terminated | October 13, 2010 | 1528 | 71512
Z-1671-2009 | 52376 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 12 French 3 Lumen 16 cm Large-Bore Central Venous Catheters. Product numbers AK-12123-H, CS-12123-E, MA-12123, and MC-12123. |  |  | May 08, 2009 | 2009 | Arrow International Inc | Wire guide can have difficulty passing through catheter juncture hub or being removed from the catheter. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 5/11/09 to their customers informing them of the problem and the potential consequences. The letter asked customers to check their stock, cease use and distribution, and quarantine the affected products. They are to contact Arrow's Customer Service Department at 1-800-523-8446 for a Return Authorization Number; complete the Recall Acknowledgement &amp; Stock Status Form and fax to 1-800-343-2935; and return any affected product. The letter states that Arrow can provide either replacement with a similar product at no charge or credit customers' accounts. Additional information or clarification can be obtained by calling Customer Service at 1-800-523-8446. |  | 46610 | 46,610 total | Worldwide Distribution -- USA, Belgium, Canada, Cyprus, Czech Republic, Estonia, Finland, France, Germany, Hungary, Indonesia, Ireland, Israel, Italy, Lithuania, Japan, Republic of Korea, Netherlands, New Zealand, Poland, Singapore, Slovakia, South Africa, Spain, Taiwan, Ukraine, and United Kingdom. | Terminated | July 29, 2010 | 447 | 46610
Z-1593-2009 | 51963 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche ACCU-Chek Aviva Meter Kit, Roche Diagnostics, Indianapolis, IN; REF 04893247001 includes the Roche ACCU-CHEK Softclix Lancing Device, Distributed by Roche Diagnostics, Indianapolis, IN. Model 3144844001 (10 lancets). |  |  | May 13, 2009 | 2009 | Roche Diagnostics Corp. | The lancet cap may be missing, which could result in an unintended lancet stick to the user. | Other | An "Urgent Medical Device Correction" letter dated May 13, 2009 was sent to all affected customers, distributors and consumer consignees who have either registered their blood glucose meter, contacted the firm about the lancets since 10/1/08, or have been shipped lancets by Roche since 10/1/08 were notified of the problem. Distributors were requested to notify each patient self-tester to whom a meter kit was distributed since October 1, 2008. The firm issued a press release on May 18, 2009. Diabetes Care customers: If you have questions regarding the recall, please contact the ACCU-CHEKÂż Softclix Lancet Hotline at 1-800-778-7057.CoaguChek meter customers: If you have questions regarding the recall, please contact the ACCU-CHEKÂż Softclix Lancet Hotline at 1-800-778-7505. |  | 449533 | 449,533 packages of lancets | Nationwide. | Terminated | January 05, 2011 | 602 | 449533
Z-1813-2009 | 52439 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Esprit Ventilator System, Model Number V1000 with Optional External Battery; Product Number: 1001456. |  |  | April 22, 2009 | 2009 | Respironics California Inc | The recall was initiated because through routine product monitoring Respironics discovered that under certain conditions, when operating on AC line power, a portion of the current from the External Battery charger is erroneously diverted to the ventilator. When this occurs, the External Battery charger may not transition correctly. Prolonged operation in this condition results in overcharging and | Device Design | Respironics California, Inc. issued an "Urgent - Medical Device Field Correction" notice dated April 22, 2009. Customers were informed of the affected device including corrective actions taken by the firm. A Respironics representative will contact customers to schedule service of the device.Further questions can be addressed by contacting Respironics US Customer Service at 1-800-345-6443. |  | 508 | 508 | Worldwide Distribution: United States, Hong Kong, India, Mexico, South Africa and Vietnam. | Terminated | July 13, 2012 | 1178 | 508
Z-1815-2009 | 52523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Phoenix Infusion System with MedNet Software, Symbiq One-Channel Infusion System, Software versions v2.1 &amp; v3.0, Manufactured by Hospira Inc., Morgan Hill, CA. |  |  | March 31, 2009 | 2009 | Hospira Inc | Potential delay/underinfusion of critical therapy-- Devices experience increased frequency of alarms that result in temporary pump stoppage. | Software design | Hospira, Inc. issued an "Urgent: Device Field Correction" dated March 31, 2009 via Federal Express to Healthcare Professionals and Valued Hospira Customers. Consignees were informed of the affected devices and will be contacted by a Hospira representative to arrange necessary device upgrades.For further questions, contact Hospira Global Product Safety and Complaints at 1-800-441-4100. |  |  |  | Product was distributed to 59 consignees to all 50 States. | Terminated | December 18, 2010 | 627
Z-1731-2009 | 51998 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Cardiac Troponin I Flex reagent cartridge (CTNI). |  |  | April 08, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | possibility of falsely elevated results | Process control | Siemens Healthcare Diagnostics issued a letter to their customers starting April 8, 2008 informing them of the problem and to discontinue use of the product. Further questions should be addressed to Siemens Healthcare Diagnostics, Inc. at 1-800-441-9250. |  | 4574 | 4,574 cartons | Worldwide Distribution -- Including countries of the United States, Bahamas, Belgium, Czech Republic, France, Germany, Italy, Poland, Portugal, Saudi Arabia, Spain, South Korea, and Slovakia. | Terminated | August 13, 2009 | 492 | 4574
Z-1811-2009 | 52601 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 60ml Syringe Luer-Lok tip, Sterile, BD, Franklin Lakes, NJ 07417, Made in USA. |  |  | July 07, 2009 | 2009 | Becton Dickinson &amp; Company | Insufficient Seal: Becton Dickinson received complaints of insufficient/open seals on one lot of 60 ml Luer-Lok Syringes. | Packaging process control | Becton Dickinson issued an "Urgent: Product Recall" notice via UPS tracking dated July 7, 2009 to all distributors informing them of the affected product. The firm asked all distributors to identify end customers in an attempt to notify all customers. The firms asks that all affected product be returned.Further questions should be directed to BD Customer Service representatives at 1-888-237-2762. |  | 133600 | 133,600 units | Nationwide Distribution -- Including states of IL, CA, VA, PA, and FL. | Terminated | October 02, 2009 | 87 | 133600
Z-1878-2009 | 52708 | No | KKO | Teething ring. | Dental | Dental | 510(K) Exempt | Gel Filled Teethers, various shapes and sizes, marketed under the names: Nuby, Cottontails, and Playskool. |  |  | July 17, 2009 | 2009 | Luv N' Care Ltd | Gel filled teethers were found to contain to contain Bacillus circulans and Bacillus subtilis. | PMA | The firm initiated their recall by telephone and email, and followed with letters to their consignees directing a recall to the consumer level. Consumers are urged to stop using the affected product and discard or return to the place of purchase. Consumers with questions about the recall may contact Luv N' Care, LTD by calling 1-800-256-2399. The firm issued a Press Release on July 17, 2009.Press Release about this recall is located on the website address: http://www.fda.gov/Safety/Recalls/ucm172689.htm. |  | 546320 | 546,320 units | Worldwide Distribution -- US (states of NY, IA, WI, IL, FL, NE, OH, MI, TX, NY, RI, and VA), Canada, Haiti, West Indies, Surinam, Antilles, Uruguay, Mexico, Bermuda, Canada, Dominican Republic, Dutch West Indies, Bahamas, North Antilles and Paraguay. | Terminated | February 28, 2013 | 1322 | 546320
Z-1685-2009 | 50097 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident Hemispherical Acetabular Shell; 72mm, Arc Deposited, Hydroxylapatite Coated, Multi holed; Sterile; Stryker, France. |  |  | May 29, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that Trident Hemispherical Shells, Catalog Number 508-11-74J, are actually marked and labeled as Catalog Number 598-11-72J, Lot Code 20282301 | Labeling mix-ups | Stryker Orthopaedics issued an "Important Market Withdrawal" notice via Federal Express dated May 29, 2007. Consignees were asked to locate and return affected product and fax the attached Product Accountability Form to 1-201-831-6069.For further questions, contact Stryker Orthopaedics at 1-201-831-5825. |  | 2 | 2 units | Distributed to 2 Stryker branches. | Terminated | August 17, 2009 | 811 | 2
Z-1734-2009 | 52079 | No | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks AutoPrep System for in vitro diagnostic use, software version 2.2 or greater. Distributed by Veridex, 33 Technology Drive, Warren, NJ 07059. Manufactured by Immunicon, Huntingdon Valley, PA. Immunicon was acquired by Veridex in August of 2008. |  |  | May 06, 2009 | 2009 | Veridex, LLC | Bottle degradation: Investigation has determined that degradation may occur in the plastic waste bottle of the CellTracks AutoPrep System. | Device Design | Veridex, LLC issued a recall notice to consignees dated March 12, 2009 detailing the affected product. The firm will provide replacement bottles to all customers in the near future.For further questions, contact Veridex Customer Technical Services at 1-877-837-4339. |  | 196 | 196 instruments | Worldwide Distribution | Terminated | October 05, 2010 | 517 | 196
Z-1720-2009 | 52022 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, CT Perfusion 4 on Advantage Workstation, CT and PET/CT consoles. |  |  | May 04, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of the need to adjust default settings associated with the use of CT Perfusion 4 on the Advantage Workstation, ST and PET/CT consoles that may impact patient safety.The default parameter settings in CT Perfusion 4 may not be optimal in the case of data acquired with longer time sampling intervals such as axial scans with inter scan delay (ISD) greater th | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 23, 2009. The letter was addressed to Hospital Administrators / Risk managers and Radiology department mangers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information instructing the user to follow the enclosed User Guide Addendum to mitigate further problems.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan); for other countries, contact your local GE Healthcare field service representative. |  | 444 | 444 | Worldwide Distribution -- Including states of AZ, CA, CT, DC, FL, HI, IL, IA, MA, MI, NJ, NY, NC, OK, PA, PR, TX, WA, WV and WI and countries of ARGENTINA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, PAKISTAN, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM and VENEZUELA. | Terminated | May 16, 2012 | 1108 | 444
Z-1735-2009 | 52602 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, BostonÂż Conditioning Solution, Original Formula, For Rigid Gas Permeable Contact Lenses, 1 fl. oz. (30 ml), Bausch &amp; Lomb Incorporated, Rochester, NY 14609. |  |  | June 30, 2009 | 2009 | Bausch &amp; Lomb Inc | The product failed to meet one of the shelf life specifications at the end of the expiry period. | Other | Bausch &amp; Lomb issued an "Urgent Voluntary Medical Device Recall" notice dated June 26, 2009. All consignees were requested to review their inventory, and return any of the affected product to Stericycle Pharmaceutical Services, Indianapolis, IN. A Business Reply Card was included to be returned to Bausch &amp; Lomb via US mail. A prepaid UPS Packing Slip was also provided to return the affected product.For further questions, contact Bausch &amp; Lomb Customer Resource Center at 1-800-828-9030. |  | 56000 | 56,000 bottles | Nationwide Distribution. | Terminated | May 10, 2012 | 1045 | 56000
Z-1719-2009 | 51120 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare MUSE Cardiology Information Systems vs. 7.0, 7.0.1, 7.0.2, 7.1, 7.1.1. |  |  | February 02, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of an issue resulting from a workflow situation related to the automated labeling of patient data in the MUSE Cardiology Information System. When a prior record has been confirmed with an invalid (default) patient identifier (PID), MUSE will accept future records with the same default PID as correct. If subsequent records don not contain a name, MUSE will pre-popu | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 13, 2008, addressed to Healthcare Administrator/Risk Manager, Cardiology Administrators and Head of Cardiology. The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Background Information and Contact Information.For further questions, contact GE Healthcare at 1-800-558-7044. Outside the US, please contact your local GE Healthcare Service Representative. |  | 327 | 327 | Worldwide Distribution -- Including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WI, WY and DC and countries of UNITED KINGDOM , TAIWAN, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, QATAR, NEW ZEALAND, NETHERLANDS, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CANADA, BELGIUM, AUSTRIA and AUSTRALIA. | Terminated | January 16, 2014 | 1809 | 327
Z-1670-2009 | 51138 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Spine Expedium SFX Cross Connector System Torque Driver Shaft, Non-Sterile, REF: 2894-10-300. |  |  | February 11, 2009 | 2009 | DePuy Spine, Inc. | Tip of Driver may break during use | Nonconforming Material/Component | DePuy Spine notified direct accounts by visit and distributors by letter titled "Urgent - Voluntary Product Recall Notification". Accounts are requested to return the notification fax form and all affected product.For further questions, contact Depuy Spine at 1-800-227-6633. |  | 117 | 117 units | Nationwide Distribution -- Including states of MD, PA, NC, MO, NV, AZ, WI and CT. | Terminated | July 09, 2015 | 2339 | 117
Z-1730-2009 | 51722 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | GE Healthcare, Drapes and Sterile Covers |  |  | August 29, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has found that the sterile products manufactured by Contour Fabricators, Inc. (CFI) for GE healthcare (GEHC) may have sterile seals that could be compromised during the shelf life of the products. This issue could result in the product not maintaining its sterile state. | Release of Material/Component prior to receiving test results | Consignees were sent 8/29/08 a GE Healthcare "Urgent Recall Notice" dated August 15, 2008. The letter was addressed to Director / Manager of Radiology and Director of Surgery. The letter described the problem and provided a list of products, solutions and requested the return of the attached fax forms. In the event of any questions or concerns, a service team should be contacted at 800-874-7378 option #8. |  | 137689 | 137,689 (123,373 US; 14,316 OUS) | Worldwide Distribution -- USA, Canada, Australia, Bahamas, China, Hong Kong, Italy, Japan, Korea, Mexico, Malaysia, Singapore, and El Salvador. | Terminated | May 15, 2012 | 1355 | 137689
Z-1733-2009 | 52111 | No | JSY | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Middlebrook 7H10 Agar, Ref R01600, packaged 10 units/package. The firm name on the label is Remel, Lenexa, KS. |  |  | May 06, 2009 | 2009 | Remel, Inc | The plate labeling is incorrect, but the outer box labeling is correct | Error in labeling | The recalling firm issued Important Medical Device Customer Recall Notice letters dated 5/6/08 via regular mail explaining the reason for recall and requesting all remaining inventory of the recalled lot be discarded. A Product Inventory Checklist was enclosed for completion to report the amount of inventory destroyed. A self-addressed postage-paid envelope was enclosed for return of the Product Inventory Checklist. Customers are to contact the firm's Technical Services Department at 800-447-3641 if they have inquiries concerning the letter. |  | 118 | 118 packages | Nationwide Distribution -- including states of FL, NJ, IN, FL, OH, KY, TX, and WI. | Terminated | October 15, 2009 | 162 | 118
Z-1718-2009 | 51322 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092. |  |  | February 02, 2009 | 2009 | GE Medical Systems, LLC | It was reported by a customer site that when using Advantage Sim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated March 11, 2009. The letter was addressed to AdvantageSim Customers, Hospital Administrators/ Risk Mangers and Oncology Department Mangers. The letter addressed the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan) and for other countries, contact your local GE Healthcare sales or field service representative. |  | 230832 | 687 units (351 US / 336 OUS) | Worldwide Distribution -- Including states of AL, AZ, CA, CO, CT, FL, GA, DC, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS,MO, MT,NE, NH, NJ, NM, NY, NC, ND, OH, OK PR, PA, SC, SD, TN, TX, UT, VA, WA and WI and AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TOKELAU, TURKEY, UNITED ARAB EMIRATES and the UNITED KINGDOM. | Terminated | December 10, 2011 | 1041 | 230832
Z-1714-2009 | 52457 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | SJM Peel Away Introducer, REF 405153, Lot 2694252, Length 14 cm, 7F (2.33 mm), Max Guidewire o.d .038" (.965 mm), contents 5, 2011-08, Manufacturer St Jude Medical 14901 DeVeau Place, Minnetonka, MN 55345-2126 USA, Sterile EO, Rx only. |  |  | June 24, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Peel Away Introducers . Isolated lots could contain some items that have been mis-labelled with the incorrect French size of the outer box label. Some 9F products have been labelled as 7F and vice versa. Only the outer box label is affected by this error, the labels on the individual product pouches correctly identify the product. | Labeling mix-ups | The Consignee was provided with a St. Jude Medical "Urgent Product Information" letter and a doctor Acknowledgement form, via e-mail on June 2009. The letter described the problem and the product involved and advised to isolate and discontinue use of the product involved.The Consignee will notify the end user and remove any remaining affected units and will notify St Jude Medical of the results. |  | 200 | 200 | OUS only: Belgium. | Terminated | March 29, 2012 | 1009 | 200
Z-1812-2009 | 52160 | No | LIO | Microbiological specimen collection and transport device. | Microbiology | Microbiology | 510(k) | Remel BactiSwab NPG Collection and Transport System, Sterile, 100 swab units/box, Catalog 12300. The firm name on the label is Remel, Lenexa, KS. |  |  | May 07, 2009 | 2009 | Remel, Inc | Incorrect expiration date on the product | Incorrect or no expiration date | The recalling firm issued Important Medical Device Product Recall Notice letters dated 5/7/09 via regular mail to their direct accounts explaining the reason for recall and requesting they discard all remaining units of the affected lot number in their inventory. A Product Inventory Checklist was enclosed to be returned in the enclosed self-addressed, stamped envelope to report the amount of product that was discarded. The Technical Services Department should be contacted at 800-447-3641 if there are any inquiries concerning the letter. |  | 103 | 103 packages | Nationwide Distribution | Terminated | January 15, 2010 | 253 | 103
Z-1821-2009 | 52525 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Dall Miles Cable Grip System; Catalog number 6704-0-210; Stryker Ireland. |  |  | June 18, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Sterilization issue: During post-sterilization inspection of the outer blister packaging of the affected lots, it was determined that the outer blister did not meet the required minimum seal width. | Process control | Urgent Product Recall letters were sent via Federal Express on June 18, 2009 to all Stryker Branches/Agencies, hospital Risk Management, hospital Chief of Orthopaedics and surgeons who may have used the recalled products. The letters stated the issue, potential hazards, and the risk mitigation. Customers may contact Colleen O'Meara at 201-831-5970 with any questions. |  | 13 | 13 total, (both products) | Nationwide Distribution | Terminated | August 27, 2012 | 1166 | 13
Z-1818-2009 | 52379 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | T2 Ankle Arthrodesis Nail Surgical Protocol, Catalog Number LT2AA-OT, Stryker Orthopaedics |  |  | November 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Operative Surgical Technique was found to have a typographical error where hyphens were inadvertently omitted in incision lengths (i.e. 1520 cm should be 15-20 cm). | Labeling Change Control | Important Product Correction Letters, dated November 17, 2006, were sent to all Stryker Branches/agencies via Federal Express. The letter stated the issue and asked consignees to examine their inventory and contact their sales reps. It also told them to destroy the affected Operative Techniques by cutting them in half and discarding them, complete the attached Acknowledgment Form, and fax a signed copy of that form to 201-831-6069. The letter stated that Stryker Orthopaedics Customer Service reps would place consignees' re-orders for the product being destroyed. Also, the letter stated that in the interim a correct revision of the Operative Technique could be found online. Customers are to contact Dave O'Dell at 201-831-5277 or Rita Intorrella at 201-831-5825 with any questions. |  | 1270 | 1270 | Nationwide Distribution | Terminated | August 20, 2009 | 1007 | 1270
Z-1933-2009 | 52809 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator, Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 66112-5798, USA. |  |  | July 20, 2009 | 2009 | Boston Scientific CRM Corp | Boston Scientific Cardiac Rhythm Management retrieved devices that could be subject to a potential for acute non-secure lead connections when implanted. Nonsecure lead connections can manifest as high impedance, or lead electrogram artifacts resulting in oversensing and/or inappropriate delivery of therapy. | Device Design | Boston Scientific Corporation field personnel began retrieving devices containing the original set screw system on Friday, July 17, 2009. If needed a letter, dated 07/17/2009 was left with the hospital administrator letting them know that product has been collected for return to Boston Scientific Corporation. Replacement devices will be provided at no charge. If the field personnel was unable to contact the hospital in the time provided, a notification letter dated July 22, 2009 was sent via Federal Express to the hospital. This letter was directed to the EP Catheter Lab Manager with a cc to the Hospital Risk Manager, Hospital Administrator and Purchasing and Central Supply. The letter asked that the affected devices not be implanted and to contact their local representative for pick-up and return. The firm will replace the affected devices. Please direct questions to your local Sales representative or Boston Scientific Corporation CRM Technical Services by calling 1-800-227-3422. |  | 104 | 104 | U.S. Distribution only (states of AZ, CA, CT, FL, GA, IA, IL, KY, MD, MO, NH, NJ, NY, OH, PA, SC, TN, TX, WI, and Washington D.C.) and Puerto Rico. | Terminated | June 26, 2012 | 1072 | 149
Z-1954-2009 | 52853 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Virtuoso VR, Model D154VWC, Single Chamber implantable cardioverter defibrillator with atrial therapies. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | July 27, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic has detected a specific pattern of MOSFET IC malfunctions in its Concerto, Virtuoso and EnRhythm family of devices. The probability of occurrence decreases with time and, to date over 90% of the malfunctions related to the pattern have occurred within the first twelve months after implant. | Process control | Medtronic, Inc. personnel retrieved 2 non-implanted devices remaining at 2 hospitals. A Performance Notes letter was left at each facility. The Performance Notes described the issue, actions to address issue and recommendations to physicians. The Performance Notes will also be posted on the firm's website. Direct questions about this recall to Medtronic, Inc. by calling 1-763-514-4000. |  | 1 | 1 | Nationwide Distribution-- (states of IA and NJ). | Terminated | May 31, 2012 | 1039 | 1
Z-1694-2009 | 51956 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Capless Li Pedicle Screw System, Diameters - 4.75mm, 5.5mm, 6.5mm, 7.5mm, 8.25mm, Lengths 30mm-55mm, labeled non-sterile. |  |  | March 24, 2009 | 2009 | X Spine Systems Inc | A defect cause was discovered following a customer complaint pertaining to a post operative dissociation of the Capless Li Screw construct. The screw assembly is made up of three parts; a cup which locks onto a 5.5mm rod, a yoke that cradles the rod during fusion, and a screw that attached to the vertebra. The screw ball on the screw was manufactured out of tolerance (undersized). | Process control | X Spine Systems Inc. issued advisory notices in March 2009 to Consignees notifying them of the problem and requesting all affected product be returned to the firm. The firm also conducted telephone calls and emails for the return of the product.For further information, contact X Spine Systems Inc. at 1-937-847-8400 extension 115. |  | 9613 | 9,613 units | Worldwide Distribution -- Including states of AZ, CA, CO, FL, ID, IN, MO, MS, NV, OK, TN, TX, UT and WY and countries of Spain, South Africa and the United Kingdom. | Terminated | April 25, 2012 | 1128 | 9613
Z-1736-2009 | 52452 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump catheter and split introducer needle, sterile, Stryker Instruments, Kalamazoo, MI; REF 500-23. |  |  | June 16, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The catheter may break when removed from the body, or the catheter may have kinks or be occluded such that the patient will not receive the intended quantity of medication. | Nonconforming Material/Component | Stryker Instruments issued a recall letter dated June 16, 2009 and instructed Consignees to cease using and destroy the product and to return a business reply form for credit via fax to Stryker Instruments Regulatory Department, 1-866-521-2762.For further questions, contact Stryker Instruments at 1-800-800-4236 extension 3808. |  | 268 | 268 | Worldwide Distribution -- Including the United States, Argentina, Australia, Brazil, Canada, England, France, Germany, Hong Kong, Italy, Japan, Latin America, Malaysia, Mexico, Netherlands, New Zealand, Singapore, South Africa, Spain, Sweden, Switzerland and Taiwan. | Terminated | October 27, 2010 | 498 | 47091
Z-1958-2009 | 50783 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Reflex Hybrid Level Anterior Cervical Plate; |  |  | June 14, 2006 | 2006 | Stryker Spine | Design Flaw: The design did not prevent over angulation of the screws when the recommended instrument, All in One Guide, was not used. | Device Design | Stryker Spine issued an "Important Market Withdrawal" notice dated June 14, 2006. Consignees were advised to complete and return a Product Accountability Form via fax to 1-201-831-6069 prior to returning the affected product. Product return was to be performed during certain phase periods when upgrades were being shipped.For further questions, contact Stryker Spine at 1-201-831-5825. |  | 38987 | 38987 plates | Nationwide Distribution -- Including states of NC, OR, AZ, TN, CA, MA, MD, IL, GA, FL, NJ, NM, KY, LA, PA, MI, UT, MO, OK, KS, OH, MN, NE, IN, MT, TX, AL, CO and WI. | Terminated | August 20, 2009 | 1163 | 38987
Z-1721-2009 | 51317 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Foot Lok Cusion, Part Number MTAFS01, reusable nonsterile (20.3 x 42 x 25.4cm) cushion. |  |  | March 06, 2009 | 2009 | Med Tec Inc | The silkscreened indexing scale was reversed. | Error in labeling | Med Tec inc. (CIVCO Medical Solutions, Radiation Oncology) issued a "Recall Notification" notice dated March 6, 2009 via certified mail explaining the reason for recall and providing the customer a choice of returning the product for replacement or not take any action. A response form was enclosed to be returned via FAX to indicate the number of units in their possession and whether or not they are being returned. Telephone calls were initiated April 20, 2009 to nonresponders or for any consignee whose letter was returned as undeliverable.Foreign customers were notified by e-mail or phone beginning March 16, 2009.For further information, contact CIVCO Medical Solutions, Radiation Oncology at 1-800-842-8688. |  | 231 | 231 units | Worldwide Distribution -- Including the United States, Czech Republic, Colombia, Thailand, Morocco, Brazil, the United Kingdom, India, Costa Rica, The Netherlands, Russia, Spain, Israel, Taiwan, Mexico, Turkey, Korea, Australia, Malaysia, Italy and Japan. | Terminated | November 05, 2009 | 244 | 231
Z-1732-2009 | 52018 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Hemochron Jr. ACT- LR Assay (JACT- LR); |  |  | May 05, 2009 | 2009 | International Technidyne Corp. | One lot of ACT-LR was released when one finished good testing criteria acceptance limit did not meet acceptance criteria. | Employee error | An "Urgent - Medical Device Recall" letter dated May 4, 2009 was sent to healthcare professionals. The recall letter described the reason for recall and the possible health risks involved. Consignees were asked to check their inventory for the affected lot number and return the completed Customer Account Tracking form via mail, e-mail or fax to ITC. ITC will make arrangements for the return of affected product. If there is no affected product in the inventory and no product or safety concerns, the letter is for information purposed only. Note: the Customer Account Tracking form should be completed in all cases.If you have questions regarding the recall information, please call ITC Technical Support at 1-800-579-2255, or email ITC at JACTLR@itcmed.com. |  | 7785 | 173 boxes of 45 cuvettes | Nationwide. | Terminated | September 30, 2009 | 148 | 7785
Z-1686-2009 | 52088 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DePuy P.F.C. Sigma Knee System, Non-Porous Cruciate Retaining Femoral Component, 73 mm M/L 69 mm A/P, 5 left, sterile, DePuy Orthopaedics, Inc., Warsaw, IN; REF 96-0005. |  |  | May 07, 2009 | 2009 | Depuy Orthopaedics, Inc. | There may be a crack on the lateral side of the condyle in the posterior chamfer region.The device is used as a Orthopedic knee implant. | Process control | The firm's sales staff was notified by email on May 7, 2009, and the recall was expanded to include additional lots on May 29, 2009, instructing them to remove the affected lots from their inventory and the inventory of any consignees, to provide the consignees with a copy of the recall letter, and to return the product to the firm. Hospital consignees were notified by letter dated May 7, 2009, as well as by updated letter on May 29, 2009. A dear doctor letter dated May 29, 2009 was issued to implanting surgeons advising them of the issue. Direct questions to Depuy Customer Service by calling 1-800-366-8143. |  | 260 | 260 | Nationwide. | Terminated | October 25, 2010 | 536 | 260
Z-1724-2009 | 52057 | No | CCN | Bourdon gauge flowmeter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Amvex Integrated Flowmeter; a nonback-pressure compensated Bourdon gauge flowmeter; Amvex Corporation, 25B East Pearce St, Richmond Hill, Ontario L4B 2M9, Canada; Model No. FI. |  |  | May 29, 2009 | 2009 | Amvex Corporation | The affected flowmeters may have a damaged o-ring which may cause the flowmeter to deliver higher amounts of gas and or not be able to shut off fully. | Employee error | An "Urgent: Medical Device Recall" letter dated May 29, 2009 was issued to the affected direct accounts on the same date, informing them that the Integrated Flowmeters shipped between April 12, 2007 and August 31, 2007 may experience damage to the o-ring which helps to control the flow of gas from the flowmeter portion of the outlet. Units with a damaged o-ring might deliver higher amounts of gas and/or not be able to shut off fully. The letter provided a list of the affected serial numbers, and asked the accounts to confirm that they have flowmeters from the recall time period. The accounts were provided with interim instructions to use the flowmeters during the replacement period, and were instructed to call an Amvex Corporation Representative at 1-866-462-6839 with at least one of the affected serial numbers and provide the number of units that need replacement. The Amvex Corporation will send new flowmeters for replacement, with instructions on how to replace, record the serial numbers and return the recalled units. If the accounts distributed the units, they were requested to forward the recall letter to the end users. The accounts were also requested to complete the enclosed reply sheet acknowledging receipt of the letter and fax it back to the Amvex Corporation at 1-905-764-7743. Direct questions to the Amvex Corporation by calling 1-866-462-6839. |  | 1424 | 1,424 units | Worldwide Distribution -- US (including Puerto Rico), Canada, Iran, Peru, Taiwan and Thailand. | Terminated | August 19, 2011 | 812 | 1424
Z-1712-2009 | 52633 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Twin-PassÂż Dual Access Catheter , REF 5200, sterilized with Ethylene Oxide, RX Only, Non-Pyrogenic, Manufactured by Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | July 02, 2009 | 2009 | Vascular Solutions, Inc. | Three occurrences of Twin-PassÂż catheters containing a piece of PTFE beading material in the OTW lumen have been reported to Vascular Solutions. The PTFE beading material is a processing aide used during hydrophilic coating. The beading is to be removed followed by a visual inspection prior to packaging | Process control | Vascular Solutions, Inc. issued an "Urgent: Medical Device Recall - Lot Specific" notice dated July 6, 2009. The letter described the problem and product involved and asked Consignees to immediately check their Twin-PassÂż catheter inventory, remove any affected units and complete the Product Recall Inventory Form and return to Customer Service Department.For further questions, contact Vascular Solutions, Inc. at 1-763-656-4300. |  | 5885 | 5885 | Worldwide Distribution -- Including Puerto Rico and states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and DC and countries of CANADA, SPAIN, GREECE, IRELAND, NORWAY, SWITZERLAND, ENGLAND, TURKEY, ISRAEL, GERMANY, AUSTRIA, TAIWAN, FINLAND, SWEDEN, FRANCE, IRELAND and PORTUGAL. | Terminated | April 03, 2012 | 1006 | 5885
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49
Z-1820-2009 | 52582 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION Pressfit Stem, Straight, 20 MM X 160 MM, REF 71424053, Sterile, Smith &amp; Nephew, Memphis TX 38116. |  |  | January 22, 2009 | 2009 | Smith &amp; Nephew Inc | Products were mispackaged. | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of the problem and the recall via email and overnight mail on January 22, 2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on January 22, 2009. They were instructed to immediately quarantine the product for return. |  | 19 | 19 units | Worldwide Distribution -- US, Canada, Australia, Germany, GB, Italy, Switzerland, Spain, and France. | Terminated | February 07, 2012 | 1111 | 35
Z-1722-2009 | 51314 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Signa Ovation 0.35T, Model Numbers: |  |  | November 03, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has identified a potential pinch point hazard on Ovation MR Scanner when using CTL Array XL Body Flex and Opened Body coil. A patient's hand may be pinched between the magnet bore ceiling cover and the CTL Array, XL Body Flex, or Open Body coil if the patient puts his/her hand on the top of the coil. 2/20/09: Recall being expanded to include the OpenSpeed Systems. | Device Design | GE Healthcare issued a "Product Safety Notification" letter dated October 23, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information for the Signa Ovation 0.35T Systems.A second GE Healthcare "Urgent Medical Device Correction" letter dated March 5, 2009 was sent 3/9/09 to consignees which addressed the expanded recall to the Signa OpenSpeed0.7T MR Systems.For further questions, contact your local GE Healthcare sales or service representative. |  | 43736 | 308 (143 US / 142 OUS / 23 TBD) | Worldwide Distribution-- US (including states of AL, ALZ, AR, CA, CO,CT, DE, FL, FA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MT, MO, NE,NH, NV, NJ, NY, NC, OH, OK, OR, PA, PR, SC SD, TN, TX, VA, WA, WV and WI), CHINA, CYPRUS, ARGENTINA, BRAZIL EGYPT, DENMARK, FEDERAL REPUBLIC OF GERMANY, FEDERAL REPUBLIC OF FRANCE, GET 7 D CHINA, GREECE, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, MALAWI, MEXICO, OMAN, PANAMA, PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SPAIN , SWITZERLAND, TURKEY, UKRAINE, UNITED ARAB EMERITUS, UNITED KINGDOM and VENEZUELA. | Terminated | May 24, 2012 | 1298 | 43750
Z-1728-2009 | 51692 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | Hologic Delphi series Bone Densitometer |  |  | April 07, 2009 | 2009 | Hologic, Inc. | System may not turn off at the end of the x-ray and cause additional x-ray exposure | Component design/selection | Hologic issued an Urgent: Medical Device Recall customer notification letter, dated April 17, 2009, by certified mail to inform customers of the problem.With the notification letter, customers will receive the following:-Software version Apex 12.6.2, 2.3.2, or 2.4.2-Instructions to upgrade their systems with the provided software and to verify the new version on the system as a confirmation that the software was successfully loaded.-New Operator's Manual for installations where the operatingcharacteristics of the system have changed due to the software upgrade.For Customers with systems with software versions other than 12.6, 2.3, or 2.4:Customers were asked to upgrade their QDR software, return their old QDR software disk, and complete and return to Hologic the enclosed Confirmation of Notification. |  | 1914 | 1914 units total | Worldwide Distribution -- USA, Canada, Europe, Austria, Belgium, Bulgaria, Czech Republic, Denmark, France, Germany, Great Britain, Greece, Hungary, Iceland, Italy, Latvia, Lithuania, Macedonia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Serbia, Slovak Republic, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine, Uzbekistan, Algeria, Arab Emirates, Argentina, Australia, Bahrain, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, Hong Kong, Indonesia, India, Iran, Iraq, Israel, Japan, Jordan, Kenya, Korea, Kuwait, Lebanon, Malaysia, Morocco, Mexico, Pakistan, Peru, Puerto Rico, South Africa, Saudia Arabia, Singapore, Syria, Taiwan, Thailand, Tunisia, Venezuela, and Vietnam. | Terminated | November 29, 2010 | 601 | 1914
Z-1902-2009 | 52524 | No | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | CLINITRONÂż CII Air Fluidized Therapy Units. |  |  | June 24, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | Incorrectly repaired power cord wires can potentially overheat resulting in smoldering sound foam in the base of the therapy bed. | Equipment maintenance | Hill-Rom issued an "Urgent Field Safety Notice" letter dated February 2, 2009 addressed to Facility Risk Managers and Facility Administrators. The notice detailed the issue and instructed users on actions to be taken.For further information, please contact Hill-Rom Customer Service at 1-800-638-2546. |  | 625 | 625 units (total both units) | Worldwide Distribution -- United States, Abu Dhabi, Brazil, Canada, Chile, Columbia, Mexico, Newfoundland, Puerto Rico, Saudi Arabia, Singapore, Taiwan, Turkey and Venezuela. | Terminated | July 11, 2012 | 1113 | 625
Z-1906-2009 | 51336 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Lab - MSMS Data Suite, Product Code: 5002-0310. |  |  | January 15, 2009 | 2009 | Perkin Elmer | If Result Codes are used and if they are manually edited, they are correct unless the laboratory specifically uses the View Assay button in theSpecimen History screen to open another assay, and then returns back to the original assay | Software design | Perkin Elmer began notifying US customers January 15, 2009. Notification included instructions for how to avoid the situation and instructions for how to obtain a correction. The firm requested consignees to send in a Response Form by fax (358-2-2678-357) or by email (TurkuQMresponse@perkinelmer.com).For further information, please contact your closest Specimen Gate support center at specimen.gate.support@perkinelmer.com for North America and specimengatesupportfi@perkinelmer.com for outside of North America. |  | 13 | 13 units | Worldwide Distribution -- United States (AL, AZ, CA, CO, CT, FL, GA, IN, KS, KY, MI, MO, OH, OK, OR, TN and TX) and Canada. | Terminated | September 09, 2011 | 967 | 39
Z-1913-2009 | 52629 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Edwards LifeSciences, Percutaneous Sheath Introducer Set, Rx only, For single use only. |  |  | May 11, 2009 | 2009 | Edwards Lifesciences, Llc | Edwards LifeSciences LLC received several complaints, in which the blue connector of the introducer assembly has broken (fracture/separation) after several hours of use. | Nonconforming Material/Component | Edwards LifeSciences LLC initiated the worldwide recall via customer letters dated May 11, 2009 and are being sent by Federal Express stating that customers cease using affected product from the indicated lots and return all unused products. The firm requested that the customer provide written verification of stock and identify any unused product to be returned. For further information, contact Edwards LifeSciences LLC Customer Service at 1-800-424-3278. |  | 67357 | 67,357 Units | Worldwide Distribution: United States (AL, AR, CA, CO, CT, DC, FL, GA, IA, IL, IN,KS, KY, MA, MD, ME, MI, MO, MS, ND, NH, NJ, NM, NV, NY, OH, PA, RI, SC, TN, TX, VA, WV, HI and PR), Austria, Canada, Switzerland, Czech,republic, Germany, Denmark, Algeria, Estonia, Egypt, Spain, Finland, France, Great Britain, Hungary, Iceland, Ireland, Italy, Kuwait, Lithuania, Luxembourg, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Sweden, Slovak republic, Tunisia, Turkey, Aruba, Hong Kong, Malaysia and Singapore. | Terminated | February 23, 2012 | 1018 | 67357
Z-1955-2009 | 52718 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Stabilet Infant Warming System, model numbers 200, 300, 1250, 1500, 200/3000, 2000, 2200/3200, 3000, and 3200. |  |  | July 17, 2009 | 2009 | Draeger Medical Systems, Inc. | Hot particles fell into the infant bassinet beneath a Radiant Infant Warmer. The resulting fire injured an infant. These devices are obsolete. | Equipment maintenance | On 7/20/09, Draeger Medical, the recalling firm, issued a press release to medical media outreach organizations and an Urgent Medical Device letter to all known customers. This information explained the incident and the need to remove and replace the units in question as soon as feasibly possible. |  | 6489 | 6489 units | Worldwide distribution: USA, Brazil, Canada, China, Dominican Republic, France, Korea, Lebanon, Mexico, Nigeria, Panama, Poland, Saudi Arabia, Taiwan, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | June 01, 2011 | 684 | 6489
Z-1934-2009 | 52802 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences "VAMP Jr" (Venous Arterial blood Management Protection system), Lot numbers (all product manufactured between 8/01/06 and 2/29/08.) |  |  | March 14, 2008 | 2008 | Edwards Lifesciences AG | Edwards Lifescience initiated a recall of on Edwards "VAMP Jr" Pressure Monitoring Products manufactured between 8/01/06 and 2/29/08 due to the potential for the cap on the reservoir of the "VAMP Jr" to become disconnected. | Equipment maintenance | Edwards initiated the recall in the U.S. via a Urgent Customer lNotification letter with attached response confirmation fax-back attachment dated March 10, 2008 and was sent by Federal Express on March 12, 2008 requesting that customers cease using all "VAMP Jr" products and that they return all unused product.Customers were instructed to fax-back the confirmation attachment to 800-422-9329 and contact Edwards Customer service at 888-570-4012 to arrange for return of any unused product and obtain information about replacement product.Edwards Lifesciences requested that the customer provide written verification of stock and identify and unused product to be returned. This will enable a 100% effectiveness check for this recall. For product that was shipped to distributors, contact was made directly from Edwards Lifesciences Corporate Office in Irvine, CA. All product that was distributed internationally will be handled regionally. (i.e. US, Canada, Europe/EMEA, Japan, Intercontinental-including Latin America and Asia Pacific) |  | 61199 | 61,199 units | Worldwide Distribution including USA and the countries of Australia, Austria, Belgium, Brazil, Canada, Canaric Island (Spain), China, Columbia, Czech Republic, Denmark, France, Germany, Hungary, Italy, Japan, Kuwait, Netherlands, New Zealand, Norway, Puerto Rico, Spain, Sweden, Turkey, and United Kingdom. | Terminated | July 03, 2012 | 1572 | 61199
Z-1899-2009 | 52712 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertier Surgical Table, an AC powered, general purpose, mobile surgical table with electro-hydraulically controlled movements which are part of a control feedback system. Part Numbers: 0788100000 Standard table, 0788100001 Standard table with 5th wheel, 0788100000S SIDNE-enabled standard table, and 0788100001S SIDNE-enabled 5th wheel table |  |  | November 13, 2008 | 2008 | Stryker Communications Corp | Surgical tables were produced using an unapproved tilt frame bushing. | Nonconforming Material/Component | Stryker Communications issued an "Urgent: Medical Device Recall" notice dated November 13, 2008 informing Consignees of the affected product. The firm intends to contact each facility to coordinate service by a trained representative.For further information, contact Stryker at 1-972-410-7100. |  | 31 | 31 | Worldwide Distribution -- United States (OR, IL, WA, CA, AR, TX, UT, SC and GA) and Canada. | Terminated | August 28, 2009 | 288 | 31
Z-1915-2009 | 52676 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Stryker Orthopaedics PAL Pelvic Alignment Level, Ref: PAL 400; Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | September 11, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The bubble in a vial of the Pelvic Alignment Level (PAL) was wider than the level marking. Upon opening the PAL and removing the vial, it was discovered that the vial was leaking | Device Design | Stryker Orthopaedics issued an "Important Product Removal" letter dated September 11, 2007 via Federal Express informing consignees of the affected devices. The firm requested that users identify and return any affected product and account for product by returning a Product Accountability Form via fax.For further information, contact Stryker Orthopaedics at 1-201-831-5825. |  | 615 | 615 | Nationwide Distribution | Terminated | October 15, 2009 | 765 | 615
Z-1916-2009 | 52641 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon AR 4:1 Cutting Blocks - Size 1-8; |  |  | February 12, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Triathlon Cutting Blocks did not assemble to the Triathlon Impactor/Extractor Handle. | Process control | Important Product Correction letters were sent to all branches/agencies on February 12, 2008 by Federal Express. The letter asked consignees to examine their inventory and hospital locations to identify the affected product. Customers were asked to perform a functional inspection by assembling the Triathlon Impactor/Extractor Handle to all of their MIS Cutting Blocks. If the cutting block does not assemble, customers are to list the affected lot codes on the Attached Product Accountability Form and fax a copy to (201) 831-6069. The letter states that product should not be returned until further instructions are received. Questions should be directed to (201) 831-5825. |  | 690 | 690 total units - both varieties | Worldwide Distribution | Terminated | October 15, 2009 | 976 | 690
Z-1882-2009 | 52455 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | MR Coaxial 10g x 118mm (VC10118MR). |  |  | March 26, 2009 | 2009 | Bard Peripheral Vas | Some Vacora 10G Plastic Coaxial Cannulas may leave a small metal filament in the breast or the sample obtained post procedure. | Process control | Bard Peripheral Vascular, Inc. issued an "Urgent Medical Device Recall Notification" dated March 27, 2009 informing Customers of the affected devices. The firm requested that all affected product be removed and returned along with returning a Recall and Effectiveness Check Form via fax.For further information, contact Bard Peripheral Vascular, Inc. at 1-800-321-4254. |  | 1440 | 1,440 | Worldwide Distribution. | Terminated | April 02, 2012 | 1103 | 10425
Z-1890-2009 | 52782 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sheridan Preformed Cuffed Tracheal Tube, Nasal 6.5, Catalog Number: 5-22313, Teleflex Medical Ltd, Research Triangle Park, NC 27709 |  |  | July 06, 2009 | 2009 | Teleflex Medical | The Nasal Preformed Tracheal Tubes were found to have incorrect insertion depth mark printings. | Process control | Consignees were notified by an Urgent Medical Device Recall letter on/about 07/06/2009. They were instructed to discontinue use of and return any of the recalled product to Teleflex Medical. Distributors were advised to forward the letter to their consignees to retrieve relevant product from those locations. A Recall Acknowledgement Form was attached to be returned to Teleflex Medical. Replacement product or credit will be provided to the consignee. |  | 1160 | 1160 total | Worldwide Distribution -- USA, including states of CA, FL, IL, KY, MA, NC, NM, NJ, OH, PA, RI, TN, TX, UT, VA, WA, and WI and countries of France and Japan. | Terminated | August 24, 2010 | 414 | 1160
Z-1693-2009 | 52835 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Heater Elements for Resuscitaire Infant Radiant Warmers model numbers RW82, RW82 VHA, WBR82, and WMRW82, 120 volt configuration only. |  |  | July 29, 2009 | 2009 | Draeger Medical Systems, Inc. | Spark, flame and smoke: A mis-positioned crimp on heater element lead wire resulted in flame and smoke aerating from the top of the warmer assembly. | Process control | Draeger Medical issued an Urgent Medical Device Recall letter via certified mail on 8/7/09. The letter explained the problem, the need to conduct an inspection of the infant warmer (instructions included), and request replacement elements. |  | 2915 | 2915 elements | Worldwide distribution: USA, Canada, Saudi Arabia, Bermuda, Virgin Islands, Mexico, Panama, Colombia, and Costa Rica. | Terminated | October 13, 2010 | 441 | 2915
Z-1809-2009 | 52654 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varis, Aria Radiation Oncology - Version 8.1.15, 8.5.11, and 8.6.07, Model number HIT-INFSYS, Manufactured by Varian Medical Systems Inc., Palo Alto, CA |  |  | May 08, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential mistreatment-- due to software issues, retired plans may be used and treated repeatedly. | Software design | An Urgent Medical Device Correction Letter/ Field Safety Letter, dated May 8, 2009, were distributed to all affected users, with a description of the problem and user corrective action steps. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm intends to produce a software fix for the problem. If further clarification is needed, users should contact their Varian Customer Support District or Regional Manager. |  | 1587 | 1,587 units | Worldwide Distribution | Terminated | December 23, 2011 | 959 | 1587
Z-1928-2009 | 52118 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AMPLILINK Software, version 3.1.0 to 3.2.2 with COBAS AmpliPrep. Catalog number 03563383001, catalog number 04855094001, catalog number 04807197001, catalog number 04862392001, catalog number 05093236001, and catalog number 05201357001. |  |  | June 02, 2009 | 2009 | Roche Molecular Systems, Inc. | Software bug: A unique series of events involving the COBAS AmpliPrep instrument running AMPLILINK software version 3.2.2 that results in specimens being amplified and detected with the wrong parameters. The error can occur on either COBAS AmpliPrep/COBAS TaqMan 48, COBAS AmpliPrep/COBAS TaqMan (docked) or cobas s 201 system (docked) running Amplilink Software versions 3.1 or 3.2 series up to and | Software design | Urgent Medical Device Correction Letters were sent on June 2, 2009 by first class mail. The letter described the issue; impact on results; clinical implications; the workaround; and the actions required, which include following the workaround until a new version of AMPLILINK software is available and filing the letter for future reference. Questions or concerns should be directed to Roche Diagnostics Technical Support at 1-800-526-1247. |  | 282 | 282 | Nationwide Distribution | Terminated | August 13, 2010 | 437 | 282
Z-1901-2009 | 52717 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific (1) Flexima Regular Biliary Catheter System with RO Marker, 8 FR (2.7 mm) with Glidex hydrogel Coating Locking Pigtail and Radiopaque Marker, REF /Catalog No. 27-260 UPN; M001272600, Use before 2011-11, Sterilized with Ethylene Oxide Gas, Single Use Only, Made in USA 780 Brookside Drive, Spencer, IN 47460-1080. |  |  | July 01, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall of a single batch/lot of this product. We received a report that the labeling on the device does not match the labeling on the product packaging. Specifically, the packaging pouch for the product and outer carton labeling reflect that the device is a Flexima Biliary Catheter, which is correct, but the heat shrink inscription directly below the hub on the | Labeling mix-ups | Boston Scientific issued an "Urgent Medical Device Recall -Immediate Action Required" letter dated July 1, 2009. The letter was addressed to Risk Manger / Field Action Contact describing the problem and units affected and that further distribution or use of affected product should cease immediately. The Reply Verification Tracking Form enclosed with the letter was requested to be completed and returned.For further information, contact Boston Scientific at 1-508-683-6036. |  | 1 | 79 units (1 box = 1 unit) | Worldwide Distribution -- United States (AZ, AR, CO, DE, FL, GA, MI, MT, NE, NC, OH and WI), Spain, Finland and Netherlands. | Terminated | December 02, 2011 | 884 | 1
Z-1885-2009 | 52761 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS Picture Archiving Communications System InteleViewer module, Model #3.X.X. |  |  | October 27, 2008 | 2008 | Intelerad Medical Systems Inc. | Product software versions 3.5.2.P1 to 3.5.2.p56 inclusive and 3.6.1.P1 to 3.6.1.P82 inclusive were defective. | Software design | Intelerad Medical Systems issued an Urgent Medical Device Filed Safety Notice" dated October 24, 2008 describing the affected product and steps required for system upgrade. Users are requested to complete and return an acknowledgement form via email to quality@intelerad.com.For further information, contact Intelerad Medical Systems at 1-514-931-6222. |  | 17 | 17 (including 11 free trials) | Worldwide Distribution -- United States (NM, AZ, CT, MN, TX, NC, TN, FL and CO), Canada, Australia, and New Zealand. | Terminated | September 24, 2010 | 697 | 18
Z-1817-2009 | 52375 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Lateral Decubitus Alignment Guide, Non Sterile; manual surgical instrument. The MIS Hip Acetabular instrument set includes the curved positioner/impactor, supine alignment guide, and decubitus alignment guide. |  |  | April 27, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The Lateral Decubitus Alignment Guide either did not assemble to its mating part, the Curved Impaction Handle, or partially engages the handle, but does not fully lock. | Nonconforming Material/Component | Important Market Withdrawal Letters were sent to Stryker Branches on April 27, 2006 vis Federal Express. The letter asked consignees to examine their inventory and identify the product. They are to retrieve all affected product lots and return to their branch or agency warehouse for reconciliation. Also, they are to reconcile all product on the attached Product Accountability Form and fax a copy. Affected should be returned and their Stryker Orthopaedics Customer Service Rep should be contacted to re-order the product that is being returned as part of this recall. Questions should be referred to Rita Intorella at 201-831-5825. |  | 61 | 61 units | Nationwide Distribution | Terminated | October 15, 2009 | 1267 | 61
Z-1904-2009 | 52725 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | ConceiveEase (TM), Fertility Friendly Lubricant, Reproductive Lab, Memphis, TN |  |  | July 10, 2009 | 2009 | Memphis Fertility Laboratory Inc | Product was marketed without a 510(k). | PMA | The firm notified its sole consignee on 07/10/2009 by telephone and followed by an Urgent Medical Device Recall letter. The letter requested that they examine their inventory and quarantine all products subject to recall and immediately contact Dr. William Kutteh for return instructions. The enclosed response form should be completed and returned. Questions should be directed to Dr. Kutteh at 901-261-2401. |  | 800 | 800 units | Nationwide Distribution -- USA, including the state of MA. | Terminated | October 05, 2009 | 87 | 800
Z-1729-2009 | 52636 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac with Version 7.x Software, manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | May 29, 2009 | 2009 | Varian Medical Systems Oncology Systems | Unexpected Movement: if the stereotactic motion disable function is turned on, the couch can be moved via float mode unexpectedly. | Device Design | Varian Medical Systems issued an "Urgent Medical Device Correction" dated May 29, 2009 distributed to all affected users, as well as Varian Sales, Marketing and Service organizations, informing them of the issue. Consignees were informed that a Varian Service representative will be in contact to schedule the proper correction.For further questions, contact your local Varian Customer Support District Manager. |  | 1933 | 1,933 units | Worldwide Distribution. | Terminated | May 18, 2011 | 719 | 1933
Z-1945-2009 | 52126 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxilift Patient Lift, Non-AC Powered; a patient lift with a four point tilting frame, adjustable chassis, and optional electronic scale, to move patients in a hospital/nursing home situation; Arjo Inc., Roselle, IL 60172; Model numbers 21102406, 21200001, 21200006, 21000-01, 210024-06, 211010-06, 211024-06, 211026, 212000-06, 212000-06R, 212000-6, 212024-06, 212026, 212624, 218100-01, KGA0200, KMA0520, KMA1003-04, KMA1004, KMA1005, KMA1006, KMA1040, KMA1501-06, KMA1513-06, KMA211024, KMA82124-24, KMB2351, KMB4800-06, KMB4900-06, KMB5501-12US, KRA0300, KRA0310, KRA1000, KRA1500-12, KRA2100, KRA3000-12, MA0500, MA0510, MA1010, MA1020, MA1100, MA1500, Maxi Pillar, Maxilift, Maxipillar, MAXIPILLER, MB0500R, MB0510, and MB0610. |  |  | April 27, 2009 | 2009 | ArjoHuntleigh | The roll pin on certain configurations of the Maxilift hanger bar could break due to a combination of severe corrosion and fatigue. Breakage of the roll pin would lead to the detachment of the hanger bar from the jib of the lifter. | Equipment maintenance | ArjoHuntleigh sent a Field Safety Notice DCHU-UK/0209/001FSN to all affected end users via UPS Ground on 4/27/09. The notice informed the accounts of the potential for the roll pin to break due to a combination of severe corrosion and fatigue, leading to the detachment of the hanger bar from the jib of the lift. The accounts were advised to identify if they have any Arjo Maxilift patient lifters in their facility, identify which version of Arjo Maxilift hanger bar they have using the photographs in Figures 1,2 and 3 in the letter. Any lifters found with a hanger bar attachment as shown in Figure 1 must be immediately and permanently taken out of use, as the replacement hanger bar assembly is no longer available. Any lifters with a hanger bar attachment as shown in Figure 2 must be taken out of use until a correction has been performed by ArjoHuntleigh. Any lifters with a hanger bar attachment design as shown in Figure 3 are not affected and can continue being used. The accounts were provided with a product upgrade offer and a MaxiLift Preventative Maintenance Schedule. The accounts were requested to complete and fax, e-mail or mail the enclosed Customer Response Form to Arjo at 630-925-7974 or quality@arjousa.com, indicating the number of Maxilifts in each category that they have and that the products have been taken out of service for replacement or correction, listing the model numbers and serial numbers of their lifts. Questions were directed to the ArjoHuntleigh Customer Support or Quality Department at 1-800-323-1245. |  | 8015 | 8,015 lifts | Worldwide Distribution -- Countries of USA and Mexico. | Terminated | February 25, 2011 | 669 | 8015
Z-1897-2009 | 52640 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Owner Manual Modifications for PlasmaPro Plasma Pump Model PP-05. Apheresis Technologies, Inc. Palm Harbor, FL. - No labeling or Packaging. |  |  | April 01, 2009 | 2009 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc. is taking the following Corrective Action: Modification of the Owners Manual Only for the PlasmaPro Plasma Pump Model PP-05. Operator's Manual will modify tubeset loading instructions so as to reduce possible loading error which might result in the tube being pulled into the pump head. | Labeling design | PlasmaPro Plasma Pump Operators Manual was modified and distributed to firms consignees. Firms received written notification of re-issuance of Operator's Manual. |  | 7 | 7 Owner's Manuals | Nationwide Distribution -- Florida, Colorado and Utah. | Terminated | July 06, 2010 | 461 | 7
Z-1872-2009 | 52199 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Automated Manifold Kit , REF Model BT2000, SKU # 014613, Sterile R, Made in Mexico, Manufacturer; ACIST LMedical Systems Inc, 7605 Fuller Road, Eden Prairie, MN .55344. |  |  | May 11, 2009 | 2009 | Acist Medical Systems | ACIST Medical Systems initiated a recall due to an increase in the incidence of field reports related to a bulge on the side and/or leaks in the bond of the 2.5 high-pressure tubing of BT2000 kits during a procedure which can result in unsatisfactory performance of the ACIST BT2000 Automated Manifold Kit. | Nonconforming Material/Component | Consignees were sent on 5/11/09 a ACIST Medical Systems "Urgent: Medical Device Recall" letter dated May 11, 2009. The letter described the problem and the product. It advised consignees to discontinue the use or further distribution of the recall product and to return the inventory and the enclosed "Medical Device Recall Form". Please use the enclosed form with firms address: 7905 Fuller Rd, Eden Prairie, MD 55344. Additional questions please contact the firm at 1-877-272-2269, option 1. |  | 62 | 62, 280 (57,110 US, 5130 Canada, 40 Hong Kong) | Worldwide Distribution including the states of AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, ID, IN, IA, KS, KY, LA, MD, MI, MN, MS, MO, NE,NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT VA, WA, and WI.The countries ofCANADA and HONG KONG | Terminated | February 04, 2012 | 999 | 62
Z-1929-2009 | 52055 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor; an automated tissue stainer |  |  | September 15, 2008 | 2008 | Abbott Molecular | Thermal runaway conditions, in that the target temperature was exceeded and the temperature display could not track the actual temperature within the heated slide basins of the VP 2000 Slide Processor. | Device Design | Abbott sent Device Recall/Urgent Field Safety Notice letters dated 9/15/08 to the VP2000 customers on 9/12/08, informing them that thermal runaway conditions may occur within the heated slide basins of the VP2000 Slide Processor. The customers were given actions to take to prevent this condition: a) VP2000 should be powered down when not in use and when the controlling computer is not in use; b) Liquid level in all heated basins should be checked periodically to ensure that there is adequate reagent present for slide processing and to keep the sensor probe immersed for accurate temperature sensing; and c) If large, rapid and random fluctuations in indicated basin temperature occurs the unit should be shut down and a Field Service Representative be contacted. Any questions were directed to their local Abbott Technical Service Representative. |  | 261 | 261 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Columbia, Czech Republic, Egypt, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Malaysia, Norway, Russia, Saudi Arabia, Singapore, Switzerland and the United Kingdom. | Terminated | January 07, 2010 | 479 | 261
Z-1900-2009 | 52722 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ArthroWand Saber 30 with Integrated Cable wand, Catalog number AC4330-01, Manufactured by ArthroCare Corporation, Sunnyvale, CA |  |  | August 05, 2008 | 2008 | ArthroCare Corporation | Potential Sterility Loss-- Due to wearing or puncture in the Tyvek lid, the product may lose sterility. | Package design/selection | Firm initiated this recall on August 5, 2008. Consignees were notified via letter asking them to contact consignees and delivered via courier. Distribution discontinued on July 9, 2008. |  | 23443 | 23,443 units (includes distribution of previously returned product) | Worldwide Distribution including USA and countries of Netherlands, Slovak Republick, Ireland, Spain, Switzerland, Singapore, Russian Federatin, Greece, Japan, Saudi Arabia, India, Slovenia, Israel, South Africa, South Korea, Poland, Thailand, Jordan, Belgium, Finland, Portugal, Cyprus, Italy, UAE, Kuwait, China, Bahrain, Denmark, Norway, Sweden, Australia, Germany, Austria, UK, New Zealand, Mexico and Canada. | Terminated | June 14, 2011 | 1043 | 23443
Z-1990-2009 | 51999 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Combi Sling, 100% Polyester Shell/100% Polyester Strap; soft polyester net fabric sling for ceiling and floor patient lifts; BHM, Magog, Canada, made in China. |  |  | April 07, 2009 | 2009 | B.H.M. Medical, Inc. | Premature failure of the stitching at the junction of the shoulder strap and sling body of the BHM/Medi-Man Combi Sling used with patient lifts. | Device Design | Arjo sent "Urgent Device Recall Customer Notification" cover letters dated April 15, 2009 to the end user accounts who received the affected BHM/Medi-Man Combi Slings, to the attention of the Administrator/Risk Manager, advising them of the potential for possible stitching failure that could lead to a hazardous situation. Enclosed with the letter was a copy of the BHM Medical Field Safety Notice, FSN 10/2009/ca, dated April 3, 2009 which was sent to their BHM customers in the U.S. on April 7, 2009. The FSN identified the part numbers of the BHM/Medi-Man Combi Slings, described the potential health hazard, and listed the actions to take. The accounts were instructed to inspect and assess their slings' suitability for use per the Slings Maintenance Instructions. If the sling is showing this kind of breakage, the account was instructed to remove it from use and discard it immediately. If the inspection reveals no deterioration of the stitching, the sling is safe to use. The accounts were requested to complete and return the enclosed FSN receipt form within 10 days, acknowledging the receipt and understanding of the letter, and indicating the number of slings in use prior to inspection, the number of slings in use after the inspection, and the number of slings rejected. The accounts will be contacted for the replacement of their eligible old slings (in use up to two years). |  | 1396 | 1,396 slings | Nationwide | Terminated | January 07, 2011 | 640 | 1396
Z-1894-2009 | 52560 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Minicap Extended Life PD Transfer Set Extra Short (4") with Twist Clamp; six sets per carton; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 5C4483 The set is used in disconnect applications and cycler applications where aseptic connections and disconnections are performed at the transfer set/cycler juncture. |  |  | July 02, 2009 | 2009 | Baxter Healthcare Corp. | The peritoneal dialysis transfer set has a defective white sleeve twist clamp. The twist clamp will not open or close completely, which could cause leakage and lead to peritonitis. | Nonconforming Material/Component | Baxter Healthcare Corporation telephoned the accounts receiving the affected lots and then faxed or e-mailed "Urgent Product Recall" letters dated July 2, 2009, to them, to the attention of the Peritoneal Dialysis Nurse, to inform them of the defective white sleeve twist clamp, which may result in leakage and possible peritonitis. The accounts were asked to immediately discontinue the use of the affected lot numbers and examine their inventory for the affected lots. If any affected product was found, they were requested to quarantine it and contact Baxter Center for Service at 1-888-229-0001 to arrange for its return to Baxter. The accounts were also requested to review their patient records to determine if any of their patients had received the affected product. If so, they were advised to contact the patients to see if they had experienced any leaking. If so, the patients should return to the facility to have the set changed per the facility's transfer set change procedure. Any questions were directed to the Center for One Baxter at 1-800-422-9837, Monday through Friday, 8:00 a.m. to 5:00 p.m. Central Standard Time (CST). The accounts were requested to complete the enclosed reply form acknowledging receipt of the letter and that they followed the instructions and disseminated the information to their staff. |  | 2436 | 2,436 sets | US (states of Arkansas, Illinois, Mississippi, North Carolina, Tennessee and Virginia) and El Salvador. | Terminated | July 08, 2010 | 371 | 3150
Z-1869-2009 | 52663 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Coronary Perfusion Cannulae with Balloon Tip, with either self-inflating balloon tip or cone distal tips, straight or right angle, available with proximal end fittings for either 1/4 inch connector or male luer fitting, Model numbers 210800, 210801, 210802, 210803, 210804, 210805, 210806, 210807, 210810, 210811, 210812, 210813, 210814, 210815, 210816, 210817, 215800, 215802, 215803, 215804, 215805, 215806, 215807, 215810, 215811, 215812, 215813, 215814, 215815, 215816, and 215817. |  |  | May 19, 2009 | 2009 | Maquet Cardiovascular | Undeclared allergen: Products are missing the warning label declaring that the products contain natural rubber latex. | Process control | Maquet Cardiovascular issued a notification letter to consignees on May 19, 2009. |  | 2226 | 2,226 devices | Nationwide distribution. | Terminated | December 21, 2010 | 581 | 2226
Z-1995-2009 | 51278 | No | MLD | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software. |  |  | February 06, 2006 | 2006 | GE Medical Systems Information Tecnology | GE Healthcare (Datex-Ohmeda) S/5 Critical Care Monitor and S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software had an issue where the Heart Rate (HR) was displaying as ---(3 dashes) when the monitor was used continuously for 49 days without a cold start. A missed alarm or missed HR value could cause a delay in recognition of a cardiac event that may result i | Software design | GE Healthcare Technologies issued a Marketing Bulletin "Ref: MB MS HKI 06007" dated February 6, 2006 to explain a software update. |  | 507 | 507 | ITALY, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRAN, ISRAEL JORDAN, LITHUANIA, ITALY, MACEDONIA, MALAYSIA, MEXICO, NETHERLAND, NICARAGUA, NORWAY, POLAND, PORTUGAL PUERTO RICO, RUSSIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SPAIN, SWEDEN, SYRIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Terminated | May 17, 2012 | 2292 | 2760
Z-1936-2009 | 52427 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Lite Blood Glucose Test Strips, CSP, 50's US, Checkpoint, FS Lite, for use with FreeStyle Lite Blood Glucose Meters and FreeStyle Freedom Lite Blood Glucose Meters, Part Number 70822-02, manufactured by Abbott Diabetes Care, Inc., Alameda, CA. |  |  | June 10, 2009 | 2009 | Abbott Diabetes Care, Inc. | Low results: Freestyle Lite Blood Glucose Test strip may generate low blood glucose readings when tested with the control solution from lot 0822524 with Freestyle Freedom Lite and Freestyle Lite Blood Glucose Meters. | Process control | An "Important: Medical Device Notification" letter dated June 10, 2009 was issued to consignees. First consignees received notifications via letters to both direct consignees and registered customers of FreeStyle Lite and FreeStyle Freedom Lite blood glucose meter users in the states where the affected test strip lot was distributed. Affected accounts and customers were instructed to discontinue distribution and use of the product. The notification will be posted to the Abbott Diabetes Care website. Direct questions about the recall to Abbott Diabetes Care, Inc. by calling 510-749-5400. |  | 9252 | 9,252 vials | US only (states AL, AZ, CA, FL, GA, IL, IN, MD, NJ, NY, OH, PA, RI, SC, TN, TX and WI). | Terminated | January 10, 2011 | 579 | 9252
Z-1914-2009 | 52168 | No | DSQ |  |  | Cardiovascular | PMA | Abiomed AB5000 Portable Driver |  |  | May 08, 2009 | 2009 | Abiomed, Inc. | The Emergency System Operation (ESO) mode may not provide adequate patient support under a specific operating condition. | Finished device change control | Abiomed notified customers via an Urgent - Voluntary Medical Device Field Removal Notice on May 8, 2009 advising users that the Emergency System Operation (ESO) mode may not provide adequate patient support under a specific operating condition. The letter stated that customers may continue to support patients with the AB Portable Driver during the process since the failure has not occurred during patient support. An Abiomed Field Service Engineer will contact customers and arrange for a temporary exchange for the affected driver with a corrected driver. The firm recommends that customers switch the patient to hand pump support and switch patient to backup driver until they can confirm that the Driver(s) has been corrected by an Abiomed Field Service. Questions are to be directed to Michael O'Hara at 978-646-1543 or a Field Service at 800-554-8666. |  | 18 | 18 units | Worldwide Distribution -- USA, including states of AL, IN, FL, NJ, NY, and OH and countries of Germany, Greece, and Sweden. | Terminated | January 26, 2012 | 993 | 18
Z-1898-2009 | 52762 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | SmartPort CT Port Access System, Catalog number CT80STPD, AngioDynamics Incorporated |  |  | July 16, 2009 | 2009 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Packaging process control | Notification of the AngioDynamics, Inc. sales force began on July 16, 2009 and they were asked to assist in the return of all affected devices to the Manchester, GA division of AngioDynamics, Inc. Urgent Recall letters were sent to all customers on July 16, 2009, requesting the completion of a FAX BACK form. Those currently in possession of any remaining devices will be provided a Return Goods Authorization (RGA) number and pre-paid shipping boxes for return of the product to AngioDynamics, Inc. All returned devices will be physically and electronically quarantined until final disposition of the devices. All devices will be inspected for damage to the devices and any damaged devices destroyed (scrapped) via the AngioDynamics procedure for Disposition of Nonconformances, procedure. Non-damaged product may be reworked to remove the incorrect catheter and introducer and replaced with the correct catheter and introducer, repackaged and re-sterilized in accordance with the AngioDynamics procedure for Rework. If necessary, follow-up letters, phone calls or visits by the AngioDynamics sales representatives may be used to complete the effective checks. Level A (100%) of consignees to be contacted).Please contact your local representative or Recall Coordinator at 1-706-846-3126, for assistance. You may e-mail your questions to customerservice@angiodynamics.com. |  | 150 | 150 kits | AL, CA, CO, FL, GA, ID, IL, IN, KS, LA, MI, MO, MN, MS, NC, NY, OH, OR, PA, TN and TX. | Terminated | October 27, 2010 | 468 | 150
Z-1910-2009 | 52758 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Jejunal Feeding Set, Product Order number FLOWJ-9-20. |  |  | July 24, 2009 | 2009 | Cook Endoscopy | The English only version of the Patient Care Manuals that accompany the device were shipped to a medical facility in Italy. | Process control | The consignee was notified by an Urgent: Field Safety Notice, dated 07/24/2009. The distributor (Cook Ireland) was instructed to contact the Italian Company Representative. The appropriate Patient Care Manual with the Italian translation was provided to the medical facility with the assistance of the Area Representative. The Area Representative instructed the medical facility staff to provide the translated Patient Care Manual to the caregiver(s) of the patient that received the Jejunal Feeding Sets. The distributor is already aware of the situation. The lower portion of the letter was to be filled out and sent to Cook Endoscopy. |  | 4 | 4 units | International Only Distribution -- Country of Italy. | Terminated | April 06, 2010 | 256 | 4
Z-1947-2009 | 52604 | No | DZA | Intraoral dental drill. | Dental | Dental | 510(K) Exempt | ACT3020 Advanced Cutting technology (ACT) Twist Drill 3.0mm(D) X 20mm(L). BIOMET 3i 455 Riverside Drive, Palm Beach Gardens, FL 33410 USA. Made in Switzerland. |  |  | June 29, 2009 | 2009 | Biomet 3i, LLC | Some of the packages of ACT3020, drills may actually contain an ACT3220 drill. | Mixed-up of materials/components | Consignees were notified of recall by telephone and by faxed an Urgent: Medical Device Recall letter, dated June 30, 2009. The letter stated the problem. The letter also asked customers to check their surgical kits and confirm the drill is placed in the correct space within the kit. Customers are to verify that the shank of the drill is marked as ACT3020 on one side and as 3.0 x 20mm on the opposite side. The letter also stated the clinical implications. Customers are to respond with the attached response fax and then return the drills to BIOMET 3i. Once the returned product is received, replacement drills will be sent to the customers. Questions or concerns should be directed to Kelly Taylor at 561-776-6906. |  | 29 | 29 units | Worldwide Distribution -- Korea, Netherlands, Taiwan, United Kingdom, and USA. | Terminated | January 05, 2010 | 190 | 29
Z-1706-2009 | 52561 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | GE Medical Systems, OEC, Vascular Drape, 00-902776-01, Sterile. 150" long x 87" wide. Distributed by GE OEC Medical Systems, Salt Lake City, UT 84116 USA. |  |  | August 15, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has found that the sterile products manufactured by Contour Fabricators, Inc. (CFI) for GE healthcare (GEHC) may have seals that could be compromised during the shelf life of the products. This issue could result in the product not maintaining its sterile state. Eleven complaints and 3 MDR's have been filed. No illness or injuries reported. | Packaging process control | GE Healthcare issued an Urgent Recall Notice dated August 15, 2008. The letter described the problem, provided a list of products, instructed customers to stop using the products and to return or dispose. To receive credit, customers must complete the attached product return form and include it in the package with returned items. All customers were asked to complete and fax back the attached response form. |  | 2760 | 2760 | Nationwide Distribution -- AZ, CA, FL, MD, NY, OR, PA, SC and TX. | Terminated | May 15, 2012 | 1369 | 5820
Z-1951-2009 | 52803 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | PowerProMax PRO5100M Battery Single Trigger, |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | There is a potential the devices may self activate for their PowerPro, PowerProMax, and MPower I Handpieces manufactured prior to June 1, 2008. | Device Design | Consignees were notified via letter by FedEx delivery or equivalent method. If the product has been further distributed the consignee is requested to notify their consignees. The letter indicated the following risk mitigation actions until affected handpieces are returned to the manufacturer for preventive maintenance: 1) Apply supplied label to the identified handpieces by specific serial number, 2) If self-activation is experienced, discontinue use immediately and return to ConMed Linvtec and 3) Advise/remind the consignee to follow the manufacturers preventive maintenance intervals specified in the instructions for use manuals. Direct questions to ConMed Linvatec at 1-800-237-0169.Recall Expansion: New amended customer and distributor letters will be sent to those identified customer/distributor accounts. The letter will address the additional units. |  | 5553 | 5,553 (4,235 INITIAL AND 1,318 EXPANDED) | Worldwide Distribution including United States. | Terminated | December 19, 2011 | 871 | 23426
Z-2007-2009 | 52812 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | MC5057 12ft/3.6 meters Universal Cable. Made in USA. NON STERILE Rx Only. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | Sufficiently worn/damaged Universal Cable with product number MC5057, when used with various ConMed Linvatec electric Handpieces, may cause the handpiece to self activate. | Component design/selection | ConMed Linvatec notified customers via letter, cautioning the users to inspect cords for signs of excessive wear or damage and urging them to discontinue use and return cables when appropriate. Direct questions to Linvatec Corporation by calling 1-800-237-0169. |  | 12205 | 12,205 units | Worldwide Distribution including United States. | Terminated | July 11, 2011 | 710 | 12205
Z-1715-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3604, O.D. Size: 36mm, Neck Length: -3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 14 | 14 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29
Z-1717-2009 | 52428 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | M-19 (Iridium 192): Source Production and Equipment Company, St. Rose, LA. |  |  | March 03, 2008 | 2008 | Source Production &amp; Equipment Inc | Radioactive source broke during use. | Other | Source Production and Equipment, Inc. notified the Distributor by phone and letter March 2008 and requested the return of all affected units. |  | 3 | 3 units | Nationwide Distribution -- GA, NV and NY. | Terminated | August 19, 2009 | 534 | 3
Z-1895-2009 | 52565 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Clipping Device, Resolution Clip, working length 155 cm, working channel 2.8 mm, Manufactured for Boston Scientific Corporation, Natick, MA; Product codes M00522600 (1 clip) and M00522602 (20 clips). |  |  | June 17, 2009 | 2009 | Boston Scientific Corporation | The clip may not deploy or, if deployed, the clip may not release from the tissue. | Process control | Consignees were notified by an "Urgent Recall Letter " dated June 17, 2009 and instructed to cease using the product because the calibration of equipment used to manufacture these products was not in control. This may potentially result in an inability to deploy or release the clip from the catheter. Consignees were instructed to return a business reply form for return authorization. Please contact Cindy Finney, Field Action Team Lead at 1-508-683-4678 for assistance. |  | 217 | 217 | Nationwide, France, Germany, Great Britain, Italy, Netherlands, Spain and Sweden. | Terminated | October 08, 2010 | 478 | 4486
Z-1889-2009 | 52679 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NX 2008 Central Monitoring System (CMS). Product code: E2FB6 |  |  | June 22, 2009 | 2009 | AGFA Corp. | After opening and closing a study on a Central Monitoring System while a study is still open on an in-room NX system, a problem of image mix-up can occur on the in-room NX system. | Software design | An "Urgent Safety Notice" was sent to all potentially affected sites on/about June 22, 2009 to make them aware of the issue. A FAX Back form is requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. A software service update to solve this issue is expected to be released by August 15, 2009. A Mandatory Service Bulletin will be deployed to correct the problem in the software. If you have questions, please contact Agfa Healthcare at 1-800-581-2432 prompts 1, 3 and then 2 for CR Support. |  | 97 | 97 units | Worldwide - IL, PA, TX, Canada, Europe, Asia Pacific and Latin America | Terminated | July 22, 2010 | 395 | 97
Z-1444-2009 | 51704 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sybaritic, NannoLight MP50. |  |  | April 12, 2008 | 2008 | Sybaritic, Inc | Sybaritic Inc. had an issue with the 585nm filter used with the NannoLight MP50. The filter may cause unwanted burning and should not be used for vascular lesion treatments. | Device Design | A "Sybaritic" notification letter dated April 15, 2008 was issued to consignees. The letter described the modifications to the protocols for the NannoLight MP50 and the temporary replacement of 585 Filter for the 560 Filter for Vascular Lesions until further notice. They requested the immediate use of the protocols replacing all other versions.A follow-up letter dated June 3, 2008 was issued to customers. The letter included an attachment " Product Notice Nannolight MP50 585 nm "special" filter from Sybaritic, Inc. The letter described the new 585 nm "bandwidth filter". Also the notification advised consignees to discontinue the use of the old 585 filter and to use 560nm filter until they receive the new 585 nm bandwidth filter.Please direct any questions to your Account Manager or to Sybaritic, Inc. by calling 1-800-445-8418 or 1-952-888-8282. |  | 59 | 59 | Worldwide Distribution -- US (states of WA, CA, MN, FL and AR), POLAND, THAILAND, CHILE, CANADA, KOREA, AZERBAIJAN, INDIA, SPAIN, ARGENTINA, RUSSIA, GREECE, PHILIPPINES, POLAND, INDIA, JORDAN, TURKEY, MEXICO, URUGUAY,BELGIUM, LEBANON, BULGARIA and SINGAPORE. | Terminated | July 05, 2012 | 1545 | 59
Z-1923-2009 | 52115 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Malibu Polyaxial Screw, 7.5 x 35mm Part Number 12-7535 |  |  | March 04, 2009 | 2009 | SeaSpine, Inc | SeaSpine has received reports relating to rust on the internal washer contained within the titanium seat assembly of certain Malibu Polyaxial and Malibu Reduction Screws. | Mixed-up of materials/components | In an Urgent Voluntary Product Recall Notification letter dated March 5th, 2009, SeaSpine informed consignees about the reason for recall and asked them to immediately check the products in their possession and hospital inventories to determine if they have any of the affected product. Alternatively, there were told if they would prefer to return to SeaSpine all Malibu Screw inventory that is in their possession, SeaSpine will sort out the affected product upon receipt and will replace their inventory with Malibu Screws that conform to all SeaSpine specifications. The part number and lot are laser marked on the outside of the seat assembly of each screw.SeaSpine asked their consignees to return affected product to them immediately. Attached was a verification form, that they were told to complete and return even if they did not have any product to return. Furthermore their local Sales Manager could assist them in completing the form. They were also encouraged to contact their Sales Manager or Steve Rybka at (760) 727-8399 x202 if they had any questions regarding the recall, any of their products, or would like assistance with the recall. |  | 1116 | 1,116 Screws in total | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, GA, IN, MD, MI, MO, MS, NM, OK, TN, TX, VA, and WI and country of Greece. | Terminated | April 28, 2011 | 785 | 1116
Z-2050-2009 | 52695 | No | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | Portex Continuous Epidural Tray, Reference 4037-20 |  |  | July 15, 2009 | 2009 | Smiths Medical ASD, Inc. | Portex Continuous Epidural Tray was distributed with a 20G closed end, three eyed, nylon catheter without a stylet instead of the 20 G open end Teflon catheter with stylet specified on the label. | Mixed-up of materials/components | Smiths Medical notified consignees by an Urgent Device Recall letter dated July 15, 2009, to locate and quarantine the affected product, stop distribution, and complete and fax back the attached Distributor Reply Letter to Smiths Medical. Credit or replacement will be arranged as requested for the return of any affected product. All distributors will be required to return any affected product in stock. All of the distributors will be required to send their customer list to Smiths Medical who will notify their customers of this recall. |  | 190 | 190 kits (19 cases) | Nationwide Distribution -- including states of New York, Texas, Ohio, Utah, Alabama, Arizona, Kentucky, &amp; Wyoming. | Terminated | September 14, 2012 | 1157 | 190
Z-2004-2009 | 52724 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Roche CoaguChek XS System, CoaguChek XS Plus Care Kit, Roche Diagnostics, Indianapolis, IN; Model Number 05021537001. |  |  | July 22, 2009 | 2009 | Roche Diagnostics Operations, inc. | The recommended cleaning procedures for the meter may cause a malfunction and falsely elevated test result reports. | Device Design | Roche Diagnostics Operations, Inc. notified Consignees by letter dated July 22, 2009 informing them of the problem and instructing them to use revised cleaning procedures.For further questions, contact Roche Diagnostics Operations, Inc. at 1-317-521-3194. |  | 1327 | 1,327 in US | Worldwide Distribution | Terminated | April 27, 2011 | 644 | 66326
Z-1998-2009 | 52817 | No | FDG | Gastroenterology-urology fiberoptic retractor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Ultra-Thin LightMat Surgical Illuminator, Disposable one time use. Model Number: UA2550, Exp Date: 2012-05. Product is packaged in individual sterile pouch w/5 pouches per box. |  |  | July 01, 2009 | 2009 | Lumitex Inc | Adhesive which is used to attach LightMat to retractor may not have been tested to Medical standards. The vendor provided an adhesive strip that did not conform to Lumitex specifications. The adhesive strip provided may not have met Lumitex requirements to be medical grade, EtO compatible, and biocompatible. | Nonconforming Material/Component | Lumitex, Inc. issued a recall letter dated July 2, 2009 to their Customers describing the affected device and requesting its return.For further information, contact Lumitex, Inc. at 1-440-243-8401. |  | 295 | 59 packs/5 individual units | Product was distributed to the following states: CA, FL, KS, MA, ME, MI , NC, NM, NV, NY, OH, and TX. | Terminated | March 09, 2012 | 982 | 295
Z-1961-2009 | 52611 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Celsite Access Port Model ST305P. |  |  | June 22, 2009 | 2009 | B. Braun Interventional Systems | B. Braun Interventional Systems was notified by B. Braun Medical France, the manufacturer of the Celsite Access port that they had received 11 reports of polyurethane (type P) catheter disconnections. B. Braun Medical France has decided to recall all CelsiteÂż ports supplied with the affected catheter lot. To date, B. Braun Interventional Systems (US) has not received any reports of catheter dis | Nonconforming Material/Component | B. Braun Interventional Systems, Inc. issued an "Urgent: Medical Device Recall" notice dated June 22, 2009. The letter identified the problem and product. The letter reminded the consignee of the general instruction for use of access ports and instructed them to contact patients who had the Celsite port system implanted.For further questions, contact B. Braun Interventional Systems, Inc. at 1-800-523-9676 extension 4818. |  | 40 | 40 in US | Nationwide Distribution -- NY. | Terminated | June 05, 2012 | 1079 | 40
Z-2010-2009 | 52898 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Transtracheal (Jet Ventilation) Catheter ACU 1060.1, labeled sterile if unopened. |  |  | July 09, 2009 | 2009 | Instrumentation Industries Inc | user instructions not updated to correspond with product change | Component change control | Instrumentation Industries, Inc. issued a notification letter dated July 9, 2009 to their customers explaining the problem and the need to replace current inventory user instructions with the revised user instructions accompanying the letter.For further questions, contact Instrumentation Industries, Inc. at 1-800-633-8577. |  | 396 | 396 units | Nationwide Distribution -- AK, AL, AR, CA, CO, FL, GA, IA, ID, IL,IN, MA, MD, ME, MI, MN, NC, ND, NH, NJ, OH, PA, SC, SD, TX, VA and WA. | Terminated | October 15, 2009 | 98 | 396
Z-1965-2009 | 52798 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | iLink Bone Void Filler. Packaged in a double Tyvek pouch, then placed in a box. Model Number: 50-05-0010, Exp. 2010-04. |  |  | August 26, 2008 | 2008 | Theken Spine LLC | Failure to label the product properly. Additionally, 15 pieces were reworked and the expiration date was misprinted as 2011-04 instead of the actual date of 2010-04. | Error in labeling | Theken Spine, LLC contacted the Consignee August 26, 2008 via telephone and instructed them to inventory their stock and return affected product to the firm.For further questions, contact Theken Spine, LLC at 1-330-475-8662. |  | 22 | 22 units | Nationwide Distribution -- TX. | Terminated | March 09, 2012 | 1291 | 22
Z-2048-2009 | 52993 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Stratus OCT Model 3000 instruments distributed with version 6.0.0 software, product number 2660021119710 (new instrument) or 2660022119710 (refurbished instrument), and Stratus 6.0.0 Software Kit, product number 2660021133164. |  |  | July 31, 2009 | 2009 | Carl Zeiss Meditec Inc | Mis-labeled image. A bug in version 6.0.0 software and 6.0.0 software kit used with Stratus OCT Model 3000 can produce a report that incorrectly labels OD or OS. | Software design | Firm initiated notification to customers on July 2, 2009, by sending a letter via FedEx to US customers, and via e-mail and FedEx to outside US consignees. Customers were informed of the issue and steps needed to correct it. A revised software version will be provided to the customers. Any questions should be directed to Carl Zeiss Meditec, Inc. Customer Care at 1-800-341-6968. |  | 244 | 244 units | Worldwide - US, Austria, UK, Japan, Singapore, Argentina, France, Hong Kong, New Zealand, Canada, Colombia, Germany, Korea, Venezuela, India and China. | Terminated | April 19, 2010 | 262 | 244
Z-2029-2009 | 50785 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Rx only. |  |  | December 15, 2008 | 2008 | Newport Medical Instruments Inc | In December 2008, Newport initiated a field correction to upgrade e360 ventilators with software versions 3.0 or higher to software version 3.3. As a result of this upgrade, some customers experienced unexpected device alert alarms. Newport's investigation into these complaints confirmed that the 3.3 software responds too quickly to certain situations, giving a false device alert. Newport condu | Software design | Newport Medical initally sent out a Field Correction Notice to a limited number consignees dated December 17, 2008 and then the notice was expanded to include all e360 consignees dated June 15, 2009. The firm notified consignees that they have revised the e360 software to version 3.8. A customer verification form was attached to verify upgrade completion. For further questions, contact Newport Customer Technical Support Department at 1-800-451-3111 or +1-714-427-5811 Extension 500. |  | 4117 | 4,117 units (214 Domestically; 3,903 Internationally) | Worldwide Distribution -- US (including AL, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, ME, MI, NJ, NM, NY, OH, OK, OR, PA, UT and VA), Albania, Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Brunei, Bulgaria, Canada, Chile, Colombia, Czech Republic, Dominican Republic, Ecuador, Egypt, England, Finland, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Japan, Jordan, Kazahkstan, Kenya, Korea, Kuwait, Lebanon, Liechtenstein, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nambia, Nepal, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Sri Lanka, Taiwan, Thailand, Tunisia, Turkey, AUE, Ukraine, United Kingdom, Yemen Republic and Zimbabwe. | Terminated | November 04, 2011 | 1054 | 4117
Z-1558-2009 | 52019 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Device Interfacing Solution, N-DISVENT 02. |  |  | May 08, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an issue in the calculations associated with Datex-Ohmeda S/5 Device Interfacing Solution N-DISVENT-02 interfaced between certain Datex-Ohmeda patient monitors (Datex-Ohmeda S/5 Anesthesia Monitor (AM), Datex-Ohmeda S/5 Critical Care Monitor (CCM), Datex-Ohmeda S/5 Compact Anesthesia Monitor (CAM), Datex-Ohmeda S/5 Compact Critical Care Monitor (CCCM)) an | Software design | An "Urgent Medical Device Correction" letter was issued to consignees on or about April 9, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A representative from GE Healthcare will contact customers to schedule implementation of the correction when it is available for installation at the customer's facility. Direct questions regarding this notice to GE Healthcare Technical Support or your local GE Distributor and refer to FMI 36042. |  | 2638 | 2,638 | Worldwide Distribution -- US (including DC and states of AL, AK, AZ, AR, CA, CO, FL, IL, IN, IA, KS, KY, LA, ME, MD, MI,MN, MO,MT, NE, NH, NJ, NY, NC, OH, OK OR, PA, SC, TX, UT,VT, VA, WA, WV and WI) and countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, CANADA, CHINA, CYPRUS, DENMARK, ESTONIA, FRANCE, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PHILIPPINES, PORTUGAL SAUDI ARABIA, SERBIA, SOUTH AFRICA, SWEDEN SWITZERLAND, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | February 04, 2012 | 1002 | 2638
Z-2016-2009 | 51966 | No | LIF | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HACH SteriChek Residual Peroxide Reagent Strips, 5/100-strip bottles per kit, Kit Part #811905, Bottle part #1905LB, Hach Co., Ames, IA. |  |  | April 21, 2009 | 2009 | Hach Co | Some of the reagent strip bottles contain the incorrect strips | Labeling mix-ups | Hach Company issued an "Urgent Medical Device Recall" letter dated April 21, 2009 explaining the reason for recall and requested product remaining in inventory be returned. Subrecall to the end-user level was requested. A Return Form was also enclosed to be completed to indicate the amount of product being returned.Between April 22-23, 2009, the firm telephoned their customers to assure they received the recall letter.For further information, contact Hach Customer Service at 1-800-548-4381 extension 3560. |  | 275 | 275 kits | Nationwide Distribution -- NJ, CA, IN, PA, NV, CA, and ME. | Terminated | October 07, 2009 | 169 | 275
Z-2008-2009 | 52209 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Model 100 enFlow Fluid/Blood Warmer |  |  | March 23, 2009 | 2009 | Enginivity LLC | 1. Fluid contamination of the electronic circuitry can cause short circuits resulting in permanent fuse failure or localized circuit board overheating with possible smoldering and discoloration. 2. Electromagnetic (EMI) Interference associated with the Warmer may produce artifact in ECG, EEG, or EMG records (cardiac or neuro monitoring). | Device Design | Vital Signs, Inc. issued Product Safety Notification on March 23, 2009 informing users of potential EMI and fluid ingress issues. All warmers in the field will be replaced beginning in May 2009. |  | 3052 | 3052 units | Worldwide Distribution -- USA, Philippines, Saudi Arabia, India, Israel, Australia, Mexico, Germany, Republic of Singapore, China, Guatemala, Switzerland, Dubai United Arab Emirates, New Zealand, Malaysia, India, Ecuador, &amp; United Kingdom. | Terminated | May 02, 2012 | 1136 | 3052
Z-1862-2009 | 52163 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker 550 mL AutoFuser PainPump, single site continuous with bolus, 5 mL/hr, 1 x 5.0 in. catheter, sterile, REF 0535-500-050. |  |  | May 14, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Failure to infuse and potential compromised sterility: The seal on the tyvek tray cover may fail, compromising sterility, and the device may fail to infuse the drug, resulting in a loss of pain mitigation. | Process control | Stryker sent all consignees a Medical Device Recall Notification letter dated 5/14/09 and instructed to locate and destroy the recalled products and to complete and return the business reply card. |  | 28755 | 28,755 of all pumps | Nationwide and Canada. | Terminated | January 20, 2010 | 251 | 28755
Z-1808-2009 | 52697 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. |  |  | March 24, 2009 | 2009 | American Optisurgical Inc | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. | Software change control | American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266. |  | 12 | 12 units | Worldwide distribution: USA (California), Brazil, Algeria, France, South Korea, India, and Peru. | Terminated | September 04, 2009 | 164 | 12
Z-1600-2009 | 52002 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center Central Station (CIC Pro version 5.1, MP100 Hardware). |  |  | March 10, 2009 | 2009 | GE Medical Systems Information Tecnology | GE Healthcare has become aware of a potential issue associated with the failure to restart CARESCAPE CIC Pro v5.1 on the MP100 hardware. This issue may impact patient safety. | Component change control | A GE Healthcare "Urgent Medical Device Correction" letter dated February 24, 2009 was sent to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Affected units will be replaced with new units byGE Field Service team members.For questions regarding this notice or experience a failure to restart the device, please contact GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative. |  | 352 | 352 | Worldwide Distribution -- US (states of AL, AZ, CA, CT, FL, GA, IL, IN, KY, LA, MD,KY, MI, MN, MS, MO, NV, NH, NJ, NY, NC, OH, OK, PA, RI, TN, TX, VA, WA, WV and WI), AUSTRIA, AUSTRALIA, CANADA, CZECH REPUBLIC, GERMANY, EGYPT, SPAIN, FRANCE, GREAT BRITAIN, ITALY, POLAND and SWEDEN. | Terminated | October 16, 2012 | 1316 | 352
Z-1992-2009 | 52701 | No | MBF | Shoulder joint metal/polymer/metal nonconstrained or semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Turon Revision Glenoid Poly Liner. |  |  | July 20, 2009 | 2009 | Encore Medical, Lp | Product devices used for testing are not identical to what was manufactured. Under certain circumstances snap feature could allow poly liner to move on baseplate reducing tightness of fit. | Release of Material/Component prior to receiving test results | Firm notified consignees by an Urgent: Medical Device Recall letter, dated July 20, 2009. The letter stated the reason for the recall, asked customers to quarantine, and return the affected product. The enclosed response form needed to be completed and returned as soon as possible. Questions were to be directed to David Harris, Manager of Customer Service at 512-834-6330. |  | 60 | 60 units | Nationwide Distribution -- Including states of TX, FL, and GA. | Terminated | January 07, 2010 | 171 | 60
Z-2026-2009 | 52773 | valve, prosthesis | ISP | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Icelock Expulsion Valve 551 Model L-551000 |  |  | June 23, 2009 | 2009 | Ossur Engineering, Inc | The plastic caps have come off or become loose when the user stepped into their prosthesis. | Process change control | Customers were sent a Safety Notice Voluntary Product Recall, dated June 23, 2009, that advised them to return any unused valves to Ossur for a free replacement. Those using valves in combination with a cushion liner and suspension sleeve should replace immediately if problems are experienced, otherwise replace at next routine service. Customers using valve in combination with an Iceross Seal-In liner should replace ASAP and return used valves to Ossur for free replacement. Customers with questions can call Darlyne Roa 1-800-233-6263 ext 1023743. |  | 4647 | 4,647 | Worldwide Distribution | Terminated | March 06, 2012 | 987 | 4647
Z-1989-2009 | 52775 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Power Supply Accessory (model AD-0660G) |  |  | June 18, 2009 | 2009 | Breg Inc, An Orthofix Company | Power supplies short-circuited at the electrical outlet upon initial use of the power supply, causing the circuit breaker for the electrical outlet to trip. | Employee error | An Urgent: Medical Device Recall Notification was sent to customers on June 18, 2009. FOR USERS: Customers are advised to immediately open all unused Kodiak cold therapy units and inspect the lot code of the Power Supply included with the Kodiak unit. If the lot code is 3508, 3908, 4308, or 4808, then remove the power supply, cut the cord using scissors and discard. Call Breg Customer Service for a replacement: 1-800-321-0607. Complete and return the response form ASAP.FOR DISTRIBUTORS: Send Breg your distribution information, including contact and quantities. Breg will notify your customers of the recall directly. |  | 20439 | 20,439 | Worldwide Distribution | Terminated | March 28, 2011 | 648 | 20439
Z-1996-2009 | 52801 | fiber, medical, absorbent | FRL | Medical absorbent fiber. | General Hospital | General Hospital | 510(K) Exempt | Coloplast, InterDry Ag 10" x 144" Roll P/N 6791001400, Manufactured for Coloplast A/S DK-3050 Humlebaek, Denmark, Distributed by Coloplast Corp Minneapolis MN 55411. |  |  | July 21, 2009 | 2009 | Coloplast Corp Skin Care Div | Coloplast determined that a portion of one lot of InterDry" Ag Lot number 1789435, did not meet the product specification for silver content. In this portion of the lot antimicrobial properties of the product could be compromised although the absorption and transportation of moisture would be as expected. | Process control | Consignees were faxed on 7/27/09 a Coloplast "Product Recall" letter dated July 2009. The letter addressed the problem and requested consignees to contact Coloplast Customer Service at 1-800-533-0464 to communicate the status of the product. Replacement product will be provided upon receipt of returned product. It also advised distributors to contact their customers immediately and advise them of the recall and to return the product them. Questions or concerns should be directed to the customer's customer care rep at 1-800-533-0464. |  | 47216.4 | USA, 876 cases (10 rolls/case) ; Canada 53.9 cases (10 rolls/case) | Worldwide Distribution -- USA, including states of AL, AZ, CA, CT, FL, GA, IL, IN, IA, KY. LA, MI, MA,MO, MN, MS, MD, NE, NV, NH, NJ, NM, NY, NC, OH, OK OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and country of CANADA. | Terminated | January 11, 2011 | 539 | 47216.4
Z-2015-2009 | 52826 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Expo, Angiographic Catheter, Sterilized with ethylene oxide gas, Made in Mexico, Parque Industrial La Mesa, Edificion G1, Suites B. H, Tijuana, Baja California 22650 MEX. |  |  | August 05, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of numerous batches/lots of Wiseguide Guide Catheter, Impulse Angiographic (Diagnostic) Catheter, Expo Angiographic (Diagnostic) Catheter. Boston Scientific has determined that the sterile barrier in the packaging may be compromised. During internal inspections, they discovered gaps between the side seals and the top seal approximately 3/8 in len | Packaging | Boston Scientific issued an "Urgent Medical Device Recall Removal" letter dated August 5, 2009. The letter was addressed to Risk Manger / Field Action Contact. The letter described the problem and product involved. The Firm recommended to segregate affected product immediately and return to Boston Scientific and requested return of the "The Reply Verification Tracking Form" enclosed with the letter.For further questions, contact your local Boston Scientific Sales Representative or call 1-763-494-1634. |  | 367898 | 367,898 (US 167,496; OUS 224,292) | Worldwide Distribution -- US (including DC and Puerto Rico), Japan, United Arab Emirates, Albania, Austria, Belgium, Bulgaria, Bahrain, Switzerland, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, Finland, France, Great Britain, Greece, Hungary, Ireland, Israel, Iraq, Iceland, Italy, Jordan, Kuwait, Lebanon, Lithuania, Libya, Morocco, Netherlands, Norway, Pakistan, Poland, West Bank Gaza Strip/ Palestine, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovakia, Syria, Turkey, Ukraine, South Africa, Argentina, Australia, Brazil, Canada, Chile, Colombia, Dominican Republic, Ecuador, Guatemala, Indonesia, India, South Korea, Mexico, Malaysia, Nicaragua, New Zealand, Panama, Philippines, Puerto Rico, Paraguay, Singapore, El Salvador, Thailand, Taiwan, Uruguay, American Virgin Islands and Vietnam. | Terminated | January 07, 2012 | 885 | 787578
Z-1948-2009 | 52768 | No | FHQ | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARDÂż Simpulse" SOLO System with Soft Splash Shield Tip |  |  | July 30, 2009 | 2009 | Davol, Inc., Sub. C. R. Bard, Inc. | Product sterility may be compromised due to lack of package integrity. | Process control | Davol issued a Recall Letter dated July 30, 2009 requesting user facilities examine inventory and immediately discontinue use of any product from the affected lot and return it to Davol. Direct questions about the recall to Davol, Inc., Sub. C. R. Bard, Inc. by calling 1-401-463-7000. |  | 1140 | 1,140 units | Nationwide. | Open, Classified |  |  | 1140
Z-1883-2009 | 52767 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System 3500, Cardinal Health, Clinical Technologies and Services, San Diego, CA. |  |  | December 24, 2008 | 2008 | Cardinal Health Manufacturing LLC | Cardinal Health has sent out a Safety Alert Notification to inform its customers that the SR2b Software for the Pyxis Anesthesia System 3500 (PAS 3500) has received several customer reports of system "non-responsiveness" (lock ups) during the use of the product since its release in September 2008. | Software design | The "Safety Alert Notification" letter was mailed to all Pyxis Anesthesia System 3500 customers on December 17, 2008 by registered return receipt mail. The letter was addressed to the Director of Pharmacy at each facility. The mailing packages contain the "Safety Alert Notification" letter and the revised pages of the Pyxis Anesthesia System 3500 User Guide and Safety Manual. The "Safety Alert Notification" letter reinforced to customers how to maintain back-up processes to allow quick access if a system lock-up condition is experienced.The letter informed customers the situation will be corrected and a solution will be implemented. If you experience any system unresponsiveness during the use of the product, please report the issue to the Cardinal Health Technical Support Center at 1-800-727-4102. If you have any questions regarding this issue, please contact your local field support representative or the Customer Advocacy Department at Customer.Advocacy@cardinalhealth.com. |  | 250 | 250 | Nationwide and Germany. | Terminated | March 28, 2012 | 1190 | 250
Z-2053-2009 | 51674 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra NeuroSciences Ventrix Ventricular Tunneling Pressure Monitoring Kit NL950-V and Integra NeuroSciences Ventrix Ventricular Tunneling Pressure Monitoring Kit with Ventricular Drainage System NL950-VC. |  |  | March 16, 2009 | 2009 | Integra Neuro Sciences | Due to a dimensional error, the diameter of the stylet provided with the device may be too large to insert into the lumen of the catheter tip. | Process design | An "Urgent-Product Recall Notification" letter dated March 23, 2009 was issued to consignees. The letter described the issue and instructed customers to examine their inventory immediately to determine whether they have any NL950V - Ventrix Ventricular Tunneling Pressure Monitoring Kits, or NL950VC - Ventrix Ventricular Tunneling Pressure Monitoring Kits with Ventricular Drainage System. Customers were further instructed not to use the above recalled product.Call Integra Customer Service at 1-800-654-2873, and then select 5 for 'Returns and Repairs' to obtain a Returned Material Authorization number, and return the product as directed.Direct questions about the recall to Integra NeuroSciences by calling 1-858-455-1115, extension185. |  | 274 | 274 kits | Worldwide Distribution -- US (states of AR, CA, OH, TX, CT, NY, LA, FL, NY, IL, and VA), Belgium and Malaysia. | Terminated | July 19, 2012 | 1221 | 274
Z-1946-2009 | 52232 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJO Carendo Multi-Purpose Hygiene Chair; Model BIB2003-01, a non-AC powered patient lift; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden. |  |  | June 11, 2009 | 2009 | ArjoHuntleigh | Male patient genitals may be pinched between the seat and the frame of the patient lift chair when the Care Raiser seat is returned to its lowest position. | Device Design | ArjoHuntleigh sent Urgent Device Field Correction letters dated June 11, 2009, along with the Urgent Field Safety Notice ID 090526AHE dated May 26, 2009, via UPS Ground delivery. Customers were informed of the affected device and were requested to complete the response form with new contact information and fax it back to ArjoHuntleigh at 630-567-5012 or 630-925-7969. For further questions, contact ArjoHuntleigh at 1-800-323-1245. |  | 1987 | 1,987 hygiene chairs | Nationwide Distribution | Terminated | August 19, 2010 | 434 | 1987
Z-1930-2009 | 52000 | No | DCN | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | Dimension Flex Reagent Cartridge C-Reactive Protein Extended Range (RCRP). |  |  | April 15, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | falsely elevated results in EDTA plasma samples | Labeling False and Misleading | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice on April 15, 2008 informing their customers that the EDTA plasma specimen claim had been rescinded. Current lots in the firm's inventory have been put on hold and reworked with an Alert Card rescinding the EDTA plasma claim. Subsequent lots use a revised instruction for use stating "Do not use EDTA plasma" in the Specimen Collection and Handling Section.For further questions, contact Siemens Healthcare Diagnostics at 1-302-631-6564. |  | 25760 | 25,760 cartons | Worldwide Distribution -- US, Austria, Bosnia, Czech Republic, Hungary, Slovakia, Serbia, Denmark, Estonia, Latvia, Lithuania, Norway, Sweden, Belgium, Cyprus, Canary Islands, France, Great Britain, Germany, Ireland, Italy, Netherlands, Poland, Russia, Portugal, Saudi Arabia, Spain, South Africa, Switzerland, Turkey, Yemen, Bahamas, Brazil, Canada, Mexico, Australia, Bangladesh, China, Hong Kong, Japan, Malaysia, Pakistan, and Taiwan. | Terminated | September 09, 2009 | 512 | 25760
Z-1949-2009 | 52672 | No | EJL | Dental bur. | Dental | Dental | 510(K) Exempt | BIOMET MICROFIXATION Trauma One SYSTEM 1.8X115 MM STOP DRILL, 1.8X115 MM WITH 26 MM STOP STAINLESS STEEL, NON-STERILE, BIOMET MICROFIXATION, INC. 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA. |  |  | February 16, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling the 46-1004, Trauma One Drill, 18.x115 mm, 26 mm Stop Lot Numbers 619450 and JC1002. The drills may be color-banded blue instead of brown. The color is a visual aid in determining what system the drill goes with. The blue indicates the drill is for the 2.0mm system. The brown indicates the drill is to be used for our 2.3mm system. The drill is etched and labe | Process control | Biomet Microfixation, Inc. issued an Urgent Medical Device Recall Notice dated February 19, 2009 via certified mail with return receipt. The notice informed users of the affected device and instructed them to immediately discontinue use and return to the firm. Consignees were also instructed to complete and return an Inventory Reconciliation sheet via fax to 1-904-741-9425.For further information, contact Biomet Microfixation, Inc. at 1-800-874-7711 or 1-904-741-4400. |  | 79 | 79 | Nationwide Distribution -- Texas, Maryland, Florida, Washington DC, Michigan, Kansas Colorado and Missouri. | Terminated | November 05, 2009 | 262 | 79
Z-1887-2009 | 52757 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Impax R6.3 DB Server. |  |  | July 16, 2009 | 2009 | AGFA Corp. | Measurements are inaccurate on multiple workstations. | Software design | Agfa Healthcare issued an "URGENT FIELD SAFETY NOTICE" dated July 16, 2009. The notice informed of the problem and provided details to identify potentially compromised data that could result in measurement inaccuracies. A feedback form was provided to be completed and returned to the firm.For further information, contact Agfa Healthcare at 1-773-592-4769. |  | 1 | One unit | Nationwide Distribution -- IL | Terminated | April 28, 2011 | 651 | 1
Z-1912-2009 | 52694 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System; |  |  | June 30, 2009 | 2009 | Vital Signs Inc | Not compatible: The luer connector on the Intraosseus needle provided does not fit securely on the luer connection of the IV Extension Set in the module. | Device Design | Vital Signs, Inc. issued an "Urgent Medical Device Correction" letter dated June 30, 2009 to all direct and indirect accounts by first class mail. Consignees were informed of the affected device, asked to complete a Confirmation form for return and to return unused product for replacement.For further questions, contact Vital Signs, Inc. at 1-973-790-1330 extension 404. |  | 23429 | 23,429 kits | Worldwide Distribution -- US, Canada, United Arab Emirate, Venezuela, Guam, Korea and United Kingdom. | Terminated | August 27, 2012 | 1154 | 23429
Z-2024-2009 | 52642 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 The Endovascular Company, Protege GPS 6mm x 80mm Self Expanding Peripheral Stent System , Rx Only, and SERP65-06-80-120. |  |  | June 10, 2009 | 2009 | Ev3, Inc. | ev3 Inc. is conducting a voluntary recall of specific lots of the Protege GPS Self-Expanding Stent System because labeling did not match the actual stent size contained in the package. The affected product was labeled as a 6mm x 80mm Protege GPS stent but actually contained an 8mm x 40mm stent. This situation can lead to vessel damage or inadequate coverage of the target lesion. | Labeling mix-ups | Consignees were emailed a letter with a subject heading of "Medical Device Recall". The e-mail was dated June 10, 2009. The letter informed the consignee of the problem and product involved. They requested that any affected devices remaining in inventory be identified and returned and to fill out and return the "Device Recall Field Action Reconciliation Form". |  | 2 | 2 | International Distribution Only -- Brazil, Uruguay, &amp; Malaysia. | Terminated | December 11, 2011 | 914 | 13
Z-2047-2009 | 52927 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Pedi-Cap CO2 Detector, 1 kg - 15 kg, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Assembled in Mexico. Sold in packages of 6 (PediCap6) or 24 (PediCap). |  |  | August 14, 2009 | 2009 | Covidien Limited | CO2 detector may increase resistance to air flow in an intubated patient. | Component change control | Covidien issued an Urgent Medical Device Recall letter dated August 14, 2009 informing Consignees of the affected devices. Users were told to stop use and arrange for return of all affected product. For further information, contact Covidien at 1-800-635-5267. |  | 276390 | 276,390 units | Worldwide Distribution -- US (FL, ID, MD, NC, VA and PA including Puerto Rico), Australia, Belgium, Canada, Japan, Singapore. | Terminated | June 01, 2012 | 1022 | 276390
Z-2042-2009 | 52885 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Biomet 8.0MM cannulated reamer for the Aimer system, Biomet Sports Medicine, Ontario, CA; REF 909618. |  |  | August 06, 2009 | 2009 | Biomet, Inc. | These surgical instruments are etched with incorrect measurement markings. | Nonconforming Material/Component | Consignees were notified by letter dated 8/6/09 to remove the instrument from use and to return the recalled instrument to Biomet. Implanting surgeons were notified by letter dated 8/6/09, advised of the problem and it was suggested that the patients be monitored for instability if deemed necessary by the surgeon. |  | 2 | 2 | Nationwide Distribution -- Including the state of Texas. | Terminated | October 06, 2010 | 426 | 2
Z-2041-2009 | 52804 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ST ULTRAPOWER DIAMOND WHEEL (RED) 25.4mm, Catalog Number 00702109200, STERILE, Single-use. Made is USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Ultrapower Bur, Diamond Wheel for possibility of a packaging breach in the sterile barrier. | Package design/selection | Firm notified international and domestic consignees by an Urgent Medical Device Recall letter on July 31, 2009. Customers were asked to check their inventory; complete, sign, and return the enclosed Reply Form; and return the devices for credit. Questions should be directed to Customer Service at 1-800-237-0169. |  | 2502 | 2502 | Worldwide Distribution | Terminated | May 11, 2010 | 284 | 2502
Z-2049-2009 | 52963 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron 5F Select Catheter, Neuro Deliver/Intracranial Access System Catheter, manufactured by Penumbra, Inc., Alameda, CA. |  |  | July 02, 2009 | 2009 | Penumbra Inc. | Leakage-- Due to a manufacturing error, some products may have a fold in the innermost layer, causing disruption in the layering process and may result in blow holds or exposed wire braids, which may result in further damage. | Process control | Penumbra, Inc. initiated Recall activities on May 29, 2009. The Firm notified consignees via telephone and email of the problem and asked them to return the product to Penumbra.For further information, contact Penumbra, Inc. at 1-510-748-3223. |  | 9 | 9 units | Worldwide Distribution -- US (IA, NJ, NY and VA), and Canada. | Terminated | January 05, 2011 | 552 | 9
Z-1877-2009 | 52659 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific, LeVeen Needle Electrode, 25 cm, 5.0 cm, sterile; REF M001262170. For tissue coagulation. |  |  | July 06, 2009 | 2009 | Boston Scientific Corp | The tines may detach from the core wire during use, which may result in bleeding or soft tissue trauma requiring medical intervention. | Process control | Boston Scientific notified consignees by letter dated 7/6/09 and instructed to cease using and to return the products, and to complete and return a reply form. |  | 347 | 347 | Nationwide. | Terminated | October 08, 2010 | 459 | 3121
Z-1911-2009 | 52652 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Adapt AO, Akreos Advanced Optics Aspheric Lens. Bausch &amp; Lomb Incorporated, Rochester, NY. Manufactured by Bausch &amp; Lomb Incorporated, Clearwater, FL. |  |  | December 08, 2008 | 2008 | Bausch &amp; Lomb Inc | Affected lenses were inadvertently shipped to Bausch &amp; Lomb's European Logistics Center (ELC) prior to release authorization. Initial testing indicated the product did not meet release criteria for endotoxin limits. | Release of Material/Component prior to receiving test results | All consignees were notified via phone, telefax and letter. Products are to be returned to Bausch &amp; Lomb's European Logistics Center (ELC) for destruction. Direct questions to Bausch and Lomb, Inc. by calling 1-585-338-6612. |  | 158 | 158 | Foreign Distribution only -- Algeria, Belgium, Bulgaria, Croatia, Germany, Hungary, Ireland, Israel,Italy, Mexico, Portugal,Spain, Thailand andUnited Kingdom. | Terminated | November 05, 2009 | 332 | 158
Z-1999-2009 | 52667 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | 300 Microliter CO-RE Tips for use with Microlab STAR, manufactured by Hamilton Company, Reno, NV. |  |  | April 09, 2009 | 2009 | Hamilton Co | Product Defect-- some tips from the mold cavity 28 are too short and do not always aspirate or dispense liquid correctly. | Process control | The firm initiated the recall on April 10, 2009, and issued notification to consignees via certified mail and e-mail. Customers were instructed to return the affected product to the Hamilton Company. A fax-back form was supplied with each notification. Direct questions about the recall to the Hamilton Company by calling 1-775-858-3000, extension 236. |  | 4719000 | 20 cases of Part 235950 and 27 cases of Part 235902. | Worldwide Distribution -- US (states of DE, NJ, PA, MA and CA), New Zealand and Japan. | Terminated | January 04, 2010 | 270 | 4719000
Z-2052-2009 | 52668 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Neuroform 3 Microdelivery Stent System, Model Number M003E3400300, Catalog model number and product upn number M00490310, manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | May 08, 2009 | 2009 | Boston Scientific Corporation | Mislabeling-- Product was placed in a carton, which has printed on it the name of another product. A correct product label was placed on the carton. | Employee error | The firm issued an Urgent Product Field Action Notice to all listed customers, along with a reply verification form, requesting the return of affected products. |  | 8 | 8 units | International Distribution Only -- Countries of Germany, Great Britain, France, Ireland, and Italy. | Terminated | October 21, 2009 | 166 | 8
Z-1993-2009 | 52423 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tibial Tray, component of Ortho Development Balanced Knee System, Nonporous, Sterile R, Size 3, REF 162-1300, Ortho Development, Draper, Utah 84020. |  |  | May 29, 2009 | 2009 | Ortho Development Corporation | Erroneous lot number on 3 of 28 units of Balanced Knee System. Three units from Lot Number 0055245 were mixed in with the recalled lot (0054891). | Mixed-up of materials/components | Consignees were notified via e-mail on May 29, 2009 and June 5, 2009 requesting return of units. Direct questions about the recall to the Ortho Development Corporation by calling 1-801-619-3450. |  | 13 | 13 units | Worldwide -- US (states of FL and TX) and Japan. | Terminated | May 11, 2010 | 347 | 13
Z-1870-2009 | 52605 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | GE Healthcare, Uni-Circuit Kit, Adult, Disposable, 1.8 m/72 in, REF M1089028, M1087972, M1087975, M1126445, M1127508, M1127509, M1129780, |  |  | June 05, 2009 | 2009 | GE Medical Systems, LLC | Breathing circuit inner tube may separate from the rigid connectors under certain conditions. A separation of the inner tube from the connector may prevent adequate gas flow (O2, N2O and anesthetic agents) to the patient. A separation can also cause CO2 re-breathing. This could lead to hypoxia and/or hypercarbia and inadequate anesthesia delivery. | Process control | GE Healthcare sent consignees 6/5/09 an "Urgent Medical Device Correction" letter dated June 5, 2009. The letter was addressed to Health Care Administrator/Risk Manager; Chief of Anesthesia; Director of Clinical Engineering; Risk Manager. The letter described the Safety Issue, Affected Product Details, Safety Instructions (included "Recall Confirmation" form fax to 973-790-4150), Product Correction and Contact Information. |  | 85097 | 85097 | Worldwide distribution: USA, AUSTRALIA, DENMARK, FRANCE, ESTONIA, GERMANY, ITALY, JORDAN, NETHERLANDS, NORWAY, SWEDEN, and UNITED ARAB EMIRATES. | Terminated | January 16, 2014 | 1686 | 85097
Z-2054-2009 | 52434 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EUB-6500/HI VISION 6500 Diagnostic Ultrasound Scanners, DICOM Option EZU-FC5, Software Versions: V01-03, V01-05, V02-04, V03-02, V04-02, V04-03, V04-04, V05-03, V05-05A, V05-07A and V05-30A. |  |  | May 12, 2008 | 2008 | Hitachi Medical Systems America Inc | The firm's parent company discovered a DICOM UID (Universal Identifier) code mix-up which can occur due to a software error. This error can cause the images from two (2) different patient scans to become intermingled and consequently, labeled incorrectly. | Software design | HMSA issued a 'DEVICE CORRECTION' notification to their customers on 5/12/2008. The notification describes the potential data transmission problem (potential to archive studies done on different patients into the same patient folder) and the conditions under which this problem can occur. The notification cautions the user/customer to verify all images stored on the PACS (Picture Archiving and Communications System) in order to ensure that they do not contain images from a different patient. The DEVICE CORRECTION letter also states that the firm will be installing updated software on all affected systems beginning in June, 2008, in order to address this issue. It also states that a Hitachi Service Rep. will be contacting each customer and making arrangements to set up a convenient time to come out and perform a software upgrade. Lastly, the letter provides a toll-free telephone # to the firm's 'Customer Service Center: 1-800-800-4925, which the customer can use to obtain additional information, if desired. |  | 10 | 10 units | The recalled device system units were distributed to customers located in the following states within the U.S ony.: AR, CA, and TX. | Terminated | July 20, 2011 | 1164 | 16
Z-2044-2009 | 52814 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Engenex Cover Strip; Product Code 446471; |  |  | July 29, 2009 | 2009 | ConvaTec | Labeling: The expiration date does not appear on the product label. | Labeling design | Urgent Product Recall letters dated 7/29/09, were hand delivered to all direct customers by ConvaTec Sales Reps. Customers were asked to return any unused products to their supplier. The letter also stated that a credit would be issued for the product. Customers are to contact Karen Krenta at the firm's Customer Interaction Center, 1-800-422-8811. |  | 128 | 128 units | Nationwide Distribution | Terminated | October 02, 2009 | 65 | 456
Z-2051-2009 | 52811 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | REF IA-2000-S, LightWave Suction Ablator, STERILE, Rx ONLY, MADE IN USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec has determined there is a possibility that the device may not shut off if the ablate button is released, which may result in users/patients suffering burns. | Device Design | ConMed Linvatec is notifying their direct consignees by an Urgent Medical Device Recall letter, requesting that customers check their inventory for the affected products; complete, sign, and return the Reply Form; and return of any affected lot numbers of the product. The product is to be returned to ConMed Linvatec. Questions should be directed to Customer Service at 1-800-237-0169. |  | 28074 | 28,074 units | Worldwide Distribution | Terminated | February 25, 2011 | 574 | 28074
Z-2000-2009 | 51896 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Roche/Hitachi Preciset TDM I, Roche Diagnostics, Indianapolis, IN; Catalog No. 03375790 190. |  |  | April 22, 2009 | 2009 | Roche Diagnostics Corp. | The bottle caps may be more difficult to remove than usual, and minor injuries may be caused by the application of the force required to open the bottles. | Packaging process control | Roche Diagnostics Corp. notified Consignees via letter dated April 22, 2009 and instructed to not use excessive force when opening bottles and contact the firm for product replacement instead. A Fax Back Form was included for customers to return via fax to 1-888-345-0480.For further information, contact your Roche Diagnostics representative or call 1-800-428-5074. |  | 3970 | 3,970 bottles of both products | Nationwide Distribution | Terminated | December 07, 2009 | 229 | 3970
Z-1963-2009 | 52107 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD Diagnostics, BD GeneOhm MRSA Assay, REF: 441242 (200 Reaction Kit) and 441244 (48 Reaction Kit). |  |  | November 11, 2008 | 2008 | BD Diagnostics (GeneOhm Sciences, Inc) | BD Diagnostics GeneOhm (BDDG) has determined that the Cepheid SmartCycler reaction tubes packaged in the BD GeneOhm MRSA test kits identified above may be defective. | Other | BD Diagnostics GeneOhm began contacting consignees by phone November 17, 2008 followed by a letter dated November 21, 2008 instructing discontinued use of the affected lots. Consignees were also asked to complete and return the enclosed Field Action Verification form.To arrange for product replacement or for further information, contact BD Diagnostics Technical Service at 1-888-436-3646 extension 2. |  | 114 | 114 kits | Worldwide Distribution -- US (CA, IL, IN, KT, MD, ME, MN, MO, MT, NC, NJ, NY, OK, PA, TN, TX, UT, VA, and WI), Belgium and Japan. | Terminated | November 29, 2011 | 1113 | 114
Z-1987-2009 | 52662 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Hip System Bipolar Head Bipolar Head, 28 mm ID, 59 mm OD |  |  | July 10, 2009 | 2009 | Omni Life Science | The APEX Hip System Bipolar Head implant may not properly capture the femoral head assembly of the bipolar head onto the femoral head. | Device Design | Omni Life Science notified distributors and implanting surgeons by an Urgent, APEX Hip System Bipolar Head Voluntary Recall letter dated June 29, 2009. The method of notification is both e-mail and FedEx overnight. The letter requires a signature acknowledging receiving the letter and the letter is to be faxed back to Omni Life Science. The Distributors/Sale Representatives are instructed to fill in the "Quantity Returned" on the recall letter and fax back to Omni. They are to ship all existing Apex Hip System Bipolar Head inventory by 07/10/2009 to the directed address. |  | 17 | 17 units | Nationwide Distribution -- Including states of AZ, CA, CO, FL, IL, MI, NJ, OK, UT, VA, and WI. | Terminated | June 28, 2012 | 1084 | 295
Z-2040-2009 | 52114 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Leckey 4-Point Pelvic Harness Belts, catalog 081326990; James Leckey Design Ltd., Kilwee Business Park, Upper Dunmurry Lane, Belfast, United Kingdom; Distributed by Patterson Medical; sold as replacement belts and as a component of the following adjustable mechanical chairs: |  |  | April 17, 2009 | 2009 | Patterson Medical Holdings, Inc. | The plastic buckles on the hip belts of the 4-Point Pelvic Harness may break. | Nonconforming Material/Component | Sammons Preston, the U.S. agent for James Leckey Design Ltd., issued recall letters dated 4/17/09 to all of their customers who purchased the Leckey 4-Point Pelvic Harness, used in the Leckey Contoured Advance Seat, Leckey Early Sitting System and the Leckey Squiggles Saddle Seat, informing them that the manufacturer, James Leckey Design Ltd., is recalling the harness because the pelvic belt buckle may not lock securely and/or break. A copy of the Urgent Field Safety Notice dated 1/28/09 issued by Leckey in the United Kingdom was included with the Sammons letter. The customers were advised to be vigilant and check the hip belts daily to ensure that the user is secure in their seating system until the replacement hip belt is received. The accounts were requested to contact Carla Dreher at 1-800-288-3693, ext. 8724 for instructions on how to receive a replacement hip belt. |  | 226 | 226 belts | Nationwide and Canada | Terminated | May 13, 2011 | 756 | 226
Z-2033-2009 | 52383 | No | ECX | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Empty D aluminum oxygen cylinders packaged in shipping containers labeled as: Product #0022-58 P, Desc CYL 02 ALUM 400 LITER, Catalog #31-10-2014. The brand and firm name on the label affixed to the cylinder is Chemetron, Allied Healthcare Products Inc., St. Louis, MO. |  |  | May 15, 2009 | 2009 | Allied Healthcare Products Inc | Medical valves do not meet quality specifications which may cause the valve to fail and result in uncontrolled release of oxygen when the valve is in the open position. | Nonconforming Material/Component | The recalling firm first notified their customers via a letter dated 5/15/09 issued regular mail explaining the reason for recall and requesting the cylinders with specific post valves be returned. A second letter dated 6/1/09 was issued via regular mail which provided specific shipping dates that each customer received the suspect product. Neither of these two letters requested subrecall. A third letter dated 6/17/09, which included copies of the first two letters, via certified mail. This letter informed the customer the recall was to be conducted to the user level and requested their customer conduct a subrecall. Requested to remove the cylinder from service and call Allied Healthcare Products for return and a replacement at 1-800-268-1661. |  | 177 | 177 cylinders | Class II Recall Worldwide Distribution including states of NY, NC, KY, PA, WI, NH, NJ, CA, CT, FL and Singapore. | Terminated | April 23, 2012 | 1074 | 288
Z-1937-2009 | 52352 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | EASYMax N. Self-Monitoring Blood Glucose System, distributed in cardboard carton kits. Each kit contains the following components: Blood Glucose Meter; Test Strip x 10 ct.; Normal Control Solution; Carrying Case; Alkaline Battery x 2 ct.; Lancets x 5 ct.; Lancing Device; AST Lancing Device Cap; User's Manual; Test Strip Instructions and Control Solution Instructions. |  |  | May 26, 2009 | 2009 | Oak Tree International Holdings, LLC, Inc. | Blood glucose meters, which were set to the mmol/L. unit of measure, rather than the accepted/labeled mg./dL unit of measure, were distributed. | Process control | EPS Technology issued a certified notification letter to the sole distributor/customer: GEMCO, Coconut Creek, FL dated May 26, 2009. The letter describes the potential preset error and its possible consequences to the user/customer. The letter goes on to inform the distributor that EPS Technology Corp. will be sending out letters to all affected user/customers during a direct mailing. If the error is found, the user/customer can either: 1) have the firm's Customer Service Dept. instruct them as to how to fix the unit, or 2) request that a replacement meter be sent to them.For further information, contact EPS Technology Customer Service at 1-866-994-3345. |  | 5396 | 5,396 units | Nationwide Distribution -- AL, AR, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MD, ME, MI, MS, NC, NE, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, VA, WA, WY and HI. | Terminated | September 27, 2012 | 1220 | 5396
Z-1997-2009 | 52614 | No | KKO | Teething ring. | Dental | Dental | 510(K) Exempt | MDI Hybrid Implant, O-Ball abutment, sizes 2.9mm x 13mm and 2.9mm x 15mm, endosseous dental implant. |  |  | July 07, 2009 | 2009 | IMTEC Corporation | Size printed on product packaging does not match size of enclosed product. | Other | A notification letter dated July 7, 2009 was issued to customers. The letter described the issue and instructed customers do not use the affected product and call IMTEC, a 3M Company, at 1-800-879-9799, extension 240 for assistance. Customers were further instructed to complete the enclosed Recall Response Fax-Back Form with information on their inventory (whether or not the recalled product is in their stock). The firm's representative will follow-up with customers if no response within three business days of receiving the notification letter. |  | 605 | 605 units | Worldwide -- US, Azerbaijan, Guatemala, Turkey, and Germany. | Terminated | December 02, 2009 | 148 | 605
Z-2073-2009 | 53129 | No | LKK |  |  | General Hospital | PMA | Medtronic Indura 1P Intrathecal Catheter, model 8709SC, Method of Sterilization: Ethylene Oxide. Medtronic, Inc., Minneapolis, MN 55432-5604 USA. The implanted infusion system components consist of a Medtronic pump and a model 8709SC Catheter. The catheter connects to the pump at the catheter port. Contents: 81.4-cm Catheter with length markers and guide wire, 15 T-guage introducer needle, Sutureless pump connector with attached 7.6-cm catheter and connector pin, Transparent strain-relief sleeves, Anchors. |  |  | August 27, 2009 | 2009 | Medtronic Neuromodulation | Current labeling for Medtronic Sutureless Connector (SC) catheters and revision kits (hereafter referred to as SC catheters ) incorrectly states that SC catheters are compatible with Medtronic IsoMedÂż constant-flow infusion pumps. SC catheters are compatible with Medtronic SynchroMedÂż II and SynchroMed EL pumps. Medtronic has determined that SC catheters are not compatible with IsoMed pumps. | Device Design | A "Medical Device Correction Notification" letter was issued on August 14, 2009 to implanting and managing physicians. The letter described the nature of the issue, scope of affected product, potential severity of the issue and provided recommendations. Physicians were instructed that Sutureless Connector (SC) catheters must not be used with IsoMed pumps for future revisions and implants. A copy of the letter was also sent to the Risk Manager at each hospital associated with physicians using SC catheters. Healthcare professionals were asked to complete and return the enclosed reply card.Direct questions to your Medtronic field representative or contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. Patient Management information is also available at http://www.professional.medtronic.com under the Heading, "Advisories". |  | 20000 | 20,000 of all products | Worldwide Distribution -- US (including Washington D.C. and Puerto Rico), Algeria, Aruba, Australia, Austria, Belgium, Canada, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Poland, Portugal, Reunion, Romania, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom and Vatican City. | Terminated | December 12, 2011 | 837 | 20000
Z-2028-2009 | 52810 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360 Orbital Atherectomy Device, 1.50 LT Classic, Model #DB-150L, CSI Part # 70025-02 |  |  | April 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is correcting (8) packages of Diamondback 360 due to a missing "Use By Date" on the device labeling. The labelling should have been a 12 month shelf life (2010-02) rather than no information for shelf life. The units affected by this correction are sterilized products, and CSI believes that there is no immediate risk to health associated with these distributed products recognizing the valid s | Incorrect or no expiration date | Cardiovascular System, Inc. issued an "URGENT: Medical Device Correction" letter dated April 16, 2009. The letter described the product and the problem. The letter informed the consignees that the "Device" re-labeling would be performed by company representatives. Devices will not be returned.For further information, contact Cardiovascular System, Inc. 1-877-274-0901. |  | 8 | 8 | Nationwide Distribution -- CA, GA, MI, NY, NY and TN. | Terminated | December 10, 2011 | 968 | 8
Z-1938-2009 | 51281 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Mirror assembly to magnetic resonance device (1.5T Atlas SPEEDER Head-Neck Coil) Coil Part #Q7000021, Quality Electrodynamics, Mayfield Village, OH 44143. |  |  | December 29, 2008 | 2008 | Quality Electrodynamics Llc | The firm received a complaint indicating that the mirror is high risk and could drop on a patient. | Device Design | Quality Electrodynamics, LLC notified Customers via email beginning December 29, 2008 requesting the affected device be returned to the firm for improvement.For further information, contact Quality Electrodynamics, LLC at 1-440-484-2228. |  | 5 | 5 units | Nationwide Distribution: IL and OH. | Terminated | April 25, 2012 | 1213 | 5
Z-2011-2009 | 51726 | No | ECX | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Post Medical Valves VO870A-Y and VO870-A-Y-T PSI 2015, PSI 2015, CGA 870, Inlet .750 STDS, Gas services (Oxygen), PSI 2015, Aluminum Cylinders, InterMed Gas Products LLC Industrial Division, 2056 Vista Parkway West Palm Beach, FL 33411. |  |  | March 27, 2009 | 2009 | Intermed Gas Products Corporation | Firm received reports from customers that valves malfunctioned when installation and/or filling process. Customers forwarded pictures of the defect where the top portion of the packing nut broke off from the threaded portion of the net; which may cause the valve to fail and result in uncontrolled release of oxygen when the valve is in the open position. | Component design/selection | Intermed Gas Products Corporation notified customers of the affected product by phone beginning March 27, 2009 and followed up with a formal "Product Recall Notice" letter dated March 31, 2009. All affected product should be located and returned to the firm for replacement.For further information, contact Intermed Gas Products at 1-561-483-1364. |  | 8767 | 8,767 | Nationwide Distribution -- FL, SC, TX, MO, TN, NC, NY, GA, MN, AZ. | Terminated | May 17, 2011 | 781 | 8767
Z-2039-2009 | 52700 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Advanced Bionics HiRes 90K Cochlear Implant, Model: CI-1400-01 and CI-1400-02H, Implant Magnet Case. |  |  | June 10, 2009 | 2009 | Advanced Bionics LLC | The recall was initiated after Advanced Bionics became aware of fourteen (14) devices manufactured between May 27, 2005 and June 30, 2005 that are confirmed to have or potentially have an issue related to the internal magnet. These devices are were made with an internal magnet where the metal case surrounding the magnet was not properly welded. Patients with these devices may experience a decreas | Process control | Advanced Bionics, LLC issued an "Important Notification" dated June 10, 2009 via Fed Ex to Consignees in the US and Canada. Outside of the US, the notification letters will be distributed in such a manner that delivery confirmation can be obtained. There were three type of notification letters distributed: one for Patients, one for Surgeons and one for Clinicians. Letters described the affected device prompting users on what to do next. An acknowledgement form was enclosed and should be completed and sent via fax to 1-661-362-7621.For further information, contact Advanced Bionics at 1-877-829-0026. |  | 18 | 18 | Worldwide Distribution -- US, Canada, France and South Korea. | Terminated | March 26, 2012 | 1020 | 18
Z-2211-2009 | 53124 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Integrated ECG Cable are supplied as part of the start-up kits provided with some Spacelabs' monitors. The kit PNs are: 718-0017-10, 718-0017-11, 718-0017-12, and 718-0017-13. The cables are also offered as replacement parts under Spacelabs PNs: 700-0008-22, 700-0008-23, 700-0008-24, and 700-0008-25. |  |  | July 23, 2009 | 2009 | Spacelabs Healthcare, Incorporated | The Spacelabs Medical Integrated ECG Cables, PN: 700-0008-22, 700-0008-23, 700-0008-24, and 700-0008-25 will not withstand the high voltage of a defibrillation shock. | Process control | Spacelabs Healthcare issued an "Urgent: Medical Device Recall" dated July 20, 2009. Consignees are being sent replacement ECG cables for each item ordered and are requested to replace the affected cables. Also, they are instructed to destroy and dispose of the recalled cables according to their own protocol or local regulations.For further information, contact Spacelabs Healthcare Technical Support at 1-800-522-7025, option 2 (US) or +86-512-87171888, extension 304 (outside US). |  | 1395 | 1,395 | International Distribution -- Bangladesh, Turkey, Trinidad and Tobago, Romania, Guatemala, Algeria, Netherlands, Egypt, Chile, Venezuela, Spain, Colombia, India, the United Kingdom, Singapore, France and Ecuador. | Terminated | July 21, 2011 | 728 | 1395
Z-2117-2009 | 52713 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone Screw; |  |  | June 26, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Raw material supplied by a third party used to manufacture various medical devices does not comply with metallurgical requirements outlines in ASTM standard for titanium surgical implants. | Nonconforming Material/Component | Stryker Orthopaedics issued "Urgent Product Recall" letters via Federal Express priority overnight beginning June 26, 2009 to Branches, Hospital Risk Management and Chief of Orthopaedics and Surgeons. All users are requested to examine their inventory, retrieve all affected product and return for reconciliation to Stryker Orthopaedics. Users are also requested to complete the enclosed Product Accountability Form and fax with necessary information to the firm at 1-201-831-6069.For further information, contact Stryker Orthopaedics at 1-201-831-5970. |  | 51113 | 51,113 total - all sizes. | Worldwide Distribution | Terminated | July 27, 2010 | 396 | 51137
Z-1810-2009 | 52643 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Delivery System, a subsystem of the CyberKnife Robotic Radiosurgery System, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | June 22, 2009 | 2009 | Accuray Inc | Targeting accuracy out of specification, Error alert does not render the system down, which may result in mistreatment in the wrong area. | Software design | Accuray, Inc. issued an "Urgent Device Correction" notice dated June 22, 2009. The consignees were informed of the defect including methods for detecting the error. The firm is in the process of developing a software patch to address the issue.For further information, contact Accuray, Inc. Customer Support at +1-877-668-8667 (US) or +1-408-716-4700 (outside US) or customersupport@accuray.com. |  | 61 | 61 units | Worldwide Distribution -- US including states of MA, FL, AL, TN, DE, CA, NJ, OK, CO, NC, IL, PA, MD, DC, MS, CN, NV, RI, SC, UT, VA, MI, NE, MO, TX and AL and countries of Canada, Russia, UK, Ukraine, Saudi Arabia, Switzerland, Turkey, Taiwan, India, China, Korea and Thailand. | Terminated | January 05, 2011 | 562 | 61
Z-2111-2009 | 50802 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow Corporation Easy Pump LT 60-24 (60 ml x 2 ml/hr), REF: 04434366, assembled in Mexico. |  |  | February 11, 2008 | 2008 | I-Flow Corporation | I-Flow Corporation received two complaints involving Easypump LT 60-24 (Lot 612996) containing a 100 ml x 2 ml/hr pump in the sterile package instead of a 60 ml x 2 ml/hr pump (Illustration was included to show consignee what both the correct and incorrect part looks like). Lot 612996 was manufactured on January 6, 2006. As of February 7, 2008, there have been no further complaints received on t | Labeling mix-ups | I-Flow Corp. issued an "Important Recall Information" letter dated February 7, 2008 informing Consignees to quarantine and return any remaining inventory and notify any end users of the affected lot.For further information, contact I-Flow Corp. at 1-949-206-2653. |  | 4450 | 4,450 Distributed | Nationwide Distribution to B. Braun, Spain, who then further distributed to Italy, Germany, Netherland, Latvia, Spain, Portugal and Singapore. | Terminated | February 04, 2010 | 724 | 4450
Z-2030-2009 | 53035 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Medtronic Physio-Control LIFEPAK CR Plus Defibrillator/Monitor. |  |  | July 31, 2009 | 2009 | Physio Control, Inc. | The K1 relay on the Analog printed circuit board assembly may short due to moisture ingress in a highly humid environment. K1 shorting may affect the ECG amplitude causing the Shock Advisory System(SAS) to not analyze the ECG rhythm correctly thus preventing the device from delivering therapy. | Nonconforming Material/Component | Physio-Control, Inc. notified Customers by phone beginning August 18, 2009. In addition, the recalling firm Issued an "Urgent Medical Device Recall" letter dated August 2009 informing Customers of the affected product and informing them that a Physio-Control representative will call and arrange for immediate device replacement.For further information, contact Physio-Control Customer Care at 1-800-442-1142. |  | 34 | 34 devices (2 domestically &amp; 32 internationally) | Worldwide Distribution -- US (DC and MN) and Japan. | Terminated | March 15, 2013 | 1323 | 34
Z-2179-2009 | 50678 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Access Systems, Inc., Gauderer Genie* II PEG (Percutaneous Endoscopic Gastrostomy) Safety System Ponsky "Pull" Placement Technique REF 000373, 20F, STERILE. |  |  | December 04, 2008 | 2008 | C.R. Bard Access Systems, Inc | Snare can detach in patient. The Polypectomy Snares and the Entake Peg System (which contain these snares) are being recalled because the looped portion of the snare can become detached and drop into the stomach or other internal canal during use, necessitating retrieval of the remaining piece. | Nonconforming Material/Component | Bard contacted all 1,883 worldwide customers who received the distributed suspect 467 product code lots, comprising 177, 614 units that were either not expired or were still available in the field. These will be returned by the customer to Bard Access Systems. 1-800-290-1689. |  | 177614 | 177,614 units | Worldwide, including USA, Belgium, Canada, Australia, and Japan. | Terminated | November 09, 2011 | 1070 | 177614
Z-2027-2009 | 52617 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Orthopaedics PIP (proximal Interphalangeal) FIXATION HINGE KIT, REF 101638, for external fixation of finger joint. |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | Hinge failure: The product is cracking or breaking during use. | Nonconforming Material/Component | The Smith &amp; Nephew Sales Representatives were notified of the problem and the recall via certified letter and e-mail on 01/23/2009. A second notice was sent to the hospitals on 06/06/2009. All affected Smith &amp; Nephew International Distributors were notified via email on 01/23/2009. They were instructed to immediately notify their accounts concerning the recall. Smith &amp; Nephew c/o Return Goods, 4287 Delp St, Memphis, TN 38118 |  | 7089 | 7,089 units | Worldwide Distribution: USA, Turkey, Denmark, Germany, Malaysia, Hong Kong, Korea, Switzerland, Japan, Singapore, Canada, New Zealand, Australia, and Italy. | Terminated | July 29, 2013 | 1648 | 7089
Z-2109-2009 | 53179 | No | NAJ | Vascular embolization device. | Cardiovascular | Obstetrics/Gynecology | 510(k) | Boston Scientific, Contour SE Microspheres, for tumor treatment. |  |  | August 17, 2009 | 2009 | Boston Scientific Corporation | Compromised sterility: Boston Scientific is initiating a Recall Removal for twenty (20) UPN codes and 638 associated lots/batches of its Contour SE Microspheres product. All affected product is packaged within a vial. Boston Scientific has determined that, for affected products, the sterile barrier in the packaging (pouch) that contains the vial may be breached. The breach of sterility could l | Process design | Boston Scientific sent consignees an "Urgent Medical Device Recall Removal" letter dated August 28, 2009. The letter was addressed to Risk Manger / Field Action Contact. The letter described the problem and the product and advised the consignee to cease immediately the distribution and use of the product. Requested consignees to return affected product to Boston Scientific. The letter enclosed the Recall Instructions, Reply Verification Tracking Form and Field Action Return Shipping Label. |  | 45724 | 45,724 (13,048 US, 45,724 OUS) | Worldwide distribution: USA, United Arab Emir., Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Colombia, Costa Rica, Czech Republic, Germany, Denmark, Algeria, Ecuador, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, South Korea, Kuwait, Lebanon, Latvia, Libya, Mexico, Namibia, Nicaragua, Netherlands, Norway, New Zealand, Oman, Panama, Poland, Puerto Rico, Portugal, Russian Fed., Saudi Arabia, Sweden, Slovakia, Tunisia, Turkey, Taiwan, Ukraine, and South Africa. | Terminated | December 18, 2011 | 853 | 45724
Z-2220-2009 | 48960 | No | MTG |  |  | Immunology | PMA | Beckman Coulter Access Hybritech free PSA Calibrators, REF 37215. |  |  | January 08, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter incorrectly shipped Access Hybritech free PSA calibrators containing dual calibration cards (Hybritech and WHO) to your lab. The WHO calibration for free PSA is marketed without a PMA. | Other | Beckman Coulter contacted customers by phone beginning January 11, 2009 and sending a Product Corrective Action letter informing them to discontinue use and discard the referenced product immediately.For further information, contact Customer Technical Support at 1-800-854-3633 or contact your local Beckman Coulter Representative. |  | 22 | 22 Kits | Nationwide Distribution -- OH, CA, TX, OR, MD, MO, FL and PA. | Terminated | November 29, 2011 | 1421 | 22
Z-1931-2009 | 52200 | No | DSQ |  |  | Cardiovascular | PMA | iPulse Circulatory Support System |  |  | April 23, 2009 | 2009 | Abiomed, Inc. | iPulse Console SC1035 displayed a "Low Pressure / Low Flow" alarm and stopped pumping during patient transport in battery operation. | Device Design | An "Urgent Voluntary Medical Device Field Correction Notice" was issued April 8, 2009 to affected customers. The customers were notified that the iPulse Console SC1032 displayed a "Low Pressure/ Low Flow" alarm and stopped pumping during patient transport in battery operation. Instructions were provided to users in case they encounter this failure mode. ABIOMED Field Service Engineers will contact customers to schedule a time to visit site and correct affected consoles. Please contact ABIOMED's Field Service at 1.800.554.8666, if you have any questions. |  | 79 | 79 | Domestic: MA, MI, SC, NY, NJ, OH, CA, PA, AR, MI, MD, CT, IL, WI, TX, NJ, TN, IN, MI, DE, FL, IA, WA, KY, and WVForeign: Germany, Greece, Italy, Latvia, Spain, and Turkey | Terminated | June 05, 2014 | 1869 | 79
Z-1886-2009 | 52658 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Trio Connector Hex Socket for Torque Wrench 3.5MM |  |  | July 01, 2009 | 2009 | Stryker Spine | The Trio 3.5MM Hex socket has been reported to break under high stresses. | Other | Product Correction letters were sent on July 1, 2009 to all Branches/Agencies, Surgeons and Hospital Risk Management OR Supervisor by Federal Express. The letter indicates the potential hazards associated with the breakage and it reinforces the appropriate surgical technique in using the hex socket. Customers are requested to complete an acknowledgement form and return it via FAX to Aminah Crawford at 201.760.8370 within five (5) days of receipt of the notice. If you have questions, please contact Tiffani Rogers, Regulatory Compliance Manager at 201.760.8206. |  | 143 | 143 | Nationwide - OR, AR, CA, MA, KY, MD, GA, FL, NY, NM, UT, FL, OK, KS, MI, MN, IN, MT, TX, WI, and TN. | Terminated | October 20, 2009 | 111 | 143
Z-2005-2009 | 52449 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche ACCU-CHEK Performa blood glucose test strips, Roche Diagnostics, Indianapolis, IN. |  |  | June 10, 2009 | 2009 | Roche Diagnostics Corp. | There may be a significant negative bias when testing neonate samples. | Process control | Roche Diagnostics notified consignees in the affected countries of the problem by "Urgent Field Safety Notice" dated June 10, 2009 and instructed to cease using the Accu-Chek Performa and Accu-Chek Inform II systems on neonates, except for systems using test strips with advanced chemistry. For questions and additional information, contact your local Roche Diagnostics sales affiliate or sales representative. |  | 1518 | 1,518 of total products | International Distribution -- Argentina, Australia, Brazil, Chili, China, Colombia, Ecuador, France, Germany, Great Britain, Hong Kong, India, Korea, Malaysia, Mexico, New Zealand, Pakistan, Panama, Peru, Philippines, Singapore, South Africa, Taiwan, Thailand, Uruguay and Venezuela. | Terminated | October 29, 2010 | 506 | 1518
Z-1892-2009 | 52704 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView Gamma Camera Radius Drive Assembly, manufactured by Philips Medical Systems (Cleveland) Inc., Cleveland, OH, 3/8" tilt, model number 4535604 34811, 3/4"tilt, model number 4535674 34801, and 3/8" fixed, model number 4535604 34781, all catalog number 882480 |  |  | July 02, 2009 | 2009 | Phillips Nuclear Medicine | Unexpected movement-- if the drive assembly is severely compromised, the detector may slide to its hardware limit. | Process control | Firm initiated its recall notification on July 6, 2009 issuing a field safety notice to its consignees via FedEx and internationally through its personnel. Responses will be tracked for effectiveness. Questions may be directed to Philips Customer Care Center at 1.800.722.9377. |  | 288 | 288 units | Worldwide | Terminated | April 12, 2010 | 284 | 288
Z-2009-2009 | 52766 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Cobham AMOGS Advanced Medical Oxygen Generating System, 70309-03900-101. Concentrator, Oxygen, Molecular Sieve, S/A P/N: 70309-03900-101, Mfg P/N 3261165-0101. The firm name on the label is Carleton Life Support Systems Inc., Davenport, IA. |  |  | July 24, 2009 | 2009 | Carleton Life Support Systems Inc | The backup oxygen supply control valve on the oxygen concentrator could fail resulting in it sticking in the open or closed positions. | Device Design | Cobham Life Support Systems, Inc. issued letters to their two customers via e-mail. One customer letter was dated July 24, 2009 and the other was dated July 28, 2009. Both letters notified the customer of the reason for the recall and that the usage of the concentrator must be restricted to exclude actual use with a patient until corrected. The letter informs customers they are redesigning the backup supply control valve and will coordinate the retrofit in the future. For further information, contact Cobham Life Support at 1-563-383-6421. |  | 23 | 23 units | Nationwide Distribution -- CT and AL. | Terminated | March 22, 2012 | 972 | 23
Z-1940-2009 | 52823 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RICAÂż Sher-I-Swiv &amp; Sher-I-Swiv F/O, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | July 31, 2009 | 2009 | Teleflex Medical | The tether attaching to the double swivel cap may partially or completely break at the attachment points, resulting in the potential of the accidental aspiration of the tether into the patient's lungs. | Process control | Teleflex Medical, Inc. notified consignees via overnight mail starting July 31, 2009 requesting acknowledgement of the recall and a request to return all of the affected product. Distributors were instructed to forward the letter to their consignees and retrieve relevant product from those locations. Affected product received by Teleflex Medical will be destroyed or reworked and replacement product or credit will be provided to the consignee. For further information, contact Teleflex Medical at 1-919-433-4986. |  | 569465 | 569,465 (total for all types) | U.S. Nationwide, Canada, Mexico, Europe, India, Asia Pacific Rim, South and Latin America. (Including: Malaysia, Hong Kong, Vietnam, Singapore, Thailand, Korea, China, Japan, Philippines, Australia, New Zealand, France, Germany, Ireland, Sweden, Costa Rica, Puerto Rico, Jamaica, Brazil, Argentina, Uruguay and Colombia). | Terminated | September 15, 2011 | 776 | 569465
Z-2212-2009 | 53135 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drager Surgical Lights Sola 500 and Sola 700. |  |  | August 26, 2009 | 2009 | Draeger Medical, Inc. | welded seam may fail | Nonconforming Material/Component | Draeger Medical, Inc. issued a "Medical Device Recall" letter dated August 2009 via certified mail with return receipt. Consignees were informed on the affected product and that a service representative will be in contact shortly to schedule a replacement.For further information, contact Draeger Medical at 1-800-543-5047, press 1 at the prompt and then 2349. |  | 20 | 20 units | Nationwide Distribution -- IL, KY, and PA. | Terminated | March 08, 2010 | 194 | 20
Z-2219-2009 | 51309 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Ventri, GE Healthcare, Waukesha, WI. |  |  | November 03, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare become aware of an incident which caused a slight patient injury during his unload following a prone-oriented scan on Ventri. Although clear safety instructions were not followed in this case, additional measures will be implemented in an effort to reduce the probability or repeat occurrence. | Process control | GE Healthcare issued a "Product Safety Notification" dated January 2009 to consignees. The letter was addressed to Hospital Administrators, Managers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter describes the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. A GE Field Engineer will visit each site and implement the correction.For further information, contact your local GE Healthcare service representative at 1-800-437-1171 (US or Canada) or 1-262-896-2890 (outside the US). |  | 308 | 308 (252 US, 57 OUS) | Worldwide Distribution -- US (AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, KY, LA, ME, MD, MA, MI, MS, MO, NE, NV, NH, NJ, NY, NC, OH,OK PA, RI,SC, TN, TX, UT, VA, WA, WI and WV), AUSTRALIA, BRAZIL CANADA, DENMARK, FRANCE, GERMANY, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY,SWEDEN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | July 12, 2012 | 1347 | 308
Z-1939-2009 | 52769 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | The BladderScan BVI 9400 is a portable ultrasound instrument that provides a noninvasive measurement of urinary bladder volume. The device consists of an ultrasound probe that scans the patient's bladder, and a battery operated console that provides measurement related information. The BVI 9400 measures ultrasonic reflections on multiple planes inside the body and produces a three dimensional image. Based on the image, the BVI 9400 calculates and displays bladder volume. |  |  | May 18, 2009 | 2009 | Verathon, Inc. | Three specific issues are being addressed:1. Calculation and use of an incorrect year in dates after 12/31/2009.2. Double scans, continuous scans, and double printing.3. Calibration sensitivity. | Software design | U.S. customers (including Government) were sent a post card on or about 5/18/2009 and were called about the upgrade. Units are to be upgraded by the customers using an online tool called ScanPoint. Customers may elect to return their units to the firm for upgrading.Canadian customers were phoned by Verathon.For foreign units sold and distributed through one of the recalling firm's related firms, Service Bulletin SB-0014 was sent to the distributors located in The Netherlands, United Kingdom, and Australia.Contact Verathon Customer Care at 1-800-331-2313 if you have any questions. |  | 1978 | 1,978 | Worldwide - US, Canada, Malaysia, Hong Kong, Singapore, Argentina, New Zealand, Barbados, Dominican Republic, and Costa Rica. | Terminated | October 07, 2010 | 507 | 1978
Z-2019-2009 | 52862 | No | KGC | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard * Button Device Continuous Feeding Tube with 90 degree Adaptor, 18F x 1.7cm, NON STERILE, REF 000282, Bard Access Systems, Inc., Salt Lake City, UT 84116 USA. |  |  | January 20, 2009 | 2009 | C R Bard Inc | Possible restriction/occlusion in adaptors/tube accessories causing difficulty to infuse during routine feeding treatment. | Equipment maintenance | Customers were notified by letter and instructed to not use or further distribute any of the affected lot numbers. Customers were further instructed to return the recalled products to Bard Access Systems (BAS) through a return goods authorization number. Questions and requests for assistance should be directed to BAS Customer Service at 800-290-1689. |  | 4381 | 13, 337 total of all products | Worldwide Distribution -- US, Canada, Austria, Australia, Belgium, Finland, France, Germany, Great Britain, Holland, Italy, Norway, Spain, Sweden and Switzerland, Japan, and Latin America. | Terminated | December 09, 2009 | 323 | 4381
Z-2225-2009 | 50528 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Sytec 6000, Sytec 8000, and Lemage/ProSeed/ProSpeed families. |  |  | December 01, 2008 | 2008 | GE Medical Systems, LLC | During internal testing, GE Healthcare has recently determined that the collimator used on our CT/e and HiSpeed X/i families of Computed Tomography systems does not comply with the specific requirements in United States Title 21 of the Code of Federal Regulations, and IEC 60601-1-3, related to leakage radiation from the diagnostic source assembly, which includes the collimator. The excess leak | Radiation Control for Health and Safety Act | GE Healthcare issued an "Important - Electronic Product Radiation Warning" to affected customers. GE Healthcare field service representative will perform a field corrective action on all affected units. All potential users of the device should be made aware of the notification and the information contained.For further information, contact GE Healthcare at 1-800-437-1171 (US) or by contacting your local GE Healthcare service manager. |  | 643 | 643 units ( 136 US; 507 OUS) | Worldwide Distribution -- UNITED STATES (AL, AZ, AR, CA, CT, DC, FL, GA, IL, IN, KS, KY, LA, MD, MA, MI, LMN, MS, MO, NE, NJ, NM, NY, NC, OH, OK, PA, PR,SC, SD, TN, TX, VA, WA, WV and WI), ALGERIA, ANTIGUA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BELGIUM, BENIN, BOLIVIA, BOSNIA &amp; HERZEGOVINIA, BRAZIL BULGARIA, BYELORUSSIAN, CAMERON, CANADA, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK , DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COST, JAPAN, JORDAN, KAZAKHSTAN, KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, , MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITANIA, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGER, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL QATAR, ROMANIA, RUSSIAN F=FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND , SYRIA ARAB REPUBLIC, TAIWAN, THAILAND, TRINIDAD &amp; TOBAGO TUNISIA, TURKEY,, UGANDA, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, YEMEN and YUGOSLAVIA. | Terminated | January 27, 2016 | 2613 | 7648
Z-2223-2009 | 52218 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BioTek Gen5CL Microplate Data Collection and Analysis Software. |  |  | March 12, 2009 | 2009 | BioTek Instruments, Inc. | Software design allowed false negatives to be read when wells are dark. | Software design | BioTek Instruments, Inc. issued a "Product Recall Notification" dated March 13, 2009 information customers of the affected product. Follow-up letters with software installation instructions were sent March 25, 2009.For further information, contact BioTek Instruments, Inc. at 1-888-451-5171 extension 405. |  | 13 | 13 | Worldwide Distribution -- United States (MN, CA, IL, MA, MI, CA, GA, FL, AZ and TX) and Peru. | Terminated | May 18, 2015 | 2258 | 13
Z-2222-2009 | 51275 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Abbott AxSYM Cancer Assay Disk Version 5.0, for in vitro diagnostic use with the AxSYM System; Abbott Laboratories, Abbott Park, IL 60064; List Number 3D50-05. |  |  | February 20, 2009 | 2009 | Abbott Laboratories | The AxSYM Cancer Assay Disk (LN 3D50-05 version 5.0) contains an AxSYM AFP assay file where the Low Concentration flag for the AFP 1:101 automated dilution protocol the instrument to report results between 0 ng/mL and 0.4 ng/mL (40.4 ng/mL accounting for the 1:101 dilution). The sensitivity of the assay is 0.4 ng/mL. | Device Design | Abbott sent "Product Correction Immediate Action Required" letters dated February 2, 2009 to all AxSYM AFP customers. Customers were requested to install the updated version of the AxSYM AFP assay using version 6.0 of the AxSYM Cancer Assay Disk provided with the letter and to retain a copy of this communication for their records. Any questions regarding the letter were directed to Customer Service at 1-877-4ABBOTT. |  | 8205 | 8,205 disks | Worldwide Distribution -- United States (including Puerto Rico), Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Cayman Islands, Chile, Columbia, Costa Rica, Curacao (Neth Ant), Dominican Republic, Germany, Guatemala, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Trinidad and Tobago, Turks and Caicos and Venezuela. | Terminated | August 23, 2010 | 549 | 8205
Z-2221-2009 | 50488 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | October 27, 2008 | 2008 | Abbott Laboratories, Inc | When an ARCHITECT c16000 System is connected to an ACCELERATOR APS System, the software on the ACCELERATOR APS System may not generate an error warning when a Sample Presentation/Sample Queue error occurs, which could result in contamination of subsequent samples. | Software design | Abbott Laboratories notified Customers of the affected device via a "Product Correction" letter dated October 27, 2008. Consignees provided with a temporary work around until a service representative upgrades the system software.For further information, contact your local Abbott Laboratories, Inc. service representative. |  | 15 | 15 units | International Distribution -- Spain, Italy, the United Kingdom, Sweden, Switzerland, the Netherlands, France, Canada and Scotland. | Terminated | December 02, 2009 | 401 | 15
Z-2226-2009 | 52774 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical 15x100mm Kii Optical Access System, Non-Threaded 6/Box, Model Number: C0R36, Labeled Sterile. |  |  | June 04, 2009 | 2009 | Applied Medical Resources Corp | Applied Medical is conducting a coluntary recall of two lot numbers, 1084426 and 1085506, of the 15mm Kii non-threaded optical access system, model C0R36, due to a potential weakness in the cannula tip, possibly resulting in tip breakage. | Component design/selection | Applied Medical Resources Corp. notified Customers via an overnight "Voluntary Recall Notification" Letter dated June 9, 2009. The letter informed customers of the reason for recall and asked that the listed product be returned. The recall letter provided customers with a contact and telephone number for product return questions (1-949-713-8652) and a separate contact for regulatory questions (1-949-713-8369). |  | 934 | 934 units | Class II Recall - Worldwide -- United States, Australia, New Zealand, France, Canada, Norway and Great Britain. | Terminated | July 14, 2011 | 770 | 934
Z-2213-2009 | 52230 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Epic II critical care bed, 115v, Stryker Medical, Portage, MI; Model 2030. |  |  | August 10, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The backrest of the bed may not stay in an upright position (may drift down when left raised). | Device Design | Stryker Medical issued an "Urgent - Medical Device Correction" letter dated August 10, 2009 informing customers to locate the affected product listed and verify whether it is functioning properly. In the case the device is not properly functioning, remove from service and contact the firm as soon as possible at 1-800-STRYKER, option 3. Customers are also asked to return an enclosed return receipt post card and also forward the notification to users who have loaned or purchased the device. A representative from Stryker field service will be contacting you to perform device corrections shortly.For further information, contact Stryker Medical at 1-269-324-6527. |  | 12599 | 12,599 total products | Worldwide Distribution -- United States, Brazil, Canada, Chile, China, India, Korea, Latin America, Mexico, Singapore and Switzerland. | Terminated | April 25, 2011 | 623 | 12599
Z-2228-2009 | 53121 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical FAST-CATH Hemostasis Introducer 7F, REF 406702, Length 12 CM, Lot 6372 Max. Guidewire O.D. .038", 2010-01, Manufacturer, St. Jude Medical, 14901 DeVeau place, Minnetonka, MN 55345-2126, Sterile EO. |  |  | August 17, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Fast-Cath Hemostasis Introducers-reorder number 406709, lot number 4219 and reorder number 406702, lot number 6372. They discovered that the product lots could contain some items which have been mis-labeled with the incorrect French size. In one case, the pouch label does not correctly identify the content of the pouch. In another | Error in labeling | Field Representatives were sent a "St Jude Medical Urgent Product Information" letter to be hand delivered to their customers. The letter listed the products and problem and advised consignees to discontinue use of the product. A form was attached for the return of products to the firm.Contact your local St Jude Medical representative for further information. |  | 2000 | 2,000 | Worldwide Distribution -- US (CA, MI, PA and TN), China and Belgium. | Terminated | December 17, 2011 | 852 | 2490
Z-2110-2009 | 53087 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Biphasic LIFEPAK 15 Monitor/Defibrillator. |  |  | August 12, 2009 | 2009 | Physio Control, Inc. | Some LIFEPAK 15 Monitor/Defibrillators have a Non-Invasive Blood Pressure (NIBP) module that was supplied by CAS Medical Systems, Branford, Connecticut who is recalling the modules. The NIBP modules may be susceptible to sudden changes in cuff pressure. Sudden changes in cuff pressure may possibly be caused by bumping or squeezing thecuff. In some instances this may result in a screen display | Software design | The firm notified customers by letter dated August 2009. Customers will have the Noninvasive Blood Pressure (NIBP) module replaced. If additional questions about the notification, contact Technical Support by calling 1-800-442-1142, option 5, Monday-Friday 6:00 a.m. to 4:00 p.m. (Pacific Standard Time).. |  | 445 | 445 devices worldwide | Worldwide Distribution -- US, AUSTRALIA, GERMANY and NETHERLANDS. | Terminated | August 10, 2012 | 1094 | 445
Z-2070-2009 | 53075 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | PortexÂż Tracheal Tube, Oral/Nasal, Murphy Eye, Clear, 3.5 mm |  |  | August 25, 2009 | 2009 | Smiths Medical | Pediatric Trachel Tube manufactured with internal diameter smaller than indicated in the label may result in inability to remove secretions and cause partial or complete obstruction of the airway | Process control | Smiths Medical issued an "Urgent Field Safety Notice" dated August 24, 2009 via UPS Second Day Mail including a Guidance Chart. Return of the enclosed Confirmation Form was also requested by fax to 1-603-355-8157. Smiths Medical issued a revised recall letter dated September 9, 2009. Customers are requested to return unused affected product to the firm.For further information, contact Smiths Medical ASD, Inc. by phone at 1-800-348-6064, prompt 4 extension 2273. |  | 24582 | 24,582 (US) 15,620 (International) | Worldwide Distribution -- United States, Andorra, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Barbados, Belgium, Bermuda, Borneo, Brazil, Brunei, Bulgaria, Canada, Chile, China, CÂżÂżte d'Ivoire, Czech Republic, Denmark, East Timor, Egypt, Estonia, Finland, France, Georgia, Germany, Gibraltar, Greece, Guatemala, Vatican City, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Liechtenstein, Lithuania, Luxembourg, Malaysia, Maldives, Malta, Mexico, New Zealand, Netherlands, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Senegal, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom and Uruguay. | Terminated | November 16, 2012 | 1179 | 655164
Z-2012-2009 | 52828 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Acuvue Advance for Astigmatism Diagnostic Product. |  |  | April 20, 2009 | 2009 | Vistakon | Vistakon is recalling Acuvue Advance for Astigmatism Diagnostic product for mislabeling. | Other | Firm notified consignees by phone from April 7, 2009 to April 27, 2009. Firm requested product disposition. Direct questions about the recall to Vistakon by calling 1-904-443-1763. |  | 73 | 73 | Nationwide. | Terminated | October 01, 2009 | 164 | 73
Z-2251-2009 | 52988 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrainLab Radiotherapy Treatment Planning Software; |  |  | August 12, 2009 | 2009 | Brainlab AG | Failure to conduct important safety checks when using the BrainLab radiotherapy treatment planning software in combination with BrainLab conical collimators could result in unintended radiation outside the conical shaped field, which may lead to serious injury of the patient. | Other | BrainLAB AG issued an "Urgent Field Safety Notice" letter dated August 12, 2009 informing all customers of the BrainLab radiotherapy planning software using conical collimators of the important safety checks to avoid the failure to apply and verify an adequate jaw-shaped field size and field position at the LINAC for the used conical collimator(s) or the failure to mount and verify the correct, intended conical collimator. Consignees were instructed on proper steps to take to avoid further issues. For further questions, contact BrainLAB AG by phone at 1-800-597-5911.*** On February 1, 2010, BrainLab sent supplemental letters to 50 of the accounts who received the August 12, 2009 letter and have Novalis Dose Delivery Systems (excluding all Novalis Tx systems). The accounts were informed that in addition to the safety checks listed in the August letter, BrainLab will be installing a mandatory software modification on their systems that limits the field size to a maximum of 4x4 cm2 when the conical collimators are inserted. The accounts would be contacted by a service representative within the next six weeks to schedule the update. Any questions were directed to BrainLab at 1-800-597-5911. *** |  | 230 | 230 units | Nationwide. | Terminated | March 16, 2011 | 581 | 230
Z-1935-2009 | 52781 | [6 empty] | Chiba Needle Marked Model: CNM-2308, Size: 23 GA x 20 CM, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | June 19, 2009 | 2009 | Remington Medical Inc. | Two unsealed pouches were identified after product release, compromising product sterility. | Process control | Initial notification to consignees was issued via telephone on June 19, 2009 and by certified mail on June 30, 2009. They were instructed to examine their inventory and quarantine product subject to the recall. If the product was further distributed, they were advised to identify their customers and notify them at once of the recall via the recall letter. A Return Authorization number was provided. Forty of 50 units in distribution were returned from the two customers. |  | 50 | 50 units | US - CA and WI. | Terminated | October 23, 2009 | 126 | 50
Z-2112-2009 | 50687 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump, REF INFP01, Zevex Enteral Nutrition Delivery Systems, Salt Lake City, UT 84123. |  |  | December 18, 2008 | 2008 | Zevex International, Inc. | Enteral feeding pumps may not accurately deliver fluid, potentially under-infusing or over-infusing. | Software design (manufacturing process) | All of Zevex' customers and distributors were notified by an Urgent Medical Device Correction letter, dated 12/18/2008. The letters explained the issue, the patient impact, and what necessary actions should be taken. Customers were instructed to contact Zevex to schedule the upgrade of their installed base of pumps at 800-970-2337 or sales@zevex.com. Distributors were to contact their customers and notify them of the letter. |  | 32545 | 32,545 | Nationwide, including 1 VA facility in CA. Foreign distribution to 1 consignee in Canada. | Terminated | October 27, 2011 | 1043 | 32863
Z-2078-2009 | 52211 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DYONICS 25 Fluid Management System Disposable Inflow/Outflow Tubeset With Forked Suction |  |  | May 27, 2009 | 2009 | Smith &amp; Nephew, Inc. Endoscopy Division | The sterility of the device cannot be assured | Package design/selection | Smith &amp; Nephew notified direct accounts by letter dated May 27, 2009 . International accounts notified by Smith &amp; Nephew Inc. Endoscopy Division to each affected Outside United States (OUS) country or distributor by electronic communication, with telephone follow-up as required. The distributor is expected to notify each of their customers that has been confirmed to have product included in this recall. |  | 67697 | 67,697 units | Worldwide Distribution -- USA, Korea, Australia, Hong Kong, India, Mexico, New Zealand, United Kingdom, Spain, Singapore, Austria, Poland, Russia, Finland, Denmark, Japan, Switzerland, China, Italy, Puerto Rico, France, Asia Pacific, Belgium, Canada, Norway, Thailand, The Netherlands, Malaysia, Brazil, Argentina, Chile, Colombia, Costa Rica, Venezuela, Turkey, Israel, Germany, UAE, Greece, and South Africa. | Terminated | May 08, 2012 | 1077 | 67697
Z-2116-2009 | 51725 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Uniflex Ti Buttress Thread Screw; |  |  | March 27, 2009 | 2009 | EBI, L.P. | Certain lots of raw material contain inclusions that could potentially reduce the strength of some devices. | Nonconforming Material/Component | Letters to Biomet/EBI reps were sent on March 29, 2009 by US mail to alert them of the product recall and the affected parts that may be in their possession. They were also asked to notify each customer and locate all affected inventory. All letters advised of the recall and the specific lot numbers involved. A letter was sent to the hospital and physician that received affected parts and asked that implanting physicians make any medical decisions. |  | 5 | 5 | Worldwide Distribution -- USA and Spain. | Terminated | September 23, 2009 | 180 | 56
Z-2085-2009 | 52683 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Single Shot Epidural Tray, Disposable, Convenience Kit, Reorder Number CUS511-08, Integra LifeSciences Corporation, Salt Lake City, Utah 84104. |  |  | July 13, 2009 | 2009 | Integra Life Sciences Corporation | Lack of packaging seal integrity on various pain management kits may compromise sterility of contents. | Packaging process control | Consignees were notified by telephone on July 09, 2009 and by "Urgent Medical Device Recall" letter on July 13, 2009. Customers were instructed to immediately examine inventory, cease use and distribution of any affected product and to contact Integra Pain Management Customer Service by calling 1-800-241-2210 to arrange for return and replacement. Refer questions to Integra Pain Management by calling 1-800-241-2210. |  | 350 | 350 units | Worldwide Distribution -- US (states of AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IN, KY, ME, MO, NC, NJ, NV, NY, OH, OK, PA, RI, SD, ST, TX and UT), Belgium and Japan. | Terminated | October 23, 2009 | 102 | 3594
Z-2232-2009 | 52959 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Medical Capnostream20 capnograph/pulse oximeter monitor and monitor with printer. Part numbers CS 08651, CS 08653, CS08657, and CS08659. |  |  | June 30, 2009 | 2009 | Oridion Medical 1987 Ltd. | 1) The monitor's ability to withstand electrostatic discharge (ESD) to exposed metal connectors on the back of the monitor which may cause the monitor's display to freeze.2) The unit generates electromagnetic radiation above the limits set by applicable standards for medical devices | Device Design | Oridian sent a letter to all its distributors June 25, 2009, instructing them to gather all the units to be exchanged for factory reconditioned units that have the hardware and software enhancements and meet all performance requirements in full. |  | 326 | 326 units | Nationwide. | Terminated | October 01, 2010 | 458 | 326
Z-2252-2009 | 53207 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Flexiflo Quantum Enteral Pump Model # 50596, 50598 and 59322. |  |  | September 04, 2009 | 2009 | Abbott Nutrition, Abbott Laboratories | The firm initiated this voluntary recall because some of its Flexiflo Quantum Enteral Pumps, have a 110-volt AC power cord plug that may crack or fail. This potential power cord plug failure may cause electrical shock and/or a delay in feeding that could represent a potential health and safety risk. | Device Design | An "Urgent Device Recall" letter dated September 4, 2009 was issued via FedEx overnight delivery to all direct consignees. The letter described the affected product, issue and immediate actions for customers. Direct consignees were further instructed to immediately contact affected customers and notify them about the Urgent Device Recall. Abbott Nutrition (AN) Device Call Center (DCC) will telephone those customers who do not responded promptly to the "Urgent Device Recall" letter with the completed AN Account Summary of the Flexiflo Quantum Enteral Pump Shipments. AN Device Call Center will make arrangements for the service or replacement of the recalled pumps. Please contact the AN Device Call Center at 1-877-457-0249, Monday through Friday between the hours of 8:30 am and 5:00 pm Eastern Standard Time (EST). An AN representative will work with you to arrange service or replacement of the recalled enteral pumps. |  | 5173 | 5,173 pumps | Nationwide Distribution (states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY). | Terminated | September 19, 2012 | 1111 | 5173
Z-2243-2009 | 52729 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, Version 3.X |  |  | August 10, 2009 | 2009 | TomoTherapy Incorporated | TomoTherapy received one customer report indicating that after starting a patient procedure from the StatusConsole, the radiation on light and the audible indicator on the Status Console came on as expected,however the Operator Station (OS) did not display the dose rate and the procedure time, or the couch position updates. Within a few minutes, the customer pressed the Stop button on the Sta | Software design | Consignees were sent a TomoTherapy "Urgent Medical Device Correction Field Safety Notice" letter dated August 10, 2009. The letter addressed the Issue, Product Affected, Investigation, Recommended Actions and Resolution.Customers were recommended to perform the following actions (in the event that the dose rate, procedure timer, and couch position updates are not present on the OS, after pressing Start on the status console and the radiation on light and audible indicators are active):1) Immediately press the Stop button on the Status Console and 2) contact your authorized TomoTherapy Service Representative for further direction. The affected issue will be corrected in the next software version, which is scheduled for release in early 2010. The TomoTherapy Customer Contact Center can be reached 24 hours a day, 7 days a week, by e-mail atsupport@tomotherapy.com or by phone at 1-866-368-4807 (North America). |  | 259 | 259 (169 US; 91 OUS) | Worldwide Distribution -- US (states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN,MO, NE,NV,NJ, NM, NY, OH, OK OR, PA, PR, SC, SC, TN, TX, UT, VA, WA, WI and WY) andBelgiumCanadaSwitzerlandChinaGermany Spain FranceEnglandHong KongIndiaItaly Japan KoreaNetherlandsPolandSaudi ArabiaSwedenSingaporeTurkey andTaiwan. | Terminated | December 17, 2011 | 859 | 259
Z-2312-2009 | 51730 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Multi-Axial Screw (TI); 8.00MM (D) x 70MM(L) W/Collar; |  |  | April 02, 2009 | 2009 | EBI, L.P. | Some devices could have been made with raw material lots that may have inclusions, and could result in decreased strength of the device. | Nonconforming Material/Component | The firm sent Urgent Recall Notices to all direct accounts, hospital Risk Management and Physicians by first class mail.Letters are dated March 29, 2009, April 2, 2009 and April 16, 2009 respectively. The letter describes the problem and that that there is a possible risk of a device failure.An additional e-mail was sent on May 22, 2009 to 5 customers due to an expansion of the recall to include 4 additional items.On June 19, 2009, the firm again expanded the recall to include 2 additional products and notified customers.Questions should be directed to William Hennig, at 1.973.299.9300. ext. 1557. |  | 1 | 1 | Worldwide - US -NY, NJ, MI, MD, FL, GA, PR, MN, SD, OK, TX,, LA, CA, PA, OH, NV, HI, ME, NC, KS, UT, CO, TN, WV, VA, and WA. Foreign: Brazil, Colombia, Denmark, Germany, Finland, Greece, Italy, Japan, The Netherlands, Switzerland, Sweden, Portugal, Spain, United Kingdom, Saudi Arabia, and Trinidad. | Terminated | September 23, 2009 | 174 | 708
Z-2209-2009 | 52780 | No | MNB |  |  | Obstetrics/Gynecology | PMA | Boston Scientific HydroThermAblator Endometrial Ablation System- |  |  | July 31, 2009 | 2009 | Boston Scientific Corporation | Cracked procedure sheaths, incorrect care/use of device, and console malfunctions may result in fluid leaks and loss of a cervical seal that are responsible for burns. | Process design | Boston Scientific issued an Urgent Medical Device Recall - Immediate Action Required notification to hospitals and distributors dated July 31, 2009 via Federal Express and requested the immediate discontinued use and segregation of Recall product for return and Complete and return the Reply Verification Tracking Form. An email was sent to Physicians to provide the recall information. Boston Scientific issued a Clarification letter dated August 6, 2009 via Federal Express as a follow-up to the July 31, 2009 Recall letter for the following reason: customers may remove product from the outer box and store it on the inventory shelf by the inner package only. If this is a practice at the facility, a customer must consider both the inner and outer packaging product codes when searching for affected recalled product as the UPN numbers on the inner and outer labeling are different. A reply verification tracking form for customers was provided in case they have located any additional product to return. |  | 22744 | 22,744 eaches | Worldwide Distribution -- USA, Austria, Denmark, Finland, France, Great Britain, Italy, Netherlands, and Sweden. | Terminated | May 03, 2011 | 641 | 24299
Z-2227-2009 | 52631 | helmet, cranial, for protective use | NDA | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | 3i Custom # 5753, OSCM2913, 2.9MMX 3.4MM X 13MM, OSSEOTITE CUSTOM, MICROMINI IMPLANT, LOT # 470258, Sterile. Implant Innovations, Inc. 4555 Riverside Drive, Palm Beach Gardens, FL 33410. |  |  | December 11, 2006 | 2006 | Biomet 3i, LLC | Healing Abutment was not made within required manufacturing specification. The product did not have the relief behind the thread, which does not allow proper seating. | Nonconforming Material/Component | Each of the US customers was contacted by phone. Following the phone contacts, Biomet 3i performed a follow- up of each call with a faxed letter. Each customer was instructed to check their respective inventory for the affected product and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i. Contact Biomet 3i, LLC at 1.561.776.6706, if there are any questions. |  | 24 | 24 | USA and Italy | Terminated | October 01, 2009 | 1025 | 24
Z-0003-2009 | 53158 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d systems with control room PC model 5117866-2 with Software collector 5194381. GE Healthcare, 3000 North Grandview, Waukesha, WI 53188. |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an error in the generator software on the Revolution XR/dconfigured with HP 8200 PC consoles that reports a mAs reading that is higher than is actually exposed to patients. | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology, and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, and Contact Information. |  | 83 | 83 (63 USA, 20 OUS) | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, ID, IL, IN, LA, MA, MI, MO, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA, and WI and countries of Canada, China, Japan, Korea, and Taiwan. | Terminated | December 18, 2011 | 1403 | 83
Z-2256-2009 | 52711 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Connecting Strap/Belt, Cotton, 36" (232036). |  |  | April 20, 2009 | 2009 | J T Posey Company | Posey has initiated a recall on certain Posey product units which were manufactured using an incorrect key lock buckle. The intended use of these products is to limit a patient's movement to the extent necessary for treatment, examination or protection of the patient or others. | Employee error | JT Posey Company issued an "Urgent Medical Device Recall (Removal)" letter dated May 14, 2009 informing customers of the affected products. Any used and/or existing inventory subject to this recall is requested to be returned to Posey including the Recall Reply Fax-Back Form.For further questions, contact your local sales representative or JT Posey Company at 1-800-447-6739 extension 111. |  | 47 | 47 | Class II Recall - Worldwide Distribution -- United States, Canada, France, Belgium and the Netherlands. | Terminated | April 11, 2012 | 1087 | 3136
Z-2230-2009 | 52824 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Move Scoop Stretcher; Note: The stretcher is an accessory to the Maxi Move Patient Lift; Made by Arjo Hospital Equipment AB, Eslov, Sweden; product number KRA2210-20. |  |  | August 18, 2009 | 2009 | ArjoHuntleigh | A Scoop Stretcher is made up of two frames that are held together with two locking devices, one on each end. Each locking device contains a safety latch. The safety latch could unintentionally release when the side frames are pulled for positioning. | Manufacturing material removal | ArjoHuntleigh sent their customers a cover letter and the Field Safety Notification, informing them of the potential problem with the safety latches on the stretcher and a pending visit from an ArjoHuntleigh Service Technician to correct the problem. These letters were sent via UPS 2nd day delivery on 8/18/09. The customers were requested to test their stretchers upon receipt of the letter following the instructions for checking the safety latches on the stretcher:* Place the stretcher on its side on a non-scratch surface.* Try to pull the two sides apart with the safety latch engaged.* Test both the head and foot ends of the stretcher.* If the stretcher pulls apart, remove it from use immediately.It is important to note that all affected stretchers will be upgraded, whether they fail the test or not. |  | 214 | 214 units | Worldwide Distribution | Terminated | January 07, 2010 | 142 | 214
Z-2239-2009 | 53132 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ILUMAVision Software, version 2.1.0 Picture Archiving and Communication system |  |  | August 03, 2009 | 2009 | IMTEC IMAGING L.L.C. | Under certain circumstances, software Version 2.1.0 has potential for incorrect data imaging and patient identification. | Software design | Firm sent notifications to all ILUMAVision customers on 8/3/2009 and offered upgraded software. The notification letter indicated the the reason for the recall and identified the potential issues associated with the device. Contact 3M IMTEC Technical Support at 1.800.879.9799, if there are any questions. |  | 223 | 223 | Worldwide - US - Foreign: Australia, Austria, Belgium, Bulgaria, Canada, Croatia, Egypt, France, Germany, Hong Kong, Italy, Korea, Malaysia, Mexico, The Netherlands, Poland, Saudi Arabia, Spain, Turkey, and United Kingdom | Terminated | January 27, 2010 | 177 | 223
Z-2231-2009 | 53045 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Proton Convolution Superposition DCF, version 8.6.15, Model H48, radiation therapy treatment planning system. |  |  | July 17, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software Issue/Mistreatment Potential-- Due to software issues, dose distribution may be incorrect and unintended radiation doses may be delivered. | Software design | Varian Medical Systems Recall initiated on July 17, 2009. A product notification letter was distributed to the affected user, with a description of the problem and user corrective action steps. The letter will be distributed to the sales, marketing and service organizations via certified mail. A software correction is planned. |  | 1 | 1 | Product was distributed to one user in Texas | Terminated | November 27, 2010 | 498 | 1
Z-2236-2009 | 53041 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | Action Products, Inc., Bol-X-II, 30x30x0.5cm, item number 32052, packaged 10/carton. Polymer pad used during radiation therapy. |  |  | June 26, 2009 | 2009 | Action Products Incorporated | Medical device intended for use by healthcare practitioners was incorrectly labeled. 1.0 cm thick pad was labeled as 0.5 cm thickness, and vice versa. | Process control | Action Products Inc. notified one affected distributor by phone and email on 06/26/09 and issued a Return Merchandise Authorization. |  | 1 | 1 carton | Nationwide distribution. | Terminated | September 29, 2009 | 95 | 1
Z-2235-2009 | 50473 | Yes | NVZ |  |  | Cardiovascular | PMA | Sorin Group Reply DR Rate responsive dual-chamber pacemaker. Made in Italy by: Sorin BioMedica CRM S.r.l., Via Crescentino s.n., 13040 Saluggia (VC) Italy. ELA Medical, Inc., 14401 West 65th Way, Arvada, CO 80004, USA. REPLY DR is a DDDR-type dual-chamber rate-responsive pacemaker. It is equipped with a physiological sensor and an accelerometer to allow adaptation of pacing to suit the patient's activity. The REPLY device can be programmed to separate pacing and sensing polarities, independent of chamber. |  |  | September 23, 2008 | 2008 | ela Medical Llc | Analysis showed that the screwdriver was not fully inserted in the setscrew's hex cavity, which damaged the setscrew. | Employee error | A Reply Technical Note, dated September 2008, titled Additional Instructions for use at Implant for REPLY pacemakers, was issued to physicians on September 24, 2008. The Technical Note describes the background for an increased incidence of complaints relating to the screwdriver and setscrew. The letter provided a reminder for inserting the screwdriver, and additional instructions for insertion of the screwdriver. |  | 233 | 233 | Nationwide Distribution -- USA, including states of AL, AZ, CA, CO, CT, FL, GA, KS, KY, LA, MA, MI, NC, NJ, OH, OK, SC, TN, and TX. | Terminated | December 17, 2011 | 1180 | 233
Z-0345-2009 | 50742 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Integris H5000F |  |  | December 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Tube failure occurs involving fluoroscopy of large patients and a partial failure of the grid switch. | Radiation Control for Health and Safety Act | An Electronic Product Radiation Warning, dated October 2008, was sent to customers. The letter identified the problem and under what circumstances that it could occur; the actions that should be taken by the customer/user to prevent risks for patients or users; and the actions planned by Philips to correct the problem. Further information or support concerning the issue could be obtained by contacting Bob Pettis at 425-487-7469 or the Philips Call Center at 800-722-9377.Firm FCO tracking numbers 72200107, 72200108, 72200111, 72200112, 72200117, 72200118, 72200119, 72200122, 72200123, 72200124, 72200125, 72200129, 72200142, 7989980, 7080085, 70600034, 70600035, 72200132, 72200133, 72200137, 72200138, 72200140 and 72200141. C &amp; R # 1217116-10/20/08-016C |  | 929 | 929 units | Worldwide Distribution | Terminated | August 17, 2012 | 1346 | 3361
Z-1155-2009 | 51856 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Fluorostar 7900 Mobile Fluoroscopy System |  |  | February 19, 2009 | 2009 | Ge Healthcare | The alarm timer is reset after each exposure and therefore never reaches the five minute alarm state unless one exposure exceeds 5 minutes in duration. | Radiation Control for Health and Safety Act | GE Healthcare sent a customer notification out to all sites that have potentially affected systems. GEHC Surgery will also install a software correction into the affected systems in order to resolve this issue at no charge to the customer. Until the GEHC OEC Fluorostar 7900 has received the corrective software installed, they are encouraging facilities to exercise caution in monitoring the amount of cumulative fluoroscopy time when the affected system is in use. |  | 554 | 554 units | Worldwide Distribution | Terminated | July 22, 2011 | 883 | 554
Z-2241-2009 | 52703 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | InnovaÂż 3100/3100IQ GE Innova 3100 / 3100 IQ, Cardiovascular Imaging System, GE Healthcare, Waukesha, WI 53188 (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). |  |  | July 31, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently became aware of a potential power supply failure with the InnovaÂż 2000, 3100 and 4100 systems that may impact patient safety. A potential power supply failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray generation and image processing. Should this occur, the system cannot be restarted. | Component design/selection | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated June 25, 2009. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology or Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.GE Healthcare will correct all affected systems by replacing the power supply located in the GE equipment cabinet. A GE Healthcare service representative will contact you to arrange for this correction. Direct questions about the recall to the firm's Call Center by calling: 1-800-437 -1171, option 4 (United States) or 0120 - 055 - 919 (Japan). For other countries, please contact your local GE Healthcare field service Representative if you have any questions or concerns regarding the notification. |  | 6 | 6 | Worldwide Distribution -- US (states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA., MI, MN, MO, NE, NV, NH, NJ, NY, NC, OH, OK , OR, PA, PR, RI, SC, TN, TX, UT VA, WA, WI and WV),TUNISIA, THAILAND , TAIWAN, SYRIAN ARAB REPUBLIC, SWEDEN SPAIN, SOUTH AFRICA SAUDI ARABIA, RUSSIAN FEDERATION , POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MARTINIQUE, MACAU, LEBANON, KUWAIT, REPUBLIC OF KOREA, JAPAN, ITALY, ISRAEL, IRAN, INDONESIA, INDIA, GREECE, GERMANY, FRANCE , CHINA, CANADA, BULGARIA, BRAZIL BELGIUM, BELARUS and AUSTRALIA. | Terminated | May 15, 2012 | 1019 | 508
Z-1888-2009 | 52723 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV Systems (Part Number 8139789), equipped with a syngo RT Therapist (Part Number 8162815) or RT Therapist Assist system (Part Number 8162807), with Software version NET VA10A for RTT4.1.122 or RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | April 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Unexpected treatment plan alteration-- Software issue may result in unintended modifications to treatment | Software design | Siemens Medical Solutions USA, Inc. notified consignees of the affected product by advisory letter via Certified Mail on April 09, 2009.For further information, contact Siemens Medical Solutions at 1-925-602-8157. |  | 53 | 53 units | Worldwide Distribution -- US (MA, UT, OH, NJ, NE, FL, OH, MI and PA), Germany, Poland, the UK, Ireland, Italy, France, South Korea, and Belgium. | Terminated | April 12, 2011 | 733 | 53
Z-2250-2009 | 53080 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | B. Braun Blistered Omnifix L.L. Syringe packaged in the Easypump LT 60-24, 60 ml x 2 ml/hr |  |  | August 01, 2008 | 2008 | I-Flow Corporation | I-Flow identified a B. Braun syringe package that had a tear/hole in the paper portion of the primary package.Used to provide an optional filling device for the Easypump. | Packaging process control | I-Flow sent a recall letter dated August 1, 2008 to consignees telling them of the above mentioned recall and asking them to Inspect their Easypump inventory to determine if they have any of the affected lots identified on this Recall Notification Letter. They were told not to distribute affected devices and quarantined remaining inventory immediately. They were asked to check the appropriate response on the attached I-Flow Fax sheet and return the form back to I-Flow within 48 hours of receipt.The I-Flow Customer Service number (949) 206-2700, and hours of operation (7:00 am - 5:00 pm PST) were provided to answer questions. |  | 1608 | 1,608 units | France | Terminated | October 07, 2009 | 432 | 2177
Z-2314-2009 | 52616 | No | LZY | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tandem Unipolar, 46 MM CO-CR, REF 126646, Sterile, Smith &amp; Nephew, Memphis TX 38116 |  |  | January 12, 2009 | 2009 | Smith &amp; Nephew Inc | 46 mm Femoral head was mismarked and packaged as being 45 mm | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and sent recall notice in overnight mail on 01/12/2009. All affected Smith &amp; Nephew International Distributors were notified via e-mail and telephone on 01/12/2009. They were instructed to immediately quarantine the product for return. The firm re-issued the letter on 04/06/2009 to address hazard to health.Contact Smith &amp; Nephew, Inc at 1.901.399.6771 if there are any questions. |  | 10 | 10 units | CA, FL, MI, and Canada | Terminated | March 25, 2011 | 802 | 10
Z-1036-2009 | 53176 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 4100/4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally, rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. |  |  | November 04, 2008 | 2008 | GE Medical Systems, LLC | One of three copper spectral filters could remain within the x-ray while the system displays no spectral filter. The displayed Air Kerma rate and thus cumulative Air Kerma are overestimated. The Air Kerma rate may be over estimated by up to 60%. | Radiation Control for Health and Safety Act | Affected customers were notified by means of a Important Electric Product Radiation Warning letter. GE field personnel will be replacing the collimators. |  | 166 | 166 | Nationwide Distribution -- AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and Puerto Rico. | Terminated | May 16, 2012 | 1289 | 166
Z-2244-2009 | 53143 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Renegade" Hi-Flo Microcatheter, Peripheral Vascular Use. Sterilized with Ethylene Oxide. Legal Manufacturer: Boston Scientific Cork Ltd., Business &amp; Technology Park, Model Farm Road, Cork, Ireland. |  |  | August 26, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific has determined that the sterile barrier in the packaging of the affected product may be compromised. It has been reported that the pouch seal (top seal) may be breached. | Packaging | Consignees were sent a Boston Scientific Urgent Medical Device Recall Removal - Immediate Action Required letter Dated August 26,2009. The letter was addressed to Risk Manager / Field Action Contact. The letter described the problem and product involved. Advised consignees to segregate the affected product and return it to Boston Scientific. Requested the return and completion of The Account Reply Verification Tracking form enclosed. Replacements will be issued for all recalled product that is properly returned to Boston Scientific. |  | 80885 | 80,885 (51,491 US; 29,394 OUS) | Worldwide Distribution -- USA, excluding New Mexico and Wyoming, and countries of Europe, the Middle East and Africa, Japan, American Virgin Islands, Argentina, Aruba, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dutch Antilles, Egypt, Finland, France, Georgia, Germany, Great Britain, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Libya, Luxembourg, Macau, Malaysia, Mexico, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Poland, Portugal, Puerto Rico, Romania, Saudia Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, and Vietnam. | Terminated | August 24, 2012 | 1094 | 118920
Z-2316-2009 | 52858 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive (EA) versions 3.0 to 3.0.7 software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | August 14, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software error: There is potential safety issue with Centricity Enterprise Archive (EA) 3.0.x software where study split operations are not correctly replicated to a secondary 'shadow' archive. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated August 7, 2009 to the attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the potential safety issue and provided safety instructions to follow until the software is updated. Users were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479. Any questions should be directed to the Remote Online Center or the GE Customer Care Center at 1-800-437-1171. |  | 212 | 212 units | Worldwide Distribution -- United States, Belgium, Canada, Cayman Islands, China, Denmark, France, Germany, Great Britain, Hong Kong, India, Israel, Italy, Korea, Malaysia, Malta, Mexico, New Zealand, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan and Turkey. | Terminated | February 25, 2011 | 560 | 212
Z-1690-2009 | 53260 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 4100, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). |  |  | October 09, 2008 | 2008 | Ge Healthcare | GE Healthcare has become aware of an Air Kerma Rate issue associated with a large Equivalent Patient Thickness (EPT) on the Innova system that my impact patient safety. During a Maximum Air Kerma Rate (AKR) calibration procedure check, GE Healthcare found the AKR calibration procedure does not exactly reflect the worst-case clinical usage of the product. This may result, when using large Equival | Radiation Control for Health and Safety Act | Consignees were sent a GE Healthcare "Important electronic Product Radiation Warning" letter dated November 3, 2008. The letter was addressed to Healthcare Administrator/Risk Managers, Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Defect Description and Related Hazards, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For further information, contact the GE Healthcare Technical Support Center at 1-800-437-1171. |  | 929 | 929 units total | Nationwide Distribution -- United States including Puerto Rico and Guam. | Terminated | January 07, 2012 | 1185 | 929
Z-0002-2009 | 53157 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d system control room PC model 5117866-2 with Software collector 5115434 |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential risk of observing a blank image after exposure onRevolution XR/d systems equipped with HP 8200 PC user consoles that may impact patient safety.It was reported that images acquired with off-center collimation for extremities with theposition selected as Digits to front or Digits to back are displayed as blank. Investigation intoth | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information. |  | 129 | 129 | Nationwide Distribution -- AL, AZ, CA, CO, FL, GA, HI, IL , IN, KY, LA, MD, MA, MI,MN,MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN, TX, VA, WA, and WV. | Terminated | January 07, 2012 | 1423 | 129
Z-2313-2009 | 52989 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron Delivery Catheter 053, 115 cm/12 cm-- MP, manufactured by Penumbra Inc., Alameda, CA |  |  | July 01, 2009 | 2009 | Penumbra Inc. | Misbranding--Some products may be identified with the wrong ID band content 'Neuron 6F 070' instead of 'Neuron 6F delivery' | Mixed-up of materials/components | Firm initiated its recall on June 30, 2009. Consignees notified on July 6 and 7, with follow up contact on July 7, 2009. Contact Penumbra Inc. at 1.510.618.3223, if there are any questions. |  | 37 | 37 units | CA, FL, NC, MI, MO, VA, AZ, CT, IA, NY, and CO, Foreign: Germany and Australia. | Terminated | November 30, 2010 | 517 | 37
Z-2234-2009 | 52794 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | S&W; Electrodes for Defibrillation, 10 Defibrillation Electrodes, Reorder #90101, S &amp; W Healthcare Corp, 1525 Flight Path Drive Brooksville, Florida 34604 |  |  | July 10, 2009 | 2009 | Leonhard Lang Gmbh | A potential risk was identified associated with the connectors of the defibrillation electrodes DF20. It was found that some of the connectors do not slip easily into the sockets of the defibrillator or an extension cord of the defibrillator. In such cases increased force or some wiggling is required to insert the connector fully. A potential risk for patients exist, if (a) such a connector (b) | Nonconforming Material/Component | A Leonhard Lang letter dated June 9, 2009 was sent June 6, 2009 to its customer. The letter identified the problem and the product. The letter instructed the consignee to stop using or distributing the product and to quarantine them. It also requested them to report all quantities of product that they had in stock using the attached form. |  | 3044 | 3044 | Worldwide Distribution - US - FL Foreign: Austria, Germany, Romania, France, Dubai, Greece, Italy, South Korea, United Kingdom, Hungary, Latvia, Saudi Arabia and Lebanon | Terminated | December 17, 2011 | 890 | 4477
Z-2248-2009 | 52849 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100, 31Gauge, 1cc Insulin Syringe, 5/16" (8mm), Short Needle, NDC 0603-7002-21, packaged 10 syringes per package, 10 packages per carton |  |  | August 21, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles separate from the barrels of the syringes | Nonconforming Material/Component | The firm initiated the recall via a press release on 08/21/2009. They began notifying their consignees of the problem and the recall by letter on 08/21/2009. email: recallresponse@qualitestrx.com |  | 1652000 | 16,520 boxes of 100 units | Nationwide | Terminated | November 24, 2010 | 460 | 5386000
Z-1666-2009 | 53034 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Diagnost 96/97 DSI |  |  | July 06, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The x-ray system fail to meet the requirements of 21 CFR 1020.32(d)(2)(ii) when a specific tube failure occurs involving fluoroscopy of large patients and a partial failure of the grid switch | Radiation Control for Health and Safety Act | A Customer Notification letter describing the potential grid switch failure and actions planned by Philips to correct the problem will be sent to all users of the affected systems during the week of July 6, 2009. |  | 17 | 17 units distributed in US. | Worldwide Distribution | Terminated | December 16, 2011 | 893 | 401
Z-0623-2009 | 50748 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems MultiDiagnost Eleva Flat Detector System with software version 4.3.1 or higher |  |  | December 30, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | System fails to meet requirements when it is used at kV levels greater than 115 kV at certain frame rates and in certain acquisition modes, as grid switch leakage may occur. | Radiation Control for Health and Safety Act | A Customer Notification letter describing the failure, the means, by which a customer can prevent the failure mode from occurring, and actions planned by Philips to correct the problem were sent to all users of the affected systems. Philips will modify the system software and will be installed at no cost to the customer. |  | 30 | 30 systems | Nationwide distribution including states CA, CO, IA, ID, KS, MN, NC, NJ, NY, OH, SC, TX, UT, VA, and WI. | Terminated | September 28, 2010 | 637 | 30
Z-1223-2009 | 51867 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MilCAM Recon III, Family of IR Imagers that incorporates a Class IIIb infared laser pointer |  |  | March 10, 2009 | 2009 | Flir Systems, Inc. | MilCAM Recon III models were sold to non-DoD customers prior to certification and reporting as required in Title 21 of the Code of Federal Regulations. Conditions of sale were not stated in variance approval letters issued for similar laser products. | Radiation Control for Health and Safety Act | FLIR Systems notified direct account by letter of the Corrective Action Plan and updated certification label that references CDRH issued variance and instructions on how to install it. |  | 58 | 58 | Nationwide Distribution -- Washington, D.C. | Terminated | May 28, 2015 | 2270 | 58
Z-1118-2009 | 51490 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Definium 5000 X-ray System, model number 5220493 and 5220393-2 |  |  | March 24, 2009 | 2009 | GE Healthcare, LLC | A defect in the AEC board in the system's control causes failure to terminate additional exposures at the exposure time selected for the initial exposure. | Radiation Control for Health and Safety Act | A notification was sent to all affected customer sites alerting them to this issue, with instructions for safe use of the affected systems. |  | 85 | 85 | Nationwide Distribution | Terminated | May 16, 2012 | 1149 | 85
Z-0937-2009 | 51488 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MilCAM Recon III Model Family of IR Imagers that incorporate a Class I infared laser for Laser Range Finding (LRF). |  |  | January 30, 2009 | 2009 | Flir Systems, Inc. | LRF models were sold to non-DoD customers prior to certification and reporting as required in Title 21 of the Code of Federal Regulations, Subchapter J, Parts 1000 - 1050 (21 CFR 1000 - 1050). A filter defect allows transmission of laser energy such that the affected LRF would be Class 3B laser product, however the product is reported to CDRH as a Class I product. | Radiation Control for Health and Safety Act | Customers will be notified by a letter instructing them to stop using the LRF, contact a Customer Service Representative, and to detach and return the Laser Range Finder. For each returned LRF, FLIR will provide a certified Class I LRF. |  | 120 | 120 | Nationwide Distribution | Terminated | June 02, 2015 | 2314 | 120
Z-1259-2009 | 51879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare CT Systems Table Model Number 2269165 used on LightSpeed 1X through 4X and HiSpeed Qxi CT Systems. |  |  | March 12, 2009 | 2009 | GE Medical Systems, LLC | Failure to provide the certification label required by 21 CFR 1010.2. | Radiation Control for Health and Safety Act | A GE Field Service Representative will visit each affected customer site and affix the certification label required. A Field Modification Instruction (FMI Number 25412) has been developed for release to all facilities in possession of the affected systems. |  | 1179 | 1179 total of all units | Nationwide Distribution | Terminated | December 10, 2011 | 1003 | 1179
Z-0005-2009 | 53451 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix fluoroscopy DFP-8000D Digital Radiography System with version 3.40 or higher software. Controller for Infinix interventional angiography systems. |  |  | September 29, 2008 | 2008 | Toshiba American Medical Systems Inc | Audible signal is not available at all times when using the HLC fluoroscopy mode. Special means of activation are not designed in for selection of High Level Controls, as required by regulation. | Software design | Toshiba issued a letter entitled Important Electronic Product Radiation Warning, September 30, 2008, explaining the issue. The customers were requested to reply to the letter via fax-back form, and a Toshiba service representative would install the corrective action. |  | 129 | 129 | Nationwide | Terminated | October 04, 2012 | 1466 | 129
Z-1658-2009 | 52173 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Celex/GE Precision MPi, REF no.: 02980000 |  |  | May 01, 2009 | 2009 | NRT - Nordisk Rontgen Teknik A/S | The GE Precision MPi is in violation with 21 CFR Part 1020.32 requiring a signal audible to the fluoroscopist to sound for each passage of 5 minutes of fluoroscopic irradiation time during an examination or procedure. | Radiation Control for Health and Safety Act | An Advisory Notice concerning GE Precision MPi, dated May 1, 2009, was issued to Healthcare Professionals. The notice discussed the problem, risk, recommended immediate action, future actio, and affected system and serial numbers. A GE Field Service Representative will visit each affected U.S. customer site and introduce a speaker connected to the generator touch screen (to which the audible tone is delivered) in the operator's room, to be placed on the wall in the x-ray room. For further information, customers were to contact their local GE field engineer. |  | 38 | 38 units (US) | Nationwide Distribution, including CA, CO, FL, IN, MA, MI, MN, NV, NY, OH, OR, PA, TN, VT, WA, WI. | Terminated | October 20, 2009 | 172 | 38
Z-2276-2009 | 52730 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedÂż II Programmable Pump, 8637, Rx Only, Sterilized by ethylene oxide gas, Medtronic inc., Minneapolis, MN 55432-5604. (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) |  |  | July 01, 2009 | 2009 | Medtronic Neuromodulation | This recall is related to the potential for reduced battery performance in a small percentage of Medtronic Model 8637 SynchroMedÂż II pumps with batteries manufactured during two distinct time periods prior to April 2005.Nature of the Device Issue: As part of ongoing analysis of returned explanted product, Medtronic has confirmed that reduced battery performance resulted in eight (8) occurrence | Nonconforming Material/Component | Consignees were sent on July 7, 2009 a Medtronic " Urgent: Medical Device Correction" letter dated July 2009. The letter was addressed to Healthcare Professionals. The letter described the nature of the Device Issue, Scope, Potential Severity of the Issue, Recommendations, Ongoing patient Management Recommendations and Additional Information. Important patient management information is also available at http://www.professional.medtronic.com under the heading Product Advisories.Direct questions to your Medtronic Field Representative or Medtronic Neuromodulation Technical Services by calling at 1-800-707-0933. |  | 18572 | 18,572 | US (including DC and Puerto Rico), Australia, India, canada, Hong Kong, Japan, Austria, Belgium Cyprus, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Luxembourg, Netherlands, Norway, Poland, Portugal, San Marino, Saudi Arabi, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom and Vatican City. | Terminated | May 10, 2012 | 1044 | 18572
Z-2315-2009 | 53162 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL R Series Defibrillator/Pacemaker/Monitor |  |  | August 31, 2009 | 2009 | ZOLL Medical Corporation, World Wide Headquarters | Failure of Device to display ECG could delay delivery of defibrillation or pacing therapy. | Nonconforming Material/Component | Zoll notified domestic consignees on August 31, 2009 and international consignees on September 1, 2009 by an Urgent Medical Device Correction letter. Customers were instructed to continue to perform daily tests per the recommendations provided in the R Series Operator's Guide. In the event of a failure, customers are to remove the device from service and contact biomedical or clinical engineering department per their facility's protocol. Units that operate as expected can remain in service and back-up monitor should be available in the event of a failure. Zoll's technical support department should be contacted at 1-800-348-9011 or +1-978-421-9460 to schedule the return of affected devices with the defect for replacement of the suspect component. |  | 247 | 247 units | Worldwide Distribution -- USA, UK, Austria, Australia,, Canada, Switzerland, Germany, France, French Guiana, Jordan, Macedonia, Netherlands, Panama, Poland, Romania, and Saudi Arabia. | Terminated | May 30, 2012 | 1003 | 247
Z-1991-2009 | 52861 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics System included with the Dimension Vista System. |  |  | July 14, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | incorrect results may be transmitted to LIS | Software design | Siemens issued an "Urgent Device Correction" letter dated July 2009. The letter informed the user of the affected product including instructions for avoiding the problem.For further information, contact Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 359 | 359 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, Portugal, Slovenia, Spain and Switzerland. | Terminated | January 05, 2010 | 175 | 359
Z-1814-2009 | 52680 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConvaTec, Engenex Tunnel Dressing, for use with Engenex Negative Pressure Wound Therapy, (1 x 15 cm) |  |  | July 13, 2009 | 2009 | ConvaTec | Labeling: the actual manufacturing date (2/2009) and actual expiration date (2/2011) are switched. | Labeling Change Control | Recall notification letters were hand delivered on July 13, 2009 to all direct customers (tier 1) by ConvaTec reps. Letters were provided to all direct customers so that they could notify all customers that received the product. |  | 764 | 764 | Nationwide distribution and ConvaTec sales rep in the UK and France. | Terminated | October 02, 2009 | 81 | 764
Z-2238-2009 | 48191 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | SYNCHRON Systems Salicylate (SALY) Reagent; Part Number: 378194. |  |  | October 22, 2007 | 2007 | Beckman Coulter Inc | The Salicylate Reagent lots listed here may fail calibration and/or fail to maintain linearity at the upper end of analytical range. | Nonconforming Material/Component | An "Urgent: Product Corrective Action" letter dated October 22, 2007 was issued to customers. Customers were instructed to do the following: 1) Discontinue use and discard appropriately the affected lots of Salicylate Reagent., 2) Contact your Customer Service Representatives at 1-800-526-3821 in the United States or local Beckman Coulter Representative to discuss replacement options and 3) Complete and return the enclosed response form within 10 days. Contact your Customer Service Representative at 1-800-526-3821 in the United States or 1-800-463-7828 in Canada or, if you are outside the United States or Canada, please contact your local Beckman Coulter Representative to discuss replacement options.Direct any questions regarding this Product Corrective Action letter to the Customer Support Center (Hotline) at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, please contact your local Beckman Coulter Representative. |  | 7842 | 7,842 total | Nationwide and Canada. | Terminated | July 16, 2012 | 1729 | 7842
Z-2326-2009 | 51156 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DCL (Diagnostic Chemicals Limited) Creatine Enzyme Reagent (R2) Catalog Number: 265-30. |  |  | February 12, 2009 | 2009 | Diagnostic Chemicals, Ltd. | Labeled with extended expiration date (24 months instead of 18 months) | Labeling Change Control | Genzyme Corp. notified consignees by Letter dated February 12, 2009 via Federal Express. Accounts are requested to discontinue use and discard units. Customers were also asked to complete and sign the Confirmation of Notification the fax to 1-902-628-6504.For further information, contact Genzyme Technical Support at 1-800-565-0265. |  | 82 | 82 kits | Worldwide Distribution -- United States (AL, CA, NJ, MO, PA, KY, FL, LA, MN, NY, NC, OH, TX and UT), Canada and Korea. | Terminated | May 23, 2012 | 1196 | 107
Z-2317-2009 | 51258 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AxSYM System Tubing Decontamination Solution; 500 mL bottle; Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064; List Number 07B05-09. |  |  | February 18, 2009 | 2009 | Abbott Laboratories | Bottle labels from lot 211974, AxSYM System Tubing Decontamination Solution (List Number 7B05-09), were found to be missing the following information: lot number, date of manufacture, and expiration date. | Manufacturing material removal | Abbott sent Product Recall Immediate Action Required letters dated February 18, 2009 to customers informing them of the affected product. Customers were instructed to examine their inventory and dispose all affected product. Accounts were requested to complete and fax back customer reply sheet acknowledging receipt of the letter, as well as the number of bottles destroyed, and to retain a copy of the letter for their records. For further information, contact Abbott Customer Service at 1-877-4ABBOTT. |  | 1965 | 1,965 kits | Worldwide Distribution -- United States including Puerto Rico and Argentina, Brazil, Canada, Chile, Columbia, Germany, Hong Kong, Japan, Mexico, Peru, South Korea, Thailand and Trinidad &amp; Tobago. | Terminated | July 08, 2010 | 505 | 1965
Z-2328-2009 | 51872 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Martel Printer Accessory to i-STAT Portable Clinical Analyzer, rechargeable battery pack. |  |  | April 22, 2009 | 2009 | Abbott Point of Care Inc. | It appears that the rechargeable battery pack is overheating with or without the presence of smoke. | Nonconforming Material/Component | Urgent Recall Notice letters dated April 2009 were sent out the third week of April by Federal Express. The letter discussed the background, required actions, and additional information. Customers are to identify all affected Martel printer rechargeable battery packs, record the number of affected packs on the attached business reply cards, and return the card along with the affected batteries as instructed. Questions should be directed to Abbott Technical Support at 800-366-8020. |  | 2552 | 2552 | Worldwide Distribution | Terminated | February 25, 2010 | 309 | 2552
Z-2325-2009 | 51928 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower. |  |  | April 24, 2009 | 2009 | Abbott Laboratories | There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59). | Software change control | Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 5713 | 1,037CD-ROMs - US, 4,676 CD-ROMs - WW | Worldwide Distribution -- United States including Puerto Rico and Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Curacao, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay and Venezuela. | Terminated | November 30, 2009 | 220 | 5713
Z-1871-2009 | 52089 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper 903 |  |  | May 06, 2009 | 2009 | Whatman, Inc. | The two portions of the card are coming apart prematurely. | Process control | Whatman issued an Urgent Medical Device Correction letter to to the direct account on 5/6/09 requesting this notice is distributed to users. The notice states the safety issue; requests users to inspect 903 cards prior to use, if the 903 Paper Portion of the card partially or completely detaches from the Demographic portion of the card, discard the card obtain a new card and inspect. On May 26, 2009 the firm expanded the notification to additional users of the 903 cards to inspect cards prior to use. |  | 2853410 | 2,853,410 cards | Nationwide Distribution -- AZ, FL, GA, NC, IL, IN, LA, KS, KY, MI, MO, NE, NY, OH, OK, PA, SC, OR, TX, TN,UT VA, WI, WI, &amp; PR. | Terminated | October 25, 2011 | 902 | 2853410
Z-2324-2009 | 52813 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | SensititreÂż Mueller-Hinton Broth w/Lysed Horse Blood (MHB w/LHB), Model Number: CP112-10; Lot number 150444SA. |  |  | June 04, 2009 | 2009 | Trek Diagnostic Systems | Thirty-one boxes of unapproved quarantine product, CP112-10 MHB w/LHB, Lot # 150444SA, were inadvertently shipped to customers on 6/2/09 instead of shipping approved Lot # 150344SA. | Employee error | On 6/4/2009 each customer was telephoned and instructed to immediately return the recall lot back to the firm. Replacement product was shipped. |  | 310 | 31 boxes/10/11 ml tubes | Nationwide Distribution -- PA, VA &amp; WA. | Terminated | February 26, 2010 | 267 | 310
Z-2253-2009 | 52961 | No | LXV |  |  | Ear Nose &amp; Throat | 510(k) | VEST 6.6 or earlier software for use with Neuro Kinetics products. |  |  | July 24, 2009 | 2009 | Neuro Kinetics | Marketed without a 510k or PMA submission to include the normative data display. | PMA | A notification letter was emailed to customers starting on August 10, 2009. The letter stated that the normative display needed to be removed pending FDA clearance.Direct questions to Neuro Kinetics, Inc. by calling 1-412-963-6649. |  | 72 | 72 units | Worldwide Distribution -- US (states of AL, AZ, CA, DC, FL, GA, HI, IN, MA, MD, MI, MN, NC, NY, OR, PA, SC, TN, UT, VA, WA, and WV), Korea, United Kingdom, and Spain. | Terminated | January 13, 2010 | 173 | 72
Z-0001-2010 | 53426 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron Delivery Catheter 070, percutaneous catheter, catalog numbers PND6F0701058, PND6F070956, PND6F0701058M, and PND6F070956M; manufactured by Penumbra, Inc., Alameda, CA. |  |  | September 22, 2009 | 2009 | Penumbra Inc. | Failure-- product can kink or ovalize, which may result in difficulty in catheter advancement or delivery of other devices and damage to catheter, which could lead to patient injury or death as a result of portions being retained in the patient. | Device Design | The firm issued an Urgent Voluntary Field Removal Notice letter on 09/22/2009 to its consignees, asking them to inspect their stocks; remove affected product and place in quarantine for return to Penumbra Inc.; and complete and return the attached product identification / return form. |  | 2246 | 2246 units | Worldwide Distribution -- USA, Canada, Hong Kong, Chile, the UK, Denmark, Norway, Sweden, Finland, Netherlands, Germany, Switzerland, Austria, and Spain. | Terminated | November 24, 2010 | 428 | 2246
Z-0010-2010 | 52202 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Multi-Drug 8 Panel Integrated iCup Drug Screen Test Cup. Part Number DUE-187-071-08A. |  |  | April 27, 2009 | 2009 | Biosite Inc | The cup associated with these lots of product may have a pin-sized hole in the floor of the cup. As a result, there is a potential for affected lots of cups to leak urine after an extended period of time. | Process control | Biosite, Inc. issued a "Notice of Correction" letter dated May 1, 2009 to consignees describing the problem and instructed users of proper use in order to avoid any cross contamination of samples until all affected product is completely used.For further information, contact Biosite, Inc. using the following information:Biosite Incorporated, an Inverness Medical Company9975 Summers Ridge RoadSan Diego, CA 92121Phone: 1-877-441-7440, option 2Fax: 1-858-695-7100Email: productadvisory@invmed.com |  | 113575 | 113,575 cups distributed in total | Nationwide Distribution | Terminated | April 02, 2012 | 1071 | 113575
Z-0018-2010 | 52866 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare PCA with Hospira Mednet, list numbers 20709-04-01, 20709-04-03, 20709-04-51, 20709-04-53, 20709-04-77; AC cord manufactured by Electri-Cord, Westfield PA, Infusion pumps manufactured by Hospira, Morgan Hill, CA. |  |  | August 11, 2009 | 2009 | Hospira Inc | Fire/Shock hazard-- A component used in the device may cause sparks/flashes and poses a fire and shock hazard. | Component design/selection | Firm initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and an Urgent: Device Field Correction letter to consignees was sent via UPS. The letter recommended that customers check their infusion pumps and spare parts in their inventory to determine if the specified AC Power Cords are in use at their facility; and, to inspect the affected cords for bent or cracked prongs and bridge connection, burnt plastic, or excessive wear and tear. Also, it was recommended that customers immediately discontinue use of the affected cords exhibiting the described characteristics and contact their local Hospira representative for assistance. The letter stated that customers may continue to use and monitor cords that are not exhibiting the described characteristics until replacement cords are sent to their facility. Hospira will contact customers to schedule replacement. Customers are to complete and return the attached Reply Form. Question should be directed to Hospira Customer Service at 1-800-241-4002. |  | 204999 | 204,999 total | Worldwide Distribution | Terminated | November 30, 2010 | 476 | 204999
Z-0036-2010 | 52760 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ROC Lumbar bolt caddy, Part Number: 81997-01-05. |  |  | December 16, 2008 | 2008 | Alphatec Spine, Inc. | Alphatec discovered that the lot of bolt caddies had the incorrect bolt length laser etched on one side of the caddy. The actual ROC bolts themselves are correctly laser etched. | Employee error | Alphatec contacted consignees via phone on December 16, 2008 and informed consignees that they will arrange to have Federal Express contact them for pick-up.Consignees were asked to check their inventory to see where this product is located, and ensure that it is put on hold for return; and what days and hours they are in the office, (Confirm address: attach copy of the pick-up request form.)If they have any follow up questions, they were told to call 1-800-922-1356. |  | 7 | 7 Screws | Nationwide Distribution -- Including states of TX, FL, &amp; NJ. | Terminated | November 13, 2009 | 332 | 7
Z-0038-2010 | 53081 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | AGI-2238 Portable 4 Cylinder Manifold Cart, for connecting medical gas cylinders. |  |  | August 06, 2009 | 2009 | Airgas Puritan Medical | Yoke may become displaced. Yoke index safety system could become compromised, allowing a user to attach the wrong gas type to the yoke connection. | Component design/selection | Airgas issued a Medical Device Removal letter dated 8/6/09 to its customers informing them of the problem and the need to have a yoke back-up plate installed on the Manifold Block Assembly. |  | 208 | 208 | Nationwide distribution. | Terminated | January 26, 2010 | 173 | 208
Z-0059-2010 | 52847 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Medical Gas Yoke Assembly. Generic 4-Cylinder Manifold Cart Model #8070. |  |  | August 10, 2009 | 2009 | Western - a Scott Fetzer Company | The yoke pins of the attached medical gas yoke assemblies could become displaced to such a degree that the yoke pin index safety system could become compromised and allow a post valve of one gas type to become interchangeable with a yoke connection of another gas type. | Component design/selection | Western Enterprises issued an "Urgent-Device Recall" notification dated August 20, 2009 informing Consignees of the affected product. Consignees are to contact the firm to request backup plates and installation instructions or request an RGA and shipping instructions for return of the manifold block assembly for modification.For further information, contact Western at 1-800-783-7890. |  | 18 | 18 cart assemblies containing 72 yoke assemblies (36 Oxygen yokes and 36 Nitrous Oxide yokes) | Nationwide Distribution -- AL, AR, CA, CO, FL, GA, HI, IL, IN, MA, OH, OR, PA, TX, WA and WI. | Terminated | April 25, 2012 | 989 | 363
Z-0064-2010 | 53033 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | 200996-Cable Assembly, Hybrid Extended Length, |  |  | July 17, 2009 | 2009 | Mako Surgical Corporation | Cables fail to provide adequate transmission. MAKO Surgical Corporation is recalling 200996 Cable Assembly, Hybrid Extended Length; 200997 Cable Assembly, Hybrid, Guidance Module; and 200998 Cable Assembly, Hybrid, UI Panel. These parts are sub components of the 20300 Rio Standard System. | Device Design | A Recall Notice, dated July 15, 2009, was sent to owners of the RIO system (hospital OR staff) with a signed acknowledgement returned to MAKO. Surgeons using the RIO system during the recall time period, regardless of whether their system has already been upgraded, are given the notification letter and asked to sign and return the acknowledgement memo. The letter stated that customers' RIO will be upgraded to replace the potentially faulty cable with a new cable. Their MAKOplasty Specialist (MPS) will notify them when their system was scheduled to have their fiber optic cable upgrade. The upgrade will be installed by a Field Service Engineer (FSE) free of charge and the RIO will be completely tested prior to being returned to service. Customers should direct questions or concerns to their MPS, FSE, or Jim Pomeroy, who can be reached at 954-927-2044. |  | 23 | 23 | Nationwide Distribution | Terminated | May 12, 2010 | 299 | 23
Z-0065-2010 | 52964 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Emerald Hematology Analyzer, model number 09H39-01. Distributed by Abbott Diagnostics Division, Santa Clara, CA. |  |  | July 16, 2009 | 2009 | Abbott Laboratories | Software bug. The CELL-DYN Emerald WBC L1 flag is configured incorrectly. The flag is not generated according to system requirements. | Software design | Abbott Diagnostics issued a "Product Correction" notice dated July 16, 2009 describing the affected product. The notice was sent to all currently active users and included a customer reply form, a software update and installation instructions.For further information, contact Abbott Customer Support at 1-877-4ABBOTT. Customers outside the U.S. should contact your local hematology customer support representative. |  | 411 | 411 units | Worldwide Distribution -- United States (AL, AR, AZ, CA, CT, FL, GA, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV), Canada, Germany, Puerto Rico, Singapore, and South Korea. | Terminated | December 06, 2010 | 508 | 411
Z-0066-2010 | 52632 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad CycleSure Biological Indicator, Product Code 14324. |  |  | June 23, 2009 | 2009 | Advanced Sterilization Products | Biological Indicator indicating a positive result does not maintain yellow color for the entire incubation period and some indicators experience media evaporation prior to the end of the labeled maximum incubation time. | Device Design | An Important Product Labeling Correction - New Instructions for Use dated, June 26, 2009, has been proposed to be sent by the firm for Cyclosure Biological Indicators to consignees. The revised instructions reduce the maximum read time for the indicators to 3 days (72 hours). The letter also states why ASP changed the Instructions. Questions are to be directed to ASP Customer Care Center. |  |  |  | Worldwide Distribution. | Terminated | November 29, 2011 | 889
Z-0067-2010 | 52808 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Arthro-Knife, ConMed, Linvatec 11311 Concept Blvd, Largo FL 33773, Made in USA. RX only. U.S. PAT NO. 5.141.517. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed has determined there is a possibility that the device may not fully retract once it is deployed. | Process control | ConvaMed Linvatec USA issued an "Urgent Medical Device Recall" notification via Fed Ex to all direct consignees of the affected product. Users were asked to inventory and return all affected product to the firm and to return a reply form by mail or fax.For further information, contact customer service at 1-800-237-0169 for US or your local ConMed Linvatec office for outside US or email at Custserv1@linvatec.com. |  | 66144 | Approximately 66,144 products. | Worldwide Distribution | Terminated | February 25, 2011 | 574 | 66144
Z-0069-2010 | 53498 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Extended PTFE Coated Blade electrosurgical electrodes, 6.5" (16.51CM) Length, 12 each &amp; 20 each per box, Labeled MEDLINE brand. |  |  | August 28, 2009 | 2009 | Unimed Surgical Products, Inc. | Unimed Surgical Products, Inc., Largo, FL initiated a voluntary recall of the Extended PTFE Coated Blade electrosurgical electrodes, product number ES0014, due to concerns regarding product sterility for Lot #042209-03. | Process control | Unimed Surgical Products, Inc. issued a Medical Device Removal letter dated August 28, 2009. The customer, Medline Industries, Inc., was instructed to quarantine any undistributed product and to notify its customers of the recall.For further information, contact Unimed Surgical Products, Inc. at 1-727-546-1900 or 1-800-886-1225. |  | 1920 | 1920 units | Nationwide Distribution -- MI | Terminated | May 12, 2010 | 257 | 1920
Z-0068-2010 | 53435 | No | MAF |  |  | Cardiovascular | PMA | Boston Scientific LiberteÂżÂż MonorailÂż and Over-the-Wire Coronary Stent Systems to be relabeled as VeriFLEX". |  |  | September 11, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific initiated a field correction for the LiberteÂż Bare-Metal coronary stent products. They have received reports from cardiac cath labs that TAXUS Liberte Paclitaxel-Eluting (TAXUSÂż LibertÂżÂż Drug-Eluting) coronary stents have been inadvertently selected when the interventional cardiologist intended to implant a LiberteÂż Bare-Metal stent, and LibertÂżÂż Bare-Metal stents have been ina | Labeling design | Boston Scientific issued an "Urgent Medical Device Field Correction" letter dated September 18, 2009. The letter was addressed to Risk Manger/Field Action Contact. The letter described the product, problem and actions which included renaming the LiberteÂżÂż Bare-Metal stent as VeriFlex" Bare-Metal stent. Enclosed with the letter included the Direction for Field Correction and a Reply Verification Tracking Form.For further information, contact you local Boston Scientific Sales Representative or call 1-800-811-3211. |  | 391769 | 391,769 | Nationwide Distribution -- United States including Puerto Rico. | Terminated | June 28, 2012 | 1021 | 391769
Z-0073-2010 | 53163 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biotrak PLLA Pin, Resorbable Fixation System, Biotrak 20 mm Pin; sterile pin packaged in sterile bag, enclosed within a box labeled Biotrak Resorbable Fixation System Biotrak 20 mm Pin REF 30170320-S Manufactured or Distributed by: ACUMED LLC 5885 NW Cornelius Pass Rd Hillsboro, OR 97124. |  |  | August 21, 2009 | 2009 | Acumed LLC | Labeling may contain 2 lot numbers and 2 model numbers.It was observed that the box top label (30170060-S, W67391) does not match the two side labels (3010320-S, W67385). The part and patient labeling inside the box match the side labels (30170320-S, W67385). | Pending | Urgent Notice: Device Recall Letters dated August 21, 2009 were sent to all consignees advising of the issue and and required action. Customers are to identify and quarantine the product having the affected lot code(s). Customers are to also return all product from the specified lot(s) to Acumed immediately. If product was further distributed, then those consignees should be notified of the recall and return of the product(s) should be requested. Questions should be directed to Acumed Customer Service at 1-888-627-9957. |  | 125 | 125 | Worldwide Distribution -- USA and South Africa. | Terminated | November 06, 2010 | 442 | 125
Z-0074-2010 | 53296 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Kerr TempBond Clear, Part # 33351, 1717 West Collins Avenue, Orange, CA 92867. |  |  | September 09, 2009 | 2009 | Kerr Corp | The recall was initiated because the catalyst portion of the syringe may have polymerized, therefore making the product unusable. | Nonconforming Material/Component | Kerr Corporation issued an "Urgent Medical Device Recall" dated September 8, 2009. The letter informed customers that the product recalled was shipped to their establishment between June and July 2009 and if they have any of the affected products listed to contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. |  | 13518 | 13,518 units | Worldwide Distribution -- United States, Canada, Mexico, Hong Kong, Israel, Singapore, Colombia, Dominican Republic and Europe. | Terminated | December 12, 2011 | 824 | 13518
Z-0080-2010 | 53327 | No | JEY | Bone plate. | Dental | Dental | 510(k) | SYNTHES TI MATRIXMANDIBLE SAGITTAL SPLIT PL 1.5MM THK/4H/45MM, Catalog # SD448.005. |  |  | September 03, 2009 | 2009 | Synthes USA (HQ), Inc. | The recalling firm has determined that the customized plates instruments do not have a cleared premarket notification. | PMA | The customer account was notified by an Urgent: Medical Device Removal letter via USPS certified mail on 9/4/2009. The letter stated that the products did not have clearance for the hospital's intended use and that if the customer has any of the identified devices, they should call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Also, the customer should complete the Verification Section, and return the letter along with any remaining product. For questions the hospital could call 610-719-5865. |  | 15 | 15 | Nationwide Distribution -- State of CA. | Terminated | March 08, 2010 | 186 | 136
Z-0086-2010 | 52048 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ortho Development Corp., Pivot BiPolar Head, 28mm, Hip Prosthesis, Femoral Head. |  |  | April 29, 2009 | 2009 | Ortho Development Corporation | A Pivot BiPolar Head, the femoral head in a hip prosthesis, might not move freely after assembly due a stacking up of manufacturing tolerances. | Device Design | Ortho sent a Field notification Memo dated 4/2/2009 to agents and distributors, giving them pre-implant examination guidelines. Decision to recall non-implanted units was made on 4/29/2009 and this was communicated by telephone to distributors. |  | 23 | 23 | AZ, CA, FL, MA, MI. No military, government or foreign distribution. | Terminated | May 25, 2010 | 391 | 23
Z-0083-2010 | 53073 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Foot Pedal, Foot Control for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and their accessories. Catalog Numbers: EMAX2-FP and FPLUS-FP. |  |  | August 13, 2009 | 2009 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling the Foot Control/ Foot Pedal for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and accessories. The Foot Control Cables lacks the shield connection from the cable to the connector body. The lack of shielding may result in radiated emissions which could possibly cause interference on other electrical equipment adjacent to the foot pedal | Device Design | Anspach issued an "Urgent Product Removal" Letter dated August 17, 2009 informing customers of the affected product. Consignees are asked to screen their inventory and return any affected product including a completed Product Replacement Form to the firm.For further questions, contact Anspach at 1-800-327-6887 or 1-561-627-1080. |  | 14 | 14 | Worldwide Distribution -- Canada, Indonesia, United Kingdom, Korea and United States (NC, OK, TN, CA, KS, MI and WA). | Terminated | January 05, 2010 | 145 | 14
Z-0035-2010 | 51228 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir 135 Thermal Regulating System. Catalog Numbers: 86186-115 V, 06186-115 V (sold in the U.S.) and #86185-100 V ( not sold in the U.S.). Intended for warming |  |  | February 16, 2009 | 2009 | Cincinnati Sub-Zero Products Inc | The blower on the WarmAir 135 Warming Unit may overheat. Resistance in a connector between circuit board and wiring harness results in excessive heat that may melt connector or deform the outer casing of the blower unit. | Process control | The firm sent an "URGENT Medical Device Recall' notice via First Class Mail on 2/16/2009 (Note: the letter is dated 2/13/2009). The notice informs the customers of the potential for the device to overheat and of the 'small number of customer complaints received by CSZ concerning this issue. In addition, the notice asks that the customers inspect their units for signs of discoloration indicative of a faulty connector. In the notice, the customers were instructed to return the product to CSZ for repair by contacting their Customer Service division and obtaining a Return Authorization number and specific instructions concerning packaging and returning of the unit(s) for repair. The customer is also instructed to complete the attached Recall Response Form which accompanies the recall notification. The notification provides that customer with a toll-free number at (800) 989-7373ext.3234 or call collect at (513) 719-3234 to reach CSZ representative between the hours of 8:30 a.m. until 4:30 P.M. Monday through Friday EST. Once the customer ships the recalled unit to CSZ, the firm will repair the unit(s) using new connectors manufactured, after which time, the repaired unit is returned to distribution for customer use. |  | 2674 | 2,674 units | Nationwide Distribution -- DE, TX, PA, NJ, LA, MI, GA, MA, OH, KY, NY, FL, NV, NC, MT, RI, WA, MN, AL, CT, MD, TN, UT, VA, CO, MO, IA, AZ, AR, IN, NE, KS, WI, ME, MS, and Puerto Rico. | Terminated | September 19, 2012 | 1311 | 2674
Z-0090-2010 | 53192 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Vacuum Collection Bottle, 1 gallon (3200 mL); Ohio Medical Company, Gurnee, IL 60031; part number 0212-0302-600 (6700-209-500); the bottles are sold alone and with various cap &amp; float assemblies: |  |  | September 21, 2009 | 2009 | Ohio Medical Corporation | The re-useable collection bottles were made with polycarbonate instead of the specified polysulfone. The recommended cleaning of the bottles include use of ammonia and autoclaving, which may have an adverse impact on polycarbonate. | Process control | Ohio Medical Corporation issued recall letters beginning September 17, 2009. Consignees were informed of the affected product and requested to return affected collection bottles to the firm. The accounts were requested to complete and return the enclosed acknowledgement form, indicating the number of affected units on hand, to obtain a returned material authorization for the return of the recalled bottles and their replacement. For further information, contact Ohio Medical Quality Assurance Manager, Anita Christ, at 847-855-6270. |  | 823 | 823 bottles | Worldwide Distribution -- United States, Canada, Chile, Costa Rica, Cyprus, Egypt, El Salvador, Greece, Hong Kong, India, Israel, Japan, Jordan, Korea, Kuwait, Malaysia, Netherlands, Oman, Paraguay, Philippines, Qatar, Saudi Arabia, Singapore, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | June 22, 2011 | 639 | 6715
Z-0088-2010 | 53086 | No | LNQ |  |  | Gastroenterology/Urology | HDE - Humanitarian Device Exemption | Medtronic Enterra Therapy, Gastric Electrical Stimulation System, consisting of model 3116 Enterra Therapy Neurostimulator and model 4351 Unipolar Intramuscular Lead. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | August 25, 2009 | 2009 | Medtronic Neuromodulation | Medtronic is in the process of updating the Enterra Therapy labeling to include information regarding the risks of bowel obstruction and/or perforation. Medtronic has received 15 reports since November 2002 of bowel obstruction and/or perforation of the bowel associated with Enterra Therapy System leads. Of the 15 reported events, 11 required surgical intervention (all patients recovered withou | Device Design | Medtronic Neuromodulation issued a "Medical Device Correction" letter beginning August 25, 2009 addressed to Healthcare Professionals. The letter described the Nature of the Issue, Scope, Potential Severity of the Issue, Enterra Therapy System labeling updating information warnings, Recommendations and other Additional Information.For further information, contact Medronic Neuromodulation Technical Services at 1-800-707-0933. |  | 1909 | 1909 (1095 US, 814 OUS) | Worldwide Distribution. | Terminated | May 22, 2012 | 1001 | 1909
Z-0094-2010 | 53436 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wallach LL50x Cryosurgical Freezer CO2 Unpinned Yoke System |  |  | September 29, 2009 | 2009 | Cooper Surgical, Inc. | CO2 Cryosystem thermal load may exceed the exhaust gas phase, resulting in ice formation in the exhaust valve. Should this occur, the Cryo tip may cool at a less than optimal rate, resulting in ice formation in the exhaust valve. | Device Design | Cooper Surgical notified customers by letter titled "Notice of Recall - Urgent" dated September 29, 2009 advising them of the problem and requesting return of the device. International accounts were notified by letter Beginning October 1, 2009.For further information, contact Cooper Surgical by fax at 1-203-799-2002 or phone at 1-203-799-2000. |  |  |  | Worldwide Distribution -- United States, Canada, Australia, Belgium, England, France, Germany, Greece, Hungary, Iceland, Italy, Ireland, Latvia, Lithuania, Netherlands, Portugal, Spain, Solvenia, Sweden, Hong Kong, Republic of China, Taiwan, Angola, Cameroon, Congo, Croatia, Egypt, Ghana, Guam, Guineae Bissau, Indonesia, Kenya, Kuwait, Lebanon, Madagascar, Mali, Mauritania, Mozambique, New Zealand, Niger, Pakistan,Qatar, Russia, Sao Tome, Saudi Arabia, Senegal, Serbia, Singapore, South Africa, St Vincents &amp; Grenadines, Oman, Thailand, Turkey, UAE and Zimbabwe. | Terminated | April 17, 2012 | 931
Z-0119-2010 | 53492 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Nucleus Non-Magnetic Plug for removable magnet implants, REF Z50100, STERILE EO, Cochlear Limited. |  |  | September 30, 2009 | 2009 | Cochlear Americas Inc. | Cochlear implant device component was mis-labeled. | Labeling mix-ups | On September 28, 2009, each surgeon who received a non-magnetic plug after December 5, 2008 was sent a notification letter via courier. The letter described the mislabeling issue and actions for consignees. Customers were requested to return affected product (and the quantity discarded) to the firm and review patient records to verify the implantation of the non-magnetic plug and not the mislabled product (sterile magnetic). Direct questions about the recall by calling at 1-800-523-5798. |  | 10 | 10 units | Worldwide Distribution -- United States (AL, AZ, CA, GA, IA, NC, MA, MI, MN, MO, NY, OH, OR, PA, TN, TX and WA), Brazil, and Argentina. | Terminated | December 22, 2009 | 83 | 10
Z-0039-2010 | 53489 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Infant Manual Pulmonary Resuscitator with reservoir bag, face mask &amp; Gas sampling elbow. Product ID # 5801, 12-8713, 732-E, and 2206. |  |  | October 06, 2009 | 2009 | Unomedical | In some Manual Pulmonary Resuscitator units the duckbill valve and retention ring inside the patient valve housing may become dislodged and fall into the MPR bag/body. If this occurs, the MPR may lose its ability to generate the positive pressure necessary necessary to ventilate the patient. | Device Design | Unomedical issued Press on October 6, 2009. Letters were sent to customers on October 7, 2009 via UPS - return receipt. Letters explained the reason for recall and the potential hazard to patients. The letter contains a list of all affected lot numbers as well as a picture which shows that units with a clearly visible blue ring are NOT affected. Distributors are asked to examine inventory and quarantine all affected product and return the response form, even if no recalled product is in inventory. If the product was further distributed, the letter requests that a copy of the letter be sent to all customers and for returns, follow Distributor actions listed in the letter. Retailers and end users/Hospitals/EMT's are asked to perform a count of recalled product, segregate it and return the response form. All customers are asked to call 1-800-634-6003 where they will be assigned a RGA number and given instructions on how to return product.Contact information: 8:00AM - 5:30 PM Central Time Monday through Thursday and 8:00AM to 1:00 PM on Fridays, Central Time. |  | 200 | over 200 total lots | Worldwide distribution: USA, Argentina, Canada, Denmark, France, Greece, India, Italy, Japan, Mexico, Paraguay, Peru, Poland, Russia and Vietnam. | Terminated | August 10, 2010 | 308 | 200
Z-0104-2010 | 53583 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Liquid Glucose (HEXO) Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog #'s G7517-120,G7517-120, G7517-500, and G7517-1L; 8-G7517-120, 8-G7517-500, 3-G7517-L, 7-G7517-500 and 7-G7517-1000. (Clinical chemistry) |  |  | October 19, 2009 | 2009 | Pointe Scientific, Inc. | There is a loss of linearity. The product fails to maintain linearity specification of 500 mg/dL. | Process design | The firm issued a press release, and an amended press release, on 10/30/09. Consignees were notified by an Urgent: Medical Device Recall letter dated 10/19/09, which identified the problem; and, instructed testing laboratories to inform patients' attending physicians, and determine, with their input, whether confirmation of the previous test results will be required. Customers are to examine their inventory for any of the lot numbers listed on the letter; and, for any product that has been further distributed, those customers need to be identified and notified of this product recall by including a copy of the recall letter. Customers are to complete and return the Recall Return Response form and upon receipt, Pointe Scientific will replace any disposed of reagent. The recall was expanded to include two additional lots by letter dated 11/2/09. |  | 474 | 400 kits plus 59 liters and 15 vials. | Worldwide Distribution -- USA, Israel, Mexico, South Africa, and West Indies. | Terminated | September 27, 2010 | 343 | 474
Z-0121-2010 | 53329 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens ARTISTE syngo RT Therapist, Therapist Assist, Therapist Express Basic, and Therapist Express Assist Systems, manufactured by Siemens Medical Solutions, Concord, CA |  |  | September 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A failed data transfer is possible that may result in data not being recorded in the interfaced system. Siemens Medical Solutions, Inc. USA became aware of this potential issue with its ARTISTE Linear Accelerator's Syngo RT Therapist v4.1 Workspace (containing the Siemens brand of Verify and Record application called PRIMEVIEW) and the LANTIS Oncology Information System product. | Software design | The firm issued a TH015/09/S Safety Modification syngo RT Therapist 4.1 Customer Advisory letter, dated September 10. 2009. the letter stated the reason for modification, the affected systems, working time, deadline, reimbursement, and reporting. The firm is preparing a software upgrade for the affected products to correct the problem. The letter was sent via certified mail. |  | 97 | 97 units | Worldwide Distribution -- USA, Spain, Germany, Poland, Australia, Malaysia, New Zealand, the UK, India, Ireland, Egypt, Italy, Belgium, Republic of Korea, and Canada. | Terminated | December 06, 2010 | 452 | 97
Z-0087-2010 | 53574 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis CROSSOVER. Cordis CROSSOVER Sheath Introducer. Thomas Medical Products, Inc. 65 Great Valley Parkway, Malvern, PA 19355 USA. |  |  | October 15, 2009 | 2009 | Cordis Corporation | Cordis has received complaints of stretching or fracture of the sheath shaft during use. | Component design/selection | A Medical Device Recall letter dated October 19, 2009 was issued to consignees by sending an overnight letter to each account. The letter described the affected device, issue and actions for customers.A Cordis Corporation representative will follow-up and contact consignees as necessary to facilitate obtaining signature, faxing the acknowledgement form to the Cordis Corporation, collecting and returning units. A press release from the Cordis Corporation announced a nationwide recall of the CROSSOVER Sheath Introducer available on (http://www.webwire.com/ViewPressRel.asp?aId=106460).In addition to the local sales representative, customers may contact the special Customer Service number at 1-866-854-1672, Monday through Friday from 8:30 a.m. to 8:00 p.m. Eastern Standard Time, if you have any additional questions relative to product return or alternate product. |  | 2362 | 2,362 units | Worldwide Distribution -- United States, Belgium, Switzerland, Germany, Finland, France, Italy, Netherlands, Sweden and United Kingdom. | Terminated | August 12, 2013 | 1397 | 2362
Z-0124-2010 | 53142 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CIVCO MTAPSID2932 Type-S Thermoplastic, Disposable nonsterile (48.3 cm) IMRT Style 29 thermoplastic (3.2 mm), Ref. MTAPSID2932, disposable frame mask, 5 per box. The firm name on the label is CIVCO, Orange City, IA. |  |  | August 31, 2009 | 2009 | Med Tec Inc | Some of the packages contain the incorrect mask style. | Process control | The recalling firm initiated their recall via letter dated 8/27/09 issued certified mail for the domestic customers and e-mail for the foreign customer. The communication informs the customer of the reason for recall and offers to replace the product. For questions or for return of the mask, a phone number is provided. A response form was enclosed with the letter for completion to indicate the amount of affected product in the possession of the customer and whether or not they plan to return it. |  | 12 | 12 boxes | Worldwide Distribution -- USA, including states of FL, TX, and OH and country of Australia. | Terminated | March 07, 2012 | 919 | 12
Z-0085-2010 | 53576 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | EndoClamp aortic catheter label reading in part: "CARDIOVATIONS EndoClamp aortic catheter 65 cm,STERILE R, REF EC65, Edwards |  |  | September 24, 2009 | 2009 | Edwards Lifesciences Research Medical, Inc. | Spontaneous balloon rupture during use of the product. | Process control | Customers were notified by an Urgent - Product Recall letter, dated 9/29/09, and instructed to check their inventory, identify any unused product and return it to the company. Customers are to contact Customer Service Organization at 1-800-424-3278 to obtain a Returned Goods Authorization number, information concerning replacement product, and if they have any questions. |  | 2737 | 2,737 total | Worldwide Distribution -- US, Europe, Canada, Australia, and Thailand. | Terminated | April 25, 2012 | 944 | 2737
Z-0151-2010 | 50763 | No | EKC | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Neos Vanadium Excavator Part Number 591/0 |  |  | November 14, 2008 | 2008 | Sybron Dental Specialties | The tips of the Neos Vanadium Excavators may prematurely release from the handle during use, and if so, the tip may be accidentally swallowed. | Component change control | Sybron Dental wrote to the distributor by letter dated November 20, 2008 and posted Nov 24, 2008 titled "URGENTMEDICAL DEVICE RECALL," informing them of the recall, and the reason for recall. Sybron informed the distributor to return affected batches of product. Any product in stock with the mentioned batch number was replaced at no charge or credited to their account.The consignee was also asked to contact KerrHawe Customer Care at +41 91 610 05 05 to receive an RGA number. The affected product should be returned to KerrHawe at the following address:KerrHawe SAVia Strecce 4CH-6934 BioggioSwitzerlandPLEASE LABEL RETURNED PRODUCT "RECALLED PRODUCT RETURN". In addition, they asked for the completed RETURN FORM to be returned by fax to +41 91 610 05 14 |  | 135 | 135 | Worldwide distribution - Austria, Belgium, Denmark, France, Germany, Italy, Netherlands, Portugal, Spain, Switzerland, China, and the United States. One US consignee. | Terminated | August 09, 2012 | 1364 | 355
Z-0145-2010 | 53257 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Power cord for Q2 Continuous Cardiac Output (CCO)/SO2 Computer, list #s 52235-04-01, 52235-04-03, 52236-04-05, 52235-04-07, 52235-04-11, AC cord manufactured by Electri-Cord, Westfield PA, Infusion pumps manufactured by Hospira, Morgan Hill, CA. |  |  | August 11, 2009 | 2009 | Hospira Inc | Fire/Shock hazard-- The power cord used in the device may cause sparks/flashes and poses a fire and shock hazard. | Component design/selection | Hospira initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and a letter to consignees was sent via UPS. The firm will replace affected cords and effectiveness checks will be conducted through tracking of servicing. |  | 204999 | 204,999 power cords distributed for all devices | Worldwide distribution. | Terminated | April 04, 2012 | 967 | 204999
Z-0149-2010 | 53251 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical Stair-PRO Stair Chair with wheels and a track system, patient transport device, Model 6252. |  |  | September 02, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The front of the handle may become detached from the chair, which could result in a patient fall or an operator injury. | Device Design | Consignees were sent a letter titled URGENT MEDICAL DEVICE CORRECTION dated 9/9/09, advising them of the problem and to remove the devices from service until they have been repaired. U.S. customers were notified that Stryker will complete the correction while international customers were sent a kit with instructions for completion of the correction. International consignees were told that Stryker will complete the correction if the consignee does not feel comfortable in doing so. Questions or comments regarding this issue should be addressed to Sarah Marie Puterbaugh of Stryker Medical at (800) 869-0770 ext. 6902. |  | 165 | 165 | Worldwide distribution - Colorado, Ohio, Wisconsin, and Germany, Italy, Netherlands, Poland and Switzerland. | Terminated | February 16, 2011 | 532 | 230
Z-0139-2010 | 53301 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | John Bunn 1392 Series Jackson Tracheostomy Stainless Steel Tube. |  |  | September 14, 2009 | 2009 | Gf Health Products, Inc. | Incorrect material. Manufacturer changed the material of the trach tube from Stainless Steel to German silver/nickel, contrary to product specifications. | Nonconforming Material/Component | GF Health Products (Graham Field) notified consignees by certified US Mail on/about 09/14/2009. The distribution to the dealers includes the announcement of the recall to the dealer the letter to send to the end users and a Proof of Compliance card for the dealer to submit to Graham Field following completion of the recall instructions. Graham-Field instructed the dealers to return any impacted products on site, as well as any delivered by end-users, customers or other impacted third parties. Dealer will be compensated for executing the recall. Graham Field does not have direct access to end user names; however it plans to reach end-users by financially supporting a mailing from dealers to their individual customers notifying them of this recall. A follow up 60 day certified mailing will be performed for non-responders. The dealers will receive financial support for this second mailing. A secure third company will be contracted to hold for destruction all products affected by this recall. The method of destruction will be to cut the products into pieces, make them not usable. The FDA will be notified prior to the destruction of product and detailed logs and records will be maintained. Periodic effectiveness status report will be maintained for this recall. |  | 396 | 396 units | Nationwide distribution. | Terminated | December 30, 2010 | 472 | 396
Z-0134-2010 | 53177 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Stryker Imaging, Guardian Services Software, Part Number 0240-009-430 with Part Numbers 0240-020-013 and 0240-020-022 (Office PACS Power). |  |  | August 31, 2009 | 2009 | Stryker Imaging | Firm has determined Guardian Services does not contain 100% data backup for 12 OfficePACS Power customers on a daily basis. | Other | Stryker Imaging issued an "Urgent: Device Correction" notice dated August 18, 2009. Customers were notified of the defective product and informed that Stryker will be in contact shortly to coordinate the product correction. Actions to be taken during the interim were also detailed.For further information, contact Stryker Imaging at 1-972-410-7100. |  | 12 | 12 devices affected | Nationwide Distribution -- AR, TN, IL, GA, TX, WA, MI, PA, and MS. | Terminated | April 14, 2010 | 226 | 12
Z-0132-2010 | 52833 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000 System and ARCHITECT i2000 SR wash buffer and wash zone mechanism grounding kit. |  |  | July 17, 2009 | 2009 | Abbott Laboratories, Inc | Under certain conditions, the product's wash buffer containing sodium azide can come in contact with the wash zone mechanism ground cables on either of the two wash zones. The copper in the ground strap may corrode and form an unstable chemical substance (metal azide) that may be sensitive to direct pressure and impact. | Other | A product correction notice was issued to consignees on July 17, 2009. Field representatives were instructed to visit consignees, inspect cables and replace them |  | 4526 | 4,526 | Worldwide Distribution | Terminated | January 26, 2010 | 193 | 4526
Z-0138-2010 | 52056 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | Potential pump failure associated with the Coolix 4000 chiller of the Innova single plane and biplane systems that may stop the chiller from working. If a chiller failure occurs, the user will be informed with a message on the n-room monitor stating that three minutes are left before the X-ray is inhibited. This could lead to a delay in treatment and possible additional X-ray exposure and contra | Device Design | Consignees were sent on May 1, 2009 a GE Healthcare " Urgent Medical Device Correction" letter dated April 16, 2009. The letter was addressed to Hospital Administrator/ Risk Manger, Mangers of Radiology/ Cardiology, and Radiologists/ Cardiologists. The letter described the problem, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 61 | 15 AND 46 UNLOCATED | Worldwide distribution: USA, ARGENTINA, BELGIUM, BRAZIL, BULGARIA, CHINA, CROATIA, EGYPT, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LATVIA, LIBYA, MEXICO, NORWAY, POLAND, RUSSIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, ROMANIA, ISRAEL, UKRAINE, YEMEN, BELARUS, and AUSTRALIA. | Terminated | December 03, 2015 | 2422 | 234
Z-0003-2010 | 52990 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx (M3535A/M3536A) Defibrillator/Monitor. |  |  | August 21, 2009 | 2009 | Philips Healthcare Inc. | incorrect default configuration settings, insufficientbattery power as the user may not be alerted that no battery is installed or that a low battery condition exists prior to disconnecting from AC or DC power | Process change control | Philips Healthcare issued a Urgent Voluntary Medical Device Correction Notification on August 21, 2009 advising users of the problem and Customers are asked to follow the "Action to be taken by customer/user" section of the UrgentVoluntary Medical Device Correction Notification. In this section, customers are given the instructions to correct their devices. For further information or support concerning this issue, please contact Philips representative at 1-800-722-9377. |  | 1523 | 1,523 units | NationwideCanadaAustralia, Austria, Beharin, BRUNEI DARUSSALAM,Chian,. Finland, France, Germany,Hong Kong, India, Indonesia, Ireland, Israel,Italy, Japan, Libya, Lithuania. Mexico, Nambia, Netherlands, Palestine, Philippines, Poland, Portugal, Qatar,Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, UAE, UK | Terminated | April 25, 2017 | 2804 | 1523
Z-0141-2010 | 53137 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cardinal Health, Craig Calibrated Cannula, overall length 5-3/8" (13.7 cm), catalog #OS5000-005, sold separately, or as a component of the Craig Vertebral Biopsy Set, catalog #OS5000 |  |  | August 28, 2009 | 2009 | Cardinal Health | The Craig Calibrated Cannula does not interface with the toothed cutter properly, which may prevent the toothed cutter from properly advancing through the cannula to the biopsy site. | Nonconforming Material/Component | Cardinal Health sent Urgent Product Recall letters dated 8/28/09 to all customers who received the product during the recalled period on the same day via UPS Next Day Air. The accounts were informed of the problem with the cannula, and were instructed to functionally test the cannula by passing the toothed cutter within it to ensure the cutter advances freely into the cannula. If any restriction is noticed, the product should be segregated and returned to Cardinal Health. The accounts were requested to complete and return by fax the enclosed acknowledgement form, indicating the number of affected cannulas being returned. Any questions were directed to V. Mueller Customer Service at 800-323-9088. |  | 19 | 19 units | Nationwide distribution. | Terminated | May 12, 2010 | 257 | 19
Z-0130-2010 | 53294 | Yes | LWS |  |  | Cardiovascular | PMA | This device is not approved in the United States. |  |  | September 09, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Disease Managment | A subset of active implanted Concerto CRT-D and Virtuoso ICD devices worldwide (estimated at 6,300) may not meet expected device longevity due to gradually increasing current drain caused by low voltage capacitor degradation. This issue may present in the affected devices as reaching the Recommended Replacement Time (RRT) earlier than projected. This issue does not compromise device functionalit | Nonconforming Material/Component | An "Important Medical Device Information" letter, dated September 2009 was hand delivered by Medtronic Representatives beginning 09/09/09. The letter describes the issue, identifies affected devices and provides recommendations for patients with devices in the affected subset.. A Patient letter was mailed to patients beginning 09/28/09.For additional information: Doctors are advised to contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636;Device tracking managers are advised to contact your Medtronic representative, or the Patient Registration Department at 1-800-328-2518 x41832; andPatients are advised to contact Medtronic Patient Services number at 800-551-5544 ext. 41835. |  | 439 | 439 | Worldwide distribution (except Idaho and Wyoming) including countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Kuwait, Netherlands, Norway, Portugal, Saudi Arabi, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | May 29, 2012 | 993 | 6294
Z-0133-2010 | 53461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VARiS 6.2, 6.1 and 6.0 when used with LINK, Model numbers H88 and H85. Manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | September 15, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software issue: Under certain circumstances, when a therapy treatment field is transferred to VARiS Vision 6.2, 6.1 and 6.0 via LINK procedures, the rotation direction will be reversed. | Software design | Recall initiated September 15, 2009. An "Urgent - Field Safety Notice" letter was distributed to all affected users, with a description of the problem and corrective action steps. The letter was also distributed to Varian sales, marketing, and service organizations, informing them of the issue. For further information in the USA and Canada, contact Varian at 1-650-424-5731 or 1-888-VARIAN5 (888- 827-4265). In Europe, contact +41 41 749 8844. |  | 100 | 100 units | Worldwide distribution to 100 consignees throughout the US in IL, NY, MI, NJ, AL, TN, MD, OH, FL, and TX, and to Brazil, Chile, China, Czech Republic, Egypt, Germany, Hungary, Israel, Italy, Japan, Malaysia, the Netherlands, Romania, Slovakia, Slovenia, South Korea, Spain, Taiwan, Thailand and the United Kingdom. | Terminated | July 29, 2011 | 682 | 100
Z-0108-2010 | 53509 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes USA, TI Synex (TM) II Central Body 14mm - 19mm, Catalog number 04.808.001. |  |  | September 14, 2009 | 2009 | Synthes USA (HQ), Inc. | loss of device height | Device Design | Synthes USA issued an "Urgent: Medical Device Recall" letter to hospital accounts, Synthes personal and sales consultants dated September 23, 2009. Users were asked to take inventory, complete and return a Verification form and return all affected product to the firm.For further information regarding the recall, contact Synthes at 1-800-479-6329 or 1-800-620-7025 extension 5453. |  | 267 | 267 units | Worldwide Distribution -- Including the United States and Canada. | Terminated | July 09, 2015 | 2124 | 2275
Z-0125-2010 | 53532 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) 1-Day (ocufilcon D) Toric 30 daily disposable contact lenses for astigmatism --- Name and Address on Label: CooperVision Hamble, SO31 4RF, UK, and Scottsville, NY 14546 USA --- Each lens is supplied sterile in a blister containing buffered saline solution. The blister is labeled with the base curve, diameter, dioptric power, manufacturing lot number, and expiration date of the lens. ClearSight contact lenses are made from a material containing 55% water and 45% ocufilcon D, a random copolymer of 2-hydroxyethylmethacrylate and methacrylic acid. The lenses have a tint (VAT Blue 6) which is added to make the lens more visible for handling. |  |  | August 28, 2009 | 2009 | CooperVision Inc. | Thirty-eight lots of Ocufilcon D 55% 1-Day Toric Contact Lens with a modified packaging solution sold under the brand name ClearSight 1-Day Toric were inadvertently shipped into the US market prior to complete evaluation of the 510(k) by the FDA. | Other | Urgent: Medical Device Recall - Contact Lens letters, dated 8/28/09, Response Forms, along with reorder forms and postage paid return labels were sent to the consignees via first class mail with instructions for returning the affected product. Customers are to examine their inventory and customer records to determine if they have any of the affected product. The affected products are to be returned and CooperVision will reimburse for all postage associated with the return and replace the lenses free of charge. Questions should be directed to Bonnie Tsymbal at (800) 341-2020 ext. 3210. |  | 1440 | 1440 cartons | Nationwide Distribution | Terminated | June 12, 2012 | 1019 | 1440
Z-0140-2010 | 53482 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Olsen Medical, MIDAS TOUCH Âż NEEDLE ELECTRODE, WITH MODIFIED TIP. Model Numbers: 30-1011, 30-1012, 30-1013, 30-1014 and 30-6011. |  |  | September 24, 2009 | 2009 | Kentucky Packaging Service LP | THERE WERE INCONSISTENCIES NOTICED ON THE INSULATION OVERLAP. THE ELECTRODE HAS A LARGE DIAMETER SHAFT THAT TRANSITIONS TO A SMALLER DIAMETER SHAFT. THE 1ST LAYER OF INSULATION WHICH COVERS THE SMALL DIAMETER SHAFT OVERLAPS WITH THE 2ND LAYER OF INSULATION THAT COVERS THE LARGER DIAMETER SHAFT. ON SOME ELECTRODES IT APPEARS THAT THE OVERLAP IS TOO SHORT TO COVER THE TRANSITION WHICH MAKES THIS | Process control | Olsen Medical issued an "Urgent: Device Recall" letter to Consignees beginning September 24, 2009 . Customers were instructed to inventory their stock and for all remaining product immediately discontinue use and return. Consignees were also instructed to contact their customers and notify them of the recall.For further information, contact Olsen Medical at 1-502-772-4272. |  | 3617 | 3,617 pieces | Nationwide Distribution -- AZ, FL, IA, MD, MO, NM, OR, NV, TN, TX and VA. | Terminated | October 05, 2010 | 376 | 3617
Z-0169-2010 | 52639 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Byrd Screw, 2.0X6MM SD, Model Number: 01-0264 |  |  | February 19, 2009 | 2009 | Biomet Microfixation, Inc. | Byrd Screw, Self Drilling, may break during usage. BIOMET has determined that the parts were not correctly processed by the vendor , i.e. that the parts were not hardened in the final processing and that lead to weaker than acceptable material. | Nonconforming Material/Component | Biomet plans to send letters to customers on February 19, 2009. Customers are told to discontinue use of the item and return it to Biomet Microfixation. The firm plans to follow up with an e-mail to their customers. Questions pertaining to this issue, should be directed to Christy Cain at 1.800.874.7711 or 904.741.4400. |  | 22 | 22 | AZ, CA, GA, IL, IN, NM, NY, TX, VA, Puerto Rico and Canada. | Terminated | May 03, 2012 | 1169 | 22
Z-0170-2010 | 53128 | No | LGW |  |  | Neurology | PMA | Eon Rechargeable Implantable Pulse Generator (IPG) System |  |  | December 22, 2006 | 2006 | Advanced Neuromodulation Systems, Inc | Product's external charging system on occasion was failing to communicate with the Implantable Pulse Generator preventing re-charging of the pain relief system. | Device Design | 12/22/2006, Marketing Update 5015 was issued to field sales representatives providing information on improving communication between the Internal Pulse Generator and its charger. May 15, 2007, Marketing Update 5022 was issued to field sales representatives to reiterate avoiding contact between the charger antenna and conductive materials and place the preferred side of the charging antenna face-down over the implant. Troubleshooting Guide was created in January 2008 for field sales representatives and technical services departments regarding isolation of the communication issue to the external devices instead of the implanted component. March 25, 2008, the Clinician's Manual was updated adding a recommendation for a minimum depth of pocket for device implantation. |  | 24750 | 24,750 units total | Worldwide | Terminated | July 23, 2010 | 1309 | 24750
Z-0166-2010 | 53353 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Koodinat M/MP Patient Table used with Angiostar Model 5294009. |  |  | July 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The table swivel locking mechanism may not completely engage adn could result in uncontrolled patient movement during transfer resulting in a patient fall. | Labeling False and Misleading | The recalling firm issued a Customer Safety Advisory Notice dated July 1, 2009 to affected customers via Update Instructions AX025/09/S to inform customer of the potential issues and provide instructions to prevent occurrence. The firm also provided an addendum to User Manual. Direct questions about the recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 107 | 107 units | Nationwide Distribution. | Terminated | January 12, 2010 | 195 | 469
Z-0168-2010 | 53133 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Guardian Services, Part Number 0240-009-430 with Part Numbers: 0240-030-004 and/or 0240-030-002, 0240-030-008 (Net Practice). |  |  | July 31, 2009 | 2009 | Stryker Imaging | Product serving as data backup protection does not contain 100% data backup for all NetPractice customers on a daily basis. | Other | Firm notified consignees by an Urgent: Device Correction letter, dated July 31, 2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also provides Stryker's recommendations for reducing risk to patients. Technical instructions on the backup storage for the NetPractice are included with the letter. Stryker will not bill customers for the Guardian Service until the problem is remediated and Stryker notifies customers. If there are questions, customers should contact Stryker at 972-410-7100. |  | 55 | 55 devices affected | Nationwide Distribution | Terminated | January 08, 2010 | 161 | 55
Z-0159-2010 | 52987 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb ReNuÂż Multi-Purpose Solution, 12 fl.oz. (355 ml), Bausch &amp; Lomb, Incorporated, Rochester, NY 14609. |  |  | August 17, 2009 | 2009 | Bausch &amp; Lomb Inc | During routine testing the product failed to meet shelf life specifications. | Other | Consignees were notified by letter on/about August 17, 2009. All consignees were asked to review their inventory and return any of the affected product to Stericycle Pharmaceutical Services.For questions about the recall process or need additional shipping labels, please contact Stericycle Inc. at 1-888-345-83 16. For questions regarding this recall, please contact Bausch &amp; Lomb's Customer Resource Center at 1-800-828-6974. |  | 5000 | 5,000 bottles | Distribution: Nationwide, Trinidad, St. Croix, St. Thomas, Aruba and Bermuda. | Terminated | May 10, 2012 | 997 | 5000
Z-0002-2010 | 52834 | [6 empty] | Edwards Lifesciences Fogaty Occlusion Catheter Ref # 62080822F. |  |  | June 22, 2009 | 2009 | Edwards Lifesciences, Llc | Several complaints were the two parts of the tubular package became loose under the shrink wrap that holds the DFU on the packaging tube. | Packaging change control | Edwards initiated a worldwide recall via customer letter dated June 29, 2009, sent via Federal express on same day, requesting that customers cease using Fogarty Occlusion Catheter, Model No. 62080822F products from the indicated lots and that they return all unused products. Edwards requested that the customers provide written verification of stock and identify any unused products to be returned to Edwards Lifesciences LLC, Att: Field Action Coordinator, ODW Logistics, Inc., 13725 Pipeline Ave., Chino, CA 91710. Additionally provided Customer Service Tel # 800.424.3278. |  | 9548 | 9548 | AL AR AZ CA CO CT DC DE FL GA HI IA ID IL IN KS KY LA MA MD MI MN MO MS MO NC ND NE NH NJ NM NV NK OH OR PA RI SC SD TN TX UT VA VT WA WI WV WY, Puerto Rico Australia, Belgium, Brazil, Bulgaria, Canada, Canary Island, China, Czech Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Ireland, Italy, Japan, Luxembourg, Netherland, New Zealand, Poland, Portugal, Rumania, Russia, Singapore, Slovak Republic, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, and Turkey. | Terminated | February 23, 2012 | 976 | 9548
Z-0150-2010 | 52791 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL 80 CO-OX SYSTEM, software versions 1.30 and 1.31 |  |  | February 20, 2009 | 2009 | Sendx Medical Inc | Inverness have recently become aware of a potential error contained in V1.30 and V1.31 software for the ABL 80 CO-OX platform. The error can occur if the analyzer is set to automatically upload patient records to Radiance or an HIS/LIS system. This can result in the reporting of incorrect results for one or more CO-OX derived parameters. Incorrect results can be displayed on the screen and the | Other | Radiometer sent out a letter to consignees informing them of public reason for recall, and telling them that they have provided a copy of the upgrade software and instructions to correct this potential problem immediately. Software upgrade instructions were included with the recall letter. Consignees were informed that if they preferred that their local service representative perform the upgrade, to follow the Instructions for disabling auto upload of patients records.SOFTWARE UPGRADE INSTRUCTIONS:The upgrade software must be installed via the Software Upgrade option under Menu&gt;Utilities&gt;Software Upgrade. It is not a "boo table" CD.~The Upgrade process will take approximately 10 minutes. The analyzer should be allowed to complete the entire process without user interaction. Do not attempt to turn the analyzer on or off manually at any time during the process.~The upgrade process will include the following general steps:1. Automatic power down and reboot of the analyzer2. File transfer process which includes a variety of progress bars to communicate the status3. Automatic power down and reboot of the analyzer a second time4. A dialog box on the main menu screen that states, "ABL80 FLEX software has been upgraded to version 1.32." This signifies a successful upgrade process.~If a solution pack and a sensor cassette are already installed on the analyzer, their installation status will not be affected by the software upgrade.~All stored data will be retained in all logsThey were also asked to complete and return the attached fax form, with their signature, to 1-800-736-0601.If their were any questions regarding the letter, or upgrade, to contact Techline at 1-800-736-0600 opt. |  | 317 | 317 | Domestically and Worldwide: Germany, France, Spain, United Kingdom, Poland, Austria, Italy, Tokyo, Australia, and DenmarkDomestically to OH | Terminated | July 16, 2012 | 1242 | 317
Z-0161-2010 | 53136 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medcomp, 10F Bioflex Tesio Kit with cutting trocar. Catalog number BFR-6CTT 5 per box. |  |  | September 02, 2009 | 2009 | Medical Components, Inc dba MedComp | mislabeled as right are actually left | Employee error | Medcomp issued a "Product Alert" letter dated September 2, 2009 to all customers informing them of the problem and the need to return the product. For further information, contact Medcomp at 215-256-4201. |  | 100 | 100 trays | Worldwide Distribution -- United States (AR, FL, KS, LA, MO, PA, TN, and TX) and Austria. | Terminated | April 16, 2010 | 226 | 100
Z-0160-2010 | 53134 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Guardian Services, Guardian Software, Part Number 0240-009-430 (Guardian Services) with Part Numbers 0250-003-016, 0250-003-017 and 0250-003-018 (OrthoPad) |  |  | August 03, 2009 | 2009 | Stryker Imaging | Guardian Services data backup does not contain 100% data for all OrthoPad customers on a daily basis. | Other | Firm notified consignees by an Urgent: Device Correction letter, dated 7/31/2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also states the firm's recommendations to reduce risk to patients. Stryker will not bill customers for the Guardian service until the problem is remediated and the firm notifies customers. If there are any questions, customers are to contact the firm at 972-410-7100. |  | 85 | 85 | Nationwide Distribution | Terminated | January 26, 2010 | 176 | 85
Z-0164-2010 | 53453 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Biograph mCT-S 64, model number 10248672. |  |  | August 11, 2009 | 2009 | Siemens Medical Solutions USA Inc. | In the unlikely event that the device is modified to operate in the tilted position, the telescoping support mechanism will not engage properly. | Device Design | Siemens Medical Solutions USA, Inc. issued an "Urgent Field Correction - Recall" notice to consignees by letter dated September 3, 2009 informing them of the problem and that a Siemens service representative will be in contact shortly to service the affected product.For further information, contact Siemens at 1-800-888-7436. |  | 2 | 2 units | Worldwide Distribution -- United States (IL and OH), Germany, France and Hong Kong. | Terminated | December 22, 2009 | 133 | 3
Z-0171-2010 | 53531 | No | FHW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Coloplast Titan OTR Inflatable Penile Prosthesis,18cm scrotal Bioflex cylinder set w/OTR Pump, P/N QSR9181400. Sterile EO. Manufacturer: Coloplast A/S, 3050 Denmark. Distributed by: Coloplast Corp., Minneapolis, MN 55411 USA. |  |  | October 06, 2009 | 2009 | Coloplast Manufacturing US, LLC | Coloplast determined that a portion of Titan OTR pumps were manufactured using a process that where validation activities were not completed. The Titan OTR pumps were made to our standard product specifications and have passed all in-process testing to ensure mechanical reliability of these pumps. Coloplast is not recommending product explant because there are no identified risks since the prod | Process control | Coloplast issued a customer notification letter dated October 2009 describing the affected product.For further information, contact Coloplast at 1-800-533-0464. |  | 3 | 3 | Nationwide Distribution -- FL, GA, NC. | Terminated | December 12, 2011 | 797 | 3
Z-0173-2010 | 53535 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Morcellex Tissue Morcellator; |  |  | September 28, 2009 | 2009 | Ethicon, Inc. | There is the potential for the duckbill seal to become dislodged. | Device Design | Recall notification letters titled "URGENT VOLUNTARY PRODUCT RECALL" were sent on 10/8/2009 via UPS next day. The letters describe the lots within expiration that are recalled. The motor drive unit is not included in the recall. Letters have photos and illustrations of the product which clearly show the location of the product code and expiration date. Recalled product is to be returned to Stericyle using an enclosed pre-paid authorized shipping label. Ethicon will notify customers when the product is available.Questions are to be directed to Ethicon at 1-888-276-5901. |  | 60930 | 60,930 in US. | Worldwide distribution - US and United Arab Emirates, Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Greece, Israel, Italy, Jordan, Republic of Korea, Mexico, Malaysia, Netherlands, Norway, New Zealand, Puerto Rico, Reunion, Russian Federation, Sweden, Slovina, Turkey, Saudi Arabia, Venezuela and S. Africa. | Terminated | August 02, 2010 | 308 | 60930
Z-0174-2010 | 53564 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with a NAS or HSM license |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incomplete amount of image data; specifically, The image count in the series table of the online and archive database will display the full amount of images received at the Syngo Imaging XS as archived. However, the zip file on the NAS or HSM will contain fewer images-those which have been archived until the time of the error/interruption. | Software design | The recalling firm issued a "Customer Safety Advisory Notice" dated 9/14/09 to all affected customers via Update Instructions IM101/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. Questions regarding this issue are directed to the local SIEMENS Uptime Service Center. |  | 2 | 2 units | The product was shipped to medical facilities in NY and OH | Terminated | July 05, 2012 | 1009 | 2
Z-0165-2010 | 53545 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | * (Box label) SUCTION COAGULATOR 10FR QTY 25 ***** |  |  | September 30, 2009 | 2009 | ConMed Electrosurgery | One production lot of Suction Coagulators received ETO sterilization instead of gamma radiation sterilization. | Process control | Recalling firm notified all distributors by phone from 9/30 - 10/1/09. All customers were notified by letter titled "URGENT: DEVICE RECALL" dated 09/30/2009 on 10/2/2009. Customers were instructed to check their inventory for the affected product, complete the reply form and return affected product. For questions, please contact Damaris Velez or Jeff Dickinson of ConMed Electrosurgery at 800-552-0138, ext 5225 or email succoagrecall AT conmed.com. |  | 2700 | 2700 units | US only. CA, FL, GA, ID, IL, IN, KY, LA, MD, MI, MO, MS, NC, NY, OH, OR, PA, SC, SD, TX, VA | Terminated | March 16, 2010 | 167 | 2700
Z-0176-2010 | 53676 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | RESPIRONICS V60 VENTILATOR |  |  | October 26, 2009 | 2009 | Respironics California Inc | The recall was initiated to correct an issue with a component on the V60 Ventilator Power Management (PM) Printed Circuit Board (PCB). A component on the PCB has the potential to break off or become damaged if the PCB is removed from the ventilator. If this occurs, the ventilator may not properly switch between back up battery power and AC power. This can result in power loss during ventilation fo | Device Design | The recall notification was initiated October 26, 2009 with the firm mailing a Urgent Medical Device Field Correction letter to the affected customers. The letter informed the customers of the reason for the notice and how Respironics Field Service Technicians were going to correct this issue. The letter also gave the customers instructions to follow pending the completion of the field inspection/replacement.For questions or additional information contact Respironics US Customer Service at 1.800.345.6443, Option #5 then Option #1. |  | 524 | 524 | Nationwide and Internationally to :Libya, Hong Kong, France, Italy, India, Spain, United Kingdom,Germany, Netherlands and Saudi Arabia | Terminated | October 13, 2011 | 717 | 524
Z-0175-2010 | 53582 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific 1.0mm x 570 mm rigid pneumatic probe for use only with Swiss LithoClast Select System model no. 840-300K, Swiss LithoClast Ultra System model nol 840-700K, or Swiss LithoClast Ultra Upgraded System model no. 840-701U, compatible with 330mm Ultrasound Probes (REF 840-714 and 840-715), Rx, sterile, no detectable latex, Packaged and Distributed by Boston Scientific Corporation, Natick, MA; REF/catalog no 840-733. |  |  | September 17, 2009 | 2009 | Boston Scientific Corp | The label incorrectly reads "compatible with 330mm ultrasound probes (Ref 840-714 and 840-715)" instead of bearing the correct statement "compatible with 403mm ultrasound probes (Ref 840-716 and 840-7170." | Other | Consignees were notified by letter titled "Urgent Medical Device Recall-Immediate Action Required" dated 9/17/09 and instructed to cease using the product and to call the firm's customer service for return authorization. For additional information, call the firm at 866-868-4004. |  | 173 | 173 | Nationwide. | Terminated | October 08, 2010 | 386 | 173
Z-0177-2010 | 53536 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular Metal Shoulder, humeral stem, 12 mm stem diameter, 130 mm stem length, 42 degrees, sterile, Zimmer, Warsaw, IN, REF 4342-112-13. |  |  | September 17, 2009 | 2009 | Zimmer Inc. | Some of the devices may possess a taper which will not mate with the humeral head, as intended. | Other | Zimmer distributors were notified by email on 9/19/09 to locate and return the products. Letter dated 9/29/09 were sent to the Zimmer salesforce and distributors with a letter to be given to each affected hospital account to inform them of the problem. Zimmer accounts were again asked to locate and return the devices. Hospital accounts were requested to consider the available facts when monitoring potentially affected patients. |  | 12 | 12 | Nationwide. | Terminated | December 22, 2009 | 96 | 12
Z-0190-2010 | 52759 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Alphatec Spine, Inc., NOVEL SPINAL SPACER SYSTEM, Part Number(s): 64817-010, 64817-012, 64817-013, 64817-014, 64817-107, 64113-012, 64113-010, 64765-109, 64733-124. |  |  | May 13, 2009 | 2009 | Alphatec Spine, Inc. | Due to a deficiency in the process validation related to the machine that manufacturedthese parts, Alphatec would like to have the parts returned to them so that they may perform anadditional inspection of each implant. | Process control | Alphatec Spine notified consignees of the affected product via telephone. The firm requested that consignees immediately remove the affected devices in their possession and return to Alphatec Spine. For further information, contact Alphatec Spine at 1-800-922-1356. |  | 37 | 37 devices | Nationwide Distribution -- CA, NV, FL, CT, TX, IA and WI. | Terminated | November 13, 2009 | 184 | 37
Z-0179-2010 | 53139 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | System 1000 family of Hemodialysis Instruments, including the System 1000, TINA, AURORA and AltraTouch; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | August 31, 2009 | 2009 | Baxter Healthcare Renal Div | When the main power is lost temporarily, a software anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's manual. | Software design | Baxter sent an Urgent Device Correction letter dated August 31, 2009 to all System 1000 Series of Hemodialysis Instruments customers via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letters informed the accounts that a software anomaly in software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 was discovered when the main power was lost temporarily during a simulated treatment. The anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's Manual. The accounts were informed that instruments with software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 will be upgraded to version 4.B.1.4 software, provided the instruments have a 486 Motherboard or Single Board Computer, and that instruments containing versions of 4.A software are not affected. The accounts were requested to complete the enclosed Customer Reply form, indicating the instrument serial number and software version for each machine at the site, and if they would be requesting Baxter Technical Services to upgrade their instruments or have the software sent to them to make the upgrades, and fax the sheet to 1-847-270-5457. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 4, option 1. |  | 8258 | 8,258 units | Worldwide Distribution -- United States (including Puerto Rico and the Virgin Islands), Algeria, Argentina, Austria, Bahrain, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, India, Indonesia, Ireland, Italy, Jordan, Korea, Latvia, Lebanon, Lithuania, Malaysia, Mauritania, Mexico, Morocco, Nigeria, Oman, Panama, Paraguay, Peru, Philippines, Poland, Romania, Russia, Saipan, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay and Venezuela. | Terminated | November 15, 2010 | 441 | 8258
Z-0180-2010 | 53083 | Yes | LWS |  |  | Cardiovascular | PMA | Guidant LATITUDEÂż Communicator, model 6481, Boston Scientific 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. The LATITUDEÂż Patient Management system enables physicians to periodically monitor both patient and device status remotely. |  |  | August 07, 2009 | 2009 | Boston Scientific CRM Corp | The L ATITUDEÂż communicator (Model 6481) for a patient may not have been able to receive measurements from their weight scale or blood pressure monitor afer initial setup. The issue has been resolved. If any patient weight or blood pressure readings are stored in the patient's sensor, they will be uploaded to LATITUDE after the patient's communicator calls in to Boston Scientific's servers. It | Software design | Boston Scientific used phone calls as the delivery method for these communications. The communication began August 6, 2009 for the original 6 patients affected. Notifications via phone calls to additional patients affected began on August 17, 2009. Direct questions about the recall to the LATITUDE Customer Service by calling1-651-582-4000. |  | 97 | 97 | Nationwide Distribution -- AZ, AR, CA, CT, FL, GA, IL, IN, MA, MO, NV, NJ, NY, OH, OR, PA, SC, TX, UT, VA and WA. | Terminated | December 04, 2012 | 1215 | 97
Z-0189-2010 | 53494 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Siregraph CF X-Ray System, Model Number: 4466033. |  |  | September 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | possible injury to hand/fingers | Device Design | Siemens Medical Solutions USA, Inc. issued a "Customer Safety Advisory Notice" dated September 11, 2009 to inform customers of the potential hazard to patients when using the Siregraph CF system.For further information, contact Siemens at 1-610-448-4500. |  | 14 | 14 units | Nationwide Distribution -- FL, IL, MI, PA and TX. | Terminated | January 12, 2010 | 111 | 14
Z-0181-2010 | 53458 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis, Inc., Apligraf (draftskin) packaging, Canton, MA 02021. |  |  | September 27, 2009 | 2009 | Organogenesis, Inc. | Unit contaminated with Staphylococcus epidermidis | Packaging process control | Organogenesis, Inc. initiated a recall notification dated September 28, 2009 via fax to Physicians asking for return of unused affected product. If the unit was used, Physicians are advised to use standard wound care precautions to assure the safety of the patient. Users were asked to acknowledge receipt by completing and returning the notification to the firm.For further information, contact Organogenesis, Inc. at 1-888-432-5232. |  | 60 | 60 units | Nationwide Distribution | Terminated | March 31, 2010 | 185 | 60
Z-0188-2010 | 53499 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XTEN LED DF cupola, P/N: 567815999. |  |  | October 07, 2009 | 2009 | MAQUET Inc. | Several reports have been received regarding one model of spring arm (Acrobat 200 series) in combination with surgical light from various Manufacturers. All incidents pertain to a crack on the front joint of the Acrobat arm. | Device Design | Urgent Device Correction letters, dated October 7, 2009, were sent out by FedEx. The letter identified the affected products and asked customers to call 1-888-627-8383 #1, if they have any questions. The letter also provided a Technical Description along with a Customer Fax Response form. The letter stated the potential hazard and recommended precaution. If customers need additional info or have any questions they should contact their local MAQUET representative. |  | 346 | 346 | Nationwide Distribution. | Terminated | June 04, 2012 | 971 | 940
Z-0191-2010 | 53581 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet 5.5 mm peek Allthread PEEK Knotless Anchor, knotless, sterile, Biomet sports Medicine, Inc., Ontario, CA; REF 904842P. |  |  | September 09, 2009 | 2009 | Biomet, Inc. | These anchors may be difficult to remove, or may not detach from the inserter shaft. | Device Design | Biomet issued an "Urgent Medical Device Recall Notice" dated September 9, 2009 informing consignees of the affected product and requesting return of the device. Distributors were instructed to notify their customers copy of the recall letter and to locate and return the products from user facilities.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983. |  | 894 | 894 | Worldwide Distribution -- United States and Belgium. | Terminated | January 19, 2011 | 497 | 894
Z-0105-2010 | 53550 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Tri-State Hospital Supply Corp., Centurion Healthcare Products Premie Packs, 50, sterile, Tri-State Hospital Supply Corp., Howell, MI; code no. LM110. |  |  | September 22, 2009 | 2009 | Centurion Medical Products | The pediatric tracheal tubes used in these kits were manufactured with an internal diameter smaller than indicated in the label, which could result in an inability to remove secretions and cause partial or complete obstruction of the airway. | Process control | The firm's sales representative was notified of the recall on September 22, 2009 via email and the sole customer was notified by letter dated September 23, 2009. All product was removed from the customer account on September 28, 2009 and returned to the recalling firm. |  | 650 | 650 | Illinois. | Terminated | August 03, 2010 | 315 | 1400
Z-0198-2010 | 52836 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer VerSys Hip System Femoral Head, 12/14 taper, 28 mm dia., neck length +0 mm, sterile, Zimmer, Warsaw, IN; REF 8018-28-02. |  |  | August 03, 2009 | 2009 | Zimmer Inc. | Manufacturing material may be present on the surface of the implant. | Manufacturing material removal | An Urgent: Device Recall letter, dated 8/3/2009, was issued to consignees. The letter identified the affected product, stated the reason for the recall, the clinical implications (risks to health), and actions to be taken. Customers are to locate any unused devices and quarantine them immediately, carry out a physical count of affected product and record the data on the Return Certification Form included with the letter, and fax a copy of the completed form back to Zimmer, Inc. at 574-372-4265. Customers are to return the recalled product along with the completed Inventory Return Certification Form. A "Dear Surgeon" letter will be issued to any hospitals where the devices have been implanted. |  | 104 | 104 | Worldwide Distribution -- USA, Brazil, Canada, and Switzerland. | Terminated | December 15, 2009 | 134 | 104
Z-0197-2010 | 53471 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Equalizer Balloon Catheter, Order No./REF 17-112 (M001171120). EQL/40/7/2/100. Sterilized using Ethylene Oxide. Single Use Only. Legal Manufacturer Boston Scientific Cork Ltd., Business and Technology Park, Model Farm Rd, Cork, Ireland. |  |  | September 29, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a Recall Removal of numerous batches/lots of Equalizer Occlusion Balloon Catheter because the sterile barrier in the packaging of the affected product may be compromised. The outer pouch seal may be breached. If a compromised seal is not detected and a sterility-compromised catheter is used clinically, there is a risk of adverse health consequences. | Other | An Urgent Medical Device Recall Removal letter dated 09/29/2009 was sent to affected facilities and addressed to Risk Manager / Field Action Contact, beginning 09/29/2009. The letter explained the issue and identified product. Customers were asked to identify affected product within their inventory, segregate it and immediately return it to Boston Scientific. Customers will receive replacements for all recalled product that is returned. Customers were also to fax back to Boston Scientific the Reply Verification Tracking Form. Questions regarding this Recall Removal are directed to the local Sales Representative. |  | 9800 | 9800 total (2373 US and 7427 OUS) | Worldwide. US distribution includes: AL, AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT. NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Washington D.C. and Puerto Rico. Worldwide distribution includes: Algeria, Argentina, Austria, Belgium, Brazil, Bulgaria, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Lithuania, Luxembourg, Malaysia, Morocco, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates. | Terminated | December 17, 2011 | 809 | 9800
Z-0260-2010 | 53517 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GEMINI PET/CT Systems Upper Patient Pallet for the GEMINI GXL, 16 Slice and 6 Slice scanning device configurations. Model Numbers: 4535-679-71891 and 4535-6797591, 510(k) #K051170. |  |  | January 26, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Upper patient pallet component of the recalling firm's Gemini PET/CT system may be thinner than normally required, which may reduce the product's load bearing capacity, resulting in cracking under less than the designed full-load weight. | Process change control | The firm notified customers with an "URGENT - Field Safety Notice" which was sent to the user/customers via FedEx on 1/26/2009. The letter informs customers of this potential safety hazard; the actions to be taken by the user/customer in order to mitigate risk to patients while undergoing scans; and the actions planned by Philips Medical's Service Engineers in order to correct the problem. The notification refers the user/customers to the applicable sections of their Systems Operations and Systems Information Manuals for a review on procedures for safely loading; unloading; and monitoring patients before, during, and after scanning studies. The notice goes on to inform the customers that replacement parts for the upper patient pallet(s) are on order and that a Philips Healthcare Field Service Engineer will be visiting each customer site within two (2) months in order to replace each affected pallet. Lastly, the notification provides a toll-free telephone number and instructions which the customer can use to contact a local Philips representative if additional information, or assistance is required. The firm's Customer Care Center contact is 1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology. |  | 2 | 2 units | The recalled products were distributed to customers in the following states within the U. S.: OH, GA, MI, and CA and to customers overseas in: Germany, France Japan, and South Korea. | Terminated | November 24, 2009 | 302 | 11
Z-0262-2010 | 53578 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Intramedullary Skeletal Kinetic Distractor (ISKD) is an internal limb lengthening device consisting of a telescoping Internal Limb Lengthener, Locking Screws, Instrumentation and an external hand-held Monitor; catalog numbers: T10-215-265, T10-215-265NS, T10-255-305, T10-255-305NS, T10-255-335, T10-300-350, T10-300-350NS, T12-215-265, T12-245-295, T12-255-305, T12-255-305NS, T12-255-335, T12-300-350, T12-380-380, F12-255-305, F12-255-305NS, F12-255-335, F12-255-335NS, F12-300-350, F12-300-350NS, F12-300-380, F12-300-380NS, F12-345-395, F12-345-395NS, F12-345-425, and F12-345-425NS. |  |  | October 09, 2009 | 2009 | Orthofix, Inc | ISKD lengthening device has the potential for not pre-distracting. This may cause or contribute to device malfunction, and if not discovered prior to use, may require surgical re-intervention or removal. | Other | Firm notified consignees by letter sent via certified mail on 10/9/2009. Follow up is planned for nonresponders. Questions are directed to the firm at (800) 527-0404. |  | 539 | 539 units | Worldwide distribution - In the US: FL, MD, TN, OH, NY, CA, GA. Also Mexico, Italy, Canada. | Terminated | September 24, 2010 | 350 | 539
Z-0199-2010 | 53399 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Natural-Knee Flex System, Prolong Articular Surface, congruent, Size 00,0, left, 9 MM height, use with baseplate size 00 or 0, sterile, Zimmer, Warsaw, IN; REF 00-5424-010-09 |  |  | October 08, 2009 | 2009 | Zimmer Inc. | The surgeon may have difficulty inserting the polyethylene into the tibial baseplate, which may result in intraoperative damage to the device. | Other | An "Urgent: Field Correction" dated October 8, 2009 was issued to Consignees. The notification letter described the affected product, issue, health risks and actions for customers to take. The letter notified consignees of a changed to the surgical technique, to include the use of a new peripheral rib impactor tool.Notifications of this field notification were sent to all affected direct accounts of Zimmer, Inc. Contact your Zimmer Sales Representative ifyou have any questions regarding this action. For shipping assistance, questions or assistance in notifying your accounts about the field correction please contact Zimmer, Inc. at 1-800-613-6131. |  | 26404 | 26,404 of all products | Worldwide Distribution -- United States, Austria, Belgium, France, Germany, Japan, Spain and Switzerland. | Terminated | May 16, 2011 | 585 | 26404
Z-0192-2010 | 53543 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Revolution Blood Pump, "revOlution Centrifugal pump Catalog# 050300000***Rx ONLY***SORIN GROUP ITALIA via Statale 12 Nord, 86 41037 MIRANDOLA (MO) (Italy) Distributed in the U.S. by: Sorin Group USA, Inc. 14401 W. 65th way Arvada (CO) 80004-3599" |  |  | September 25, 2009 | 2009 | Sorin Group USA, Inc. | A weakness in the plastic packaging tray could cause the tray to crack; compromising product sterility. | Packaging change control | Customers were notified by an Urgent: Field Safety Notice, dated 9/25/09, and instructed to: check their inventory for the affected product, complete the customer response form and return or visually inspect all affected product before use. For additional information contact Sorin Group Customer Service at 1-800-650-2623. |  | 808 | 808 units | Worldwide Distribution -- US, Canada, Japan, and Europe. | Terminated | December 14, 2009 | 80 | 1064
Z-0259-2010 | 53603 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Medtronic Integrated Power Console System, IP X1 |  |  | July 17, 2009 | 2009 | Medtronic Xomed, Inc. | An error code (Code 13) is displayed on some of the consoles when a Medtronic Stylus or EHS handpiece is connected to the IPC indicating that the handpiece has not been recognized by the console. | Software design | A Product Correction Notification was issued to each consignee on June 15, 2009. Medtronic sales representatives will visit each facility to advise consignees of the issue and to install a software upgrade that will correct the handpiece recognition issue on the console(s). |  | 493 | 493 | Worldwide Distribution | Terminated | January 20, 2010 | 187 | 493
Z-0268-2010 | 52546 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GMD UNIVERSAL URINARY INCONTINENCE SLING, Rx ONLY Product Code 1020-18. |  |  | June 18, 2009 | 2009 | Generic Medical Devices | Product that is approved only for use outside the U.S. was shipped to two customers in the U.S. | Employee error | On June 17, 2009 the recalling firm called the consignees and arranged for the unused product to be returned. On July 13, 2009 the firm sent a notification letter to the consignee that had implanted 1 of the devices. |  | 15 | 15 devices | USA - NJ, and PA | Terminated | March 14, 2012 | 1000 | 15
Z-0264-2010 | 53131 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Programming M250 Programming System, model number 250, version 4.6, 6.1, 7.0 and 7.1. The Model 250 Handheld software is provided on flashcards packaged in software envelopes or pre-installed on handheld computers. |  |  | July 24, 2009 | 2009 | Cyberonics, Inc | Some VNS Therapy System replacement Demipulse generators reporting low lead impedance readings. In rare instances, a system diagnostic test using Model 250 Programming Software (versions 7.1 and earlier) may report "Lead Impedance: OK" when a short-circuit condition exists. | Device Design | Cyberonics sent an Important Safety Notice on July 24, 2009, to all physicians explaining the event, risks presented and recommended actions to take. Labeling updates are being developed concurrently with the notification. The firm will mail or hand-deliver the notifications via firm representatives. |  | 9500 | 9,500 programming devices | Worldwide: USA, Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, Norway, Poland, Portugal, Puerto Rico (US Territory), Russia, San Marino, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | September 26, 2012 | 1160 | 9500
Z-0266-2010 | 53497 | No | CCT | Laryngotracheal topical anesthesia applicator. | Anesthesiology | Anesthesiology | 510(k) | Nasal Mucosal Atomizer Device with syringe, LATEX FREE For Single Patient Use Only Packaged Clean -Non-Sterile- MADE IN CANADA. |  |  | March 05, 2009 | 2009 | Sharn, Inc. | Sharn Anesthesia Inc., Tampa, FL received reports of leaking containers following use of their Sharn Anesthesia Microblast nasal and pediatric laryngeal atomizer device. | Device Design | Recalling firm is contacting all direct accounts (distributors &amp; user facilities) by telephone and letter. They request completion and return of the of the Return Goods Form and of any affected product by UPS using the firm's account number. |  | 550 | 550 | Nationwide Distribution | Terminated | July 13, 2012 | 1226 | 3225
Z-0269-2010 | 53496 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Norian CRS Rotary Mixer Cement 10cc sterile |  |  | September 14, 2009 | 2009 | Synthes USA (HQ), Inc. | Non-conforming material - increased set time. Potential clinical ramifications associated with increased set time and decreased strength include extrusion of unset material and fragmentation of cured material. The risk of Norian fragmentation post cure increases with a larger defect and a thicker layer of product.Customers were advised to call Synthes at 1-800-479-6329 to obtain a Return Autho | Nonconforming Material/Component | The firm issued an Urgent Medical Device Recall letter dated 9/17/09 via USPS certified mail. The letter informed customers of the problem and the need to return the product. Sales Consultants that were shipped the device were notified by teleconference on 9/14/09. Consultants were instructed to go to their accounts to remove the product from inventory. Customers were advised to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Questions are directed to Synthes at 800-620-7025 x 5453 or 610-719-5453 or to your Synthes CMF Sales Consultant. |  | 15 | 15 units | The product was shipped to medical facilities in CO, KS, MA, MO, NY, SC, TX, UT, and WA. | Terminated | March 08, 2010 | 175 | 15
Z-0263-2010 | 53549 | No | DZI | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Stryker Instruments 2.1 mm Cross-Cut Fissure Bur, carbide extended, sterile; REF 2296-101-621. |  |  | October 05, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The cross-cuts are missing on the device, which will result in poor cutting and overheating and/or smoking of the device. | Vendor change control | Consignees were notified by letter dated 10/5/09 and requested to return the devices to Stryker. The international account was notified by email. |  | 93 | 93 | USA (California, Georgia, Louisiana, Ohio, Wisconsin) and Brazil. | Terminated | November 30, 2009 | 56 | 93
Z-0270-2010 | 53563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with JPEG Compression. |  |  | September 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Distorted image. The "Distorted" image cannot be used for diagnostic findings, and, depending on the customer scenario, a repeated exam could be necessary if the image cannot be reported on any other workplace or recovered from another source. | Software design | The recalling firm issued a Customer Safety Advisory notice dated 9/14/09 to all affected customers via Update Instructions IM100/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. To avoid the potential risk of the issue, the local compression function can be disabled as described in the letter. Depending on the image data, the disabling of compression will result in increased storage usage.The issue will be resolved in the release of version VA70B.For all potentially affected customers with "JPEG Compression" license an update instruction will be released which defines the necessary corrective actions in case of "Distorted" images. The update instruction will be performed on request by the customer service group. Further questions regarding this issue are directed to the local SIEMENS Uptime Service Center or +49 (9131) 84-2056. |  | 28 | 28 units | The product was shipped to medical facilities in CA, FL, GA, IL, MD, MI, MO, NE, NC, NY, OH, PA, TN, TX, and WV. | Terminated | July 05, 2012 | 1010 | 58
Z-0286-2010 | 50135 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Respiratory Gating System, Catalog #NPTA540, Model #4535 674 26701, which includes: Pulmonary digital assembly kit and Pulmonary toolkit oncology Pro available on the following GEMINI PET/CT systems: GEMINI GXL 16 slice configuration, Model Number: 4535 679 75691. |  |  | November 24, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | A leak between the Tube Interface and the Outlet Tube of the transducer for the Respiratory Gating System component may result in a failure of the device to produce respiratory correlated images. | Process design | Philips Medical Systems Nuclear Medicine division initiated this recall by sending a 'Customer Advisory Notification to their customer accounts on 11/24/2008. This notification informs the customer of the potential for the device equipped with the Pulmonary Gating Option to develop an air leak which could result in the device being unable to produce the desired respiratory correlated images. The notification provides field test instructions and recommends that the customer test the bellows on each day of use in order to assess that the unit is functioning properly. If the bellows device fails the test the recalling firm recommends that the unit not be used and that the recalling firm be contacted via telephone or E-mail. The firm intends to conduct a field correction of the affected units in which their Field Service Engineers will install upgrade kits on the affected devices tentatively scheduled to begin in September, 2009. |  | 20 | 20 units | Worldwide distribution: USA, Puerto Rico, Canada, Belgium, China, Denmark, Egypt, France, Germany, Hong Kong, India, Japan, Netherlands, Pakistan, Philippines, Poland, Saudi Arabia, Singapore, South Korea, Spain, and Sweden. | Terminated | September 24, 2010 | 669 | 71
Z-0277-2010 | 52229 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical transport stretcher, Stryker Medical, Portage, MI; Model 734. |  |  | August 10, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The brake/steer pedal may seize, resulting in the brakes not working. | Nonconforming Material/Component | Consignees were notified by letter dated August 10, 2009 that Stryker will contact them to upgrade the stretchers within 6 months. The recall was expanded in scope to include additional serial numbers by a second letter dated September 17, 2009, which was issued to additional consignees who had not been previously notified of the recall, as well as the consignees who had been issued the August 10, 2009 letter. |  | 241 | 241 | Nationwide, Argentina, Australia, Brazil, Canada, Chile, China, France, Germany, Greece, India, Italy, Japan, Korea, Latin America, Malaysia, Mexico, Netherlands, Philippines, Poland, Portugal, Singapore, South Africa, Spain, Switzerland, Taiwan, and United Kingdom, | Terminated | October 29, 2010 | 445 | 17807
Z-0332-2010 | 53605 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 32CM ALTA LR GOLD STD KIT |  |  | October 23, 2009 | 2009 | Spire Biomedical, Inc. | There is a possibility that the catheter tubing could separate from the hub/bifurcation of the device. If this occurs, bleeding is the primary clinical complication likely to result from catheter tubing/hub separation. A second potential clinical complication is air embolism, particularly for the segment of the patient population that develops negative central venous pressure during normal respira | Process control | Spire Medical issued an "Urgent device Recall" letter dated 10/22/09 by FEDEX to end users. Users are requested to return product. For questions, contact Spire Biomedical Customer Service at 1-877-876-8635. |  | 666 | 666 units | Worldwide distribution - Nationwide, Government. Foreign: Canada, France. | Terminated | June 09, 2010 | 229 | 55183
Z-0273-2010 | 53615 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Magnum Disposable Biopsy Core Needle 16 gauge, Catalog Number: MN1610. |  |  | October 08, 2009 | 2009 | Bard Peripheral Vascular Inc | The recall was initiated after the firm determined that some Magnum Disposable Biopsy Core Needle 16 gauge lots may exhibit an exposed sample chamber, due to extended stylet length resulting from a loose hub connection. The exposed sample chamber will typically result in an inability to obtain the biopsy sample. In the majority of reported cases, there was no injury to the patient or user. When us | Process design | The recall notification was initiated by the firm issuing an urgent medical device notification to it's consignees on October 08, 2009 via USPS requesting them to check all inventory locations for the affected products, remove and return affected products to Bard. The consignees were instructed to complete and return a recall effectiveness check form. Bard provided this 800 number 1-800-321-4254 for any questions consignees may have. |  | 127169 | 127,169 units | Nationwide and internationally to CanadaMexicoAlgeria Andorra Austria Bahrain Belarus Belgium Croatia Cyprus Czech Republic Denmark Finland France French Guiana Germany Gibraltar Greece Hungary Ireland Israel Italy Ivory Coast Kenya Kuwait Latvia Lebanon Luxembourg Malta Martinique Mauritius Morocco Netherlands New Caledonia Nigeria Norway Oman Poland Portugal Republic of Kosovo Republic of Macedonia Reunion Romania Russia Saudi Arabia Senegal Serbia Seychelles Slovenia South Africa Spain Sweden Switzerland Syria Tunisia Turkey Ukraine United Arab Emirates United Kingdom AustraliaNew ZealandCanadaJapanBangladeshChinaHong KongIndiaIndonesiaKoreaMalaysiaPakistanSingaporeSri LankaTaiwanThailandVietnamArgentinaBahamasBarbadosBrazilChileColombiaDominican RepublicMexicoPeru andSuriname | Terminated | April 02, 2012 | 907 | 334701
Z-0272-2010 | 53586 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Richard Wolf Hulka Clip; for single patient use only; a sterile tubal occluding device; 2 clips per procedure tray/pouch, 10 pouches per shipping carton; Richard Wolf Medical Instruments Corp., Vernon Hills, IL 60061; item 4986.90. |  |  | October 16, 2009 | 2009 | Richard Wolf Medical Instruments Corp. | The sterile pouch seal may fail to remain intact during handling rendering the Hulka clip non-sterile. | Packaging | Richard Wolfe sent recall letters dated October 16, 2009 to the direct accounts via certified mail on the same date. The accounts were informed that the lot of Hulka clips did not meet their sterile seal burst test requirements, and that the sterile seal may fail to remain intact during handling, resulting in the clips becoming contaminated or non-sterile. They were instructed to inspect their inventory for the affected lot, quarantine any product found and call Richard Wolf at 800-323-9653, ext 225 or 380 to arrange for the return of the affected product for replacement. The accounts were also requested to complete and return the enclosed response form to Richard Wolf, even if they do not have any of the affected lot in stock. |  | 150 | 150 cartons | Nationwide | Terminated | March 16, 2011 | 516 | 150
Z-0338-2010 | 51389 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 100 Low Bed, Model P3930. AC-powered adjustable hospital bed. |  |  | April 20, 2009 | 2009 | Hill-Rom, Inc. | The siderails may exhibit false latching, or unintentional lowering, or may become inoperable. | Process control | Hill Rom notified consignees by recall letter dated 4/20/09 and advised them of the problem and requested them to locate the beds. Consignees with patients in need of siderails were directed to be moved to a different model bed with siderails and, if the patient did not need siderails; consignees were requested to remove the siderails, move the siderails to a stowed position, to tie the siderail to the bed frame or to tag the siderail as out of service. The letter stated the consignee would be contacted within 90 days by a firm representative to do a correction on all affected units. |  | 4017 | 4017 beds | Worldwide distribution: USA, Canada, France, India, Singapore and Spain. | Terminated | October 15, 2010 | 543 | 4017
Z-0365-2010 | 51876 | No | NSX |  |  | General Hospital | Contact ODE | PICIS, CareSuite Family- OR Manager (4235), Preop Manager (4250), Anesthesia Manager (4220), PACU Manager (4245) &amp; Critical Care Manager (1110), patient information systems. |  |  | April 08, 2009 | 2009 | Picis Inc. | Demographic data, most notably allergy and precaution data, can be overwritten with incomplete data or blanks by the interface between non-Picis clinical systems (your "HIS" or "CIS") and certain Picis applications, if that interface is not configured properly. | Software design | On April 8th 2009, Picis initiated a customer notification letter to the installed base of potentially affected users, explaining the potential source of patient data conflict and to provide support in choosing the appropriate configuration option and implementation. The notification letters included with this report package contain additional details regarding configuration options. |  | 325 | 325 | Worldwide distribution: USA, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and United Kingdom | Terminated | December 22, 2014 | 2084 | 325
Z-0366-2010 | 53604 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AGFA IMPAX HeartStation ECG Management System. |  |  | May 07, 2009 | 2009 | AGFA Corp. | HeartStation does not refresh content changes to the PDF reports that are printed emailed or faxed from the Index page. This may lead to a discrepancy in patient demographics between what is displayed in the HeartStation Client interface and the report delivered to the consumer. Questions are being directed to the local Agfa representative; the contact information was included in the letter. | Other | An Urgent Safety Notice was sent via FED-EX to all the sites with the affected units (a total of 6) on May 7, 2009. The three foreign accounts were notified by letter on May 14, 2009. The letter describes the potential issue and mitigation. Acknowledgment via FAX-Back that the information was received and understood was requested from the sites. Customers who use the print, email, fax functionality of this product were verbally notified of the defect and they have been advised of a workaround. In the notice letter, Agfa HealthCare is also providing customers with the recommended workaround. The workaround is to only print from the Viewer screen or to print the ECG once confirmed. The Viewer screen, however, does not allow the user to print batches of reports as does the Index screen. Questions about this matter are being directed to the local Agfa HealthCare representative. |  | 9 | 9 units | CA, KY, MA, NY, OH, SC, Belgium, Italy and Singapore | Terminated | July 27, 2010 | 446 | 9
Z-0370-2010 | 51933 | No | NQP | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Zimmer Dynesys top-loading cutter, spacer, Zimmer Spine, Minneapolis, MN; REF 07.01284.001. |  |  | May 28, 2009 | 2009 | Zimmer Inc. | The spacer cutter may jam during surgery, preventing use of the instrument and causing surgical delay. | Other | Consignees were notified by letter dated 5/20/09 titled "Urgent Medical Device Correction" that instructed customers that they may continue to use the recalled product until adequate replacements become available, but that they should always ensure they have a backup spacer cutter (of a different model which does not have this issue) available in case this model should jam during use. Users were also advised to ensure the spacer cutter is properly lubricated and is sharp. For questions, please contact Zimmer Spine at 1-800-777-7505. |  | 177 | 177 | Nationwide, Austria, Belgium, Canada, Italy, Netherlands, Singapore, Spain, Switzerland and United Kingdom. | Terminated | October 05, 2011 | 860 | 177
Z-0369-2010 | 50784 | No | LPB |  |  | Cardiovascular | PMA | Biosense Webster NaviStar RMT Diagnostic/Ablation Steerable Tip 8mm DS Catheters; Part Numbers: D-1259-01 &amp; D-1259-02 |  |  | October 02, 2008 | 2008 | BioSense Webster, Inc. | Biosense Webster initiated the recall because of the potential concern with one of the Printed Circuit Board (PCB) components used in the NaviStar RMT Catheters, which may cause the ablation catheter icon to shift on the map viewer when RF energy is being applied. Therefore the user, attempting a linear ablation in the location displayed on the CARTO RMT screen, may ultimately apply RF energy to u | Process control | The firm sent an "Urgent Medical Device Correction" letter to consignees on October 6, 2008 informing consignees. Users were advised to pay "close attention to the location of the ablation catheter icon on the map viewer when applying RF energy."If consignees suspect that they have a defective NAVISTAR RMT Catheter, they were instructed to notify Biosense Webster and to return the defective catheter for evaluation and refund/replacement.Customer were instructed to contact their local Biosense Webster representative for any questions or further information. No return reply was requested. |  | 2180 | 2180 Units | CA, CT, FL, GA, IA, IL, IN, KS, KY, MA, MI, MN, MO, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI | Terminated | November 23, 2009 | 417 | 4739
Z-0367-2010 | 53580 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biomet K-Wire/Trocar Point, both ends/Dia. 045, length: 9 in., certified stainless steel, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 35-351534. |  |  | August 06, 2009 | 2009 | Biomet, Inc. | Packages labeled as having a .045 inch diameter trocar point actually contain a diamond point with a diameter of .062 inches. | Process control | An "Urgent Medical Device Recall Notice" dated August 6, 2009 was issued to consignees. The letter described the affected product, issue and actions for consignees. Customers were instructed to complete and return the "FAX back Response Form" (fax number 1-574-372-1683) and all affected product in inventory to Biomet Orthopedics.Direct questions related to the recall notification by calling at 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8AM to 5PM. |  | 19 | 19 | United States (Florida and Pennsylvania). | Terminated | October 06, 2010 | 426 | 19
Z-0389-2010 | 52615 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxinium Femoral Head, 36 MM OD-3, 12/14 Taper, REF 71343603, Sterile, Smith &amp; Nephew, Memphis TN 38116 |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | Radius at the end of the taper of the femoral implant is smaller than specification. This error could result in the tapers not locking fully or properly on femoral stems or in causing impingement and/or increased wear debris. The deficiency may not be noticed by the user. Impingement and/or wear debris may lead to revision. | Other | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on 01/23/2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. Further questions are directed to the contacts at Smith and Nephew listed in the letter. |  | 925 | 925 units | Nationwide; 17 international: Switzerland, Dubai, Brazil, Canada, South Africa, Australia, Belgium, Spain, Thailand Turkey, Costa Rica, Denmark, India, Italy, and France. | Terminated | February 07, 2012 | 1110 | 925
Z-0371-2010 | 53772 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Unicondylar Knee Femoral Component SC2462-1. |  |  | May 05, 2009 | 2009 | Stelkast Co | incorrect size markings on component | Nonconforming Material/Component | The recalling firm telephoned the medical supply companies on 5/5-19/09 to inform them of the problem and the need to return the products. If there are questions regarding this issue, please call 1.888.273.1583. |  | 20 | 20 units | Nationwide | Terminated | November 20, 2009 | 199 | 20
Z-0107-2010 | 53579 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Guider XF Softip Guide Catheter, manufactured by Boston Scientific Corporation, Maple Grove, MN, distributed by Boston Scientific Corporation, Fremont, CA. |  |  | October 06, 2009 | 2009 | Boston Scientific Corporation | Product degradation-- a polymer portion of the distal section of the catheter may degrade prematurely, leading to potential systemic embolization of particulates with ischemic complications. | Environmental control | The firm initiated its 'Safety Alert' on October 6, 2009. A notification letter was sent to each customer reminding them to follow storage and handling instructions. The firm has not ceased distribution. Direct questions about this notification to the Boston Scientific Corporation at 1-510-624-1734. |  | 94296 | 94,296 units | Worldwide Distribution -- United States, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macau, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, Tobago, Tunisia, Turkey, Ukraine, UAE, Uruguay, Venezuela, Vietnam and Yemen. | Terminated | December 19, 2011 | 804 | 94296
Z-0122-2010 | 53614 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Instruments Navigation System II- CART, Stryker Navigation, Kalamazoo, MI; Stryker REF 7700-100-000. |  |  | October 26, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Multiple software-related malfunctions may occur. The device may suddenly stop working, the screen may freeze, the screen may only be updated very slowly, the unit may not respond to user input, or the unit may only respond to user very slowly. | Software design | The firm issued a press release on 11/4/09. Consignees were notified by an Urgent: Medical Device Recall Notification letter dated 10/23/09, which advised of the affected products, the reason for the recall, the recall plan, and the actions to be taken. Customers were to immediately review this notification and distribute to all departments affected. Also, customers were to immediately quarantine any affected product and label as recalled product; complete and sign the enclosed Business Reply Form and fax a copy to 866-521-2762. Upon receipt of the form, a Stryker Navigation Field Service Rep would contact them to schedule the repair of their equipment. All U.S. units were corrected by 10/30/09. |  | 21 | 21 | Worldwide Distribution -- USA, including states of Florida, Kansas, Kentucky, Massachusetts, Michigan, New Jersey, Pennsylvania, South Carolina, Texas, and Washington and countries of Australia, China, Germany, Japan, Korea, Switzerland, and the United Kingdom. | Terminated | December 31, 2009 | 66 | 23
Z-0413-2010 | 53619 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Merit Medical Systems, Inc., Convenience Trays, custom packaged under the following Merit MedicalÂż labels: (1) Prentice PICC Kit, Catalog number K12T-03698; (2) Insertion Kit, catalog number K12T-03050C; and (3) Angio Procedure Pack, catalog number K09T-09530A. |  |  | May 19, 2009 | 2009 | Merit Medical Systems, Inc. | Surgical gloves in hospital convenience kits may be degraded and tear during use. | Nonconforming Material/Component | Merit Medical Systems notified customers by phone, fax, or visit by sales representatives on 05/19/09. The notification advised of the convenience kit correction due to the recall of the Derma Prene Ultra Surgical Gloves, and that use of the product should cease immediately. Sales representatives were instructed to provide replacement gloves and "warning" stickers to be placed on identified kits in inventory. The warning stickers advise the customer to discard the surgical gloves in the kit and replace. Customers were requested to verify product disposition by signing a response form. |  | 378 | 378 units | USA distribution: nine hospitals in Texas and Colorado. | Terminated | December 03, 2009 | 198 | 378
Z-0414-2010 | 53475 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS Harmony LL 500/700 Surgical Light, Vertical Spring-Loaded Arm. |  |  | September 11, 2009 | 2009 | Steris Corp | DrÂżger Medical AG &amp; Co., KGaA informed STERIS Corporation that they received (4) reported international cases where a welded seam at the joint of the spring-loaded arm of the DrÂżger SOLA 500/700 surgical light broke as a result of a crack that had slowly developed over time. The STERIS Harmony LL 500/700 surgical light is a Private Label version of the DrÂżger SOLA 500/700 surgical light. | Device Design | The firm sent Urgent Product Correction letters, dated Sept. 11, 2009, to their customers. The letter identified the problem and stated that Drager had supplied STERIS the replacement spring arms and that a STERIS Service Technician would perform the replacement. The letter also stated that the correction would be performed at no charge to the customers. According to the letter, customers would be contacted by their STERIS Service Representative to arrange a convenient time for replacement of potentially affected spring arms. In the meantime, the letter states that customers should use their surgical lights in accordance with the guidelines presented in the Operator's Manual. Questions or concerns regarding the letter should be directed to Barbara Rich at 440-392-7318 or your local STERIS account manager. |  | 388 | 388 units | Worldwide Distribution -- USA, including the states of AL, AR, CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, WA, WI &amp; WY and the country of Canada. | Terminated | March 01, 2012 | 902 | 388
Z-0374-2010 | 53688 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SureSigns Vital Signs Viewer (VSV):Spot Check Vital Signs Monitor with NBP (options to include Predictive Temperature only; SPo2 only; SPo2 and Recorder only; SPo2 and Predictive Temperature only; or SPo2, Predictive Temperature, and Recorder |  |  | October 30, 2009 | 2009 | Philips Healthcare Inc. | Use of these devices at their maximum volume setting may result in the premature failure of the internal speaker | Software design | Customer notifications dated October 2009, in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be instructed to lower the volume setting while using existing units during the first notification. As part of the second notification to the customers, the field action correction will be a customer installable software upgrade. Customers may contact Philips Response Center with questions or concerns at 1.800.722.9377.On December 14, the firm issued the follow-up notification letter implementing the action plan to provide the software upgrade kit for installation and instructions. |  | 221 | 221 units | Nationwide Foreign: CanadaAfghanistan, Albania, Algeria, Angola, Argentina, Austria, Australia, Bahrain, Bosnia &amp; Herzegowina, Bangladesh, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Switzerland, Chile, China, Colombia, Costa Rica, Croatia/Hrvatska, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Gabon, Ghana, Great Britain, French Guiana, Guadeloupe, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Iceland, Italy, Iraq, Jordan, Japan, Jamaica Republic of Korea, Kenya, Kosovo, Kuwait, Lebanon, Sri Lanka, Lithuania, Luxembourg, Latvia, Libyan Arab Jamahiriya, Morocco, Norway, New Zealand, Republic of Moldova, Macau, Malawi, Mali, Malta, Martinique, Mauritius, Mexico, Malaysia, Namibia, Nepal, Netherlands,Nicaragua, Nigeria, Norway, Philippines, Pakistan, Oman, Panama, Peru, Poland, Puerto Rico, Palestinian Territories, Portugal, Qatar, Romania, Republic of Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Spain, Sri Lanka, Switzerland, El Salvador, Tanzania, Thailand, Tunisia, Turkey, Taiwan, Uganda, Venezuela, Vietnam, South Africa, and Zimbabwe United Arab Emirates | Terminated | February 20, 2013 | 1209 | 49648
Z-0397-2010 | 53694 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FlexShaft, product code FS14; FlexShaft II, product code FS214; FlexShaft III, product code FS312; FlexShaft III Sterilization Connector Guard, product code FS3SCG; FlexShaft II Reconditioned, product code FS214RC; Manual Release for FlexShaft II &amp; III, product code MR14. |  |  | September 23, 2009 | 2009 | Surgical Devices/Covidien | Lack of assurance that products were manufactured under GMPs. This action is being undertaken due to certain products potentially exhibiting compromised staple formation or cutting which in some instances may result in extended surgical time for the patient. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 9/22/09 to all customers. The letter informed customers that Covidien acquired PMI, about the problem with the products, and about the need to return any products remaining inventory. Customers were asked to return the entire inventory of Power Medical Interventions product, regardless of its expiration date. This process can be initiated by contacting Power Medical Innovations customer service at 1-866-769-3763. |  | 1100 | 1100 units | The products were shipped medical facilities nationwide. The product was also shipped to distributors in Canada, Japan, Germany, Ireland, Italy, Spain, Taiwan, Belgium, Sweden, Netherlands, Israel, Portugal, India, Austria, Turkey, and Greece. | Terminated | March 08, 2010 | 166 | 240432
Z-0391-2010 | 53434 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Howmedica Osteonics, Skirtless 5 40' LFIT V40 Femoral Head (32 mm +0 mm offset); hip prosthesis component. |  |  | September 04, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Packaging mix up between the Taper V40 Femoral Head and the LFIT V40 Femoral Head. | Process control | Stryker sent Notification letters and Product Accountability Forms via FedEx on September 4, 2009 with return receipt. Corrected letters were sent on September 9, 2009. |  | 10 | 10 units | Nationwide distribution. | Terminated | June 26, 2013 | 1391 | 17
Z-0383-2010 | 52806 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail System, Antegrade Femoral Recon Module, standard, nonsterile, Zimmer, Warsaw, IN; REF 00-2490-001-08. |  |  | July 01, 2009 | 2009 | Zimmer Inc. | The guides may come apart during use, some guides may experience surface cracking, guide handles and modules may not target the femoral nail holes accurately and the instruments may not be able to be assemble together due to a change in the threading method used. | Other | On July 1, 2009, U.S. distributors were notified by e-mail to return two instruments. The recall was expanded with international accounts notified by e-mail on August 5, 2009, and U.S. accounts being notified by e-mail on August 6, 2009, to return 7 instruments. The scope of the recall was then expanded again to include 14 instruments and another letter dated September 14, 2009 to distributors and user accounts.The notifications described the affected products, problem, health risks and actions for consignees. Direct questions about the recall to Zimmer, Inc. by calling 1-800-613-6131. |  | 327 | 327 of total product | Worldwide Distribution -- United States, France, Germany, Italy, Switzerland and United Kingdom. | Terminated | October 05, 2011 | 826 | 327
Z-0415-2010 | 53299 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | AlMER ENDOFEMORAL 4MM OFFSET (PURPLE) |  |  | September 03, 2009 | 2009 | Smith &amp; Nephew, Inc., Endoscopy Div. | Device may have improper Laser Marking, 4mm device may be marked as 6mm and 6mm device may be marked as 4mm. | Labeling design | A notification letter was issued via Fedex on September 3, 2009 to domestic consignees and international consignees were notified via e-mail. The letter described the issue as the affected device may have improper laser marking. The 4mm devices are marked as 6mm and the 6mm devices are marked at 4mm. Customers were instructed to return the completed response form and affected product in inventory to Smith &amp; Nephew. Direct questions about the recall to your local Smith &amp; Nephew Sales Representatives or call 1-508-261-3655. |  | 134 | 134 individual units | Worldwide Distribution -- United States including Puerto Rico, Australia, Austria, Belgium, Canada, Costa Rica, Dubai, Finland, France, Germany, Great Britain, Netherlands, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | May 08, 2012 | 978 | 134
Z-0455-2010 | 52886 | No | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier, LMP-50 (Lesser Metatarsal Phalangeal Implant), Size 50, Sterile EO, , Manufactured by Tornier, Inc, Edina, MN 55435. |  |  | April 06, 2009 | 2009 | Tornier, Inc. | Tornier, Inc. has received a report of a mislabeled product due to a product mix-up. The package label indicated that the product was a model LMP-50 however the product inside the package was Tornier FGT-50 implant (Primus Flexible Great Toe Implant With Grommetts). | Labeling mix-ups | Consignees were sent on April 6, 2009 a Tornier "Urgent-Product Recall" letter dated April 3, 2009. The letter was addressed to Healthcare professionals and distributors. The letter described the problem and product involved. Advised consignees to immediately discontinue use of the affected product and to return any unused affected product to their facility. Advised their distributors to notify all accounts. They also attached a questionnaire needed for inventory status. |  | 77 | 77 | Worldwide Distribution -- USA, including states of CA, FL, MD, MI, NJ, OH, and PA, and country of France. | Terminated | December 17, 2011 | 985 | 77
Z-0456-2010 | 53462 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) RA1000 Workstation software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There are potential safety issues, including misdiagnosis or delay in treatment, associated with the use of GE Centricity PACS RA1000 Workstation software that may impact patient safety. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated September 23, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers, to the attention of the Hospital Risk Manager, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator, on 9/25/09. The letters advised the users of the patient safety issues associated with the use of GE Centricity PACS related to unintentionally adding new images to an approved exam; deletion of images from PACS if Cancel Checked or Cancel Selected is chosen in the CA Tool - Archive Queue; inadvertently getting two or more exams out of sync with the corresponding reports resulting in the RIS-IC reports being associated with the wrong patients exam; and when rejecting a duplicate image(s) from an unspecified exam that is still in the Arrived or Verified status, the duplicate image(s) is also rejected from the Dictated, Transcribed, or Completed status exam. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issues and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution -- United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 973 | 1188
Z-0120-2010 | 53672 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Reinforced Catheter Introducer System, 5 per carton. The product is sold under the Cordis Crossover and Cardiovascular Systems Inc (CSI) ViperSheath labels. |  |  | October 21, 2009 | 2009 | Thomas Medical Products Inc | Sheath may break or fracture which may cause vessel damage or require surgical intervention to retrieve catheter pieces. | Component design/selection | An "Urgent Medical Device Correction" letter dated October 22, 2009 was issued to the direct accounts. The letter described the problem and actions for consignees. Consignees were instructed to cease distribution of the recalled product and to notify their subaccounts of the problem. In addition, customers were instructed to return the affected, unused product to Thomas Medical Product, Inc.'s contract warehouse, Inner-Space Warehouse, Inc.Contact your local sales representative for questions about the recall, or Customer Service by calling 1-866-446-3003, Monday through Friday from 8:30AM to 5:OOPM Eastern Standard Time, for questions about product return. |  | 17236 | 17,236 units | United States (FL and MN). | Terminated | November 08, 2010 | 383 | 17236
Z-0420-2010 | 53681 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | The Natus Blue Light Phototherapy is also called the "neoBLUE LED Phototherapy System". The "neoBLUE mini LED Phototherapy System" |  |  | October 01, 2009 | 2009 | Natus Medical Incorporated | The clamp connecting the neoBLUE mini phototherapy light to the accessory pole stand can sustain stress fractures or break when excessive force is used when tightening the clamp. No injuries have been reported. | Device Design | On 10/1/09 Natus notified domestic consignees by Urgent Product Advisory Notice letter that advised them of the issue, instructed them to inspect the clamps, and included pictures of the affected clamp. Consignees are instructed to return the recalled clamps for a new replacement to be shipped by Natus.International distribution partners were notified by letter on 10/1/2009 and instructed to identify and contact their consignees. Replacements will be sent to the distribution partner for their consignees. |  | 2127 | 2127 | Worldwide distribution: USA, Argentina, Australia, Austria, Belgium, Bermuda, Canada, Chile, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Kenya, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Philippines, Poland, Portugal, PR, Qatar, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, and UK | Terminated | September 22, 2011 | 721 | 2127
Z-0422-2010 | 53561 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate 12 volt sealed lead acid Battery Clips that are components of the HeartMate II LVAS (PMA P060040) and the HeartMate XVE LVAS (PMA P920014) Systems. |  |  | October 14, 2009 | 2009 | Thoratec Corp | Patient reports alerted Thoratec that during the process of disconnecting power leads from the battery clips, the threaded connector of the battery clip housing was observed to be loose, and in some cases dislodged from the battery clip housing. The problem affects HeartMate 12V SLA Battery Clips including those used for the HeartMate II LVAS and HeartMate XVE systems. If not detected and handled | Other | Firm has issued a notification letter to consignees, dated October 14, 2009 asking them to return proof of receipt, and to remove all affected unused devices from stock, and asking ongoing patients to return to hospitals for inspection, replacement if needed, and retraining on the process of switching power sources and safely switching to a back-up system controller. Contact Thoratec Customer Services at 1-800-456-1477 for any questions or concerns about this issue. |  | 8194 | 8,194 battery clips | Worldwide - USA- AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and the District Columbia. Internationally to Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, France, Germany, Greece, Ireland, Israel, Italy, Lebanon, Lithuania, Luxembourg, Monaco, Netherlands, Singapore, South Africa, Sweden, Switzerland, Taiwan and the UK. | Terminated | April 18, 2011 | 551 | 8194
Z-0418-2010 | 53593 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio (ICIP) Critical Care Release D, Version D.02. |  |  | October 21, 2009 | 2009 | Philips Healthcare Inc. | A software defect exists in the ICIP Release 0.02 scheduling function for order management such that if a user edits either the Schedule Time or the Stop Time for an existing order, the corresponding intervention(s) for the order may not be created. | Software design | Philips Medical Systems notified accounts by letter on October 21, 2009, of the release ICIP version D.03, which contains a software correction . A Philips representative will schedule the implementation of this upgrade. Please contact the Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377. |  | 53 | 53 units | MD, RI, WV, and TXAustraliaBelgiumDenmarkFranceItalyMalaysiaSingaporeSpainSwitzerlandTaiwanUnited Arab EmiratesUnited KingdomandIreland | Completed |  |  | 53
Z-0443-2010 | 53606 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Thor Anterior Plating System; |  |  | October 02, 2009 | 2009 | Stryker Spine | Surgeons may apply an accidental axial torque to and/or over-angulate the All-In One Guide while implanting the screw. This may cause the screw not to seat properly in the locking ring and may cause the Thor screwdrivers to twist or break at the tip. | Device Design | Recall notification letters were sent to all Stryker branches, Hospital Risk Management and Surgeons on October 19, 2009 by Federal Express. Questions are to be directed to Tiffani Rogers, Regulatory Compliance Manager at 201-760-8206. |  | 262 | 262 | Nationwide | Terminated | January 23, 2013 | 1209 | 9550
Z-0421-2010 | 52445 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Kuschall K Junior Manual wheelchair |  |  | June 15, 2009 | 2009 | Invacare Corporation | The handle has been breaking at a calculated weight of 65 kg which may have serious consequences (broken bones) depending on the circumstances of the break (e.g., user on ramps, hills, etc.). | Device Design | The firm contacted their customers by telephone and then sent a recall notification letter to each customer on 6/15/2009. Questions are directed to email: Recalls@Invacare.com or phone: 800-333-6900. |  | 20 | 20 Wheelchairs | Product was distributed to the following states: CA, CO, MN, NJ, OR, PA, TX &amp; VA and one to Canada. | Terminated | November 21, 2011 | 889 | 20
Z-0457-2010 | 52720 | No | IZZ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | GE Medical Systems, COMPAX 40E X-ray Systems, Models; 2208203, 2208205, 2244282, 2244284, 2225863, 2225864 , 2225865, 2225870, 2225881, and 2225882. |  |  | July 30, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware that table lock safety features, which are provided with Compax40E X-ray systems, are not always being utilized and has resulted in unintended lateral tabletop movementthat may impact patient safety. | Other | Consignees in US and Canada were sent on July 31, 2009 a GE Healthcare "Urgent Medical device Correction" letter dated July 30, 2009. The letter was addressed to Hospital Administrators/ Risk mangers; Radiology Department Mangers; Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product correction, and Contact Information. The letter instructs customers that the table manual mechanical lock is provided with the system and should be engaged when loading or unloading patients to ensure the tabletop is stabilized and not moving. For the rest of the world, the customer letter was dated October 22, 2009 and was mailed October 30, 2009. |  | 1046 | 1046 | Worldwide Distribution -- USA, including states of AL, AK, AZ, AR, CA, CO, CT, DC,FL, GA, HI, ID, IL , IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH,NJ, NM, NY, NC, OH, OK OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of UNITED KINGDOM, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SAUDI ARABIA, PUERTO RICO, PORTUGAL NEW CALEDONIA, MARTINIQUE, KOREA, JAPAN, JAMAICA, ITALY, ISRAEL, HONG KONG, HONDURAS, GREECE, GERMANY, FRANCE, EGYPT, DOMINICAN, CHINA, CANDA, BRAZIL AUSTRIA, and AUSTRALIA. | Open, Classified |  |  | 1046
Z-0419-2010 | 53261 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips Lifeline Personal Help Button Pendant Classic and Slimline Style. |  |  | September 11, 2009 | 2009 | Philips Lifeline Inc | Pendant Personal Help Button neck cord may not break away and if not will present a potential choking risk. | Other | Philips Lifeline initiated Safety Notification letters on 9/21/09 to users and advised of the potential choking risk of the Pendant Personal Help Button The firm recommends that users consult their health care providers to determine which style of emergency button, including those that are worn on the wrist, is most beneficial for them. A postage-paid reply envelope had been provided. Questions should be addressed to Philips at 1-877-221-8756 (Monday  Friday, 8 a.m.  8 p.m. EST). Firm issued Press on 9/11/09. FDA Press issued 9/22/09 |  | 715000 | 715,000 | NationwideCanada | Terminated | October 02, 2012 | 1117 | 715000
Z-0458-2010 | 53801 | No | MUD | Oximeter. | Cardiovascular | Cardiovascular | 510(k) | Nonin Model 7600B cerebral oximeter pod. Used with the Nonin Model 7600 Regional Oximeter with Equanox Technology and Bluetooth Wireless Technology. Nonin Medical, Inc., 13700 First Avenue North, Plymouth, Minnesota 55441-5443 USA. |  |  | October 23, 2009 | 2009 | Nonin Medical, Inc | A malfunction was detected with one of the components that, under rare circumstances, may result in the overheating of the POD. This overheating, while unlikely to occur, has the potential to cause a burn if the POD is in direct contact with the patient. | Device Design | An Urgent: Medical Device Recall letter was emailed to customers on 10/23/2009. The letter described the issue and hazard and asked that the use of the device be discontinued and to immediately contact Nonin for return and repair/replacement instructions. Questions or concerns should be directed to John Dalpee at 763-577-3166. |  | 62 | 62 | Worldwide Distribution -- USA, including states of LA, MN, NC, and NY and countries of Canada, Germany, Italy, Netherlands, and United Kingdom. | Terminated | March 29, 2012 | 888 | 62
Z-0479-2010 | 52121 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cargille Modified Kingsbury-Clark Albumin Standards For in vitro Diagnostics (Set of 8 Standards: 5, 10, 20, 30, 40, 50, 75, 100 mg/ml); Catalog number: 52064; |  |  | July 07, 2006 | 2006 | Cargille-Sacher Laboratories Inc. | The 10 ml size standard of Cargille modified Kingsbury Clark Albumin Standards of lot Jan-12-05 appear shrunken or liquefied. | Storage | Cargill sent recall letters mailed July 7, 2006 by first class mail. |  | 5 | 5 kits | The product was distributed to 5 customers in OH, Washington, DC, PA, WI, and VA. | Terminated | December 02, 2009 | 1244 | 5
Z-0459-2010 | 53161 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Bath Safe Adjustable Shower Seat, Without Back and Arms, Catalog number 727142001 |  |  | July 09, 2009 | 2009 | Nutech Manufacturing Corp | Safety chair used in shower/bath may not meet specifications for weight load. | Device Design | The subsidiary distributor, Maddak, Inc. notified customers by A "Voluntary Product Recall Notice", phone and e-mail on 07/09/2009. The own label distributor instructed customers to return their inventory of affected products received between May 1, 2009 and July 7, 2009. Distributors were instructed to immediately quarantine then return all affected inventory to Maddak and notify their accounts. Distributors were requested to instruct dealers to display a "storefront" notice and provide consumers with a customer response form. Distributors were also requested to forward a request to dealers to post a website notice if they have one. Recovered products will be replaced. |  | 3627 | 3,627 total units | Worldwide Distribution -- USA and Canada. | Terminated | May 29, 2013 | 1420 | 3627
Z-0467-2010 | 53588 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio CR Waffle Femur Lfit with posts; |  |  | October 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | It was determined that a change made in the dimensions of the collar component of the packaging resulted in the inner blister material being overstretched causing a potential hole in the sterile barrier system. | Other | An "Urgent Product Recall" letter was sent on October 2, 2009 by Federal Express to Stryker branches and Hospital Risk Management Departments in the United States as well as foreign consignees. The letter described the issue, potential hazard and actions for affected accounts. Consignees were instructed to examine their inventory and return recalled product as instructed in the notification letter to Stryker Orthopaedics. In addition, consignees were instructed to reconcile affected product by completing the attached Product Accountability form (with spreadsheet) and faxing to 1-201-831-6069. Direct questions to Stryker Orthopaedics by calling 1-201-972-2100. |  | 77 | 77 of total product | Worldwide Distribution -- United States (MA, NJ, IN, MI, CA and PA), Hong Kong, Germany, Netherlands, Japan, Korea, Switzerland, Malaysia, and the United Kingdom. | Terminated | April 05, 2010 | 185 | 77
Z-0478-2010 | 53867 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ProtÂżgÂż EverFlex" Self-Expanding Biliary Stent System (for United States distribution). 6 x 150mm, Model # PRB35-06-150-080. Sterile EO, ev3 Inc, 4600 Nathan lane North, Plymouth, MN 55442-2920, Rx only. |  |  | November 06, 2009 | 2009 | Ev3, Inc. | ev3 has determined that a lot of 150mm ProtÂżgÂż EverFlex Biliary Stent System contains a 100mm length Self-Expanding stent. The implantation of a stent 50mm shorter than expected may not cover the full length of the target lesion and may require subsequent intervention. | Labeling mix-ups | Two separate letters were sent to United States and international consignees on November 9, 2009. The letter to US consignees was addressed to Cath Lab Managers while the letter to foreign consignees was addressed to the Risk Manager or Cath Lab Manager. The letters described the affected product, gave an Issue Summary and Required Action, which advised consignees to locate and remove the recalled product. An Ev3, Inc. sales representative will contact consignees to arrange return of all unused product to Ev3, Inc. Direct questions regarding the recall to Ev3, Inc. by calling Customer Service at 1-800-716-6700. |  | 6 | 6 (United States) and 6 (Outside United States) | Worldwide Distribution -- United States (TX and FL), GERMANY, FRANCE and UNITED KINGDOM. | Terminated | March 08, 2012 | 853 | 6
Z-0480-2010 | 53612 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Accessories, Connection Bracket. Catalogue # 1180.36A0. |  |  | October 22, 2009 | 2009 | Maquet Inc. | Product monitoring has determined that the connection bracket shows potential mechanical play at the OR table interface for some procedures. This is due to the hole in the lock washer, which is responsible for locking the accessory to the OR table, being too large. | Device Design | Urgent Device Correction Letters were sent by Federal Express to affected US hospitals on October 22, 2009. The letter stated the technical description, potential hazard, and the recommended precaution. The letter asked customers to immediately verify whether their Connection Bracket is affected thru the FaxBack provided. Questions or need for additional information can be directed to customers' local MAQUET representative. |  | 9 | 9 units in US. | Nationwide Distribution -- USA, including states of SC and PA. | Terminated | May 18, 2010 | 208 | 9
Z-0486-2010 | 53555 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | E-Z CLEAN Non Stick Cautery Tip. Label reads in part: "MEGADYNE 11506 SOUTH STATE STREET DRAPER, UTAH 84020 USA***CAT/REF 0029M E-Z CLEAN NON STICK CAUTERY TIP***STERILE *** Expiration Date: 2104-07***" |  |  | October 13, 2009 | 2009 | Megadyne Medical Products, Inc. | Product labels reference an incorrect expiration date. Expiration date referenced on labels is 2104-07 instead of 2014-07. | Error in labeling | A notification letter dated October 13, 2009 was sent to the first level consignees. Consignees were instructed to examine their stock, remove affected product and contact Megadyne Customer Service at 1-800-747-6110 (United States) or 1-801-576-9669 for product return instructions. |  | 324 | 324 units | Worldwide Distribution -- United States (CA, FL, NE, LA, MT, OK, OR, PA, SC and UT) and New Zealand. | Terminated | February 17, 2010 | 127 | 324
Z-0482-2010 | 53673 | No | EBG | Temporary crown and bridge resin. | Dental | Dental | 510(k) | TempSpan Dual-Cure Temporary Crown and Bridge Material, PART NUMBER N69AD, |  |  | October 22, 2009 | 2009 | Jeneric/Pentron | Pentron Clinical Technologies LLC received two (2) complaints alleging that the TempSpan Dual-Cure Temporary Crown and Bridge Material was not setting.Investigation of product in inventory was conducted and it was confirmed that some of the bottles had been mispackaged with an incorrect catalyst. | Employee error | The recall customer notification was initiated on October 22, 2009 with the firm forwarding an urgent medical device letter with recall return form via USPS first class mail to the affected consignees. The letter informed the consignees of the reason for recall and instructed them to return any used or unused product of the affected lot. The firm provided an 800 number for the consignees to call to arrange for product return and replacement. |  | 58 | 58 units | Nationwide | Terminated | July 19, 2012 | 1001 | 58
Z-0481-2010 | 53748 | No | MRZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas ezManager Max Diabetes Management System version 2.0.9 |  |  | November 02, 2009 | 2009 | Animas Corporation | Printed hard copy of results is not the same as the system screens- carbohydrate to insulin ratio incorrect. | Software design | Animas issued an Urgent Medical Device Correction letter dated 11/2/09 to the pump user and another letter to the health care professional that prescribed the pump. The letters explain the issue. The patient letter is accompanied with a CD that contains the upgraded software and instructions to follow. |  | 20648 | 20,648 units | Nationwide distribution, and Canada, Cayman Islands, and Denmark. | Terminated | October 27, 2010 | 359 | 20648
Z-0483-2010 | 53844 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis MP, Model number 5904466 |  |  | October 26, 2009 | 2009 | Siemens Medical Solutions USA, Inc | c-arm may tilt | Other | The firm has issued a Customer Safety Advisory Notice via Update Instructions AX051/09/S to affected customers. This letter informs customers of the potential malfunction and of the possible danger to patients, users or other persons. |  | 70 | 70 units | Nationwide | Terminated | January 26, 2010 | 92 | 113
Z-0491-2010 | 50685 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zodiac Polyaxial Spinal Fixation System. A spinal interlaminal fixation orthosis is a device intended to be implanted made of an alloy, such as stainless steel, that consists of various hooks and a posteriorly placed compression or distraction rod. The device is implanted, usually across three adjacent vertebrae, to straighten and immobilize the spine to allow bone grafts to unite and fuse the vertebrae together. The device is used primarily in the treatment of scoliosis (a lateral curvature of the spine), but it also may be used in the treatment of fracture or dislocation of the spine, grades 3 and 4 of spondylolisthesis (a dislocation of the spinal column), and lower back syndrome. |  |  | November 03, 2008 | 2008 | Alphatec Spine, Inc. | Alphatec Spine recently discovered that this lot of screws has the incorrect length laser etched on one side of the screw body. As a result, they are conducting a voluntary recall of the Zodiac Titanium cannulated polyaxial screw 6.5 mm x 50 mm L. | Error in labeling | Alphatec Spine contacted consignees via phone on October 16, 2008 and requested that consignees immediately remove the Zodiac Titanium cannulated polyaxial screws (6.5 mm x 50 mm L) from the surgical set to help ensure that they will not be used in surgery. There is a script the company developed to assure consistent responses to customers, and a phone number for more information.The strategy calls for products to be returned to Alphatec Spine. The company will arrange to have Federal Express contact the consignees for pick-up of the screws. Follow up questions are directed to the company at 1-800-922-1356. |  | 56 | 56 | AZ, CA, FL, TX | Terminated | December 08, 2009 | 400 | 56
Z-0063-2010 | 53383 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Heartstart FR2+ automated external defibrillator Models M3860A and M3840A are equipped with an ECG Module. M3860A is under the Philips label and M3840A is under the Laerdal label. |  |  | October 01, 2009 | 2009 | Philips Medical Systems | A potential exists for a component failure, which, if it occurs, could result in the inability to deliver therapy. | Nonconforming Material/Component | On September 28, 2009 Philips issued a nationwide press release. On October 1, 2009 the firm issued notification letters to consignees arranging for the return and replacement of the affected units. The firm also set up a page on their website with a serial number look-up tool.For further information or support concerning this issue, customers are asked to contact Philips at 1.800.263.3342 or visit www.philips.com/FR2PlusAction for the serial number look-up list, answers to frequently asked questions, directions for the battery insertion test, and the replacement AED selection form. |  | 5808 | 5,808 | Domestic and Worldwide distribution to:United States, Canada, Australia, Austria, Bulgaria, China, Finland, France, Germany, Gibraltar, Great Britain, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Jordan, Lebanon, Malaysia, Mexico, Netherlands, Norway, Portugal, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Arab Emirates. | Terminated | September 27, 2013 | 1457 | 5808
Z-0492-2010 | 53538 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE (R) CANAL FILLING STEM EXTENSION, REF: KSP1-0140, DIAMETER 11mm, LENGTH 140mm, 1 EACH, STERILE R, (FOR USE WITH CEMENT ONLY), SINGLE USE ONLY, (Ti6A14V), Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | September 22, 2009 | 2009 | Wright Medical Technology Inc | Titanium used in the recalled devices was discovered to have internal microstructure conditions of alloy segregation, primary alpha and/or a transformed beta matrix in some of the heat lots. The absence of alloying elements results in an area with less strength. For those instances where the product has been implanted, there is the potential increased risk of failure for these patients. | Process control | Wright Medical initiated the recall of the product on September 22, 2009 by contacting the distributors by telephone followed with letters addressed to the hospitals and surgeons dated October 9, 2009. At the firm, Customer Service is the contact for return instructions and replacement inventory. This phone number is 800-238-7117. Other questions regarding this matter are directed to 800-874-5630. |  | 37 | 37 units | Nationwide, Columbia, Brazil, Portugal, UK, Belgium, Austria, and Greece | Terminated | December 16, 2010 | 450 | 62
Z-0488-2010 | 53812 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Elements Gutta Percha and Elements RealSeal Cartridges; Part Number: 972-1002, 972-1003, 972-2001, 972-2002 |  |  | October 27, 2009 | 2009 | Sybronendo | The recall was initiated because during the manufacturing process of the affected units, the molding process resulted in lock nuts that did not have the level of acceptability required, and therefore, may not withstand the pressures exhibited by the Elements Obturation Unit with which the cartridges are used. More Specifically, there exists a possibility that if the lock nut were to fail, the cart | Component design/selection | The recall communications was initiated on November 06, 2009 with the firm forwarding a Recall Letter to the affected consignees via US Postal Service 1st class mail. The recall letter informed the consignees of the reason for recall, the hazard and the affected products. Consignees were instructed to call SybronEndo Customer Care at 1-800-346-3636 directly to handle returns &amp; replacements. In addition, consignees were asked to complete the enclosed Recall Return Form and return it by fax to 909-962-5605.The recall notice also instructed consignees that are an authorized SybronEndo distributor to identify and recover the affected product lots listed above that may have been shipped to their customers. |  | 5046 | 5046 units (3720 units in US) | Nationwide and Worldwide to: USA, Canada, Chile, China, Ecuador, Guatemala, Hong Kong, Japan, Malaysia, Singapore, Taiwan and Venezuela | Terminated | March 22, 2012 | 877 | 5046
Z-0487-2010 | 53751 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Buretrol Solution Sets, a sterile fluid pathway of 105 (2.7 m), 60 drops/mL, with 150 mL Burette Drip Chamber Filter Valve, 2 injection Y-sites, 3 Luer Activated Valves and Male Luer Lock Adapter with Retractable Collar; 48 individually pouched sterile sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | November 03, 2009 | 2009 | Baxter Healthcare Corp. | The tubing may separate from the distal Clearlink Y-site of the set, which could result in patient injury if the separation occurs during patient use. | Employee error | An "Urgent Product Recall" letter dated November 3, 2009 was sent to the direct accounts via first class mail on the same date, to the attention of the Director of Nursing and the Director of Materials Management. The letter informed the accounts that there have been reports of a separation of the tubing from the distal Clearlink Y-site of the set, with one instance resulting in a serious injury. The accounts were instructed to examine their inventory and remove any sets of Product code 2C8864, Lot number UR09E21238 from inventory and return the affected product for credit or replacement product by calling Baxter Healthcare Center for Service at 1-888-229-0001, Prompt 2, during hours of operation 7:00 a.m. to 6:00 p.m. Central Standard Time. Questions regarding the communication were directed to the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 a.m. to 5:00 p.m. Central Standard Time. If the accounts further distributed the Buretrol Sets, they were requested to forward the communication to those accounts. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter. |  | 173 | 173 cases | Worldwide Distribution -- United States (Alabama, California, Colorado, Florida, Georgia, Illinois, Kentucky, Louisiana, Michigan, Mississippi, Missouri, Nevada, New Mexico, Ohio, Tennessee, Texas, Utah and Wisconsin), Canada and New Zealand. | Terminated | July 08, 2010 | 247 | 173
Z-0499-2010 | 53813 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental AdVent Implant System, Abutment, Tapered, Rx, sterile, Zimmer Dental, Inc., Carlsbad, CA; REF AVACT. |  |  | November 06, 2009 | 2009 | Zimmer Inc. | The devices may not meet specifications, which may result in fracturing during tightening of the component. | Process control | User accounts were notified of the issue by phone on or about 11/6/09. Clinicians were requested to schedule a follow-up appointment and to remove and replace the recalled components for any patients who had this device placed without fracturing during the procedure. Questions are directed to the company at 1-800-854-7019. |  | 738 | 738 | Ohio, Canada and France. | Terminated | October 08, 2010 | 336 | 738
Z-0498-2010 | 51734 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ILLICO Posterior Thoracolumbar Access Retractor, Part Number: 73520 |  |  | January 28, 2009 | 2009 | Alphatec Spine, Inc. | The recall was initiated after Alphatec Spine received a complaint on the ILLICO Posterior Thoracolumbar access retractor. It was reported that one of the retractor arms in the retractor broke during surgery. | Device Design | On January 28, 2009 Alphatec Spine contacted their domestic consignees via telephone and the international consignee was contacted via telephone and electronic mail. Alphatec Spine's telephone call consisted of informing the consignees of the product affected, the problem, what actions needed to be taken and what action Alphatec Spine would take to arrange for pick-up of the affected products. |  | 14 | 14 | AZ, CA, OR, TX, and Tokyo, Japan. | Terminated | December 22, 2009 | 328 | 14
Z-0500-2010 | 53689 | No | LGW |  |  | Neurology | PMA | Medtronic Accessory Kit , TitanÂż Anchor Accessory Kit, 3550-39, Single Use, RX Only, Contents of the inner package are STERILE., Medtronic, Inc, Minneapolis, MN 55432-5604 |  |  | October 27, 2009 | 2009 | Medtronic Neuromodulation | Medtronic is issuing a medical device correction for the Model 3550-39 TitanÂż anchor due to the potential for lead migration as a result of insert separation within the anchor. The Medtronic Model 3550-39 TitanÂż anchor accessory kit is designed for use with Medtronic percutaneous in-lineconnector spinal cord stimulation leads. Medtronic has received reports of separation of the titaniuminsert | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent: Medical Device Correction" letter on 10/27/09. The letter was addressed to "Dear Healthcare Professional" and was dated October 2009. The letter described the problem and the product as well as the Scope and Potential Severity and gave Recommendations. Enclosures with the letter included: Use By Date (UBD) location on the Titan anchor accessory kit product packaging, Titan anchor patient letter, Reply card and Return envelope. Questions are direct to Medtronic via email at rs.neuroqa1a@medtronic.com. |  | 55600 | 55,600 | All states in the United States including Puerto Rico.Australia, New Zealand, Canada, Andurra, Belgium, Czech Republic, Denmark, Finland, France, Germany, Greece, Netherlands, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Norway, Portugal, South Africa, Spain, Sweden, Switzerland, Tukey, United Kingdom. | Terminated | December 24, 2011 | 788 | 55600
Z-0502-2010 | 53144 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Alphatec Spine Solanas Titanium Pedicle Screw, Part Number 63035-12 |  |  | August 05, 2009 | 2009 | Alphatec Spine, Inc. | Alphatec Spine discovered that the affected lots were not manufactured correctly in that the screw body portion of the assembly has only 22.5 degrees of angulation in the east-west plane instead of the designed 38-40 degrees. | Process control | Alphatec Spine telephoned consignees informing them of the above mentioned Public Reason for Recall on August 5, 2009. They were asked to check their inventory to see where this product is located.If the product had not already been used Alphatec arranged to have Federal Express contact consignee for pick-up of the product(s). They were told to expect to receive the replacement(s) within 2-3 days. |  | 163 | 163 Units | Nationwide distribution. | Terminated | June 07, 2010 | 306 | 321
Z-0511-2010 | 53557 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Systems, Inc. Connection tubes label reading in part: CONNECTING TUBE W/ STOPCOCK 12" CATALOG NUMBER K10-04131 LOT F720653 Connection tubing can be used with any drainage catheter and drainage bag |  |  | September 29, 2009 | 2009 | Merit Medical Systems, Inc. | Potential sterility breach caused by excessive shipment damage. | Other | Merit Medical Customers were contacted directly by phone and letter on 9/29/09. Customers were instructed to identify and quarantine all unused affected product. Further information was available at 801-208-4365. |  | 1 | 1 connection tube | US (AL, LA, MO, KY, TN, GA). No military, government or foreign distribution. | Terminated | December 15, 2009 | 77 | 352
Z-0525-2010 | 53454 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Abbott Spine Thoracolumbar Common, SpdLnk 5.5 Trans Cntr Med 43-51 mm, non-sterile, Manufactured by Abbott Spine, Inc., Austin, TX; Ref. 1302-55M. Note: This lot was not distributed in the United States. |  |  | September 21, 2009 | 2009 | Zimmer Inc. | The titanium stock used to manufacturer the product is reported to exhibit areas of alloy segregation, which could result in a weakening of the strength of the metal. | Nonconforming Material/Component | Consignees were notified by letter dated September 21, 2009 and instructed to return the affected product to Zimmer, Inc. |  | 13 | 13 | Worldwide Distribution -- Canada, Mexico and South Africa. (Product not distributed in the United States) | Terminated | October 08, 2010 | 382 | 13
Z-0501-2010 | 53860 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Nexgen Complete Knee Solution, Legacy Knee- Posterior Stabilized, Gender Solutions Female, Femoral Component, option, LPS-flex, size C, right, for cemented use only; REF 00-5764-013-52. |  |  | November 13, 2009 | 2009 | Zimmer Inc. | The inner package may interfere with the outer seal, resulting in lack of assurance of sterility. | Packaging process control | Zimmer sent a recall letter, dated 11/12/09, to each initial consignee on 11/13/09. Zimmer field staff is to remove the product from affected hospitals, and to provide them with a copy of the recall letter. |  | 22 | 22 | Nationwide, Korea and Switzerland. | Terminated | October 15, 2010 | 336 | 22
Z-0535-2010 | 52837 | No | LRK | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Repose G2 Bone Screw System, Sterile, REF 76353200; Lot # 62450300. |  |  | July 27, 2009 | 2009 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. is recalling the Repose Bone Screw System product and Bone Screw product due to improper labeling. The labeling for the Repose System was inadvertently placed on the Repose Screw and vice versa. | Labeling mix-ups | Firm notified customers by phone on 7/27/2009. Firm asked customers to check inventories for the affected products listed on the enclosed Checklist.1. Fill-in the "quantity on-hand "column on the checklist.2. If you have any this product lot, contact Medtronic ENT Customer Service at 800-874-5797 (select Option 1 for Customer Service) and ask for the REPOSEField Action Contact Person for help to arrange for returns and replacements or credit.3. Fax your Checklist to Medtronic ENT at the number provided.4. Return products clearly marked. |  | 7 | 7 | 9 consignees in the continental United States. | Terminated | March 15, 2010 | 231 | 20
Z-0509-2010 | 53495 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | IMPACT Catheters, Balloon Dilatation and Esophageal Dilatation Catheter, Model Number 307, STERILE, Rx only |  |  | September 23, 2009 | 2009 | Numed Inc | The firm received a complaint from the sole direct account that while performing an inventory review it was observed that the packaging of the IMPACT catheter was cracked and deteriorating. The product was still in its packaging and never used. The customer referred to it as looking like "paint crackle." The risk is that for continued sterility to be assured, the packaging must remain intact. | Labeling design | An informal e-mail was sent to B. Braun to give them notice not to distribute any Impact devices until further notice. This was followed up by a formal URGENT MEDICAL DEVICE RECALL letter (dated 9/24/09) to the sole consignee. The products are to be returned to the manufacturer, including any products further distributed. Customer questions should be directed to the recall contact (replace " AT " with @) nlaflesh ATnumedusa.com |  | 7367 | 7,367 units domestically | One distributor in Pennsylvania | Terminated | May 23, 2012 | 973 | 7367
Z-0530-2010 | 53584 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Recalled component is labeled: |  |  | January 09, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc. is recalling their OnPoint TMJ Kit (Part Numbers 24-3050, 24-3055) and Arthrosimplicity TMJ Arthocentesis Kit (Part Numbers 922070, 922090, 922075). These two kits contain a BD sterile disposable 60 ml Luer-Lok syringe, Part number 309653 which is under recall by Becton Dickinson (Recall Number Z-2347-2008). Becton Dickinson has determined that unit package seal integ | Device Design | Consignees were notified by a combination of telephone, fax, e-mail and certified mail. Questions related to this notice should be directed to Christy Cain at 1-800-874-7711. |  | 121 | 121 Kits | Product was distributed Internationally to Argentina, Denmark, Norway, and UK. Product was distributed domestically to AL, CA, CT, FL, GA, HI, IA, IL, IN, KS, LA, MA, MD, MN, MO, MS, NC, OH, NY, PA, SC, TN, TX, UT, and VA. | Terminated | March 17, 2010 | 432 | 235
Z-0529-2010 | 53858 | No | HTS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Sterile Knee Instrument, Patella Cutter; |  |  | November 11, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Disposable Patella Cutters - 30 mm (catalog number 8000-0030, lot V31257001) were packaged and labeled as 33 mm Disposable Patella Cutters (catalog number 8000-0033, lot V31257001 and vice versa. | Labeling mix-ups | An "Urgent Product Recall" letter dated November 11, 2009 was sent to all affected Stryker branches, hospital risk management, hospital Chief or Orthopaedics and surgeons by Federal Express. The letter described the issue, potential hazards and risk mitigation factors and requested that the Recall Acknowledgement form be completed and returned. The Recall Acknowledgement form has information regarding the return of affected product to the firm. Direct questions to Stryker Orthopaedics by calling 1-201-831-5832. |  | 12 | 12 | Nationwide Distribution (MA, NJ, FL, TN, OR and NY). | Terminated | August 30, 2010 | 292 | 12
Z-0537-2010 | 50486 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Modular Knee Joint, 3R66, Otto Bock HealthCare GmbH, Max-Nader-StraBe 15 . 37115 Duderstadt /Germany. |  |  | October 28, 2008 | 2008 | Bock,Otto,Orthopedic Ind,Inc | Otto Bock Health Care LP has discovered that in some cases, the securing screws of the joint axes may not have been properly mounted. This can cause the axes to gradually move out of the joint and result in a risk of falling for the user. | Employee error | Consignees were sent on 10/28/08 a Otto Bock "Urgent: Correction Notice" letter dated October 27, 2008. The letter described the product and problem. It provided instructions if the product was in stock and instructed the consignee to contact the patients (a patient letter was provided) and to complete and return the reply card. Consignees were to contact Otto Bock at their Customer Service to request a return authorization. |  | 9 | 9 | The product was only distributed within the US by the US company. Any OUS distribution was handled by the manufacturer in Germany. | Terminated | December 10, 2011 | 1138 | 9
Z-0526-2010 | 52010 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a. Model # 2259988. |  |  | August 13, 2009 | 2009 | GE Medical Systems, LLC | Certain GE Radiographic fixed X-ray tables (Proteus XR/a, Revolution XR/d, and Definium 8000) were found to have tabletops that move unexpectedly in a bi-directional mode. In rare cases, a malfunction may occur, which could result in the tabletop brakes falling to engage when the control foot pedal is released to lock the floating tabletop in place. | Nonconforming Material/Component | A GE Healthcare "Urgent Medical Device Correction" letter dated April 9, 2009 was issued to customers. The letter was addressed to Hospital Administrators/Managers of Radiology and Radiologists. The letter described the Product Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. Operators should verify that the tabletop is locked before the patient is loaded on or off the table. A Field Modification will be performed on the affected products. A GE Healthcare Field Engineer willschedule a field visit to make this modification. Direct questions or concerns regarding the recall to your local GE Healthcare service representative or by calling 1-262-853-7685. |  | 5943 | 5,943 of total product | Worldwide Distribution -- UNITED STATES (including Puerto Rico), VIETNAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKS AND CAICOS ISLANDS, TURKEY, TUNISIA, TRINIDAD AND TOBAGO, THAILAND, TAIWAN, SWITZERLAND, SWEDEN SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL POLAND, PANAMA, PAKISTAN, NORWAY NEW ZEALAND, NETHERLANDS , MEXICO MARTINIQUE, MALTA, MALAYSIA, LITHUANIA, LABA NON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KAZAKSTAN, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, GUAM, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE FINLAND , ECUADOR, DENMARK CZECH REPUBLIC, CROATIA, COLUMBIAN, CHINA, CHILE , CANDA, BULGARIA, BRAZIL, BOSNIA, BELGIUM AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | May 24, 2012 | 1015 | 5943
Z-0172-2010 | 53809 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MX Series Products, Model numbers MX 20, MX 20 High Definition (HD), MX 15, MX 15i (integrated) and MX 15 HDi. |  |  | August 01, 2009 | 2009 | Wescam Inc | Products do not comply with 21 CFR 1040.11. | Radiation Control for Health and Safety Act | L-3 Wescam has issued notification of the MX Series Products describing a non-compliance with 1040.11(b). The notice instructs customers to refrain from using the infrared laser illuminator, based on a serious eye hazard from viewing the direct or reflected laser emission. Products will be repaired or removed if they cannot be repaired and a refund will be implemented. |  | 306 | 306 total | Nationwide Distribution | Terminated | January 15, 2013 | 1263 | 306
Z-0541-2010 | 53519 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Healthcare Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems |  |  | October 22, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an incorrect operating procedure associated with software versions 9.0.0, 9.0.1, and 9.0.2 of the Voluson E6 or Voluson E8 Diagnostic Ultrasound Systems that may impact patient safety.When using the combination of 2D-CRI and PW with a steered Doppler angle, the displayed velocity scale is incorrect. The consequence can be an underestimation of flow v | Software design | Consignees were sent on 10/23/09 a GE healthcare "Urgent Medical Device Correction" letter dated 10/22/09. The letter was addressed to Hospital Administration/Risk manger, Primary Physician and Biomedical Dept. Manger. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Questions and concerns regarding this issue are directed to the local GE Healthcare service representative or to 1-800-437-1171.GE also plans to update all of the affected GE Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems that were distributed. |  | 715 | 715 | AL, AR, CA, CT, FL, GA, IL, KS, KY, LA ME, MD, MA, MI, MN,MS, MO, NE NJ, NY, NC, OH, OK PA, SC, TN,TX, VA,WI.AUSTRALIA, AUSTRIA , BELGIUM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ITALY, JAMAICA, JAPAN, KOREA, KUWAIT, LITHUANIA, MALAYSIA. NETHERLANDS, NE ZEALAND, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, UNITED ARAB, EMIRATES, UNITED KINGDOM, MEXICO, VENEZUELA. | Terminated | December 17, 2011 | 786 | 715
Z-0540-2010 | 53596 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2121-1Q/3131-1Q Biplane Cardiovascular Imaging Systems. |  |  | October 14, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has become aware of a potential limit switch failure with Innova 2121IQ and 3131IQ (Biplane) systems that may impact patient safety.A potential limit switch failure could occur due to collision between the lateral plane of Innova Positioner and other equipment in the room. This collision may break the x-ray and positioning switches, and may not allow exam completion. To date, no in | Employee error | GE Healthcare issued an "Urgent Medical Device Correction" notice dated October 14, 2009. The letter was addressed to Hospital Administrators/Risk mangers, Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For further information, contact GE Healthcare Systems at 1-800-437-1171, option 4 (U.S.) or contact your local GE Healthcare Service Representative. |  | 114 | 114 | Worldwide Distribution -- Including the UNITED STATES, YEMEN, UNITED KINGDOM, TAIWAN, SINGAPORE, RUSSIA, ROMANIA, MALAYSIA, LYB IA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, GERMANY, FRANCE, EGYPT, CHILE , CANADA and BAHRAIN. | Terminated | February 04, 2012 | 843 | 114
Z-0538-2010 | 53943 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Biomet 1.1 X 50MM, 3.5MM STOP STAINLESS STEEL, 1.5MM SYSTEM TWIST DRILL WITH J NOTCH, REF 01-7141. |  |  | November 06, 2009 | 2009 | Biomet Microfixation, Inc. | The Twist Drill 1.1x50mm, 3.5mm stop, with J notch Lots #'s 537407, 536306, and 535933 which were packaged and shipped without the instructions for use (IFU). | Process change control | Consignees were notified of the recall by a combination of telephone calls, letters, and emails during the week of November 16, 2009. Customers were instructed to complete the Inventory Reconciliation Sheet and fax (1-904-741-9425) to the firm. Customers who further distributed the affected product was instructed to notify recipients of the recall. Direct questions about the recall to Biomet Microfixation, Inc. by calling 1-800-874-7711 or 1-904-741-4400, extension 468, Monday through Friday, 8 a.m. to 4 p.m. Eastern Standard Time. |  | 283 | 283 pieces | Worldwide Distribution -- United States (including Puerto Rico), Australia, Brazil, Canada, Chile, Europe, France, Germany, Greece, Japan, Lebanon, Spain, and the United Arab Emirates. | Terminated | March 02, 2010 | 116 | 283
Z-0539-2010 | 52384 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Holding Forceps for Trident Acetabular System; non Sterile; |  |  | January 24, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The tips of the forceps were larger than the positioning holes located on the trials and the implant liners. | Pending | Important Product Correction letters, dated January 24, 2007, were sent to all consignees on via Federal Express. The letter stated the problem and asked customers to examine their inventory and hospital locations to identify the affected product. All affected forceps product should be retrieved from hospitals and returned to their branch or agency warehouse for reconciliation. Customers are also to reconcile all affected product on the attached Product Accountability Form and fax a copy of that form to 201-831-6069 within 5 days of receipt of the letter. All affected forceps product is to be destroyed and discarded at their branch location. Questions should be directed to their Stryker Orthopaedics Customer Service rep at 1-866-OR-ASSIST. |  | 351 | 351 units | Worldwide Distribution -- Including countries of Hong Kong, Australia, Japan, New Zealand, Hong Kong, Singapore, Germany, Sweden, The Netherlands, France, Spain, Portugal, and the United Kingdom. | Terminated | December 21, 2009 | 1062 | 351
Z-0543-2010 | 53771 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-Cat Classic, 3D Dental Imaging System manufactured by Imaging Sciences International, Hatfield, PA. |  |  | October 30, 2009 | 2009 | Imaging Sciences International, LLC | overhead carriage drops quickly | Device Design | Imaging Sciences International issued an "Urgent Medical Device Recall" notification dated October 30, 2009 to all customers informing them of the problem and that a field service representative would schedule a visit to install an upgrade kit.For further information, contact Imaging Sciences International at 1-800-205-3570, option 5. |  | 98 | 98 units | Worldwide Distribution -- Including the United States, Austria, Australia, Belgium, Brazil, Canada, France, Germany, India, Israel, Italy, Jamaica, Spain, South Korea and Taiwan. | Terminated | September 30, 2010 | 335 | 98
Z-0544-2010 | 53577 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump (with X-coating), catalog 164275X, packaged in 8 piece shelf boxes, and convenience kits with single pump units. |  |  | July 06, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The tubing- to- pump connection in a medical device used in surgical procedures may disengage resulting in blood loss. | Device Design | Terumo Cardiovascular Systems Corporation notified consignees of an "Urgent Medical Device Correction" by letter on July 06, 2009. The letter was addressed to hospital Perfusion, Cardiovascular Surgery, and Operating Room departments and contained specific measures to be taken to ensure proper use of the affected device. The firm recommended that users continue to use the devices and follow the corrective measures enclosed with the notification. Consignees were requested to route the notification to all users and confirm receipt of the notification by returning a customer response form.For further information, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 20702 | 20,702 units | Nationwide Distribution | Terminated | November 09, 2011 | 856 | 34455
Z-0546-2010 | 53996 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Legend Footed Attachment, AF01. |  |  | November 30, 2009 | 2009 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Component of product attachment may be outside established tolerance with potential for footed portion to break during normal use. | Nonconforming Material/Component | Medtronic issued an "Urgent: Medical Device Recall" letter dated November 30, 2009. Consignees were notified of the issued and asked to quarantine any affected product and return affected product to the firm for replacement.For further information, contact Medtronic Powered Surgical Solutions Customer Service at 1-800-433-7080 or 1-817-788-6400. |  | 64 | 64 | Worldwide Distribution -- Including the United States (DC, TN, NC, KY, AL, FL, NJ, OH, MI, IN, NY, WA, CA, NV, NM, IA, MO, NE and UT), Great Britain, Brazil, Turkey, Singapore and the Netherlands. | Terminated | May 12, 2010 | 163 | 64
Z-0548-2010 | 53987 | No | HDD | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Four Birthing Bed, Hill-Rom Company, Inc., Batesville, Indiana; Model P3700B. |  |  | November 16, 2009 | 2009 | Hill-Rom, Inc. | The brakes may not hold or lock. | Device Design | Hill-Rom issued an "Urgent Safety Notice" notification dated November 16, 2009 to customers. The letter advised users of the problem and asking them to locate the beds and to remove them from use until corrected if possible and, if not possible to inspect them for the problem and remove them from use if it is present. The firm will visit each account to make corrections.For further information, contact Hill-Rom Technical Support at 1-800-445-3720. |  | 2551 | 2551 (US), 4171 (OUS) | Worldwide Distribution -- United States, Canada, Argentina, Australia, Bahamas, Brazil, Chile, China, Denmark, Dominican Republic, France, Great Britain, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libiya, Netherlands, Poland, Puerto Rico, Russia, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Syria, Thailand, Turkey and United Arab Emirates. | Terminated | July 02, 2012 | 959 | 2551
Z-0549-2010 | 53889 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Fetal Monitor FM 20 |  |  | November 23, 2009 | 2009 | Philips Healthcare Inc. | Potential for inaccurate ultra-sound derived fetal heart rate readings in Philips Avalon Fetal Monitors. Although the types of inaccuracies described in these complaints reflect known limitations inherent to ultrasound fetal heart rate monitoring, the frequency of such complaints is greater for these devices than for the previous generation Philips Series 50 fetal monitors. Clinical decisions bas | Process design | Philips notified customers with an Urgent Medical Device Recall notice and Instructions for Use Addendum on 11/23/09 via UPS. Philips is providing the Addendum to reinforce the Avalon Instructions for Use regarding:"Recognizing these phenomena as they relate to the Avalon Fetal monitor and fetal monitors in general; and"Understanding the differences between the Philips Avalon Fetal Monitors and their predecessors, the Philips Series 50 Fetal Monitors, as they relate to the display of ultrasound-derived FHR.Philips is providing access to application specialists who are on call 24 hours a day to assist users with questions related to device use by calling 1-800-722-9377, press &lt;3&gt; then &lt;1&gt; for patient monitoring application issues and ask to speak with an application specialist. Customers will be requested to incorporate the information contained in the Urgent Medical Device Recall notice and the Addendum to the Instructions for Use with the existing Avalon labeling (Instructions for Use) provided with the product. Customers will be asked to complete an enclosed Confirmation Form indicating that they have received and disseminated this information, and then return the completed form to Philips. Questions about this field action or reports of product problems should be directed to the Philips Response Center at 1-800-722-9377. |  | 10128 | 10,128 units | Nationwide (US and Puerto Rico).Foreign: AfghanistanAlgeriaArgentinaAustraliaAustriaBahrainBelgiumBoliviaBosnia and HerzegovinaBotswanaBrazilBulgariaCanadaChileChinaColombiaCosta RicaCroatiaCzech RepublicDenmarkDominican RepublicEcuadorEgyptEl SalvadorEstoniaFinlandFranceFrench GuianaGabonGermanyGreeceGuadeloupeGuatemalaHong KongHungaryIcelandIndiaIndonesiaIrelandIsraelItalyJapanJordanKazakhstanKenyaKorea, Republic ofKosovoKuwaitLatviaLebanonLibyan Arab JamahiriyaLithuaniaLuxembourgMacauMalaysiaMartiniqueMauritiusMexicoMoldovaMoroccoNamibiaNetherlandsNetherlands AntillesNew CaledoniaNew ZealandNorwayOmanPakistanPanamaPeruPhilippinesPolandPortugalPuerto RicoQatarReunionRomaniaRussiaSaudi ArabiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSwedenSwitzerlandSyrian Arab RepublicTaiwanThailandTunisiaTurkeyTurks and Caicos IslandsUkraineUnited Arab EmiratesUnited KingdomUzbekistanVenezuelaViet NamZambia | Completed |  |  | 18940
Z-0547-2010 | 53938 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions Pronto-Short Extraction Catheter. Model 5030. The PRONTO-Short Extraction catheter is a dual lumen, over-the-wire (OTW) catheter with related accessories. The catheter is designed to be delivered through a 6F or larger introducer sheath over the included 0.018" (0.45mm) guidewire. Sterilized with ethylene oxide. Manufactured by: Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369 USA. |  |  | November 20, 2009 | 2009 | Vascular Solutions, Inc. | It was found that three Pronto Short (model 5030) device lots were manufactured and labeled with an expiration date which is inconsistent than the expiration of the included guidewire component. The guidewire is sterilized separately and was realized to have an expiration date of October 31, 2009. | Process control | Vascular Solutions issued an "Urgent: Medical Device Recall - Lot Specific" notification dated November 20, 2009, via USPS. The letter described the issue, identified affected product, asked for affected devices to be removed from inventory, quarantine, and call customer service for a Return Authorization number. Units will be replaced upon Vascular Solutions (VSI) receiving the returned devices. A Product Recall Inventory form was asked to be completed and faxed back to VSI.For further information, contact your local sales representative or Vascular Solutions Director of Marketing at 1-763-656-4300. |  | 274 | 274 (150 US, 124 OUS) | Worldwide Distribution -- United States (CA, CO, FL, GA, IA, IL, IN, KS, LA, MI, MO, NC, NE, NH, NV, NY, OH, OR, PA, SC, TX, VA, VT, WA and WI), Canada, France, Ireland, Israel, Italy, Netherlands, Spain, Sweden, Taiwan and Turkey. | Terminated | April 05, 2012 | 867 | 274
Z-0557-2010 | 51180 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Howmedica Osteonics, Primary Super SecurFit Stems, hip prosthesis, Sterile; |  |  | January 14, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there may be difficulty in mating stem to implantation/extraction instrument. | Device Design | Stryker sent Urgent product Correction Letters on January 14, 2009 by Federal Express to Stryker branches/agencies, hospital risk management, chief of orthopaedics and surgeon. |  | 786 | 786 total stems | Nationwide. Product was distributed to 3 Stryker branches/agencies and 2 hospitals in PA and CA. | Terminated | July 18, 2012 | 1281 | 786
Z-0555-2010 | 51219 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer, Insall/Burstein II Modular Knee System, Tibial Wedge and 2 Screws, Size 59, 7 Deg. x 9.5 mm Height, Tivanium TI-6AL-4V Alloy, sterile; REF 5220-62-01. |  |  | February 26, 2009 | 2009 | Zimmer Inc. | Smaller diameter screws than required were included in the package. | Process control | Zimmer notified consignees by recall letter dated 2/25/09 instructing them to cease use of the device and to return it. Zimmer distributors and international accounts were notified by email on 2/26/09. |  | 13 | 13 | USA (Connecticut, Indiana, Pennsylvania), Australia, Serbia, Spain, Switzerland and the United Kingdom. | Terminated | February 28, 2011 | 732 | 13
Z-0554-2010 | 52805 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Zimmer / Abbott Spine Sequoia dorsal height &amp; revision tool, non-sterile; REF 3367-1. Surgical instrument for pedicle screw system. |  |  | August 20, 2009 | 2009 | Zimmer Inc. | The tip of the instrument may fracture during use. | Device Design | Zimmer sent a letter dated 8/17/09 to distributors and consignees on or about 8/20/09. Users were notified of the potential problem via the letter and advised that the instruments will be removed and replaced when a new version of the instrument becomes available (around 10/09). |  | 215 | 215 | Nationwide, Belgium and New Zealand. | Terminated | April 13, 2011 | 601 | 215
Z-0558-2010 | 53687 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Zimmer/CAS Power Cord, Sesamoid Plasty, NA, Zimmer/CAS, Montreal (Quebec), Canada; REF 20-8000-070-12. The device is the power cord component for the Sesamoid Plasty CAS workstation which connects the workstation to the power mains. |  |  | November 02, 2009 | 2009 | Zimmer Inc. | The power cord female receptacle is not recessed sufficiently and may present a shock hazard. | Device Design | Zimmer sales staff were notified by letter dated 11/2/09 and instructed to locate the units and to upgrade the cords and to notify consignees by copy of a letter addressed to risk managers and dated 11/2/09. |  | 71 | 71 | Nationwide. | Terminated | October 15, 2010 | 347 | 71
Z-0560-2010 | 53878 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Toggleloc System, Biomet Ziptight Ankle Syndesmosis Fixation System - SS with Ziploop Technology, polyethylene/polypropylene/ polyester/ stainless steel, sterile Biomet Sports Medicine, Warsaw, IN.; REF 909857. |  |  | November 10, 2009 | 2009 | Biomet, Inc. | Packages labeled as containing a stainless steel device actually contain a titanium device, and vice versa. | Labeling mix-ups | Distributors, hospital consignees and an implanting surgeon were notified of the recall by an Urgent Medical Device Recall Notice, dated 11/10/09. Distributors were instructed to locate and remove the identified products from circulation. The instructions on the enclosed "FAX Back Response Form" should be followed and product from their accounts should be returned. Hospital consignees and the implanting surgeon were informed of the recall and advised of the potential health hazards presented by the mix-up. |  | 24 | 24 of both products. | Nationwide Distribution -- Including states of California, Florida, South Carolina, and Virginia. | Terminated | November 03, 2010 | 358 | 24
Z-0542-2010 | 53946 | No | HQY | Sunglasses (nonprescription). | Ophthalmic | Ophthalmic | 510(K) Exempt | Polycarbonate prescription lenses (single vision, progressive, and trifocal) SKU #'s: 3703154, 4286102, 4286103, 4286104 and 4286106 |  |  | November 24, 2009 | 2009 | Luxottica | The lenses made for sunglass frames (Liberty Sport Slam or MaxMorpheus ) do not meet the lens thickness specification established by the frame manufacturer. The lens thickness range from 1.0 to 1.9 millimeters instead of the required 2.0 millimeter minimum. There is a potential risk that the lenses could break or disengage from the frame in the event of an impact. | Employee error | On 11/24/2009 the firm sent recall notification letters to their customers indicating the problem with the lenses. they instructed their customers return the product to the place of purchase for replacement or refund. Questions are directed to the firm at 1-877-488-2407 from 8:00 a.m. to 5:00 p.m. ET. |  | 276 | 276 polycarbonate prescription lenses (138 pairs of lenses) | Product was distributed to the following states: AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MT, NC, NH, NJ, NM, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WI, &amp; WV. | Terminated | October 27, 2010 | 337 | 276
Z-0553-2010 | 53873 | No | MRC | Rigid gas permeable contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb BostonÂż Conditioning Solution, Original Formula, Bausch &amp; Lomb Incorporated, Rochester, NY 14609. The product is packed in 1 fl oz (30 mL), 3 fl oz (90 mL), 3.5 fl oz (105 mL) and 4 fl oz (120 ml) bottles. |  |  | October 30, 2009 | 2009 | Bausch &amp; Lomb Inc | The product failed to meet Out of Specification shelf life at the 9 months time point. This could result in increased risk of an ocular infection, especially in the presence of other contributing or risk factors. | Other | Consignees were notified by mail on/about 11/06/2009. All consignees were asked to review their inventory, and return any of the affected products to Stericycle Pharmaceutical Services. A Business Reply Card was included to be returned via US mail within 5 days. A packing slip was provided to be affixed to a pre-paid USP shipping label for return. |  | 159672 | 159,672 bottles | U.S. Nationwide, Canada and Hong Kong | Terminated | August 29, 2012 | 1034 | 159672
Z-0561-2010 | 47370 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI. |  |  | September 17, 2008 | 2008 | Xoran Technologies, Inc. | The system may overheat, causing poor image quality, interruption in image acquisition and/or damage to the system. Also, if the system is not allowed to cycle down prior to being restarted after an emergency stop, the x-ray and gantry motions will both resume. | Device Design | A letter dated 9/17/08 was sent to each account telling them that the PLC would be upgraded. The firm visited each account and made corrections. |  | 82 | 82 | Nationwide Distribution | Terminated | December 23, 2009 | 462 | 82
Z-0566-2010 | 53817 | No | MMZ | Urological table and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stille Sonesta Gynecology and Urology Table 6303 ; an electric cystometric table; Stille AB, Solna, Sweden. |  |  | November 03, 2009 | 2009 | Stille AB | There is a potential hazard to the patient if the bolts holding the side rail break, causing the accessories attached at the side rails of the tables and chairs to fall off during a medical procedure. | Nonconforming Material/Component | Stille sent Field Safety Notices dated 10/30/09 to the direct accounts, informing them that there is a potential hazard with the Sonesta tables if the bolts holding the side rail break, causing the accessories attached at the side rails of the tables and chairs to fall off during a medical procedure. Stille developed a retrofit kit consisting of six side rail bolts and instructions for installing the bolts, which were sent to the end users with the safety notice. The accounts were requested to replace the side rail bolts with the new bolts provided, and to complete and return the enclosed acknowledgement card to Stille by fax or e-mail as soon as possible and no later than 12/11/09. Contact Stille Surgical Inc. at 224-612-5404 for assistance. |  | 68 | 68 units total | Worldwide Distribution -- Including countries of USA and Canada. | Terminated | December 10, 2010 | 402 | 68
Z-0562-2010 | 53820 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart Vehicle Wall Mount, Model M5528A |  |  | November 12, 2009 | 2009 | Philips Healthcare Inc. | HeartStart Vehicle Wall Mount (Model M5528A) with DC Power Module (Model M5529A)--Incorrect wiring during installation by user may result in a smoke or fire hazard. | Device Design | Philips Healthcare, Inc. notified customers by an "URGENT - Voluntary Medical Device Correction" letter, plus revised Instructions For Use (IFU) for the Vehicle Wall Mount accessory and the DC Power Module accessory of the HeartStart MRx which includes a revised description of correct installation and verification of performance. Customers were asked to follow the Actions in the customer/user section of the Urgent - Voluntary Medical Device Correction Notification. In this section, customers were given the instructions to correct their devices. Distribution of the Urgent - Voluntary Medical Device Correction Notification along with the IFU will be managed by Philips representatives in each affected geography. Direct questions or concerns about the recall to your local Philips representative. |  | 3265 | 3,265 units | Class II Recall -- Worldwide Distribution -- United States, Canada, Australia, Brazil, China, France, Germany, Hong Kong, Iceland, India, Indonesia, Italy, Japan, Jordan, Kenya, Libyan Arab Jamahiriya, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Romania, Saudi Arabia, Spain, Sri Lanka, Switzerland, Trinidad, Tobago, United Arab Emirates, and United Kingdom. | Terminated | January 25, 2012 | 804 | 3265
Z-0563-2010 | 53837 | No | ETP | Air or water caloric stimulator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | AquaStar Water Caloric Irrigator; water caloric stimulator; Manufactured for/Distributed by Micromedical Technologies, Inc., 10 Kemp Drive, Chatham, IL 62629 USA, Instrumentation Difra, Made in Belgium. |  |  | November 16, 2009 | 2009 | Micromedical Technologies, Inc | Some of the filters on both the display board and the main electronic board (PCB) become oxidized causing the LCD display to be garbled and the timer will not operate. | Nonconforming Material/Component | Micromedical Technologies, Inc. sent letters dated November 16, 2009 via e-mail (with read receipt) to each customer and their representative. If they did not have the customer's e-mail address, the sales representative was requested to forward the letter to the customer and copy Micromedical. The letters advised the customers that one or more electronic components in the AquaStar caloric irrigator may fail due to contact corrosion. In the event of this failure, unheated tap water will continue to flow through the device back to the water drain, but the LCD display will be garbled and the timer will not operate. If this problem develops, the account should stop using the irrigator and contact Micromedical Technologies, Inc. at 1-217-483-2122 or via e-mail at mmtinfo@micromedical.com to report the problem for repair of the unit. |  | 10 | 10 units | Worldwide Distribution -- United States (New York and Pennsylvania), Canada, Egypt, Korea, Poland and South Africa. | Terminated | April 18, 2011 | 518 | 10
Z-0567-2010 | 53952 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | SeQual Technologies Inc. Eclipse Oxygen System, Model 1000 |  |  | October 01, 2009 | 2009 | Sequal Technologies Inc | Unit may shut down when operating from the Power Cartridge. A temporary power interruption would cause the unit to sound an alarm and shut down. If the patient does not have access to an alternate supply of oxygen or access to AC or DC power, this may lead to shortness of breath, hypoxia, or in some patients may cause serious health consequences. | Pending | The recall notification was initiated on November 12, 2009 with the firm forwarding a "Medical Device Correction" letter to all of the affected consignees via US Postal Service for US Consignees and FedEx for Foreign Consignees. The Medical Device Correction Letter contained the following:(1) Medical Device Correction Letter; (2) List of Serial Number affected that were distributed to the Provider; (3) Reply Card that the Provider will be requested to send back regarding receipt notification of the Medical Device Correction Letter and confirming that they have the affected product; (4) Additional Medical Device Correction Letters for Providers to distribute to their patients.The Medical Device Correction Letter informed the consignees of the reason for recall, the hazard involved and the actions that were required. For further details or information, consignees were instructed to contact SeQual Technical Support at 1-800-826-4610 from 8:00 AM to 5:00 PM PST. |  | 6588 | 6,588 units | Worldwide Distribution -- Including countries of USA, ARGENTINA, AUSTRIA, BAHAMAS, BERMUDA, CANADA, CHILE, COLUMBIA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, ITALY, ISRAEL,JAPAN, JORDAN, KUWAIT, MALTA, NETHERLANDS, POLAND, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | June 18, 2012 | 991 | 6588
Z-0497-2010 | 53879 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | TrailBlazer Support Catheter, ev3 inc, 4600 Nathan Lane North, Plymouth, MN 55442-2920, Rx Only, Sterile EO. |  |  | November 06, 2009 | 2009 | Ev3, Inc. | Catheter may crack near the radiopaque marker band at the distal end of the device during use. Cracking of the catheter shaft in this location can result in material separation, and potential embolization. | Component design/selection | Consignees were sent a Medical Device Recall letter, dated 11/6/09. The letter was addressed to Risk Manger or Cath Lab Manage. The letter provided the Issue Summary, identified the Affected Product, and stated the Required Action, which asked consignees to locate and remove the affected product. Questions should be directed to ev3 Customer Service at 800-716-6700. |  | 70 | 70 | Nationwide Distribution -- USA, including states of CA, FL, GA, IL, IN, IA, KS, KY, MI, MT, NJ, PA, SD, TN, TX, and WI. | Terminated | January 05, 2011 | 425 | 70
Z-0568-2010 | 53339 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Paramed Polypropylene evacuation sled for non-ambulatory patients. |  |  | October 29, 2008 | 2008 | Paramed Systems | Possible partial board disassembly due to inadequate webbing attachment. | Process change control | Customers were notified by letter and phone during 10/20/2008 and 10/29/2008 to schedule visits by Paramed to correct the devices. |  | 722 | 722 Units | Nationwide Distribution -- Including states of AZ, CA, FL, GA, IL, IN, KS, KY, MA, MI, MT, NY, PA, SC, TN and UT. | Terminated | January 28, 2010 | 456 | 722
Z-0574-2010 | 53709 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100 30 Gauge, 1cc Insulin Syringe, 5/16" (8mm) Needles, Short Needle, packaged 10 syringes per package, 10 packages per carton, NDC 0603-7000-21 |  |  | October 27, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles are becoming dislodged from the barrels. | Other | The recall is being conducted to the Consumer level. A Press Release was issued on October 27, 2009 stating that Qualitest Pharmaceuticals was issuing a nationwide recall of all Accusure ÂżÂż Insulin Syringes. The consignees were asked to post a copy of the press release in their retail stores and to alert patients of this action. Also, Recall Notifications were sent to all of the firm's consignees by First Class Mail on October 30, 2009. The consignees were instructed to notify their customers or sub accounts who may have been sold or dispensed the affected syringes starting in 2002, and to return the affected product. Questions are directed to the firm at telephone number 800-444-4011. |  | 17142100 | 171,421 cartons of 100 | Nationwide | Terminated | November 24, 2010 | 393 | 237867600
Z-0580-2010 | 53777 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Enpath Lead Adapter, Ref: 501214. The contents of the inner package have been sterilized by ethylene-oxide gas. Enpath Medical, Inc., 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. |  |  | September 01, 2009 | 2009 | Greatbatch Medical | The silicone tubing of affected leads was processed outside of validated parameters. Although the material remains biocompatible, test data regarding the performance characteristics are not available. | Component change control | Customers were contacted by phone beginning 09/01/09. They were asked to hold product and determine if they had distributed any product. A written notification followed which identified the product, described the issue and asked for product to be returned. Product was distributed to two consignees and contained the following day. Questions are directed to the company at 612-559-2613. |  | 45 | 45 units | CA and MN | Terminated | December 17, 2011 | 837 | 80
Z-0586-2010 | 53901 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 30 ml Syringe; |  |  | November 23, 2009 | 2009 | Becton Dickinson &amp; Company | Reports received of open seals on lot # 9175046, BD 30 ml Luer-Lok Syringes. This indicates product sterility may be compromised. | Other | Recall notification letters were sent to all distributors on 11/23/2009 by Federal Express. BD called all distributors to get customer lists and will start sending letter to customers commencing on December 8, 2009. Questions are directed to BD Customer Service at 1-888-237-2762, Monday  Friday, 8:00 AM  5:00 PM Eastern Time. |  | 166240 | 166240 units | Distributors nationwide. | Terminated | April 06, 2010 | 134 | 166240
Z-0582-2010 | 53875 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Delta/Delta XL Monitors with Omega S Package |  |  | November 09, 2009 | 2009 | Draeger Medical, Inc. | monitor is rebooting/resetting | Software design | Draeger Medical Systems, Inc. issued and Urgent Medical Device Recall letter dated November 2009 to their affected customers informing them of the problem and a work around until software can be validated. A Technical Service Bulletin has been generated and distributed globally that instructs the affected International Drager subsidiaries to inform customers in their countries.For further information, contact Draeger Medical Systems, Inc. at 1-800-543-5047, extension 1-2349. |  | 4188 | 4,188 units | The product was shipped to medical facilities in CA, CO, CT, ID, IN, KY, MS, NY, OK, and WI. The product was also shipped to United Arab Emirates, Argentina, Austria, Australia, Bosnia Herz, Belgium, Benin, Brunei Darussalam, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Croatia, Hungary, Ireland, India, Iran, Italy, Japan, Kenya, Korea, Kuwait, Kazakhstan, Lebanon, Lithuania, Latvia, Libya, Morocco, Mexico, Malaysia, Netherlands, Norway, Nepal, Oman, Poland, Portugal, Qatar, Reunion, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Thailand, Turkey, Taiwan, Ukraine, Venezuela, Vietnam, and South Africa. | Terminated | August 23, 2011 | 652 | 4188
Z-0579-2010 | 53871 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX SR Power Wheelchair with Stability Lock Model numbers: TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-HD, TDXSR-MCG, TDXSR-MCG-HD, TDXSRV-HD |  |  | September 15, 2009 | 2009 | Invacare Corporation | For some Invacare chairs, the stability lock feature may not be engaging properly or consistently. If the stability lock feature does not engage properly, the wheelchair may do one or more of the following: (1) veer to one side; (2) rock forward onto its front riggings; (3) drive in an unintended circular motion; and/or (4) fall forward or to one side. In these cases there is a risk to injury for | Device Design | On 9/15/2009 the firm sent recall notification letter to their customers instructing them to immediately contact the dealer to have the stability lock inspected and/or adjusted. If the dealer is unavailable or the repair will take longer than 60 days, they are to contact the firm and arrange for an inspection and/or adjustment. Invacare requested contact names to whom it can send additional information. Invacare plans to mail, via FEDEX, information regarding the recall to each consignee. Initial consignees received a response card and instructions to forward cards to any customer to whom they have distributed the product. The products are to be inspected at the cylinder label and if the label does not have a long alpha-numeric string demonstrated in the label, the cylinders on the chair have to be replaced. Questions are directed to Recalls@Invacare.com or to 800-333-6900. |  | 36 | 36 | Product was distributed to the following states: AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, MA, ME, MI, MN, MO, MS, NC, ND, NM, NY, OH, SC &amp; TX. Product was also distributed in Canada. | Terminated | August 25, 2014 | 1805 | 36
Z-0585-2010 | 53937 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oxygen Sensor #5804, Distributed by: Teleflex Medical, Research Triangle Park, NC 27709 |  |  | November 25, 2009 | 2009 | Teleflex Medical | Sensor causes the oxygen monitor to display prematurely a Low Sensor warning which indicates an early maintenance requirement for the instrument. This may result in interrupted treatment. | Vendor change control | Consignees were notified via next day delivery on or about 11/25/2009. Each customer received a letter requesting acknowledgement of the recall and a request to return the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second and third mailing to non-responding customers on the consignee list will be made, or as an alternative, the third contact may be by phone, e-mail, or fax. Affected products received by Teleflex Medical will be destroyed or reworked and replacement product or credit will be provided to the consignee. Records of contacts and of destruction or rework will be maintained. Questions are directed to the company at 1-866-246-6990. |  | 51 | 51 eaches | FL, KY, MI, OH, PA, SD, TN, Australia and Canada. | Terminated | April 07, 2011 | 498 | 51
Z-0583-2010 | 53919 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Siemens Medical Solutions USA, Inc., Hoffman Estates, IL 60195-2061. |  |  | December 11, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Some of the hand controllers of the Symbia S &amp; T systems are missing a resistor switch, and the absence of the switch may lead to unintended system motion during camera set up activities. | Process control | Siemens issued an Urgent Field Correction Recall letters dated November 18, 2009 and December 11, 2009 to the affected users of the Symbia S and Symbia T Systems, respectivley. The customers were advised of precautionary measures and that Siemens will replace the affected device. They also were advised that a Siemens Customer Service Representative would contact them within a few days to schedule a replacement of the device. For further information, contact Siemens Medical Solutions USA, Inc. at 1-800-888-7436 (USA). |  | 135 | 135 units | Worldwide Distribution -- United States, Algeria, Australia, Austria, Belgium, Canada, Chile, China, Finland, France, Germany, Hong Kong, India, Indonesia, Italy, Korea, Mexico, Netherlands, Norway, Philippines, Portugal, Qatar, Russia, Saudi Arabia, Spain, Sweden, Taiwan, Turkey, Venezuela and Vietnam. | Terminated | August 27, 2012 | 990 | 232